## Sheet1
Web Address | Report Number | Event Date | Event Type | Manufacturer | Date Received | Product Code | Brand Name | Device Problem | Patient Problem | PMA/PMN Number | Exemption Number | Number of Events | Event Text
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21613422&pc=NIO | 9681442-2025-00078 | 2017-08-27 04:00 | Death | ANGIOMED GMBH & CO. MEDIZINTECHNIK KG | 2025-03-15 04:00 | NIO | E-LUMINEXX VASCULAR STENT | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P08007 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESULTS OF THE CLINICAL TRIAL THAT APPROXIMATELY ONE YEAR FIVE MONTHS AND SIX DAYS POST A STENT PLACEMENT ON THE RIGHT LEG FOR THE DE-NOVO STENOSIS ON THE COMMON AND EXTERNAL ILIAC ARTERY, THE SUBJECT HAD AN ADVERSE EVENT OF DEATH. THE PRIMARY CAUSE OF DEATH WAS NOT PROVIDED. REPORTEDLY, THE ADVERSE EVENT DEATH WAS POSSIBLY RELATED TO THE STUDY DEVICE AND NOT RELATED TO THE PROCEDURE.Event Description: IT WAS REPORTED THROUGH THE RESULTS OF THE CLINICAL TRIAL THAT APPROXIMATELY ONE YEAR FIVE MONTHS AND SIX DAYS POST A STENT PLACEMENT ON THE RIGHT LEG FOR THE DE-NOVO STENOSIS ON THE COMMON AND EXTERNAL ILIAC ARTERY, THE SUBJECT HAD AN ADVERSE EVENT OF DEATH. THE PRIMARY CAUSE OF DEATH WAS NOT PROVIDED. REPORTEDLY, THE ADVERSE EVENT DEATH WAS POSSIBLY RELATED TO THE STUDY DEVICE AND NOT RELATED TO THE PROCEDURE. Manufacturer Narrative: H11: MANUFACTURING REVIEW: THE LOT NUMBER OF THE PLACED STENT WAS NOT REPORTED. THEREFORE, A DHR REVIEW COULD NOT BE REVIEWED. INVESTIGATIONS SUMMARY: THE STENT REMAINS IMPLANTED. X-RAY IMAGES WERE NOT PROVIDED FOR EVALUATION. THE EVALUATION OF THE MATERIAL AVAILABLE IS SUBJECT TO THE RETROSPECTIVE ASSESSMENT AS PART OF POST-MARKET CLINICAL FOLLOW-UP (PMCF) ACTIVITIES. NO IRREGULAR STENT PLACEMENT DURING INDEX PROCEDURE AND NO DEFICIENCY OF THE PLACED STENT, SUCH AS DEFORMATION OR FRACTURE WAS REPORTED. THE DEATH OF THE PATIENT WAS DOCUMENTED NOT BEING RELATED TO THE PLACED STENT OR THE STENT PLACEMENT PROCEDURE. THE PATIENT DIED ABOUT ONE YEAR AND FIVE MONTHS AFTER THE STENT PLACEMENT PROCEDURE. THE PATIENT WAS A CURRENT SMOKER AND HAD A CLINICAL HISTORY OF HYPERTENSION, DYSLIPIDEMIA, CEREBRAL ARTERY DISEASE, POLYNEUROPATHY, SUBCORTICAL ARTERIOSCLEROTIC ENCEPHALOPATHY AND C2 ABUSE. BASED ON INFORMATION AVAILABLE, THERE IS NO INDICATION THAT THE DEATH OF THE PATIENT MIGHT BE RELATED TO THE PLACED STENT OR THE STENT PLACEMENT PROCEDURE. BASED ON THE INFORMATION AVAILABLE, THE INVESTIGATION IS CLOSED WITH INCONCLUSIVE RESULT. BASED ON INFORMATION AVAILABLE, THERE IS NO INDICATION THAT THE DEATH OF THE PATIENT MIGHT BE RELATED TO THE PLACED STENT OR THE STENT PLACEMENT PROCEDURE. LABELING REVIEW: RELEVANT LABELING SUPPLIED WITH THIS PRODUCT WAS REVIEWED. POTENTIAL COMPLICATIONS AND ADVERSE EVENTS WHICH MAY OCCUR DURING USE OF THE DEVICE WERE FOUND TO BE REFERENCED IN THE IFU. USERS AND/OR PATIENTS SHOULD REPORT ANY SERIOUS INCIDENT THAT HAS OCCURRED IN RELATION TO THE DEVICE TO THE MANUFACTURER AND TO EITHER THE EUROPEAN UNION NATIONAL COMPETENT AUTHORITY OR TO THE REGULATORY AUTHORITY OF THE COUNTRY IN WHICH THE USER AND/OR PATIENT IS ESTABLISHED. IN THIS CASE THE DEATH OF THE PATIENT WAS DOCUMENTED NOT BEING RELATED TO THE PLACED STENT OR THE STENT PLACEMENT PROCEDURE. G3, H6 (METHOD). SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. Manufacturer Narrative: H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21613440&pc=NIO | 9681442-2025-00079 | 2016-08-16 04:00 | Death | ANGIOMED GMBH & CO. MEDIZINTECHNIK KG | 2025-03-15 04:00 | NIO | E-LUMINEXX VASCULAR STENT | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P080007 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESULTS OF THE CLINICAL TRIAL THAT APPROXIMATELY ONE YEAR FIVE MONTHS AND SEVEN DAYS LATER POST A STENT PLACEMENT ON THE RIGHT LEG FOR THE DE-NOVO STENOSIS ON THE EXTERNAL ILIAC ARTERY, THE SUBJECT HAD AN ADVERSE EVENT OF DEATH. THE PRIMARY CAUSE OF DEATH WAS NOT PROVIDED. REPORTEDLY, THE ADVERSE EVENT DEATH WAS POSSIBLY RELATED TO THE STUDY DEVICE AND NOT RELATED TO THE PROCEDURE.Event Description: IT WAS REPORTED THROUGH THE RESULTS OF THE CLINICAL TRIAL THAT APPROXIMATELY ONE YEAR FIVE MONTHS AND SEVEN DAYS LATER POST A STENT PLACEMENT ON THE RIGHT LEG FOR THE DE-NOVO STENOSIS ON THE EXTERNAL ILIAC ARTERY, THE SUBJECT HAD AN ADVERSE EVENT OF DEATH. THE PRIMARY CAUSE OF DEATH WAS NOT PROVIDED. REPORTEDLY, THE ADVERSE EVENT DEATH WAS POSSIBLY RELATED TO THE STUDY DEVICE AND NOT RELATED TO THE PROCEDURE. Manufacturer Narrative: H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. Manufacturer Narrative: H11: MANUFACTURING REVIEW: BASED ON THE INFORMATION AVAILABLE IT IS NOT REASONABLY SUGGESTED THAT A MANUFACTURING PROCESS MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED ISSUE. THE LOT NUMBER OF THE PLACED STENT WAS NOT REPORTED. THEREFORE, A DHR REVIEW COULD NOT BE REVIEWED. INVESTIGATION SUMMARY: THE STENT REMAINS IMPLANTED. X-RAY IMAGES WERE NOT PROVIDED FOR EVALUATION. THE EVALUATION OF THE MATERIAL AVAILABLE IS SUBJECT TO THE RETROSPECTIVE ASSESSMENT AS PART OF POST-MARKET CLINICAL FOLLOW-UP (PMCF) ACTIVITIES. NO IRREGULAR STENT PLACEMENT DURING INDEX PROCEDURE AND NO DEFICIENCY OF THE PLACED STENT, SUCH AS DEFORMATION OR FRACTURE WAS REPORTED. THE DEATH OF THE PATIENT WAS DOCUMENTED NOT BEING RELATED TO THE PLACED STENT OR THE STENT PLACEMENT PROCEDURE. THE PATIENT DIED ABOUT ONE YEAR AND FIVE MONTHS AFTER THE STENT PLACEMENT PROCEDURE. THE PATIENT WAS A CURRENT SMOKER AND HAD A CLINICAL HISTORY OF CONGESTIVE HEART FAILURE, STROKE, HYPERTENSION, CORONARY ARTERY DISEASE, ATRIAL FIBRILLATION, ARRHYTHMIA, VENTRICULAR EXTRASYSTOLE, AV BLOCK, CHRONIC ISCHEMIC HEART DISEASE, COLON CARCINOMA, AND A OTHER. BASED ON INFORMATION AVAILABLE, THERE IS NO INDICATION THAT THE DEATH OF THE PATIENT MIGHT BE RELATED TO THE PLACED STENT OR THE STENT PLACEMENT PROCEDURE. BASED ON THE INFORMATION AVAILABLE, THE INVESTIGATION IS CLOSED WITH INCONCLUSIVE RESULT. BASED ON INFORMATION AVAILABLE, THERE IS NO INDICATION THAT THE DEATH OF THE PATIENT MIGHT BE RELATED TO THE PLACED STENT OR THE STENT PLACEMENT PROCEDURE. LABELING REVIEW: RELEVANT LABELING SUPPLIED WITH THIS PRODUCT WAS REVIEWED. POTENTIAL COMPLICATIONS AND ADVERSE EVENTS WHICH MAY OCCUR DURING USE OF THE DEVICE WERE FOUND TO BE REFERENCED IN THE IFU. USERS AND/OR PATIENTS SHOULD REPORT ANY SERIOUS INCIDENT THAT HAS OCCURRED IN RELATION TO THE DEVICE TO THE MANUFACTURER AND TO EITHER THE EUROPEAN UNION NATIONAL COMPETENT AUTHORITY OR TO THE REGULATORY AUTHORITY OF THE COUNTRY IN WHICH THE USER AND/OR PATIENT IS ESTABLISHED. IN THIS CASE THE DEATH OF THE PATIENT WAS DOCUMENTED NOT BEING RELATED TO THE PLACED STENT OR THE STENT PLACEMENT PROCEDURE. G3, H6 (METHOD) SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21613442&pc=NIO | 9681442-2025-00080 | 2022-09-22 04:00 | Death | ANGIOMED GMBH & CO. MEDIZINTECHNIK KG | 2025-03-15 04:00 | NIO | E-LUMINEXX VASCULAR STENT | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P080007 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESULTS OF THE CLINICAL TRIAL THAT APPROXIMATELY THREE YEARS FIVE MONTHS AND TWENTY FIVE DAYS LATER POST A STENT PLACEMENT ON THE RIGHT LEG FOR THE DE-NOVO STENOSIS ON THE EXTERNAL ILIAC ARTERY, THE SUBJECT HAD AN ADVERSE EVENT OF DEATH. THE PRIMARY CAUSE OF DEATH WAS NOT PROVIDED. REPORTEDLY, THE ADVERSE EVENT DEATH WAS POSSIBLY RELATED TO THE STUDY DEVICE AND NOT RELATED TO THE PROCEDURE.Event Description: IT WAS REPORTED THROUGH THE RESULTS OF THE CLINICAL TRIAL THAT APPROXIMATELY THREE YEARS FIVE MONTHS AND TWENTY FIVE DAYS LATER POST A STENT PLACEMENT ON THE RIGHT LEG FOR THE DE-NOVO STENOSIS ON THE EXTERNAL ILIAC ARTERY, THE SUBJECT HAD AN ADVERSE EVENT OF DEATH. THE PRIMARY CAUSE OF DEATH WAS NOT PROVIDED. REPORTEDLY, THE ADVERSE EVENT DEATH WAS POSSIBLY RELATED TO THE STUDY DEVICE AND NOT RELATED TO THE PROCEDURE. Manufacturer Narrative: H11: MANUFACTURING REVIEW: BASED ON THE EVENT INFORMATION PROVIDED THERE IS NO INDICATION THAT A MANUFACTURING PROCESS MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED ISSUE. THEREFORE, A DHR REVIEW IS CONSIDERED TO BE NOT REQUIRED. INVESTIGATION SUMMARY: THE STENT REMAINS IMPLANTED. X-RAY IMAGES WERE NOT PROVIDED FOR EVALUATION. THE EVALUATION OF THE MATERIAL AVAILABLE IS SUBJECT TO THE RETROSPECTIVE ASSESSMENT AS PART OF POST-MARKET CLINICAL FOLLOW-UP (PMCF) ACTIVITIES. NO IRREGULAR STENT PLACEMENT DURING INDEX PROCEDURE AND NO DEFICIENCY OF THE PLACED STENT, SUCH AS DEFORMATION OR FRACTURE WAS REPORTED. THE DEATH OF THE PATIENT WAS DOCUMENTED NOT BEING RELATED TO THE PLACED STENT OR THE STENT PLACEMENT PROCEDURE. THE PATIENT DIED ABOUT THREE YEARS AND SIX MONTHS AFTER THE STENT PLACEMENT PROCEDURE. THE PATIENT WAS A CURRENT SMOKER AND HAD A CLINICAL HISTORY OF HYPERTENSION, DYSLIPIDEMIA, CEREBRAL ARTERY DISEASE, POLYNEUROPATHY, SUBCORTICAL ARTERIOSCLEROTIC ENCEPHALOPATHY AND C2 ABUSE. BASED ON INFORMATION AVAILABLE, THERE IS NO INDICATION THAT THE DEATH OF THE PATIENT MIGHT BE RELATED TO THE PLACED STENT OR THE STENT PLACEMENT PROCEDURE. BASED ON THE INFORMATION AVAILABLE, THE INVESTIGATION IS CLOSED WITH INCONCLUSIVE RESULT. BASED ON INFORMATION AVAILABLE, THERE IS NO INDICATION THAT THE DEATH OF THE PATIENT MIGHT BE RELATED TO THE PLACED STENT OR THE STENT PLACEMENT PROCEDURE. LABELLING REVIEW: RELEVANT LABELING SUPPLIED WITH THIS PRODUCT WAS REVIEWED. POTENTIAL COMPLICATIONS AND ADVERSE EVENTS WHICH MAY OCCUR DURING USE OF THE DEVICE WERE FOUND TO BE REFERENCED IN THE IFU. USERS AND/OR PATIENTS SHOULD REPORT ANY SERIOUS INCIDENT THAT HAS OCCURRED IN RELATION TO THE DEVICE TO THE MANUFACTURER AND TO EITHER THE EUROPEAN UNION NATIONAL COMPETENT AUTHORITY OR TO THE REGULATORY AUTHORITY OF THE COUNTRY IN WHICH THE USER AND/OR PATIENT IS ESTABLISHED. IN THIS CASE THE DEATH OF THE PATIENT WAS DOCUMENTED NOT BEING RELATED TO THE PLACED STENT OR THE STENT PLACEMENT PROCEDURE. G3, H6 (METHOD). SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. Manufacturer Narrative: H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21613514&pc=MVK | 3008642652-2025-02742 | 2025-02-27 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-03-15 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED AN INAPPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 07:50:55 ON (B)(6) 2025. AT 13:19:10, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY. AT 13:19:59, THE PATIENT RECEIVED THE INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. THE RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM WAS AF FROM 25-80 BPM. THE ELECTRODE BELT WAS DISCONNECTED AT 15:30:10 ON (B)(6) 2025. Manufacturer Narrative: THE ELECTRODE BELT AND MONITOR HAVE BEEN RETURNED AND THE EVALUATION IS UNDERWAY. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21613591&pc=BZD | 2518422-2025-103064 | 2025-01-13 05:00 | Death | RESPIRONICS, INC. | 2025-03-15 04:00 | BZD | REP DREAMSTATION AUTO CPAP, DOM - RECRT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A NOTIFICATION RECEIVED REQUESTING THE RETURN OF A DEVICE. THE MANUFACTURER WAS INFORMED THAT THE PATIENT WAS DECEASED. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21613594&pc=BZD | 2518422-2025-103065 | 2025-01-14 05:00 | Death | RESPIRONICS, INC. | 2025-03-15 04:00 | BZD | REMSTAR SYSTEM ONE 60 SERIES | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A CUSTOMER REQUESTING INFORMATION ON RETURNING A DEVICE DUE TO A PATIENT PASSING. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21613595&pc=BZD | 2518422-2025-103066 | 2025-01-14 05:00 | Death | RESPIRONICS, INC. | 2025-03-15 04:00 | BZD | REP DREAMSTATION AUTO CPAP, DOM - RECRT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A NOTIFICATION RECEIVED REQUESTING THE RETURN OF A DEVICE. THE MANUFACTURER WAS INFORMED THAT THE PATIENT WAS DECEASED. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21613599&pc=OZP | 2032227-2025-136139 | 2024-12-21 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-03-15 04:00 | OZP | MINIMED 780G US SYSTEM BLE CONNECT 3.0 MG/DL | Adverse Event Without Identified Device or Use Problem; Appropriate Term/Code Not Available | Pneumonia | P160017 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2024 AND THE CAUSE OF DEATH WAS PNEUMONIA. IT WAS ALSO REPORTED THAT THE CUSTOMER EXPERIENCED HYPERGLYCEMIA WITH THE BLOOD GLUCOSE VALUE OF 500 MG/DL AT THE TIME OF HOSPITALIZATION WHILE ADMITTED FOR PNEUMONIA. THE LAST KNOWN PRODUCT FOR THIS CUSTOMER ARE AS FOLLOWS: MMT-7040A, MMT-1884, MMT-7841ZN. TROUBLESHOOTING WAS PERFORMED. THE CUSTOMER WAS ILL AND HAD A HIGH BLOOD PRESSURE. THE INSULIN PUMP WAS IN USE AT THE TIME OF PASSING AND CUSTOMER WAS WORN SENSOR AT TIME OF EVENT. IT WAS UNKNOWN WHETHER THE CUSTOMER HAD USED THE INSULIN PUMP SYSTEM WITHIN 48 HOURS OF THE REPORTED EVENT AND UNKNOWN WHETHER THE CUSTOMER USED THE SMARTGUARD/AUTO MODE FEATURE OF THE INSULIN PUMP. NO PRODUCT RETURN IS REQUIRED FOR MMT-7040A. NO PRODUCT RETURN IS REQUIRED FOR MMT-1884. NO PRODUCT RETURN IS REQUIRED FOR MMT-7841ZN. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21613602&pc=MNS | 2518422-2025-103067 | 2025-01-23 05:00 | Death | RESPIRONICS, INC. | 2025-03-15 04:00 | MNS | BIPAP AUTOSV ADVANCED SYSTEM ONE 60 SERIES | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K090539 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO AN INFORMATION REQUEST CONCERNING DISPOSING OF A DEVICE. THE MANUFACTURER WAS INFORMED THAT THE PATIENT WAS DECEASED. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21613620&pc=BZD | 2518422-2025-103068 | 2025-03-11 04:00 | Death | RESPIRONICS, INC. | 2025-03-15 04:00 | BZD | RENTAL-DREAMSTATION AUTO CPAP, GB | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A CUSTOMER REQUESTING INFORMATION ON RETURNING A DEVICE DUE TO A PATIENT PASSING. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21603110&pc=PZX | 2184009-2025-00424 | 2025-03-03 05:00 | Death | PERFUSION SYSTEMS | 2025-03-14 04:00 | PZX | PENDITURE LAA EXCLUSION SYSTEM | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K232295 |  | 1 | Event Description: MEDTRONIC RECEIVED INFORMATION THROUGH THE CLIP-IT CLINICAL POST-MARKET STUDY, THAT FOLLOWING A PROCEDURE INVOLVING USE OF THE PENDITURE CLIP ON (B)(6) 2025, IT WAS REPORTED THAT THE PATIENT HAD DIED ON (B)(6) 2025. THE ADVERSE EVENT WAS DEEMED BY THE SITE AS POSSIBLY RELATED TO THE CLIP AND AS NOT RELATED TO THE PENDITURE DELIVERY SYSTEM DEVICE. Manufacturer Narrative: ADDITIONAL INFORMATION B5: MEDTRONIC RECEIVED ADDITIONAL INFORMATION THAT PER SITE, THE DEATH WAS NOT RELATED TO THE PENDITURE DEVICE. CORRECTION A5A: THIS FIELD HAS BEEN UPDATED. CORRECTION B5: THE CAUSE OF DEATH IS UNKNOWN. THE ADVERSE EVENT WAS DEEMED BY THE SPONSOR AS POSSIBLY RELATED TO THE PENDITURE CLIP AND AS NOT RELATED TO THE PENDITURE DELIVERY SYSTEM DEVICE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC RECEIVED ADDITIONAL INFORMATION THAT THE ADVERSE EVENT IS POSSIBLY RELATED TO THE PENDITURE CLIP AND POSSIBLY RELATED TO THE PENDITURE DELIVERY SYSTEM. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21603862&pc=MIH | 3002808486-2025-00045 | 2020-02-10 05:00 | Death | WILLIAM COOK EUROPE | 2025-03-14 04:00 | MIH | ZENITH ALPHA THORACIC ENDOVASCULAR GRAFT PROXIMAL TAPERED COMPONENTS | Adverse Event Without Identified Device or Use Problem |  |  |  | 1 | Event Description: DESCRIPTION OF EVENT ACCORDING TO STUDY: WCE-1713: ZENITH ALPHA THORACIC POST MARKET RETROSPECTIVE STUDY ON THE DAY OF THE PROCEDURE, A TYPE IA ENDOLEAK OCCURRED WHICH LED TO THE PATIENT'S DEATH. ON (B)(6) 2020, THE PATIENT WAS ROUTINELY TREATED FOR FUSIFORM AORTIC ANEURYSM WITH PLACEMENT OF A ZTA-PT-42-38-225 AND 3 NON-COOK STENTS. NO MODIFICATIONS WERE REPORTED TO THE ZTAS, HOWEVER WITH 3 STENTS PLACED AND LANDING ZONE BEGINNING AT ZONE 2. A LEFT CAROTID-SUBCLAVIAN BYPASS HAD BEEN PREVIOUSLY PERFORMED. PROCEDURE ANGIOGRAPHY REVEALED PATENT STUDY DEVICES, NO ENDOLEAK, AND NO DEVICE ISSUE. ON THE SAME DAY AS THE PROCEDURE, THE PATIENT EXPERIENCED A TYPE IA PROXIMAL ENDOLEAK. THE ADVERSE EVENT FORM INDICATES THIS WAS TREATED WITH ANTIBIOTICS, REINTUBATION, VENTILATION, AND TRACHEOSTOMY. THE EXIT FORM INDICATES AN ENDOVASCULAR PROCEDURE WAS ALSO COMPLETED.THE SITE ASSESSED THE EVENT AS NOT RELATED TO THE STUDY DEVICE OR PROCEDURE, RELATED TO THE TREATED AORTIC DISEASE. THE EVENT LED TO THE PATIENT¿S DEATH. PATIENT OUTCOME: ON (B)(6) 2020, 7 DAYS POST-PROCEDURE, THE PATIENT DIED. DESCRIPTION OF CIRCUMSTANCES SURROUNDING THE DEATH IS NOTED AS ¿MULTIPLY COMPLICATED POSTOERATIVELY WITH A PERSISTENT ENDOLEAK, DESPITE AN ENDOVASULAR REINTERVENTION." THIS CLINICAL REVIEW NURSE CANNOT RULE OUT DEVICE AND/OR PROCEDURAL RELATIONSHIP TO ENDOLEAK OR DEATH. Manufacturer Narrative: MANUFACTURERS REF# (B)(4). G4) SIMILAR TO DEVICE MARKETED UNDER PMA/510(K): P140016. INVESTIGATION IS STILL IN PROGRESS. THIS REPORT IS REQUIRED BY THE FDA UNDER 21CFR PART 803. THIS REPORT IS BASED ON UNCONFIRMED INFORMATION SUBMITTED BY OTHERS. NEITHER THE SUBMISSION OF THIS REPORT NOR ANY STATEMENT MADE IN IT IS INTENDED TO BE AN ADMISSION THAT ANY COOK DEVICE IS DEFECTIVE OR MALFUNCTIONED; THAT A DEATH OR SERIOUS INJURY OCCURRED; OR THAT ANY COOK DEVICE CAUSED OR CONTRIBUTED TO; OR IS LIKELY TO CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY IF A MALFUNCTION OCCURRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21604249&pc=DWC | 9611109-2025-00105 | 2015-02-20 05:00 | Death | LIVANOVA DEUTSCHLAND | 2025-03-14 04:00 | DWC | HEATER COOLER SYSTEM | Microbial Contamination of Device | Bacterial Infection |  |  | 1 | Event Description: LIVANOVA RECEIVED A REPORT THAT A HEATER-COOLER SYSTEM 3T DEVICE WAS USED DURING A MITRAL VALVE REPAIR AND TRICUSPID VALVE ANNULOPLASTY SURGERY PERFORMED ON (B)(6) 2015. PATIENT, SUBSEQUENTLY DEVELOPED SYMPTOMS INCLUDING NIGHT SWEATS AND FEVERS. IN 2017, PATIENT RECEIVED CORRESPONDENCE FROM LGI ADVISING RISK OF DEVELOPING M. CHIMAERA FOLLOWING SURGERY IN (B)(6) 2015. VARIOUS INVESTIGATIONS WERE PERFORMED AND IT WAS NOT UNTIL 2023 THAT M. CHIMAERA WAS IDENTIFIED FOLLOWING A SPINAL BIOPSY. IN (B)(6) 2023, PATIENT WAS ADMITTED TO (B)(6) HOSPITAL (IN (B)(6)), WHERE PATIENT WAS INFORMED OF THE NEED OF TREATMENT FOR M. CHIMAERA. TREATMENT WAS NOT SUCCESSFUL, AND PATIENT DIED ON (B)(B) 2024. IT HAS SINCE BEEN CONFIRMED BY (B)(6) HOSPITAL THAT THE HCU IN QUESTION WAS CONTAMINATED WITH M. CHIMAERA, AND IT WAS TAKEN OUT OF SERVICE IT SOME POINT IN 2015/2016. HOSPITAL NOTICE: JUNE 2015. Manufacturer Narrative: A.1.-A.5. PATIENT INFORMATION NOT AVAILABLE. D.4 UNIQUE DEVICE IDENTIFIER (UDI) NUMBER IS NOT AVAILABLE FOR THIS PRODUCT AS IT IS A CLASS II MEDICAL DEVICE MANUFACTURED BEFORE SEPTEMBER 24, 2016 WHICH IS THE FDA COMPLIANCE DATE TO IMPLEMENT UDI CODE. G.5. THE HEATER-COOLER 16-02-80 IS NOT DISTRIBUTED IN THE USA, AND IT IS SIMILAR TO HEATER-COOLER 16-02-85, WHICH IS DISTRIBUTED IN THE USA (510(K) NUMBER: K220635). H.9. LIVANOVA DEUTSCHLAND IMPLEMENTED A FIELD SAFETY NOTICE FOR DISINFECTION AND CLEANING OF HEATER-COOLER DEVICES. THE Z NUMBER IS Z-2076/2081-2015. H11: LIVANOVA DEUTSCHLAND MANUFACTURES THE HEATER-COOLER SYSTEM 3T. DHR REVIEW OF INVOLVED DEVICE SERIAL NUMBER HAS BEEN COMPLETED AND DID NOT IDENTIFY ANY DEVIATIONS OR NON-CONFORMITIES RELEVANT TO THE REPORTED ISSUE. ANALYSIS OF COMPLAINTS REPORTED BY SAME HOSPITAL DIDN'T IDENTIFY ANY EVIDENCE OF CLINICAL PRACTICES DEVIATIONS. THE INVOLVED DEVICE IN USE AT THE HOSPITAL AT THE TIME OF SURGERY ((B)(6) 2015) WAS NOT YET EQUIPPED WITH VACUUM AND SEALING KIT AND SUBSEQUENTLY SCRAPPED IN (B)(6) 2019. BASED ON ALL KNOWN FACTS, SOURCE OF PATIENT CONTAMINATION REMAINS UNKNOWN AND THE LIVANOVA DEVICE INVOLVEMENT AS WELL. LIVANOVA ALREADY IMPLEMENTED A STRATEGY TO DECREASE THE PROBABILITY OF BACTERIA GROW IN THE HC DEVICE BY APPLYING MULTIPLE MEASURES IMPLEMENTED OVER THE PAST FEW YEARS THROUGH DEDICATED CAPA AND FIELD ACTION. LIVANOVA BECAME AWARE OF THESE CASES THROUGH LEGAL ACTIONS. ACCORDING TO THE LIVANOVA LEGAL COUNSEL, THESE ARE THE ONLY AVAILABLE DETAILS AT THE MOMENT, AND WE DO NOT EXPECT TO RECEIVE ADDITIONAL INFORMATION IN THE NEAR FUTURE. THEREFORE, THE COMPANY WILL PROCEED WITH CLOSING THE CASE. HOWEVER, IF NEW INFORMATION BECOMES AVAILABLE THROUGH THE DISCOVERY PROCESS, THE COMPLAINT WILL BE REOPENED AND UPDATED ACCORDINGLY. LIVANOVA HAS A PERIODIC PROCESS FOR UPDATING LEGAL CASES IN COLLABORATION WITH ITS LEGAL COUNSEL, ENSURING THAT ANY NEW INFORMATION WILL BE COLLECTED AND MANAGED EVEN AFTER THE CASE IS CLOSED. THE RISK IS IN THE ACCEPTABLE REGION. NO SPECIFIC ACTION WAS CURRENTLY DEEMED NECESSARY. LIVANOVA WILL KEEP MONITORING THE MARKET.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21604356&pc=MIH | 3002808486-2025-00046 | 2021-05-27 04:00 | Death | WILLIAM COOK EUROPE | 2025-03-14 04:00 | MIH | ZENITH ALPHA THORACIC ENDOVASCULAR GRAFT PROXIMAL COMPONENTS | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Hematoma |  |  | 1 | Event Description: DESCRIPTION OF EVENT ACCORDING TO STUDY: (B)(6): ZENITH ALPHA THORACIC POST MARKET RETROSPECTIVE STUDY. THE PATIENT IS REPORTED TO HAVE DIED 5 DAY POST-PROCEDURE AFTER RUPTURE OF LARGE PERIPROSTHETIC HAEMATOMA. ON (B)(6) 2021, THE PATIENT WAS EMERGENTLY TREATED FOR A RUPTURED THORACIC AORTIC ANEURYSM WITH PLACEMENT OF 3 PROXIMAL ZTA COMPONENTS, IN ORDER FROM MOST PROXIMAL TO MOST DISTAL: ZTA-P-36-209, ZTA-P-38-167, ZTA-P-42-225. PROCEDURE ANGIOGRAPHY REVEALED PATENT STUDY DEVICES, NO ENDOLEAKS, AND NO DEVICE ISSUES. ON (B)(6) 2021, THE PATIENT EXPERIENCED ¿A LARGE RUPTURED PERIPROSTHETIC HAEMATOMA IN THE POSTERIOR MEDIASTINUM.¿ THE TREATMENT WAS CONVERSION TO OPEN SURGICAL REPAIR. ADDITIONAL INFORMATION PROVIDED: ¿CARDIORESPIRATORY ARREST DURING THE FIRST NIGHT OF HOSPITALISATION WITH RETURN OF SPONTANEOUS CIRCULATION AFTER 6 MINUTES OF SPECIALISED CARDIOPULMONARY REANIMATION. THE CT SCAN ON (B)(6) 2021 REVEALED A LEFT HAEMOTHORAX, ASSOCIATED WITH A LARGE PERIPROSTHETIC HAEMATOMA AROUND THE STUDY ZTA, WITH COMPRESSION OF THE LEFT LUNG. THE ZTA APPEARED TO BE PERMEABLE WITH NO LEAK OF CONTRAST MEDIUM SEEN ON THE CT SCAN. EMERGENCY OPERATIONS ON (B)(6) 2021: SUTURE OF THE DISTAL END OF THE ZTA GRAFT, FLATTENING OF THE THORACIC ANEURYSM AND SURGICAL HAEMOSTASIS, POLYTRANSFUSION. UNFAVOURABLE POST-OPERATIVE OUTCOME WITH CARDIAC, RESPIRATORY AND SEPTIC COMPLICATIONS. IN THIS CONTEXT OF MULTIVISCERAL DYSFUNCTION, A NEW EPISODE OF CARDIAC ARREST OCCURRED ON (B)(6) 2021, WHICH LED TO THE PATIENTS DEATH AT 3:37 AM.¿ THE SITE DID NOT ANSWER THE QUESTIONS ABOUT THE RELATIONSHIP OF THE EVENT TO THE STUDY DEVICE OR PROCEDURE. ADDITIONAL INFORMATION RECEIVED 10MAR2025: IMAGING DATA FOR THE (B)(6) 2021 CT SCAN (DAY AFTER PROCEDURE). THE CT SHOWS THAT AT LEAST ONE STUDY DEVICE WAS NOT PATENT (THE SITE CANNOT PROVIDE WHICH), NO EVIDENCE OF THROMBUS, NO DEVICE ISSUE, AND AN UNKNOWN TYPE OF ENDOLEAK THAT REQUIRED SECONDARY INTERVENTION. SECONDARY INTERVENTION OCCURRED ON (B)(6) 2021. IT WAS NOT CONSIDERED SUCCESSFUL, NOTING ¿POST-OPERATIVE CT SCAN SHOWS PERSISTENT HAEMATOMA.¿ THE REASON FOR THE SECONDARY INTERVENTION WAS CLINICAL SIGNS/SYMPTOMS, SPECIFICALLY CARDIORESPIRATORY ARREST, AND IMAGING RESULTS. THE INDICATION FOR SECONDARY INTERVENTION WAS ¿LEFT HAEMOTHORAX AND RUPTURED PERIPROSTHETIC HAEMATOMA IN THE POSTERIOR MEDIASTINUM.¿ TYPE OF INTERVENTION WAS SURGICAL CONVERSION TO OPEN REPAIR BECAUSE OF AN ENDOLEAK. THE STUDY DEVICES WERE NOT EXPLANTED. CARDIORESPIRATORY ARREST DURING THE FIRST NIGHT OF HOSPITALISATION WITH RETURN OF SPONTANEOUS CIRCULATION AFTER 6 MINUTES OF SPECIALISED CARDIOPULMONARY REANIMATION. THE CT SCAN ON (B)(6) 2021 REVEALED A LEFT HAEMOTHORAX, ASSOCIATED WITH A LARGE PERIPROSTHETIC HAEMATOMA AROUND THE STUDY ZTA, THIS IS DUE TO A PROSTHETIC LEAK. EMERGENCY OPERATIONS ON (B)(6) 2021. PATIENT OUTCOME: THE PATIENT DIED ON (B)(6) 2021. THE CAUSE OF DEATH WAS ¿COMPLICATIONS OF A RUPTURED ANEURYSM OF THE DESCENDING THORACIC AORTA (INITIAL PATHOLOGY).¿ THE DEATH WAS NOTED AS RELATED TO PRIMARY DISEASE PROGRESSION AND COMPLICATIONS OF THE IMPLANT PROCEDURE.Event Description: NO ADDITIONAL INFORMATION REGARDING THE PATIENT AND/OR EVENT HAS BEEN RECEIVED SINCE THE PREVIOUS MEDWATCH REPORT WAS SENT. Manufacturer Narrative: MANUFACTURERS REF# (B)(4). BLANK FIELDS ON THIS FORM INDICATE THE INFORMATION IS UNKNOWN OR UNAVAILABLE. D4) CATALOG# IS UNKNOWN BUT REFERRED TO AS EITHER ZTA-P-36-209, ZTA-P-42-225 OR ZTA-P-38-167. G4) SIMILAR TO DEVICE MARKETED UNDER PMA/510(K): P140016. INVESTIGATION IS STILL IN PROGRESS. THIS REPORT IS REQUIRED BY THE FDA UNDER 21CFR PART 803. THIS REPORT IS BASED ON UNCONFIRMED INFORMATION SUBMITTED BY OTHERS. NEITHER THE SUBMISSION OF THIS REPORT NOR ANY STATEMENT MADE IN IT IS INTENDED TO BE AN ADMISSION THAT ANY COOK DEVICE IS DEFECTIVE OR MALFUNCTIONED; THAT A DEATH OR SERIOUS INJURY OCCURRED; OR THAT ANY COOK DEVICE CAUSED OR CONTRIBUTED TO; OR IS LIKELY TO CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY IF A MALFUNCTION OCCURRED. Manufacturer Narrative: MANUFACTURER¿S REF# (B)(4). AFTER INVESTIGATION THE EVENT IS NO LONGER CONSIDERED REPORTABLE TO FDA. INVESTIGATION IS STILL IN PROGRESS. COOK MEDICAL WILL CONTINUE TO MONITOR FOR SIMILAR EVENTS. THIS REPORT IS REQUIRED BY THE FDA UNDER 21 CFR PART 803. THIS REPORT IS BASED ON UNCONFIRMED INFORMATION SUBMITTED BY OTHERS. NEITHER THE SUBMISSION OF THIS REPORT NOR ANY STATEMENT MADE IN IT IS INTENDED TO BE AN ADMISSION THAT ANY COOK DEVICE IS DEFECTIVE OR MALFUNCTIONED; THAT A DEATH OR SERIOUS INJURY OCCURRED; OR THAT ANY COOK DEVICE CAUSED OR CONTRIBUTED TO; OR IS LIKELY TO CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY IF A MALFUNCTION OCCURRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21605517&pc=MVK | 3008642652-2025-02692 | 2024-11-29 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-03-14 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Signal Artifact/Noise | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2024 WHILE REPORTEDLY WEARING THE LIFEVEST. THE DEVICE WAS STARTED UP AT 11:44:00 ON (B)(6) 2024. AT 11:56:31 ON (B)(6) 2024, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF WITH VARYING AMPLITUDES AND MOTION ARTIFACT. THE DETECTIVE PROPERLY DETECTED VF. VARYING AMPLITUDE AND MOTION ARTIFACT PREVENTED THE LIFEVEST FROM TREATING THE PATIENT. THE ELECTRODE BELT WAS DISCONNECTED AT 12:04:45 ON (B)(6) 2024. Manufacturer Narrative: INCOMING TESTING OF THE ELECTRODE BELT HAS BEEN COMPLETED. THE ELECTRODE BELT ECG ACQUISITION AND PULSE DELIVERY CIRCUITRY WAS TESTED AND FOUND TO BE FULLY FUNCTIONAL. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION. THE MONITOR HAS BEEN RETURNED AND THE EVALUATION IS UNDERWAY. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21605522&pc=OZD | 1220648-2025-26722 | 2025-02-17 05:00 | Death | ABIOMED, INC. | 2025-03-14 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Hematoma; Hemorrhage/Bleeding; Tachycardia; Cardiac Tamponade; Pseudoaneurysm; Vascular Dissection | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WAS ELECTIVELY IMPLANTED WITH AN IMPELLA 5.5 DEVICE FOR MECHANICAL CIRCULATORY SUPPORT AND EXPERIENCED A VASCULAR DISSECTION WITH A DEMISED OUTCOME. A COMPUTED TOMOGRAPHIC ANGIOGRAPHY OF THE CHEST WAS COMPLETED AFTER THE DEVELOPMENT OF SPONTANEOUS TAMPONADE APPROXIMATELY 17 DAYS AFTER IMPELLA IMPLANT. FINDINGS REVEALED A 2.2 CENTIMETER CONTAINED LEAK ARISING FROM THE ANTERIOR WALL OF THE DISTAL ASCENDING AORTA IMMEDIATELY SUPERIOR TO THE ENTRY POINT OF THE IMPELLA DEVICE WITH A SURROUNDING HEMATOMA. THE DECISION MADE TO TAKE THE PATIENT TO THE OPERATING ROOM FOR CHEST EXPLORATION AND EXPLANT OF THE IMPELLA 5.5 DEVICE. UPON RETRACTING THE CHEST OPEN, A LEAK WAS NOTED IN THE AORTA. THE PATIENT WAS EMERGENTLY CANNULATED FOR CARDIOPULMONARY BYPASS AND WAS NOTED TO HAVE AN EXTENSIVE AORTIC DISSECTION. THE PATIENT EXPIRED ON THE OPERATING ROOM TABLE. THE PHYSICIAN'S COMMENTS AFTERWARDS NOTED UNKNOWN CAUSE WHY THIS EVENT OCCURRED. THE PHYSICIAN NOTED THE POSSIBILITY OF INFECTION AT THE SITE THOUGH NO OBVIOUS SIGNS OF INFECTION WERE NOTED. EXTENSIVE REPAIR WAS REQUIRED WHEN THE PATIENT WAS PLACED ON CARDIOPULMONARY BYPASS. THE DISSECTION EXTENDED TO WHERE NO FLOW WAS GETTING TO HEAD VESSELS WHILE ATTEMPTING TO PUT ON CIRCULATORY ARREST.Event Description: A NOTIFICATION WAS RECEIVED VIA ABIOMED¿S LONG-TERM OUTCOME AND QUALITY INDICATOR (LOQI) IMPELLA REGISTRY REGARDING A PATIENT THAT ENDURED VENTRICULAR TACHYCARDIA WITH A "POSSIBLE" RELATIONSHIP TO THE IMPELLA DEVICE USED: ¿POST CABG AND IMPELLA IMPLANT, SUBJECT HAD MULTIPLE RUNS OF VT REQUIRING CARDIOVERSION X2.¿ THE PATIENT RECOVERED WITH NO SEQUELAE. ADDITIONAL INFORMATION RECEIVED: LOQI REPORTED THAT A CTA CHEST SHOWED A LEAK FROM THE IMPELLA'S GRAFT'S ENTRY POINT. THE PATIENT TAKEN TO OR FOR IMPELLA REMOVAL. ¿DURING IMPELLA REMOVAL, SUBJECT'S AORTA BEGAN BLEEDING PROFUSELY. IMPELLA REMOVED AND A LARGE PSEUDOANEURYSM WAS FOUND IN AORTA. THIS DISSECTED DOWN TOWARDS THE ROOT OF THE AORTA. SUBJECT EXPIRED IN THE OR.¿ THE PATIENT ENDURED CARDIAC PSEUDOANEURYSM WITH A "POSSIBLE" RELATIONSHIP TO THE IMPELLA DEVICE USED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. Manufacturer Narrative: THE INVESTIGATION OF VASCULAR DAMAGE - GREAT VESSEL HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. Manufacturer Narrative: B.5 ADDITIONAL INFORMATION WAS RECEIVED. G.2 ADDED REPORT SOURCE DUE TO NEW INFORMATION RECEIVED. H.6 ADDITIONAL CODED WERE ADDED DUE TO NEW INFORMATION RECEIVED. H.10 ADDED INVESTIGATION RESULTS FOR VENTRICULAR TACHYCARDIA ISSUE. THE INVESTIGATION FOR VENTRICULAR TACHYCARDIA ISSUE HAS BEEN COMPLETED. THE ROOT CAUSE OF THE VENTRICULAR TACHYCARDIA ISSUE WAS NOT DETERMINED SINCE THE PRODUCT WAS NOT RETURNED AND INSUFFICIENT CLINICAL INFORMATION WAS PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21605708&pc=NPT | 2025587-2025-01897 | 2024-08-13 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-14 04:00 | NPT | EVOLUT R TRANSCATHETER AORTIC VALVE | Insufficient Information | Insufficient Information | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 5 YEARS, 6 MONTHS, AND 19 DAYS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT DIED. THE CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21605877&pc=CBK | 2518422-2025-103022 | 2025-03-10 04:00 | Death | RESPIRONICS, INC. | 2025-03-14 04:00 | CBK | TRILOGY EVO, JAPAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K181166 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT PASSED AWAY WHILE USING A TRILOGY EVO DEVICE. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ALTHOUGH THE CAUSE OF THE DEATH WAS DETERMINED TO BE PROGRESSION OF ALS, IT IS UNKNOWN ABOUT THE CAUSAL RELATIONSHIP WITH THIS DEVICE. THE PATIENT USED THE DEVICE FOR APPROXIMATELY 8 MONTHS. PRELIMINARY EVALUATION FROM THE SALES REPRESENTATIVE REVIEW OF CARE ORCHESTRATOR, SHOWS "THE LAST USE DATE/TIME WAS FROM THE EVENING OF (B)(6) 2025 TO THE MORNING OF (B)(6) (2025). (THE SALES REPRESENTATIVE'S OPINION IS) PROBABLY, THE DEVICE REMAINED STANDBY, AND THEY DID NOT PRESS START THERAPY". THE COMPLAINT REVIEW HAS DETERMINED THAT DUE TO INSUFFICIENT INFORMATION THIS WILL BE REPORTED. ALTHOUGH THERE IS NOT AN ALLEGATION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE MENTIONED "DEATH". THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER PREVIOUSLY REPORTED INFORMATION RECEIVED THAT A PATIENT PASSED AWAY WHILE USING A TRILOGY EVO DEVICE. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ALTHOUGH THE CAUSE OF THE DEATH WAS DETERMINED TO BE PROGRESSION OF ALS, IT IS UNKNOWN ABOUT THE CAUSAL RELATIONSHIP WITH THIS DEVICE. THE PATIENT USED THE DEVICE FOR APPROXIMATELY 8 MONTHS. PRELIMINARY EVALUATION FROM THE SALES REPRESENTATIVE REVIEW OF CARE ORCHESTRATOR, SHOWS "THE LAST USE DATE/TIME WAS FROM THE EVENING OF (B)(6) TO THE MORNING OF (B)(6) (2025). (THE SALES REPRESENTATIVE'S OPINION IS) PROBABLY, THE DEVICE REMAINED STANDBY, AND THEY DID NOT PRESS START THERAPY." THE COMPLAINT REVIEW HAS DETERMINED THAT DUE TO INSUFFICIENT INFORMATION THIS WILL BE REPORTED. ALTHOUGH THERE IS NOT AN ALLEGATION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE MENTIONED "DEATH." UPON EVALUATION OF THE DEVICE THE EVENT LOG CONFIRMED THAT THE DEVICE WAS TURNED OFF AT 08:13 ON (B)(6) AND REMAINED INACTIVE UNTIL IT ENTERED STANDBY MODE AT 03:36 ON (B)(6). AN OPERATIONAL CHECK WAS PERFORMED USING A BREATHING CIRCUIT, BUT NO FAILURES WERE CONFIRMED. FUNCTIONAL TESTS OF THE PRESSURE SENSOR AND FLOW SENSOR WERE PERFORMED, AND NO FAILURES WERE CONFIRMED. THE DEVICE WAS DISASSEMBLED, AND AN INTERNAL INSPECTION WAS CARRIED OUT, BUT NO FAILURES WERE CONFIRMED. BASED ON THE ABOVE RESULTS, NO FAILURES WERE CONFIRMED WITH THE DEVICE, AND IT HAS BEEN DETERMINED TO BE OPERATING NORMALLY. PERFORMED FINAL TEST, BATTERY CHECK, VALVE CHECK, RUN IN TESTS AND CLEANING THEN CONFIRMED THE UNIT OPERATED PROPERLY. SOFTWARE VERSION WAS CHECKED (VER.1.06.10.00).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21605893&pc=NPT | 2025587-2025-01899 | 2024-03-07 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-14 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Intracranial Hemorrhage | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 9 MONTHS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, A CEREBRAL HEMORRHAGE WAS NOTED. THE SAME DAY, THE PATIENT DIED. THE CAUSE OF DEATH WAS CEREBRAL HEMORRHAGE. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21605898&pc=NPT | 2025587-2025-01900 | 2024-12-24 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-14 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Insufficient Information | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 11 MONTHS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT DIED. THE CAUSE OF DEATH WAS UNKNOWN. PER THE PHYSICIAN, THERE WAS NO RELATION BETWEEN THE VALVE NOR THE IMPLANT PROCEDURE AND THE DEATH. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21605928&pc=MIH | 3002808486-2025-00047 | 2021-03-21 04:00 | Death | WILLIAM COOK EUROPE | 2025-03-14 04:00 | MIH | ZENITH ALPHA THORACIC ENDOVASCULAR GRAFT PROXIMAL COMPONENTS | Adverse Event Without Identified Device or Use Problem | Shock; Ruptured Aneurysm |  |  | 1 | Event Description: DESCRIPTION OF EVENT ACCORDING TO STUDY: (B)(6): ZENITH ALPHA THORACIC POST MARKET RETROSPECTIVE STUDY ON THE DAY OF THE PROCEDURE, THE PATIENT DIED AFTER EXPERIENCING HEMORRHAGIC SHOCK. ON (B)(6) 2021, THIS 68-YEAR-OLD FEMALE PATIENT WAS URGENTLY TREATED FOR FUSIFORM THORACIC AORTIC ANEURYSM WITH PLACEMENT OF A ZTA-P-36-161 AND A ZDEG-PT-38-30-199-PF. THE PATIENT WAS HEMODYNAMICALLY STABLE PRIOR TO THE PROCEDURE. NO PRE-PROCEDURE OR PROCEDURE IMAGING DATA HAVE BEEN ENTERED. ON THE DAY OF THE PROCEDURE, THE PATIENT EXPERIENCED HEMORRHAGIC SHOCK, SUSPECTED AORTIC RUPTURE DUE TO THORACO ABDOMINAL AORTIC ANEURYSM (TAAA). RESUSCITATION MEASURES WERE UNSUCCESSFUL AND THE PATIENT DIED. PATIENT OUTCOME: ON (B)(6) 2021, THE PATIENT DIED. THE CAUSE OF DEATH IS NOTED AS ¿HEMORRHAGIC SHOCK SOME HOURS AFTER PROCEDURE. SUSPECTED AORTIC RUPTURE DUE TO TAAA.¿ CLINICAL REVIEW NURSE CANNOT RULE OUT RELATIONSHIP OF DEATH TO DEVICE AND/OR PROCEDURE. Manufacturer Narrative: MANUFACTURERS REF# (B)(4). BLANK FIELDS ON THIS FORM INDICATE THE INFORMATION IS UNKNOWN OR UNAVAILABLE. (G4) SIMILAR TO DEVICE MARKETED UNDER PMA/510(K): P140016. INVESTIGATION IS STILL IN PROGRESS. THIS REPORT IS REQUIRED BY THE FDA UNDER 21CFR PART 803. THIS REPORT IS BASED ON UNCONFIRMED INFORMATION SUBMITTED BY OTHERS. NEITHER THE SUBMISSION OF THIS REPORT NOR ANY STATEMENT MADE IN IT IS INTENDED TO BE AN ADMISSION THAT ANY COOK DEVICE IS DEFECTIVE OR MALFUNCTIONED; THAT A DEATH OR SERIOUS INJURY OCCURRED; OR THAT ANY COOK DEVICE CAUSED OR CONTRIBUTED TO; OR IS LIKELY TO CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY IF A MALFUNCTION OCCURRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21605965&pc=NPT | 2025587-2025-01903 | 2023-12-04 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-14 04:00 | NPT | EVOLUT PRO PLUS VALVE | Insufficient Information | Insufficient Information | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 10 MONTHS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT DIED. THE CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21606235&pc=OZD | 1220648-2025-26724 | 2025-02-18 05:00 | Death | ABIOMED, INC. | 2025-03-14 04:00 | OZD | IMPELLA CP | Positioning Failure | Cardiovascular Insufficiency | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. THE DEVICE WAS REMOVED BECAUSE IT HAD BECOME DISLODGED INTO THE AORTA. A NEW IMPELLA CP WAS PLACED. THE ISSUE OCCURRED UPON MOVING THE BED FROM THE CATHETER TABLE. IT WAS NOTED THAT THE FIXING RING MAY HAVE BEEN FORGOTTEN TO BE CLOSED. THE PATIENT LATER EXPIRED AFTER EIGHT DAYS OF SUPPORT. MEDICAL SAFETY REVIEW OF THE EVENT NOTED THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA DEVICE AS AN ASSOCIATED FACTOR IN THE PATIENT'S DEMISE. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21606271&pc=LYJ | 1644487-2025-00288 | 2025-02-01 05:00 | Death | CYBERONICS - HOUSTON | 2025-03-14 04:00 | LYJ | PULSE GEN MODEL 103 | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P970003 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD PASSED AWAY DUE TO AN UNKNOWN REASON. THE PATIENT DEVICE WAS NOTED TO BE AT END OF SERVICE, PRIOR TO THE DATE OF DEATH. NO OTHER RELEVANT INFORMATION HAS BEEN RECEIVED TO DATE. Manufacturer Narrative: LIVANOVA USA, INC. SUBMITS THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION, BASED ON INFORMATION THAT LIVANOVA HAS OBTAINED, BUT MAY NOT HAVE BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION, OR A CONCLUSION BY FDA OR ANYONE ELSE, THAT THE DEVICE, LIVANOVA, OR LIVANOVA'S EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS¿ OR "MALFUNCTIONS¿. THESE WORDS ARE INCORPORATED INTO THE FDA 3500A MEDWATCH FORM BY THE FDA, AND LIVANOVA OBJECTS TO THEIR USE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21606397&pc=NIQ | 2124215-2025-14998 | 2025-02-23 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-03-14 04:00 | NIQ | PROMUS ELITE | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Ventilator Dependent; Coma |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. THE TARGET LESION WAS LOCATED IN THE LEFT ANTERIOR DESCENDING ARTERY (LAD) AND LEFT CIRCUMFLEX CORONARY ARTERY (LCX). A 38 X 3.00MM PROMUS ELITE MR DRUG-ELUTING STENT WAS ADVANCED FOR TREATMENT AND DEPLOYED IN THE LAD AND THE SECOND 2.75 X 48MM SYNERGY DRUG-ELUTING STENT WAS DEPLOYED IN THE LCX. AFTER STENTING, PATIENT WAS TRANSFERRED TO CCU. AFTER SOME TIME, THEY OBSERVED DEVELOPING HYPOTENSION AND THE PATIENT BECOMES UNCONSCIOUS THEN THEY PUT THE PATIENT ON A VENTILATOR AND AFTER A WHILE PATIENT EXPIRED.Event Description: IT WAS REPORTED THAT THE PATIENT DIED. THE TARGET LESION WAS LOCATED IN THE LEFT ANTERIOR DESCENDING ARTERY (LAD) AND LEFT CIRCUMFLEX CORONARY ARTERY (LCX). A 38 X 3.00MM PROMUS ELITE MR DRUG-ELUTING STENT WAS ADVANCED FOR TREATMENT AND DEPLOYED IN THE LAD AND THE SECOND 2.75 X 48MM SYNERGY DRUG-ELUTING STENT WAS DEPLOYED IN THE LCX. AFTER STENTING, THE PATIENT WAS TRANSFERRED TO CORONARY CARE UNIT. AFTER SOME TIME, THEY DEVELOPED HYPOTENSION AND BECAME UNCONSCIOUS. THE PATIENT WAS PLACED ON A VENTILATOR AND SUBSEQUENTLY EXPIRED. IT WAS FURTHER REPORTED THAT THE TARGET LESION WAS 95% STENOSED, MILDLY CALCIFIED AND NON-TORTUOUS. THE STENT WAS SUCCESSFULLY DEPLOYED. NO AUTOPSY WAS PERFORMED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21606407&pc=NIQ | 2124215-2025-15001 | 2025-02-23 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-03-14 04:00 | NIQ | SYNERGY | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Ventilator Dependent; Coma |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. THE TARGET LESION WAS LOCATED IN THE LEFT ANTERIOR DESCENDING ARTERY (LAD) AND LEFT CIRCUMFLEX CORONARY ARTERY (LCX). A 38 X 3.00MM PROMUS ELITE MR DRUG-ELUTING STENT WAS ADVANCED FOR TREATMENT AND DEPLOYED IN THE LAD AND THE SECOND 2.75 X 48MM SYNERGY DRUG-ELUTING STENT WAS DEPLOYED IN THE LCX. AFTER STENTING, THE PATIENT WAS TRANSFERRED TO CORONARY CARE UNIT. AFTER SOME TIME, THEY DEVELOPED HYPOTENSION AND BECAME UNCONSCIOUS. THE PATIENT WAS PLACED ON A VENTILATOR AND SUBSEQUENTLY EXPIRED.Event Description: IT WAS REPORTED THAT THE PATIENT DIED. THE TARGET LESION WAS LOCATED IN THE LEFT ANTERIOR DESCENDING ARTERY (LAD) AND LEFT CIRCUMFLEX CORONARY ARTERY (LCX). A 38 X 3.00MM PROMUS ELITE MR DRUG-ELUTING STENT WAS ADVANCED FOR TREATMENT AND DEPLOYED IN THE LAD AND THE SECOND 2.75 X 48MM SYNERGY DRUG-ELUTING STENT WAS DEPLOYED IN THE LCX. AFTER STENTING, THE PATIENT WAS TRANSFERRED TO CORONARY CARE UNIT. AFTER SOME TIME, THEY DEVELOPED HYPOTENSION AND BECAME UNCONSCIOUS. THE PATIENT WAS PLACED ON A VENTILATOR AND SUBSEQUENTLY EXPIRED. IT WAS FURTHER REPORTED THAT THE TARGET LESION WAS 95% STENOSED, MILDLY CALCIFIED AND NON-TORTUOUS. THE STENT WAS SUCCESSFULLY DEPLOYED. NO AUTOPSY WAS PERFORMED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21606445&pc=NKM | 2135147-2025-01277 | 2016-01-01 05:00 | Death | ABBOTT MEDICAL | 2025-03-14 04:00 | NKM | MITRACLIP® SYSTEM CLIP DELIVERY SYSTEM | Off-Label Use | Stroke/CVA; Hemorrhage/Bleeding; Intracranial Hemorrhage; Renal Failure; Cardiac Tamponade; Cardiogenic Shock; Respiratory Failure; Pericardial Effusion; Heart Failure/Congestive Heart Failure; Tricuspid Valve Insufficiency/ Regurgitation | P100009 |  | 1 | Event Description: THE ARTICLE "COMPARISON OF TRANSCATHETER VERSUS SURGICAL TRICUSPID REPAIR AMONG PATIENTS WITH TRICUSPID REGURGITATION: TWO-YEAR RESULTS" WAS REVIEWED. THE ARTICLE PRESENTED A PROSPECTIVE MULTI CENTER STUDY, TO COMPARE OUTCOMES OF PATIENTS WITH TRICUSPID REGURGITATION (TR) TREATED WITH TRICUSPID TRANSCATHETER EDGE-TO-EDGE REPAIR (T-TEER) VERSUS SURGICAL TRICUSPID VALVE REPAIR. DEVICES MENTIONED INCLUDE TRICLIP, MITRACLIP, AND PASCAL. THE ARTICLE CONCLUDED THAT AMONG MEDICARE BENEFICIARIES, 2-YEAR MORTALITY AND HEART FAILURE HOSPITALIZATION RATES WERE COMPARABLE BETWEEN THE 2 GROUPS. NOTABLY, T-TEER SHOWED ADVANTAGES IN PERIOPERATIVE OUTCOMES, INCLUDING IN-HOSPITAL MORTALITY AND PACEMAKER IMPLANTATION RATES WHILE TRICUSPID VALVE REINTERVENTIONS WERE MORE FREQUENT IN THE T-TEER GROUP. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS TOSHIKI KUNO, AT N, MASSACHUSETTS GENERAL HOSPITAL AND HARVARD MEDICAL SCHOOL, BOSTON, USA WITH CORRESPONDING EMAIL TKUNO@MGH.HARVARD.EDU]. THE TIME FRAME OF THE STUDY WAS JULY 2016 TO DECEMBER 2020. A TOTAL OF 1143 (409 T-TEER) PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH AN UNKNOWN AMOUNT RECEIVED AN ABBOTT DEVICE. FOR T-TEER GROUP, AVERAGE AGE WAS 81 AND MAJORITY GENDER WAS FEMALE. COMORBIDITIES INCLUDED HYPERTENSION, HYPERLIPIDEMIA, DIABETES, MYOCARDIAL INFARCTION, ISCHEMIC HEART DISEASE, CONGESTIVE HEART FAILURE, ATRIAL FIBRILLATION, CHRONIC KIDNEY DISEASE, STROKE/TIA, COPD, ASTHMA, LIVER DISEASE, ANEMIA, PERIPHERAL VASCULAR DISEASE, DEMENTIA, DEPRESSION, CANCER, ARTHRITIS, OBESITY, FUNCTIONAL TRICUSPID REGURGITATION, SECONDARY TRICUSPID REGURGITATION, RHEUMATIC DISEASE, PACEMAKER, PULMONARY HYPERTENSION, PRIOR CARDIAC SURGERY. PERI AND POST-PROCEDURAL COMPLICATIONS INCLUDED HOSPITALIZATION, REINTERVENTION WITH T-TEER, RECURRENT TRICUSPID REGURGITATION, DEATH, ISCHEMIC STROKE, INTRACEREBRAL HEMORRHAGE, ACUTE KIDNEY INJURY, CARDIOGENIC SHOCK, PERMANENT PACEMAKER, RESPIRATORY COMPLICATIONS, POSTOPERATIVE HEMORRHAGE, TRANSFUSION, PERICARDIAL EFFUSION, CARDIAC TAMPONADE, HEART FAILURE.Event Description: THE ARTICLE "COMPARISON OF TRANSCATHETER VERSUS SURGICAL TRICUSPID REPAIR AMONG PATIENTS WITH TRICUSPID REGURGITATION: TWO-YEAR RESULTS" WAS REVIEWED. THE ARTICLE PRESENTED A PROSPECTIVE MULTI CENTER STUDY, TO COMPARE OUTCOMES OF PATIENTS WITH TRICUSPID REGURGITATION (TR) TREATED WITH TRICUSPID TRANSCATHETER EDGE-TO-EDGE REPAIR (T-TEER) VERSUS SURGICAL TRICUSPID VALVE REPAIR. DEVICES MENTIONED INCLUDE TRICLIP, MITRACLIP, AND PASCAL. THE ARTICLE CONCLUDED THAT AMONG MEDICARE BENEFICIARIES, 2-YEAR MORTALITY AND HEART FAILURE HOSPITALIZATION RATES WERE COMPARABLE BETWEEN THE 2 GROUPS. NOTABLY, T-TEER SHOWED ADVANTAGES IN PERIOPERATIVE OUTCOMES, INCLUDING IN-HOSPITAL MORTALITY AND PACEMAKER IMPLANTATION RATES WHILE TRICUSPID VALVE REINTERVENTIONS WERE MORE FREQUENT IN THE T-TEER GROUP. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS TOSHIKI KUNO, AT N, MASSACHUSETTS GENERAL HOSPITAL AND HARVARD MEDICAL SCHOOL, BOSTON, USA WITH CORRESPONDING EMAIL TKUNO@MGH.HARVARD.EDU]. THE TIME FRAME OF THE STUDY WAS JULY 2016 TO DECEMBER 2020. A TOTAL OF 1143 (409 T-TEER) PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH AN UNKNOWN AMOUNT RECEIVED AN ABBOTT DEVICE. FOR T-TEER GROUP, AVERAGE AGE WAS 81 AND MAJORITY GENDER WAS FEMALE. COMORBIDITIES INCLUDED HYPERTENSION, HYPERLIPIDEMIA, DIABETES, MYOCARDIAL INFARCTION, ISCHEMIC HEART DISEASE, CONGESTIVE HEART FAILURE, ATRIAL FIBRILLATION, CHRONIC KIDNEY DISEASE, STROKE/TIA, COPD, ASTHMA, LIVER DISEASE, ANEMIA, PERIPHERAL VASCULAR DISEASE, DEMENTIA, DEPRESSION, CANCER, ARTHRITIS, OBESITY, FUNCTIONAL TRICUSPID REGURGITATION, SECONDARY TRICUSPID REGURGITATION, RHEUMATIC DISEASE, PACEMAKER, PULMONARY HYPERTENSION, PRIOR CARDIAC SURGERY. PERI AND POST-PROCEDURAL COMPLICATIONS INCLUDED HOSPITALIZATION, REINTERVENTION WITH T-TEER, RECURRENT TRICUSPID REGURGITATION, DEATH, ISCHEMIC STROKE, INTRACEREBRAL HEMORRHAGE, ACUTE KIDNEY INJURY, CARDIOGENIC SHOCK, PERMANENT PACEMAKER, RESPIRATORY COMPLICATIONS, POSTOPERATIVE HEMORRHAGE, TRANSFUSION, PERICARDIAL EFFUSION, CARDIAC TAMPONADE, HEART FAILURE. Manufacturer Narrative: IT IS UNKNOWN IF THE DEVICE IS RETURNING FOR ANALYSIS. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. B3 EVENT DATE IS ESTIMATED. THE ADDITIONAL PATIENT OUTCOME OF DEATH IS FILED IN A SEPARATE REPORT. LITERATURE ATTACHMENT: COMPARISON OF TRANSCATHETER VERSUS SURGICAL TRICUSPID REPAIR AMONG PATIENTS WITH TRICUSPID REGURGITATION: TWO-YEAR RESULTS Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF MITRACLIP WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING HYPERTENSION, HYPERLIPIDEMIA, DIABETES, MYOCARDIAL INFARCTION, ISCHEMIC HEART DISEASE, CONGESTIVE HEART FAILURE, ATRIAL FIBRILLATION, CHRONIC KIDNEY DISEASE, STROKE/TIA, COPD, ASTHMA, LIVER DISEASE, ANEMIA, PERIPHERAL VASCULAR DISEASE, DEMENTIA, DEPRESSION, CANCER, ARTHRITIS, OBESITY, FUNCTIONAL TRICUSPID REGURGITATION, SECONDARY TRICUSPID REGURGITATION, RHEUMATIC DISEASE, PACEMAKER, PULMONARY HYPERTENSION, PRIOR CARDIAC SURGERY. COMPLICATIONS REPORTED HOSPITALIZATION, REINTERVENTION WITH T-TEER, RECURRENT TRICUSPID REGURGITATION, DEATH, ISCHEMIC STROKE, INTRACEREBRAL HEMORRHAGE, ACUTE KIDNEY INJURY, CARDIOGENIC SHOCK, PERMANENT PACEMAKER, RESPIRATORY COMPLICATIONS, POSTOPERATIVE HEMORRHAGE, TRANSFUSION, PERICARDIAL EFFUSION, CARDIAC TAMPONADE, HEART FAILURE, OFF-LABEL USE; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. THE REPORTED OFF-LABEL USE WAS ASSOCIATED WITH USING A MITRACLIP ON A TRICUSPID VALVE. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. THEREFORE, NO CORRECTIVE ACTION IS REQUIRED. LITERATURE: ARTICLE TITLE: COMPARISON OF TRANSCATHETER VERSUS SURGICAL TRICUSPID REPAIR AMONG PATIENTS WITH TRICUSPID REGURGITATION: TWO-YEAR RESULTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21606479&pc=NPS | 2135147-2025-01280 | 2016-01-01 05:00 | Death | ABBOTT MEDICAL | 2025-03-14 04:00 | NPS | TRICLIP DELIVERY SYSTEM (TCDS) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P230007 |  | 1 | Event Description: THE ARTICLE "COMPARISON OF TRANSCATHETER VERSUS SURGICAL TRICUSPID REPAIR AMONG PATIENTS WITH TRICUSPID REGURGITATION: TWO-YEAR RESULTS" WAS REVIEWED. THE ARTICLE PRESENTED A PROSPECTIVE MULTI CENTER STUDY, TO COMPARE OUTCOMES OF PATIENTS WITH TRICUSPID REGURGITATION (TR) TREATED WITH TRICUSPID TRANSCATHETER EDGE-TO-EDGE REPAIR (T-TEER) VERSUS SURGICAL TRICUSPID VALVE REPAIR. DEVICES MENTIONED INCLUDE TRICLIP, MITRACLIP, AND PASCAL. THE ARTICLE CONCLUDED THAT AMONG MEDICARE BENEFICIARIES, 2-YEAR MORTALITY AND HEART FAILURE HOSPITALIZATION RATES WERE COMPARABLE BETWEEN THE 2 GROUPS. NOTABLY, T-TEER SHOWED ADVANTAGES IN PERIOPERATIVE OUTCOMES, INCLUDING IN-HOSPITAL MORTALITY AND PACEMAKER IMPLANTATION RATES WHILE TRICUSPID VALVE REINTERVENTIONS WERE MORE FREQUENT IN THE T-TEER GROUP. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS TOSHIKI KUNO, AT N, MASSACHUSETTS GENERAL HOSPITAL AND HARVARD MEDICAL SCHOOL, BOSTON, USA WITH CORRESPONDING EMAIL TKUNO@MGH.HARVARD.EDU]. THE TIME FRAME OF THE STUDY WAS JULY 2016 TO DECEMBER 2020. A TOTAL OF 1143 (409 T-TEER) PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH AN UNKNOWN AMOUNT RECEIVED AN ABBOTT DEVICE. FOR T-TEER GROUP, AVERAGE AGE WAS 81 AND MAJORITY GENDER WAS FEMALE. COMORBIDITIES INCLUDED HYPERTENSION, HYPERLIPIDEMIA, DIABETES, MYOCARDIAL INFARCTION, ISCHEMIC HEART DISEASE, CONGESTIVE HEART FAILURE, ATRIAL FIBRILLATION, CHRONIC KIDNEY DISEASE, STROKE/TIA, COPD, ASTHMA, LIVER DISEASE, ANEMIA, PERIPHERAL VASCULAR DISEASE, DEMENTIA, DEPRESSION, CANCER, ARTHRITIS, OBESITY, FUNCTIONAL TRICUSPID REGURGITATION, SECONDARY TRICUSPID REGURGITATION, RHEUMATIC DISEASE, PACEMAKER, PULMONARY HYPERTENSION, PRIOR CARDIAC SURGERY. PERI AND POST-PROCEDURAL COMPLICATIONS INCLUDED HOSPITALIZATION, REINTERVENTION WITH T-TEER, RECURRENT TRICUSPID REGURGITATION, DEATH, ISCHEMIC STROKE, INTRACEREBRAL HEMORRHAGE, ACUTE KIDNEY INJURY, CARDIOGENIC SHOCK, PERMANENT PACEMAKER, RESPIRATORY COMPLICATIONS, POSTOPERATIVE HEMORRHAGE, TRANSFUSION, PERICARDIAL EFFUSION, CARDIAC TAMPONADE, HEART FAILURE.Event Description: THE ARTICLE "COMPARISON OF TRANSCATHETER VERSUS SURGICAL TRICUSPID REPAIR AMONG PATIENTS WITH TRICUSPID REGURGITATION: TWO-YEAR RESULTS" WAS REVIEWED. THE ARTICLE PRESENTED A PROSPECTIVE MULTI CENTER STUDY, TO COMPARE OUTCOMES OF PATIENTS WITH TRICUSPID REGURGITATION (TR) TREATED WITH TRICUSPID TRANSCATHETER EDGE-TO-EDGE REPAIR (T-TEER) VERSUS SURGICAL TRICUSPID VALVE REPAIR. DEVICES MENTIONED INCLUDE TRICLIP, MITRACLIP, AND PASCAL. THE ARTICLE CONCLUDED THAT AMONG MEDICARE BENEFICIARIES, 2-YEAR MORTALITY AND HEART FAILURE HOSPITALIZATION RATES WERE COMPARABLE BETWEEN THE 2 GROUPS. NOTABLY, T-TEER SHOWED ADVANTAGES IN PERIOPERATIVE OUTCOMES, INCLUDING IN-HOSPITAL MORTALITY AND PACEMAKER IMPLANTATION RATES WHILE TRICUSPID VALVE REINTERVENTIONS WERE MORE FREQUENT IN THE T-TEER GROUP. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS TOSHIKI KUNO, AT N, MASSACHUSETTS GENERAL HOSPITAL AND HARVARD MEDICAL SCHOOL, BOSTON, USA WITH CORRESPONDING EMAIL TKUNO@MGH.HARVARD.EDU]. THE TIME FRAME OF THE STUDY WAS JULY 2016 TO DECEMBER 2020. A TOTAL OF 1143 (409 T-TEER) PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH AN UNKNOWN AMOUNT RECEIVED AN ABBOTT DEVICE. FOR T-TEER GROUP, AVERAGE AGE WAS 81 AND MAJORITY GENDER WAS FEMALE. COMORBIDITIES INCLUDED HYPERTENSION, HYPERLIPIDEMIA, DIABETES, MYOCARDIAL INFARCTION, ISCHEMIC HEART DISEASE, CONGESTIVE HEART FAILURE, ATRIAL FIBRILLATION, CHRONIC KIDNEY DISEASE, STROKE/TIA, COPD, ASTHMA, LIVER DISEASE, ANEMIA, PERIPHERAL VASCULAR DISEASE, DEMENTIA, DEPRESSION, CANCER, ARTHRITIS, OBESITY, FUNCTIONAL TRICUSPID REGURGITATION, SECONDARY TRICUSPID REGURGITATION, RHEUMATIC DISEASE, PACEMAKER, PULMONARY HYPERTENSION, PRIOR CARDIAC SURGERY. PERI AND POST-PROCEDURAL COMPLICATIONS INCLUDED HOSPITALIZATION, REINTERVENTION WITH T-TEER, RECURRENT TRICUSPID REGURGITATION, DEATH, ISCHEMIC STROKE, INTRACEREBRAL HEMORRHAGE, ACUTE KIDNEY INJURY, CARDIOGENIC SHOCK, PERMANENT PACEMAKER, RESPIRATORY COMPLICATIONS, POSTOPERATIVE HEMORRHAGE, TRANSFUSION, PERICARDIAL EFFUSION, CARDIAC TAMPONADE, HEART FAILURE. Manufacturer Narrative: IT IS UNKNOWN IF THE DEVICE IS RETURNING FOR ANALYSIS. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION ADDITIONAL PATIENT EFFECTS MENTIONED IN B5 WILL BE FILED UNDER A SEPARATE REPORT NUMBER. B2 DEATH DATE IS ESTIMATED. B3 EVENT DATE IS ESTIMATED. LITERATURE ATTACHMENT: COMPARISON OF TRANSCATHETER VERSUS SURGICAL TRICUSPID REPAIR AMONG PATIENTS WITH TRICUSPID REGURGITATION: TWO-YEAR RESULTS. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF TRICLIP WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING HYPERTENSION, HYPERLIPIDEMIA, DIABETES, MYOCARDIAL INFARCTION, ISCHEMIC HEART DISEASE, CONGESTIVE HEART FAILURE, ATRIAL FIBRILLATION, CHRONIC KIDNEY DISEASE, STROKE/TIA, COPD, ASTHMA, LIVER DISEASE, ANEMIA, PERIPHERAL VASCULAR DISEASE, DEMENTIA, DEPRESSION, CANCER, ARTHRITIS, OBESITY, FUNCTIONAL TRICUSPID REGURGITATION, SECONDARY TRICUSPID REGURGITATION, RHEUMATIC DISEASE, PACEMAKER, PULMONARY HYPERTENSION, PRIOR CARDIAC SURGERY. COMPLICATIONS REPORTED INCLUDED HOSPITALIZATION, REINTERVENTION WITH T-TEER, RECURRENT TRICUSPID REGURGITATION, DEATH, ISCHEMIC STROKE, INTRACEREBRAL HEMORRHAGE, ACUTE KIDNEY INJURY, CARDIOGENIC SHOCK, PERMANENT PACEMAKER, RESPIRATORY COMPLICATIONS, POSTOPERATIVE HEMORRHAGE, TRANSFUSION, PERICARDIAL EFFUSION, CARDIAC TAMPONADE, HEART FAILURE; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21606551&pc=BZD | 2518422-2025-103031 | 2025-03-11 04:00 | Death | RESPIRONICS, INC. | 2025-03-14 04:00 | BZD | RENTAL-DREAMSTATION AUTO CPAP, GB | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A DREAMSTATION AUTO CPAP. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT HAS PASSED AWAY. NO ADDITIONAL DETAILS WERE PROVIDED, AND THERE ARE NO ALLEGATIONS SUGGESTING THAT THE DEVICE CONTRIBUTED TO OR CAUSED THE PATIENT'S DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21606574&pc=NPT | 2025587-2025-01904 | 2025-03-07 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-14 04:00 | NPT | EVOLUT FX VALVE | Patient Device Interaction Problem | Cardiac Arrest; Ventricular Fibrillation | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE IMPLANT OF THE TRANSCATHETER BIOPROSTHETIC VALVE INTO A FRAGILE PATIENT WITH TYPE IV LUNG CANCER, THE RIGHT CAROTID ACCESS WAS USED. ONCE THE GUIDEWIRE ENTERED THE VENTRICLE, THE PATIENT WENT INTO CARDIAC ARREST. AFTER 20 MINUTES OF RESUSCITATION, THE IMPLANT TEAM DECIDED TO IMPLANT THE VALVE TO SEE IF THE PATIENT WOULD RECOVER. THE PATIENT DIED.Event Description: IT WAS REPORTED THAT DURING THE IMPLANT OF THE TRANSCATHETER BIOPROSTHETIC VALVE INTO A FRAGILE PATIENT WITH TYPE IV LUNG CANCER, THE RIGHT CAROTID ACCESS WAS USED. ONCE THE GUIDEWIRE ENTERED THE VENTRICLE, THE PATIENT WENT INTO CARDIAC ARREST. AFTER 20 MINUTES OF RESUSCITATION, THE IMPLANT TEAM DECIDED TO IMPLANT THE VALVE TO SEE IF THE PATIENT WOULD RECOVER. THE PATIENT DIED. ADDITIONAL INFORMATION WAS RECEIVED THAT A MEDTRONIC (CONFIDA) GUIDEWIRE WAS USED DURING THE PROCEDURE. PER THE PHYSICIAN, THE CAUSE OF DEATH WAS VENTRICULAR FIBRILLATION AND CARDIORESPIRATORY ARREST, WHICH THE PATIENT'S UNDERLYING MEDICAL HISTORY CONTRIBUTED TO DUE TO BEING HIGH RISK FROM COMORBIDITY, SEVERE LUNG DISEASE, AND SEVERE MALNUTRITION. AFTER THE VALVE WAS IMPLANTED, NO OBSTRUCTION OF THE CORONARY ARTERIES OR RUPTURE WAS NOTED.Event Description: ADDITIONAL INFORMATION WAS RECEIVED WHICH INDICATED THAT PER THE PHYSICIAN, NEITHER THE GUIDEWIRE, VALVE, NOR THE DELIVERY CATHETER SYSTEM CAUSED OR CONTRIBUTED TO THE ADVERSE EVENTS.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE PATIENT WAS IN EXTREMELY FRAGILE CONDITION WITH STAGE IV PULMONARY ARTERY DISEASE PRIOR TO THE VALVE IMPLANT PROCEDURE. THEY WERE EVALUATED FOR A VALVE DUE TO SEVERE AORTIC STENOSIS. IT WAS NOT POSSIBLE TO USE VIABLE FEMORAL ACCESS, THEREFORE RIGHT CAROTID ACCESS WAS USED AS AN ALTERNATIVE STRATEGY. PRIOR TO THE 23 MILLIMETER (MM) VALVE BEING INSERTED INTO THE PATIENT, AS THE GUIDEWIRE CROSSED INTO THE LEFT VENTRICLE, THE PATIENT SUFFERED CARDIAC ARREST. ADVANCED CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED IMMEDIATELY, WHICH LASTED MORE THAN 20 MINUTES WITHOUT RESTORING SPONTANEOUS CIRCULATION. UP TO THIS POINT, THE VALVE HAD NOT BEEN INSERTED, IMPLANTED OR INTERACTED WITH THE PATIENT IN ANY WAY. IN AN ATTEMPT TO REVERSE THE PATIENT'S CRITICAL CONDITION, THE MEDICAL TEAM DECIDED TO PROCEED WITH THE IMPLANTATION OF THE VALVE AS A LAST RESORT HOPING TO RESTORE CARDIAC FUNCTION. HOWEVER, THE PATIENT NEVER REGAINED EFFECTIVE CIRCULATION AND THEY DIED DURING THE PROCEDURE. THE FATAL OUTCOME WAS NOTED TO BE DUE TO THE PATIENT'S UNDERLYING CRITICAL ILLNESS AND EXTREME INSTABILITY. THEREFORE, THE DEVICE DID NOT CONTRIBUTE TO OR WORSEN THE REPORTED EVENT. DURING THE PROCEDURE AND AFTERWARDS, THE LEAD IMPLANT PHYSICIAN STATED " THE VALVE DID NOT CAUSE ANY HARM TO THE PATIENT. THE DEVICE WAS NOT INVOLVED IN THE CARDIAC ARREST OR THE PATIENT'S HEMODYNAMIC DETERIORATION." Manufacturer Narrative: UPDATED DATA: B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID D-EVOLUTFX-2329 (LOT: 0012526061); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE (B)(6) 2025-; EXPLANT DATE (B)(6) 2025, PRODUCT ID L-EVOLUTFX-2329 (LOT: 0012526084); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: (B)(6) 2025; EXPLANT DATE: (B)(6) 2025. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME: D-EVOLUTFX DELIVERY CATHETER SYSTEM (DCS); PRODUCT ID: D-EVOLUTFX-2329; LOT: 0012526061; USE BY DATE: 2026-11-07; UDI: (B)(4). OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: PRODUCT ID: L-EVOLUTFX-2329 (LOT: 0012526084); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: N/A; EXPLANT DATE: N/A. PRODUCT ID: GWBC30 (LOT: 92206879); PRODUCT TYPE: 0193-GUIDEWIRE; IMPLANT DATE: N/A; EXPLANT DATE: N/A. UPDATED: B5, H6, H11. CORRECTED: H11. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21607080&pc=NKM | 2135147-2025-01300 | 2016-01-01 05:00 | Death | ABBOTT MEDICAL | 2025-03-14 04:00 | NKM | MITRACLIP® SYSTEM CLIP DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P100009 |  | 1 | Event Description: THE ARTICLE "TRANSCATHETER EDGE-TO-EDGE REPAIR VERSUS ANNULOPLASTY IN FUNCTIONAL MITRAL VALVE REGURGITATION: A COMPARISON OF CARDIOVASCULAR OUTCOMES" WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE MULTI CENTER STUDY, TO EVALUATE IN-PATIENT MORTALITY AND CARDIOVASCULAR COMPLICATIONS IN PATIENTS WITH HEART FAILURE WITH REDUCED EJECTION FRACTION (HFREF) WHO DEVELOPED SEVERE FMVR REQUIRING TREATMENT WITH MITRACLIP G4 VERSUS ANNULOPLASTY. COMPARISONS OF OUTCOMES TO PREVIOUS ITERATIONS OF THE MITRACLIP WERE INCLUDED IN THE ANALYSIS. DEVICES MENTIONED INCLUDE MITRACLIP. THE ARTICLE CONCLUDED THAT THE MITRACLIP GROUP WAS ASSOCIATED WITH A DECREASED RISK OF IN-HOSPITAL MORTALITY (ODDS RATIO (OR): 0.38, CONFIDENCE INTERVAL (CI): 0.18 - 0.77), ISCHEMIC STROKE (OR: 0.29, CI: 0.13 - 0.61), AND MYOCARDIAL INFARCTION (OR: 0.15, CI: 0.08 - 0.28). THE MITRACLIP G4 COHORT DID NOT OUTPERFORM EARLIER CLIP VERSIONS IN REDUCING COMPLICATIONS. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS ANDREW SAGALOV AT SIU SCHOOL OF MEDICINE INSTITUTION WITH CORRESPONDING EMAIL ASAGALOV56@SIUMED.EDU]. THE TIME FRAME OF THE STUDY WAS JANUARY 2016 TO DECEMBER 2020. A TOTAL OF 19,475 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH 19,475 RECEIVED AN ABBOTT DEVICE. AS THIS EVENT IS FROM A LITERATURE REVIEW, THERE IS NO RELEVANT PATIENT INFORMATION (DATE OF BIRTH, AGE, GENDER, WEIGHT, AND MEDICAL HISTORY) TO REPORT. COMORBIDITIES INCLUDED CEREBROVASCULAR DISEASE, CORONARY AND PERIPHERAL ARTERY DISEASE, ATRIAL FIBRILLATION, HYPERTENSION, DIABETES. (B)(6), UNK MITRACLIP: PERI- AND POST-PROCEDURAL COMPLICATIONS INCLUDED SLDA, EMBOLIZATION, ISCHEMIC STROKE, MYOCARDIAL INFARCTION, CARDIAC TAMPONADE, PULMONARY EMBOLISM, BLEEDING, PERICARDIAL EFFUSION, TISSUE DAMAGE, DEATH. Manufacturer Narrative: IT IS UNKNOWN IF THE DEVICE IS RETURNING FOR ANALYSIS. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF MITRACLIP DEVICE WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING CEREBROVASCULAR DISEASE, CORONARY AND PERIPHERAL ARTERY DISEASE, ATRIAL FIBRILLATION, HYPERTENSION, DIABETES. COMPLICATIONS REPORTED INCLUDED SLDA, EMBOLIZATION, ISCHEMIC STROKE, MYOCARDIAL INFARCTION, CARDIAC TAMPONADE, PULMONARY EMBOLISM, BLEEDING, PERICARDIAL EFFUSION, TISSUE DAMAGE, DEATH; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21607093&pc=OZD | 1220648-2025-26730 | 2025-02-16 05:00 | Death | ABIOMED, INC. | 2025-03-14 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT EXPERIENCED LIMB ISCHEMIA REQUIRING INTERVENTION. THE PATIENT WOULD LATER EXPIRE. THE PATIENT PRESENTED WITH ST-SEGMENT ELEVATION MYOCARDIAL INFARCTION. PERCUTANEOUS CORONARY INTERVENTION WAS COMPLETED TO THE LEFT ANTERIOR DESCENDING ARTERY WITH ONE STENT. LEFT VENTRICULAR DIASTOLIC PRESSURE WAS 30 AND LEFT VENTRICULAR EJECTION FRACTION WAS 30%. POST PCI THE IMPELLA CP DEVICE WAS IMPLANTED, AND THE PATIENT WAS SENT TO THE UNIT ON SUPPORT. APPROXIMATELY TWELVE HOURS LATER INTO THE NEXT DAY THE PATIENT DEVELOPED LIMB ISCHEMIA. CREATINE LEVEL WAS RISING, AND LACTATE LEVEL WAS INCREASING. THERE WERE NO PULSES DISTAL TO THE IMPELLA CP. THE DECISION WAS MADE TO EXPLANT AND REPLACE WITH ESCALATION TO AN IMPELLA 5.5. ADDITIONAL INFORMATION RECEIVED INDICATING THAT SUTURES WERE USED WHEN THEY DID A FEMORAL VASCULAR CUT DOWN TO REMOVE THE IMPELLA CP AND THEN DID A FASCIOTOMY TO THE LEG. THE PATIENT WAS PLACED ON THE IMPELLA 5.5. THIS SAME THE PATIENT WAS PREEMPTIVELY STARTED ON CONTINUOUS RENAL REPLACEMENT THERAPY AND THE PATIENT HAD WASHOUT LACTIC FROM THE ISCHEMIC LEG. THE PATIENT REMAINED CRITICALLY ILL AND WAS PLANNED FOR DIALYSIS FOR LACTIC ACIDOSIS. BLEEDING ISSUES FROM ALL SITES (IMPELLA SITE, FEMORAL SITE, FEMORAL FASCIOTOMY, AND ORALLY) WAS OBSERVED. IMPELLA 5.5 FLOWS AND WAVEFORMS WERE NOTED WITHIN NORMAL LIMITS. HEPARIN WAS HELD POST PROCEDURE AND BLOOD PRODUCTS WERE ADMINISTERED. LATER IN THE NIGHT IT WAS NOTED THAT THE PATIENT WAS UNABLE TO TOLERATE PROLONGED INTERMITTENT RENAL REPLACEMENT THERAPY. URINE OUTPUT WAS POOR, AND THE NURSE REPORTED UNCHANGED BLEEDING FROM LINE SITES AND INCISIONS. THE PATIENT REMAINED CRITICALLY ILL AND WOULD EXPIRE THE FOLLOWING DAY. A TRANSESOPHAGEAL ECHOCARDIOGRAM REVEALED A LARGE VENTRICULAR SEPTAL DEFECT AND CARE WAS WITHDRAWN. MEDICAL SAFETY REVIEW OF THE EVENT NOTED THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA CP AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME (DEMISE). Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THIS IS ONE OF THREE DEVICE MANUFACTURING REPORTS ASSOCIATED WITH THE EVENT AND/OR IMPLANT EXPERIENCE. REFER TO THE FOLLOWING: MEDICAL DEVICE REPORT FOR THE IMPELLA CP SERIAL NUMBER (B)(6) WITH DATE OF THE EVENT 16-FEB-2025 AND PATIENT IDENTIFIER ENDING IN (B)(6). (THIS REPORT). MEDICAL DEVICE REPORT FOR THE IMPELLA 5.5 SERIAL NUMBER (B)(6) WITH DATE OF THE EVENT 16-FEB-2025 AND PATIENT IDENTIFIER ENDING IN (B)(6). MEDICAL DEVICE REPORT FOR THE AUTOMATED IMPELLA CONTROLLER SERIAL NUMBER (B)(6) WITH DATE OF THE EVENT 16-FEB-2025 AND PATIENT IDENTIFIER (B)(6). INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: IMPELLA CP WITH SMARTASSIST SYSTEM. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): "ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY" (INCLUDING VENTRICULAR PERFORATION).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21607164&pc=NPT | 2015691-2025-01936 | 2025-02-19 05:00 | Death | EDWARDS LIFESCIENCES | 2025-03-14 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Thrombosis/Thrombus | P140031 |  | 1 | Event Description: AS REPORTED BY OUR EDWARDS LIFESCIENCES AFFILIATE IN SPAIN REGARDING A 26 MM SAPIEN 3 ULTRA TRANSCATHETER HEART AORTIC VALVE, AFTER 3 YEARS AND 3 MONTHS OF IMPLANT PROCEDURE, PATIENT UNDERWENT A CARDIAC SURGERY FOR A MITRAL-AORTIC REPLACEMENT DUE TO AORTIC VALVE THROMBOSIS AND NATIVE MITRAL ISSUE NOT RELATED WITH SAPIEN 3 VALVE. THE PATIENT PASSED AWAY DUE TO COMPLICATIONS POST-SURGERY. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), VALVE THROMBOSIS IS A POTENTIAL RISK ASSOCIATED WITH THE USE OF THE TRANSCATHETER HEART VALVES (THV). USE OF THE SAPIEN TRANSCATHETER HEART VALVES (ALL MODELS) IS CONTRAINDICATED IN PATIENTS WHO CANNOT TOLERATE AN ANTICOAGULATION/ANTIPLATELET REGIMEN. IN ADDITION, IT WARNS THAT VALVE RECIPIENTS SHOULD BE MAINTAINED ON ANTICOAGULANT/ANTIPLATELET THERAPY, EXCEPT WHEN CONTRAINDICATED, AS DETERMINED BY THEIR PHYSICIAN. THROMBOSIS IS A RARE AND WELL-RECOGNIZED COMPLICATION ASSOCIATED WITH THE USE OF BIOPROSTHETIC HEART VALVES. VALVE THROMBOSIS IS THE FORMATION OF SIGNIFICANT BLOOD CLOTS FORMING ON THE VALVE. POTENTIAL PATIENT RISK FACTORS SUCH AS ATRIAL FIBRILLATION, SYSTEMIC DISEASE (E.G., SYSTEMIC LUPUS ERYTHEMATOSUS, INFLAMMATION, AND DAMAGE TO VARIOUS BODY TISSUES, INCLUDING JOINTS, SKIN, KIDNEYS, HEART, LUNGS, BLOOD VESSELS, AND BRAIN), THE VALVULAR DISEASE ITSELF, AND REDUCED CARDIAC EJECTION FRACTION CAN CONTRIBUTE TO INCREASED RISK OF THROMBUS/THROMBOSIS. SUB-OPTIMAL ANTICOAGULATION DURING THE PROCEDURE AND UNDERLYING PATIENT CONDITIONS CAN ALSO RESULT IN INCREASED THROMBOGENICITY. THESE PATIENTS ARE ANTICOAGULATED FOR THE PROCEDURE AND INTERVENTIONAL BEST PRACTICES MANDATE METICULOUS PREPARATION, ASPIRATION, AND FLUSHING OF THE DEVICES TO PREVENT AND/OR REMOVE CLOT(S). SHORT-TERM ANTICOAGULATION THERAPY MAY ALSO BE NECESSARY AFTER VALVE REPAIR, ANTICOAGULATION IS PRESCRIBED PER INSTITUTIONAL GUIDELINES. INTRA-PROCEDURAL INTRA-CARDIAC THROMBUS AND POST-PROCEDURE OR LATE VALVE THROMBOSIS ARE COMPLEX PROCESSES TRIGGERED BY THE INTERACTION BETWEEN THE HOST AND THE DEVICE WHICH ARE HIGHLY VARIABLE AMONG PATIENTS. IT IS THE NATURAL TENDENCY OF THE BODY TO FORM A CLOT ON FOREIGN OBJECTS IN THE VASCULAR SPACE AND A DEFINITIVE ROOT CAUSE CANNOT ALWAYS BE CONFIRMED. THE DEVICE TRAINING MANUALS INSTRUCT THE OPERATOR TO CONSIDER ALL PROCEDURAL AND ANATOMICAL FACTORS. EDWARDS EXTENSIVELY TRAINS PHYSICIANS BEFORE THEY ARE QUALIFIED TO USE THE DEVICE SYSTEM. TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES. OPERATORS ARE ALSO INSTRUCTED TO USE FLUOROSCOPY IN CONJUNCTION WITH ECHOCARDIOGRAPHY FOR OPTIMAL VISUALIZATION DURING POSITIONING AND DEPLOYMENT AND CAUTIONS INCLUDE THE MAINTENANCE OF THE PATIENT'S ANTICOAGULATION STATUS (ACT AT >250 SECONDS) DURING THE PROCEDURE IN ORDER TO PREVENT THROMBOSIS. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST/INDICATE A DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED DUE TO LIMITED INFORMATION PROVIDED. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21607404&pc=OZD | 1220648-2025-26731 | 2025-02-16 05:00 | Death | ABIOMED, INC. | 2025-03-14 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS INITIALLY IMPLANTED WITH AN IMPELLA CP DEVICE AND WAS ESCALATED TO AN IMPELLA 5.5 DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. THE PATIENT DEVELOPED LIMB ISCHEMIA WHILE ON IMPELLA CP MCS (ON DAY AFTER START OF THIS SUPPORT). THE DECISION WAS MADE TO ESCALATE TO AN IMPELLA 5.5 WHICH WAS SUCCESSFULLY COMPLETED. ON THE SAME DAY OF THE IMPELLA 5.5 IMPLANT, THE PATIENT WAS PREEMPTIVELY STARTED ON CONTINUOUS RENAL REPLACEMENT THERAPY AND THE PATIENT HAD WASHOUT LACTIC FROM THE ISCHEMIC LEG. THE PATIENT REMAINED CRITICALLY ILL AND WAS PLANNED FOR DIALYSIS FOR LACTIC ACIDOSIS. BLEEDING ISSUES FROM ALL SITES (IMPELLA SITE, FEMORAL SITE, FEMORAL FASCIOTOMY, AND ORALLY) WERE OBSERVED. IMPELLA FLOWS AND WAVEFORMS WERE NOTED WITHIN NORMAL LIMITS. HEPARIN WAS HELD POST PROCEDURE AND BLOOD PRODUCTS WERE ADMINISTERED. LATER IN THE NIGHT IT WAS NOTED THAT THE PATIENT WAS UNABLE TO TOLERATE PROLONGED INTERMITTENT RENAL REPLACEMENT THERAPY. URINE OUTPUT WAS POOR, AND THE NURSE REPORTED UNCHANGED BLEEDING FROM LINE SITES AND INCISIONS. THE PATIENT REMAINED CRITICALLY ILL AND WOULD EXPIRE THE NEXT DAY. A TRANSESOPHAGEAL ECHOCARDIOGRAM REVEALED A LARGE VENTRICULAR SEPTAL DEFECT AND CARE WAS WITHDRAWN. MEDICAL SAFETY REVIEW OF THE EVENT NOTED THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA CP AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME (DEMISE). RELATED REPORTS ARE NOTED IN H.10 ACCORDINGLY. Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING AND VASCULAR DAMAGE - PERIPHERAL VESSEL HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. B2 SELECTION OF LIFE-THREATENING WAS ERRONEOUSLY SELECTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26731. H1 TYPE OF REPORTABLE EVENT WAS ERRONEOUSLY SELECTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26731. H6 HEALTH EFFECT - IMPACT CODE 1802 DEATH WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26731. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THIS IS ONE OF THREE DEVICE MANUFACTURING REPORTS ASSOCIATED WITH THE EVENT AND/OR IMPLANT EXPERIENCE. REFER TO THE FOLLOWING: MEDICAL DEVICE REPORT FOR THE IMPELLA CP SERIAL NUMBER (B)(6) WITH DATE OF THE EVENT 16-FEB-2025 AND PATIENT IDENTIFIER ENDING IN (B)(6). MEDICAL DEVICE REPORT FOR THE IMPELLA 5.5 SERIAL NUMBER (B)(6) WITH DATE OF THE EVENT 16-FEB-2025 AND PATIENT IDENTIFIER ENDING IN (B)(6). (THIS REPORT). MEDICAL DEVICE REPORT FOR THE AUTOMATED IMPELLA CONTROLLER SERIAL NUMBER (B)(6) WITH DATE OF THE EVENT 16-FEB-2025 AND PATIENT IDENTIFIER (B)(6).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21607684&pc=OUT | 2029214-2025-00671 | 2022-11-03 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-03-14 04:00 | OUT | PIPELINE FLEX W/SHIELD TECHNOLOGY | Adverse Event Without Identified Device or Use Problem | Ruptured Aneurysm | P100018 |  | 1 | Event Description: MADJIDYAR, J., KELLER, E., WINKLHOFER, S., TOTH, D., BARNAURE, I., SCHUBERT, T., THURNER, P., FIERSTRA, J., WILLMS, J. F., REGLI, L., KULCSAR, Z. (2023). SINGLE-ANTIPLATELET REGIMEN IN RUPTURED CEREBRAL BLOOD BLISTER AND DISSECTING ANEURYSMS TREATED WITH F LOW-DIVERTER STENT RECONSTRUCTION. JOURNAL OF NEUROINTERVENTIONAL SURGERY, 15(10), 953¿957. HTTPS://DOI.ORG/10.1136/JNIS-2022-019361. MEDTRONIC REVIEW OF THE LITERATURE ARTICLE FOUND A SINGLE-CENTER CASE SERIES OF NINE PATIENTS TREATED FOR RUPTURED BLOOD BLISTER ANEURYSMS (BBA) OR DISSECTING ANEURYSMS WITH ASSOCIATED SUBARACHNOID HEMORRHAGE (SAH) USING PIPELINE FLEX WITH SHIELD TECHNOLOGY FLOW DIVERSION STENTS FROM 2019 TO 2022. THE CASES WERE OFF-LABEL USE FOR THE PIPELINE EMBOLIZATION DEVICES DUE TO USE OF SINGLE ANTIPLATELET TREATMENT AS DUAL ANTIPLATELET TREATMENT IS RECOMMENDED PER THE INSTRUCTIONS FOR USE (IFU). FURTHERMORE, PER THE IFU, THE PIPELINE¿ FLEX EMBOLIZATION DEVICE WITH SHIELD TECHNOLOGY¿ SHOULD NOT BE USED ALONE AS SOLE THERAPY FOR ACUTELY RUPTURED ANEURYSMS. THE PURPOSE OF THE ARTICLE WAS TO INVESTIGATE THE EFFICACY OF A SINGLE ANTIPLATELET REGIMEN TO PREVENT ANTIPLATELET TREATMENT RELATED HEMORRHAGIC COMPLICATIONS. ALL PIPELINES WERE IMPLANTED WITHOUT PROCEDURAL ADVERSE EVENTS AND NO DEVICE MALFUNCTION WAS REPORTED IN THE ARTICLE. TWO PATIENTS ALSO HAD COIL EMBOLIZATION IN ADDITIONAL TO THE STENT IMPLANTATION. ONE PATIENT HAD TWO PIPELINES IMPLANTED TELESCOPING IMPLANTED IN THE INITIAL PROCEDURE FOR FULLY AFFECTIVE RECONSTRUCTION OF THE CAROTID SEGMENT. IT SHOULD BE NOTED THAT THE PHENOM 27 MICROCATHETER WAS USED IN THE FLOW DIVERSION STENTING PROCEDURES. ONE PATIENT DEATH WAS REPORTED: - PATIENT 7 WAS TREATED FOR A FUSIFORM DISSECTION ANEURYSM OF THE RIGHT VERTEBRAL ARTERY AND SUFFERED A LETHAL RE-RUPTURE SIX DAYS AFTER THE INDEX PROCEDURE. Manufacturer Narrative: SEPARATE REPORTS WILL BE SUBMITTED FOR THE PHENOM 27 CATHETERS AND SERIOUS INJURIES REPORTED IN THE ARTICLE THAT DID NOT HAVE AN OUTCOME OF MORTALITY. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21607690&pc=LPB | 2029046-2025-00790 | 2025-02-14 05:00 | Death | BIOSENSE WEBSTER INC | 2025-03-14 04:00 | LPB | THERMOCOOL SMARTTOUCH SF | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Cardiac Perforation | P030031 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A PAROXYSMAL ATRIAL FIBRILLATION ABLATION WITH A THERMOCOOL SMARTTOUCH SF AND THE PATIENT EXPERIENCED A PERICARDIAL EFFUSION AND CARDIAC ARREST THAT REQUIRED PERICARDIOCENTESIS AND CPR (CARDIOPULMONARY RESUSCITATION). THE PATIENT DIED. IT WAS REPORTED THAT THE PATIENT SUFFERED A PERICARDIAL EFFUSION DURING AN AFIB ABLATION. WHILE THE PHYSICIAN WAS MAPPING AND ABLATING IN THE LEFT ATRIUM WITH THE THERMOCOOL SMARTTOUCH SF CATHETER THEY BELIEVED A PERFORATION OCCURRED NEAR THE ANTERIOR ROOF AREA. THIS WAS DISCOVERED WHEN THE PATIENT'S BLOOD PRESSURE DROPPED AND EFFUSION CONFIRMED USING INTRACARDIAC ECHO. THE PATIENT LOST DISCERNABLE HEART RHYTHM. THE PATIENT CODED, CPR WAS ADMINISTERED. A PERICARDIOCENTESIS WAS PERFORMED AND VENOUS SHEATHS WERE SECURED IN PLACE. THE PATIENT STABILIZED TEMPORARILY BUT CODED MULTIPLE TIMES BEFORE LEAVING THE EP LAB. PATIENT STATUS WAS UNKNOWN AT THE TIME OF THE CALL. ADDITIONAL INFORMATION WAS RECEIVED ON 24-FEB-2025. PATIENT REQUIRED CARDIAC SURGERY. THE PATIENT REQUIRED EXTENDED HOSPITALIZATION; HOWEVER, THE OUTCOME OF THE ADVERSE EVENT WAS DEATH. THE PHYSICIAN'S OPINION ON THE CAUSE OF THE ADVERSE EVENT IS THE PROCEDURE. IT WAS REPORTED THAT THE PATIENT HAD LOW HEMOGLOBIN, NEEDED BREATHING TREATMENT AT BEGINNING OF CASE. THERE WAS NO EVIDENCE OF STEAM POP. NO ISSUES WITH EQUIPMENT. ONE TRANSSEPTAL PUNCTURE SITE WAS PERFORMED DURING THE PROCEDURE. ADDITIONAL INFORMATION WAS RECEIVED 12-MAR-2025. PATIENT WAS IN THE EP LAB FOR AN ATRIAL FIBRILLATION ABLATION, BUT HAD NO SURGERY POST PROCEDURE. Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED ON 14-MAR-2025. PATIENT DID NOT HAVE SURGERY. DOES NOT KNOW THE DATE THE PATIENT PASSED AWAY. THE DEVICE EVALUATION WAS COMPLETED ON 20-MAR-2025. IT WAS REPORTED THAT A PATIENT UNDERWENT A PAROXYSMAL ATRIAL FIBRILLATION ABLATION WITH A THERMOCOOL SMARTTOUCH SF AND THE PATIENT EXPERIENCED A PERICARDIAL EFFUSION AND CARDIAC ARREST THAT REQUIRED PERICARDIOCENTESIS AND CARDIOPULMONARY RESUSCITATION. THE PATIENT REQUIRED EXTENDED HOSPITALIZATION; HOWEVER, THE OUTCOME OF THE ADVERSE EVENT WAS DEATH. THE PRODUCT WAS RETURNED TO JOHNSON & JOHNSON MEDTECH (J&J MEDTECH) FOR EVALUATION. A VISUAL INSPECTION AND REVISION OF ALL FEATURES WERE PERFORMED FOLLOWING J&J MEDTECH PROCEDURES. VISUAL ANALYSIS REVEALED NO DAMAGE OR ANOMALIES. THE DEVICE FEATURES WERE REVIEWED, AND NO FORCE, MAGNETIC, OR DEFLECTION ISSUES WERE OBSERVED; HOWEVER, A HIGH TEMPERATURE ALERT WAS DISPLAYED DUE TO NO IRRIGATION RELATED TO THE IRRIGATION TUBE BEING DAMAGED AND BLOCKED WITH REDDISH MATERIAL IN THE LUER HUB AREA. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH NUMBER, AND NO INTERNAL ACTIONS WERE IDENTIFIED. THE HIGH TEMPERATURE AND IRRIGATION ISSUES ARE UNRELATED TO THE REPORTED EVENT; THEREFORE, ADVERSE EVENT ISSUE REPORTED COULD NOT BE CONFIRMED AS PRODUCT RELATED, THERE MAY HAVE BEEN OTHER CIRCUMSTANCES OR ISSUES THAT OCCURRED DURING THE USE OF THE DEVICE THAT COULD NOT BE REPLICATED DURING THE ANALYSIS. THE POTENTIAL CAUSE OF THE ADVERSE EVENT REMAINS UNKNOWN. THE POTENTIAL CAUSE OF THE IRRIGATION AND CONSEQUENTIAL HIGH TEMPERATURE ALERT COULD BE RELATED TO THE PROCEDURE; HOWEVER, THIS COULD NOT BE ESTABLISHED CONCLUSIVELY. THE INSTRUCTION FOR USE (IFU) STATES THAT CAREFUL CATHETER MANIPULATION MUST BE PERFORMED TO AVOID CARDIAC DAMAGE, PERFORATION, OR TAMPONADE. PURGE THE CATHETER AND THE IRRIGATION TUBING WITH HEPARINIZED NORMAL SALINE PRIOR TO INSERTION OF THE CATHETER INTO THE PATIENT. ALWAYS MAINTAIN A CONSTANT HEPARINIZED NORMAL SALINE INFUSION TO PREVENT COAGULATION WITHIN THE LUMEN OF THE CATHETER. DO NOT USE THE CATHETER WITHOUT IRRIGATION FLOW. AS PART OF JOHNSON & JOHNSON MEDTECH'S QUALITY PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND RELEASED TO APPROVED SPECIFICATIONS. EXPLANATION OF CODES: INVESTIGATION FINDINGS: OPERATIONAL PROBLEM IDENTIFIED (C13)/ INVESTIGATION CONCLUSIONS: CAUSE NOT ESTABLISHED (D15) / COMPONENT CODE: HOSE (G04069) WERE SELECTED AS RELATED TO THE BIOSENSE WEBSTER INC. ANALYSIS FINDING OF THE ¿IRRIGATION TUBE BEING DAMAGED¿. INVESTIGATION FINDINGS: NO DEVICE PROBLEM FOUND (C19) / INVESTIGATION CONCLUSIONS: NO PROBLEM DETECTED (D14) WERE SELECTED AS RELATED TO THE CUSTOMER¿S REPORTED ¿ADVERSE EVENT¿ ISSUE. OCC - IRRIG TUBING BLOCK AT LUER HUB-INVESTIGATION FINDINGS: OPERATIONAL PROBLEM IDENTIFIED (C13)/ INVESTIGATION CONCLUSIONS: CAUSE NOT ESTABLISHED (D15) / COMPONENT CODE: HOSE (G04069) WERE SELECTED AS RELATED TO THE BIOSENSE WEBSTER INC. ANALYSIS FINDING OF ¿BLOCKED WITH REDDISH MATERIAL IN THE LUER HUB AREA¿. IF ADDITIONAL INFORMATION IS RECEIVED REGARDING THIS EVENT, A SUPPLEMENTAL 3500A REPORT WILL BE SUBMITTED TO THE FDA. MANUFACTURER'S REFERENCE NUMBER: (B)(4). Manufacturer Narrative: PER THE ADDITIONAL INFORMATION RECEIVED ON 24-FEB-2025 STATING THAT THE OUTCOME OF THE ADVERSE EVENT WAS DEATH; HOWEVER, THE DATE OF DEATH WAS NOT PROVIDED, PROCESSED THE "DATE OF DEATH" TO (B)(6) 2025. ADDITIONAL CLARIFICATION TO THIS DATE WAS REQUESTED. HOWEVER, NO FURTHER INFORMATION HAS BEEN MADE AVAILABLE. THE BWI PRODUCT ANALYSIS LAB RECEIVED THE DEVICE FOR EVALUATION ON 05-MAR-2025. THE ANALYSIS HAS BEGUN BUT IS NOT COMPLETED AT THIS TIME. WHEN THE INVESTIGATIONAL ANALYSIS HAS BEEN COMPLETED, A SUPPLEMENTAL 3500A REPORT WILL BE SUBMITTED. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. MANUFACTURER'S REFERENCE NUMBER: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21607699&pc=DSQ | 2916596-2025-01563 | 2025-02-26 05:00 | Death | THORATEC CORPORATION | 2025-03-14 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO AN INTRACRANIAL HEMORRHAGE (ICH). THIS WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED AND WOULD NOT RETURN. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: THE PUMP REMAINS IN USE SUPPORTING THE PATIENT. A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF STROKE COULD NOT BE CONCLUSIVELY DETERMINED. STROKE IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. STROKE HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: CORRECTION TO SECTION H11: THE PATIENT PASSED AWAY AND THE PUMP WAS NO LONGER IN USE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21607703&pc=OZD | 1220648-2025-26732 | 2025-02-17 05:00 | Death | ABIOMED, INC. | 2025-03-14 04:00 | OZD | AUTOMATED IMPELLA CONTROLLER | Device Sensing Problem | Cardiac Arrest | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED SHORTLY AFTER START OF THE IMPELLA 5.5 SUPPORT TO A PATIENT WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK, THE PLACEMENT AND LEFT VENTRICLE SIGNALS BOTH WOULD INTERMITTENTLY GO OUT IN THE ABSENCE OF ANY ALARMS SUCH AS PLACEMENT UNRELIABLE. THE DECISION WAS MADE TO SWITCH OUT THE AUTOMATED IMPELLA CONTROLLER (AIC) MID-SUPPORT WITH REPORT OF NO HARM TO THE PATIENT. THE PATIENT REMAINED ON THE IMPELLA MECHANICAL CIRCULATORY SUPPORT FOR NINE DAYS BEFORE SUPPORT WAS REMOVED. THE PATIENT WAS NOT A CANDIDATE FOR ANY OTHER SUPPORT, EXPRESSED TO DISCONTINUE THE SUPPORT, AND SUBSEQUENTLY EXPIRED. THE DIRECT CAUSE OF DEATH IS UNKNOWN. ABIOMED'S POSITION IS THE AIC IS NOT A MAIN FACTOR IN THE DEMISE. HOWEVER, THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE DEVICE AS AN ASSOCIATED FACTOR TO THE EVENT PER MEDICAL SAFETY REVIEW. Manufacturer Narrative: THE INVESTIGATION OF OPTICAL SIGNAL ISSUE HAS BEEN COMPLETED. DATA ANALYSIS: THIS FAILURE MODE REPRODUCED ON THE COMPLAINT DATE WITH THE OPTICAL BENCH CONSTANTLY RESETTING AS OPTICAL COMMUNICATION WAS MALFUNCTIONING. THIS RESULTED IN SOMETIMES A ONE SECOND DROP TO -1638.4 MMHG IMMEDIATELY FOLLOWED BY RISE TO +1638.4MMHG FOR TWO SECONDS BEFORE RETURNING TO PREVIOUS RANGE OF SIGNAL. ADDITIONALLY, THE PLACEMENT SIGNAL SPIKED TO 3000 MMHG DURING THIS TIME WHICH IS WHAT THE USER WOULD HAVE SEEN. PUMP 570404 WAS MOVED OVER TO IC1755 AND THE FAILURE MODE DID NOT REPRODUCE ON THAT CONSOLE. DEVICE ANALYSIS: THE FAILURE MODE WAS CONSISTENTLY REPRODUCED DURING TESTING WHENEVER A PUMP WAS CONNECTED TO THE CONSOLE. THE OPTICAL BENCH WAS REPLACED WITH A KNOWN GOOD BENCH AND THE MALFUNCTION DID NOT REPRODUCE. ROOT CAUSE: THE LOSS OF PLACEMENT SIGNAL WAS CAUSED BY A CONSISTENTLY RESETTING OPTICAL BENCH THAT WAS MALFUNCTIONING. B1 REVISED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26732 TO REFLECT BOTH ADVERSE EVENT AND PRODUCT PROBLEM. Manufacturer Narrative: THE AUTOMATED IMPELLA CONTROLLER DEVICE WAS RETURNED, AND AN EVALUATION IS UNDERWAY. UPON COMPLETION OF OUR ANALYSIS, A SUPPLEMENTAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21607716&pc=KRA | 2029214-2025-00672 | 2022-11-03 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-03-14 04:00 | KRA | CATHERA | Adverse Event Without Identified Device or Use Problem | Ruptured Aneurysm | K151638 |  | 1 | Event Description: MADJIDYAR, J., KELLER, E., WINKLHOFER, S., TOTH, D., BARNAURE, I., SCHUBERT, T., THURNER, P., FIERSTRA, J., WILLMS, J. F., REGLI, L., & KULCSAR, Z. (2023). SINGLE-ANTIPLATELET REGIMEN IN RUPTURED CEREBRAL BLOOD BLISTER AND DISSECTING ANEURYSMS TREATED WITH F LOW-DIVERTER STENT RECONSTRUCTION. JOURNAL OF NEUROINTERVENTIONAL SURGERY, 15(10), 953¿957. HTTPS://DOI.ORG/10.1136/JNIS-2022-019361. MEDTRONIC REVIEW OF THE LITERATURE ARTICLE FOUND A SINGLE-CENTER CASE SERIES OF NINE PATIENTS TREATED FOR RUPTURED BLOOD BLISTER ANEURYSMS (BBA) OR DISSECTING ANEURYSMS WITH ASSOCIATED SUBARACHNOID HEMORRHAGE (SAH) USING PIPELINE FLEX WITH SHIELD TECHNOLOGY FLOW DIVERSION STENTS FROM 2019 TO 2022. THE CASES WERE OFF-LABEL USE FOR THE PIPELINE EMBOLIZATION DEVICES DUE TO USE OF SINGLE ANTIPLATELET TREATMENT AS DUAL ANTIPLATELET TREATMENT IS RECOMMENDED PER THE INSTRUCTIONS FOR USE (IFU). FURTHERMORE, PER THE IFU, THE PIPELINE¿ FLEX EMBOLIZATION DEVICE WITH SHIELD TECHNOLOGY¿ SHOULD NOT BE USED ALONE AS SOLE THERAPY FOR ACUTELY RUPTURED ANEURYSMS. THE PURPOSE OF THE ARTICLE WAS TO INVESTIGATE THE EFFICACY OF A SINGLE ANTIPLATELET REGIMEN TO PREVENT ANTIPLATELET TREATMENT RELATED HEMORRHAGIC COMPLICATIONS. ALL PIPELINES WERE IMPLANTED WITHOUT PROCEDURAL ADVERSE EVENTS AND NO DEVICE MALFUNCTION WAS REPORTED IN THE ARTICLE. TWO PATIENTS ALSO HAD COIL EMBOLIZATION IN ADDITIONAL TO THE STENT IMPLANTATION. ONE PATIENT HAD TWO PIPELINES IMPLANTED TELESCOPING IMPLANTED IN THE INITIAL PROCEDURE FOR FULLY AFFECTIVE RECONSTRUCTION OF THE CAROTID SEGMENT. IT SHOULD BE NOTED THAT THE PHENOM 27 MICROCATHETER WAS USED IN THE FLOW DIVERSION STENTING PROCEDURES. ONE PATIENT DEATH WAS REPORTED: PATIENT 7 WAS TREATED FOR A FUSIFORM DISSECTION ANEURYSM OF THE RIGHT VERTEBRAL ARTERY AND SUFFERED A LETHAL RE-RUPTURE SIX DAYS AFTER THE INDEX PROCEDURE. THE FOLLOWING NON-LETHAL ADVERSE EVENTS WERE REPORTED: IN PATIENT 3, A SECOND PIPELINE WAS IMPLANTED 14 DAYS AFTER THE INITIAL TREATMENT PROCEDURE DUE TO INCREASING SIZE OF THE ANEURYSM. THREE PATIENTS HAD NEW NEUROLOGICAL DISABILITY REPORTED DURING FOLLOW-UP. TWO PATIENTS HAD A NEW MODERATE DISABILITY (MRS 3 AT 6¿10 MONTHS) AND ONE PATIENT HAD A NEW MODERATELY SEVERE DISABILITY (MRS 4 AT 3 MONTHS). THE CAUSE OF THE NEW NEUROLOGICAL DEFICIT WAS NOT SPECIFIED IN THE ARTICLE. TWO PATIENT WERE REPORTED TO HAVE SUBTOTAL ANEURYSM OCCLUSION AT 3 AND 6 MONTHS WITH NO LATER FOLLOW UP AVAILABLE AND NO ADDITIONAL TREATMENT, INTERVENTION, OR ASSOCIATED SYMPTOMS REPORTED IN THE ARTICLE. IT SHOULD BE NOTED THAT THERE WERE ADVERSE EVENTS REPORTED IN THE ARTICLE WHICH WERE NON-COMPLAINT EVENTS PER 027-WI185, DUE TO CAUSE BEING ATTRIBUTED TO THE PATIENTS' PRESENTING CONDITION OF SAH, NOT WITH THE MEDTRONIC DEVICE OR THE PROCEDURE. THE NON-COMPLAINT EVENTS INCLUDED: SEVERE VASOSPASM ASSOCIATED WITH THE SAH OCCURRED IN TWO PATIENTS, LEADING TO DELAYED CEREBRAL ISCHEMIA. FOUR PATIENTS NEEDED VENTRICULOPERITONEAL SHUNTING DUE TO SAH-RELATED HYDROCEPHALUS. Manufacturer Narrative: SEPARATE REPORTS WILL BE SUBMITTED FOR THE SERIOUS INJURIES REPORTED IN THE ARTICLE THAT DID NOT HAVE AN OUTCOME OF MORTALITY. A2. REPORTED PATIENT AGE IS REPRESENTATIVE OF THE MEAN AGE OF ALL PATIENTS INCLUDED IN THE STUDY GROUP. A3. REPORTED PATIENT SEX IS REPRESENTATIVE OF THE MAJORITY (5/9) PATIENTS INCLUDED IN THE STUDY GROUP. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21608255&pc=DSQ | 2916596-2025-01524 | 2025-02-22 05:00 | Death | THORATEC CORPORATION | 2025-03-14 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Cardiac Tamponade; Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED POST-OPERATIVE CARDIAC TAMPONADE WITH MULTIPLE RE-ENTRIES IN THE CHEST THAT ULTIMATELY SUCCUMBED TO MULTISYSTEM ORGAN FAILURE (MSOF). PER REQUEST FROM THE PATIENT'S FAMILY, THE PUMP WAS DEACTIVATED AND THE PATIENT PASSED AWAY. IT WAS UNKNOWN IF AN AUTOPSY WAS PERFORMED. THE DEVICE WAS NOT EXPLANTED. THE DEVICE OPERATED AS INTENDED AND THE PATIENT OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION,¿ LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH, MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION, AND PERICARDIAL FLUID COLLECTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21608375&pc=LPB | 2029046-2025-00793 | 2025-02-17 05:00 | Death | BIOSENSE WEBSTER INC | 2025-03-14 04:00 | LPB | THERMOCOOL SMARTTOUCH SF | Adverse Event Without Identified Device or Use Problem | Sepsis; Cardiac Tamponade; Cardiogenic Shock | P030031 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A PULMONARY VEIN ISOLATION (PVI) ABLATION WITH A THERMOCOOL SMARTTOUCH SF AND THE PATIENT EXPERIENCED CARDIAC TAMPONADE THAT REQUIRED PERICARDIOCENTESIS. PATIENT HAD A PERICARDIAL EFFUSION AFTER A CARTO CASE. SMARTABLATE POWER CONTROL 50 W WAS USED. ADDITIONAL INFORMATION WAS RECEIVED ON 03-MAR-2025. PUNCTURE OF THE CARDIAC TAMPONADE WAS PERFORMED FOR INTERVENTION. THE OUTCOME WAS DEATH. THE PHYSICIAN'S OPINION ON THE CAUSE OF THE ADVERSE EVENT IS THE PROCEDURE. ADDITIONAL INFORMATION WAS RECEIVED ON 04-MAR-2025. NO MALFUNCTION REPORTED ON THE GENERATOR. ONE TRANSSEPTAL PUNCTURE SITE PERFORMED DURING THE PROCEDURE. PRIOR TO NOTING THE PERICARDIAL EFFUSION, ABLATION WAS PERFORMED. THE EVENT OCCURRED AFTER THE PATIENT WAS OUT OF THE LAB. NO ERROR MESSAGES OBSERVED ON BIOSENSE WEBSTER EQUIPMENT DURING THE PROCEDURE. THE CATHETER SETTINGS SELECTED ON THE GENERATOR WERE 50 W/20 ML/MIN. ADDITIONAL INFORMATION WAS RECEIVED ON 12-MAR-2025. PHYSICIAN DID NOT FOLLOW THE INSTRUCTIONS FOR USE (IFU) FOR FLOW RATE AND POWER, HE USED 20 ML/MIN. THE DURATION OF THE APPLICATION WITH 50 W WAS UNKNOWN AS THE PERICARDIAL EFFUSION WAS OBSERVED AFTER THE CASE. Manufacturer Narrative: D4: UDI: AS THE LOT NUMBER FOR THE DEVICE INVOLVED IN THE EVENT WAS NOT PROVIDED, THE FULL UDI IS CURRENTLY NOT AVAILABLE. AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. AN EVALUATION OF THE MANUFACTURING RECORD COULD NOT BE PERFORMED AS THE REQUIRED PRODUCT IDENTIFICATION NUMBER WAS NOT PROVIDED TO COMPLETE THE EVALUATION. HOWEVER, IF THE PRODUCT OR PRODUCT ID NUMBERS ARE RECEIVED AT A LATER DATE, THE INVESTIGATION WILL BE UPDATED AS APPLICABLE. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. MANUFACTURER'S REFERENCE NUMBER: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED ON 15-APR-2025 INDICATED THAT THE PHYSICIAN¿S OPINION OF THE CAUSE OF DEATH WAS THERAPY-REFRACTORY COMBINED SEPTIC AND CARDIOGENIC SHOCK. AFTER THE PVI PROCEDURE, THE PATIENT DEVELOPED A CARDIAC TAMPONADE WHICH WAS TREATED BY PERICARDIOCENTESIS ON (B)(6) 2025. INITIALLY THE PATIENT FULLY RECOVERED, BUT SEVERAL HOURS LATER THE PATIENT¿S CONDITION WORSENED WITH POSITIVE SHOCK INDEX. EVENTUALLY, WE EXAMINED THE PATIENT ONCE AGAIN AND MADE ANOTHER PERICARDIOCENTESIS PERFORMED IN THE AFTERNOON. THE PATIENT¿S RESPIRATORY CONDITION ALSO WORSENED, SO THAT THE PATIENT HAD TO BE INTUBATED BY THE ANESTHESIOLOGISTS. AFTERWARDS THE PATIENT WENT INTO A SEPTIC AND CARDIOGENIC SHOCK AND DIED UNDER MAXIMAL INVASIVE MANAGEMENT AND MAXIMAL MEDICAL THERAPY SEVERAL HOURS LATER ON (B)(6) 2025, AFTER THE PATIENT WAS REANIMATED SEVERAL TIMES WITHOUT GETTING A RETURN OF SPONTANEOUS CIRCULATION (ROSC) FINALLY. THEREFORE, UPDATED THE B2. DATE OF DEATH FIELD. IN ADDITION, ADDED UNDER H6. HEALTH EFFECT - CLINICAL CODE FIELD ¿CARDIOGENIC SHOCK (E233601)¿ AND ¿SEPSIS (E0306)¿. IF ADDITIONAL INFORMATION IS RECEIVED REGARDING THIS EVENT, A SUPPLEMENTAL 3500A REPORT WILL BE SUBMITTED TO THE FDA. MANUFACTURER'S REFERENCE NUMBER: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21608394&pc=DSQ | 2916596-2025-01545 | 2025-02-25 05:00 | Death | THORATEC CORPORATION | 2025-03-14 04:00 | DSQ | THORATEC® HEARTMATE®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Asystole | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD PSEUDOMONAS BACTEREMIA. THE PATIENT WAS MADE DO NOT RESUSCITATE (DNR), EXPERIENCED ASYSTOLE, AND ULTIMATELY PASSED AWAY ON (B)(6) 2025. THE VENTRICULAR ASSIST DEVICE (VAD) WAS STATED TO HAVE FUNCTIONED AS INTENDED. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: THE PUMP REMAINS IN USE SUPPORTING THE PATIENT. A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF INFECTION COULD NOT BE CONCLUSIVELY DETERMINED. INFECTION IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM. INFECTION HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21608435&pc=OZD | 1220648-2024-17106 | 2021-05-23 04:00 | Death | ABIOMED, INC. | 2025-03-14 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Perforation | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. WHILE ON SUPPORT, THE PHYSICIAN DISCOVERED A WIRE PERFORATION TO THE LEFT VENTRICLE AFTER THE IMPELLA WAS PLACED. THIS PERFORATION WAS SUCCESSFULLY REPAIRED. THE PATIENT WAS TRANSFUSED WITH MULTIPLE BLOOD PRODUCTS. ADDITIONALLY, THE PATIENT EXPERIENCED A GASTROINTESTINAL (GI) BLEED; HOWEVER, THE ACTIONS TAKEN TO RESOLVE THE ISSUE ARE UNKNOWN. ADDITIONAL INFORMATION NOTED CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. MEDICAL SAFETY REVIEW OF THE EVENT NOTED THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA DEVICE AS AN ASSOCIATED FACTOR IN THE PATIENT'S DEMISE. Manufacturer Narrative: THIS COMPLAINT HAS BEEN IDENTIFIED AS PART OF A RETROSPECTIVE REVIEW. DUE TO THE LIMITED CLINICAL INFORMATION AND LACK OF AVAILABLE DATA/PRODUCT THE ROOT CAUSE OF THIS INVESTIGATION IS NOT DETERMINED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA IFU: IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿ SECTION: WARNINGS & CAUTIONS: WARNINGS SECTION: PRE-SUPPORT EVALUATION SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER SECTION: DIRECT AORTIC INSERTION SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION. [ADDITION FOR 5.0, 5.5 ONLY] IN INSTANCES WHERE THE IMPELLA PUMP HAS BEEN PLACED PRIOR TO PERFORMING CARDIAC SURGERY WITH AORTIC CROSS CLAMPING AND CARDIOPLEGIC ARREST, CARE SHOULD BE TAKEN WHEN MANIPULATING THE HEART WHEN THE PUMP POSITION IS FIXED WITH APPLICATION OF THE AORTIC CROSS CLAMP ACROSS THE PUMP CATHETER." ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE. SECTION: WARNINGS & CAUTIONS: WARNINGS SECTION: PRE-SUPPORT EVALUATION SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER SECTION: DIRECT AORTIC INSERTION SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES ¿TO REDUCE THE RISK OF CARDIAC INJURY (INCLUDING VENTRICULAR PERFORATION), PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY IN THE SETTINGS OF ACUTE INFARCTION WITH TISSUE NECROSIS.¿ ¿TO REDUCE THE RISK OF VASCULAR INJURY, PHYSICIANS SHOULD EXERCISE CAUTION WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX PERIPHERAL VASCULAR ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: UNREPAIRED ABDOMINAL AORTIC ANEURYSM, SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM, DISSECTION OF THE ASCENDING/ TRANSVERSE/DESCENDING AORTA, CHRONIC ANATOMICAL CHANGES IN THE RELATIONSHIP OF THE AORTA/AORTIC VALVE/VENTRICULAR ALIGNMENT, SIGNIFICANT MOBILE ATHEROMATOUS DISEASE IN THE THORACIC OR ABDOMINAL AORTA OR PERIPHERAL VESSELS.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING A RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY IN THE LEFT VENTRICLE AFTER CPR WITH ECHOCARDIOGRAPHY GUIDANCE. SECTION: WARNINGS & CAUTIONS: CAUTIONS ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ THIS REPORT IS BEING FILED AS PART OF A RETROSPECTIVE REVIEW OF HISTORICAL RECORDS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21608443&pc=KDI | 3003464075-2025-00054 | 2025-02-18 05:00 | Death | NXSTAGE MEDICAL, INC. | 2025-03-14 04:00 | KDI | NXSTAGE SYSTEM ONE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | K122051 |  | 1 | Event Description: A REPORT WAS RECEIVED ON 19 FEB 2025 FROM THE HOME THERAPY NURSE (HTN) OF A 55-YEAR-OLD MALE PATIENT WITH A MEDICAL HISTORY INCLUDING END STAGE RENAL DISEASE, WHO STATED THE PATIENT EXPIRED DURING A HOME HEMODIALYSIS TREATMENT ON (B)(6) 2025. ALTHOUGH REQUESTED, ADDITIONAL INFORMATION WAS NOT PROVIDED. Manufacturer Narrative: THE DEVICE WAS NOT RECEIVED FOR EVALUATION. A REVIEW OF THE DEVICE HISTORY RECORD (DHR) WAS CONDUCTED WHICH CONFIRMED THAT THE DEVICE MET ALL QUALITY CRITERIA AND MANUFACTURING SPECIFICATIONS PRIOR TO RELEASE. THERE WAS NO INDICATION OF A DEVICE MALFUNCTION FROM THE AVAILABLE INFORMATION. THE NXSTAGE SYSTEM ONE USER GUIDE OUTLINES RISKS ASSOCIATED WITH PERFORMING HEMODIALYSIS THERAPY AND WARNS ALL TREATMENTS MUST BE ADMINISTERED UNDER A PHYSICIAN'S PRESCRIPTION AND MUST BE PERFORMED BY A TRAINED AND QUALIFIED PERSON, CONSIDERED TO BE COMPETENT IN THE USE OF THIS DEVICE BY THE PRESCRIBING PHYSICIAN. UDI: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21608658&pc=MNT | 2518422-2025-034865 | 2025-03-01 05:00 | Death | RESPIRONICS, INC. | 2025-03-14 04:00 | MNT | RESPIRONICS | No Audible Alarm | Insufficient Information | K102985 |  | 1 | Event Description: PHILIPS RECEIVED A COMPLAINT FROM THE CUSTOMER, REPORTING THAT THE V60 VENTILATOR DID NOT ALARM WHEN THE DEVICE WAS DISCONNECTED. THE DEVICE WAS IN CLINICAL USE WHEN THE ISSUE OCCURRED. IT WAS REPORTED THAT THE EVENT WAS FATAL TO THE PATIENT. A FOLLOW UP WAS PERFORMED WITH THE KEY MARKET (KM), AND THE RESPONDER WAS UNABLE TO PROVIDE ANY DETAILS REGARDING THE PATIENT STATUS. THE KM RESPONDER REPORTED THAT THEY HAVE REACHED OUT TO THE HOSPITAL SEVERAL TIMES; HOWEVER, THE HOSPITAL STAFF HAVE NOT PROVIDED ANY DETAILS. CLINICAL ASSESSMENT HAS FOUND NO CAUSAL RELATIONSHIP OF THE DEVICE TO THE PATIENT EXPIRATION, HOWEVER CONTRIBUTION OF THE DEVICE AND ITS USE TO THE PATIENT EXPIRATION HAS BEEN CONFIRMED: FORESEEABLE UNINTENTIONAL USE ERROR - ADDITION OF RESISTIVE COMPONENTS TO THE PATIENT SIDE OF THE PROXIMAL PRESSURE LINE RESULTING IN COMPROMISED PATIENT DISCONNECTION ALARM FUNCTIONALITY. THE USER REPORTED THAT THE PATIENT WAS CONNECTED TO THE V60 VENTILATOR, BUT AFTER THE FOLLOWING MORNING, THE USER FOUND THAT THE VENTILATOR HAD BEEN DISCONNECTED FROM THE PATIENT, AND THAT THE DEVICE DID NOT ALARM (PATIENT DISCONNECT). AFTER SEVERAL INQUIRIES, AND A VIDEO FOOTAGE OF THE DEVICE, IT WAS OBSERVED THAT USER HAD USED TWO BACTERIA FILTER (MAKE/MODEL UNSPECIFIED) CONNECTED TO THE DEVICE. PER THE INSTRUCTIONS FOR USE (IFU), IT IS ADVISED TO "AVOID ADDING ANY COMPONENTS TO THE PATIENT CIRCUIT THAT ARE NOT ABSOLUTELY NECESSARY. ADDITIONAL COMPONENTS INSTALLED IN THE PATIENT CIRCUIT CAN CHANGE THE PRESSURE GRADIENT ACROSS THE VENTILATOR BREATHING SYSTEM, INCREASE THE DEAD SPACE, AND ADVERSELY AFFECT THE VENTILATOR PERFORMANCE." IN ADDITION, THE IFU STATES, "AVOID ADDING RESISTIVE CIRCUIT COMPONENTS ON THE PATIENT SIDE OF THE PROXIMAL PRESSURE LINE. SUCH COMPONENTS MAY DEFEAT THE DISCONNECT ALARM." THE DEVICE WAS THEN EVALUATED BY PHILIPS SERVICE ENGINEER (SE), AND IT WAS REPORTED THE RESPIRATORY CIRCUIT WAS CHECKED, AND IT WAS CONFIRMED THAT A FILTER BETWEEN THE PATIENT AND EXHALATION PORT, AND A NEBULIZATION DEVICE BETWEEN THE FILTER AND EXHALATION PORT. WHEN THE SE DISCONNECTED THE CIRCUIT CONSUMABLE, THE PATIENT DISCONNECT ALARM IS RE-ESTABLISHED. THIS INVESTIGATION IS ONGOING. Manufacturer Narrative: (B)(6). Manufacturer Narrative: BASED ON THE INFORMATION PROVIDED, THE EVENT HAS SHOWN THAT THE DEVICE PERFORMED AS INTENDED WITH NO FAILURE OR MALFUNCTION TO PERFORM TO MANUFACTURER DECLARE SPECIFICATIONS. THE ADDITIONAL COMPONENTS ADDED TO THE PATIENT SIDE OF THE PROXIMAL PRESSURE LINE RESULTED IN THE DEVICE PATIENT DISCONNECTION ALARM BEING DEFEATED. FURTHER GOOD FAITH EFFORTS YIELDED NO ADDITIONAL INFORMATION REGARDING THE PATIENT STATUS. THE INVESTIGATION CONCLUDES THAT NO FURTHER ACTION IS REQUIRED AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED, THE COMPLAINT FILE WILL BE REOPENED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21609896&pc=NPT | 2135147-2025-01315 | 2025-02-17 05:00 | Death | ABBOTT MEDICAL | 2025-03-14 04:00 | NPT | NAVITOR TRANSCATHETER AORTIC VALVE | Migration or Expulsion of Device; Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P190023 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 23MM NAVITOR VALVE WAS SELECTED FOR IMPLANT USING A SMALL FLEXNAV DELIVERY SYSTEM. THE PATIENT'S CONDITION PRIOR TO PROCEDURE WAS "GOOD BUT FRAGILE." PRE-DILATATION WAS NOT PERFORMED. DURING PROCEDURE, THE VALVE WAS IMPLANTED AS PART OF A VALVE-IN-VALVE INTO A 21MM TRIFECTA VALVE AND POST-DILATATION WAS PERFORMED. AFTER POST-DILATATION, THE NAVITOR VALVE "WAS CARRIED INTO THE AORTIC ARCH WHEN THE GUIDE WAS REMOVED." THE IMPLANT DEPTH AT 80% WAS 0.5MM AND THE IMPLANT DEPTH AT RELEASE PRIOR TO MIGRATION WAS 0.5MM. THE DS WAS IN NEUTRAL POSITION AND THE GUIDEWIRE WAS IN MIDVENTRICULAR POSITION PRIOR TO FINAL DEPLOYMENT. THE NAVITOR VALVE WAS NOT SNARED INTO POSITION AND IT DID NOT BLOCK ANY ARTERIES. THERE WAS NO ENTANGLEMENT WITH THE DS. THERE ARE NO ALLEGATIONS OF MALFUNCTION WITH THE NAVITOR VALVE. ON (B)(6) 2025, THE PATIENT PASSED AWAY DUE TO UNKNOWN CAUSES. THE USER SUSPECTS THAT THE PATIENT PASSING WAS RELATED TO THE PROCEDURE.Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 23MM NAVITOR VALVE WAS SELECTED FOR IMPLANT USING A SMALL FLEXNAV DELIVERY SYSTEM. THE VALVE WAS IMPLANTED AS PART OF A VALVE-IN-VALVE INTO A 21MM TRIFECTA VALVE DUE TO VALVE DEGENERATION. THE PATIENT'S CONDITION PRIOR TO PROCEDURE WAS "GOOD BUT FRAGILE." THE PATIENT'S AORTIC VALVE ANGLE WAS NOTED AS NORMAL. PRE-DILATATION WAS NOT PERFORMED. THE IMPLANT DEPTH AT 80% WAS 5MM AND THE IMPLANT DEPTH AT RELEASE PRIOR TO MIGRATION WAS 5MM. THE DS WAS IN NEUTRAL POSITION AND THE GUIDEWIRE WAS IN MIDVENTRICULAR POSITION PRIOR TO FINAL DEPLOYMENT. POST-DILATATION WAS PERFORMED. AFTER POST-DILATATION, THE NAVITOR VALVE "WAS CARRIED INTO THE AORTIC ARCH WHEN THE GUIDE WAS REMOVED." THE NAVITOR VALVE WAS NOT SNARED INTO POSITION AND IT DID NOT BLOCK ANY ARTERIES. THERE WAS NO ENTANGLEMENT WITH THE DS. THERE ARE NO ALLEGATIONS OF MALFUNCTION WITH THE NAVITOR VALVE. ON (B)(6) 2025, THE PATIENT PASSED AWAY DUE TO UNKNOWN CAUSES. THE USER SUSPECTS THAT THE PATIENT PASSING WAS RELATED TO THE PROCEDURE. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. Manufacturer Narrative: AN EVENT OF MIGRATION OR EXPULSION OF DEVICE (AORTIC DIRECTION) WAS REPORTED. A RETURNED DEVICE ASSESSMENT COULD NOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED FOR ANALYSIS. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED, AND THE PRODUCT MET ALL SPECIFICATIONS. BASED ON THE INFORMATION RECEIVED, THE CAUSE FOR THE REPORTED DEVICE MIGRATION IN AORTIC AND DEATH COULD NOT BE DETERMINED. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH REGARDS TO MANUFACTURE, DESIGN, OR LABELING. CODES REMOVED: MEDICAL DEVICE PROBLEM CODE: 2993 ADVERSE EVENT WITHOUT IDENTIFIED DEVICE OR USE PROBLEM.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21610044&pc=DSQ | 2916596-2025-01651 | 2025-01-06 05:00 | Death | THORATEC CORPORATION | 2025-03-14 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Mechanical Problem; Electrical Power Problem; Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT ARRIVED AT THE EMERGENCY DEPARTMENT (ED) WITH NO POWER AND DEAD BATTERIES ON THEIR LEFT VENTRICULAR ASSIST DEVICE (LVAD). THE PATIENT RECEIVED CARDIOPULMONARY RESUSCITATION (CPR). WITH A NEW SET OF BATTERIES, THE EMERGENCY ROOM (ER) WAS ABLE TO POWER THE DEVICE BACK ON BUT HAD ACTIVATED THE LOW FLOW ALARM. THE ER WAS UNABLE TO GET A RETURN OF SPONTANEOUS CIRCULATION (ROSC). THE ER HAD CALLED TIME OF DEATH AFTER 45 MINUTES, AND THEY HAD NOTICED THE PATIENT HAD BLOWN PUPILS. THE PATIENT¿S FAMILY HAD RETURNED THE PRIMARY SYSTEM CONTROLLER TO THE HEALTHCARE PROVIDER. THE LOG FILES WERE SUBMITTED FOR REVIEW TO CONFIRM IF THE PATIENT¿S PASSING WAS DUE TO ISSUES WITH THE PUMP. THE LOG FILES WERE REVIEWED AND REVEALED VARIOUS LOSS OF EXTERNAL POWER ALARMS ON (B)(6) 2025. THE LOG FILES ALSO CONTAINED VARIOUS DRIVELINE DISCONNECT/RECONNECT ALARMS. WHEN THE SYSTEM CONTROLLER WAS POWERED BACK ON, A CONTROLLER CLOCK CORRUPT ALARM HAD ACTIVATED. THE MOST RECENT DATA IN THE LOG FILES INDICATED THAT THE PUMP WAS NOT CONNECTED WHEN THE DATA WAS DOWNLOADED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE PATIENT'S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE CONTROLLER EVENT LOG FILE CONTAINED EVENTS FROM 04JAN2025 THROUGH 06JAN2025, AS WELL AS EVENTS WITH THE CLOCK NOT SYNCED, DISPLAYING A DEFAULT TIMESTAMP OF 01JAN2000. NO OTHER NOTABLE EVENTS OR ALARMS WERE CAPTURED. THE LOG FILE CAPTURED A NO EXTERNAL POWER ALARM ON 06JAN2025 AT 4:35:06 DUE TO FULL BATTERY DEPLETION FOLLOWING EXTENDED USE. THE BACKUP BATTERY SUPPORTED THE SYSTEM FOR THE NEXT TWO HOURS UNTIL IT ALSO BECAME FULLY DEPLETED; THEN, THE PATIENT¿S PUMP STOPPED ON 06JAN2025 AT 6:32:56. THE SYSTEM CONTROLLER FULLY SHUT DOWN A FEW MINUTES LATER, AND THE PUMP WAS OFF FOR AN UNKNOWN AMOUNT OF TIME. AFTER THE PUMP WAS RESTARTED, TWO DRIVELINE DISCONNECTIONS AND RECONNECTIONS WERE OBSERVED. THE PUMP RAMPED BACK UP TO THE SET SPEED FOLLOWING EACH DRIVELINE RECONNECTION; HOWEVER, LOW FLOW ALARMS WERE CAPTURED, CONSISTENT WITH THE REPORTED EVENTS AND PATIENT DEATH. NO OTHER NOTABLE EVENTS WERE CAPTURED. THE PUMP APPEARED TO FUNCTION AS INTENDED FOR THE DURATION OF THE LOG FILE. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED BY THE ACCOUNT. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 7, ¿ALARMS AND TROUBLESHOOTING,¿ EXPLAINS SYSTEM ALARMS AND THE RECOMMENDED ACTIONS ASSOCIATED WITH THEM. THE HEARTMATE 3 LVAS PATIENT HANDBOOK IS CURRENTLY AVAILABLE. THE PATIENT HANDBOOK CAUTIONS USERS TO CALL THEIR HOSPITAL CONTACTS IF THEY THINK, FOR ANY REASON, ANY PORTION OF THEIR EQUIPMENT IS NOT FUNCTIONING AS USUAL, IS BROKEN, OR THEY ARE UNCOMFORTABLE WITH THE OPERATION OF THE EQUIPMENT. THE PATIENT HANDBOOK ALSO INSTRUCTS USERS TO REGULARLY ENSURE THAT THEIR DRIVELINE IS CONNECTED AND SECURED WITHIN THE SYSTEM CONTROLLER, AND TO RECONNECT IT IMMEDIATELY IN THE CASE THAT IT BECOMES DISCONNECTED. USERS ARE WARNED THAT THE PUMP WILL STOP RUNNING WHEN THE DRIVELINE IS DISCONNECTED. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21610055&pc=PYX | 1220648-2025-26738 | 2025-02-17 05:00 | Death | ABIOMED, INC. | 2025-03-14 04:00 | PYX | IMPELLA RP | Restricted Flow rate | Hemolysis; Renal Failure; Tachycardia | P170011 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT IN POSTCARDIOTOMY CARDIOGENIC SHOCK WAS INITIALLY IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS) AND EXPERIENCED LOW PUMP FLOW AND BECAME HEMODYNAMICALLY UNSTABLE. THE PATIENT WAS ESCALATED TO AN IMPELLA 5.5 DEVICE AND ALSO PLACED ON AN IMPELLA RP DEVICE FOR BIPELLA MCS SUPPORT. THE PATIENT WOULD EXPERIENCE HEMOLYSIS, LOW PUMP FLOW AND SUBSEQUENTLY EXPIRE. THE FOLLOWING DAY AFTER BIPELLA PLACEMENT, CONTINUOUS RENAL REPLACEMENT THERAPY WAS STARTED WITH THE PLAN TO INCREASE FLOW AS PATIENT TOLERATES AND NOT REMOVING VOLUME AT SUCH TIME. THE PATIENT WAS NOTED TO BE PINK WARM AND DRY. IMPELLA RP DRESSING WAS DRY AND INTACT IN RIGHT FEMORAL VEIN AND THE LEFT SUBCLAVIAN IMPELLA 5.5 DRESSING SEROSANGUINEOUS. THE PATIENT'S URINE WAS NOTED TO BE PURPLE AND PINK TINGED. PATIENT DID RECEIVE METHYLENE BLUE THE PRIOR NIGHT IT WAS NOTED. THE POSSIBLE HEMOLYSIS WAS DISCUSSED AND SUSPECTED. RENAL FAILURE WAS NOTED AND ALSO THE PATIENT WAS REPORTED TO HAVE SYSTEMIC ORGAN FAILURE. TWO UNITS OF RED BLOOD CELLS WERE GIVEN. THE FOLLOWING DAY IT WAS NOTED THE PATIENT REQUIRED GREATER VASOPRESSOR SUPPORT OVERNIGHT AFTER CRRT STARTED PULLING ~100/HR. THE PATIENT WAS REQUIRED ONE SHOCK DUE TO ABNORMAL RHYTHM. CHEST X-RAY SHOWED THE IMPELLA RP SHALLOW BUT APPEARED TO BE ACROSS PULMONIC VALVE. THE IMPELLA RP WAS AT PERFORMANCE LEVEL P-9 AND FLOWING 3.4 L/MIN, IMPELLA 5.5 WAS AT P4 AND FLOWING 2.5. LOW PURGE PRESSURE ALARMS FROM THE IMPELLA RP DEVICE OCCURRED AND WERE RESOLVED. DISCUSSION WAS MADE WITH FAMILY, AND DECISION WAS MADE TO MOVE THE PATIENT TO COMFORT CARE. THE PATIENT WOULD EXPIRE THEREAFTER. MEDICAL SAFETY REVIEW OF THE EVENT NOTED THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA RP AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME (DEMISE). THE IMPELLA RP FLOWS DROPPED; HEMODYNAMIC CONSEQUENCE WITH THE IMPELLA RP. Manufacturer Narrative: THE INVESTIGATION OF HEMOLYSIS HAS BEEN COMPLETED. THE PRODUCT AND DATA WERE NOT RETURNED FOR REVIEW. BASED ON CLINICAL DESCRIPTION, THE ROOT CAUSE OF HEMOLYSIS WAS MOST LIKELY DUE TO PUMP WAS NOT IN THE PROPER POSITION AS THE RP PUMP WAS UNABLE TO ADVANCE INTO LEFT PULMONARY ARTERY. ON 2/18, PATIENT WENT INTO VENTRICULAR TACHYCARDIA AT 1003, IMPELLAS TURNED TO P2 AND TIME OF DEATH CALLED AT 1013. THE CAUSE OF THE VT STORM WAS UNABLE TO BE DETERMINED DUE TO INSUFFICIENT CLINICAL DETAILS. B1 PRODUCT PROBLEM WAS ERRONEOUSLY SELECTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26738. H6 HEALTH EFFECT - CLINICAL CODE 2095 WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26738. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THIS IS ONE OF THREE DEVICE MANUFACTURING REPORTS ASSOCIATED WITH THE EVENT AND/OR IMPLANT EXPERIENCE. REFER TO THE FOLLOWING: MEDICAL DEVICE REPORT FOR THE IMPELLA CP SERIAL NUMBER (B)(6) WITH DATE OF THE EVENT 16-FEB-2025 AND PATIENT IDENTIFIER ENDING IN (B)(6). MEDICAL DEVICE REPORT FOR THE IMPELLA 5.5 SERIAL NUMBER (B)(6) WITH DATE OF THE EVENT 17-FEB-2025 AND PATIENT IDENTIFIER ENDING IN (B)(6). MEDICAL DEVICE REPORT FOR THE IMPELLA RP SERIAL NUMBER (B)(6) WITH DATE OF THE EVENT 17-FEB-2025 AND PATIENT IDENTIFIER ENDING IN (B)(6). (THIS REPORT)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21610146&pc=LWS | 9614453-2025-01046 | 2024-07-25 04:00 | Death | MEDTRONIC EUROPE SARL | 2025-03-14 04:00 | LWS | EVERA MRI XT DR SURESCAN | Inappropriate/Inadequate Shock/Stimulation | Shock from Patient Lead(s); Insufficient Information | P980016 |  | 1 | Event Description: IT WAS REPORTED BY THE PATIENT¿S SPOUSE THAT THE PATIENT IS DECEASED AND DIED LAST YEAR ON (B)(6) 2024. THE PATIENT¿S SPOUSE STATED THAT THE PATIENT WAS IN HOSPICE AND HAD RECEIVED SHOCKS FOR ABOUT 30 MINUTES. THE PATIENT¿S SPOUSE NOTED THAT THERE WERE ORDERS IN FOR DEVICE DEACTIVATION. HOWEVER, THE HOSPICE/NURSING DID NOT COORDINATE THE DEVICE DEACTIVATION AND THE ORDERS WERE CANCELLED DUE TO THE PATIENT PASSING AWAY. ADDITIONAL INFORMATION WAS REQUESTED FROM THE CLINIC BUT HAS NOT BEEN RECEIVED.Event Description: IT WAS FURTHER REPORTED THAT, DUE TO THE TONES OF THE DEVICE, THE HOSPICE TEAM WAS ABLE TO PLACE A MAGNET ON THE DEVICE AND THE PATIENT WAS ABLE TO PASS. Manufacturer Narrative: CONTINUATION OF D10: 6935M62 LEAD IMPLANTED: (B)(6)2022; 5076-52 LEAD IMPLANTED: (B)(6)2022. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21610153&pc=NPT | 2025587-2025-01923 | 2025-03-11 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-14 04:00 | NPT | EVOLUT FX PLUS VALVE | Perivalvular Leak; Activation Failure; Patient Device Interaction Problem | Rupture; Valvular Insufficiency/ Regurgitation | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT PRIOR TO THE IMPLANT OF A TRANSCATHETER VALVE, A PRE-IMPLANT BALLOON AORTIC VALVULOPLASTY (BAV) WAS PERFORMED USING A 20 MILLIMETER (MM) NON-MEDTRONIC BALLOON, WHICH WAS STANDARD FOR THE IMPLANTING PHYSICIAN. IT WAS NOTED THAT THE PATIENT HAD A PRE-IMPLANT PRESSURE GRADIENT OF 110 MILLIMETERS OF MERCURY (MMHG) AND CALCIFIED ANATOMY. FOLLOWING THE IMPLANT OF A TRANSCATHETER VALVE, MODERATE PARAVALVULAR LEAK (PVL) AND UNDER-EXPANSION OF THE VALVE FRAME WAS OBSERVED. SUBSEQUENTLY, A POST-IMPLANT BALLOON AORTIC VALVULOPLASTY (BAV) WAS PERFORMED USING A 24 MILLIMETER (MM) NON-MEDTRONIC BALLOON. FOLLOWING THE POST-IMPLANT BAV, THE PATIENT BEGAN TO HEMODYNAMICALLY COMPROMISE AND AN ANNULAR RUPTURE OCCURRED. IT WAS REPORTED THAT THE PATIENT WAS STABLE UPON LEAVING THE HOSPITAL; HOWEVER, IT WAS NOTED THE PATIENT WAS TAKEN OFF SUPPORT A FEW HOURS LATER AND DIED. IT WAS NOTED THAT THE CAUSE OF THE ANNULAR RUPTURE WAS DUE TO THE POST-IMPLANT BAV. PER THE PHYSICIAN, THE PATIENT¿S CALCIFIED ANATOMY CONTRIBUTED TO THE EVENT, AND THE CAUSE OF DEATH WAS DUE TO THE ANNULAR RUPTURE.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT PERICARDIOCENTESIS AND VASOPRESSOR SUPPORT WERE PERFORMED TO TREAT THE ANNULUS RUPTURE AND STABILIZE THE PATIENT. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID {D-EVOLUTFX-34}; PRODUCT LOT/SERIAL NUMBER {UNKNOWN}; PRODUCT TYPE: {1095 HEART VALVES}; PRODUCT ID {24 TRUE BALLOON}; PRODUCT LOT/SERIAL NUMBER {UNKNOWN}; PRODUCT TYPE: {BALLOON AORTIC VALVULOPLASTY}. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21610495&pc=MIH | 2247858-2025-00056 | 2023-05-19 04:00 | Death | BOLTON MEDICAL, INC. | 2025-03-14 04:00 | MIH | RELAY PLUS NBS THORACIC STENT-GRAFT SYSTEM | Adverse Event Without Identified Device or Use Problem; Insufficient Information; Patient Device Interaction Problem; Failure to Seal | Insufficient Information | P110038 |  | 1 | Event Description: "AN NON-SMOKING 46-YEAR-OLD MALE 181CM TALL, WEIGHING 92KG. WAS ADMITTED TO HOSPITAL ON (B)(6) 2023 AND TREATED ON (B)(6) 2023. THE PATIENT HAD A HISTORY OF HYPERTENSION (ONSET DATE UNKNOWN- ONGOING). THE PATIENT WAS PREVIOUSLY TREATED WITH THORAFLEX HYBRID AND GELWEAVE BIFURCATE DEVICE ON (B)(6) 2023) AND GELWEAVE TUBE ON (B)(6) 2013). THE PATIENT SUFFERS FALLS UNDER ASA CLASS III PATIENT WITH SYSTEMIC DISEASE AND CO-MORBIDITY SCORE OF 0.2. THE PATIENT IS ON ASPIRIN. THE PATIENT WAS TREATED VIA ELECTIVE SURGERY WITH A RELAY PLUS NBS DEVICE TO TREAT CHRONIC AORTIC DISSECTION>90 DAYS (ELONGATED THORACOABDOMINAL AORTIC DISSECTION) AS PART OF A PRIMARY EXTENSION PROCEDURE (DISTALLY). ACCESS WAS GAINED VIA PERCUTANEOUS CUT DOWN OF LEFT AND RIGHT COMMON ARTERIES- THERE WAS NO ARTERY COVERAGE. NO ENDOLEAKS, ADVERSE EVENTS OR DEVICE DEFICIENCIES WERE REPORTED AT THE END OF THE PROCEDURE. ON (B)(6) 2023 AN ENDOLEAK TYPE IV AT CONNECTION STENT GRAFT (RADIOLOGICALLY PROVEN WAS IDENTIFIED) IT WAS REPORTED AS UNDETERMINED WHETHER RELATED TO DEVICE OR PROCEDURE. NO ACTIONS WERE TAKEN- THE EVENT IS ONGOING. THE PATIENT WAS DISCHARGED ON (B)(6)2023 FOLLOWING 10 DAYS IN HOSPITAL TO PRE-OP LIVING CONDITIONS. ON (B)(6) 2023 THE SITE IDENTIFIED AN INCREASE IN SIZE OF THE DESCENDING AORTA WITH CONTINUED PERFUSION OF THE FALSE LUMEN DD ENDOLEAK TYPE IV, SURGERY WAS PLANNED. THE EVENT IS REPORTED BY SITE AS SERIOUS, RELATED TO THE PROCEDURE, UNDETERMINED WHETHER RELATED TO THE DEVICE AND ONGOING." PATIENT OUTCOME: "PATIENT DIED ON (B)(6) 2024 - NO FURTHER INFORMATION PROVIDED BY SITE"Event Description: "A NON-SMOKING 46-YEAR-OLD MALE 181CM TALL, WEIGHING 92KG. WAS ADMITTED TO HOSPITAL ON (B)(6) 2023 AND TREATED ON (B)(6) 2023. THE PATIENT HAD A HISTORY OF HYPERTENSION (ONSET DATE UNKNOWN- ONGOING). THE PATIENT WAS PREVIOUSLY TREATED WITH THORAFLEX HYBRID AND GELWEAVE BIFURCATE DEVICE ON (B)(6) 2023) AND GELWEAVE TUBE ON (B)(6) 2013). THE PATIENT SUFFERS FALLS UNDER ASA CLASS III PATIENT WITH SYSTEMIC DISEASE AND CO-MORBIDITY SCORE OF 0.2. THE PATIENT IS ON ASPIRIN. THE PATIENT WAS TREATED VIA ELECTIVE SURGERY WITH A RELAY PLUS NBS DEVICE TO TREAT CHRONIC AORTIC DISSECTION >90 DAYS (ELONGATED THORACOABDOMINAL AORTIC DISSECTION) AS PART OF A PRIMARY EXTENSION PROCEDURE (DISTALLY). ACCESS WAS GAINED VIA PERCUTANEOUS CUT DOWN OF LEFT AND RIGHT COMMON ARTERIES- THERE WAS NO ARTERY COVERAGE. NO ENDOLEAKS, ADVERSE EVENTS OR DEVICE DEFICIENCIES WERE REPORTED AT THE END OF THE PROCEDURE. ON (B)(6) 2023, AN ENDOLEAK TYPE IV AT CONNECTION STENT GRAFT (RADIOLOGICALLY PROVEN WAS IDENTIFIED) IT WAS REPORTED AS UNDETERMINED WHETHER RELATED TO DEVICE OR PROCEDURE. NO ACTIONS WERE TAKEN- THE EVENT IS ONGOING. THE PATIENT WAS DISCHARGED ON (B)(6) 2023 FOLLOWING 10 DAYS IN HOSPITAL TO PRE-OP LIVING CONDITIONS. ON (B)(6) 2023, THE SITE IDENTIFIED AN INCREASE IN SIZE OF THE DESCENDING AORTA WITH CONTINUED PERFUSION OF THE FALSE LUMEN DD ENDOLEAK TYPE IV, SURGERY WAS PLANNED. THE EVENT IS REPORTED BY SITE AS SERIOUS, RELATED TO THE PROCEDURE, UNDETERMINED WHETHER RELATED TO THE DEVICE AND ONGOING." PATIENT OUTCOME: "PATIENT DIED ON (B)(6) 2024 - NO FURTHER INFORMATION PROVIDED BY SITE." Manufacturer Narrative: BOLTON MEDICAL, INC. (D/B/A TERUMO AORTIC), HEREIN KNOWN AS THE "COMPANY", IS SUBMITTING THIS REPORT PURSUANT TO 21 CFR PART 803, HAS MADE REASONABLE EFFORTS TO OBTAIN COMPLETE INFORMATION, AND HAS PROVIDED AS MUCH AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT IS BASED ON INFORMATION OBTAINED BY THE COMPANY, WHICH MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE REGULATIONS. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN, UNAVAILABLE, OR NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY ANYONE THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE CERTAIN REGULATIONS, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. Manufacturer Narrative: BOLTON MEDICAL, INC. (D/B/A TERUMO AORTIC), HEREIN KNOWN AS THE "COMPANY", IS SUBMITTING THIS REPORT PURSUANT TO 21 CFR PART 803, HAS MADE REASONABLE EFFORTS TO OBTAIN COMPLETE INFORMATION, AND HAS PROVIDED AS MUCH AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT IS BASED ON INFORMATION OBTAINED BY THE COMPANY, WHICH MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE REGULATIONS. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN, UNAVAILABLE, OR NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY ANYONE THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE CERTAIN REGULATIONS, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21590785&pc=GDW | 1219930-2025-01208 | 2025-02-17 05:00 | Death | COVIDIEN LP LLC NORTH HAVEN | 2025-03-13 04:00 | GDW | SIGNIA | Insufficient Information; Complete Loss of Power | Insufficient Information; No Clinical Signs, Symptoms or Conditions | K160176 |  | 1 | Event Description: ACCORDING TO THE REPORTER, DURING THE PROCEDURE, THE POWERED STAPLER WHICH WAS ASSEMBLED WITH THE SHELL TURNED OFF AND DID NOT TURN ON AGAIN. A MANUAL STAPLER WAS USED INSTEAD. THERE WAS NO PATIENT INJURY. Manufacturer Narrative: ADDITIONAL REVIEW OF THE EVENT SHOWS THAT THERE IS NO INFORMATION TO REASONABLY SUGGEST THAT THE DEVICE IN THIS REPORT MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH OR SERIOUS INJURY OR THAT THE DEVICE IN THIS REPORT HAS MALFUNCTIONED. THEREFORE, THIS EVENT DID NOT AND DOES NOT MEET THE REPORTING REQUIREMENTS IN 21 CFR 803. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: D10 CONCOMITANT MEDICAL PRODUCT: SIGPHANDLE, SIG POWER SIGPHANDLE HANDLE, (SERIAL# UNKNOWN) MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21590929&pc=NIP | 2017233-2025-05932 | 2025-02-27 05:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-03-13 04:00 | NIP | GORE® VIABAHN® ENDOPROSTHESIS WITH HEPARIN BIOACTIVE SURFACE | Adverse Event Without Identified Device or Use Problem | Hematoma; Hemorrhage/Bleeding; Unspecified Infection; Pseudoaneurysm; Ruptured Aneurysm | P040037 |  | 1 | Event Description: THE FOLLOWING INFORMATION WAS REPORTED TO GORE: ON AN UNKNOWN DATE IN (B)(6) 2023, THE PATIENT VISITED THE HOSPITAL COMPLAINING OF COLDNESS AND PAIN IN BOTH LOWER LIMBS. WOUNDS ON THE RIGHT LOWER LIMB, EXTENSIVE OCCLUSION, AND LEFT SUPERFICIAL FEMORAL ARTERY OCCLUSION WERE OBSERVED. ON AN UNKNOWN DATE, THE PATIENT UNDERWENT TREATMENT. A GORE® VIABAHN® ENDOPROSTHESIS WITH HEPARIN BIOACTIVE SURFACE (VIABAHN) WAS PLACED IN THE RIGHT EXTERNAL ILIAC ARTERY, AND A BOVINE PERICARDIAL PATCH WAS PLACED IN THE COMMON FEMORAL ARTERY. ON AN UNKNOWN DATE, AFTER REHABILITATION, THE PATIENT WAS DISCHARGED HOME. ANTIBIOTIC TREATMENT CONTINUED FOR MORE THAN THREE MONTHS. AFTER DISCHARGE, HOME MEDICAL CARE CONTINUED, AND THE WOUNDS IMPROVED, BUT SWELLING AND PAIN IN THE RIGHT GROIN WERE NOTED. AN ULTRASOUND EXAMINATION REVEALED BLOOD FLOW INTO THE ANEURYSM AT THE DISTAL END OF THE VIABAHN PLACED IN THE RIGHT ILIAC ARTERY, SUGGESTING A RUPTURED INFECTIOUS ANEURYSM. ON AN UNKNOWN DATE, THE PATIENT UNDERWENT EMERGENCY TREATMENT. TO CONTROL THE INFECTION, TRIMMING OF THE TISSUE AROUND THE WOUND WAS CARRIED OUT. SURGICAL TREATMENT WAS DEEMED DIFFICULT, SO AN ADDITIONAL VIABAHN WAS PLACED IN THE RIGHT COMMON FEMORAL ARTERY, AND THE WOUND WAS LEFT OPEN FOR DRAINAGE. ALTHOUGH THE ABI (ANKLE-BRACHIAL PRESSURE INDEX) IMPROVED, CONTINUED WOUND MANAGEMENT AND ANTIBIOTIC TREATMENT WERE NECESSARY. ON AN UNKNOWN DATE, DESPITE REGULAR CLEANING, GOOD GRANULATION FORMATION WAS NOT OBSERVED, AND EXPOSURE OF THE VIABAHN WAS CONFIRMED. LONG-TERM WOUND MANAGEMENT CONTINUED, BUT INFECTION CONTROL WAS DIFFICULT. ON AN UNKNOWN DATE, THE PATIENT DEVELOPED HEMORRHAGIC SHOCK FROM THE WOUND AND EXPIRED. REPORTEDLY, THE CAUSE OF DEATH WAS A PSEUDOANEURYSM, HEMATOMA, AND RUPTURED INFECTIOUS ANEURYSM. THE PHYSICIAN STATED THE FOLLOWING: EARLY REOPERATION, HEMATOMA REMOVAL, AND APPROPRIATE ANTIBIOTIC TREATMENT MIGHT HAVE PREVENTED THE RUPTURE OF THE INFECTIOUS ANEURYSM. Manufacturer Narrative: HEPARIN SURFACE INCORPORATES CARMEDA HEPARIN MANUFACTURED FROM HEPARIN SODIUM API, WHICH IS COVALENTLY BOUND TO THE DEVICE SURFACE AND IS ESSENTIALLY NON-ELUTING. H6: A REVIEW OF THE MANUFACTURING RECORDS FOR THE DEVICE COULD NOT BE PERFORMED AS A VALID LOT NUMBER WAS NOT PROVIDED. A1: NO PATIENT SPECIFIC DETAILS HAVE BEEN PROVIDED. THEREFORE, THE PATIENT INITIALS REFLECT THE W. L. GORE INTERNAL CASE NUMBER. B2: THE DATE OF PATIENT DEATH IS UNKNOWN. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21591268&pc=OZP | 2032227-2025-134623 | 2025-02-09 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-03-13 04:00 | OZP | MINIMED 780G US SYSTEM BLE CONNECT 3.0 MG/DL | Adverse Event Without Identified Device or Use Problem | Hypoglycemia; Heart Failure/Congestive Heart Failure | P160017 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER'S NEXT OF KIN SAID THAT CUSTOMER PASSED AWAY ON (B)(6) 2025. CAUSE OF THE DEATH WAS LOW BLOOD SUGAR EVENT AND HEART COMPLICATION. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: MMT-332A, MMT-397A, MMT-1884, MMT-7911NA, MMT-7040A. TROUBLESHOOTING WAS PERFORMED FOR DECEASED REPORTING. THE PUMP WAS WORN AT THE TIME OF PASSING. CUSTOMER WAS USING THE INSULIN PUMP SYSTEM AT THE TIME OF REPORTED EVENT. THERE IS NO MENTION OF THE AUTO MODE FEATURE AT THE TIME OF THE EVENT AND THE CUSTOMER WAS USING SENSOR AT THE TIME OF EVENT. NO PRODUCT RETURN IS REQUIRED FOR MMT-332A. NO PRODUCT RETURN IS REQUIRED FOR MMT-397A. MMT-1884 WAS REQUESTED AND RESPONSE WAS THE DEVICE WILL BE RETURNED. NO PRODUCT RETURN IS REQUIRED FOR MMT-7911NA. NO PRODUCT RETURN IS REQUIRED FOR MMT-7040A. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. THE DEVICE WILL BE RETURNED FOR ANALYSIS AND FURTHER INFORMATION WILL FOLLOW ONCE THE ANALYSIS HAS BEEN COMPLETED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: THE PUMP PASSED THE FUNCTIONAL TEST, INCLUDING THE SELF TEST, SLEEP CURRENT MEASUREMENT, ACTIVE CURRENT MEASUREMENT, REWIND TEST, PRIME/SEATING TEST, BASIC OCCLUSION TEST, OCCLUSION TEST, FORCE SENSOR TEST, DISPLACEMENT TEST, AND THE DAT AT 0.0874 INCHES. THE FOLLOWING WERE NOTED DURING VISUAL INSPECTION: SCRATCHED CASE AND PILLOWING KEYPAD OVERLAY. THE SC1 CAP AND RESERVOIR LOCKED PROPERLY INTO RESERVOIR COMPARTMENT DURING TESTING. HISTORY DOWNLOAD WAS SUCCESSFUL USING THUMP AND CARELINK UPLOAD WAS SUCCESSFUL. THE PUMP WAS RECEIVED WITHOUT A BATTERY. THE PUMP WAS RECEIVED WITHOUT THE BATTERY CAP. UNABLE TO VERIFY CUSTOMER'S DAILY TOTAL OF ALL INSULIN DELIVERED, DAILY TOTAL OF BASAL INSULIN DELIVERED AND DAILY TOTAL OF BOLUS INSULIN DELIVERED SURROUNDING THE EVENT DATE OF 09-FEB-2025 LISTED ON SMARTSOLVE DUE TO INSUFFICIENT DATA IN THE TRACE/HISTORY FILES. PERFORM THE HISTORY REVIEW 1 WEEK PRIOR TO THE EVENT DATE 09-FEB-2025 IN THE PUMP HISTORY FILE AND FOUND 2 LOSTSENSOR1ALERT (780) ON (B)(6) 2025 AND A FAILEDBATTTEST (58) ALARM ON (B)(6) 2025 13:31:43.000. THE POWER MANAGEMENT TOOL CONFIRMED THE UNLOADED VOLTAGE (UL VLITH) AND LOADED VOLTAGE (LOADED VLITH) WAS WITHIN SPEC RANGE. FAILED BATTERY TEST ALARM WAS EXPECTED DUE TO THE CUSTOMER'S BATTERY INSERTED ON A BATTERY CHANGE DOES NOT HAVE SUFFICIENT VOLTAGE. THE PUMP PROPERLY PAIRED WITH THE GUARDIAN LINK 3 TRANSMITTER. AFTER THE PUMP COMPLETED WARM UP AND CALIBRATION, THE PUMP WAS ABLE TO DISPLAY THE TEST BG VALUE OF 239 MG/DL ON THE PUMP'S HOME SCREEN. NO COMMUNICATION ANOMALY, CALIBRATION ANOMALY OR LOST SENSOR ALERT NOTED. PUMP WAS CUT OPEN TO PERFORM VISUAL INSPECTION AND FOUND NO EVIDENCE OF PHYSICAL OR MOISTURE DAMAGE ON THE PCBA1, PCBA2, FORCE SENSOR, MOTOR AND VIBRATOR ASSEMBLY NOTED. FORCE SENSOR ZERO OFFSET WITHIN SPECIFICATION (24.1 MV). THE PUMP PASSED THE FUNCTIONAL TESTING. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21591558&pc=NIO | 9681442-2025-00071 | 2025-02-13 05:00 | Death | ANGIOMED GMBH & CO. MEDIZINTECHNIK KG | 2025-03-13 04:00 | NIO | E-LUMINEXX VASCULAR STENT | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P080007 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESULTS OF THE CLINICAL TRIAL THAT APPROXIMATELY THREE YEARS AND EIGHT MONTHS LATER POST A STENT PLACEMENT ON THE LEFT LEG FOR THE DE-NOVO STENOSIS ON THE COMMON AND EXTERNAL ILIAC ARTERY, THE SUBJECT HAD AN ADVERSE EVENT OF DEATH. THE PRIMARY CAUSE OF DEATH WAS NOT PROVIDED. REPORTEDLY, THE ADVERSE EVENT DEATH WAS POSSIBLY RELATED TO THE STUDY DEVICE AND NOT RELATED TO THE PROCEDURE.Event Description: IT WAS REPORTED THROUGH THE RESULTS OF THE CLINICAL TRIAL THAT THREE YEARS AND EIGHT MONTHS POST A STENT PLACEMENT ON THE LEFT LEG FOR THE DE-NOVO STENOSIS ON THE COMMON AND EXTERNAL ILIAC ARTERY, THE SUBJECT HAD AN ADVERSE EVENT OF DEATH. THE PRIMARY CAUSE OF DEATH WAS NOT PROVIDED. REPORTEDLY, THE ADVERSE EVENT DEATH WAS POSSIBLY RELATED TO THE STUDY DEVICE AND NOT RELATED TO THE PROCEDURE. Manufacturer Narrative: H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE E-LUMINEXX VASCULAR STENT PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE E-LUMINEXX VASCULAR STENT PRODUCTS ARE IDENTIFIED IN D2 AND G4. H10: DATE OF DEATH FOR THIS PATIENT WAS NOT PROVIDED, DATE OF DEATH UPDATED AS 01-JAN-1900 FOR THE MDR SUBMISSION REQUIREMENT. H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. Manufacturer Narrative: H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE E-LUMINEXX VASCULAR STENT PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE E-LUMINEXX VASCULAR STENT PRODUCTS ARE IDENTIFIED IN D2 AND G4. H10: DATE OF DEATH FOR THIS PATIENT WAS NOT PROVIDED, DATE OF DEATH UPDATED AS (B)(6) 1900 FOR THE MDR SUBMISSION REQUIREMENT. H10: MANUFACTURING REVIEW: BASED ON THE EVENT INFORMATION PROVIDED THERE IS NO INDICATION THAT A MANUFACTURING PROCESS MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED ISSUE. THEREFORE, A DHR REVIEW IS CONSIDERED TO BE NOT REQUIRED. INVESTIGATION SUMMARY: THE STENT REMAINS IMPLANTED. X-RAY IMAGES WERE NOT PROVIDED FOR EVALUATION. THE EVALUATION OF THE MATERIAL AVAILABLE IS SUBJECT TO THE RETROSPECTIVE ASSESSMENT AS PART OF POST-MARKET CLINICAL FOLLOW-UP (PMCF) ACTIVITIES. BASED ON THE INFORMATION AVAILABLE, THE INVESTIGATION IS CLOSED WITH INCONCLUSIVE RESULT. BASED ON INFORMATION AVAILABLE, THERE IS NO INDICATION THAT THE DEATH OF THE PATIENT MIGHT BE RELATED TO THE PLACED STENT OR THE STENT PLACEMENT PROCEDURE. LABELING REVIEW: RELEVANT LABELING SUPPLIED WITH THIS PRODUCT WAS REVIEWED. POTENTIAL COMPLICATIONS AND ADVERSE EVENTS WHICH MAY OCCUR DURING USE OF THE DEVICE WERE FOUND TO BE REFERENCED IN THE INSTRUCTIONS FOR USE. USERS AND/OR PATIENTS SHOULD REPORT ANY SERIOUS INCIDENT THAT HAS OCCURRED IN RELATION TO THE DEVICE TO THE MANUFACTURER AND TO EITHER THE EUROPEAN UNION NATIONAL COMPETENT AUTHORITY OR TO THE REGULATORY AUTHORITY OF THE COUNTRY IN WHICH THE USER AND/OR PATIENT IS ESTABLISHED. IN THIS CASE THE DEATH OF THE PATIENT WAS DOCUMENTED NOT BEING RELATED TO THE PLACED STENT OR THE STENT PLACEMENT PROCEDURE. H10: G3. H11: H6 (METHOD). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21592247&pc=MHX | 9610816-2025-000197 | 2024-11-27 05:00 | Death | PHILIPS MEDIZIN SYSTEME BÖBLINGEN GMBH | 2025-03-13 04:00 | MHX | PATIENT INFORMATION CENTER IX | Wireless Communication Problem | Asystole | K211900 |  | 1 | Event Description: PHILIPS RECEIVED A COMPLAINT ON A PATIENT INFORMATION CENTER IX INDICATING THAT ALL USERS THAT WERE INTENDED TO GET ALARM NOTIFICATIONS DID NOT GET THESE NOTIFICATIONS. THE DEVICE WAS IN CLINICAL USE AT THE TIME THE ISSUE WAS DISCOVERED. DURING THIS INCIDENT, THE PATIENT PASSED AWAY. Manufacturer Narrative: THIS COMPLAINT IS A SUPPLEMENTAL TO 1218950-2024-000886 DUE TO INCORRECT REGISTRATION NUMBER USED. LOGS WERE ABLE TO BE OBTAINED AND WERE REVIEWED BY THE PHILIPS REMOTE SERVICE ENGINEER (RSE) AND A PHILIPS PRODUCT SUPPORT ENGINEER (PSE). AN ALARM-CHAIN CONFIGURATION REQUIRES ASSIGNMENT OF A ¿LEVEL¿, ¿DEVICE¿, AND ¿USER¿ TO BE SUCCESSFUL. REVIEW OF THE AVAILABLE LOGS REVEALED THAT THE SETTINGS WERE COMPLETE AND ACCURATE FOR USERS ¿SSK A¿ AND ¿SSK B¿ WHO RECEIVED THE MOBILE ALARM VIA THE SYSTEM AS EXPECTED. FOR USER(S) WHO DID NOT RECEIVE THE MOBILE ALARMS, THE CONFIGURATION WAS FOUND TO BE INCOMPLETE. THE AVAILABLE LOGS START AT (B)(6) 2024 AND SHOW THE SAME DEVICE CONFIGURATION UP UNTIL THE EVENT DATE, WHICH DID NOT INCLUDE MOBILE ALARMS CONFIGURED FOR ALL USERS. THUS, IT IS NOT POSSIBLE TO DETERMINE WHEN THIS CONFIGURATION WAS CHOSEN BY THE CUSTOMER. A POSSIBLE DELAY IN THE HEALTHCARE PROFESSIONAL¿S AWARENESS OR TREATMENT OF THE PATIENT¿S CONDITION, WHICH WOULD NOT BE CAPTURED IN EVENT LOGS. AS THE ALARMS OCCURRED CLOSE TOGETHER, IT IS UNCLEAR IF TREATMENT WAS PROVIDED AFTER THE FIRST INSTANCE OF THE ALARM AND BEFORE THE SECOND ALARM OCCURRED. BOX E: REPORTING ADDRESS STATE: (B)(6). REPORTING INSTITUTION PHONE #: +\(B)(6). REPORTER PHONE #: +(B)(6).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21592714&pc=PNJ | 9612164-2025-01325 | 2025-02-20 05:00 | Death | MEDTRONIC IRELAND | 2025-03-13 04:00 | PNJ | MICRA | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT FOUR DAYS POST IMPLANT OF THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) THE PATIENT UNDERWENT A NINE HOUR COLON CANCER RESECTION SURGERY. AFTER THE SURGERY, THE PATIENT'S HEART AND BREATHING STOPPED SUDDENLY. RESCUE WAS INEFFECTIVE AND THE PATIENT PASSED AWAY. Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [MICRA] PRODUCT ID [MC1VR01-DELSYS] (SERIAL: (B)(6)). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21592824&pc=PYX | 1220648-2025-26687 | 2025-02-17 05:00 | Death | ABIOMED, INC. | 2025-03-13 04:00 | PYX | IMPELLA RP FLEX | Restricted Flow rate | Cardiac Perforation | P170011 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA RP FLEX DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. IT WAS NOTED THAT THE PHYSICIAN MADE SEVERAL ATTEMPTS TO ACCESS THE PULMONARY ARTERY (PA), HOWEVER DIFFICULTY WAS NOTED TO BE CAUSED BY VENOUS CANNULA FROM PERFUSION. LEFT PULMONARY ARTERY ACCESS WAS EVENTUALLY OBTAINED AND IMPELLA 0.027 WIRE WAS POSITIONED. AN IMPELLA RP FLEX WAS THEN INSERTED OVER THE WIRE BUT WAS UNABLE TO PASS THE CANNULA. ACCESS TO THE PA CONTINUED TO BE DIFFICULT BUT WAS OBTAINED AND IMPELLA 0.027 WIRE WAS POSITIONED IN RIGHT PA. THE IMPELLA RP FLEX WAS INSERTED AND POSITIONED IN RIGHT PA UNDER FLUOROSCOPY. THE WIRE WAS REMOVED, AND THE IMPELLA WAS SLOWLY STARTED. FLOWS WERE APPROXIMATELY 3.5 LITERS AT P7 FOR APPROXIMATELY FOUR TO FIVE MINUTES BEFORE IT WAS NOTED THAT FLOWS DROPPED TO 1.0 LITER. NO ALARMS WERE NOTED. MOTOR CURRENT WAS NOTED TO BE FLAT AND PLACEMENT SIGNAL WAS SEVERELY DAMPENED. ADDITIONALLY, IT WAS NOTED THAT THE RIGHT VENTRICLE SHOWED A PERFORATION. ATTENTION WAS SWITCHED TO ATTEMPT TO CLOSE PERFORATION. THE IMPELLA FLOW NOTED TO REMAIN AT 1.0 LITER DESPITE INCREASING P LEVEL. THE PHYSICIAN ATTEMPTED TO CLOSE THE PERFORATION BUT ULTIMATELY WAS NOT SUCCESSFUL AND THE PATIENT EXPIRED. ADDITIONAL INFORMATION WAS RECEIVED DURING FOLLOW-UP THAT INDICATED THAT THE PHYSICIAN DOES NOT BLAME THE IMPELLA FOR THE RIGHT VENTRICLE PERFORATION. THE PUMP WAS UNABLE TO BE SENT BACK DUE TO THE PATIENT PASSING. OF NOTE, THERE WAS A SUSPECTED VOLUME ISSUE WITH THE PATIENT RESULTING IN POTENTIAL SUCTION AND SUBOPTIMAL PUMP FLOWS. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE,¿AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA RP FLEX WITH SMARTASSIST SYSTEM WITH AUTOMATED IMPELLA CONTROLLER. SECTION: WARNINGS & CAUTIONS: CAUTIONS: ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ARRHYTHMIA, ATRIAL FIBRILLATION, BLEEDING, CARDIAC TAMPONADE, CARDIOGENIC SHOCK, DEATH, DEVICE MALFUNCTION, HEMOLYSIS, HEPATIC FAILURE, INSERTION SITE INFECTION, PERFORATION, PHLEGMASIA CERULEA DOLENS (A SEVERE FORM OF DEEP VENOUS THROMBOSIS), PULMONARY VALVE INSUFFICIENCY, RESPIRATORY DYSFUNCTION, SEPSIS, THROMBOCYTOPENIA, THROMBOTIC VASCULAR (NON-CENTRAL NERVOUS SYSTEM COMPLICATION, TRICUSPID VALVE INJURY, CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION), VENOUS THROMBOSIS, VENTRICULAR FIBRILLATION AND/OR TACHYCARDIA.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿DO NOT ADVANCE OR WITHDRAW THE IMPELLA RP FLEX WITH SMARTASSIST SYSTEM CATHETER AGAINST RESISTANCE WITHOUT USING FLUOROSCOPY TO DETERMINE THE CAUSE OF THE RESISTANCE. DOING SO COULD RESULT IN SEPARATION OF THE CATHETER OR GUIDEWIRE TIP, DAMAGE TO THE CATHETER OR VESSEL, OR PERFORATION.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21593768&pc=BTI | 2518422-2025-033792 | 2023-08-14 04:00 | Death | RESPIRONICS, INC. | 2025-03-13 04:00 | BTI | ESSENCE | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K042655 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH AN ESSENCE DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE DEVICE IS MISSING AND UNAVAILABLE FOR RETURN TO THE MANUFACTURER FOR INVESTIGATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21593861&pc=FTM | MW5167617 |  | Death | DAVOL INC., SUB. C.R. BARD, INC. | 2025-03-13 04:00 | FTM | XENMATRIX SURGICAL GRAFT |  | Abdominal Pain; Adhesion(s); Hernia; Obstruction/Occlusion; Fluid Discharge |  |  | 1 | Event Description: ALLEGED EVENT: LIMITED INFORMATION WAS REPORTED THROUGH A LEGAL EVENT THAT A PATIENT WAS IMPLANTED WITH 1ST STRATTICE ON (B)(6) 2016 AND 2ND STRATTICE ON (B)(6) 2017. THE FORM ALSO INDICATES THAT ON AN UNSPECIFIED DATE, THE STRATTICE IMPLANT HAD BEEN ¿REMOVED OR REVISED¿. NO INFORMATION WAS PROVIDED ON THE INDICATION FOR THE INITIAL SURGERY OR THE REASON FOR THE REMOVAL OR REVISION. PATIENT HAS DIED ON (B)(6) 2022 AND THERE IS NO FURTHER DETAILS ON DEATH CAUSALITY. THIS RECORD CAPTURED 1ST STRATTICE. ON (B)(6) 2025, PMQA RECEIVED NOTIFICATION FROM LEGAL THAT THE PLAINTIFF PROFILE FORM WAS RECEIVED ASSOCIATED WITH THIS EVENT. AS PER THE PPF FORM, THE PATIENT WAS BORN WITH AN ABDOMINAL WALL DEFECT (GASTROSCHISIS), WHOM UNDERWENT A ¿VENTRAL¿ HERNIA SURGERY AND WAS IMPLANTED WITH A STRATTICE DEVICE LOT ¿SP100381-101¿ ON (B)(6) 2016. THE PATIENT UNDERWENT AN ¿OPEN VENTRAL INCISIONAL HERNIA REPAIR WITH BIOLOGICAL MESH¿ ON (B)(6) 2017. THE PATIENT WAS IMPLANTED WITH A SECOND STRATTICE DEVICE LOT ¿SP100537-078¿ ON THE SAME DATE. THE PATIENT THEN UNDERWENT AN ¿INCISIONAL/VENTRAL HERNIA REPAIR WITH BIOLOGIC MESH¿ ON (B)(6) 2018 AND IMPLANTED WITH A THIRD STRATTICE DEVICE LOT ¿SP100595." THE PATIENT THEN UNDERWENT AN ¿OPEN VENTRAL INCISIONAL HERNIA REPAIR WITH MESH¿ ON (B)(6) 2018 AND IMPLANTED WITH A NON- ABBVIE DEVICE, ¿XENMATRIX LOT HUCR0748.¿ THE MESH WAS REVISED DUE TO ¿RECURRENT VENTRAL HERNIA REPAIR WITH BIOLOGIC MESH¿ ON (B)(6) 2018 AND IMPLANTED WAS A SECOND NON-ABBVIE DEVICE, ¿BARD MESH LOT #HUCQ0919.¿ THE MESH WAS REVISED ON (B)(6) 2022 DUE TO ¿OPEN REPAIR OF RECURRENT VENTRAL INCISIONAL HERNIA WITH BIOLOGICAL MESH, ABDOMINAL RE-ENTRY FOR CORRECTION OF RE-HERNIATION¿ AND IMPLANTED WITH A FOURTH STRATTICE DEVICE LOT ¿ SP200312-045." THE FORM LISTS THE CONDITIONS THAT WERE TREATED WERE ¿OPEN VENTRAL INCISIONAL HERNIA REPAIR WITH BIOLOGIC STRATTICE MESH¿ ON OR ABOUT (B)(6) 2016. THE PATIENT ALSO UNDERWENT A ¿OPEN VENTRAL INCISIONAL HERNIA REPAIR WITH BIOLOGIC MESH (STRATTICE)¿ ON OR ABOUT (B)(6) 2017. THE PATIENT WAS TREATED FOR ¿INCISIONAL/ VENTRAL HERNIA REPAIR WITH BIOLOGIC MESH (STRATTICE)¿ BETWEEN (B)(6) 2018. THE PATIENT UNDERWENT ¿OPEN VENTRAL INCISIONAL HERNIA REPAIR WITH MESH (RECURRENT); AT BEDSIDE SEROUS SANGUINOUS OOZING AROUND LEFT JP DRAIN¿ BETWEEN (B)(6) 2018 AND (B)(6) 2018. THE PATIENT WAS TREATED FOR ¿RECURRENT VENTRAL HERNIA REPAIR WITH BIOLOGIC MESH¿ ON OR ABOUT (B)(6) 2018. THE PATIENT WAS SEEN FOR ¿ABDOMINAL PAIN¿ BETWEEN 2018 AND (B)(6) 2022. THE PATIENT WAS TREATED FOR ¿BOWEL OBSTRUCTION¿ BETWEEN APPROXIMATELY 2017-2018. THE PATIENT RECEIVED TREATMENT FOR ¿ABDOMINAL PAIN/ADHESIONS¿ BETWEEN (B)(6) 2016 AND (B)(6) 2022. THE PATIENT WAS SEEN FOR ¿BOWEL OBSTRUCTION, ABDOMINAL PAIN, ADHESIONS¿ BETWEEN (B)(6) 2016 AND (B)(6) 2022. THE PATIENT WAS SEEN FOR ¿HERNIA SYMPTOMS¿ FROM APPROXIMATELY 2018 TO (B)(6) 2022. THE PATIENT WAS TREATED ¿OPEN REPAIR OF RECURRENT INCISIONAL HERNIA WITH BIOLOGIC MESH (STRATTICE); ABDOMINAL RE-ENTRY FOR CORRECTION OF RE-HERNIATION¿ BETWEEN (B)(6) 2022 AND (B)(6) 2022. THE PATIENT RECEIVED ¿POST-OP REHAB¿ BETWEEN 2018-2019. "THIS REPORT REFLECTS INFORMATION RECEIVED BY FDA IN THE FORM OF A NOTIFICATION PER 803.22 (B)(2)". REFERENCE REPORT: MW5167618.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21593940&pc=FTL | MW5167618 |  | Death | DAVOL INC., SUB. C.R. BARD, INC. | 2025-03-13 04:00 | FTL | BARD MESH |  | Abdominal Pain; Adhesion(s); Hernia; Obstruction/Occlusion; Fluid Discharge |  |  | 1 | Event Description: ALLEGED EVENT: LIMITED INFORMATION WAS REPORTED THROUGH A LEGAL EVENT THAT A PATIENT WAS IMPLANTED WITH 1ST STRATTICE ON (B)(6) 2016 AND 2ND STRATTICE ON (B)(6) 2017. THE FORM ALSO INDICATES THAT ON AN UNSPECIFIED DATE, THE STRATTICE IMPLANT HAD BEEN ¿REMOVED OR REVISED¿. NO INFORMATION WAS PROVIDED ON THE INDICATION FOR THE INITIAL SURGERY OR THE REASON FOR THE REMOVAL OR REVISION. PATIENT HAS DIED ON (B)(6) 2022 AND THERE IS NO FURTHER DETAILS ON DEATH CAUSALITY. THIS RECORD CAPTURED 1ST STRATTICE. ON (B)(6) 2025, PMQA RECEIVED NOTIFICATION FROM LEGAL THAT THE PLAINTIFF PROFILE FORM WAS RECEIVED ASSOCIATED WITH THIS EVENT. AS PER THE PPF FORM, THE PATIENT WAS BORN WITH AN ABDOMINAL WALL DEFECT (GASTROSCHISIS), WHOM UNDERWENT A ¿VENTRAL¿ HERNIA SURGERY AND WAS IMPLANTED WITH A STRATTICE DEVICE LOT ¿SP100381-101¿ ON (B)(6) 2016. THE PATIENT UNDERWENT AN ¿OPEN VENTRAL INCISIONAL HERNIA REPAIR WITH BIOLOGICAL MESH¿ ON (B)(6) 2017. THE PATIENT WAS IMPLANTED WITH A SECOND STRATTICE DEVICE LOT ¿SP100537-078¿ ON THE SAME DATE. THE PATIENT THEN UNDERWENT AN ¿INCISIONAL/VENTRAL HERNIA REPAIR WITH BIOLOGIC MESH¿ ON (B)(6) 2018 AND IMPLANTED WITH A THIRD STRATTICE DEVICE LOT ¿SP100595." THE PATIENT THEN UNDERWENT AN ¿OPEN VENTRAL INCISIONAL HERNIA REPAIR WITH MESH¿ ON (B)(6) 2018 AND IMPLANTED WITH A NON- ABBVIE DEVICE, ¿XENMATRIX LOT HUCR0748.¿ THE MESH WAS REVISED DUE TO ¿RECURRENT VENTRAL HERNIA REPAIR WITH BIOLOGIC MESH¿ ON (B)(6) 2018 AND IMPLANTED WAS A SECOND NON-ABBVIE DEVICE, ¿BARD MESH LOT #HUCQ0919.¿ THE MESH WAS REVISED ON (B)(6) 2022 DUE TO ¿OPEN REPAIR OF RECURRENT VENTRAL INCISIONAL HERNIA WITH BIOLOGICAL MESH, ABDOMINAL RE-ENTRY FOR CORRECTION OF RE-HERNIATION¿ AND IMPLANTED WITH A FOURTH STRATTICE DEVICE LOT ¿ SP200312-045." THE FORM LISTS THE CONDITIONS THAT WERE TREATED WERE ¿OPEN VENTRAL INCISIONAL HERNIA REPAIR WITH BIOLOGIC STRATTICE MESH¿ ON OR ABOUT (B)(6) 2016. THE PATIENT ALSO UNDERWENT A ¿OPEN VENTRAL INCISIONAL HERNIA REPAIR WITH BIOLOGIC MESH (STRATTICE)¿ ON OR ABOUT (B)(6) 2017. THE PATIENT WAS TREATED FOR ¿INCISIONAL/ VENTRAL HERNIA REPAIR WITH BIOLOGIC MESH (STRATTICE)¿ BETWEEN (B)(6) 2018. THE PATIENT UNDERWENT ¿OPEN VENTRAL INCISIONAL HERNIA REPAIR WITH MESH (RECURRENT); AT BEDSIDE SEROUS SANGUINOUS OOZING AROUND LEFT JP DRAIN¿ BETWEEN (B)(6) 2018. THE PATIENT WAS TREATED FOR ¿RECURRENT VENTRAL HERNIA REPAIR WITH BIOLOGIC MESH¿ ON OR ABOUT (B)(6) 2018. THE PATIENT WAS SEEN FOR ¿ABDOMINAL PAIN¿ BETWEEN 2018 AND (B)(6) 2022. THE PATIENT WAS TREATED FOR ¿BOWEL OBSTRUCTION¿ BETWEEN APPROXIMATELY 2017-2018. THE PATIENT RECEIVED TREATMENT FOR ¿ABDOMINAL PAIN/ADHESIONS¿ BETWEEN (B)(6) 2016 AND (B)(6) 2022. THE PATIENT WAS SEEN FOR ¿BOWEL OBSTRUCTION, ABDOMINAL PAIN, ADHESIONS¿ BETWEEN (B)(6) 2016 AND (B)(6) 2022. THE PATIENT WAS SEEN FOR ¿HERNIA SYMPTOMS¿ FROM APPROXIMATELY 2018 TO (B)(6) 2022. THE PATIENT WAS TREATED ¿OPEN REPAIR OF RECURRENT INCISIONAL HERNIA WITH BIOLOGIC MESH (STRATTICE); ABDOMINAL RE-ENTRY FOR CORRECTION OF RE-HERNIATION¿ BETWEEN (B)(6) 2022. THE PATIENT RECEIVED ¿POST-OP REHAB¿ BETWEEN 2018-2019. "THIS REPORT REFLECTS INFORMATION RECEIVED BY FDA IN THE FORM OF A NOTIFICATION PER 803.22 (B)(2)". REFERENCE REPORT: MW5167617.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21593943&pc=LWQ | 1649833-2025-00006 | 2024-08-21 04:00 | Death | ON-X LIFE TECHNOLOGIES, INC. | 2025-03-13 04:00 | LWQ | ON-X ASCENDING AORTIC PROSTHESIS WITH THE VASCUTEK GELWEAVE VALSALVA GRAFT (ONXA | Adverse Event Without Identified Device or Use Problem | Endocarditis; Hemorrhage/Bleeding; Embolism/Embolus; Thrombosis/Thrombus | P000037 |  | 1 | Event Description: ACCORDING TO THE INITIAL NOTIFICATION ON 02/28/2025, " THIS PUBLICATION ¿DECELLULARIZED AORTIC HOMOGRAFT¿S VERSUS MECHANICAL COMPOSITE GRAFTS FOR AORTIC ROOT REPLACEMENT¿ BY ANDREEVA, ET ALL PUBLISHED ON 21AUG 2024 IN THE EUROPEAN JOURNAL OF CARDIO-THORACIC SURGERY WAS DISCOVERED DURING THE MOST RECENT ON-X AAP EU MDR CYCLE. THIS STUDY INCLUDED 289 PATIENTS UNDERGOING AAR (AORTIC ROOT REPLACEMENT) IN AN ELECTIVE SETTING WITH DAH (DECELLULARIZED AORTIC HOMOGRAFT) OR MCVG (MECHANICAL COMPOSITE VALVE GRAFT) BETWEEN 2000 AND 2022. DATA WAS COLLECTED RETROSPECTIVELY FROM THE INSTITUTIONAL HOSPITAL DATABASE AND PROSPECTIVELY THROUGH OUTPATIENT VISITS AND PHONE CALLS. ALL-CAUSE MORTALITY WAS THE PRIMARY ENDPOINT AND A COMPOSITE END POINT OF SVD, NSVD (NON- STRUCTURAL VALVE DETERIORATION), BLEEDING, THROMBOTIC AND EMBOLIC EVENTS WAS CREATED FOR ASSESSMENT OF EVENT-FREE SURVIVAL. THE FOLLOWING CONDUITS WERE USED AS MCVG: SJM REGENT (ABBOTT, ABBOTT PARK, IL, USA) N = 107, 49.5%, CARBOMEDICS CARBOSEAL (LIVANOVA, SALUGGIA, ITALY) N = 83, 38.4%, ON-X (ARTIVION, KENNESAW, GA, USA) N =24, 11.1%, MEDTRONIC OPEN PIVOT (MEDTRONIC, MINNEAPOLIS, MN, USA) N = 2, 0.9%. FOUR DEATHS WERE REPORTED FOR 30-DAY MORTALITY THEY WERE FULMINANT SEPSIS, MASSIVE INTRACRANIAL HEMORRHAGE, MASSIVE PERICARDIAL EFFUSION AND MULTIORGAN FAILURE WITH MYOCARDIAL INFARCTION. OTHER EVENTS REPORTED WERE BLEEDING, THROMBOSIS, EMBOLIZATION, ENDOCARDITIS, STRUCTURAL VALVE DETERIORATION, PERMANENT PACEMAKER IMPLANTATION AND RE-INTERVENTION. THE CONCLUSION PRESENTED WAS ¿AORTIC ROOT REPLACEMENT WITH EITHER A MCVG OR A DECELLULARIZED AORTIC HOMOGRAFTS SHOWED VERY LOW MORTALITY IN AN ELECTIVE SETTING. PATIENTS AFTER DAH IMPLANTATION DEMONSTRATED SIGNIFICANTLY HIGHER FREEDOM FROM VALVE-RELATED ADVERSE EVENTS DURING MID- TERM FOLLOW-UP. DAH PRESENT A PROMISING TREATMENT OPTION FOR YOUNG PATIENTS REQUIRING ARR; HOWEVER, FURTHER DATA ON LONG- TERM DURABILITY ARE NEEDED¿." Manufacturer Narrative: ACCORDING TO THE INITIAL NOTIFICATION ON 02/28/2025, " THIS PUBLICATION ¿DECELLULARIZED AORTIC HOMOGRAFT¿S VERSUS MECHANICAL COMPOSITE GRAFTS FOR AORTIC ROOT REPLACEMENT¿ BY ANDREEVA, ET ALL PUBLISHED ON 21AUG 2024 IN THE EUROPEAN JOURNAL OF CARDIO-THORACIC SURGERY WAS DISCOVERED DURING THE MOST RECENT ON-X AAP EU MDR CYCLE. THIS STUDY INCLUDED 289 PATIENTS UNDERGOING AAR (AORTIC ROOT REPLACEMENT) IN AN ELECTIVE SETTING WITH DAH (DECELLULARIZED AORTIC HOMOGRAFT) OR MCVG (MECHANICAL COMPOSITE VALVE GRAFT) BETWEEN 2000 AND 2022. DATA WAS COLLECTED RETROSPECTIVELY FROM THE INSTITUTIONAL HOSPITAL DATABASE AND PROSPECTIVELY THROUGH OUTPATIENT VISITS AND PHONE CALLS. ALL-CAUSE MORTALITY WAS THE PRIMARY ENDPOINT AND A COMPOSITE END POINT OF SVD, NSVD (NON- STRUCTURAL VALVE DETERIORATION), BLEEDING, THROMBOTIC AND EMBOLIC EVENTS WAS CREATED FOR ASSESSMENT OF EVENT-FREE SURVIVAL. THE FOLLOWING CONDUITS WERE USED AS MCVG: SJM REGENT (ABBOTT, ABBOTT PARK, IL, USA) N = 107, 49.5%, CARBOMEDICS CARBOSEAL (LIVANOVA, SALUGGIA, ITALY) N = 83, 38.4%, ON-X (ARTIVION, KENNESAW, GA, USA) N =24, 11.1%, MEDTRONIC OPEN PIVOT (MEDTRONIC, MINNEAPOLIS, MN, USA) N = 2, 0.9%. FOUR DEATHS WERE REPORTED FOR 30-DAY MORTALITY THEY WERE FULMINANT SEPSIS, MASSIVE INTRACRANIAL HEMORRHAGE, MASSIVE PERICARDIAL EFFUSION AND MULTIORGAN FAILURE WITH MYOCARDIAL INFARCTION. OTHER EVENTS REPORTED WERE BLEEDING, THROMBOSIS, EMBOLIZATION, ENDOCARDITIS, STRUCTURAL VALVE DETERIORATION, PERMANENT PACEMAKER IMPLANTATION AND RE-INTERVENTION. THE CONCLUSION PRESENTED WAS ¿AORTIC ROOT REPLACEMENT WITH EITHER A MCVG OR A DECELLULARIZED AORTIC HOMOGRAFTS SHOWED VERY LOW MORTALITY IN AN ELECTIVE SETTING. PATIENTS AFTER DAH IMPLANTATION DEMONSTRATED SIGNIFICANTLY HIGHER FREEDOM FROM VALVE-RELATED ADVERSE EVENTS DURING MID- TERM FOLLOW-UP. DAH PRESENT A PROMISING TREATMENT OPTION FOR YOUNG PATIENTS REQUIRING ARR; HOWEVER, FURTHER DATA ON LONG- TERM DURABILITY ARE NEEDED¿." MULTIPLE ATTEMPTS FOR ADDITIONAL INFORMATION HAVE GONE UNMET. NO ADDITIONAL INFORMATION FORTHCOMING. THE MANUFACTURING RECORDS COULD NOT BE REVIEWED AS PATIENT INFORMATION, SERIAL NUMBERS, NOR DATES OF IMPLANT WERE PROVIDED. THE INFORMATION WAS REVIEWED. THIS INVESTIGATION REVIEWS THE CLINICAL OUTCOMES PRESENTED IN THE 2024 PUBLICATION ¿DECELLULARIZED AORTIC HOMOGRAFTS VERSUS MECHANICAL COMPOSITE GRAFTS FOR AORTIC ROOT REPLACEMENT¿ BY ANDREEVA ET AL. THE STUDY IS A SINGLE-CENTER RETROSPECTIVE COHORT ANALYSIS CONDUCTED AT THE MEDICAL UNIVERSITY OF VIENNA, ENCOMPASSING 289 ADULT PATIENTS WHO UNDERWENT ELECTIVE AORTIC ROOT REPLACEMENT (ARR) BETWEEN 2000 AND 2022. PATIENTS INCLUDED IN THE STUDY RECEIVED EITHER A MECHANICAL COMPOSITE VALVE GRAFT (MCVG) OR A DECELLULARIZED AORTIC HOMOGRAFT (DAH), WHILE THOSE UNDERGOING URGENT OR EMERGENT PROCEDURES WERE EXCLUDED. OF THE 289 PATIENTS, 216 RECEIVED AN MCVG AND 73 RECEIVED A DAH. THE MEAN AGE WAS 48.5 ± 11.5 YEARS, WITH MCVG PATIENTS AVERAGING 49 YEARS AND DAH PATIENTS 47 YEARS. THE MAJORITY OF PATIENTS WERE MALE (79.6%), WITH 83.3% IN THE MCVG GROUP AND 68.4% IN THE DAH GROUP. AMONG THOSE TREATED WITH MCVG, THE FOLLOWING CONDUITS WERE USED: SJM REGENT (49.5%, N=107), CARBOMEDICS CARBOSEAL (38.4%, N=83), ON-X (11.1%, N=24), AND MEDTRONIC OPEN PIVOT (0.9%, N=2). THE STUDY REPORTED A 30-DAY MORTALITY RATE OF 1% (N=3), ALL OCCURRING IN THE MCVG GROUP. OPERATIVE MORTALITY WAS 1.4% (N=4), ALSO CONFINED TO THE MCVG GROUP. ONE-YEAR SURVIVAL WAS 100% IN THE DAH GROUP AND 95.8% IN THE MCVG GROUP, WHILE THREE-YEAR SURVIVAL RATES WERE 98.6% AND 94.4%, RESPECTIVELY. IN TERMS OF VALVE-RELATED COMPLICATIONS IN THE MCVG COHORT, BLEEDING EVENTS WERE OBSERVED IN 14 PATIENTS (6.5%) AT 1 YEAR AND 2.6% AT 3 YEARS. THESE INCLUDED THREE CASES OF CEREBRAL BLEEDING (1.4%) AND 11 PERIPHERAL (5.1%). VALVE THROMBOSIS OCCURRED IN TWO PATIENTS (0.9%) AT BOTH 1 AND 3 YEARS. EMBOLIZATION EVENTS WERE REPORTED IN 14 PATIENTS (6.5%) AT 1 YEAR AND 4.4% AT 3 YEARS, OF WHICH 12 (5.6%) WERE CEREBRAL EMBOLISMS. THESE LED TO STROKE IN 10 CASES (4.6%) AND TRANSIENT ISCHEMIC ATTACK IN TWO CASES (0.9%). PERMANENT PACEMAKER IMPLANTATION WAS REQUIRED IN 15 PATIENTS (6.9%) AT 1 YEAR AND 4.7% AT 3 YEARS. AORTIC VALVE ENDOCARDITIS WAS OBSERVED IN 10 PATIENTS (4.6%), WITH AN INCIDENCE OF 1.2% AT 1 YEAR AND 1.9% AT 3 YEARS; SIX PATIENTS (2.8%) REQUIRED REOPERATION. NOTABLY, NO CASES OF STRUCTURAL VALVE DETERIORATION (SVD) WERE REPORTED IN THE MCVG GROUP. REINTERVENTION WAS NECESSARY IN 12 MCVG PATIENTS (5.6%) AT 1 YEAR AND 3.1% AT 3 YEARS, WITH NO REINTERVENTIONS REPORTED IN THE DAH GROUP. SEVEN OF THESE REINTERVENTIONS (3.2%) WERE VALVE-RELATED, INCLUDING SIX CASES OF ENDOCARDITIS AND ONE DUE TO VALVE THROMBOSIS. FOUR REINTERVENTIONS (1.9%) WERE RELATED TO THE AORTA AND INCLUDED PROCEDURES SUCH AS AORTIC ARCH REPLACEMENT, THORACIC ENDOVASCULAR AORTIC REPAIR (TEVAR) FOR TYPE B DISSECTIONS, AND RE-ARR FOR PSEUDOANEURYSM. ONE REINTERVENTION (0.4%) WAS PERFORMED FOR POSTOPERATIVE LEFT MAIN DISSECTION, WHICH REQUIRED CORONARY ARTERY BYPASS GRAFTING. ACCORDING TO THE STUDY AUTHORS, ELECTIVE ARR USING EITHER MCVG OR DAH WAS ASSOCIATED WITH LOW EARLY MORTALITY. HOWEVER, PATIENTS IN THE DAH GROUP DEMONSTRATED SIGNIFICANTLY GREATER FREEDOM FROM VALVE RELATED ADVERSE EVENTS DURING MID-TERM FOLLOW-UP. THE AUTHORS SUGGEST THAT DAH REPRESENTS A PROMISING TREATMENT OPTION FOR YOUNG PATIENTS REQUIRING ARR, THOUGH FURTHER DATA ARE NEEDED TO ASSESS LONG-TERM DURABILITY. ALL REPORTED COMPLICATIONS AND MORTALITY EVENTS ARE CONSISTENT WITH THE KNOWN RISKS OUTLINED IN THE ON-X® VALVE¿S INSTRUCTIONS FOR USE (IFU). THIS CLINICAL OUTCOME ANALYSIS IS BASED ON A SINGLE-CENTER EXPERIENCE INVOLVING 289 PATIENTS WHO UNDERWENT ELECTIVE AORTIC ROOT REPLACEMENT (ARR) USING EITHER A MECHANICAL COMPOSITE VALVE GRAFT (MCVG) OR A DECELLULARIZED AORTIC HOMOGRAFT (DAF). HOWEVER, BASED ON THE INFORMATION PROVIDED IN THE PUBLICATION, IT IS NOT POSSIBLE TO IDENTIFY WHICH PATIENTS RECEIVED THE ON-X AORTIC AAP IMPLANT VERSUS A MCVG FROM ANOTHER MANUFACTURER. CONSEQUENTLY, IT IS NOT FEASIBLE TO ATTRIBUTE SPECIFIC ADVERSE EVENTS OR MORTALITY OUTCOMES TO THE AAP DEVICE. THEREFORE, WE CANNOT ASSESS THE AAP DEVICE¿S CONTRIBUTION, IF ANY, TO THE REPORTED EVENTS. IT IS IMPORTANT TO NOTE THAT ALL REPORTED COMPLICATIONS AND MORTALITY OUTCOMES ARE IDENTIFIED IN THE ON-X AAP INSTRUCTIONS FOR USE (IFU) AS KNOWN RISKS. NO FURTHER ACTION IS REQUIRED WITHOUT FURTHER INFORMATION. BASED ON THE INFORMATION PROVIDED IN THE PUBLICATION, IT IS NOT POSSIBLE TO IDENTIFY WHICH PATIENTS RECEIVED THE ON-X AORTIC AAP IMPLANT VERSUS A MCVG FROM ANOTHER MANUFACTURER. CONSEQUENTLY, IT IS NOT FEASIBLE TO ATTRIBUTE SPECIFIC ADVERSE EVENTS OR MORTALITY OUTCOMES TO THE AAP DEVICE. THEREFORE, WE CANNOT ASSESS THE AAP DEVICE¿S CONTRIBUTION, IF ANY, TO THE REPORTED EVENTS. THERE IS NO INDICATION THAT AN ERROR OR DEFICIENCY OCCURRED AT ARTIVION AND THE IFU ADEQUATELY COMMUNICATES RISK. ARTIVION WILL CONTINUE TO MONITOR SIMILAR COMPLAINTS TO DETERMINE IF ADDITIONAL ACTIONS ARE WARRANTED; HOWEVER, AT THIS TIME NO FURTHER ACTIONS ARE NECESSARY. THIS REPORT IS BEING SUBMITTED AS REQUIRED BY FEDERAL REGULATIONS AND DOES NOT CONSTITUTE AN ADMISSION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. FURTHERMORE, THIS REPORT REFLECTS THE EVENT AS ALLEGED BY THE COMPLAINANT AND DOES NOT IMPLY THAT THE INFORMATION REPORTED TO ARTIVION IS ACCURATE OR HAS BEEN CONFIRMED BY ARTIVION. Manufacturer Narrative: THIS INVESTIGATION IS CURRENTLY ONGOING. ANY ADDITIONAL INFORMATION WILL BE PROVIDED IN THE FOLLOW-UP REPORT. THIS REPORT IS BEING SUBMITTED AS REQUIRED BY FEDERAL REGULATIONS AND DOES NOT CONSTITUTE AN ADMISSION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. FURTHERMORE, THIS REPORT REFLECTS THE EVENT AS ALLEGED BY THE COMPLAINANT AND DOES NOT IMPLY THAT THE INFORMATION REPORTED TO ARTIVION IS ACCURATE OR HAS BEEN CONFIRMED BY ARTIVION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21594435&pc=DQY | 2124215-2025-15324 | 2025-02-03 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-03-13 04:00 | DQY | WATCHMAN TRUSEAL ACCESS SYSTEM | Adverse Event Without Identified Device or Use Problem | Air Embolism; Cardiac Arrest; Stroke/CVA; Low Blood Pressure/ Hypotension |  |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT DEATH OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED. A WATCHMAN TRUSEAL ACCESS SYSTEM (WAS) AND A WATCHMAN FLX LAA CLOSURE DEVICE AND DELIVERY SYSTEM (WDS) WERE SELECTED TO BE USED. THE PHYSICIAN SUCCESSFULLY PERFORMED THE TRANSSEPTAL PUNCTURE AND THEN INSERTED THE WAS INTO THE LEFT ATRIUM OF THE PATIENT. WHILE ADVANCING THE PIGTAIL CATHETER THROUGH THE WAS, THE PHYSICIAN NOTICED THAT THERE WAS AIR PRESENT IN THE WAS. THE AIR ENTERED THE PATIENT LEADING TO THE PATIENT GOING INTO CARDIAC ARREST AND EXPERIENCING A CEREBRAL VASCULAR ACCIDENT. THE PATIENT RECEIVED CARDIOPULMONARY RESUSCITATION LEADING TO RETURN OF SPONTANEOUS CIRCULATION AFTER THREE MINUTES. THE PATIENT WAS GIVEN NORADRENALINE FOR HYPOTENSION. A COMPUTED TOMOGRAPHY (CT) ANGIOGRAM SHOWED THE CORONARY ARTERIES WITHOUT STENOSIS OR AIR PRESENCE. THE PATIENT WAS INTUBATED AND RECEIVED MECHANICAL VENTILATION DUE TO PERSISTENT CONSCIOUSNESS LOSS BEFORE THEY WERE TRANSFERRED TO CRANIAL CT TO RULE OUT A CEREBROVASCULAR EVENT. THE PATIENT DID NOT RECOVER FROM THIS EVENT AND DIED. THE OFFICIAL CAUSE OF DEATH WAS CONSIDERED MASSIVE AIR EMBOLISM LEADING TO MASSIVE STROKE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21594443&pc=CBK | 2518422-2025-033805 | 2025-01-03 05:00 | Death | RESPIRONICS, INC. | 2025-03-13 04:00 | CBK | TRILOGY EVO | Device Alarm System | No Clinical Signs, Symptoms or Conditions | K181166 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT PASSED AWAY IN A HOME ENVIRONMENT. THE INFORMATION INDICATES THE VENTILATOR WAS FAILED TO ALARM. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: AFTER RECEIVING FURTHER INFORMATION. SECTION ADVERSE EVENT/PRODUCT PROBLEM IN BOX B: ADVERSE EVENT OR PRODUCT PROBLEM HAS BEEN UPDATED/CORRECTED IN THIS REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21595063&pc=MOZ | MW5167631 |  | Death | IBSA FARMACEUTICI ITALIA SRL | 2025-03-13 04:00 | MOZ | GELSYN-3 |  | Insufficient Information |  |  | 1 | Event Description: PATIENT PASSED AWAY ON (B)(6) 2025.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21595097&pc=DSP | MW5167632 | 2025-03-05 05:00 | Death | MAQUET /DATASCOPE CORP. | 2025-03-13 04:00 | DSP | IABP BALLOON |  | Insufficient Information |  |  | 1 | Event Description: SUSPECTED TEAR IN THE IABP(INTRA-AORTIC BALLOON PUMP) BALLOON ASSOCIATED WITH PATIENT DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21595213&pc=NGV | 2124215-2025-15495 | 2024-11-10 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-03-13 04:00 | NGV | WATCHMAN LAA CLOSURE DEVICE & DELIVERY SYSTEM | Difficult to Open or Close | Stroke/CVA; Hemorrhage/Bleeding; Transient Ischemic Attack; Embolism/Embolus; Thrombosis/Thrombus | P130013 |  | 1 | Event Description: REPORTED VIA JOURNAL ARTICLE: IT WAS REPORTED THAT PATIENT DEATH(S) OCCURRED. THIS WAS A MULTICENTER, PROSPECTIVE COHORT STUDY THAT INCLUDED 3,096 PATIENTS FROM 39 CENTERS IN CHINA, CONDUCTED BETWEEN APRIL 1, 2019, AND OCTOBER 31, 2020. CONSECUTIVE PATIENTS FROM EACH PARTICIPATING CENTER WHO RECEIVED THE LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE AND WERE OF LEGAL AGE TO PROVIDE INFORMED CONSENT WERE RECRUITED IN THIS REGISTRY. PATIENTS WHO WERE PARTICIPATING IN OTHER TRIALS, DECLINED TO PROVIDE INFORMED CONSENT, OR REFUSED TO PARTICIPATE WERE EXCLUDED. PERI- AND INTRAPROCEDURAL TECHNIQUES AND POSTPROCEDURAL MEDICATIONS WERE LEFT TO THE OPERATOR'S DISCRETION. A TOTAL OF 159 OPERATORS, WITH VARYING LEVELS OF EXPERIENCE IMPLANTING THE DEVICE, PARTICIPATED IN THE STUDY. RESULTS: AT 1 YEAR, THE COMPOSITE ENDPOINT OF DEATH, STROKE, AND SYSTEMIC EMBOLISM OCCURRED IN 133 PATIENTS, CONSISTING OF 68 DEATHS, 52 ISCHEMIC STROKES, 24 HEMORRHAGIC STROKES, AND 4 SYSTEMIC EMBOLISMS. DEVICE RELATED THROMBUS OCCURRED IN 45 PATIENTS. 10 PATIENTS EXPERIENCED A TRANSIENT ISCHEMIC ATTACK. 10 PATIENTS HAD A PERIDEVICE LEAK OF GREATER THAN 5MM. 39 PATIENTS WERE REPORTED TO EXPERIENCE A MAJOR BLEEDING EVENT, AND 47 PATIENTS WERE REPORTED TO HAVE EXPERIENCED A MINOR BLEEDING EVENT. Manufacturer Narrative: B2 DATE OF DEATH - ARTICLE PUBLISH DATE USED AS DEATH DATE IS UNKNOWN. B3 DATE OF EVENT - ARTICLE PUBLISH DATE USED AS EVENT DATE IS UNKNOWN. D3 MODEL NUMBER, D3 LOT NUMBER, D4 EXPIRATION DATE, D4 UNIQUE IDENTIFIER (UDI) # ARE UNKNOWN AS THE INFORMATION WAS NOT PROVIDED IN THE LITERATURE ARTICLE. D6A IMPLANT DATE WAS REPORTED AS (B)(6) 2019, THROUGH (B)(6) 2020. LITERATURE CITATION: GAO, C., SU, F., LIU, J., ZHANG, T., NING, Z., YANG, B., ... & RECORD INVESTIGATORS. (2024). 1-YEAR CLINICAL OUTCOMES AND THE IMPACT OF PROCEDURAL CONFIGURATIONS IN LEFT ATRIAL APPENDAGE OCCLUSION PATIENTS. JACC: ASIA, 4(10), 777-790. DOI.ORG/10.1016/J.JACASI.2024.07.013.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21595242&pc=OZD | 1220648-2025-26702 | 2025-02-17 05:00 | Death | ABIOMED, INC. | 2025-03-13 04:00 | OZD | IMPELLA 5.5 | Increase in Pressure | Cardiovascular Insufficiency | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. DURING SUPPORT THERE WAS HIGH PURGE PRESSURE. TISSUE PLASMA ACTIVATOR WAS ADDED TO THE PURGE. SUPPORT CONTINUED. SURVIVAL OUTCOME AT EXPLANT NOTED THE PATIENT EXPIRED. THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA DEVICE AS AN ASSOCIATED FACTOR IN THE PATIENT'S DEMISE. Manufacturer Narrative: THE INVESTIGATION OF HIGH PURGE PRESSURE HAS BEEN COMPLETED. THE PUMP WAS NOT RETURNED FOR INVESTIGATION. THE DATA LOG INDICATES A GRADUAL INCREASE IN PURGE PRESSURE OVER FOUR DAYS, ACCOMPANIED BY A HIGH ALARM THAT WAS RESOLVED BY A DECREASING TREND OVER THE FOLLOWING TWO DAYS. ADDITIONALLY, ON THE EXPLANT DAY, A PURGE SYSTEM BLOCK ALARM WAS OBSERVED, WITH A THREE-MINUTE RISE IN PURGE PRESSURE AND A SMALL SPIKE IN MOTOR CURRENT. THIS TREND ALSO RESOLVED WITH A GRADUAL DECREASE OVER THE COURSE OF TEN HOURS. THE CLINICAL TEAM ALSO MENTIONED THAT THE TISSUE PLASMINOGEN ACTIVATOR (TPA) WAS ADMINISTERED WHILE A HIGH PURGE PRESSURE ALARM WAS OBSERVED, ALONG WITH MONITORING OF THE MOTOR CURRENT TREND. THE CAUSE OF THE HIGH PURGE PRESSURE ISSUE WAS MOST LIKELY BIOMATERIAL BUILDUP, BASED ON DATA LOG REVIEW. E4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26702 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA IFU: IMPELLA 5.5 WITH SMARTASSIST: SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA CATHETER: ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP. CLINICIANS SHOULD MONITOR MOTOR CURRENT AND CONSIDER REPLACING THE IMPELLA CATHETER WHENEVER A RISE IN MOTOR CURRENT IS SEEN.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21595716&pc=CBK | 1220908-2025-00888 | 2025-02-07 05:00 | Death | ZOLL MEDICAL CORPORATION | 2025-03-13 04:00 | CBK | Z VENT | Communication or Transmission Problem | Insufficient Information | K162832 |  | 1 | Event Description: COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A PATIENT (AGE AND GENDER UNKNOWN), THE DEVICE DISPLAYED "INTERNAL COMM FAILURE- 1172"AND "INTERNAL COMM FAILURE- 1174" ERROR MESSAGES. COMPLAINANT INDICATED THAT THE CLINICIAN BAG VENTILATED TO CONTINUE TREATING THE PATIENT. COMPLAINANT INDICATED THAT THE PATIENT SUBSEQUENTLY EXPIRED. Manufacturer Narrative: THE DEVICE WAS RETURNED TO ZOLL MEDICAL CORPORATION FOR EVALUATION. THE CUSTOMER'S REPORT WAS OBSERVED DURING REVIEW OF THE DEVICE DATA LOGS. LOGS SUGGEST THE EVENT OCCURRED ON 02/07/2025. ERRORS INCLUDED LOW O2 SUPPLY FAULTS (2020), AND PATIENT DISCONNECTS (2100), POTENTIALLY INFLUENCED BY PATIENT COUGHING OR ASYNCHRONOUS BREATHING. THE DEVICE PASSED STRESS TESTING AND INTERNAL INSPECTION WITH NO INTERNAL COMMUNICATION TYPE FAULTS. THE DEVICE WAS RECERTIFIED AND RETURNED TO THE CUSTOMER. ANALYSIS OF REPORTS OF THIS TYPE HAS NOT IDENTIFIED AN INCREASE IN TREND. Manufacturer Narrative: ZOLL MEDICAL CORPORATION HAS RECEIVED THE PRODUCT AND WILL BE PROVIDING A SUPPLEMENTAL REPORT WHEN OUR INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21595903&pc=DSQ | 2916596-2025-01542 | 2025-02-24 05:00 | Death | THORATEC CORPORATION | 2025-03-13 04:00 | DSQ | HEARTMATE II LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Mechanical Problem; Pumping Stopped; Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure; No Clinical Signs, Symptoms or Conditions | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED TO THE CLINIC DESCRIBING ALARMS AND A PUMP STOP DURING THE PREVIOUS DAY. THE LOG FILE INDICATED A SHORT TO SHIELD (STS). A PUMP EXCHANGE WAS RECOMMENDED. THE ALARMS INCLUDED LOW FLOW, NO EXTERNAL POWER, AND PUMP OFF.Event Description: IT WAS NOTED THAT THE PATIENT EXPERIENCED SYMPTOMS/ADVERSE EVENTS RELATED TO THE PUMP STOPS. NO X-RAYS WERE TAKEN OF THE DRIVELINE. THE PATIENT HAD MORE EPISODES OF PUMP STOPS AND THE PATIENT, PATIENTS' FAMILY, AND HEALTH CARE PROVIDERS DECIDED THAT THE PUMP WOULD NOT BE EXCHANGED. DURING THE FOURTH EPISODE OF PUMP STOPS, THE PUMP HAD NOT RESTARTED AND WAS DISCONNECTED AFTER 10 MINUTES. THE PATIENT SUBSEQUENTLY PASSED AWAY ON (B)(6) 2025. THE CAUSE OF DEATH WAS DUE TO HEART FAILURE. THE OUTCOME WAS CONSIDERED DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: SECTION B1: TYPE OF REPORT CORRECTED. SECTION H1: TYPE OF REPORTABLE EVENT CORRECTED. SECTION H6: HEALTH EFFECT - IMPACT CODE CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: REVIEW OF THE SUBMITTED LOG FILE CONFIRMED THE REPORTED PUMP STOPS WHICH WERE CONSISTENT WITH SUSPECTED DRIVELINE DAMAGE. THE SUBMITTED LOG FILE CONTAINED EVENTS SPANNING FROM (B)(6) 2025 THROUGH (B)(6) 2025, PER THE TIMESTAMPS. PUMP STOP AND LOW SPEED EVENTS WERE CAPTURED ON (B)(6) 2025 WHILE THE SYSTEM WAS CONNECTED TO BATTERY POWER AND POWER MODULE. BASED ON PREVIOUS COMPLAINT HISTORY, THE ABNORMAL PUMP OPERATION APPEARED TO BE CONSISTENT WITH POTENTIAL DRIVELINE WIRE COMPROMISE. NO OTHER NOTABLE EVENTS OR ALARMS WERE CAPTURED. THE PUMP WAS NOT EXPLANTED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) AS WELL AS DRIVELINE, SERIAL NUMBER (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), AND PATIENT HANDBOOK, ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTIONS 6 AND 8 OF THE HMII IFU, AS WELL AS SECTIONS 4 AND 6 OF THE HMII PATIENT HANDBOOK, PROVIDE INFORMATION REGARDING HOW TO CARE FOR THE DRIVELINE. THESE DOCUMENTS INSTRUCT THE USER TO CHECK THEIR DRIVELINE DAILY FOR SIGNS OF DAMAGE, SUCH AS CUTS, HOLES, OR TEARS. ADDITIONALLY, THESE SECTIONS OUTLINE INDICATIONS OF DRIVELINE DAMAGE AS WELL AS THE HOW TO RESPOND TO SUCH EVENTS, AND ADDRESS DAMAGE DUE TO WEAR AND FATIGUE OF THE DRIVELINE. THE PATIENT HANDBOOK ALSO INSTRUCTS THE USER TO CALL THEIR HOSPITAL CONTACT RIGHT AWAY IF THE DRIVELINE IS DAMAGED (OR MIGHT BE DAMAGED). SECTION 5 OF THE PATIENT HANDBOOK, ¿ALARMS AND TROUBLESHOOTING¿, AND SECTION 7 OF THE IFU, ¿ALARMS AND TROUBLESHOOTING¿, PROVIDES INFORMATION ON SYSTEM ALARM CONDITIONS AS WELL AS THE APPROPRIATE ACTIONS ASSOCIATED WITH EACH CONDITION. THE PATIENT HANDBOOK ALSO CONTAINS A SECTION ON HANDLING EMERGENCIES AND FURTHER INSTRUCTS THE USER TO CALL THEIR HOSPITAL CONTACT IF THE PATIENT THINKS THAT, FOR ANY REASON, ANY PORTION OF THEIR EQUIPMENT IS NOT FUNCTIONING AS USUAL, IS BROKEN, OR THEY ARE UNCOMFORTABLE WITH THE OPERATION OF THE EQUIPMENT. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21596152&pc=DSQ | 2916596-2025-01611 | 2025-02-20 05:00 | Death | THORATEC CORPORATION | 2025-03-13 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Diarrhea; Encephalopathy; Hypoxia; Vomiting; Coma; Loss of consciousness; Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED EMESIS AND DIARRHEA WHILE AT HOME AND WAS EMERGENTLY TRANSPORTED DUE TO EXTRINSIC OUTFLOW GRAFT OBSTRUCTION (EOGO). EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) WAS INSERTED. THE PATIENT EXPERIENCED A LOSS OF CONSCIOUSNESS WHILE ADMITTED AND HAD LOW FLOW ALARMS DURING THE EVENT. THE PATIENT WAS TAKEN BACK TO THE OPERATING ROOM AND EOGO WAS REMOVED SURGICALLY AND DETERMINED TO BE THE CAUSE OF THE EVENT. FOLLOWING THE PROCEDURE, THE PATIENTS FLOW WAS RECOVERING AND ECMO WAS BEING WEANED, BUT THE PATIENT REMAINED UNCONSCIOUS AT THE TIME. FLOW RECOVERED AND ECMO WAS WITHDRAWN, BUT THE PATIENT REMAINED UNCONSCIOUS. ON (B)(6) 2025 THE PATIENT PASSED AWAY DUE TO HYPOXIC ENCEPHALOPATHY AND THE PUMP WAS REMOVED.Event Description: IT WAS ADDITIONALLY REPORTED THAT THE PATIENT WAS EMERGENTLY TRANSPORTED AND ADMITTED ON (B)(6) 2025. IT WAS NOTED THAT DURING AN OUTPATIENT VISIT ON (B)(6) 2025 THERE WERE NO PARTICULAR PROBLEMS. THE ARTIFICIAL BLOOD VESSEL OF THE VENTRICULAR ASSIST DEVICE (VAD) WAS BLOCKED DUE TO DEHYDRATION CAUSED BY CLOSTRIDIUM DIFFICILE (C. DIFF) ENTERITIS, LEADING TO A COLLAPSE OF CIRCULATORY DYNAMICS. SINCE THE LOW FLOW ALARM HAD BEEN SOUNDING SINCE THE TIME THE PATIENT WAS AT HOME, IT WAS BELIEVED THAT SUFFICIENT BLOOD FLOW WAS NOT ENSURED THROUGHOUT THE BODY AT AN EARLY STAGE. DURING THE EMERGENCY TRANSPORT, THE PATIENT'S LEVEL OF CONSCIOUSNESS DECREASED, AND THEY STOPPED BREATHING. THERE WERE SIGNS OF HYPOXIC ENCEPHALOPATHY, ACCOMPANIED BY MULTIPLE ORGAN FAILURE, THE CAUSE OF DEATH WAS RELATED TO EQUIPMENT FAILURE AND CLOSTRIDIUM DIFFICILE (C. DIFF) ENTERITIS DEHYDRATION.Event Description: A CAUSAL RELATIONSHIP OF THE EVENT TO THE DEVICE WAS REPORTEDLY PRESENT. IT WAS NOTED THAT EOGO WAS PREVIOUSLY SUSPECTED AND WAS BEING FOLLOWED ANNUALLY. DUE TO DEHYDRATION AND OTHER FACTORS, THE GRAFT BECOME OCCLUDED, RESULTING IN THE LOW FLOW. AS OF 2024, THE STENOSIS WAS AT 50%. IT WAS BELIEVED THAT WHEN THE LOW FLOW OCCURRED DUE TO DEHYDRATION, THE PRESSURE BALANCE BETWEEN THE BIO-DEBRIS HAD ACCUMULATED INSIDE THE BEND RELIEF AND THE PRESSURE WITHIN THE GRAFT WAS DISRUPTED, LEADING TO OCCLUSION. Manufacturer Narrative: H6: HEALTH EFFECT - CLINICAL CODE 2144 - VOMITING. NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21596261&pc=NPT | 2025587-2025-01874 | 2024-07-02 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-13 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Heart Failure/Congestive Heart Failure | P130021 |  | 1 | Event Description: I WAS REPORTED THAT APPROXIMATELY 4 MONTHS AND 10 DAYS AFTER THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT WAS HOSPITALIZED FOR HEART FAILURE. APPROXIMATELY 1 MONTH LATER, THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS NOTED TO BE WORSENING HEART FAILURE. Manufacturer Narrative: UPDATED DATA: D4. H4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21596291&pc=DSQ | 2916596-2025-01577 | 2024-10-11 04:00 | Death | THORATEC CORPORATION | 2025-03-13 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE DEVICE WAS NOT RETURNED FOR EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE EVENT; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED BY THE CUSTOMER. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21596408&pc=QZI | 2124215-2025-15483 | 2025-02-17 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-03-13 04:00 | QZI | FARAWAVE | Off-Label Use | Bradycardia; Cardiac Arrest |  |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING A PULSE FIELD ABLATION (PFA) PROCEDURE USING A FARAWAVE PFA CATHETER THE PATIENT PASSED AWAY. DURING THE PROCEDURE PULMONARY VEIN (PVA) AND POSTERIOR WALL (PW) ABLATIONS WERE PERFORMED, THE POSTERIOR WALL ABLATIONS ARE CONSIDERED OFF LABEL USE AS THE FARAWAVE IS ONLY INDICATED FOR TREATMENT OF THE PULMONARY VEINS FOR PAROXYSMAL ATRIAL FIBRILLATION. ABLATIONS TO THE CRISTA TERMINALIS WERE PERFORMED WITH A NON-BOSTON SCIENTIFIC CATHETER. IT WAS NOTED THAT THEY HAD DIFFICULTY ADVANCING A NON-BOSTON SCIENTIFIC INTRACARDIAC ECHOCARDIOGRAPHY (ICE) CATHETER THROUGH THE PATIENT'S LEFT FEMORAL VASCULATURE BUT WERE ABLE TO INTRODUCE IT INTO THE RIGHT ATRIUM TO ASSIST WITH THE TRANSSEPTAL PUNCTURE. AFTER THE TRANSSEPTAL PUNCTURE ABLATIONS THE PVI AND PW ABLATIONS WERE PERFORMED WITH THE FARAWAVE CATHETER, EIGHT ABLATIONS WERE APPLIED TO EACH PULMONARY VEIN WITH ANO ADDITIONAL 4 ABLATIONS TO THE POSTERIOR WALL. IT WAS NOTED THAT THE PATIENT HAD A SMALLER THAN EXPECTED LEFT ATRIUM, BUT THESE ABLATIONS WERE COMPLETED WITH NO OBVIOUS COMPLICATIONS AT THE TIME. THEN A NON-BOSTON SCIENTIFIC LASSO MAPPING CATHETER WAS USED TO PERFORM THE POST TREATMENT MAPPING. APPROXIMATELY 5 MINUTES INTO THIS 'POST MAP' "THE PATIENT BEGAN TO LOSE PACING IN THE CS CATHETER AND ITS DR PACEMAKER DID NOT WORK AS WELL". THEY OBSERVED A "WENCKEBAH" BRADYCARDIA PATTERN AND EVENTUALLY ASYSTOLE. CPR WAS STARTED AT THIS TIME AND THEN SHOCKED HER AFTER "4 CYCLES IN VF", WHICH RESTORED HER PULSE AND HEART'S CONTRACTION. APPROXIMATELY AN HOUR LATER THE BRADYCARDIA LEADING TO ASYSTOLE HAPPENED AGAIN SO CPR AND THEN ANOTHER SHOCK WERE DELIVERED BUT THIS TIME HER PULSE DID NOT RETURN. A BLOOD GAS ANALYSIS WAS PERFORMED AND THE PATIENT'S PH RETURNED AS "A LITTLE ACID". THE ANESTHESIOLOGIST "BEGAN TO FIX IT" AND THE NEXT BLOOD TEST RETURNED NORMAL. AFTER APPROXIMATELY 25 MINUTES OF CPR THE PHYSICIAN'S DECLARED THE PATIENT AS DECEASED. POST HUMOROUS FLUOROSCOPIC IMAGING FROM THE PATIENT'S LEGS TO HER LUNGS WERE PERFORMED BUT NO ABNORMALITIES WERE FOUND. THERE WAS NO PERFORATION OR OTHER HEART DAMAGE OBSERVED SO CARDIAC TAMPONADE WAS RULED OUT. PULMONARY EDEMA WAS RULED OUT DUE TO A LACK OF FLUID IN THE LUNGS. NO TRAUMA TO THE VASCULATURE TRAVERSED BY THE SHEATHS WAS PRESENT. IN THE PHYSICIAN'S OPINION "THEY HAD NOTHING TO SUSPECT THAT THE FARAPULSE SINGLE USE DEVICES COULD BE RELATED TO THE ABRUPT DECLINE IN HER GLASGOW SCALE". ULTIMATELY THE CAUSE OF THE CARDIAC ARREST WAS NOT DETERMINED AT THIS TIME. THE CATHETER WAS DISPOSED BY THE FACILITY AND WILL NOT BE RETURNED AS THE DEVICE PERFORMED AS EXPECTED DURING THE PROCEDURE AND THERE WERE NO REPORTS OF ANY PERFORMANCE CONCERNS. Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21596611&pc=NPT | 2015691-2025-01872 | 2025-02-19 05:00 | Death | EDWARDS LIFESCIENCES | 2025-03-13 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Low Blood Pressure/ Hypotension; Obstruction/Occlusion | P140031 |  | 1 | Event Description: EDWARDS RECEIVED NOTIFICATION FROM OUR AFFILIATE IN CANADA. AS REPORTED, THIS WAS A CASE OF AN IMPLANT OF A 26MM SAPIEN 3 ULTRA RESILIA IN AORTIC POSITION BY TRANSFEMORAL APPROACH. THE PATIENT HAD NATIVE BICUSPID VALVE WITH MODERATE ANNULAR AND LVOT CALCIUM. THE VALVE WAS DEPLOYED WITH MINUS 2ML OF VOLUME. AFTER DEPLOYMENT, AN ANGIOGRAM SHOWED GOOD POSITIONING AND FLOW TO BOTH LEFT AND RIGHT CORONARY ARTERIES. AFTER THAT, PATIENT DEVELOPED HYPOTENSION PROGRESSING TO PULSELESS ELECTRICAL ACTIVITY (PEA). IT WAS DECIDED TO START RESUSCITATION AND LEFT MAIN, LEFT CIRCUMFLEX ARTERY AND LEFT ANTERIOR DESCENDING ARTERY WERE STENTED. THE RESUSCITATION CONTINUED FOR FORTY-FIVE MINUTES WITH INTUBATION FOR AIRWAY MANAGEMENT AND PHARMACOLOGICAL SUPPORT. THE PATIENT REMAINED IN PEA, AND HAD NOT RETURN OF SPONTANEOUS CIRCULATION OR CHANGE IN RHYTHM. DUE TO THE LENGTH OF CPR IT WAS DECIDED TO STOP RESUSCITATION AND UNFORTUNATELY THE PATIENT PASSED AWAY. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), CORONARY FLOW OBSTRUCTION IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT PROCEDURE. THE IFU CAUTIONS THAT SAFETY AND EFFECTIVENESS HAVE NOT BEEN ESTABLISHED FOR PATIENTS WITH BULKY CALCIFIED AORTIC VALVE LEAFLETS NEAR CORONARY OSTIA. IN ADDITION, IT WARNS THAT CAUTION SHOULD BE EXERCISED IN IMPLANTING A BIOPROSTHESIS IN PATIENTS WITH CLINICALLY SIGNIFICANT CORONARY ARTERY DISEASE. CORONARY OBSTRUCTION CAN RESULT IN MYOCARDIAL ISCHEMIA OR INFARCTION DUE TO OBSTRUCTION OF THE CORONARY BLOOD FLOW AND MAY REQUIRE INTERVENTION (E.G., PERCUTANEOUS CORONARY INTERVENTION (PCI)). MULTIPLE PATIENT FACTORS COULD PUT THE PATIENT AT RISK FOR CORONARY FLOW OBSTRUCTION DURING THE THV PROCEDURE, INCLUDING SIGNIFICANT UNDERLYING CORONARY ARTERY DISEASE AND BULKY CALCIFICATION OF THE NATIVE LEAFLETS AND ROOT. DISPLACEMENT OF CALCIUM DEPOSITS WITH EMBOLIZATION OF DEBRIS INTO ONE OF THE ARTERIES, OR AORTIC DISSECTION WITH CONTINUITY OF THE RUPTURE INTO THE INTIMA OF ONE OF THE CORONARY OSTIA, CAN RESULT IN THIS COMPLICATION. ADDITIONALLY, MINIMAL DISTANCE BETWEEN THE NATIVE ANNULUS AND THE CORONARY OSTIA, BULKY CALCIFICATION, LONG NATIVE LEAFLETS, AND OBLITERATED CORONARY SINUSES. PROCEDURAL FACTORS SUCH AS TORN NATIVE LEAFLET DURING BALLOON VALVULOPLASTY, PLAQUE SHIFT, DEPLOYMENT OF THE BIOPROSTHETIC HEART VALVE TOO AORTIC, AND SIGNIFICANT VALVE OVER SIZING COULD ALSO CONTRIBUTE TO THIS COMPLICATION. THE EDWARDS THV TRAINING MANUALS ADVISE THE OPERATOR ON PRE-PROCEDURE ASSESSMENT OF THE AORTIC VALVE, ROOT, AND CORONARY ANATOMY. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE TRANSCATHETER HEART VALVE (THV). TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES. THE PHYSICIAN IS INSTRUCTED TO EVALUATE THIS RISK EARLY IN THE PATIENT SCREENING PROCESS IN ALL PATIENTS THE FOLLOWING FACTORS SHOULD BE CONSIDERED: DEGREE OF CALCIFICATION ON LEAFLETS, ANNULUS TO CORONARY OSTIA DISTANCE, LENGTH OF THE VALVE LEAFLET, WIDTH OF THE VALSALVA SINUSES, MOVEMENT OF THE LEAFLETS DURING BAV, PATENCY OF CORONARIES DURING BAV, AND EXPANDED HEIGHT OF THE INTENDED THV. THE TRAINING MANUALS ALSO PROVIDE THE FOLLOWING TIPS FOR DETECTING RISK FOR LEFT MAIN OCCLUSION: (1) AORTOGRAM OR TEE BEFORE THV IMPLANTATION TO REVEAL BULKY CALCIFIED LEAFLETS; (2) DURING PRE-DILATATION, NOTE BULKY CALCIFICATION ON VALVE MOVING TOWARDS OSTIUM ON LEFT MAIN; AND (3) CONSIDER AORTOGRAM DURING VALVULOPLASTY TO ASSESS CORONARY FLOW. THE EDWARDS THV MANUALS ALSO ADVISE THE OPERATOR ON PRE-PROCEDURE ASSESSMENT OF THE AORTIC VALVE, ROOT, AND CORONARY ANATOMY. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE TRANSCATHETER HEART VALVE (THV). TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST PATIENT RELATED FACTORS (BICUSPID VALVE WITH MODERATE ANNULAR AND LVOT CALCIUM) CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21596653&pc=DSQ | 3007042319-2025-00652 | 2025-03-04 05:00 | Death | HEARTWARE, INC. | 2025-03-13 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP/DRIVELINE | Pumping Problem | Aneurysm; Atrial Fibrillation; Bacterial Infection; Fever; Hematoma; Unspecified Infection; Paresis; Respiratory Failure; Unspecified Nervous System Problem | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WITH A VENTRICULAR ASSIST DEVICE EXPERIENCED A DRIVELINE INFECTION, LOW-GRADE FEVER, AND RESPIRATORY SYMPTOMS. THE INFECTION WAS ASSOCIATED WITH THE EXTERNAL DRIVELINE AND THE PUMP RELATED TO THE VENTRICULAR ASSIST DEVICE. THE PATIENT RECEIVED INTRAVENOUS ANTIBIOTICS AS TREATMENT, AND A RESPIRATORY PANEL WAS CONDUCTED, WHICH RETURNED NEGATIVE RESULTS. THE PATIENT HAD BEEN ON CHRONIC SUPPRESSIVE DUAL ANTIBIOTICS THERAPY SINCE (B)(6) 2024. THE POWER CONSUMPTION OF THE VENTRICULAR ASSIST DEVICE WAS SLIGHTLY HIGHER THAN THE NORMAL OPERATING RANGE, PROMPTING MONITORING FOR A POSSIBLE THROMBUS. THE PATIENT WAS ADMITTED TO THE ER FOR WORSENING DRIVELINE INFECTION AND STARTED ON INTRAVENOUS ANTIBIOTICS.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT EXPERIENCED A CHANGE IN MENTAL STATUS. COMPUTERIZED TOMOGRAPHY (CT) REVEALED ACUTE PARENCHYMAL HEMATOMA WITH ASSOCIATED LEFT SIDE HEMIPARESIS. PROGRESSIVE CTS SHOWED WORSENING HEMATOMA. IT WAS ALSO FOUND THAT THE PATIENT HAD SERRATIA BACTEREMIA INSIDE OF CHEST WALL AND LEFT BREAST ABSCESSES. THE PATIENT HAD POSITIVE BLOOD AND WOUND CULTURES. THE PATIENT WAS TREATED FOR INFECTION AND HAD STABILIZATION OF BLOOD PRESSURE WITH PRESSORS. IT WAS ALSO REPORTED THAT THE PATIENT EXPERIENCED ATRIAL FIBRILLATION WITH RAPID VENTRICULAR RESPONSE REQUIRING AMIODARONE. NO THROMBUS WAS SEEN. IT WAS SUSPECTED THAT THE PATIENT WAS LIKELY TO HAVE EXPERIENCED A MYCOTIC ANEURYSM AS THE INFECTION WAS QUITE SEVERE AT TIME OF PRESENTATION WITH HIGH WHITE BLOOD CELL COUNTS. CARE WAS WITHDRAWN AND THE PATIENT SUBSEQUENTLY PASSED AWAY. NO AUTOPSY WILL BE PERFORMED AND THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IMPLANTED IN THE PATIENT. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED AS ADDITIONAL INFORMATION HAS BEING RECEIVED FOR THIS EVENT AND SECTIONS HAVE BEEN UPDATED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR ANALYSIS AND INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: THE VENTRICULAR ASSIST DEVICE (VAD) (B)(6) WAS NOT RETURNED FOR EVALUATION AS IT REMAINS IMPLANTED. A REVIEW OF THE MANUFACTURING DOCUMENTATION WAS NOT PERFORMED AS THE REPORTED EVENT AND ANALYSIS ARE NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER EXPERIENCING DRIVELINE INFECTION, FEVER, ACUTE PARENCHYMAL HEMATOMA, ATRIAL FIBRILLATION AND A MYCOTIC ANEURYSM. LOG FILE ANALYSIS REVEALED INTERMITTENT INCREASES IN POWER CONSUMPTION AND ESTIMATED FLOWS TO PARAMETERS ABOVE NORMAL OPERATING RANGE THROUGHOUT THE ANALYZED PERIOD; HOWEVER, NO ALARMS WERE LOGGED WITHIN THE ANALYZED PERIOD. AS A RESULT, THE REPORTED EVENT WAS CONFIRMED. BASED ON THE INFORMATION AVAILABLE, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. BASED ON THE RISK DOCUMENTATION, POSSIBLE CAUSES OF THE HIGH POWER MAY BE ATTRIBUTED TO MULTIPLE FACTORS INCLUDING BUT NOT LIMITED TO EXTERNAL FACTORS SUCH AS THROMBUS FORMATION/INGESTION, INAPPROPRIATE PUMP ROTATIONAL SPEED, AND/OR PATIENT RELATED FACTORS. PER THE INSTRUCTIONS FOR USE, RESPIRATORY DYSFUNCTION, INFECTION, NEUROLOGICAL DYSFUNCTION, CARDIAC ARRHYTHMIA, AND DEATH ARE KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. BASED ON REVIEW OF PAST ADVERSE EVENTS FOR THIS PATIENT, IT WAS NOTED THAT THE PATIENT HAD A HISTORY OF INFECTION AND CARDIAC ARRHYTHMIA. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE, INCLUDING CARE OF THE DRIVELINE EXIT SITE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR ADDITIONAL INFORMATION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21596907&pc=HGM | 9610816-2025-000208 | 2024-12-27 05:00 | Death | PHILIPS MEDIZIN SYSTEME BÖBLINGEN GMBH | 2025-03-13 04:00 | HGM | AVALON FM30 FETAL MONITOR | Incorrect Measurement | Insufficient Information | K140535 |  | 1 | Event Description: IT WAS REPORTED THE USER HAD DIFFICULTY DIFFERENTIATING BETWEEN THE MATERNAL HEART RATE (MHR) AND FETAL HEART RATE (FHR) AND THE BABY WAS STILLBORN. THE 37 WEEK MOTHER PRESENTED TO THE LABOR WARD IN SEVERE ABDOMINAL PAIN WITH A BASE MHR OF 160 BPM. A DOPPLER WAS PERFORMED TO DETERMINE FETAL LIFE PRIOR TO CONNECTING THE MOTHER TO FETAL MONITOR. THE STAFF HAD DIFFICULTY DETERMINING WHICH RATE WAS THE MHR AND FHR, WHICH COINCIDENCE ALARMS BEING GENERATED BY THE MONITOR. THE BABY WAS STILLBORN; HOWEVER, THE FETAL MOVEMENT PROFILE (FMP) WAS STILL PRESENT. OVERALL, THE STAFF WAS CONFUSED AS TO FUNCTIONALITY OF THE MONITOR AND REQUESTED ASSISTANCE FOR EDUCATION AND EXPLANATION OF THE BEHAVIOR. ADDITIONAL INFORMATION RECEIVED INDICATED THE CAUSE OF DEATH WAS NOT KNOWN AS NO AUTOPSY WAS PERFORMED. THE FHR ALARMS WERE SOUNDING FROM THE BEGINNING OF THE CASE ONWARD BUT THE MOTHER WAS NOT CONNECTED TO SPO2 OR HR MONITORING DURING THE BIRTH. THE BABY WAS BORN VIA C-SECTION AND THE MIDWIVES WERE UNSURE WHETHER CONFUSION OVER THE ALARMS CONTRIBUTED TO THE OUTCOME. IT WAS ALSO INDICATED THERE WAS NO DEVICE FAULT. FURTHER INFORMATION RECEIVED INDICATED THE FHR AND MHR WERE TOO CLOSE, WITH THE MHR BASELINE BEING 160BPM. THE COINCIDENCE CAME FROM TOCO AS BOTH MHR AND FHR CAN BE MONITORED VIA THIS DEVICE. THE BABY WAS DELIVERED VIA C-SECTION. THE CLINICAL APPLICATION SPECIALIST INDICATED THERE WAS NO SIGN OF MALFUNCTION. NO ADDITIONAL INFORMATION CLARIFYING THE EVENT WAS RECEIVED. BASED ON THIS INFORMATION, IT REMAINS UNCLEAR WHETHER THE DEVICE CAUSED OR CONTRIBUTED TO THE REPORTED DEATH. Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED WHICH INDICATED ON ADMISSION AT 0245, DOPPLER WAS USED ALONG WITH A TOCO TRANSDUCER, INDICATING THE MHR WAS 142 BPM AND FHR WAS 130BPM. AN ULTRASOUND WAS NOT PERFORMED BUT MHR WAS MEASURED USING A SPOT CHECK MONITOR, CONFIRMING THE HEART RATE OF 142BPM. AT 0500, THE FHR WAS 135BPM AND MHR WAS 138BPM. AT 0630, FHR WAS 145BPM AND MHR WAS 138BPM. WHEN THE PATIENT COMPLAINED OF ABDOMINAL AND SHOULD PAIN, THE PHYSICIAN SUSPECTED PLACENTAL ABRUPTION, SO A C-SECTION WAS PERFORMED. AT 0720, PRIOR TO THE C-SECTION, THE FHR COULD NOT BE DETECTED WITH DOPPLER. A PHILIPS FIELD SERVICE ENGINEER (FSE) EVALUATED THE DEVICE AND COULD NOT CONFIRM THE REPORTED, THE EVENT DATA WAS REVIEWED WHICH SHOWED COINCIDENCE ALARM WERE GOING OFF AND ACKNOWLEDGED. THERE IS NO MACHINE FAULT AS IT WAS WORKING AS INTENDED, THE DEVICE IS BACK IN USE IN THE IN THE LABOR AND DELIVERY SUITE. A PHILIPS CLINICAL SCIENTIST REVIEWED THE INFORMATION PROVIDED AND CONCLUDED THAT THE CTG-RECORDING RAN FOR APPROXIMATELY 3 HOURS AND 20 MINUTES. DURING APPROXIMATELY 2 HOURS AND 20 MINUTES, COINCIDENCE ALARMING¿S WERE ACTIVE AND ACKNOWLEDGED BY THE STAFF. THIS MEANS, THE FHR1 US TRANSDUCER AND TOCO MP TRANSDUCER WERE CONTINUOUSLY REGISTERING AND DISPLAYING THE SIGNAL OF THE MATERNAL PULSE. ACCORDING TO THE INFORMATION PROVIDED BY THE HOSPITAL, A COUNTERCHECK OF THE MATERNAL PULSE BY THE SPO2 FINGER SENSOR OR THE APPLICATION OF MECG HAS NOT BEEN PERFORMED. THIS WAS CONFIRMED BY THE RECORDINGS OF THE FETAL MONITOR. DURING THE PERIOD OF CCV ALARMING¿S AND SIGNAL LOSS ALARMING¿S, AN OBSTETRIC ULTRASONOGRAPHY OF THE FETAL CONDITION HAS NOT BEEN PERFORMED ACCORDING TO THE INFORMATION PROVIDED FROM THE HOSPITAL. THE DEVICE WAS CONFIRMED TO BE OPERATING PER SPECIFICATIONS AND NO FAILURE WAS IDENTIFIED. BASED ON THIS INFORMATION, THE DEVICE DID NOT MALFUNCTION AND THUS, DOES NOT APPEAR TO HAVE CAUSED OR CONTRIBUTED TO THE REPORTED DEATH. THE CAUSE OF DEATH REMAINS UNKNOWN. THE INVESTIGATION CONCLUDES THAT NO FURTHER ACTION IS REQUIRED AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED THE COMPLAINT FILE WILL BE REOPENED. THE INSTRUCTIONS FOR USE (IFU) WARN THAT A MISIDENTIFICATION OF THE MHR AS THE FHR CAN OCCUR. THE FHR DETECTION BY THE MONITOR MAY NOT ALWAYS INDICATE THAT THE FETUS IS ALIVE. THE IFU STATE THE IMPORTANCE TO CONFIRM FETAL LIFE BEFORE MONITORING, AND TO CONTINUE TO CONFIRM THAT THE FETUS IS THE SIGNAL SOURCE FOR THE RECORDED HEART RATE. TO ASSIST THE USER AND RECOGNIZING POSSIBLE MISIDENTIFICATION OF THE MHR INSTEAD OF THE FHR, THE FETAL MONITOR GENERATES A COINCIDENCE INOP WITH A PRINTED QUESTION MARK ON THE TRACE. FURTHER INFORMATION CAN BE READ IN THE CHAPTER ABOUT ¿CONFIRM FETAL LIFE BEFORE USING THE MONITOR¿ AND ¿CROSS-CHANNEL VERIFICATION (CCV)¿. IF A COINCIDENCE OF HEART RATES IS DETECTED, PHILIPS RECOMMENDS FOLLOWING THE APPLICABLE SECTION OF THE INSTRUCTIONS FOR USE. 1. CONFIRM FETAL LIFE BY PALPATION OF FETAL MOVEMENT OR AUSCULTATION OF FETAL HEART SOUNDS USING A FETOSCOPE, STETHOSCOPE, OR PINARD STETHOSCOPE. 2. MANUALLY DETERMINE THE MATERNAL PULSE AND COMPARE WITH THE FETAL HEART RATE SOUND SIGNALS. 3. REPOSITION THE TRANSDUCER OR ENSURE THAT THE FETAL SCALP ELECTRODE IS PLACED CORRECTLY, UNTIL YOU RECEIVE A CLEAR SIGNAL AND THE MONITOR IS NO LONGER ISSUING THE ¿COINCIDENCE INOP¿. 4. IN CASE OF DIFFICULTIES DERIVING A STABLE MATERNAL PULSE READING USING THE TOCO MP OR CL TOCO MP TRANSDUCER, USE THE SPO2 FINGER SENSOR INSTEAD. IN CASE OF SIMILAR PROBLEMS WITH THE PULSE MEASUREMENT FROM SPO2 FINGER SENSOR, USE MECG INSTEAD. REASONS TO SWITCH THE METHOD FOR DERIVING A MATERNAL PULSE OR HEART RATE INCLUDE MOTION ARTIFACTS, ARRHYTHMIA, AND INDIVIDUAL DIFFERENCES IN PULSE SIGNAL QUALITY ON THE ABDOMINAL SKIN (VIA TOCO MP). 5. IF YOU CANNOT HEAR THE FETAL HEART SOUNDS, AND YOU CANNOT CONFIRM FETAL MOVEMENT BY PALPATION, CONFIRM FETAL LIFE USING OBSTETRIC ULTRASONOGRAPHY. Manufacturer Narrative: E1: REPORTING INSTITUTION PHONE #: (B)(6). E1: REPORTER PHONE #: (B)(6). PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21597219&pc=NPT | 2025587-2025-01881 | 2025-02-06 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-13 04:00 | NPT | EVOLUT FX PLUS VALVE | Perivalvular Leak; Patient Device Interaction Problem | Cardiomyopathy; Low Blood Pressure/ Hypotension; Heart Failure/Congestive Heart Failure; Valvular Insufficiency/ Regurgitation | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT PRIOR TO A TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE INVOLVING THE TRANSCATHETER AORTIC VALVE, THE PROCEDURE WAS INITIALLY DELAYED TWICE DUE TO ANEMIA AND NON-ISCHEMIC ST ELEVATIONS. A PRE-IMPLANT BALLOON VALVULOPLASTY WAS PERFORMED DUE TO CALCIFICATION. DURING THE DEPLOYMENT OF THE VALVE, THE PATIENT EXPERIENCED HYPOTENSION, WHICH WAS MANAGED WITH MEDICATION. AFTER THE FINAL DEPLOYMENT, THE PATIENT'S BLOOD PRESSURE NORMALIZED, BUT A MODERATE TO SEVERE PARAVALVULAR LEAK WAS NOTED, WHICH WAS REDUCED TO MILD TO MODERATE AFTER TWO POST-DILATIONS USING A 24 MM BALLOON. FOLLOWING THE PROCEDURE, THE PATIENT EXPERIENCED ACUTE LEFT VENTRICLE FAILURE. A TEMPORARY HEART PUMP WAS PLACED. SUBSEQUENTLY, RIGHT VENTRICLE FAILURE OCCURRED, AND THE PATIENT WAS UNABLE TO BE FURTHER SUPPORTED, LEADING TO THEIR DEATH. THE PATIENT WAS DIAGNOSED WITH CARDIOMYOPATHY. PER THE PHYSICIAN, BOTH THE DEVICE AND THE PROCEDURE CONTRIBUTED TO THE PATIENT'S DEATH. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID: D-EVOLUTFX-2329 (UNKNOWN); PRODUCT TYPE: 0195-HEART VALVES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME: DELIV SYS D-EVOLUTFX-2329; PRODUCT ID: D-EVOLUTFX-2329; LOT #: 0012443476 CORRECTED DATA: G3. UPDATED DATA: D10. H6. H11. ADDITIONAL CODES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21597389&pc=DSQ | 2916596-2025-01774 | 2025-02-17 05:00 | Death | THORATEC CORPORATION | 2025-03-13 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Headache; Intracranial Hemorrhage | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO A LARGE HEAD BLEED. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER.Event Description: IT WAS LATER REPORTED THAT THE CAUSE FOR THE HEAD BLEED WAS UNKNOWN AND THAT THERE WERE NO CHANGES TO PATIENTS' ANTICOAGULATION STATUS LEADING UP TO THE EVENT. IT WAS STATED THAT THE PATIENT REPORTED A HEADACHE AT HOME LEADING UP TO THE EVENT. ADDITIONALLY. IT WAS REPORTED THAT THE DEVICE WAS OPERATING AS EXPECTED AND DEATH WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF STROKE COULD NOT BE CONCLUSIVELY DETERMINED. STROKE IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. STROKE HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21597458&pc=NIK | 2017865-2025-17406 | 2025-02-20 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-03-13 04:00 | NIK | GALLANT DR | Failure to Deliver Shock/Stimulation | Arrhythmia; No Clinical Signs, Symptoms or Conditions | P030054 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-17408. IT WAS REPORTED THAT THE PATIENT HAD DECEASED DUE TO AN UNKNOWN CAUSE. NO ADDITIONAL INFORMATION WAS REPORTED.Event Description: IT WAS REPORTED THAT THE PATIENT HAD DECEASED DUE TO AN UNKNOWN CAUSE AND FAILURE TO DELIVER SHOCK WAS NOTED ON THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) AND THE RIGHT VENTRICULAR (RV) LEAD. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: CORRECTION: B5 FOR MISSING FAILURE TO DELIVER THERAPY STATEMENT AND H6 FOR ARRHYTHMIA CODING. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21597466&pc=NVY | 2017865-2025-17408 | 2025-02-20 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-03-13 04:00 | NVY | DURATA STS OPTIM PASSIVE FIXATION, DF-4 CONNECTOR | Failure to Deliver Shock/Stimulation | Arrhythmia; No Clinical Signs, Symptoms or Conditions | P950022 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-17406. IT WAS REPORTED THAT THE PATIENT HAD DECEASED DUE TO AN UNKNOWN CAUSE. NO ADDITIONAL INFORMATION WAS REPORTED.Event Description: IT WAS REPORTED THAT THE PATIENT HAD DECEASED DUE TO AN UNKNOWN CAUSE AND FAILURE TO DELIVER SHOCK WAS NOTED ON THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) AND THE RIGHT VENTRICULAR (RV) LEAD. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES. Manufacturer Narrative: CORRECTION: B5 FOR MISSING FAILURE TO DELIVER THERAPY STATEMENT AND H6 FOR ARRHYTHMIA CODING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21599050&pc=DSQ | 2916596-2025-01505 | 2025-02-25 05:00 | Death | THORATEC CORPORATION | 2025-03-13 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Hematoma | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO OCCIPITAL HEMATOMA WITH MIDLINE SHIFT. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. AN AUTOPSY WOULD NOT BE PERFORMED AND THE PUMP WOULD NOT BE EXPLANTED OR RETURNED. ADDITIONAL INFORMATION PROVIDED THAT THE FAMILY DECIDED TO WITHDRAW CARE FROM THE PATIENT. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: CORRECTED DATA: REPORT SUBTYPE CORRECTED TO DEATH REPORT. THE PUMP REMAINS IN USE SUPPORTING THE PATIENT. A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF STROKE COULD NOT BE CONCLUSIVELY DETERMINED. STROKE IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. STROKE HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21599431&pc=NQX | 3004530258-2025-00009 | 2025-01-26 05:00 | Death | CEPHEID | 2025-03-13 04:00 | NQX | XPERT MRSA NXG | False Negative Result | Cardiac Arrest; Insufficient Information |  |  | 1 | Event Description: US CUSTOMER CONTACTED CEPHEID TO DISCUSS DISCREPANT RESULTS FOR ONE PATIENT USING MRSA NXG FOR THE MRSA ANALYTE. ON (B)(6) 2025, A NASAL SWAB WAS COLLECTED ON LABEL USING THE COPAN ESWAB COLLECTION DEVICE AND PROCESSED PER THE IFU AND RUN ON XPERT MRSA NXG RESULTED IN MRSA NOT DETECTED. THIS RESULT WAS REPORTED TO THE PHYSICIAN. THE SAMPLE WAS PROCESSED PER THE IFU. THE SAMPLE WAS STORED REFRIGERATED PRIOR TO TESTING. THE CUSTOMER UTILIZED THE TEST TO SCREEN PATIENT FOR MRSA COLONIZATION. A SEPARATE BLOOD CULTURE SAMPLE WAS COLLECTED FROM THE PATIENT ON (B)(6) 2025. THE DIRECT BLOOD SAMPLE WAS PLATED AND THE ISOLATE GROWN WAS IDENTIFIED AS MRSA. THIS RESULT PROMPTED THE CUSTOMER TO PLATE THE NASAL ESWAB SAMPLE 1 ON BAP, CHOC, AND MACCONKEY PLATES. AN ORGANISM GREW AND WAS ISOLATED AND IT WAS IDENTIFIED AS MRSA BY VITEK 2 SUSCEPTIBILITY TESTING, WHICH SHOWED AMOXICILLIN, AMPICILLIN, OXACILLIN, CEFAZOLIN, CIPROFLOXACIN, AND LEVOFLOXACIN RESISTANCE. THE CUSTOMER REPORTED PBP2A TESTING WAS ALSO PERFORMED ON SPECIMEN WHICH RESULTED AS POSITIVE. THE CUSTOMER DID NOT REPORT ANY KNOWN DEATH, INJURY OR DETERIORATION IN HEALTH RELATED TO THE DISCREPANCY AT THE TIME OF THE INITIAL CALL WITH CEPHEID TECHNICAL SUPPORT.Event Description: US CUSTOMER CONTACTED CEPHEID TO DISCUSS DISCREPANT RESULTS FOR ONE PATIENT USING MRSA NXG FOR THE MRSA ANALYTE. ON (B)(6) 2025, PATIENT WAS ADMITTED IN A STATE OF SHOCK. ON THE SAME DAY, A NASAL SWAB (SAMPLE 1) WAS COLLECTED ON LABEL USING THE COPAN ESWAB COLLECTION DEVICE AND PROCESSED PER THE IFU. SAMPLE WAS RUN ON (B)(6) 2025 ON XPERT MRSA NXG RESULTED IN MRSA NOT DETECTED. THIS RESULT WAS REPORTED TO THE PHYSICIAN. A SEPARATE BLOOD CULTURE SAMPLE (SAMPLE 2) WAS COLLECTED FROM THE PATIENT ON (B)(6) 2025. THE DIRECT BLOOD SAMPLE WAS PLATED AND THE ISOLATE GROWN WAS IDENTIFIED AS MRSA. THIS RESULT PROMPTED THE CUSTOMER TO PLATE THE NASAL ESWAB SAMPLE 1 ON BAP, CHOC, AND MACCONKEY PLATES. AN ORGANISM GREW AND WAS ISOLATED AND IT WAS IDENTIFIED AS MRSA BY VITEK 2 SUSCEPTIBILITY TESTING, WHICH SHOWED AMOXICILLIN, AMPICILLIN, OXACILLIN, CEFAZOLIN, CIPROFLOXACIN, AND LEVOFLOXACIN RESISTANCE. THE CUSTOMER REPORTED PBP2A TESTING WAS ALSO PERFORMED ON SPECIMEN WHICH RESULTED AS POSITIVE. DATE OF RESULTS UNKNOWN. ON (B)(6) 2025, SAMPLE 2 WAS RAN ON BIOFIRE BCID WHICH RESULTED IN S. AUREUS, NEGATIVE FOR MRSA. THE PATIENT RECEIVED A SINGLE DOSE OF VANCOMYCIN ON (B)(6) 2025. IT WAS STOPPED ACCORDING TO THE HOSPITAL PROTOCOLS FOR ANTIBIOTIC STEWARDSHIP AND WAS NOT RESTARTED. CEPHEID DOES NOT KNOW WHAT GUIDED THE ANTIBIOTIC THERAPY. THE FOLLOWING ANTIBIOTICS WERE ADMINISTERED: PIPERACILLIN/TAZOBACTAM, CLINDAMYCIN, CEFTRIAXONE, AZITHROMYCIN, CEFAZOLIN. THE PATIENT PASSED AWAY ON (B)(6) 2025. Manufacturer Narrative: THE CASE WAS FURTHER REVIEWED, AND THE LAB MANAGER WAS CONTACTED BY PSB. THE LAB MANAGER DID NOT HAVE ANY INFORMATION ABOUT HOW THE FALSE NEGATIVE RESULT AFFECTED THE PATIENT BUT INFORMED US THAT THE PATIENT IS NOW DECEASED. THE LAB MANAGER COMMUNICATED THAT THE BIOFIRE TEST WAS ALSO NEGATIVE FOR MRSA, ALTHOUGH AST AND PBP2A WERE POSITIVE. THE LAB MANAGER REFERRED CEPHEID TO THE MEDICAL DIRECTOR FOR MORE INFORMATION. THE MEDICAL DIRECTOR STATED THAT ONLY ONE DOSE OF VANCOMYCIN WAS GIVEN TO THE PATIENT PRIOR TO THE XPERT MRSA NXG RESULT. VANCOMYCIN THERAPY WAS STOPPED IN WHICH CASE APPROPRIATE THERAPY WAS DISCONTINUED. THE PATIENT WAS SWITCHED TO EITHER NAFCILLIN OR CEFAZOLIN (MEDICAL DIRECTOR COULD NOT SPECIFY WHICH) FOLLOWING THE XPERT MRSA NXG AND BIOFIRE BCID RESULTS. THE MEDICAL DIRECTOR DID NOT PROVIDE THE EXACT DATE THE BIOFIRE TEST WAS PERFORMED. THE INTENDED USE OF XPERT MRSA NXG IS TO DETECT THE PRESENCE OF MRSA DNA IN PATIENTS AT RISK FOR MRSA COLONIZATION, SO IT IS UNLIKELY TO HAVE GUIDED THE ANTIBIOTIC THERAPY. ROOT CAUSE: THE RESULTS ALL POINT TO A POSSIBLE MUTATION IN THE REGION TARGETED BY THE SCC REVERSE PRIMERS. PER REPORT, THE PATIENT PASSED AWAY. THE CHRONOLOGY OF EVENTS, WITH XPERT MRSA NXG TEST ON (B)(6) 2025, BLOOD DRAW ON (B)(6) 2025, SO CULTURE RESULTS AND LIKELY BIOFIRE BCID TEST ON (B)(6) 2025 SUGGEST THE ANTIBIOTIC THERAPY WOULD HAVE BEEN CORRECTED WITHIN A FEW HOURS. AT THIS POINT, THE INVESTIGATION IS STILL ONGOING. AT THIS TIME, PATIENT HARM HAS BEEN REPORTED, HOWEVER MORE INFORMATION NEEDS TO BE OBTAINED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE ADDITIONAL INFORMATION IS RECEIVED. INTENDED USE PER IFU: THE XPERT MRSA NXG TEST, PERFORMED ON THE GENEXPERT® INSTRUMENT SYSTEM, IS A QUALITATIVE IN VITRO DIAGNOSTIC TEST INTENDED FOR THE DETECTION OF METHICILLIN-RESISTANT STAPHYLOCOCCUS AUREUS (MRSA) DNA DIRECTLY FROM NASAL SWABS IN PATIENTS AT RISK FOR NASAL COLONIZATION. THE TEST UTILIZES AUTOMATED REAL-TIME POLYMERASE CHAIN REACTION (PCR) FOR THE AMPLIFICATION OF MRSA-SPECIFIC DNA TARGETS AND FLUOROGENIC TARGET-SPECIFIC HYBRIDIZATION PROBES FOR THE REAL-TIME DETECTION OF THE AMPLIFIED DNA. THE XPERT MRSA NXG TEST IS INTENDED TO AID IN THE PREVENTION AND CONTROL OF MRSA INFECTIONS IN HEALTHCARE SETTINGS. THE XPERT MRSA NXG TEST IS NOT INTENDED TO DIAGNOSE, GUIDE, OR MONITOR TREATMENT FOR MRSA INFECTIONS, OR PROVIDE RESULTS OF SUSCEPTIBILITY TO METHICILLIN. A NEGATIVE RESULT DOES NOT PRECLUDE MRSA NASAL COLONIZATION. CONCOMITANT CULTURES ARE NECESSARY TO RECOVER ORGANISMS FOR EPIDEMIOLOGICAL TYPING OR FOR FURTHER SUSCEPTIBILITY TESTING. NOTE FOR SECTION H3 DEVICE EVALUATED BY MANUFACTURER - ANSWER OF NO IS DUE TO THE SINGLE USE OF THE XPERT MRSA NXG TEST AND THE UNAVAILABILITY OF THAT PRODUCT LOT TO BE RETURNED TO CEPHEID. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED FROM THE CUSTOMER IN B5, B6 AND H11 IS IN RESPONSE TO FDA QUESTIONS AS WELL AS CLARIFICATION ON TESTS PERFORMED. THE CORRECTION APPLIES TO SECTION B5 IN ADDING THE BIOFIRE TEST WHICH WAS INADVERTENTLY LEFT OUT OF THE B5 SECTION IN THE INITIAL REPORT. THE CORRECTION ALSO APPLIES TO SECTION B6 TO CORRECT SAMPLE NUMBER AND UNKNOWN DATE. AS REPORTED TO CEPHEID BY CUSTOMER, THE HOSPITAL WAS FOLLOWING THEIR OWN PROTOCOLS FOR ANTIBIOTIC STEWARDSHIP. CEPHEID DOES NOT KNOW WHAT GUIDED THE ANTIBIOTIC THERAPY. HOWEVER, BASED UPON CEPHEID MEDICAL AFFAIRS STAFF'S EXPERIENCE, WE BELIEVE THAT IT IS LIKELY THAT THE BIOFIRE BCID RESULTS PERFORMED ON A POSITIVE BLOOD CULTURE (SAMPLE 2) WHICH RESULTED NEGATIVE FOR MRSA PROMPTED DECISION ON WHICH ANTIBIOTIC TO USE. ON THE OTHER HAND, THE XPERT MRSA NXG TEST IS NOT INTENDED TO DIAGNOSE, GUIDE, OR MONITOR TREATMENT FOR MRSA INFECTIONS, OR PROVIDE RESULTS OF SUSCEPTIBILITY TO METHICILLIN. IN THIS SPECIFIC CASE, THE USE OF XPERT MRSA NXG TEST RESULTS FROM SAMPLE 1 (NASAL SWAB) WHICH RESULTED NEGATIVE FOR MRSA TO GUIDE ANTIBIOTIC TREATMENT WOULD BE A USER ERROR. THE PATIENT WAS ADMITTED ON (B)(6) IN A STATE OF SHOCK AFTER CARDIOPULMONARY RESUSCITATION DUE TO PULSELESSNESS, WITH ACTIVE SUPPORT FROM VASOACTIVE AMINES. AGGRESSIVE RESUSCITATION COMBINING VASOACTIVE AMINES AND DIALYSIS LED TO TEMPORARY IMPROVEMENT, WHICH HAD TO BE INTERRUPTED DUE TO CARDIAC ARRESTS OCCURRING ON (B)(6), LEADING TO THE FAMILY DECISION TO WITHDRAW THERAPEUTIC INTERVENTION. THE ETIOLOGY OF THE INITIAL SHOCK IS UNKNOWN. CEPHEID MEDICAL AFFAIRS BELIEVE THAT THE AVAILABLE ANAMNESIS ELEMENTS SUGGEST A SEVERE CONDITION AT ADMISSION WITH A HIGHLY CRITICAL SHORT-TERM PROGNOSIS, OCCURRING IN A CONTEXT WITH RISK FACTORS FOR MORTALITY. THE DATE OF DEATH OF THE PATIENT WAS (B)(6) 2025. THE CUSTOMER DOES NOT HAVE ACCESS TO THE PATIENT'S OFFICIAL DEATH CERTIFICATE. ROOT CAUSE OF THE FALSE NEGATIVE MRSA RESULT WITH XPERT MRSA MXG: THE ISOLATED MRSA FROM PATIENT'S NASAL SWAB HAS BEEN SENT TO CEPHEID FOR ANALYSIS. SEQUENCING ANALYSIS OF THE REGION TARGETED BY OUR SCC PRIMERS, THE JUNCTION REGION BETWEEN SCCMEC AND THE S. AUREUS CHROMOSOME, REVEALED THE PRESENCE OF A LARGE INSERTED SEQUENCE THAT SEPARATED THE FORWARD FROM THE REVERSE PRIMER, THUS PREVENTING AMPLIFICATION OF THE SCC TARGET AND THE MRSA DETECTED CALLOUT. ANNEX CODES HAVE BEEN UPDATED TO REFLECT THE COMPLETION OF THE INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21599590&pc=NGV | 2124215-2025-15467 | 2022-06-01 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-03-13 04:00 | NGV | WATCHMAN FLX LEFT ATRIAL APPENDAGE CLOSURE DEVICE WITH DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P130013 |  | 1 | Event Description: REPORTED VIA JOURNAL ARTICLE IT WAS REPORTED THAT DEATH(S) OCCURRED. THE FOLLOWING EVENT WAS OBTAINED VIA A RETROSPECTIVE ARTICLE FOLLOWING AN OBSERVATIONAL COHORT STUDY OF 135 CONSECUTIVE PATIENTS WHO WERE REFERRED FOR LEFT ATRIAL APPENDAGE CLOSURE PROCEDURES WHERE WATCHMAN FLX CLOSURE DEVICES WERE IMPLANTED USING THE WATCHMAN FXD CURVE ACCESS SYSTEM DURING THE TIME FRAME BETWEEN JUNE 2022 AND APRIL 2024. IT WAS REPORTED THAT THE FOLLOWING OCCURRED: COMPLICATIONS AT 45 DAYS AND 6 MONTHS INCLUDED CARDIOVASCULAR MORTALITY, AND ALL-CAUSE MORTALITY. Manufacturer Narrative: B3: USED 06/01/2022 AS EVENT DATE AS PROCEDURES OCCURRED FROM JUNE 2022 TO APRIL 2024. D3 MODEL NUMBER, D3 LOT NUMBER, D4 EXPIRATION DATE, D4 UNIQUE IDENTIFIER (UDI) # ARE UNKNOWN AS THE INFORMATION WAS NOT PROVIDED IN THE LITERATURE ARTICLE. LITERATURE CITATION: HUSSAIN K ET AL. IMPACT OF MODERATE SEDATION ON ELECTROPHYSIOLOGY LAB TIME FOR LEFT ATRIAL APPENDAGE OCCLUSION USING 4D INTRACARDIAC ECHOCARDIOGRAPHY. J CARDIOVASC ELECTROPHYSIOL. 2024;1-9. DOI:10.1111/JCE.16445.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21599599&pc=DSQ | 2916596-2025-01509 | 2025-02-03 05:00 | Death | THORATEC CORPORATION | 2025-03-13 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THERE WERE NO DEVICE ISSUES, AND THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. IT WAS COMMUNICATED THAT AN AUTOPSY WAS NOT PERFORMED, AND THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. IT WAS COMMUNICATED THAT DUE TO PATIENT PRIVACY LAWS, THE ACCOUNT WILL NOT PROVIDE ANY FURTHER PATIENT OUTCOME INFORMATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. THE IFU LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: A1-A4: PATIENT INFORMATION CANNOT BE PROVIDED DUE TO PERSONAL DATA PRIVACY LEGISLATION/POLICY. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21599876&pc=BZD | 2518422-2025-034005 | 2024-04-25 04:00 | Death | RESPIRONICS, INC. | 2025-03-13 04:00 | BZD | PHILIPS CPAP DEVICE | Degraded | Insufficient Information |  |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT PASSED AWAY. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972,Z-1973, AND Z-1974.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21600284&pc=OPR | 2032493-2025-00219 | 2025-02-17 05:00 | Death | MICROVENTION, INC. | 2025-03-13 04:00 | OPR | WEB | Device Dislodged or Dislocated; Migration | Ischemia; Paralysis; Obstruction/Occlusion | P170032 |  | 1 | Event Description: IT WAS INITIALLY REPORTED THAT A WEB 6X3 DEVICE WAS IMPLANTED FOR AN ANV RUPTURED CAROTID END. EASILY DETACHED. THE PATIENT WOKE UP WITH HEMIPLEGIA. PATIENT TAKEN BACK TO THE OR. WEB EXITED ANEURYSM AND MIGRATED TO M1, CAUSING OCCLUSION. UNSUCCESSFUL ATTEMPT TO RECOVER WEB. THE PATIENT WAS THEN REPORTED TO BE IN A ¿CLINICALLY SERIOUS¿ CONDITION. ADDITIONAL INFORMATION LATER REPORTED TO MICROVENTION STATED THAT ON (B)(6) 2025 IN THE EARLY AFTERNOON, A PATIENT WAS TAKEN TO THE NEURORADIOLOGY UNIT FOR EMBOLIZATION OF A RUPTURED OF A RUPTURED LEFT T-CAROTID ANEURYSM. UNDER COVER OF A 4 X15 BALLOON, A WEB 6X3 IN THE ANEURYSM, CONTROL SHOWED GOOD OCCLUSION OF THE ANEURYSM, AND THE ADJACENT BRANCHES PERMEABLE. THE FINAL CONTROL WAS PERFECT, THE M1 AND A1 BRANCHES REMAIN PERMEABLE. WHEN THE PATIENT WOKE UP, PRESENTED WITH SEVERE RIGHT HEMIPLEGIA AND A COMPLETE LEFT SYLVIAN PICTURE. AS SOON AS SYMPTOMS WERE NOTED, MRI EXAMINATION [WAS PERFORMED] TO [LOCATE] THE ORIGIN OF THE HEMIPLEGIA. ON MRI, PERFORMED IMMEDIATELY AFTERWARDS OCCLUSION OF LEFT M1 WITH VERY LARGE DOWNSTREAM ISCHEMIA. THE WEB HAD MIGRATED TO M1 LEFT AND HAD OCCLUDED THE ARTERY. THE PATIENT WAS TAKEN BACK TO THE NEURORADIOLOGY UNIT FOR THROMBECTOMY IN THE LATE AFTERNOON TO ATTEMPT TO WEB REMOVAL WITH A GOOSE-NECK SNARE DEVICE BUT WITHOUT SUCCESS. COMPLICATED PROCEDURE INVOLVING MULTIPLE TECHNIQUES AND MATERIALS TO ATTEMPT TO REMOVE THE DMI. UNSUCCESSFUL CAPTURE OF WEB, WHICH MIGRATED TO M1 DESPITE THREE HOURS' ATTEMPT. PERSISTENT OCCLUSION OF THE AVC M1 MIDDLE CEREBRAL ARTERY WAS OCCLUDED BY THE WEB MIGRATION. CLINICAL CONSEQUENCES: PATIENT DEATH THE DATE OF THE PROCEDURE WAS (B)(6) 2025, AT 3:00 P.M. AND THE DATE OF THE PATIENT'S DEATH WAS (B)(6) 2025, AT 4:00 A.M. THE CAUSE OF THE PATIENT¿S DEATH WAS A VERY SIGNIFICANT STROKE DUE TO OCCLUSION OF THE M1 SEGMENT OF THE LEFT CEREBRAL ARTERY CAUSED BY WEB MIGRATION. Manufacturer Narrative: INVESTIGATION FINDINGS: ITEMS RETURNED: - N/A; THE DEVICE REMAINS IMPLANTED IN THE PATIENT AND IS NOT AVAILABLE FOR RETURN. VISUAL ANALYSIS: A VISUAL INSPECTION OF THE DEVICE CAPTURED IN THIS FILE COULD NOT BE PERFORMED AS A PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION, NOR WERE ANY IMAGES OF THE DEVICE PROVIDED IN PLACE OF A DEVICE RETURN. PROCEDURE AND MEDICAL IMAGING WERE NOT PROVIDED FOR THIS INVESTIGATION. INVESTIGATION FINDINGS: WITHOUT THE RETURN AND PHYSICAL EVALUATION OF THE DEVICE, THE INVESTIGATION CANNOT FURTHER EXAMINE IF A CONDITION EXISTED THAT WOULD HAVE CONTRIBUTED TO THE REPORTED EVENT. WITHOUT IMAGING, THE INVESTIGATION CANNOT FURTHER VERIFY IF THE EVENT OCCURRED AS DESCRIBED, NOR COULD THE INVESTIGATION FURTHER EVALUATE THE CAUSE OF THE REPORTED EVENT. BASED ON A REVIEW OF THE DEVICE¿S RISK DOCUMENTATION, THE REPORTED EVENT DID NOT INDICATE THERE WAS THE PRESENCE OF ANY POTENTIAL OR NEW MANUFACTURING, DESIGN, QUALITY, OR OTHER SYSTEMIC ISSUES, OR NON-CONFORMANCES. THE COMPLAINT CODE IS MONITORED THROUGH THE TRENDING PROCESS; CORRECTIVE ACTION IS DETERMINED, AS NEEDED, THROUGH THIS PROCESS. INVESTIGATIONS OF HISTORIC COMPLAINT FILES WITH SIMILAR COMPLAINT CATEGORY CODING ARE RECORDED IN THE COMPLAINT HANDLING SYSTEM; WITHOUT THE ABILITY TO PERFORM AN ANALYSIS OF THE DEVICE, THIS INVESTIGATION CANNOT FURTHER EXAMINE ANY POTENTIAL ROOT CAUSES. BATCH REVIEW: A SEARCH FOR NON-CONFORMANCES ASSOCIATED WITH THIS PART / LOT NUMBER COMBINATION DID NOT REVEAL ANY PRODUCTION-RELATED ISSUES RELEVANT TO THE COMPLAINT THAT OCCURRED DURING MANUFACTURING OF THE DEVICE. COMPLAINT SYSTEM REVIEW: THERE ARE NO SIMILAR COMPLAINTS BASED ON THE COMPLAINT CATEGORY REGARDING THIS BATCH NUMBER FROM THE LAST TWO YEARS RECORDED IN THE COMPLAINT SYSTEM AT THE TIME OF THIS INVESTIGATION. IFU REVIEW (ADDITIONAL INFORMATION CAN BE FOUND IN THE IFU): POTENTIAL COMPLICATIONS: POTENTIAL COMPLICATIONS INCLUDE BUT ARE NOT LIMITED TO THE FOLLOWING: HEMATOMA AT THE SITE OF ENTRY, ANEURYSM RUPTURE, EMBOLI, VESSEL PERFORATION, PARENT ARTERY OCCLUSION, HEMORRHAGE, ISCHEMIA, VASOSPASM, CLOT FORMATION, DEVICE MIGRATION OR MISPLACEMENT, PREMATURE OR DIFFICULT DEVICE DETACHMENT, NON-DETACHMENT, INCOMPLETE ANEURYSM FILLING, REVASCULARIZATION, POST-EMBOLIZATION SYNDROME, AND NEUROLOGICAL DEFICITS INCLUDING STROKE AND DEATH. WARNINGS AND PRECAUTIONS: THE WEB ANEURYSM EMBOLIZATION SYSTEM IS INTENDED FOR SINGLE USE ONLY. THE DETACHMENT CONTROL DEVICE IS INTENDED TO BE USED FOR ONE PATIENT. DO NOT RE-STERILIZE AND / OR REUSE THE DEVICE. REUSE AND/OR RE-STERILIZATION CAN INCREASE RISK OF INFECTION, CAUSE A PYROGENIC RESPONSE OR OTHER LIFE-THREATENING COMPLICATIONS. REUSE AND / OR RE-STERILIZATION CAN DEGRADE PRODUCT PERFORMANCE, LEADING TO DEVICE MALFUNCTION. DISPOSE OF ALL DEVICES IN ACCORDANCE WITH APPLICABLE HOSPITAL, ADMINISTRATIVE AND/OR LOCAL GOVERNMENT POLICY. ADVANCE AND RETRACT THE DEVICE SLOWLY. DO NOT ADVANCE THE DELIVERY DEVICE WITH EXCESSIVE FORCE. DETERMINE THE CAUSE OF ANY UNUSUAL RESISTANCE. REMOVE THE DEVICE IF EXCESSIVE FRICTION IS NOTED AND CHECK FOR DAMAGE. DO NOT ROTATE THE DELIVERY DEVICE DURING OR AFTER DELIVERY OF THE EMBOLIZATION DEVICE. ROTATING THE DEVICE MAY RESULT IN DAMAGE OR PREMATURE DETACHMENT. THE EMBOLIZATION DEVICE CANNOT BE DETACHED WITH ANY OTHER POWER SOURCE OTHER THAN A MICROVENTION INC. DETACHMENT CONTROL DEVICE. ENSURE THAT AT LEAST TWO DETACHMENT CONTROL DEVICES ARE AVAILABLE BEFORE INITIATING AN EMBOLIZATION PROCEDURE. PROCEDURE: DETACHMENT OF THE DEVICE (¿)31. THE DETACHMENT CONTROL DEVICE IS PRE-LOADED WITH BATTERIES AND WILL ACTIVATE WHEN THE DELIVERY DEVICE IS PROPERLY CONNECTED. 32. VERIFY THAT THE RHV IS FIRMLY LOCKED AROUND THE DELIVERY DEVICE BEFORE ATTACHING THE DETACHMENT CONTROL DEVICE TO ENSURE THAT THE EMBOLIZATION DEVICE DOES NOT MOVE DURING THE CONNECTION PROCESS. 33. ENSURE THAT THE DELIVERY DEVICE GOLD CONNECTORS ARE CLEAN AND FREE FROM BLOOD OR CONTRAST. IF NECESSARY, WIPE THE CONNECTORS WITH STERILE WATER AND DRY BEFORE CONNECTING. 34. INSERT THE PROXIMAL END OF THE DELIVERY DEVICE INTO THE DETACHMENT CONTROL DEVICE. WHEN THE DELIVERY DEVICE IS PROPERLY CONNECTED, THE LIGHT WILL FLASH GREEN AND AN INTERMITTENT TONE WILL BE HEARD. 35. VERIFY THE EMBOLIZATION DEVICE POSITION BEFORE PRESSING THE DETACHMENT BUTTON. 36. PUSH THE DETACHMENT BUTTON. DURING FIRING, THE LIGHT SHOULD BE SOLID GREEN AND THE BEEP SHOULD BE CONTINUOUS. 37. VERIFY DETACHMENT BY FIRST LOOSENING THE RHV VALVE, THEN PULLING BACK SLOWLY ON THE DELIVERY DEVICE AND VERIFYING THAT THERE IS NOT EMBOLIZATION DEVICE MOVEMENT. IF THE EMBOLIZATION DEVICE DOES NOT DETACH, PUSH THE DETACHMENT BUTTON AGAIN. IF THE DEVICE IS STILL NOT DETACHED, OBTAIN A NEW DETACHMENT CONTROL DEVICE AND ATTEMPT DETACHMENT UP TO TWO ADDITIONAL TIMES. IF IT DOES NOT DETACH, REMOVE THE DELIVERY DEVICE. 38. VERIFY THE POSITION OF THE EMBOLIZATION DEVICE ANGIOGRAPHICALLY THROUGH THE GUIDE CATHETER. 39. PRIOR TO REMOVING THE MICROCATHETER FROM THE TREATMENT SITE, PLACE AN APPROPRIATELY SIZED GUIDEWIRE COMPLETELY THROUGH THE MICROCATHETER LUMEN TO ENSURE THAT NO PART OF THE EMBOLIZATION DEVICE REMAINS WITHIN THE MICROCATHETER. INVESTIGATION CONCLUSION: THE PHYSICAL DEVICE WAS NOT AVAILABLE FOR EVALUATION TO INVESTIGATE IF A CONDITION EXISTED THAT WOULD HAVE CONTRIBUTED TO THE ADVERSE EVENT. SUPPLEMENTAL IMAGING WAS ALSO UNAVAILABLE FOR REVIEW; WITHOUT IMAGING, THE INVESTIGATION CANNOT VERIFY THE EVENT OCCURRED AS DESCRIBED, NOR COULD THE INVESTIGATION FURTHER EXAMINE THE CAUSE OF THE REPORTED EVENT. THIS INFORMATION MAY BE UPDATED IF ADDITIONAL INFORMATION IS PROVIDED AT A LATER DATE. MICROVENTION IS SUBMITTING THIS REPORT TO COMPLY WITH FDA REPORTING REGULATIONS UNDER 21 CFR PARTS 4 AND 803. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MICROVENTION, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MICROVENTION HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION AND HAS PROVIDED AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MICROVENTION, OR ITS EMPLOYEES THAT THE DEVICE, MICROVENTION, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. A GOOD FAITH EFFORT WILL BE MADE TO OBTAIN THE APPLICABLE INFORMATION RELEVANT TO THE REPORT. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21601534&pc=DSQ | 2916596-2025-01751 | 2025-03-03 05:00 | Death | THORATEC CORPORATION | 2025-03-13 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. ADDITIONAL INFORMATION WAS UNKNOWN. DUE TO HOSPITAL POLICY, NO FURTHER INFORMATION WOULD BE PROVIDED. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT'S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU), REV. C, IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21601892&pc=MFA | 3007284006-2025-00046 | 2025-02-14 05:00 | Death | SPECTRANETICS CORPORATION | 2025-03-13 04:00 | MFA | GLIDELIGHT LASER SHEATH | Adverse Event Without Identified Device or Use Problem | Great Vessel Perforation | P960042 |  | 1 | Event Description: A LEAD EXTRACTION PROCEDURE COMMENCED TO REMOVE A RIGHT ATRIAL (RA) LEAD DUE TO OCCLUSION. A RIGHT VENTRICULAR (RV) LEAD WAS PRESENT WITHIN THE PATIENT, BUT WAS NOT TARGETED FOR EXTRACTION. A SPECTRANETICS LEAD LOCKING DEVICE (LLD) WAS INSERTED INTO THE LEAD TO PROVIDE TRACTION, AND SECURED WITH COOK MEDICAL ONE-TIE DEVICES. FIRST, MULTIPLE SPECTRANETICS DEVICES (14F GLIDELIGHT LASER SHEATH, VISISHEATH DILATOR SHEATH, AND 11F TIGHTRAIL ROTATING DILATOR SHEATH) WERE USED TO AID IN THE EXTRACTION. THEN, SWITCHING TO A 16F GLIDELIGHT, AN SUPERIOR VENA CAVA (SVC) PERFORATION OCCURRED. THE PATIENT'S BLOOD PRESSURE DROPPED, AND RESCUE EFFORTS BEGAN, INCLUDING STERNOTOMY. DESPITE RESCUE EFFORTS, THE PATIENT DID NOT SURVIVE THE PROCEDURE. THIS REPORT CAPTURES THE 16F GLIDELIGHT DEVICE IN USE WHEN THE PERFORATION OCCURRED, REQUIRING INTERVENTION BUT RESULTING IN DEATH. THERE WAS NO ALLEGED MALFUNCTION OF ANY SPECTRANETICS DEVICES IN USE DURING THE PROCEDURE. Manufacturer Narrative: A2) PATIENT DATE OF BIRTH - NOT AVAILABLE FROM FACILITY. A4) PATIENT WEIGHT - NOT AVAILABLE FROM FACILITY. PATIENT INFORMATION REGARDING RELEVANT TESTS/LABORATORY DATA OR MEDICAL HISTORY - NOT AVAILABLE FROM FACILITY. D4) DEVICE SERIAL NUMBER - NOT AVAILABLE FROM FACILITY. H3) THE GLIDELIGHT WAS DISCARDED, THUS NO RETURNED PRODUCT INVESTIGATION WAS PERFORMED. H6) GREAT VESSEL PERFORATION AND DEATH ARE KNOWN RISKS OF COMPLICATION WITH USE OF THE GLIDELIGHT DEVICE. SUBMISSION OF THIS REPORT DOES NOT, IN ITSELF, REPRESENT A CONCLUSION BY THE MANUFACTURER AND/OR AUTHORIZED REPRESENTATIVE OR THE NATIONAL COMPETENT AUTHORITY THAT THE CONTENT OF THIS REPORT IS COMPLETE OR ACCURATE, THAT THE MEDICAL DEVICE(S) LISTED FAILED IN ANY MANNER AND/OR THAT THE MEDICAL DEVICE(S) CAUSED OR CONTRIBUTED TO AN ALLEGED DEATH OR DETERIORATION IN THE STATE OF THE HEALTH OF ANY PERSON.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21581656&pc=FGE | 2183870-2025-00128 | 2025-02-22 05:00 | Death | COVIDIEN | 2025-03-12 04:00 | FGE | PROTEGE EVERFLEX SELF-EXPANDING STENT SYSTEM LONG | Activation, Positioning or Separation Problem | Cardiac Arrest; No Clinical Signs, Symptoms or Conditions | P110023 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE INVOLVING A SELF-EXPANDING STENT WOULD NOT EXPAND OR DEPLOY ONCE UNSHEATHED FROM THE CATHETER. THE PROCEDURE WAS CONDUCTED ON THE LEFT SUPERFICIAL FEMORAL ARTERY, PROXIMAL, WHICH HAD A SEVERE CALCIFIED LESION WITH 100% STENOSIS AND A LENGTH OF 150 MILLIMETRES. THE ARTERY DIAMETER WAS 6 MILLIMETRES. IN RELATION TO THE PATIENT¿S ANATOMY AND PRIOR TO THE PROCEDURE, THE ARTERY WAS CLOTTED OFF WITH A CHRONIC TOTAL OCCLUSION, AND A PNEUMBRA AND LYTIC CATHETER WERE USED OVERNIGHT THE DAY BEFORE. MODERATE RESISTANCE WAS ENCOUNTERED WHEN ADVANCING THE DEVICE. THE DEVICE WAS SAFELY REMOVED, AND THE PROCEDURE WAS COMPLETED USING A NEW DEVICE. 10 MINUTES DELAY IN SURGERY DUE TO PRODUCT MALFUNCTION. THE PATIENT DIED ON THE NEXT DAY AND IT IS REPORTED THAT THE CAUSE OF DEATH WAS CARDIAC ARREST. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: THE CTO DID NOT INHIBIT THE DEPLOYMENT/EXPANSION OF THE FIRST PROTEGE. THE FIELDER GUIDEWIRE WAS NOT KINKED PRIOR TO THE DELIVERY OF THE FIRST PROTEGE AND THE GUIDEWIRE WAS EXCHANGED FOR USE WITH THE NEW DEVICE. THE NEW DEVICE USED WAS THE SAME SIZE/TYPE MEDTRONIC PROTEGE EF. IT WAS CONFIRMED CONFIRM THAT THE STENT UNSHEATHED WITHOUT EXPANSION THEN RE-SHEATHED. PRODUCT ANALYSIS THE DEVICE WAS RETURNED WITH THE LOCK PIN LOOSE, AND THE BLUE OUTER SHAFT DETACHED FROM THE DEPLOYMENT HANDLE AND LOCK PIN, THERE IS APPROXIMATELY 1.5MM OF THE STENT EXPOSED, MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21582433&pc=DSQ | 2916596-2025-01430 | 2025-02-06 05:00 | Death | THORATEC CORPORATION | 2025-03-12 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Sepsis; Skin Infection | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT WAS SCHEDULED FOR A HEARTMATE 3 TO HEARTMATE 3 PUMP EXCHANGE FOR (B)(6) 2025 DUE TO INFECTION. IT WAS LATER REPORTED THAT ON (B)(6) 2024 2 COLONIES OF CUTIBACTERIUM ACNES (FORMERLY PROPRIONIBACTERIUM ACNES) WERE IDENTIFIED IN CULTURES. ON (B)(6) 2024 BLOOD CULTURES CAME BACK NEGATIVE AND A NASAL SWAB OF MRSA CAME BACK POSITIVE ON (B)(6) 2024. THE PATIENT WAS INITIALLY DIAGNOSED WITH A DRIVELINE EXIT SITE (DLES) INFECTION ON (B)(6) 2024 WITH A POSITIVE DLES SWAB OF RARE -1 COLONY STAPHYLOCOCCUS AUREUS. PATIENT EXPERIENCED SYMPTOMS OF PAIN, ERYTHEMA, AND PURULENT DRAINAGE AT THE DRIVELINE EXIT SITE. THE PATIENT WAS TREATED WITH INTRAVENOUS (IV) VANCO INITIALLY INPATIENT, THEN IV DAPTOMYCIN AS OUTPATIENT. ON (B)(6) 2024 A DLES SWAB SHOW NO GROWTH, AND THE PATIENT WAS SWITCHED TO DOXYCYCLINE TREATMENT. IN (B)(6) 2024 BLOOD CULTURES CAME BACK NEGATIVE. THE PATIENT WAS ADMITTED (B)(6) 2024 WITH A NEGATIVE BLOOD CULTURE RESULT AND A DLES SWAB OF A COLONY OF STAPHYLOCOCCUS PETTENKOFER. THEY HAD ABNORMAL STAPHYLOCOCCUS AUREUS AND WERE TESTED ON (B)(6) 2024 FOR DLES FUNGAL CULTURE WHICH WAS NEGATIVE. THE PATIENT WAS THEN DISCHARGED ON (B)(6) 2024. PATIENT WAS READMITTED (B)(6) 2024 THROUGH (B)(6) 2024 AND (B)(6) 2024 THROUGH (B)(6) 2024 WITH DLES STAPHYLOCOCCUS AUREUS. FROM (B)(6) 2024 THROUGH (B)(6) 2024 PATIENT WAS FOUND TO HAVE MANY COLONIES OF STAPHYLOCOCCUS AUREUS AROUND THEIR DLES. PATIENT AS FINALLY ADMITTED (B)(6) 2025 FOR THE INFECTION WITH A POSITIVE CULTURE ON (B)(6) 2025 OF DLES STAPHYLOCOCCUS AUREUS. IT WAS REPORTED THAT PATIENT HAD THEIR PUMP EXCHANGE ON (B)(6) 2025. PATIENT WAS ADMITTED UNTIL (B)(6) 2025 WHEN THEY WERE SAID TO HAVE PASSED.Event Description: IT WAS LATER COMMUNICATED THAT THE CAUSE OF DEATH WAS DUE TO SEPSIS. THE PATIENT'S DEATH WAS NOT CONSIDERED DEVICE OR THERAPY RELATED AND THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF INFECTION COULD NOT BE CONCLUSIVELY DETERMINED. INFECTION IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. INFECTION HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SECTION B3: DATE OF EVENT CORRECTED. SECTION H6: HEALTH EFFECT - CLINICAL CODE AND HEALTH EFFECT - IMPACT CODE CORRECTED. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21582448&pc=LTI | 3012638928-2025-03288 | 2025-02-13 05:00 | Death | SPATZ FGIA INC. | 2025-03-12 04:00 | LTI | SPATZ3 ADJUSTABLE BALLOON SYSTEM | Patient-Device Incompatibility; Air/Gas in Device | Perforation | P190012 |  | 1 | Event Description: (B)(6) 2024 - A SPATZ3 BALLOON WAS IMPLANTED AT (B)(6), BY DR. (B)(6).. (B)(6) 2025 - THE PATIENT EXPERIENCED SPONTANEOUS BALLOON HYPERINFLATION, LEADING TO GASTRIC PERFORATION. THE PATIENT WAS HOSPITALIZED AT (B)(6) HOSPITAL, WHERE A LAPAROSCOPY WAS PERFORMED, INCLUDING BALLOON REMOVAL AND GASTRIC SUTURING OF THE PERFORATION. THE PATIENT REMAINED HOSPITALIZED IN THE ICU THEREAFTER. (B)(6) 2025 - WE WERE NOTIFIED THAT THE PATIENT PASSED AWAY. Manufacturer Narrative: THE INVESTIGATION IS STILL ON GOING AND THE FOLLOW UP WILL BE NOTIFIED. TO DATE, SPATZ FGIA INC. HAS NOT RECEIVED THE PRODUCT FOR EVALUATION, THEREFORE NO ANALYSIS OR TESTING HAS BEEN DONE. THE RISK IS SPECIFIED IN RISK MANAGEMENT AND DEVICE LABELLING WAS UPDATE WITH THE FOLLOWING: SPONTANEOUS HYPERINFLATION IS THE ENLARGEMENT OF THE BALLOON WITH EXTRA AIR THAT CAN OCCUR SPONTANEOUSLY. THIS CAN LEAD TO SYMPTOMS SUCH AS PAIN, NAUSEA, VOMITING, DEHYDRATION, ULCERATION, PERFORATION, AND COULD REQUIRE A DOWN ADJUSTMENT OR REMOVAL OF THE BALLOON.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21582533&pc=MVK | 3008642652-2025-02553 | 2025-01-29 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-03-12 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED TEN INAPPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 19:34:47 ON (B)(6) 2025. AT 06:13:03 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH CPR/MOTION ARTIFACT. BETWEEN 06:14:27 AND 06:22:11, THE PATIENT RECEIVED TEN INAPPROPRIATE TREATMENTS. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENTS. THE PATIENT'S RHYTHM AT THE TIME OF THE TREATMENTS WAS ASYSTOLE WITH CPR/MOTION ARTIFACT. THE RESULTING RHYTHM OF THE TREATMENTS WAS ASYSTOLE WITH CPR/MOTION ARTIFACT/ELECTRODE LEAD FALL OFF. THE ELECTRODE BELT WAS DISCONNECTED AT 07:02:47 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR HAS BEEN COMPLETED. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION. THE ELECTRODE BELT HAS NOT BEEN RECOVERED. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21583288&pc=OZD | 1220648-2025-26660 | 2025-02-13 05:00 | Death | ABIOMED, INC. | 2025-03-12 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A 60 YEAR OLD MALE STEMI PATIENT PRESENTED TO THE ER AND TAKEN TO CATH WERE REVEALED SEVERE MITRAL VALVE DISEASE. DECISION WAS MADE TO PROCEED WITH IMPELLA CP PLACEMENT TO UNLOAD THEN REEVALUATE RIGHT VENTRICLE. IT WAS NOTED THAT THE PATIENT BEGAN TO DECOMPENSATE AND REQUIRED 1 ROUND CPR REQUIRING INTUBATION DURING IMPELLA PLACEMENT. AFTER IMPELLA PLACEMENT THE PATIENT BECAME MORE HEMODYNAMICALLY STABLE. DURING SUPPORT IT WAS REPORTED THAT THE PATIENT WENT INTO VENTRICULAR TACHYCARDIA REQUIRING TWO SHOCKS AND RETURNED TO SINUS TACH. ULTIMATELY, CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMART ASSIST SYSTEM STATES THE FOLLOWING: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21583405&pc=OZD | 1220648-2025-26662 | 2025-02-15 05:00 | Death | ABIOMED, INC. | 2025-03-12 04:00 | OZD | IMPELLA CP | Infusion or Flow Problem | Ischemia; Thrombosis/Thrombus | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WITH CARDIOMYOPATHY IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT EXPERIENCED LIMB ISCHEMIA FOLLOWED BY SATURATED PUMP FLOWS. THE PATIENT WOULD SUBSEQUENTLY EXPIRE. THE PATIENT TRANSFERRED FROM AN OUTSIDE HOSPITAL FOR ACUTE CHRONIC HEART FAILURE EXACERBATION. THE PATIENT WAS STARTED ON DOBUTAMINE AND NITROPRUSSIDE. THE FOLLOWING DAY A RIGHT HEART CATHETERIZATION REVEALED A REDUCED CARDIAC INDEX. AFTER DISCUSSION BETWEEN PHYSICIANS, THE DECISION WAS MADE TO PLACE IMPELLA CP TO UNLOAD AND BRIDGE TO A DECISION. AFTER THE RIGHT HEART CATHETERIZATION, THE RIGHT FEMORAL ARTERY WAS ACCESSED FOR IMPELLA CP IMPLANT. IT WAS NOTED HEPARIN 5000 UNITS WAS GIVEN. ACTIVATED CLOTTING TIME WAS 210 SECONDS AND MORE HEPARIN 3000 UNITS WAS GIVEN. THE PUMP WAS STARTED ON AUTO 3.5LPM. PEEL-AWAY WAS REMOVED AND REPOSITIONING SHEATH WAS ADVANCED. SLEEVE AND TUOHY BOST VALVE WAS LOCKED INTO PLACED. BLUE SUTURE HUB ANCHORED WITH FORWARD TENSION SUTURES. AFTER THE CASE WAS COMPLETED, THE PATIENT COMPLAINED RIGHT LOWER EXTREMITY (RLE) PAIN. THE EXTREMITY WAS COOL TO TOUCH AND PALER THAN THE OTHER LEG. DOPPLER TO THE RLE WAS NEGATIVE. THE RE-ACCESS PORT WAS USED TO INJECT CONTRAST INTO THE ARTERY AND IMAGING SHOWED THE ARTERY WAS OCCLUDED. AT THIS POINT, THE TEAM DECIDED TO ATTEMPT AN EXTERNAL FEM-FEM BYPASS. WIRING THE RIGHT SUPERFICIAL ARTERY PROVED UNSUCCESSFUL. THE PATIENT STILL DID NOT HAVE DOPPLERED PULSES, BUT THE LEG WAS WARMER THAN BEFORE. DECISION WAS MADE TO TAKE THE PATIENT TO THE UNIT AND DISCUSS WITH THE TEAM ABOUT NEXT STEPS. VASCULAR SURGERY WAS CONSULTED ABOUT THE PATIENT'S POTENTIAL ISCHEMIC LEG. THE NEXT DAY, THE PATIENT WAS TAKEN BACK TO THE CATHETERIZATION LAB TO ASSESS RLE FLOW. AORTOGRAM SHOWED SLUGGISH FLOW IN THE DISTAL RLE. DISCUSSION WITH INTENSIVE CARE UNIT, CARDIOLOGY AND VASCULAR TEAMS WAS MADE WITH THE DECISION TO LEAVE THE IMPELLA CP IN PLACE AS THERE IS SOME FLOW DISTAL FROM COLLATERALS. IF PATIENT' S CONDITION WORSENS, THEN THE PLAN WAS NOTED TO REMOVE THE IMPELLA CP AND PLACE PERCUTANEOUS SUBCLAVIAN IMPELLA CP. LATER IN THE DAY IT WAS NOTED THERE WERE NO PULSES IN RLE, BUT LEG WAS RELATIVELY WARM. THE PATIENT WAS GOING TO BE WORKED UP FOR A LEFT VENTRICULAR ASSIST DEVICE, AND POSSIBLE IMPELLA 5.5 IMPLANT IN A COUPLE DAYS AFTER A DECISION WAS MADE. TWO DAYS LATER, THE RLE WAS NOW WITH PALPABLE PULSE AND THREE DAYS THEREAFTER, IT WAS NOTED PATIENT WOULD GO FOR DURABLE LVAD. THE PATIENT WAS NOTED COUGHING UP BLOODY SPUTUM/CLOTS AND DID HAVE SOME BLOOD IN THE BOWEL MOVEMENT. DAYS LATER, NOW EIGHT DAYS AFTER START OF THE SUPPORT, HEPARIN WAS STOPPED AT 1700 DUE TO CONCERNS OF HEPARIN-INDUCED THROMBOCYTOPENIA. ARGATROBAN WAS STARTED. LATER IN THE EVENING, THE IMPELLA CP MOTOR CURRENT SPIKED AND LEFT VENTRICLE WAVEFORM BECAME ABNORMAL. MIN/MAX AND MEAN FLOWS ALL 4.2; FLOW SATURATION CONCERNS WERE HIGHLIGHTED. THE FOLLOWING DAY, DISCUSSION WAS MADE REGARDING PUMP FAILURE AND THE NEED TO URGENTLY EXCHANGE THE IMPELLA CP DEVICE. THE DECISION WAS MADE TO PLAN FOR DEVICE ESCALATION TO AN IMPELLA 5.5 CONSIDERING DURABILITY AND BENEFIT OF AXILLARY PLACEMENT. IT WAS ADDITIONALLY NOTED THERE WAS CONCERNS FOR INFECTION- PROBABLE PNEUMONIA AND RULING OUT ASPIRATION AS CAUSE. THE PATIENT WAS PAN CULTURED TO COVER ALL BASES. HEMOGLOBIN CONTINUED TO BE LOW, AND PATIENT CONTINUED TO HAVE BLOODY STOOLS (ALTHOUGH LESS FREQUENT PER THE NURSE). PATIENT WAS CONSIDERED FOR GASTROINTESTINAL (GI) SCOPE PROCEDURE, BUT THE GI PHYSICIAN WAS RELUCTANT AS THE PATIENT WAS ON ARGATROBAN. HOWEVER, THE FOLLOWING DAY, THE PATIENT WAS NOTED TO HAVE EXPIRED; IT WAS INDICATED THAT CARE WAS WITHDRAWN. FURTHER DETAILS AROUND THE PATIENT'S DEMISE WERE NOT PROVIDED AND IT DOES NOT APPEAR THE PATIENT WAS ABLE TO BE ESCALATED TO THE IMPELLA 5.5 AND THE FLOW SATURATION REMAINED UNRESOLVED. DETAILS SURROUNDING THE PATIENT'S DEMISE ARE UNCLEAR OTHER THAN CARE WAS WITHDRAWN. HOWEVER, GIVEN THE IMPELLA-RELATED EVENTS ([IT APPEARS] PARTIALLY RESOLVED LIMB ISCHEMIA AND SATURATED FLOW) THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA CP AS AN ASSOCIATED FACTOR TO THE EVENT. Manufacturer Narrative: THE INVESTIGATION OF LIMB ISCHEMIA AND SATURATED FLOW HAS BEEN COMPLETED. NO PRODUCT RETURNED. REPORTED SPIKE IN MOTOR CURRENT WITH SATURATED FLOWS AROUND 2/23 19:30. HEPARIN WAS STOPPED PRIOR AT 2/23 17:00 DUE TO CONCERNS OF HIT AND ARGATROBAN STARTED SHORTLY AFTER. DATA LOGS SHOW DIPS IN PURGE PRESSURE AND SPIKES IN PUMP AND PURGE FLOW BEGINNING AROUND 2/23 2:00AM AND CONTINUE TILL A SPIKE IN MOTOR CURRENT AT 2/23 19:30. THE MOTOR CURRENT SPIKE IS FOLLOWED BY AN UPWARD SHIFT IN FLOW, MOTOR CURRENT, AND PURGE PRESSURE WITH A DOWNWARD SHIFT IN PURGE FLOW. THE ROOT CAUSE OF THE SATURATED FLOW WAS MOST LIKELY BIOMATERIAL SINCE THERE WAS A MOTOR CURRENT SPIKE FOLLOWED BY AN UPWARD SHIFT IN FLOW, MOTOR CURRENT, AND PURGE PRESSURE WITH A DOWNWARD SHIFT IN PURGE FLOW IN DATA LOGS. THE ROOT CAUSE OF THE LIMB ISCHEMIA WAS NOT DETERMINED SINCE PRODUCT WAS NOT RETURNED AND INSUFFICIENT CLINICAL INFORMATION. G1 MANUFACTURING SITE NAME, ADDRESS, COUNTRY, AND FAX NUMBER WERE ERRONEOUSLY ENTERED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26662. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: IMPELLA CP WITH SMARTASSIST SYSTEM. SECTION: POTENTIAL ADVERSE EVENTS: ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY¿ (INCLUDING VENTRICULAR PERFORATION).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21583540&pc=GKT | 1219343-2025-00003 | 2025-02-08 05:00 | Death | HAEMONETICS CORPORATION | 2025-03-12 04:00 | GKT | NEXSYS PLASMA COLLECTION SYSTEM | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | BK180185 |  | 1 | Event Description: THIS REPORT IS BEING SUBMITTED TO ALIGN WITH CUSTOMER REPORTING REQUIREMENTS. THE USE OF THE HAEMONETICS MEDICAL DEVICE IN THE DONATION PROCEDURE IS NOT SUSPECTED TO HAVE RESULTED IN THE ADVERSE OUTCOME OF THE DONOR. DONOR WAS A 34-YEAR-OLD FEMALE WHO WAS NOTED TO HAVE EXPIRED DUE TO AN UNKNOWN CAUSE. DONOR HAS DONATED PLASMA SINCE (B)(6) 2022. SHE HAS DONATED 16 TIMES IN A LIFETIME AND 6 TIMES THIS YEAR. DONORS LAST DONATION WAS (B)(6) 2025. NO SIGNIFICANT INFORMATION WAS NOTED ON LAST DATE OF DONATION. REVIEW OF CHART NOTED DONOR HAD NO MEDICAL PROBLEMS EXCEPT A SMALL PAINLESS UMBILICAL HERNIA. THE DONOR WAS NOT TAKING ANY MEDICATIONS AND HAD NORMAL PHYSICAL EXAMS AND NORMAL LAB RESULTS DURING THEIR TIME DONATING. SCREENING TEST RESULTS WERE NORMAL EXCEPT DONOR HAS 2 OUT OF LIMIT HEMATOCRITS OF 35% (B)(6) 2025) AND 25% ((B)(6) 2022). REVIEW OF VITALS WERE NORMAL THROUGHOUT PLASMA DONATION. DONOR HAD 2 TEMPORARY DEFERRALS DUE TO OUT OF LIMIT HEMATOCRITS, 1 TEMPORARY DEFERRAL DUE TO BEING A LAPSED DONOR, 1 TEMPORARY MEDICAL DEFERRAL DUE TO MINOR ADVERSE EVENT LABELED AS HYPOTENSIVE/VASOVAGAL REACTION. IT IS UNKNOWN IF AN AUTOPSY WAS PERFORMED. THE CUSTOMER HAS NO INFORMATION FROM THE ATTENDING PHYSICIAN, SURGEON HOSPITAL REPRESENTATIVE OR HEALTHCARE PROFESSIONAL. THE DONOR DID NOT EXPERIENCE A REACTION OR ADVERSE EVENT AND THERE WERE NO ERRORS ON THE MACHINE OR ISSUES REPORTED WITH THE DISPOSABLES DURING THE DONOR'S LAST DONATION ON (B)(6) 2025. THE DEATH WAS FOUND ON A SOCIAL MEDIA POST AND TOOK PLACE SOMETIME AFTER THE DONATION; DEATH DID NOT OCCUR AT THE DONATION SITE. Manufacturer Narrative: A HAEMONETICS FIELD SERVICE ENGINEER EVALUATED THE EQUIPMENT USED DURING THE DONATION. DEVICE WAS TESTED AND NO PROBLEM WAS FOUND. ALL OTHER PARAMETERS VERIFIED. FUNCTION TESTS COMPLETED. MACHINE MEETS MANUFACTURER'S SPECIFICATIONS AND IS READY TO USE. BASED ON THIS DESCRIPTION THE DEVICE IS EVALUATED WITH NO DEFECT. THERE ARE NO NONCONFORMANCES AGAINST THE DEVICE SERIAL NUMBER AND NO CAPAS RELATED TO THIS COMPLAINT. A REVIEW OF THE DHR SHOWS NO ISSUES DURING MANUFACTURING AND ALL TESTING PASSED. THE DISPOSABLES USED WITH THE SYSTEM WERE DISCARDED, HOWEVER THERE WERE NO RECALLS OR ADVERSE TRENDS RELATED TO THE PRODUCT LOTS USED IN THE PROCEDURE. THERE IS NO EVIDENCE TO SUGGEST THAT THE DONOR EVENT WAS RELATED TO THE DEVICE OR DISPOSABLES USED DURING THE PLASMAPHERESIS PROCEDURE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21583629&pc=NIO | 9681442-2025-00069 | 2018-09-05 04:00 | Death | ANGIOMED GMBH & CO. MEDIZINTECHNIK KG | 2025-03-12 04:00 | NIO | E-LUMINEXX VASCULAR STENT | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P080007 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESULTS OF THE CLINICAL TRIAL THAT APPROXIMATELY THREE YEARS NINE MONTHS AND TWO DAYS LATER POST A STENT PLACEMENT ON THE LEFT LEG FOR THE DE-NOVO STENOSIS ON THE COMMON ILIAC ARTERY, THE SUBJECT HAD AN ADVERSE EVENT OF DEATH. THE PRIMARY CAUSE OF DEATH WAS NOT PROVIDED. REPORTEDLY, THE ADVERSE EVENT DEATH WAS POSSIBLY RELATED TO THE STUDY DEVICE AND NOT RELATED TO THE PROCEDURE.Event Description: IT WAS REPORTED THROUGH THE RESULTS OF THE CLINICAL TRIAL THAT APPROXIMATELY THREE YEARS NINE MONTHS AND TWO DAYS LATER POST A STENT PLACEMENT ON THE LEFT LEG FOR THE DE-NOVO STENOSIS ON THE COMMON ILIAC ARTERY, THE SUBJECT HAD AN ADVERSE EVENT OF DEATH. THE PRIMARY CAUSE OF DEATH WAS NOT PROVIDED. REPORTEDLY, THE ADVERSE EVENT DEATH WAS POSSIBLY RELATED TO THE STUDY DEVICE AND NOT RELATED TO THE PROCEDURE. Manufacturer Narrative: H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE E-LUMINEXX VASCULAR STENT PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE E-LUMINEXX VASCULAR STENT PRODUCTS ARE IDENTIFIED IN D2 AND G4. H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. Manufacturer Narrative: H11: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE E-LUMINEXX VASCULAR STENT PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE E-LUMINEXX VASCULAR STENT PRODUCTS ARE IDENTIFIED IN D2 AND G4. MANUFACTURING REVIEW: BASED ON THE EVENT INFORMATION PROVIDED THERE IS NO INDICATION THAT A MANUFACTURING PROCESS MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED ISSUE. THEREFORE, A DHR REVIEW IS CONSIDERED TO BE NOT REQUIRED. INVESTIGATION SUMMARY: THE STENT REMAINS IMPLANTED. X-RAY IMAGES WERE NOT PROVIDED FOR EVALUATION. THE EVALUATION OF THE MATERIAL AVAILABLE IS SUBJECT TO THE RETROSPECTIVE ASSESSMENT AS PART OF POST-MARKET CLINICAL FOLLOW-UP (PMCF) ACTIVITIES. POTENTIAL FACTORS THAT COULD HAVE LED OR CONTRIBUTED TO REPORTED EVENTS ARE EVALUATED. IN THIS CASE THERE WAS NO DOCUMENTED DEFICIENCY OF THE PLACED STENT. NO IRREGULAR STENT PLACEMENT DURING INDEX PROCEDURE AND NO DEFICIENCY OF THE PLACED STENT, SUCH AS DEFORMATION OR FRACTURE WAS REPORTED. THE DEATH OF THE PATIENT WAS DOCUMENTED NOT BEING RELATED TO THE PLACED STENT OR THE STENT PLACEMENT PROCEDURE. THE PATIENT DIED ABOUT THREE YEARS AND NINE MONTHS AFTER THE STENT PLACEMENT PROCEDURE. THE PATIENT WAS A CURRENT SMOKER AND HAD A CLINICAL HISTORY OF HYPERTENSION, DYSLIPIDEMIA, RENAL INSUFFICIENCY, TYPE II DIABETES MELLITUS. BASED ON INFORMATION AVAILABLE, THERE IS NO INDICATION THAT THE DEATH OF THE PATIENT MIGHT BE RELATED TO THE PLACED STENT OR THE STENT PLACEMENT PROCEDURE. LABELING REVIEW: RELEVANT LABELING SUPPLIED WITH THIS PRODUCT WAS REVIEWED. POTENTIAL COMPLICATIONS AND ADVERSE EVENTS WHICH MAY OCCUR DURING USE OF THE DEVICE WERE FOUND TO BE REFERENCED IN THE IFU. USERS AND/OR PATIENTS SHOULD REPORT ANY SERIOUS INCIDENT THAT HAS OCCURRED IN RELATION TO THE DEVICE TO THE MANUFACTURER AND TO EITHER THE EUROPEAN UNION NATIONAL COMPETENT AUTHORITY OR TO THE REGULATORY AUTHORITY OF THE COUNTRY IN WHICH THE USER AND/OR PATIENT IS ESTABLISHED. IN THIS CASE THE DEATH OF THE PATIENT WAS DOCUMENTED NOT BEING RELATED TO THE PLACED STENT OR THE STENT PLACEMENT PROCEDURE G3, H6 (METHOD). SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21583632&pc=NIO | 9681442-2025-00068 | 2019-12-12 05:00 | Death | ANGIOMED GMBH & CO. MEDIZINTECHNIK KG | 2025-03-12 04:00 | NIO | E-LUMINEXX VASCULAR STENT | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P080007 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESULTS OF THE CLINICAL TRIAL THAT APPROXIMATELY THREE YEARS SIX MONTHS AND FOUR DAYS LATER POST A STENT PLACEMENT ON THE RIGHT LEG FOR THE DE-NOVO STENOSIS ON THE COMMON AND EXTERNAL ILIAC ARTERY, THE SUBJECT HAD AN ADVERSE EVENT OF DEATH. THE PRIMARY CAUSE OF DEATH WAS NOT PROVIDED. REPORTEDLY, THE ADVERSE EVENT DEATH WAS POSSIBLY RELATED TO THE STUDY DEVICE AND NOT RELATED TO THE PROCEDURE.Event Description: IT WAS REPORTED THROUGH THE RESULTS OF THE CLINICAL TRIAL THAT APPROXIMATELY THREE YEARS SIX MONTHS AND FOUR DAYS LATER POST A STENT PLACEMENT ON THE RIGHT LEG FOR THE DE-NOVO STENOSIS ON THE COMMON AND EXTERNAL ILIAC ARTERY, THE SUBJECT HAD AN ADVERSE EVENT OF DEATH. THE PRIMARY CAUSE OF DEATH WAS NOT PROVIDED. REPORTEDLY, THE ADVERSE EVENT DEATH WAS POSSIBLY RELATED TO THE STUDY DEVICE AND NOT RELATED TO THE PROCEDURE. Manufacturer Narrative: H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE E-LUMINEXX VASCULAR STENT PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE E-LUMINEXX VASCULAR STENT PRODUCTS ARE IDENTIFIED IN D2 AND G4. H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. Manufacturer Narrative: H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE E-LUMINEXX VASCULAR STENT PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE E-LUMINEXX VASCULAR STENT PRODUCTS ARE IDENTIFIED IN D2 AND G4. H10: MANUFACTURING REVIEW: BASED ON THE EVENT INFORMATION PROVIDED THERE IS NO INDICATION THAT A MANUFACTURING PROCESS MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED ISSUE. THEREFORE, A DHR REVIEW IS CONSIDERED TO BE NOT REQUIRED. INVESTIGATION SUMMARY: NEITHER SAMPLE NOR IMAGES PROVIDED WHICH LEADS TO INCONCLUSIVE RESULT. THE EVALUATION OF THE MATERIAL AVAILABLE IS SUBJECT TO THE RETROSPECTIVE ASSESSMENT AS PART OF POST-MARKET CLINICAL FOLLOW-UP (PMCF) ACTIVITIES. BASED ON THE INVESTIGATION OF THE PROVIDED INFORMATION, THE INVESTIGATION IS CLOSED AS INCONCLUSIVE. BASED ON INFORMATION AVAILABLE, THERE IS NO INDICATION THAT THE DEATH OF THE PATIENT MIGHT BE RELATED TO THE PLACED STENT OR THE STENT PLACEMENT PROCEDURE. LABELING REVIEW: RELEVANT LABELING SUPPLIED WITH THIS PRODUCT WAS REVIEWED. POTENTIAL COMPLICATIONS AND ADVERSE EVENTS WHICH MAY OCCUR DURING USE OF THE DEVICE WERE FOUND TO BE REFERENCED IN THE IFU. USERS AND/OR PATIENTS SHOULD REPORT ANY SERIOUS INCIDENT THAT HAS OCCURRED IN RELATION TO THE DEVICE TO THE MANUFACTURER AND TO EITHER THE EUROPEAN UNION NATIONAL COMPETENT AUTHORITY OR TO THE REGULATORY AUTHORITY OF THE COUNTRY IN WHICH THE USER AND/OR PATIENT IS ESTABLISHED. IN THIS CASE THE DEATH OF THE PATIENT WAS DOCUMENTED NOT BEING RELATED TO THE PLACED STENT OR THE STENT PLACEMENT PROCEDURE. H10: G2, G3. H11: H6(METHOD). H11: SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21583928&pc=NIO | 9681442-2025-00070 | 2021-06-16 04:00 | Death | ANGIOMED GMBH & CO. MEDIZINTECHNIK KG | 2025-03-12 04:00 | NIO | E-LUMINEXX VASCULAR STENT | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P080007 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESULTS OF THE CLINICAL TRIAL THAT APPROXIMATELY FOUR YEARS SIX MONTHS AND TWENTY FOUR DAYS LATER POST A STENT PLACEMENT ON THE LEFT LEG FOR THE DE-NOVO STENOSIS ON THE EXTERNAL ILIAC ARTERY, THE SUBJECT HAD AN ADVERSE EVENT OF DEATH. THE PRIMARY CAUSE OF DEATH WAS NOT PROVIDED. REPORTEDLY, THE ADVERSE EVENT DEATH WAS POSSIBLY RELATED TO THE STUDY DEVICE AND NOT RELATED TO THE PROCEDURE.Event Description: IT WAS REPORTED THROUGH THE RESULTS OF THE CLINICAL TRIAL THAT FOUR YEARS, SIX MONTHS AND TWENTY FOUR DAYS POST A STENT PLACEMENT ON THE LEFT LEG FOR THE DE-NOVO STENOSIS ON THE EXTERNAL ILIAC ARTERY, THE SUBJECT HAD AN ADVERSE EVENT OF DEATH. THE PRIMARY CAUSE OF DEATH WAS NOT PROVIDED. REPORTEDLY, THE ADVERSE EVENT DEATH WAS POSSIBLY RELATED TO THE STUDY DEVICE AND NOT RELATED TO THE PROCEDURE. Manufacturer Narrative: H11: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE E-LUMINEXX VASCULAR STENT PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE E-LUMINEXX VASCULAR STENT PRODUCTS ARE IDENTIFIED IN D2 AND G4. MANUFACTURING REVIEW: BASED ON THE EVENT INFORMATION PROVIDED THERE IS NO INDICATION THAT A MANUFACTURING PROCESS MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED ISSUE. THEREFORE, A DHR REVIEW IS CONSIDERED TO BE NOT REQUIRED. INVESTIGATION SUMMARY: THE STENT REMAINS IMPLANTED. X-RAY IMAGES WERE NOT PROVIDED FOR EVALUATION. THE EVALUATION OF THE MATERIAL AVAILABLE IS SUBJECT TO THE RETROSPECTIVE ASSESSMENT AS PART OF POST-MARKET CLINICAL FOLLOW-UP (PMCF) ACTIVITIES. BASED ON THE INFORMATION AVAILABLE, THE INVESTIGATION IS CLOSED WITH INCONCLUSIVE RESULT. BASED ON INFORMATION AVAILABLE, THERE IS NO INDICATION THAT THE DEATH OF THE PATIENT MIGHT BE RELATED TO THE PLACED STENT OR THE STENT PLACEMENT PROCEDURE. LABELING REVIEW: RELEVANT LABELING SUPPLIED WITH THIS PRODUCT WAS REVIEWED. POTENTIAL COMPLICATIONS AND ADVERSE EVENTS WHICH MAY OCCUR DURING USE OF THE DEVICE WERE FOUND TO BE REFERENCED IN THE INSTRUCTIONS FOR USE. USERS AND/OR PATIENTS SHOULD REPORT ANY SERIOUS INCIDENT THAT HAS OCCURRED IN RELATION TO THE DEVICE TO THE MANUFACTURER AND TO EITHER THE EUROPEAN UNION NATIONAL COMPETENT AUTHORITY OR TO THE REGULATORY AUTHORITY OF THE COUNTRY IN WHICH THE USER AND/OR PATIENT IS ESTABLISHED. IN THIS CASE THE DEATH OF THE PATIENT WAS DOCUMENTED NOT BEING RELATED TO THE PLACED STENT OR THE STENT PLACEMENT PROCEDURE. G3, H6: (COMPONENT, METHOD). SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. Manufacturer Narrative: H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE E-LUMINEXX VASCULAR STENT PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE E-LUMINEXX VASCULAR STENT PRODUCTS ARE IDENTIFIED IN D2 AND G4. H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21584041&pc=QFG | MW5167559 | 2024-08-31 04:00 | Death | TANDEM DIABETES CARE, INC. | 2025-03-12 04:00 | QFG | T-SLIM X2 INSULIN PUMP |  | Diabetic Ketoacidosis |  |  | 1 | Event Description: PATIENT HAD REPORTEDLY COMPLAINED OF MALFUNCTIONING INSULIN PUMP WITH RECURRENT DKA (DIABETIC KETOACIDOSIS). EXPIRED WITH DKA. CALLED TANDEM MULTIPLE TIMES FOR PUMP INTERROGATION AND THEY HAVE NO FOLLOWED UP.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21584080&pc=NVZ | MW5167560 |  | Death | ST. JUDE MEDICAL, INC./ ABBOTT MEDICAL | 2025-03-12 04:00 | NVZ | ASSURITY RF DR |  | Cardiac Arrest; Respiratory Arrest |  |  | 1 | Event Description: CLINICIAN REPORTED THAT THE DEVICE WAS NOT CAPTURING ON THE RIGHT VENTRICULAR LEAD. PATIENT PASSED AWAY. THIS REPORT REFLECTS INFORMATION RECEIVED BY FDA IN THE FORM OF A NOTIFICATION PER 803.22 (B)(2).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21584097&pc=MVK | 3008642652-2025-02565 | 2024-08-10 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-03-12 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2024 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED AN INAPPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 07:14:35 ON (B)(6) 2024. AT 12:42:38, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS SINUS TACHYCARDIA @ 120 BPM WITH MOTION ARTIFACT. AT 12:43:14, THE PATIENT RECEIVED THE INAPPROPRIATE TREATMENT. MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE RHYTHM AT THE TIME OF TREATMENT WAS SINUS TACHYCARDIA @ 120 BPM WITH MOTION ARTIFACT. POST SHOCK RHYTHM WAS AF @ 90 BPM WITH PVCS AND MOTION ARTIFACT. THE ELECTRODE BELT WAS DISCONNECTED AT 13:31:15 ON (B)(6) 2024. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR HAS BEEN COMPLETED. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION. THE ELECTRODE BELT HAS BEEN RETURNED AND THE EVALUATION IS UNDERWAY. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21584235&pc=NPT | 2015691-2025-01667 | 2024-12-31 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-12 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P140031 |  | 1 | Event Description: AS REPORTED BY OUR EDWARDS LIFESCIENCES JAPANESE AFFILIATE, ON POST-OPERATIVE DAY (POD) 21 OF A TRANSFEMORAL TAVR PROCEDURE WITH A 23 MM SAPIEN 3 ULTRA RESILIA VALVE, THE PATIENT EXPIRED. THE CAUSE OF DEATH WAS REPORTED AS CEREBRAL INFARCTION. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), PERMANENT OR TRANSIENT NEUROLOGICAL EVENTS SUCH AS TRANSIENT ISCHEMIC ATTACK (TIA) AND STROKE ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND THE USE OF THE EDWARDS TRANSCATHETER HEART VALVE (THV) DEVICES. ACCORDING TO THE LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EW, STROKE IS RECOGNIZED IN THE LITERATURE AS A WELL-KNOWN COMPLICATION IN A SMALL NUMBER OF PATIENTS UNDERGOING THV. RISK FACTORS CORRELATING WITH SEVERAL PATIENT CO-MORBIDITIES HAVE BEEN IDENTIFIED. ALTHOUGH IN MANY CASES THE ROOT CAUSE OF THE EVENT IS UNABLE TO BE DETERMINED, STROKES DURING THV ARE UNDOUBTEDLY MULTIFACTORIAL, THE DOMINANT ETIOLOGY LIKELY BEING INTRA PROCEDURE EMBOLIC EVENTS. A TRANSCRANIAL DOPPLER STUDY DURING THV DEMONSTRATED THAT MOST PROCEDURAL EMBOLIC EVENTS OCCURRED DURING BALLOON VALVULOPLASTY, MANIPULATION OF CATHETERS ACROSS THE AORTIC VALVE, AND VALVE IMPLANTATION. AN ANALYSIS IN PATIENTS UNDERGOING VALVE SURGERY REVEALED FOUR BASELINE CHARACTERISTICS AND TWO PROCEDURAL EVENTS THAT WERE ASSOCIATED WITH EARLY POST-PROCEDURE STROKE: FEMALE SEX, EF < 30%, DIABETES, AGE OLDER THAN 70 YEARS, BYPASS PROCEDURE TIME> 120 MIN, AND CALCIFICATION OF THE ASCENDING AORTA. PREDICTORS OF LATE STROKE HAVE INCLUDED FEMALE SEX, AGE OLDER THAN 75 YEARS, ATRIAL FIBRILLATION, AND A HISTORY OF OR CURRENT SMOKING. THERE WERE NO SIGNIFICANT DIFFERENCES IN THE FREQUENCY OF LATE STROKES BETWEEN THV AND AVR PATIENTS. AFTER THV, THERE APPEARS TO BE A MORE CONSIDERABLE PROPORTION OF EARLY STROKES OCCURRING < 24 H POST-PROCEDURE, BUT THV PATIENTS WITH MULTIPLE CO-MORBIDITIES ARE PROBABLY AT HIGHER RISK OF BOTH EARLY AND LATE STROKES. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST/INDICATE THAT PATIENT FACTORS - FEMALE, AND 89 YEARS OLD, MAY HAVE CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21584749&pc=PXU | 1125782-2025-00002 | 2025-02-23 05:00 | Death | MERIT MEDICAL SYSTEMS INC. | 2025-03-12 04:00 | PXU | PERICARDIOCENTESIS KIT.8.3F (2.77 MM).40 CM (15.7").PIGTAIL. | Break; Material Separation | Cardiac Arrest | EXEMPT |  | 1 | Event Description: MEDWATCH RECEIVED MW5167050 - THE ACCOUNT ALLEGES THAT IN THE CATH LAB WHILE PERFORMING A PERICARDIAL FLUID DRAINAGE PROCEDURE, THE PATIENT INITIALLY BECAME HEMODYNAMICALLY UNSTABLE AND A CODE BLUE WAS CALLED. THE CODE BLUE WAS TERMINATED PER THE FAMILIES REQUEST AND THE PATIENT EXPIRED. A GUIDEWIRE FROM PERICARDIOCENTESIS KIT WAS FOUND DAMAGED POST PROCEDURE. Manufacturer Narrative: THE SUSPECT MEDICAL DEVICE IS NOT EXPECTED TO RETURN FOR EVALUATION. THE COMPLAINT COULD NOT BE CONFIRMED. THE ROOT CAUSE COULD NOT BE DETERMINED. THE DEVICE HISTORY RECORD WAS REVIEWED, AND NO EXCEPTION DOCUMENTS WERE FOUND. A SEARCH OF THE COMPLAINT DATABASE WAS PERFORMED AND NO SIMILAR COMPLAINTS FOR THIS LOT NUMBER WERE IDENTIFIED. SHOULD THE DEVICE BE RETURNED LATER, THE INVESTIGATION WILL BE REOPENED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21585148&pc=DSQ | 2916596-2025-01782 | 2023-11-18 05:00 | Death | THORATEC CORPORATION | 2025-03-12 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE ELECTRONIC IFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REV. D IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21585175&pc=DSQ | 2916596-2025-01453 | 2025-02-20 05:00 | Death | THORATEC CORPORATION | 2025-03-12 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE DEATH WAS AT THE TIME BEING INVESTIGATED AS A HOMICIDE. AN UNKNOWN ALARM REPORT OCCURRED FROM (B)(6) 2025 TO (B)(6) 2025 WHICH THE CENTER WAS UNABLE TO PROCURE. A LOG FILE REVIEW REVEALED LOW FLOW EVENTS AT 01:44 AM ON (B)(6) 2025. AT 01:48 AM THE FLOW DROPPED TO BETWEEN 0 - 0.5 LITER PER MINUTE (LPM) AND DID NOT RECOVER. NO OTHER UNUSUAL EVENTS WERE RECORDED IN THE LOG FILE EVENT HISTORY. THE MECHANICAL CIRCULATORY SUPPORT EQUIPMENT APPEARED TO HAVE BEEN OPERATING AS INTENDED. Manufacturer Narrative: CORRECTION: SECTION H6. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. FURTHERMORE, REVIEW OF THE SUBMITTED LOG FILES CONFIRMED LOW FLOW EVENTS; HOWEVER, A SPECIFIC CAUSE FOR THESE FINDINGS COULD NOT BE CONCLUSIVELY DETERMINED. IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE DEATH WAS BEING INVESTIGATED AS A HOMICIDE. NO FURTHER DETAILS REGARDING THE PATIENT'S EXPIRATION WERE PROVIDED. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE SUBMITTED SYSTEM CONTROLLER EVENT LOG FILE CONTAINED EVENTS FROM (B)(6) 2025 ¿ (B)(6) 2025. A CONTINUOUS LOW FLOW HAZARD ALARM WAS CAPTURED BEGINNING AT 01:46:26 ON (B)(6) 2025, FOLLOWED BY A PERIOD OF ZERO FLOW FROM APPROXIMATELY 01:57:34 THROUGH THE END OF THE FILE AT 17:43:31. AT 10:43:31, BOTH POWER CABLES WERE DISCONNECTED RESULTING IN NO EXTERNAL POWER ALARMS. THIS WAS FOLLOWED BY THE DISCONNECTION OF THE DRIVELINE AT 10:56:50, AND THE SYSTEM CONTROLLER WAS SHUTDOWN AT 10:57:00. THESE EVENTS APPEAR CONSISTENT WITH THE DISCONTINUATION OF SUPPORT AFTER THE PATIENT'S DEATH. NO OTHER NOTABLE EVENTS OR ALARMS WERE CAPTURED, AND THE SYSTEM APPEARED TO BE OPERATING AS INTENDED PRIOR TO THE DRIVELINE DISCONNECT. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. THIS SECTION ALSO PROVIDES AN EXPLANATION OF PUMP PARAMETERS, INCLUDING FLOW. THIS SECTION EXPLAINS THAT PUMP FLOW IS A CALCULATED VALUE THAT IS ESTIMATED BASED ON PUMP POWER. SECTION 4 OF THE IFU, ¿SYSTEM MONITOR¿, PROVIDES INFORMATION ABOUT THE PUMP FLOW DISPLAY AND THE LOW FLOW HAZARD ALARM CONDITION. THIS SECTION STATES THAT THE LOW FLOW HAZARD ALARM WILL BE TRIGGERED WHEN THE ESTIMATED PUMP FLOW IS LESS THAN 2.5 LPM AND EXPLAINS THAT CHANGES IN PATIENT CONDITIONS CAN RESULT IN LOW FLOW. SECTION 7 OF THE IFU, ¿ALARMS AND TROUBLESHOOTING¿, AND SECTION 5 OF THE PATIENT HANDBOOK, "ALARMS AND TROUBLESHOOTING" OUTLINE SYSTEM CONTROLLER ALARMS, AS WELL AS HOW TO RESPOND TO EACH ALARM CONDITION. THESE DOCUMENTS INSTRUCT THE USER THAT IN THE EVENT OF A LOW FLOW HAZARD ALARM, CALL YOUR HOSPITAL CONTACT IMMEDIATELY FOR DIAGNOSIS AND INSTRUCTIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21585212&pc=LFL | 3005075853-2025-01936 | 2024-04-19 04:00 | Death | ETHICON ENDO-SURGERY, LLC. | 2025-03-12 04:00 | LFL | HARMONIC ACE 5MM SHEAR | Adverse Event Without Identified Device or Use Problem; Patient Device Interaction Problem | Hemorrhage/Bleeding | K120729 |  | 1 | Event Description: TITLE: A COMPARATIVE STUDY BETWEEN LAPAROSCOPIC SPLENECTOMY VERSUS OPEN SPLENECTOMY. AUTHOR:1JAGDISH PRASAD, 2PARVINDER SINGH SANDHU. CITATION: ONLINE ISSN: 2250-3137. PRINT ISSN: 2977-0122. THE STUDY WAS A RANDOMIZED COMPARATIVE STUDY CONDUCTED IN THE DEPARTMENT OF GENERAL SURGERY OF SAWAI MAN SINGH HOSPITAL, JAIPUR. INITIALLY, SPLENECTOMY WAS DONE BY OPEN TECHNIQUE, INDICATIONS BEING MAINLY HAEMATOLOGICAL DISORDERS. CHRONIC ITP. WITH GROWING EXPERIENCE AND ADVANCEMENT IN TECHNOLOGY THE LAPAROSCOPIC APPROACH EXTENDED TO CYSTS,TUMOURS, STORAGE DISEASES AND EVEN TRAUMA IT WAS ALSO USED FOR HAEMATOLOGICAL DISORDERS MAINLY CHRONIC ITP. WITH GROWING EXPERIENCE AND ADVANCEMENT IN TECHNOLOGY THE LAPAROSCOPIC APPROACH EXTENDED TO CYSTS, TUMORS, STORAGE DISEASES AND EVEN TRAUMA. BETWEEN JUNE 2011 TO JUNE 2012. TOTAL 25 PATIENTS WERE INCLUDED IN STUDY PROSPECTIVELY INTO TWO GROUPS, GROUP A & GROUP B. GROUP A(N= 12 AND N=13) CASES IN GROUP B RESPECTIVELY. THE PATIENTS IN GROUP A UNDERWENT LAPAROSCOPIC SPLENECTOMY. WHILE GROUP B UNDERWENT OPEN SPLENECTOMY. THE MEAN AGE IN GROUP A WAS 35.41 YEARS WHILE IN GROUP B WAS 29.53 YEARS. SPLENIC SPAN OF 11-20 CM WAS FOUND IN 33% PATIENTS IN GROUP A AND 46% OF PATIENTS IN GROUP B .MINIMUM PREOPERATIVE PLATELET COUNT IN GROUP A WAS 7000/L IN A CASE OF I.T.P. AND THAT IN GROUP B WAS 30000/L. ACCESSORY SPLEEN WAS DETECTED IN 1 OF THE PATIENT (8%) IN LAPAROSCOPIC GROUP WHEREAS NONE OF THE PATIENT WAS FOUND TO BE HAVING ACCESSORY SPLEEN IN OPEN GROUP. GROUP A DEVELOPED STEROID TOXICITY IN THE FORM OF HYPERTENSION, DIABETES AND PAN GASTRITIS .THE OPERATIVE TIME IN GROUP A WAS IN THE RANGE OF 90-360 MINUTES WITH MEAN OPERATING TIME OF 18. MEAN INTRA OPERATIVE BLOOD LOSS IN GROUP A WAS 60 ML ±38 ML AND FOUND TO BE LESS THAN GROUP B MEAN DRAIN OUTPUT WAS 134ML± 145ML IN GROUP A WITH RANGE OF 50-550ML IN GROUP-A MEAN FOLLOW UP PERIOD WAS 9 MONTHS WITH A RANGE OF 4-18 MONTHS. IN GROUP-B MEAN FOLLOW UP PERIOD WAS 7 MONTHS WITH A RANGE OF 1-18 MONTHS. COMPLICATION DURING THE PERIOD OF FOLLOW UP WAS FOUND IN 17% OF PATIENTS IN GROUP-A AND 15% IN GROUP B. ENDOSTAPLER (ENDOLINEAR CUTTER ATS FLEX 45MM; ECHELON 60M.) WAS USED IN GROUP-A HILAR CONTROL WAS ACHIEVED. ETHILON 3-0 SUTURES WAS USED WHEN THE SKIN SUTURES WERE APPLIED. HARMONIC SCALPEL USED WHEN THE PERITONEUM OVERLYING THE SPLENIC VESSELS WAS DIVIDED REPORTED COMPLICATIONS: HARMONIC DEVICE (ETHICON ENDO-SURGERY). DEATH ON SECOND POST-OPERATIVE DAY DUE TO INTRA ABDOMINAL HEMORRHAGE LEADING TO SHOCK WHICH WAS SEEN DURING RE-EXPLORATION AS DIFFUSE PERITONEAL BLEEDING (N=1). IN CONCLUSION, LAPAROSCOPIC SPLENECTOMY IS SAFE AND EFFICACIOUS FOR MILD TO MODERATE SIZE SPLEENS. THE OPERATIVE TIME FOR LAPAROSCOPIC SPLENECTOMY IS LONGER THAN OPEN SPLENECTOMY BUT AFTER ACHIEVING LEARNING CURVE IT CAN BE SHORTENED SIGNIFICANTLY. Manufacturer Narrative: (B)(4). DATE SENT: 3/12/2025. B3: PUBLICATION YEAR OF 2024. D4: BATCH # UNK. THIS REPORT IS RELATED TO A JOURNAL ARTICLE; THEREFORE, NO PRODUCT WILL BE RETURNED FOR ANALYSIS AND THE MANUFACTURING RECORDS CANNOT BE REVIEWED AS THE LOT/BATCH NUMBER HAS NOT BEEN PROVIDED. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. DOES THE AUTHOR/SURGEON BELIEVE THAT THE ETHICON DEVICE CAUSED OR CONTRIBUTED TO THE PATIENT COMPLICATIONS AND/OR DEATH MENTIONED IN THE ARTICLE? IF YES, PLEASE EXPLAIN IS THE AUTHOR/SURGEON INTERESTED IN SPEAKING WITH AN ETHICON MEDICAL PROFESSIONAL? THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21585372&pc=DSQ | 2916596-2025-01582 | 2025-02-23 05:00 | Death | THORATEC CORPORATION | 2025-03-12 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Pumping Stopped; Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Vomiting; Unspecified Gastrointestinal Problem | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO SUSPECTED ACUTE GASTROINTESTINAL (GI) PROBLEMS THAT LED TO SEVERE VOMITING AND CARDIOPULMONARY ARREST BEFORE THE HOSPITAL. THE OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. ACCORDING TO THE DOCTOR'S DESCRIPTION, THE PATIENT'S FAMILY MEMBERS CONFIRMED AT HOME THAT THE DEVICE WAS STILL FUNCTIONING NORMALLY. IT WAS NOTED THAT THE DEVICE HAD POWER-RELATED ALARMS AND THAT THE PUMP WAS RESTARTED, CAUSING THE DATE AND TIME TO BE RESET. LOG FILES CAPTURED SYSTEM CONTROLLER CLOCK CORRUPT EVENTS. LOG FILES ALSO CAPTURED ONE PUMP STOP. ONCE POWER WAS RESTORED THE PUMP RETURNED TO OPERATING AT THE SET SPEED. THE LAST GOOD TIME STAMP IN THE EVENT LOG FILE WAS (B)(6) 2025 AT 13:35:25. THE LAST GOOD TIME STAMP IN THE PERIODIC LOG FILE WAS (B)(6) 2025 AT 12:55:17. THE FILES SHOWED THE CLOCK WAS NEVER RESET. THE LOSS OF POWER OCCURRED 3 DAYS PRIOR TO LOG FILE REVIEW. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. FURTHERMORE, REVIEW OF THE SUBMITTED LOG FILES CONFIRMED A PUMP STOP DUE TO A NO EXTERNAL POWER EVENT FOLLOWED BY FULL DEPLETION OF THE BACKUP BATTERY. THE SUBMITTED CONTROLLER EVENT LOG FILE CONTAINED EVENTS ON 23FEB2025 FROM 12:31:14 - 13:35:25, UNTIL THE CONTROLLER CLOCK WAS RESET TO THE DEFAULT TIMESTAMP. TWO DOUBLE POWER CABLE DISCONNECT EVENTS WERE CAPTURED ON 23FEB2025, RESULTING IN NO EXTERNAL POWER ALARMS FOLLOWED BY THE SYSTEM OPERATING ON THE CONTROLLER'S INTERNAL BACKUP BATTERY. THE FIRST INSTANCE OF NO EXTERNAL POWER ALARMS RESOLVED ONCE EXTERNAL POWER WAS RESTORED TO THE SYSTEM. THE SECOND INSTANCE RESULTED IN THE PUMP ULTIMATELY STOPPING AT 13:35:20 DUE TO THE BACKUP BATTERY BECOMING FULLY DEPLETED. THE CAUSE OF THE BACKUP BATTERY DEPLETION CANNOT BE CONCLUSIVELY DETERMINED, AS IT IS UNKNOWN IF THE PATIENT WAS IN A CONDITION TO RESPOND TO THE CONTROLLER ALARMS LEADING UP TO THE PUMP STOP. THE PUMP REMAINED OFF FOR AN UNKNOWN AMOUNT OF TIME, AND RESTARTED AFTER THE SYSTEM CONTROLLER WAS RE-CONNECTED TO EXTERNAL BATTERY POWER AND THE TIME CLOCK HAD SUBSEQUENTLY RESET TO THE DEFAULT TIMESTAMP OF (B)(6) 2000 AT 00:00:00. THE PUMP THEN RAMPED UP TO THE STORED PATIENT SPEED WITHOUT ISSUE, HOWEVER FLOW APPEARED TO STRUGGLE TO REMAIN ABOVE THE LOW FLOW ALARM THRESHOLD OF 2.5 LITERS PER MINUTE (LPM). AT THE TIMESTAMP OF 00:08:32, A SUSTAINED LOW FLOW ALARM ACTIVATED FOR THE REMAINDER OF THE FILE, AND FLOW BEGAN TO GRADUALLY DECREASE UNTIL REACHING 0.4 LPM AT 03:32:48 WHEN THE DRIVELINE WAS DISCONNECTED AND THE PUMP SUBSEQUENTLY STOPPED. THE DISCONNECTION OF THE DRIVELINE APPEARS CONSISTENT WITH THE DISCONTINUATION OF SUPPORT DUE TO THE PATIENT'S EXPIRATION. NO OTHER NOTABLE EVENTS OR ALARMS WERE CAPTURED, AND THE PUMP APPEARED TO FUNCTION AS INTENDED AT THE STORED PATIENT SPEED OUTSIDE THE PUMP STOP AND PRIOR TO THE DRIVELINE DISCONNECT. IT WAS ADDITIONALLY REPORTED THAT THE FAMILY MEMBERS DECLINED ANY CLINICAL EXAMINATION, AND NO FURTHER INFORMATION COULD BE PROVIDED. IT WAS COMMUNICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. B, AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK, REV. B, ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 2 OF THE IFU, ¿SYSTEM OPERATIONS¿ EXPLAINS THAT AT LEAST ONE SYSTEM CONTROLLER POWER CABLE MUST BE CONNECTED TO A POWER SOURCE (POWER MODULE OR TWO HEARTMATE 14 VOLT LITHIUM-ION BATTERIES) AT ALL TIMES. THIS SECTION STATES THAT THE 11 VOLT LITHIUM-ION BACKUP BATTERY SHOULD BE USED ONLY FOR TEMPORARY SUPPORT DURING A POWER-LOSS EMERGENCY. THE 11 VOLT LITHIUM-ION BACKUP BATTERY INSIDE THE HEARTMATE 3 SYSTEM CONTROLLER PROVIDES ENOUGH POWER TO RUN THE IMPLANTED HEARTMATE 3 PUMP FOR AT LEAST 15 MINUTES IF THE MAIN POWER SOURCE (EITHER THE POWER MODULE, MOBILE POWER UNIT (MPU), OR TWO HEARTMATE 14 VOLT LITHIUM-ION BATTERIES) IS DISCONNECTED OR FAILS. INAPPROPRIATE USE OF THE 11 VOLT LITHIUM-ION BACKUP BATTERY MAY RESULT IN DIMINISHED RUN TIME DURING A POWER-LOSS EMERGENCY. ADDITIONALLY, EVERY HEARTMATE 14 VOLT LITHIUM-ION BATTERY ALSO HAS ITS OWN ON-BATTERY GAUGE. IT SHOWS THE POWER LEVEL FOR THAT BATTERY. 4 GREEN BARS = 75¿100% OF BATTERY POWER REMAINS. 3 GREEN BARS = 50¿75% OF BATTERY POWER REMAINS. 2 GREEN BARS = 25¿50% OF BATTERY POWER REMAINS. 1 GREEN BAR = LESS THAN 25% OF BATTERY POWER REMAINS. SECTION 3 OF THE IFU, ¿POWERING THE SYSTEM¿ (UNDER ¿USING HEARTMATE 14 VOLT LITHIUM-ION BATTERIES¿) EXPLAINS HOW TO REPLACE DEPLETED BATTERIES WITH CHARGED BATTERIES. THIS SECTION (UNDER ¿SWITCHING POWER SOURCES¿) ALSO EXPLAINS HOW TO SWITCH FROM BATTERY POWER TO THE POWER MODULE AND MPU. SECTION 3 AND SECTION 6, ¿PATIENT CARE AND MANAGEMENT¿, WARN THAT THE PATIENT MUST ALWAYS CONNECT TO THE POWER MODULE OR MOBILE POWER UNIT FOR SLEEPING OR WHEN THERE IS A CHANCE OF SLEEP. A SLEEPING PATIENT MAY NOT HEAR SYSTEM CONTROLLER ALARMS. THESE SECTIONS ALSO WARN NOT TO USE BATTERIES TO POWER THE SYSTEM WHEN THE PATIENT IS SLEEPING. FURTHERMORE, SECTION 5 OF THE PATIENT HANDBOOK, "ALARMS AND TROUBLESHOOTING", AND SECTION 7 OF THE IFU, "ALARMS AND TROUBLESHOOTING", ADDRESS SYSTEM ALARM CONDITIONS AS WELL AS THE APPROPRIATE ACTIONS ASSOCIATED WITH EACH CONDITION. SECTION 8 OF THE PATIENT HANDBOOK, ¿HANDLING EMERGENCIES,¿ ALSO PROVIDES EXAMPLES OF EMERGENCIES AND THE PROPER ACTIONS TO TAKE IN THE EVENT AN EMERGENCY OCCURS. THE PATIENT HANDBOOK ALSO INSTRUCTS USERS TO CALL THEIR HOSPITAL CONTACTS IF THEY THINK, FOR ANY REASON, ANY PORTION OF THEIR EQUIPMENT IS NOT FUNCTIONING AS USUAL, IS BROKEN, OR THEY ARE UNCOMFORTABLE WITH THE OPERATION OF THE EQUIPMENT. SECTION 3 OF THE PATIENT HANDBOOK, ¿POWERING THE SYSTEM¿ DETAILS HOW TO CHECK A BATTERY'S CHARGE LEVEL, REPLACING LOW BATTERIES WITH FULLY CHARGED BATTERIES, AND SWITCHING POWER SOURCES. TWO, NEW, FULLY CHARGED LI-ION BATTERIES PROVIDE 17 HOURS OF SUPPORT. BATTERIES LAST FOR LESS TIME IF THE USER IS ACTIVE OR EMOTIONALLY STRESSED. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21585415&pc=DSQ | 2916596-2025-01393 | 2023-07-05 04:00 | Death | THORATEC CORPORATION | 2025-03-12 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. ADDITIONAL INFORMATION WAS NOT PROVIDED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE CAUSE OF DEATH WAS NOT REPORTED, AND THE DEVICE WAS NOT RETURNED FOR EVALUATION. MULTIPLE REQUESTS FOR ADDITIONAL INFORMATION REGARDING THIS EVENT WERE SUBMITTED TO THE ACCOUNT; HOWEVER, NO RESPONSE HAS BEEN RECEIVED AT THIS TIME. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C IS CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21585558&pc=DSQ | 2916596-2025-01394 | 2025-02-19 05:00 | Death | THORATEC CORPORATION | 2025-03-12 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY AND THE PUMP WOULD BE RETURNING FOR EVALUATION. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE EVALUATION OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. (B)(6) WAS RETURNED ASSEMBLED WITH THE PUMP CABLE INTACT. THE MODULAR CABLE WAS ALSO RETURNED. THE OUTFLOW GRAFT WAS RETURNED ATTACHED TO THE PUMP COVER OUTLET PORT WITH THE OUTFLOW GRAFT BEND RELIEF ENGAGED TO THE GRAFT ATTACHMENT HARDWARE. THE APICAL CUFF WAS RETURNED SECURED WITH THE CUFF LOCK FULLY ENGAGED; HOWEVER, THE APICAL CUFF FELT WAS NOT RETURNED. UPON DISASSEMBLY OF THE RETURNED PUMP, EXAMINATION OF THE BLOOD-CONTACTING SURFACES REVEALED NO EVIDENCE OF ADHERED DEPOSITIONS OR THROMBUS FORMATIONS THAT WOULD HAVE CONTRIBUTED TO A FLOW OR FUNCTIONAL ISSUE. VISUAL INSPECTION OF THE PUMP ROTOR AND ROTOR WELL DID NOT REVEAL ANY OBVIOUS SURFACE SCRATCHES OR DEFECTS. THE LEFT VENTRICULAR ASSIST DEVICE EVENT AND PERIODIC LOG FILES RETRIEVED FROM THE RETURNED DEVICE CAPTURED THE DEVICE FUNCTIONING AS INTENDED. (B)(6) WAS CLEANED, REBUILT, AND FUNCTIONALLY TESTED ON A MOCK CIRCULATORY LOOP. THE DEVICE OPERATED AS INTENDED AND IN ACCORDANCE WITH MANUFACTURING SPECIFICATIONS. INCIDENTAL FINDINGS: VISUAL INSPECTION OF THE RETURNED DEVICE REVEALED A TWIST IN THE OUTFLOW GRAFT. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REV. D AND THE HEARTMATE 3 PATIENT HANDBOOK, REV. D ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 5 OF THE IFU, ¿SURGICAL PROCEDURES", PROVIDES INSTRUCTIONS REGARDING THE PREPARATION, INSTALLATION, AND ORIENTATION OF THE SEALED OUTFLOW GRAFT. THIS SECTION FURTHER EXPLAINS THAT PRIOR TO IMPLANTATION, THE BEND RELIEF SHOULD BE DISENGAGED FROM THE GRAFT FOR THE DE-AIRING PROCEDURE. SECTION 5 ALSO CONTAINS A SUB-SECTION ON "ATTACHING THE SEALED OUTFLOW GRAFT TO THE AORTA", WHICH INSTRUCTS THE USER TO STRETCH THE GRAFT COMPLETELY AND THEN MEASURE AND CUT THE SEALED OUTFLOW GRAFT TO THE APPROPRIATE LENGTH. SECTION 5 FURTHER CAUTIONS THAT THE SEALED OUTFLOW GRAFT MUST NOT BE KINKED OR POSITIONED WHERE IT COULD ABRADE AGAINST A PUMP COMPONENT OR BODY STRUCTURE. SECTION 5, UNDER "ATTACHING THE SEALED OUTFLOW GRAFT TO THE PUMP", INSTRUCTS THE USER TO VERIFY THAT THE GRAFT IS NOT TWISTED OR KINKED BY CHECKING THE POSITION OF THE BLACK LINE ON THE GRAFT ABOVE AND BELOW THE BEND RELIEF AND ENSURING THAT THE LINE IS STRAIGHT. SECTION 5, UNDER "DE-AIRING THE PUMP", CAUTIONS THE USER: "DO NOT ROTATE/TWIST THE SEALED GRAFT. CHECK THE ALIGNMENT OF THE BLACK LINE ON THE GRAFT TO VERIFY THAT THE SEALED GRAFT IS NOT TWISTED OR KINKED." THIS SECTION ALSO EXPLAINS HOW TO ATTACH THE BEND RELIEF ONCE THE VENT NEEDLE HAS BEEN REMOVED FROM THE SEALED OUTFLOW GRAFT AND LEAKS HAVE BEEN RULED OUT. SECTION 5 OF THE IFU ALSO CAUTIONS THE USER THAT SHARP BENDS, TWISTS, OR KINKS IN THE DRIVELINE MAY MAKE IT MORE SUSCEPTIBLE TO WEAR AND FATIGUE OVER TIME. SECTION 6, "PATIENT CARE AND MANAGEMENT", CAUTIONS THE USER TO AVOID PULLING ON OR MOVING THE DRIVELINE. THIS SECTION FURTHER CAUTIONS: "DO NOT TWIST, KINK, OR SHARPLY BEND THE DRIVELINE, SYSTEM CONTROLLER POWER CABLES, THE POWER MODULE PATIENT CABLE, OR THE MOBILE POWER UNIT PATIENT CABLE, WHICH MAY CAUSE DAMAGE TO THE WIRES INSIDE, EVEN IF EXTERNAL DAMAGE IS NOT VISIBLE. DAMAGE TO THE DRIVELINE OR CABLES COULD CAUSE THE LEFT VENTRICULAR ASSIST DEVICE TO STOP. IF THE DRIVELINE OR CABLES BECOME TWISTED, KINKED, OR BENT, CAREFULLY UNRAVEL AND STRAIGHTEN." THE IFU AND PATIENT HANDBOOK ALSO CONTAIN INFORMATION ON HOW TO CLEAN AND CARE FOR THE DRIVELINE. BOTH DOCUMENTS INSTRUCT THE USER TO KEEP THE DRIVELINE CLEAN AND CHECK THE DRIVELINE DAILY FOR SIGNS OF DAMAGE, SUCH AS CUTS, HOLES, OR TEARS. THE PATIENT HANDBOOK ALSO CAUTIONS USERS TO CALL THEIR HOSPITAL CONTACTS IF THEY THINK THE DRIVELINE IS DAMAGED (OR MIGHT BE DAMAGED). THE IFU AND THE PATIENT HANDBOOK PROVIDE ADDITIONAL INSTRUCTIONS REGARDING DRIVELINE CARE AND INFORM THE USER TO CLEAN EXTERIOR SURFACES OF THE DRIVELINE CABLES WITH A DAMP, LINT-FREE CLOTH AND IF MORE AGGRESSIVE CLEANING IS NEEDED, TO USE WARM WATER AND MILD DISH SOAP. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21586101&pc=DSI | 9610816-2025-000186 | 2025-02-13 05:00 | Death | PHILIPS MEDIZIN SYSTEME BÖBLINGEN GMBH | 2025-03-12 04:00 | DSI | INTELLIVUE MP70 | No Audible Alarm | Bradycardia; Sepsis; Low Oxygen Saturation | K151681 |  | 1 | Event Description: IT WAS REPORTED THE MONITOR IN THE NICU DID NOT GENERATE A DESATURATION ALARM. THE CUSTOMER INDICATED THE PATIENT WAS FOUND WITH AN SPO2 OF 14% AND NO ALARM WAS GENERATED. AFTER CHECKING THE PATIENT DATA, IT WAS DETERMINED THE PATIENT HAD BEEN DESATURATING FOR 30 MINUTES WITH NO ALARM, WITH ALARM LIMITS SET AT 85% FOR YELLOW AND 80% FOR RED ALARMS FOR PREMATURE PATIENTS. THE PATIENT WAS FOUND WITH A GRAY COMPLEXION AND BRADYCARDIA. THE PATIENT WAS STIMULATED AND UNTIL THEY RECOVERED THEIR OXYGEN SATURATION. THE PATIENT DIED THE NEXT DAY DUE TO SEPSIS SECONDARY TO ENTEROCOLITIS. Manufacturer Narrative: PATIENT INVOLVEMENT WAS REPORTED, THE PATIENT WAS FOUND WITH A GRAY COMPLEXION AND BRADYCARDIA. THE PATIENT WAS STIMULATED UNTIL THEY RECOVERED THEIR OXYGEN SATURATION. THE PATIENT DIED THE NEXT DAY DUE TO SEPSIS SECONDARY TO ENTEROCOLITIS. THE RSE REVIEWED THE LOGS PROVIDED AND DETERMINED THAT THE PATIENT ENTERED THE ALARM CONDITION AT 2:49 A.M (B)(6) 2025, FOR WHICH THE DEVICE GENERATED AN ALARM WHICH WAS ACKNOWLEDGED. AS THE ALARM WAS REPEATED, ACKNOWLEDGEMENTS OCCURRED. THE PATIENT DID NOT GET OUT OF THE ALARM CONDITION. THIS CAN BE SEEN ON THE TREND CURVES. THIS IS WHY THE ALARM REVIEW DOES NOT INDICATE NEW ALARMS BUT INDICATES THE ALARM REPETITION ACKNOWLEDGEMENTS. IT IS STILL THE SAME ALARM. NO NEW ALARMS ARE CAPTURED IN THE PATIENT DATA AS THE ORIGINAL ALARM CONDITION NEVER RESOLVED. THEREFORE, THE DEVICE FUNCTIONED AS INTENDED. A PHILIPS PRODUCT SUPPORT ENGINEER (PSE) REVIEWED THE INFORMATION AND LOGS PROVIDED. THE PSE STATED THAT THE INFORMATION CONFIRMS THE DETERMINATION BY THE RSE. A REVIEW OF THE LOGS CONFIRMED THE DEVICE FUNCTIONED AS INTENDED. THE DEVICE GENERATED AN ALARM WHICH WAS ACKNOWLEDGED. THE ALARM WAS REPEATED AND ACKNOWLEDGEMENTS OCCURRED. ATTEMPTS WERE MADE TO CLARIFY SOME ADDITIONAL CLINICAL INFORMATION, BUT THE CUSTOMER HAS NOT RESPONDED TO ANY REQUESTS. Manufacturer Narrative: IT WAS REPORTED PER THE LOGS THERE WAS A DESAT ALARM AT 2:49 A.M. FOLLOWED BY THE ALARM ACKNOWLEDGMENT. THE BIOMED TESTED THE ALARMS IN DEMO MODE AND SURVEILLANCE MODE AND IT WORKED CORRECTLY. PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE. THE MANUFACTURING DATE FOR THE REPORTED DEVICE IS PRIOR TO 24SEPT2016 SO NO UDI REQUIRED. E1: REPORTING INSTITUTION PHONE #: (B)(6). E1: REPORTER PHONE #: (B)(6).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21586175&pc=CAW | 2515872-2025-00008 | 2025-01-02 05:00 | Death | DEVILBISS HEALTHCARE | 2025-03-12 04:00 | CAW | DEVILBISS 525 OXYGEN CONCENTRATOR | Insufficient Information | Burn(s) | K071397 |  | 1 | Event Description: DEVILBISS HEALTHCARE WAS NOTIFIED OF AN INCIDENT INVOLVING AN OXYGEN CONCENTRATOR BY THE LOCAL AUTHORITIES REGARDING A FIRE THAT OCCURRED IN AN APARTMENT OF A NURSING HOME. THE RESIDENT WAS FOUND TO HAVE DIED IN THE FIRE. THE END USER WAS A KNOWN SMOKER. DRIVE DEVILBISS HEALTHCARE IS CURRENTLY INVESTIGATING THE INCIDENT, INCLUDING ATTEMPTING TO INSPECT THE DEVICE IN COOPERATION WITH THE LOCAL AUTHORITIES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21586532&pc=DSQ | 2916596-2025-01506 | 2025-02-24 05:00 | Death | THORATEC CORPORATION | 2025-03-12 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Congenital Defect/Deformity; Hypoxia; Heart Failure/Congestive Heart Failure; Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2025.Event Description: THE CAUSE OF DEATH WAS CONGESTIVE HEART FAILURE. THE PATIENT WAS MANAGED WITH A RESPIRATOR, BUT IT WAS BELIEVED THAT HYPOXIA CONTINUED DUE TO A VENTRICULAR SEPTAL DEFECT. THE DEATH WAS NOT CONSIDERED DEVICE RELATED AND THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. (B)(6) WAS RETURNED ASSEMBLED WITH THE PUMP CABLE AND THE MODULAR CABLE CONNECTED TO THE DISTAL PORTION OF THE PUMP CABLE VIA THE INLINE CONNECTOR. THE OUTFLOW GRAFT WAS RETURNED ATTACHED TO THE PUMP COVER OUTLET PORT WITH THE OUTFLOW GRAFT BEND RELIEF HARDWARE ENGAGED TO THE GRAFT ATTACHMENT HARDWARE. THE MINI APICAL CUFF WAS RETURNED SECURED WITH THE CUFF LOCK FULLY ENGAGED. UPON DISASSEMBLY OF THE RETURNED DEVICE, EXAMINATION OF THE BLOOD-CONTACTING SURFACES REVEALED NO EVIDENCE OF DEVELOPED OR ADHERED DEPOSITIONS OR THROMBUS FORMATIONS. VISUAL INSPECTION OF THE PUMP ROTOR AND ROTOR WELL DID NOT REVEAL ANY OBVIOUS SURFACE SCRATCHES OR DEFECTS. THE RETRIEVED CONTROLLER EVENT LOG FILES FROM THE RETURNED DEVICE CAPTURED RELEVANT EVENTS FROM 17JAN2025 THROUGH 24FEB2025. THE FILE CAPTURED A SUSTAINED LOW FLOW HAZARD ALARM ON 24FEB2025 THAT REMAINED ACTIVE UNTIL THE PUMP WAS ULTIMATELY STOPPED. THE DRIVELINE WAS DISCONNECTED ON 24FEB2025, AND THE CONTROLLER WAS LATER SHUTDOWN. THESE EVENTS WERE CONSISTENT WITH THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME. (B)(6) WAS CLEANED, REBUILT, AND FUNCTIONALLY TESTED ON A MOCK CIRCULATORY LOOP. THE DEVICE OPERATED AS INTENDED AND IN ACCORDANCE WITH MANUFACTURING SPECIFICATIONS. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. THE IFU LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE IFU AND PATIENT HANDBOOK PROVIDE INFORMATION REGARDING HOW TO CLEAN AND CARE FOR THE DRIVELINE. THESE DOCUMENTS INSTRUCT THE USER TO INSPECT THE DRIVELINE DAILY FOR SIGNS OF DAMAGE. FURTHERMORE, THE IFU AND THE PATIENT HANDBOOK INSTRUCT THE USER TO CLEAN EXTERIOR SURFACES OF THE DRIVELINE CABLES WITH A DAMP, LINT-FREE CLOTH AND IF MORE AGGRESSIVE CLEANING IS NEEDED, TO USE WARM WATER AND MILD DISH SOAP. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21586569&pc=FMI | 3006260740-2025-01436 | 2025-02-20 05:00 | Death | C.R. BARD, INC. (BASD) -3006260740 | 2025-03-12 04:00 | FMI | INTRAOSSEOUS, INTRAOSSEOUS NEEDLE KITS | Material Twisted/Bent | Cardiac Arrest | K203193 |  | 1 | Event Description: IT WAS REPORTED BY THE CUSTOMER THAT WHILE USING THE 55MM NEEDLE IT BENT. THE PATIENT WAS A FULL CODE AND REQUIRED A SECONDARY SITE AT THE HUMORAL HEAD THAT WAS UTILIZED WITH SUCCESS. AFTER RESUSCITATION EFFORTS, PT WAS CONFIRMED DEAD ON SCENE FOLLOWING OUR MEDICAL DIRECTOR OFF LINES. NO FURTHER INFORMATION WAS PROVIDED. Manufacturer Narrative: H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. EACH EVENT REPORTED TO BD IS EVALUATED AND INVESTIGATED IN ACCORDANCE WITH OUR COMPLAINT INVESTIGATION PROCEDURES. THE INVESTIGATION PROCESS INCLUDES, BUT IS NOT LIMITED TO, EVALUATION OF THE EVENT DETAILS PROVIDED BY THE COMPLAINT FACILITY, A REVIEW OF COMPLAINT HISTORY, MANUFACTURING RECORDS AND RISK DOCUMENT WHERE APPLICABLE, AND AN EVALUATION OF THE SUBJECT DEVICE WHEN AVAILABLE TO IDENTIFY POTENTIAL CONTRIBUTING FACTORS. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21586596&pc=DSQ | 2916596-2025-01317 | 2025-02-17 05:00 | Death | THORATEC CORPORATION | 2025-03-12 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Sepsis; Multiple Organ Failure; Skin Infection; Drug Resistant Bacterial Infection | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY DUE TO MULTIPLE ORGAN FAILURE AND SEPSIS. POSITRON EMISSION TOMOGRAPHY-COMPUTED TOMOGRAPHY (PET-CT) SHOWED INFECTION AROUND THE PUMP AND ALONG THE DRIVELINE. METHICILLIN-RESISTANT STAPHYLOCOCCUS AUREUS (MRSA) WAS FOUND IN DRIVELINE CULTURE. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: DATE OF DEATH: CORRECTED. B2 - OUTCOMES ATTRIBUTED TO ADVERSE: CORRECTED. B7 - OTHER RELEVANT HISTORY, INCLUDING PREEXISTING MEDICAL CONDITIONS: CORRECTED. H6 HEALTH EFFECT - CLINICAL CODE: CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A SPECIFIC CAUSE FOR THE REPORTED INFECTION COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. ADDITIONALLY, A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REV. B IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RESPIRATORY FAILURE, RIGHT HEART FAILURE, RENAL FAILURE, HEPATIC DYSFUNCTION), INFECTION (LOCAL, DRIVELINE, AND PUMP POCKET), SEPSIS, AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, LISTS INFECTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. FURTHERMORE, SEVERAL SECTIONS OF THE HM3 IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21586604&pc=DSQ | 2916596-2025-01606 | 2025-02-18 05:00 | Death | THORATEC CORPORATION | 2025-03-12 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTIPLE SYSTEM ORGAN FAILURE. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. AN AUTOPSY WAS NOT PERFORMED, AND THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED THROUGH THE MANUFACTURING EXECUTION SYSTEM (MES) AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ OF THIS DOCUMENT LISTS DEATH AS AN ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. THE IFU ALSO LISTS MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RIGHT HEART FAILURE, RESPIRATORY FAILURE, RENAL DYSFUNCTION, AND HEPATIC DYSFUNCTION) AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: A1-A4: PATIENT INFORMATION CANNOT BE PROVIDED DUE TO PERSONAL DATA PRIVACY LEGISLATION/POLICY. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21586646&pc=LOX | 2024168-2025-02589 | 2025-02-17 05:00 | Death | ABBOTT VASCULAR INC. | 2025-03-12 04:00 | LOX | NC TREK | Difficult to Remove; Material Separation; Improper or Incorrect Procedure or Method; Failure to Deflate | Foreign Body In Patient | K110134 |  | 1 | Event Description: IT WAS REPORTED THAT THE PROCEDURE WAS TO TREAT THE LEFT ANTERIOR DESCENDING (LAD) CORONARY ARTERY WITH MILD CALCIFICATION, MILD TORTUOSITY AND 80% STENOSIS. THE 4.0X15 MM NC TREK WAS USED FOR POST-DILATATION AFTER STENT IMPLANTATION. THE BALLOON WAS INFLATED TWICE TO 10-12 ATMOSPHERES BUT WOULD NOT DEFLATE. NEGATIVE PRESSURE WAS HELD FOR 2 MINUTES BUT THE BALLOON COMPLETELY FAILED TO DEFLATE. SEVERAL GUIDE WIRES AND A MICROCATHETER WERE ADVANCED TO ATTEMPT TO REMOVE THE DEVICE AS WELL AS AN ATTEMPT TO ADVANCE A 1X10 MM BALLOON. EVENTUALLY THE BALLOON WAS REMOVED STILL INFLATED TOGETHER WITH THE GUIDING CATHETER AND FORCE WAS APPLIED BUT THE DISTAL END (TIP) OF THE DEVICE AND ITS SHEATH DETACHED AND REMAINED IN THE ANATOMY. THE PATIENT WAS TRANSFERRED TO ANOTHER HOSPITAL FOR BYPASS SURGERY AND THE PATIENT ULTIMATELY EXPIRED ON 02/18/2025. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS; HOWEVER, A PICTURE WAS PROVIDED BY THE ACCOUNT. THE PICTURE SHOWS A SEPARATION ON THE DISTAL SHAFT, CONFIRMING THE REPORTED SEPARATION. PRODUCTION RECORD AND CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A REVIEW OF THE COMPLAINT HISTORY IDENTIFIED NO OTHER SIMILAR COMPLAINTS FROM THIS LOT. A MEDICAL REVIEW WAS CONDUCTED BY AN ABBOTT VASCULAR CLINICAL SPECIALIST. THIS WAS A CASE OF A PATIENT WHO UNDERWENT AN INTERVENTIONAL PROCEDURE ON (B)(6) 2025. AFTER STENT IMPLANTATION IN THE PROXIMAL LAD, A 4.00*15 NC TREK BALLOON WAS INFLATED TO A NORMAL PRESSURE OF 12 ATM. UPON DEFLATION, THE BALLOON COULD NOT BE RE- FOLDED. INCREASING THE PRESSURE AND ATTEMPTING DEFLATION AGAIN, THE BALLOON STILL COULD NOT BE RETRACTED. EVENTUALLY, THE BALLOON WAS WITHDRAWN TOGETHER WITH THE GUIDING CATHETER (GC), BUT THE DISTAL END OF THE BALLOON AND ITS SHEATH REMAINED IN THE PATIENT'S BODY. THE PATIENT WAS TRANSFERRED TO ANOTHER HOSPITAL FOR OPEN-HEART SURGERY AND UNFORTUNATELY EXPIRED 1-DAY POST-PROCEDURE. DURING THE PROCEDURE, THE PRESSURE PUMP OF THE BALLOON EXPERIENCED PRESSURE DECAY, AND THE BALLOON COULD NOT BE RETRACTED. A NEW GUIDE WIRE (A TOTAL OF 4) WAS IMMEDIATELY SENT UNDER THE SUPPORT OF A MICROCATHETER. ATTEMPTS TO RETRACT THE BALLOON WITH THE GUIDE WIRES WERE UNSUCCESSFUL. AN ATTEMPT TO PASS A 1.0*10MM PRE-DILATION BALLOON ALSO FAILED. AFTER SEVERAL REPORTEDLY FORCEFUL ATTEMPTS TO REMOVE THE BALLOON CATHETER WERE MADE, THE ENTIRE SYSTEM WAS REMOVED AS A SINGLE UNIT. HOWEVER, IT APPEARS THAT THE BALLOON FRACTURED AND SEPARATED AT SOME POINT DURING THE RETRIEVAL ATTEMPTS. UPON REMOVAL OF THE ENTIRE SYSTEM TOGETHER, THE DISTAL TIP AND THE SHAFT OF THE BALLOON REPORTEDLY REMAINED IN THE LM ARTERY. FORCE WAS APPLIED UPON REMOVAL CAUSING SEPARATION OF THE NC TREK AND PORTION OF THE BALLOON WAS RETAINED IN THE CORONARY ARTERY. A BALLOON THAT REMAINS INFLATED IN A CORONARY ARTERY WILL COMPLETELY OCCLUDE THE VESSEL AND THIS CAN SOON LEAD TO CRITICAL COMPLICATIONS, SO BALLOON ENTRAPMENT REQUIRES EMERGENCY BAILOUT, UNLIKE OTHER ENTRAPMENT OF OTHER DEVICES. ENTRAPMENT OF BALLOON WITH A FULLY OR PARTIALLY DEPLOYED BALLOON CAN ALSO RESULT IN ACUTE OCCLUSION OF CORONARY ARTERY. THESE CATHETER REMNANTS CAN BE READILY RECOVERED FROM THE CORONARY ARTERIES WITH THE USE OF VARIOUS INTERVENTIONAL DEVICES WITH A HIGH SUCCESS RATE. IN THE EVENT OF FAILED INTERVENTIONAL RETRIEVAL AND PERSISTENT SIGNS OF ISCHEMIA, PATIENTS ARE URGENTLY REFERRED TO EMERGENT CARDIAC SURGERY. THE EMERGENCY BYPASS SURGICAL PROCEDURE IS DETERMINED PRIMARILY BY THE CLINICAL STATUS OF THE PATIENT BEFORE THE OPERATION. FOR PATIENTS IN CARDIOGENIC SHOCK, THE RISK OF DEATH REMAINS HIGH AROUND 20% TO 50%. ADDITIONAL RISK FACTORS THAT CAN LEAD TO COMPLICATIONS DEPEND UPON THE DURATION, EXTENT, AND LOCATION OF ISCHEMIA. WITHOUT THE CINE IMAGES, THE NC TREK DEVICE, OR ADDITIONAL INFORMATION AVAILABLE, IT CANNOT BE DEFINITIVELY STATED THAT THE NC TREK WAS THE DIRECT CAUSE OF DEATH IN THIS PATIENT. HOWEVER, WITH THE INFORMATION PRESENTED, IT APPEARS THAT IT MOST LIKELY CONTRIBUTED TO THE PATIENT'S DEATH. IT WAS REPORTED BY THE ACCOUNT THAT THE BALLOON WAS REMOVED FULLY INFLATED. IT SHOULD BE NOTED THE CORONARY DILATATION CATHETERS (CDC), NC TREK RX, GLOBAL, INSTRUCTION FOR USE (IFU, INSTRUCTIONS FOR USE SECTION) STATES: WITHDRAW THE DEFLATED DILATATION CATHETER AND GUIDE WIRE FROM THE GUIDING CATHETER THROUGH THE HEMOSTATIC VALVE. IN THIS CASE, IT IS LIKELY THAT THE IFU VIOLATION CONTRIBUTED TO THE REPORTED DIFFICULTY REMOVING THE DEVICE. ADDITIONALLY, THE ACCOUNT REPORTED THAT FORCE WAS USED DURING REMOVAL, AS DIFFICULTY WAS EXPERIENCED. IT SHOULD BE NOTED THAT THE CORONARY DILATATION CATHETERS (CDC), NC TREK RX, GLOBAL, INSTRUCTION FOR USE (IFU, WARNING SECTION) STATED: IF RESISTANCE IS FELT, DETERMINE THE CAUSE BEFORE PROCEEDING. CONTINUING TO ADVANCE OR RETRACT THE CATHETER WHILE UNDER RESISTANCE MAY RESULT IN DAMAGE TO THE VESSELS AND / OR DAMAGE / SEPARATION OF THE CATHETER. IN THIS CASE, IT IS LIKELY THAT THE COMBINATION OF THE IFU DEVIATION AND THE REPORTED DIFFICULTY REMOVING THE DEVICE CONTRIBUTED TO THE REPORTED SEPARATION. THE REPORTED PATIENT EFFECT OF DEATH IS LISTED IN THE CORONARY DILATATION CATHETERS (CDC), NC TREK RX, GLOBAL, INSTRUCTIONS FOR USE AS A KNOWN PATIENT EFFECT. BASED ON THE INFORMATION PROVIDED, A DEFINITIVE CAUSE FOR THE REPORTED FAILURE TO DEFLATE COULD NOT BE DETERMINED. POSSIBLE FACTORS THAT MAY CONTRIBUTE TO FAILURE TO DEFLATE THE BALLOON INCLUDE, BUT ARE NOT LIMITED TO, DEFLATION TECHNIQUE, LOOSE CONNECTION WITH THE INDEFLATOR, CONTAMINATION IN THE INFLATION LUMEN OR DAMAGE TO THE GUIDE WIRE AND/OR INFLATION LUMEN. THE REPORTED DIFFICULTY REMOVING THE DEVICE AND SEPARATION APPEAR TO BE RELATED TO OPERATIONAL CONTEXT. IN THIS CASE, IT IS LIKELY THAT SINCE THE BALLOON WOULD NOT DEFLATE, THE BDC BECAME STUCK, RESULTING IN THE REPORTED DIFFICULTY REMOVING THE DEVICE FROM THE GUIDING CATHETER. ADDITIONAL TREATMENT WITH SEVERAL GUIDE WIRES AND A MICROCATHETER WAS ATTEMPTED TO REMOVE THE INFLATED BALLOON BUT WERE UNSUCCESSFUL. UPON FURTHER MANIPULATION OF THE DEVICE, AGAINST RESISTANCE AND WITH FORCE, THE DEVICE ULTIMATELY SEPARATED AND REMAINED IN THE PATIENT ANATOMY. THE REPORTED UNEXPECTED MEDICAL INTERVENTION, FOREIGN BODY IN PATIENT, SURGICAL INTERVENTION AND HOSPITALIZATION APPEAR TO BE RELATED TO CIRCUMSTANCES OF THE PROCEDURE. IT CANNOT BE DEFINITIVELY STATED THAT THE NC TREK WAS THE DIRECT CAUSE OF DEATH IN THIS PATIENT; HOWEVER, WITH THE INFORMATION PRESENTED, IT APPEARS THAT IT MOST LIKELY CONTRIBUTED TO THE PATIENT'S DEATH. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO THE DESIGN, MANUFACTURE, OR LABELING OF THE DEVICE. H6: MEDICAL DEVICE PROBLEM CODE 2017: EXCESSIVE FORCE. H6: MEDICAL DEVICE PROBLEM CODE 2017: FAILURE TO FOLLOW STEPS / INSTRUCTIONS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21586799&pc=BZD | 2518422-2025-102945 | 2022-03-21 04:00 | Death | RESPIRONICS, INC. | 2025-03-12 04:00 | BZD | DREAMSTATION | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Dyspnea; High Blood Pressure/ Hypertension | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION REGARDING A DREAMSTATION DEVICE. THERE IS AN ALLEGATION THAT THE DEVICE WAS NOT WORKING CORRECTLY, AND THE PATIENT WAS HAVING PROBLEMS BREATHING. THE PATIENT WAS TAKEN TO AN EMERGENCY ROOM WHERE THEY PASSED FROM HYPERTENSION AND CARDIAC ARREST. THE PATIENT¿S FAMILY ALSO REPORTED THAT THE UNIT WAS BEING CLEANED WITH AN OZONE-BASED DISINFECTION DEVICE. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21586840&pc=OAE | 2029046-2025-00765 | 2025-02-13 05:00 | Death | BIOSENSE WEBSTER INC | 2025-03-12 04:00 | OAE | QDOT MICRO | Adverse Event Without Identified Device or Use Problem | Cardiogenic Shock | P210027 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A VENTRICULAR TACHYCARDIA (VT) ABLATION AND THE PATIENT EXPERIENCED CARDIOGENETIC SHOCK THAT REQUIRED CARDIOPULMONARY RESUSCITATION (CPR) / EXTERNAL CHEST COMPRESSIONS USING CORPULSE MACHINE. THE PATIENT DIED. THE PATIENT HAD RUNS OF VT, POORLY TOLERATED HEMODYNAMICALLY. PATIENT WENT INTO CARDIOGENIC SHOCK. CPR AND EXTERNAL CHEST COMPRESSIONS USING CORPULSE MACHINE WAS DONE. THE PATIENT DID NOT RECOVER AFTER FORTY-FIVE (45) MINUTES AND THE PATIENT EXPIRED. THERE WERE NO ISSUES WITH THE PRODUCTS OR COMPLICATIONS. THE PHYSICIAN'S OPINION ON THE CAUSE OF DEATH WAS CARDIOGENIC SHOCK DUE TO UNDERLYING DILATED CARDIOMYOPATHY AND VERY LOW EJECTION FRACTION (APPROXIMATELY 10%), AND RUNS OF VT (VENTRICULAR TACHYCARDIA) DURING THE PROCEDURE WHICH WERE HEMODYNAMICALLY NOT TOLERATED. Manufacturer Narrative: D4: UDI: AS THE CATALOG/MODEL NUMBER WAS NOT PROVIDED, THE (01)GTIN IS NOT AVAILABLE. E1. INITIAL REPORTER PHONE: (B)(6). THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. MANUFACTURER'S REFERENCE NUMBER: (B)(4). Manufacturer Narrative: AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. AN EVALUATION OF THE MANUFACTURING RECORD COULD NOT BE PERFORMED AS THE REQUIRED PRODUCT IDENTIFICATION NUMBER WAS NOT PROVIDED TO COMPLETE THE EVALUATION. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. IF ADDITIONAL INFORMATION IS RECEIVED REGARDING THIS EVENT, A SUPPLEMENTAL 3500A REPORT WILL BE SUBMITTED TO THE FDA. MANUFACTURER'S REFERENCE NUMBER: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21587020&pc=NPT | 2025587-2025-01859 | 2025-02-17 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-12 04:00 | NPT | EVOLUT FX DCS | Insufficient Information; Patient Device Interaction Problem | Hemorrhage/Bleeding; Low Blood Pressure/ Hypotension; Perforation; Rupture; Pericardial Effusion | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE GUIDEWIRE PERFORATED THE LEFT VENTRICLE. PERI CARDIAL EFFUSION WAS NOTED AND THE PATIENT'S BLOOD PRESSURE DROPPED RAPIDLY. PERCUTANEOUS CARDIOPULMONARY SUPPORT (PCPS) WAS INTRODUCED, A THORACOTOMY WAS PERFORMED TO REPAIR THE LEFT VENTRICLE, AND THE PERFORATION SITE WAS REPAIRED. SUBSEQUENTLY, AN AORTIC VALVE REPLACEMENT WAS PERFORMED. THE PATIENT LEFT THE ROOM WITH PCPS IN PLACE.Event Description: IT WAS REPORTED THAT DURING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE GUIDEWIRE PERFORATED THE LEFT VENTRICLE. PERI CARDIAL EFFUSION WAS NOTED AND THE PATIENT'S BLOOD PRESSURE DROPPED RAPIDLY. PERCUTANEOUS CARDIOPULMONARY SUPPORT (PCPS) WAS INTRODUCED, A THORACOTOMY WAS PERFORMED TO REPAIR THE LEFT VENTRICLE, AND THE PERFORATION SITE WAS REPAIRED. SUBSEQUENTLY, AN AORTIC VALVE REPLACEMENT WAS PERFORMED. THE PATIENT LEFT THE ROOM WITH PCPS IN PLACE. ADDITIONAL INFORMATION WAS RECEIVED TO REPORT THE GUIDEWIRE USED WAS A NON-MEDTRONIC PRODUCT (MANUFACTURER UNKNOWN). THE MEDTRONIC VALVE WAS EXPLANTED DURING THE THORACOTOMY. THE PHYSICIAN REPORTED IT WAS NOT POSSIBLE TO DETERMINE WHEN THE PERFORATION OCCURRED, WHETHER CAUSED BY THE GUIDEWIRE BEFORE THE VALVE WAS INSERTED OR IF IT OCCURRED DURING VALVE DELIVERY.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT APPROXIMATELY 19 DAYS FOLLOWING THE IMPLANT OF THE VALVE, THE PATIENT DIED. THE CAUSE OF DEATH WAS BLEEDING DUE TO DISSEMINATED INTRAVASCULAR COAGULATION (DIC). PER THE PHYSICIAN, THE DEATH WAS NOT RELATED TO THE VALVE, BUT THE DEATH WAS RELATED TO THE PROCEDURE. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID EVOLUTFX-26 (SERIAL: (B)(6)); PRODUCT TYPE: 0195-HEART VALVES. PRODUCT ID L-EVOLUTFX-2329 (LOT: 0012373646); PRODUCT TYPE: 0195-HEART VALVES. GUIDEWIRE (MANUFACTURER: UNKNOWN); (LOT: UNKNOWN). SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. A SECOND PARAGRAPH WAS ADDED. D. THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: PRODUCT ID: EVOLUTFX-26 (SERIAL: (B)(6)); PRODUCT TYPE: 0195-HEART VALVES. MANUFACTURED DATE: 2024-07-26; USE BY DATE: 2026-07-26; FULL UDI# (B)(4). CORRECTED DATA: D. SUSPECT MEDICAL DEVICE FULL UDI# ADDITIONAL CODES. ANNEX F AND G WERE INADVERTENTLY OMITTED FROM THE INITIAL MEDWATCH REPORT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B2. B5. D4. H1. H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21587618&pc=KDJ | 1416980-2025-01141 |  | Death | VANTIVE US HEALTHCARE LLC | 2025-03-12 04:00 | KDJ | MINICAP WITH POVIDONE-IODINE SOLUTION | Adverse Event Without Identified Device or Use Problem | Unspecified Infection | K152129 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED AN UNSPECIFIED INFECTION AFTER USING BAXTER MINICAPS AND SUBSEQUENTLY PASSED AWAY. THE CAUSE OF THE UNSPECIFIED INFECTION WAS NOT REPORTED. HOSPITALIZATION AND TREATMENT FOR THE UNSPECIFIED INFECTION WAS NOT REPORTED. THE CAUSE OF DEATH WAS NOT REPORTED. IT WAS NOT REPORTED IF AN AUTOPSY WAS PERFORMED. IT WAS NOT REPORTED IF PD THERAPY WAS ONGOING PRIOR TO DEATH OR AT THE TIME OF DEATH. NO ADDITIONAL INFORMATION WAS AVAILABLE. Manufacturer Narrative: ADDITIONAL INFORMATION: H2, H6, AND H11. H11: THE DEVICE WAS NOT RETURNED AND THE LOT NUMBER IS UNKNOWN; THEREFORE, A DEVICE ANALYSIS COULD NOT BE COMPLETED. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. Manufacturer Narrative: SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21587621&pc=DSQ | 2916596-2025-01500 | 2025-02-21 05:00 | Death | THORATEC CORPORATION | 2025-03-12 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Mechanical Problem | Hypoxia; Loss of Vision; Ischemia Stroke; Thrombosis/Thrombus; Respiratory Insufficiency | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED TO THE EMERGENCY DEPARTMENT UNRESPONSIVE. THE PATIENT WAS HYPOXIC HYPERCARBIA, HAD COOL EXTREMITIES AND MOTTLING. THEY TESTED POSITIVE FOR INFLUENZA A. A HEAD COMPUTED TOMOGRAPHY (CT) REVEALED MULTIPLE EMBOLIC STROKES. THE PATIENT ALSO EXPERIENCED PERSISTENT VISION LOSS. THEIR INTERNATIONAL NORMALIZED RATIO HAD BEEN THERAPEUTIC BETWEEN THE GOAL OF 2 AND 3 FOR OVER A YEAR. A REVIEW OF THE LOG FILES REVEALED NO UNUSUAL EVENTS. THERE WERE NO CHANGES TO PATIENT CONDITION OR ANTICOAGULATION STATUS THAT MAY HAVE CONTRIBUTED TO THE EVENT. THE PATIENT WAS INITIALLY STARTED ON A HEPARIN INFUSION. IT WAS NOTED THAT A TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) WAS ATTEMPTED BUT THE PROBE WAS UNABLE TO BE PASSED. A CT ANGIOGRAPHY (CTA) SHOWED THROMBUS IN THE OUTFLOW TRACT WHICH WAS EXPECTED TO BE THE SOURCE OF CLOTS TO BRAIN. THE PATIENT AND FAMILY CHOSE TO DISCONNECT LEFT VENTRICULAR ASSIST DEVICE (LVAD) THERAPY. THE PATIENT PASSED AWAY.Event Description: THE PATIENT PASSED AWAY ON (B)(6) 2025. Manufacturer Narrative: B2 - OUTCOMES ATTRIBUTED TO ADVERSE: CORRECTED. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: REVIEW OF THE SUBMITTED IMAGE FOUND THAT THE OUTFLOW GRAFT LUMEN APPEARED SLIGHTLY REDUCED, WHICH THE CUSTOMER REPORTED WAS INDICATIVE OF OUTFLOW GRAFT THROMBUS. HOWEVER, THROMBUS WAS UNABLE TO BE CONCLUSIVELY CONFIRMED THROUGH THIS EVALUATION AS THE DEVICE WAS NOT RETURNED FOR EVALUATION. A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. A CT IMAGE WAS SUBMITTED WHICH SHOWED A CROSS SECTION OF THE OUTFLOW GRAFT. REVIEW OF THE IMAGE NOTED THAT THE LUMEN OF THE OUTFLOW GRAFT APPEARED SOMEWHAT REDUCED AT THE ANTERIOR PORTION, WITH A POTENTIAL ARTIFACT PRESENT AT THE POSTERIOR. THE CUSTOMER COMMUNICATED THAT THE CT DEMONSTRATED THROMBUS IN THE OUTFLOW TRACT; HOWEVER, THROMBUS WAS UNABLE TO BE CONCLUSIVELY CONFIRMED THROUGH REVIEW OF THE SUBMITTED CT AND THE DEVICE WAS NOT RETURNED FOR EVALUATION. REVIEW OF THE SUBMITTED LOG FILES FOUND THAT THE SYSTEM APPEARED TO BE OPERATING AS INTENDED AT THE SET SPEED, AND THERE WERE NO NOTABLE ALARMS ACTIVE IN THE LOG FILES. IT WAS REPORTED THAT THE HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), WILL NOT BE RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ LISTS PUMP THROMBOSIS, STROKE, AND DEATH AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS. SECTION 5, "SURGICAL PROCEDURES" (UNDER ¿PREPARING THE VENTRICULAR APEX SITE¿), INSTRUCTS TO INSPECT THE VENTRICULAR CHAMBER FOR MURAL THROMBI AND CROSSING TRABECULAE FOLLOWING REMOVAL OF THE CORE AND TO ADDRESS ONE OR BOTH, AS NEEDED. FURTHERMORE, SECTION 5, UNDER ¿IMPLANT PROCEDURES¿, WARNS TO INSPECT THE VENTRICLE AND REMOVE ANY PREVIOUSLY FORMED CLOTS THAT MAY CAUSE EMBOLISM OR ANY TRABECULAE THAT MAY IMPEDE FLOW. SECTION 6 ¿PATIENT CARE AND MANAGEMENT¿ LISTS THROMBOEMBOLISM AS A POTENTIAL LATE POSTIMPLANT COMPLICATION THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. THIS SECTION ALSO CONTAINS INFORMATION REGARDING THE RECOMMENDED ANTICOAGULATION THERAPY AND INR (INTERNATIONAL NORMALIZED RATIO) RANGE. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SECTION B2: DEATH DATE WAS NOT PROVIDED. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21587667&pc=DSQ | 2916596-2025-01446 | 2025-02-21 05:00 | Death | THORATEC CORPORATION | 2025-03-12 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Infusion or Flow Problem | Liver Damage/Dysfunction; Multiple Organ Failure; Heart Failure/Congestive Heart Failure; Lactate Dehydrogenase Increased | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT LOG FILES WERE SUBMITTED FOR EVALUATION. THE PATIENT DEVELOPED VASOPLEGIC SHOCK AND UNDERWENT VENO-ARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION (VA ECMO) PLACEMENT. THE PATIENT'S LACTATE DEHYDROGENASE (LDH) WAS NOTED TO BE 1135. THE EVENT LOG FILE CAPTURED LOW FLOW ALARM EVENTS ON (B)(6) 2025. THERE WERE ALSO SEVERAL MOMENTARY LOW FLOW EVENTS RECORDED. SOME OF THESE EVENTS WERE VERY BRIEF AND DID NOT GENERATE AN ALARM. THESE OCCURRED AT TIMES WHEN PI WAS ELEVATED. THERE DID NOT APPEAR TO BE ANY TYPE OF EXTERNAL MECHANICAL CIRCULATORY SUPPORT (MCS) EQUIPMENT ISSUES THAT CAUSED THESE EVENTS. THE LOW FLOW EVENTS MAY BE DUE TO A PATIENT RELATED ISSUE. LABS TAKEN ON (B)(6) 2025 SHOWED ALANINE AMINOTRANSFERASE (ALT) AT 325, ASPARTATE AMINOTRANSFERASE (AST) AT 697, AND LDH AT 1135. THE CAUSE OF LIVER DYSFUNCTION AND ELEVATED LDH WAS RIGHT VENTRICULAR (RV) FAILURE. THE PATIENT WAS IN POOR CONDITION, AND THEY WERE UNDERGOING FULL SUPPORT.Event Description: ADDITIONAL INFORMATION RECEIVED STATED THAT THE PATIENT PASSED AWAY DUE TO MULTIPLE SYSTEM ORGAN FAILURE (MSOF) ON (B)(6) 2025. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: ANALYSIS OF THE SUBMITTED LOG FILES CONFIRMED THE REPORTED LOW FLOW ALARMS; HOWEVER, A SPECIFIC CAUSE FOR THESE EVENTS COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. ADDITIONALLY, A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE CONTROLLER EVENT LOG FILE CONTAINED EVENTS FROM 21FEB2025 THROUGH 24FEB2025. SEVERAL LOW FLOW HAZARD ALARMS WERE CAPTURED THROUGHOUT THE FILE AND WERE ASSOCIATED WITH ELEVATED PULSATILITY INDEX (PI) VALUES. NO OTHER NOTABLE EVENTS OR ALARMS WERE CAPTURED, AND THE PUMP APPEARED TO OPERATE AS INTENDED AT THE FIXED SPEED. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU AND THE HEARTMATE 3 PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (INCLUDING RIGHT HEART FAILURE AND HEPATIC DYSFUNCTION), HEMOLYSIS, AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THIS SECTION ALSO ADDRESSES ALL PUMP PARAMETERS, INCLUDING PUMP FLOW AND PULSATILITY INDEX (PI). IN REFERENCE TO PI, SECTIONS 1 AND 4 OF THE IFU STATE THAT THE PI CALCULATION REPRESENTS CARDIAC PULSATILITY AND VALUES TYPICALLY RANGE FROM 1 TO 10. IN GENERAL, THE MAGNITUDE OF THE PI VALUE IS RELATED TO THE AMOUNT OF ASSISTANCE PROVIDED BY THE PUMP. HIGHER VALUES INDICATE MORE VENTRICULAR FILLING AND HIGHER PULSATILITY (I.E., THE PUMP IS PROVIDING LESS SUPPORT TO THE LEFT VENTRICLE). SECTION 4 OF THE IFU, ¿SYSTEM MONITOR¿, DESCRIBES THE PUMP FLOW DISPLAY AND THE HAZARD ALARMS AND EXPLAINS THAT CHANGES IN PATIENT CONDITION CAN RESULT IN LOW FLOW. PER DESIGN, WHEN THE ESTIMATED FLOW VALUE IS CALCULATED AT LESS THAN 2.5 LITERS PER MINUTE (LPM), A LOW FLOW STATUS IS POSTED TO THE LOG FILE. IF THE FLOW REMAINS BELOW 2.5 LPM FOR 10 SECONDS, A LOW FLOW HAZARD ALARM IS TRIGGERED. SECTION 6 "PATIENT CARE AND MANAGEMENT" (UNDER "RIGHT HEART FAILURE") DISCUSSES THE POTENTIAL DEVELOPMENT OF RIGHT HEART FAILURE FOLLOWING IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS, INCLUDING INOTROPES AND RIGHT VENTRICULAR ASSIST DEVICE (RVAD) PLACEMENT. SECTION 6, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿ANTICOAGULATION¿), PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) RANGE, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. SECTION 5 OF THE PATIENT HANDBOOK, ¿ALARMS AND TROUBLESHOOTING¿, AND SECTION 7 OF THE IFU, ¿ALARMS AND TROUBLESHOOTING¿, PROVIDES INFORMATION ON ALL SYSTEM ALARM CONDITIONS AS WELL AS THE APPROPRIATE ACTIONS ASSOCIATED WITH EACH CONDITION. FURTHERMORE, SECTION 8 OF THE PATIENT HANDBOOK, ¿HANDLING EMERGENCIES¿, ALSO PROVIDES EXAMPLES OF EMERGENCIES AND THE PROPER ACTIONS TO TAKE IN THE EVENT AN EMERGENCY OCCURS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21588182&pc=NPT | 2135147-2025-01238 | 2025-02-12 05:00 | Death | ABBOTT MEDICAL | 2025-03-12 04:00 | NPT | NAVITOR TRANSCATHETER AORTIC VALVE | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Atrial Fibrillation; Fatigue; Fever; Perforation; Shock; Respiratory Tract Infection | P190023 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 23MM NAVITOR VALVE WAS SUCCESSFULLY IMPLANTED AT A DEPTH OF 3MM USING A SMALL FLEXNAV DELIVERY SYSTEM (DS) WITHOUT ISSUE. THERE WAS NO CALCIFICATION EXTENDING BENEATH THE AORTIC ANNULAR PLANE IN THE INTERVENTRICULAR SEPTUM. ON (B)(6) 2025, THE PATIENT PRESENTED WITH ATRIAL FIBRILLATION (AF), WHICH WAS REVERTED. ON (B)(6) 2025, THE PATIENT¿S HEART RATE WAS ABLE TO BE CONTROLLED. ON (B)(6) 2025, THE PATIENT WAS NOTED TO HAVE A FEVER, LETHARGY, AND ANOTHER OCCURRENCE OF AF. ANTIBIOTICS AND AMIODARONE WERE ADMINISTERED TO TREAT THE ISSUE. ON (B)(6) 2025, THE PATIENT WAS ABLE TO BE STABILIZED. ON (B)(6) 2025, AN ECG WAS REPEATED, SHOWING A WIDER QRS. IT WAS DECIDED TO CONTINUE TREATING THE PATIENT WITH ANTIBIOTICS, AND A PACEMAKER WAS ALSO IMPLANTED. AFTER THIS, THE PATIENT ENDED UP GOING INTO SHOCK AND AN X-RAY INDICATED POSSIBLE VENTRICULAR PERFORATION. THE PATIENT THEN BECAME UNSTABLE, AND A PERICARDIAL PUNCTURE AND DRAIN WAS ATTEMPTED. THE PATIENT¿S HEMODYNAMICS DIDN'T IMPROVE AND THEY PASSED AWAY (B)(6) 2025. THE PHYSICIAN BELIEVES THAT IT WAS LIKELY DUE TO A MULTI-RESISTANT PSEUDOMONAS INFECTION.Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 23MM NAVITOR VALVE WAS SUCCESSFULLY IMPLANTED AT A DEPTH OF 3MM USING A SMALL FLEXNAV DELIVERY SYSTEM (DS) WITHOUT ISSUE. THERE WAS NO CALCIFICATION EXTENDING BENEATH THE AORTIC ANNULAR PLANE IN THE INTERVENTRICULAR SEPTUM. ON (B)(6) 2025, THE PATIENT PRESENTED WITH ATRIAL FIBRILLATION (AF), WHICH WAS REVERTED. ON (B)(6) 2025, THE PATIENT¿S HEART RATE WAS ABLE TO BE CONTROLLED. ON (B)(6) 2025, THE PATIENT WAS NOTED TO HAVE A FEVER, LETHARGY, AND ANOTHER OCCURRENCE OF AF. ANTIBIOTICS AND AMIODARONE WERE ADMINISTERED TO TREAT THE ISSUE. ON (B)(6) 2025, THE PATIENT WAS ABLE TO BE STABILIZED. ON (B)(6) 2025, AN ECG WAS REPEATED, SHOWING A WIDER QRS. IT WAS DECIDED TO CONTINUE TREATING THE PATIENT WITH ANTIBIOTICS, AND A PACEMAKER WAS ALSO IMPLANTED. AFTER THIS, THE PATIENT ENDED UP GOING INTO SHOCK AND AN X-RAY INDICATED POSSIBLE VENTRICULAR PERFORATION. THE PATIENT THEN BECAME UNSTABLE, AND A PERICARDIAL PUNCTURE AND DRAIN WAS ATTEMPTED. THE PATIENT¿S HEMODYNAMICS DIDN'T IMPROVE AND THEY PASSED AWAY (B)(6) 2025. THE PHYSICIAN BELIEVES THAT IT WAS LIKELY DUE TO A MULTI-RESISTANT PSEUDOMONAS INFECTION. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. AN EVENT OF FATIGUE, RESPIRATORY TRACT INFECTION, SHOCK, ATRIAL FIBRILLATION AND PATIENT DEATH WAS REPORTED. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED, AND THE PRODUCT MET ALL SPECIFICATIONS. BASED ON ALL AVAILABLE INFORMATION, CAUSES FOR THE REPORTED FATIGUE, RESPIRATORY TRACT INFECTION, SHOCK, ATRIAL FIBRILLATION AND PATIENT DEATH COULD NOT BE CONCLUSIVELY DETERMINED. IT IS POSSIBLE THAT THEY ARE RELATED TO THE IMPLANT PROCEDURE. HOWEVER, THIS CANNOT BE CONFIRMED. THE REPORTED DEATH APPEARS TO BE RELATED TO THE MULTI-RESISTANT PSEUDOMONAS INFECTION. THE REPORTED UNEXPECTED MEDICAL INTERVENTION AND SURGICAL INTERVENTION WERE THE RESULT OF CASE-SPECIFIC CIRCUMSTANCES, AS PERICARDIOCENTESIS WAS PERFORMED AND PERMANENT PACEMAKER IMPLANTED. THE EVENT AND IMAGING WAS REVIEWED BY AN ABBOTT SENIOR DIRECTOR OF MEDICAL AFFAIRS. THE REVIEWER STATED, THE CAUSE OF DEATH CANNOT BE DETERMINED BASED ON THE INFORMATION PROVIDED. HOWEVER, WE SUSPECT ITS RELATED TO THE RIGHT VENTRICULAR (RV) PUNCTURE FROM THE TEMPORARY PACEMAKER PLACEMENT. NO DEVICE MALFUNCTION IS SUSPECTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21588954&pc=OZD | 1220648-2025-26685 | 2025-02-14 05:00 | Death | ABIOMED, INC. | 2025-03-12 04:00 | OZD | IMPELLA CP | Malposition of Device | Cardiac Arrest | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT A PATIENT IN CARDIOGENIC SHOCK FROM AN ACUTE MYOCARDIAL INFARCTION IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS) EXPERIENCED POSITIONING ISSUES THAT RESULTED IN EXPLANT WITH REPLACEMENT OF THE DEVICE. THE PATIENT WOULD SUBSEQUENTLY EXPIRE. THE PATIENT PRESENTED TO THE EMERGENCY DEPARTMENT WITH CHEST PAIN, SHORTNESS OF BREATH AND HYPOTENSIVE. THE PATIENT WAS ON TWO PRESSORS AND HAD AN EJECTION FRACTION OF 15%. OF NOTE, THE PATIENT DECLINED CARDIOTHORACIC SURGERY THREE YEARS PRIOR. THE PATIENT WAS TRANSFERRED TO THE CATHETERIZATION LAB FOR AN IMPELLA CP DEVICE PLACEMENT AND POSSIBLE PERCUTANEOUS CORONARY INTERVENTION. THE IMPELLA CP DEVICE WAS IMPLANTED VIA THE RIGHT FEMORAL ARTERY WITHOUT INCIDENT. CATHETERIZATION REVEALED SEVERE LEFT MAIN ARTERY DISEASE. THE PATIENT NOW ON IMPELLA MCS WAS CHEST PAIN FREE AND NORMAL ELECTROCARDIOGRAM. THE PHYSICIAN MADE THE DECISION NOT TO INTERVENE AND CONSULT WITH CARDIOTHORACIC SURGERY. THE PATIENT WAS SENT TO THE UNIT ON SUPPORT AND VOMITED ON ARRIVAL. IT WAS NOTED APPROXIMATELY TWO DAYS AFTER START OF THE SUPPORT, THE PUMP WAS PULLED BACK INTO AORTA AND REPOSITIONED AT BEDSIDE. THE POSITION WAS QUESTIONABLE, HOWEVER. THE PATIENT WAS TRANSPORTED OUT TO A HEART FAILURE HOSPITAL AND ON ARRIVAL, THE PUMP POSITIONING WAS STILL QUESTIONABLE. MULTIPLE ATTEMPTS, THEREFORE, WERE MADE TO REPOSITION, AND ALL WERE UNSUCCESSFUL. THE PATIENT WAS BROUGHT TO THE CATHETERIZATION LAB FOR REPOSITIONING UNDER FLUOROSCOPY; HOWEVER, THIS ATTEMPT TOO WAS UNSUCCESSFUL. THE DECISION WAS MADE TO EXPLANT AND REPLACE WITH A NEW IMPELLA CP DEVICE WHICH WAS COMPLETED FOR CONTINUATION OF THERAPY. HOWEVER APPROXIMATELY 1.5 DAYS LATER THE PATIENT EXPIRED ON SUPPORT. NO FURTHER DETAILS AROUND THE DEMISE WERE PROVIDED. ALTHOUGH THE IMPELLA CP DEVICE WAS EXPLANT AND REPLACED PRIOR TO THE PATIENT'S DEMISE ASSOCIATION CANNOT BE EXCLUDED GIVEN THE POSITIONING DIFFICULTY ENCOUNTERED AND POTENTIAL COMPROMISE TO CIRCULATORY SUPPORT. Manufacturer Narrative: THE INVESTIGATION OF POSITIONING ISSUES HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. REPORTED THE PATIENT VOMITED WHICH MOVED THE PUMP OUT OF POSITION WITH IMPELLA IN AORTA ALARMS. THE PUMP WAS PUSHED BACK INTO THE VENTRICLE WITHOUT THE GUIDEWIRE BUT UNABLE TO REPOSITION PUMP USING ECHO AND UNDER FLUORO. PUMP WAS REMOVED. THE ROOT CAUSE OF THE POSITIONING ISSUE WAS MOST LIKELY PATIENT MOVEMENT SINCE THE IMPROPER POSITIONING OCCURRED WHEN THE PATIENT VOMITED. E4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26685 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21589786&pc=PLO | 3002773840-2025-00013 | 2025-01-02 05:00 | Death | BIOFIRE DIAGNOSTICS, LLC | 2025-03-12 04:00 | PLO | FILMARRAY® MENINGITIS/ENCEPHALITIS (ME) PANEL | False Negative Result | Insufficient Information | K160462 |  | 1 | Event Description: SUMMARY: (B)(6) CANADA) REPORTED A POTENTIAL FALSE NEGATIVE CRYPTOCOCCUS NEOFORMANS/GATTII RESULT ON THE FILMARRAY MENINGITIS/ENCEPHALITIS (ME) PANEL AFTER TESTING A PATIENT'S CEREBROSPINAL FLUID (CSF) SAMPLE. DUE TO THE FILMARRAY ME PANEL RESULT, THE PATIENT'S DIAGNOSIS AND TREATMENT WAS DELAYED. THE CUSTOMER REPORTED THAT THE PATIENT PASSED AWAY AND THEIR DEATH APPEARED TO BE MULTIFACTORIAL, BUT THE DIAGNOSTIC DELAY POSSIBLY CONTRIBUTED TO THE UNFAVORABLE OUTCOME FOR THE PATIENT. BIOFIRE HAS REQUESTED FURTHER INFORMATION FROM THE CUSTOMER AND IS CURRENTLY INVESTIGATING THIS EVENT. NO REMEDIAL ACTION, CORRECTIVE ACTION, PREVENTIVE ACTION, OR FIELD SAFETY CORRECTIVE ACTION (FSCA) HAS BEEN DEEMED NECESSARY AT THIS TIME.Event Description: SUMMARY: (B)(6) HÔPITAL (QUEBEC, CANADA) REPORTED A POTENTIAL FALSE NEGATIVE CRYPTOCOCCUS NEOFORMANS/GATTII RESULT ON THE FILMARRAY MENINGITIS/ENCEPHALITIS (ME) PANEL AFTER TESTING A PATIENT'S CEREBROSPINAL FLUID (CSF) SAMPLE. DUE TO THE FILMARRAY ME PANEL RESULT, THE PATIENT'S DIAGNOSIS AND TREATMENT WAS DELAYED. THE CUSTOMER REPORTED THAT THE PATIENT PASSED AWAY AND THEIR DEATH APPEARED TO BE MULTIFACTORIAL, BUT THE DIAGNOSTIC DELAY POSSIBLY CONTRIBUTED TO THE UNFAVORABLE OUTCOME FOR THE PATIENT. UPON INVESTIGATION, NO MALFUNCTION OCCURRED AND THE FILMARRAY ME PANEL WAS WORKING WITHIN SPECIFICATION. THE INVESTIGATION CONCLUDED THAT THE MOST LIKELY CAUSE FOR THE DISCREPANT C. NEOFORMANS/GATTII RESULT WAS SENSITIVITY/SPECIFICITY DIFFERENCES BETWEEN THE FILMARRAY ME PANEL AND TRADITIONAL LABORATORY METHODS. Manufacturer Narrative: INVESTIGATION: (B)(6) CANADA) REPORTED A POTENTIAL FALSE NEGATIVE CRYPTOCOCCUS NEOFORMANS/GATTII RESULT ON THE FILMARRAY MENINGITIS/ENCEPHALITIS (ME) PANEL AFTER TESTING A PATIENT'S CEREBROSPINAL FLUID (CSF) SAMPLE. DUE TO THE FILMARRAY ME PANEL RESULT, THE PATIENT'S DIAGNOSIS AND TREATMENT WERE DELAYED. THE PATIENT PASSED AWAY AND THE CUSTOMER REPORTED THAT THE PATIENT'S DEATH APPEARED TO BE MULTIFACTORIAL, BUT THE DIAGNOSTIC DELAY POSSIBLY CONTRIBUTED TO THE UNFAVORABLE OUTCOME FOR THE PATIENT. BIOFIRE'S INVESTIGATION INTO THIS EVENT IS ONGOING AND FURTHER INFORMATION HAS BEEN REQUESTED FROM THE CUSTOMER. THE FULL INVESTIGATION AND ASSOCIATED CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. NO REMEDIAL ACTION, CORRECTIVE ACTION, PREVENTIVE ACTION, OR FSCA HAS BEEN DEEMED NECESSARY AT THIS TIME. CONCLUSION: INVESTIGATION ONGOING. Manufacturer Narrative: INVESTIGATION: THE PATIENT WAS A 71-YEAR-OLD MALE WHO WAS ADMITTED TO THE HOSPITAL ON (B)(6) 2024, WITH NONSPECIFIC SYMPTOMS. ON (B)(6) 2025, MORE PRONOUNCED NEUROLOGICAL SYMPTOMS APPEARED, LEADING TO CSF COLLECTION BY LUMBAR PUNCTURE. ON (B)(6) 2025, THE CSF SAMPLE WAS TESTED ON THE FILMARRAY ME PANEL. THE FILMARRAY ME PANEL REPORTED ALL ANALYTES AS NOT DETECTED. ON (B)(6) 2025, THE CSF CULTURE WAS POSITIVE FOR CRYPTOCOCCUS, WHICH PROMPTED THE CUSTOMER TO PERFORM CRYPTOCOCCAL ANTIGEN (CRAG) TEST ON THE CSF SAMPLE, CONFIRMING CRYPTOCOCCUS. TREATMENT WAS INITIATED FOLLOWING THE POSITIVE CRYPTOCOCCUS RESULT. THE CUSTOMER REPORTED THAT THIS RESULTED IN A PROCESSING DELAY OF 9 DAYS ((B)(6) 2024 TO (B)(6) 2025). PLEASE NOTE THAT THE DELAY ASSOCIATED WITH THE FILMARRAY ME PANEL RESULT WAS 4 DAYS AS THE TEST WAS PERFORMED ON (B)(6) 2025. ACCORDING TO THE CUSTOMER, CRYPTOCOCCUS WAS INITIALLY RULED OUT DUE TO THE NEGATIVE FILMARRAY ME PANEL RESULT, THE ABSENCE OF RISK FACTORS, AND A NON-SUGGESTIVE CLINICAL PICTURE. THE CUSTOMER REPORTED THAT DUE TO THE COMPLEXITY OF THE CASE, THE PATIENT WAS TRANSFERRED TO A TERTIARY CENTER ON (B)(6) 2025, AND DIED ON (B)(6) 2025. WHEN ASKED IF THE PATIENT WAS AFFECTED BY THE FILMARRAY ME PANEL RESULT, THE CUSTOMER STATED THAT IT WAS DIFFICULT TO ASSESS WITH CERTAINTY WHETHER THE PATIENT WAS DIRECTLY HARMED BY THE DIAGNOSTIC DELAY, HOWEVER, A DELAY IN TREATMENT IN A SERIOUS INFECTION SUCH AS CRYPTOCOCCOSIS, CLINICAL COURSE CAN BE INFLUENCED. THEY STATED, "IN THIS CASE, THE PATIENT'S DEATH APPEARS TO BE MULTIFACTORIAL, BUT THIS DIAGNOSTIC DELAY POSSIBLY CONTRIBUTED TO THE UNFAVORABLE OUTCOME." BIOFIRE MEDICAL AFFAIRS ALSO PROVIDED AN ASSESSMENT OF THE CASE. MEDICAL AFFAIRS STATED THAT THE GOLD STANDARD FOR DETECTING CRYPTOCOCCUS IS CULTURE AND CRYPTOCOCCAL ANTIGEN (CRAG). MOLECULAR METHODOLOGIES CAN PROVIDE VALUE IN ESTABLISHING THE ETIOLOGY OF DISEASE; HOWEVER, PERFORMANCE IS INSUFFICIENT TO RULE OUT INFECTION. A GLOBAL GUIDELINE FOR THE DIAGNOSIS OF CRYPTOCOCCOSIS RECOMMENDS THE DIAGNOSTIC WORKUP FOR A PATIENT WITH SUSPECTED INFECTION TO REQUIRE ASSESSMENT FOR CENTRAL NERVOUS SYSTEM (CNS), PULMONARY, AND OTHER SYSTEM INVOLVEMENT. SPECIFIC TO CNS INFECTION, THE RECOMMENDATIONS ARE FOR THE MEASUREMENT OF CSF OPENING PRESSURE, GLUCOSE, PROTEIN, CELL COUNTS, MICROSCOPY, CULTURE, AND QUANTIFICATION OF CSF CRAG.1 THE MAJOR CONTRIBUTING FACTORS FOR THE DIAGNOSTIC DELAY WAS: DEVIATION FROM CLINICAL GUIDELINES FOR EVALUATING PATIENTS SUSPECTED FOR CRYPTOCOCCAL INFECTION CUSTOMER MISCONCEPTION OF THE PERFORMANCE OF MOLECULAR METHODS FOR CRYPTOCOCCUS DETECTION. REFERENCE: 1. GLOBAL GUIDELINE FOR THE DIAGNOSIS AND MANAGEMENT OF CRYPTOCOCCOSIS: AN INITIATIVE OF THE ECMM AND ISHAM IN COOPERATION WITH THE ASM - THE LANCET INFECTIOUS DISEASES. ACCESSED (B)(6) 2025. HTTPS://WWW.THELANCET.COM/JOURNALS/LANINF/ARTICLE/PIIS1473-3099(23)00731-4/FULLTEXT. IN ADDITION TO THE CUSTOMER INFORMATION AND MEDICAL AFFAIRS ASSESSMENT, ANALYSIS OF THE FILMARRAY ME PANEL RUN FILE SHOWED NO AMPLIFICATION OR MELT SIGNATURE FOR THE C. NEOFORMANS/GATTII ASSAY. QUALITY CONTROL (QC) RECORDS FOR POUCH LOT# 3FBH24 (KIT LOT# 1625324) WERE REVIEWED. THIS POUCH LOT PASSED QC CRITERIA AND WAS FOUND WITHIN SPECIFICATIONS. NO RUN MALFUNCTION OCCURRED AT THE CUSTOMER SITE AND THE FILMARRAY INSTRUMENT (SERIAL NUMBER# (B)(6)) WAS WORKING WITHIN DESIGNED SPECIFICATIONS. CONCLUSION: THE INVESTIGATION CONCLUDED THAT THE MOST LIKELY CAUSE FOR THE DISCREPANT C. NEOFORMANS/GATTII RESULT WAS SENSITIVITY/SPECIFICITY DIFFERENCES BETWEEN THE FILMARRAY ME PANEL AND TRADITIONAL LABORATORY METHODS. ACCORDING TO THE FILMARRAY ME PANEL INSTRUCTIONS BOOKLET (WWW.ONLINE-IFU.COM/ITI0035), PATIENTS WITH A SUSPICION OF CRYPTOCOCCAL MENINGITIS AND A NEGATIVE CRYPTOCOCCAL PCR RESULT, SUCH AS BY THE FILMARRAY ME PANEL, SHOULD BE TESTED FOR CRYPTOCOCCAL ANTIGEN (CRAG). ACCORDING TO THE "CRYPTOCOCCUS DETECTION BY THE BIOFIRE® FILMARRAY® MENINGITIS/ENCEPHALITIS (ME) PANEL" TECHNICAL NOTE [FLM1-PRT-0278] , THERE HAVE BEEN REPORTS OF NEGATIVE RESULTS BY THE FILMARRAY ME PANEL IN PATIENTS WITH NEWLY DIAGNOSED CRYPTOCOCCAL MENINGITIS AND POSITIVE CRAG AND/OR CULTURE. ADDITIONALLY, DISCORDANT CRAG AND FILMARRAY ME PANEL RESULTS HAVE BEEN OBSERVED IN PERSONS EITHER ON ANTIFUNGAL TREATMENT AT THE TIME OF TESTING OR WITH A PAST HISTORY OF TREATED DISEASE IN WHICH THE CRAG REMAINS POSITIVE. THE FILMARRAY ME PANEL DETECTS NUCLEIC ACID WITHIN THE CSF SAMPLE, WHILE CRAG TEST DETECTS THE CRYPTOCOCCUS ANTIGEN IN THE CSF SAMPLE. PCR-BASED DIAGNOSIS OF CRYPTOCOCCAL MENINGITIS HAS NOT BEEN WIDELY DEVELOPED GIVEN THE HIGH SENSITIVITY, WIDE AVAILABILITY, AND LOW COST OF CRAG TESTING. A NEGATIVE FILMARRAY ME PANEL RESULT DOES NOT EXCLUDE THE POSSIBILITY OF CNS INFECTION AND SHOULD NOT BE USED AS THE SOLE BASIS FOR DIAGNOSIS, TREATMENT, OR OTHER MANAGEMENT DECISIONS. CLINICAL PERFORMANCE: ACCORDING TO TABLE 9. FILMARRAY ME PROSPECTIVE CLINICAL PERFORMANCE SUMMARY OF THE FILMARRAY ME PANEL INSTRUCTION BOOKLET (WWW.ONLINE-IFU.COM/ITI0035), THE PERFORMANCE CLAIM FOR THE C. NEOFORMANS/GATTII ASSAY COMPARED TO PCR WITH BI-DIRECTIONAL SEQUENCING SHOWED AN OVERALL POSITIVE PERCENT AGREEMENT (PPA) OF 100% (1/1) AND AN OVERALL NEGATIVE PERCENT AGREEMENT (NPA) OF 99.7% (95% CI 99.3-99.9%) C. NEOFORMANS/GATTII WAS DETECTED IN 2/4 FALSE POSITIVE SPECIMENS USING A COMMERCIALLY AVAILABLE ANTIGEN TEST. ONE FALSE POSITIVE SPECIMEN WAS POSITIVE BY STANDARD CULTURE. TO SUPPLEMENT THE RESULTS OF THE PROSPECTIVE CLINICAL STUDY, AN EVALUATION OF 235 PRESELECTED ARCHIVED SPECIMENS (OF WHICH 25 WERE NEGATIVE) WAS PERFORMED. ACCORDING TO TABLE 14. FILMARRAY ME PANEL ARCHIVED SPECIMEN PERFORMANCE DATA SUMMARY OF THE FILMARRAY ME PANEL INSTRUCTION BOOKLET, THE C. NEOFORMANS/GATTII ASSAY COMPARED TO A CONFIRMATORY MOLECULAR TEST (E.G., PCR WITH BI-DIRECTIONAL SEQUENCING) SHOWED A PPA OF 100% (95% CI 83.2-100%) AND AN NPA OF 100% (95% CI 97.8-100%). ONE SPECIMEN WAS SEQUENCED AND IDENTIFIED AS C. GATTII AND 18 WERE C. NEOFORMANS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21590192&pc=QMG | 3015053838-2025-00048 | 2025-02-18 05:00 | Death | SHOCKWAVE MEDICAL, INC. | 2025-03-12 04:00 | QMG | SHOCKWAVE C2+ CORONARY INTRAVASCULAR LITHOTRIPSY (IVL) CATHETER | Material Split, Cut or Torn | Arrhythmia; Cardiac Arrest; Ventricular Fibrillation | P200039/S008 |  | 1 | Event Description: A SHOCKWAVE C2+ CORONARY INTRAVASCULAR LITHOTRIPSY (IVL) CATHETER WAS USED TO TREAT A HEAVILY CALCIFIED LESION IN THE LEFT CIRCUMFLEX ARTERY (LCX). THE PROCEDURE FACED INITIAL ACCESSIBILITY CHALLENGES, REQUIRING SMALL BALLOONS AND SIGNIFICANT BALLOONING WITH A NON-COMPLIANT (NC) BALLOON. INTRAVASCULAR ULTRASOUND (IVUS) WAS DIFFICULT BUT EVENTUALLY SUCCESSFUL BEFORE IVL. DURING THE PROCEDURE, 50 IVL PULSES WERE DELIVERED AT 4 ATM BEFORE THE BALLOON RUPTURED. IT WAS REPORTED THAT THE TECH MAY HAVE INJECTED AIR FOLLOWING THE RUPTURE. THE PATIENT WENT INTO VENTRICULAR FIBRILLATION (V FIB) MULTIPLE TIMES, WAS IMMEDIATELY SHOCKED, A CODE WAS CALLED, AND THE PATIENT WAS INTUBATED. EMERGENCY MEDICATIONS WERE ADMINISTERED. THE PATIENT HAD A RETURN OF SPONTANEOUS CIRCULATION (ROSC) BUT CONTINUED TO GO BACK INTO V FIB OVER THE NEXT 45 MINUTES. THE PATIENT PASSED AWAY FOLLOWING THE CODE, APPROXIMATELY 45 MINUTES LATER. THE LAST IMAGE SHOWED THE VESSEL WAS OPEN, BUT IMAGING STOPPED WHEN V FIB OCCURRED, AND THE CODE BEGAN. THE PHYSICIAN NOTED THE POSSIBILITY OF AIR INJECTION POST-RUPTURE BUT DID NOT DIRECTLY ATTRIBUTE THE ISSUES TO THE IVL DEVICE. Manufacturer Narrative: THE DEVICE REFERENCED IN THIS REPORT WAS NOT RETURNED TO SHOCKWAVE MEDICAL, INC. THEREFORE, A PHYSICAL EXAMINATION CANNOT BE PERFORMED. THE CAUSE OF THE VENTRICULAR FIBRILLATION, CARDIAC ARREST, AND SUBSEQUENT PATIENT DEATH COULD NOT BE DEFINITIVELY DETERMINED BASED ON THE INFORMATION AVAILABLE. THE PATIENT WAS REPORTED TO HAVE A LONG HISTORY OF CARDIAC ISSUES. THE PHYSICIAN NOTED THE POSSIBILITY OF AIR INJECTION POST-RUPTURE BUT DID NOT DIRECTLY ATTRIBUTE THE ISSUES TO THE IVL DEVICE. SHOCKWAVE MEDICAL HAS CONTROLS IN PLACE TO ENSURE DEVICES ARE BUILT TO APPROVED PROCEDURES AND MEET LOT RELEASE ACCEPTANCE CRITERIA PRIOR TO BEING DISTRIBUTED. A REVIEW OF THE MANUFACTURING AND TEST DOCUMENTATION FOR SUBJECT LOT DOES NOT REVEAL ANY ISSUES WITH THE MANUFACTURING OF THE DEVICE. THE DEVICE PASSED ALL OF SHOCKWAVE MEDICAL, INC. ACCEPTANCE CRITERIA PRIOR TO SHIPPING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21590319&pc=MHX | 8030229-2025-04716 | 2025-02-10 05:00 | Death | NIHON KOHDEN CORPORATION | 2025-03-12 04:00 | MHX | LIFE SCOPE PT | No Audible Alarm; Protective Measures Problem | Sepsis | K220976 |  | 1 | Event Description: THE CUSTOMER REPORTED THAT A PATIENT PASSED AWAY.Event Description: THE CUSTOMER REPORTED THAT A PATIENT PASSED AWAY. Manufacturer Narrative: THE CUSTOMER REPORTED THAT A PATIENT PASSED AWAY. THE CUSTOMER DISCOVERED A 20-MINUTE GAP WITH NO CLINICAL ALARMS TRIGGER TO INDICATE THAT THE PATIENT WAS CONTINUALLY DESATURATING OR THAT THE EKG LEADS WERE OFF FOR NURSING TO RESPOND. ACCORDING TO THE CUSTOMER, THE PATIENT HAD ALREADY PASSED AWAY DURING THOSE 20 MINUTES. IT IS POSSIBLE THAT THE PATIENT WAS NEVER DIRECTLY ADMITTED TO THE BEDSIDE MONITOR. PER THE CUSTOMER, IN THE ER IT'S COMMON FOR PATIENTS NOT TO BE FORMALLY ADMITTED TO THE BEDSIDE MONITOR AND THE MONITOR MAY BE RUNNING CONTINUOUSLY. THE CUSTOMER LATER CONFIRMED THAT THE PATIENT WAS NOT ADMITTED BY NAME INTO THE SYSTEM, THE PATIENT WAS JUST BEING MONITORED WITHOUT A NAME IN THE ER. NIHON KOHDEN CONTINUES TO INVESTIGATE THE REPORTED EVENT. NIHON KOHDEN WILL SUBMIT A SUPPLEMENTAL REPORT IN ACCORDANCE WITH 21 CFR SECTION 803.56 WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. ADDITIONAL DEVICE INFORMATION: D10 CONCOMITANT MEDICAL DEVICE: THE FOLLOWING DEVICE(S) WERE USED IN CONJUNCTION WITH THE BSM: CNS: MODEL #:CU-152R SERIAL #: (B)(6). DEVICE MANUFACTURER DATA: 11/26/2021 UNIQUE IDENTIFIER (UDI) #: (B)(4). RETURNED TO NIHON KOHDEN: NO MODEL #:PU-681RA SERIAL #: (B)(6). DEVICE MANUFACTURER DATA: 02/04/2022. UNIQUE IDENTIFIER (UDI) #: (B)(4). RETURNED TO NIHON KOHDEN: NO. Manufacturer Narrative: DETAILS OF COMPLAINT: THE CUSTOMER REPORTED THAT A PATIENT PASSED AWAY. THE CUSTOMER DISCOVERED A 20-MINUTE GAP WITH NO CLINICAL ALARMS TRIGGER TO INDICATE THAT THE PATIENT WAS CONTINUALLY DESATURATING OR THAT THE EKG LEADS WERE OFF FOR NURSING TO RESPOND. ACCORDING TO THE CUSTOMER, THE PATIENT HAD ALREADY PASSED AWAY DURING THOSE 20 MINUTES. IT IS POSSIBLE THAT THE PATIENT WAS NEVER DIRECTLY ADMITTED TO THE BEDSIDE MONITOR. PER THE CUSTOMER, IN THE ER IT'S COMMON FOR PATIENTS NOT TO BE FORMALLY ADMITTED TO THE BEDSIDE MONITOR AND THE MONITOR MAY BE RUNNING CONTINUOUSLY. THE CUSTOMER LATER CONFIRMED THAT THE PATIENT WAS NOT ADMITTED BY NAME INTO THE SYSTEM, THE PATIENT WAS JUST BEING MONITORED WITHOUT A NAME IN THE ER. INVESTIGATION SUMMARY: NKC REVIEWED THE LOGS PROVIDED BY THE CUSTOMER AND DETERMINED THAT ROOT CAUSE WAS A USER ERROR. ACCORDING TO NKC, IT APPEARED THAT THE CUSTOMER WAS USING THE 20 SECONDS DEFAULT SETTING INTERVAL SO THE SPO2 ALARM WOULD HAVE OCCURED IF THE SPO2 VALUE BELOW THE LOWER LIMIT HAD LASTED 20 SECONDS OR LONGER. PER NKC, THE CUSTOMER MISSED THE ALARM DUE TO THIS DEFAULT SETTING. ADDITIONAL DEVICE INFORMATION: D10 CONCOMITANT MEDICAL DEVICE: THE FOLLOWING DEVICE(S) WERE USED IN CONJUNCTION WITH THE BSM: CNS: MODEL #:CU-152R. SERIAL #: (B)(6) DEVICE MANUFACTURER DATA: 11/26/2021. UNIQUE IDENTIFIER (UDI) #: (B)(4). RETURNED TO NIHON KOHDEN: NO. MODEL #:PU-681RA. SERIAL #: (B)(6) DEVICE MANUFACTURER DATA: 02/04/2022 UNIQUE IDENTIFIER (UDI) #: (B)(4). RETURNED TO NIHON KOHDEN: NO.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21590671&pc=DSQ | 2916596-2025-01489 | 2023-07-07 04:00 | Death | THORATEC CORPORATION | 2025-03-12 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY. ADDITIONAL INFORMATION WAS NOT PROVIDED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: SECTION A2: PATIENT AGE CORRECTED MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C, IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21569727&pc=NIO | 9681442-2025-00057 | 2022-02-08 05:00 | Death | ANGIOMED GMBH & CO. MEDIZINTECHNIK KG | 2025-03-11 04:00 | NIO | E-LUMINEXX VASCULAR STENT | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P080007 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESULTS OF THE CLINICAL TRIAL THAT FIVE YEARS, EIGHT MONTHS AND TWENTY-ONE DAYS POST A STENT PLACEMENT ON THE LEFT LEG FOR THE DE-NOVO STENOSIS ON THE COMMON AND EXTERNAL ILIAC ARTERY, THE SUBJECT HAD AN ADVERSE EVENT OF DEATH. IT WAS FURTHER REPORTED THAT THE PRIMARY CAUSE OF DEATH WAS NOT PROVIDED. REPORTEDLY, THE ADVERSE EVENT DEATH WAS POSSIBLY RELATED TO THE STUDY DEVICE AND NOT RELATED TO THE PROCEDURE.Event Description: IT WAS REPORTED THROUGH THE RESULTS OF THE CLINICAL TRIAL THAT FIVE YEARS, EIGHT MONTHS AND TWENTY-ONE DAYS POST A STENT PLACEMENT ON THE LEFT LEG FOR THE DE-NOVO STENOSIS ON THE COMMON AND EXTERNAL ILIAC ARTERY, THE SUBJECT HAD AN ADVERSE EVENT OF DEATH. IT WAS FURTHER REPORTED THAT THE PRIMARY CAUSE OF DEATH WAS NOT PROVIDED. REPORTEDLY, THE ADVERSE EVENT DEATH WAS POSSIBLY RELATED TO THE STUDY DEVICE AND NOT RELATED TO THE PROCEDURE. Manufacturer Narrative: H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. Manufacturer Narrative: H10: MANUFACTURING REVIEW: BASED ON THE EVENT INFORMATION PROVIDED THERE IS NO INDICATION THAT A MANUFACTURING PROCESS MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED ISSUE. THEREFORE, A DHR REVIEW IS CONSIDERED TO BE NOT REQUIRED. INVESTIGATION SUMMARY: THE STENT REMAINS IMPLANTED. X-RAY IMAGES WERE NOT PROVIDED FOR EVALUATION. THE EVALUATION OF THE MATERIAL AVAILABLE IS SUBJECT TO THE RETROSPECTIVE ASSESSMENT AS PART OF POST-MARKET CLINICAL FOLLOW-UP (PMCF) ACTIVITIES. BASED ON THE INFORMATION AVAILABLE, THE INVESTIGATION IS CLOSED WITH INCONCLUSIVE RESULT. BASED ON INFORMATION AVAILABLE, THERE IS NO INDICATION THAT THE DEATH OF THE PATIENT MIGHT BE RELATED TO THE PLACED STENT OR THE STENT PLACEMENT PROCEDURE. LABELING REVIEW: RELEVANT LABELING SUPPLIED WITH THIS PRODUCT WAS REVIEWED. POTENTIAL COMPLICATIONS AND ADVERSE EVENTS WHICH MAY OCCUR DURING USE OF THE DEVICE WERE FOUND TO BE REFERENCED IN THE INSTRUCTIONS FOR USE. USERS AND/OR PATIENTS SHOULD REPORT ANY SERIOUS INCIDENT THAT HAS OCCURRED IN RELATION TO THE DEVICE TO THE MANUFACTURER AND TO EITHER THE EUROPEAN UNION NATIONAL COMPETENT AUTHORITY OR TO THE REGULATORY AUTHORITY OF THE COUNTRY IN WHICH THE USER AND/OR PATIENT IS ESTABLISHED. IN THIS CASE THE DEATH OF THE PATIENT WAS DOCUMENTED NOT BEING RELATED TO THE PLACED STENT OR THE STENT PLACEMENT PROCEDURE. H10: G3. H11: H6 (METHOD). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21570072&pc=NIO | 9681442-2025-00058 | 2020-10-24 04:00 | Death | ANGIOMED GMBH & CO. MEDIZINTECHNIK KG | 2025-03-11 04:00 | NIO | E-LUMINEXX VASCULAR STENT | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P080007 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESULTS OF THE CLINICAL TRIAL THAT FOUR YEARS, SEVEN MONTHS AND EIGHT DAYS POST A STENT PLACEMENT ON THE LEFT LEG FOR THE DE-NOVO STENOSIS ON THE COMMON ILIAC ARTERY, THE SUBJECT HAD AN ADVERSE EVENT OF DEATH. IT WAS FURTHER REPORTED THAT THE PRIMARY CAUSE OF DEATH WAS NOT PROVIDED. REPORTEDLY, THE ADVERSE EVENT DEATH WAS POSSIBLY RELATED TO THE STUDY DEVICE AND NOT RELATED TO THE PROCEDURE.Event Description: IT WAS REPORTED THROUGH THE RESULTS OF THE CLINICAL TRIAL THAT FOUR YEARS, SEVEN MONTHS AND EIGHT DAYS POST A STENT PLACEMENT ON THE LEFT LEG FOR THE DE-NOVO STENOSIS ON THE COMMON ILIAC ARTERY, THE SUBJECT HAD AN ADVERSE EVENT OF DEATH. IT WAS FURTHER REPORTED THAT THE PRIMARY CAUSE OF DEATH WAS NOT PROVIDED. REPORTEDLY, THE ADVERSE EVENT DEATH WAS POSSIBLY RELATED TO THE STUDY DEVICE AND NOT RELATED TO THE PROCEDURE. Manufacturer Narrative: H11: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE E-LUMINEXX VASCULAR STENT PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE E-LUMINEXX VASCULAR STENT PRODUCTS ARE IDENTIFIED IN D2 AND G4. MANUFACTURING REVIEW: BASED ON THE EVENT INFORMATION PROVIDED THERE IS NO INDICATION THAT A MANUFACTURING PROCESS MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED ISSUE. THEREFORE, A DHR REVIEW IS CONSIDERED TO BE NOT REQUIRED. INVESTIGATION SUMMARY: THE STENT REMAINS IMPLANTED. X-RAY IMAGES WERE NOT PROVIDED FOR EVALUATION. THE EVALUATION OF THE MATERIAL AVAILABLE IS SUBJECT TO THE RETROSPECTIVE ASSESSMENT AS PART OF POST-MARKET CLINICAL FOLLOW-UP (PMCF) ACTIVITIES. POTENTIAL FACTORS THAT COULD HAVE LED OR CONTRIBUTED TO REPORTED EVENTS ARE EVALUATED. IN THIS CASE THERE WAS NO DOCUMENTED DEFICIENCY OF THE PLACED STENT. NO IRREGULAR STENT PLACEMENT DURING INDEX PROCEDURE AND NO DEFICIENCY OF THE PLACED STENT, SUCH AS DEFORMATION OR FRACTURE WAS REPORTED. THE DEATH OF THE PATIENT WAS DOCUMENTED NOT BEING RELATED TO THE PLACED STENT OR THE STENT PLACEMENT PROCEDURE. THE PATIENT DIED ABOUT THREE YEARS AND NINE MONTHS AFTER THE STENT PLACEMENT PROCEDURE. THE PATIENT WAS A FORMER SMOKER AND HAD A CLINICAL HISTORY OF HYPERTENSION, DYSLIPIDEMIA, TYPE II DIABETES MELLITUS, OBESITY. BASED ON INFORMATION AVAILABLE, THERE IS NO INDICATION THAT THE DEATH OF THE PATIENT MIGHT BE RELATED TO THE PLACED STENT OR THE STENT PLACEMENT PROCEDURE. LABELING REVIEW: RELEVANT LABELING SUPPLIED WITH THIS PRODUCT WAS REVIEWED. POTENTIAL COMPLICATIONS AND ADVERSE EVENTS WHICH MAY OCCUR DURING USE OF THE DEVICE WERE FOUND TO BE REFERENCED IN THE IFU. USERS AND/OR PATIENTS SHOULD REPORT ANY SERIOUS INCIDENT THAT HAS OCCURRED IN RELATION TO THE DEVICE TO THE MANUFACTURER AND TO EITHER THE EUROPEAN UNION NATIONAL COMPETENT AUTHORITY OR TO THE REGULATORY AUTHORITY OF THE COUNTRY IN WHICH THE USER AND/OR PATIENT IS ESTABLISHED. IN THIS CASE THE DEATH OF THE PATIENT WAS DOCUMENTED NOT BEING RELATED TO THE PLACED STENT OR THE STENT PLACEMENT PROCEDURE. G3, H6 (METHOD). SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. Manufacturer Narrative: H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE E-LUMINEXX VASCULAR STENT PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE E-LUMINEXX VASCULAR STENT PRODUCTS ARE IDENTIFIED IN D2 AND G4. H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21570191&pc=NIO | 9681442-2025-00059 | 2020-11-29 05:00 | Death | ANGIOMED GMBH & CO. MEDIZINTECHNIK KG | 2025-03-11 04:00 | NIO | E-LUMINEXX VASCULAR STENT | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P080007 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESULTS OF THE CLINICAL TRIAL THAT THREE YEARS AND TWENTY DAYS POST A STENT PLACEMENT ON THE LEFT LEG FOR THE DE-NOVO STENOSIS ON THE COMMON AND EXTERNAL ILIAC ARTERY, THE SUBJECT HAD AN ADVERSE EVENT OF DEATH. IT WAS FURTHER REPORTED THAT THE PRIMARY CAUSE OF DEATH WAS NOT PROVIDED. REPORTEDLY, THE ADVERSE EVENT DEATH WAS POSSIBLY RELATED TO THE STUDY DEVICE AND NOT RELATED TO THE PROCEDURE.Event Description: IT WAS REPORTED THROUGH THE RESULTS OF THE CLINICAL TRIAL THAT THREE YEARS AND TWENTY DAYS POST A STENT PLACEMENT ON THE LEFT LEG FOR THE DE-NOVO STENOSIS ON THE COMMON AND EXTERNAL ILIAC ARTERY, THE SUBJECT HAD AN ADVERSE EVENT OF DEATH. IT WAS FURTHER REPORTED THAT THE PRIMARY CAUSE OF DEATH WAS NOT PROVIDED. REPORTEDLY, THE ADVERSE EVENT DEATH WAS POSSIBLY RELATED TO THE STUDY DEVICE AND NOT RELATED TO THE PROCEDURE. Manufacturer Narrative: H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE E-LUMINEXX VASCULAR STENT PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE E-LUMINEXX VASCULAR STENT PRODUCTS ARE IDENTIFIED IN D2 AND G4. H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. Manufacturer Narrative: H11: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE E-LUMINEXX VASCULAR STENT PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE E-LUMINEXX VASCULAR STENT PRODUCTS ARE IDENTIFIED IN D2 AND G4. MANUFACTURING REVIEW: BASED ON THE EVENT INFORMATION PROVIDED THERE IS NO INDICATION THAT A MANUFACTURING PROCESS MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED ISSUE. THEREFORE, A DHR REVIEW IS CONSIDERED TO BE NOT REQUIRED. INVESTIGATION SUMMARY: THE STENT REMAINS IMPLANTED. X-RAY IMAGES WERE NOT PROVIDED FOR EVALUATION. THE EVALUATION OF THE MATERIAL AVAILABLE IS SUBJECT TO THE RETROSPECTIVE ASSESSMENT AS PART OF POST-MARKET CLINICAL FOLLOW-UP (PMCF) ACTIVITIES. BASED ON THE INFORMATION AVAILABLE, THE INVESTIGATION IS CLOSED WITH INCONCLUSIVE RESULT. BASED ON INFORMATION AVAILABLE, THERE IS NO INDICATION THAT THE DEATH OF THE PATIENT MIGHT BE RELATED TO THE PLACED STENT OR THE STENT PLACEMENT PROCEDURE. LABELING REVIEW: RELEVANT LABELING SUPPLIED WITH THIS PRODUCT WAS REVIEWED. POTENTIAL COMPLICATIONS AND ADVERSE EVENTS WHICH MAY OCCUR DURING USE OF THE DEVICE WERE FOUND TO BE REFERENCED IN THE INSTRUCTIONS FOR USE. USERS AND/OR PATIENTS SHOULD REPORT ANY SERIOUS INCIDENT THAT HAS OCCURRED IN RELATION TO THE DEVICE TO THE MANUFACTURER AND TO EITHER THE EUROPEAN UNION NATIONAL COMPETENT AUTHORITY OR TO THE REGULATORY AUTHORITY OF THE COUNTRY IN WHICH THE USER AND/OR PATIENT IS ESTABLISHED. IN THIS CASE THE DEATH OF THE PATIENT WAS DOCUMENTED NOT BEING RELATED TO THE PLACED STENT OR THE STENT PLACEMENT PROCEDURE. G3, H6 (COMPONENT, METHOD) SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21570483&pc=DYE | 2025587-2025-01800 | 2024-12-09 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-11 04:00 | DYE | MOSAIC MITRAL BIOPROSTHETIC HEART VALVE | Insufficient Information | Insufficient Information | P990064 |  | 1 | Event Description: MEDTRONIC RECEIVED INFORMATION THAT 13 YEARS AND 10 MONTHS POST IMPLANT OF THIS 31MM MITRAL BIOPROSTHETIC VALVE, THE PATIENT DIED. THE CAUSE OF DEATH WAS NOT RECEIVED. THEREFORE, RELATEDNESS OF THE DEATH TO THE VALVE COULD NOT BE EXCLUDED. IT IS UNKNOWN WHETHER AN AUTOPSY WAS PERFORMED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21570723&pc=QSK | 9612515-2025-00023 | 2024-07-29 04:00 | Death | VASCUTEK LTD | 2025-03-11 04:00 | QSK | THORAFLEX HYBRID PLEXUS | Deformation Due to Compressive Stress | Adult Respiratory Distress Syndrome; Ruptured Aneurysm; Insufficient Information | P210006 |  | 1 | Event Description: THE COMPONENT THAT THE DEVICE DEFICIENCY WAS REPORTED IS TO BE RELATED TO A THORAFLEX HYBRID DEVICE. THORAFLEX HYBRID PROCEDURE TOOK PLACE ON (B)(6)2024. THE ASCENDING AORTIC GRAFT FROM THE PREVIOUS SURGERY WAS THEN TRIMMED, AND THE PROXIMAL THORAFLEX GRAFT (30MM) TRIMMED AND BEVELLED TO FIT THIS. A PLANNED AORTIC VALVE PROCEDURE WAS ALSO CARRIED OUT. AN ENDOLEAK WAS REPORTED AS BEING PRESENT AT THE END OF THE PROCEDURE. NO ADVERSE EVENTS OR DEVICE DEFICIENCIES WERE REPORTED AT THE TIME OF THE PROCEDURE. NO ISSUES WERE REPORTED ON THE POST THORAFLEX HYBRID PROCEDURE ASSESSMENT. THE SUBJECT UNDERWENT THE EXTENSION PROCEDURE DURING THE SAME HOSPITAL ADMISSION. THE EXTENSION PROCEDURE WAS REPORTED TO HAVE TAKEN PLACE ON THE (B)(6)2024 (VARIOUS DEVICE ARE LISTED: 2 X GORE CTAG, 3 X GORE VIABAHN). NO ENDOLEAK WAS REPORTED AT THE END OF THE PROCEDURE BUT THE DEVICE DEFICIENCY KINKING/ TWISTING WAS REPORTED. PATIENT OUTCOME: THE SUBJECT IS REPORTED TO HAVE DIED ON (B)(6)2024. NO AUTOPSY WAS PERFORMED. THE PRIMARY CAUSE OF DEATH WAS CATEGORISED AS "DISEASE UNDER STUDY".Event Description: EXTEND-001 (B)(6) PATIENT NO. (B)(6) HAD AN EVENT OF DEVICE DEFICIENCY- GRAFT KINKING/ TWISTING. THE COMPONENT THAT THE DEVICE DEFICIENCY WAS REPORTED IS TO BE RELATED TO A THORAFLEX HYBRID DEVICE. THORAFLEX HYBRID PROCEDURE TOOK PLACE ON (B)(6) 2024. THE ASCENDING AORTIC GRAFT FROM THE PREVIOUS SURGERY WAS THEN TRIMMED, AND THE PROXIMAL THORAFLEX GRAFT (30MM) TRIMMED AND BEVELLED TO FIT THIS. A PLANNED AORTIC VALVE PROCEDURE WAS ALSO CARRIED OUT. AN ENDOLEAK WAS REPORTED AS BEING PRESENT AT THE END OF THE PROCEDURE. NO ADVERSE EVENTS OR DEVICE DEFICIENCIES WERE REPORTED AT THE TIME OF THE PROCEDURE. NO ISSUES WERE REPORTED ON THE POST THORAFLEX HYBRID PROCEDURE ASSESSMENT. THE SUBJECT UNDERWENT THE EXTENSION PROCEDURE DURING THE SAME HOSPITAL ADMISSION. THE EXTENSION PROCEDURE WAS REPORTED TO HAVE TAKEN PLACE ON THE (B)(6) 2024 (VARIOUS DEVICE ARE LISTED: 2 X GORE CTAG, 3 X GORE VIABAHN). NO ENDOLEAK WAS REPORTED AT THE END OF THE PROCEDURE BUT THE DEVICE DEFICIENCY KINKING/ TWISTING WAS REPORTED. PATIENT OUTCOME: THE SUBJECT IS REPORTED TO HAVE DIED ON (B)(6) 2024. NO AUTOPSY WAS PERFORMED. THE PRIMARY CAUSE OF DEATH WAS CATEGORISED AS "DISEASE UNDER STUDY". THIS REPORT IS BEING SUBMITTED AS FOLLOW UP #1 FOR MANUFACTURING REPORT NUMBER 9612515-2025-00023 TO PROVIDE EVENT CLOSURE INFORMATION FOR (B)(4). Manufacturer Narrative: MANUFACTURER NARRATIVE: CLINICAL CODE: 4436- RUPTURED ANEURYSM - RUPTURE OF THE PARA VISCERAL AORTIC SEGMENT. 1696- ADULT RESPIRATORY DISTRESS SYNDROME- ACUTE RESPIRATORY DISTRESS SYNDROME. IMPACT CODE: 1802- DEATH - THE SUBJECT IS REPORTED TO HAVE DIED ON (B)(6) 2024. NO AUTOPSY WAS PERFORMED. THE PRIMARY CAUSE OF DEATH WAS CATEGORISED AS "DISEASE UNDER STUDY". 4625- ADDITIONAL SURGERY- THE EXTENSION PROCEDURE WAS REPORTED TO HAVE TAKEN PLACE ON THE (B)(6) 2024 (DEVICE LISTED: 2 X GORE CTAG, 3 X GORE VIABAHN). DEVICE PROBLEM CODE: 2889- DEFORMATION DUE TO COMPRESSIVE STRESS - THE LOCATION OF THE COMPRESSED RINGS CORRELATES WITH THE PREVIOUSLY DEMONSTRATED PROTRUDING STENT OF THE EXISTING THORACIC ENDOVASCULAR AORTIC REPAIR (TEVAR). COMPONENT CODE: 4755 - PART/COMPONENT/SUB-ASSEMBLY TERM NOT APPLICABLE. TYPE OF INVESTIGATION: 4110 - TREND ANALYSIS: A 4-YEAR REVIEW OF SIMILAR COMPLAINTS THORAFLEX HYBRID SALES JAN 21 - JAN 25 VS KINK TYPE JAN 21 - MAR 25 AN OCCURRENCE RATE OF (B)(4) NO TREND REQUIRING ACTION HAS BEEN IDENTIFIED. 4111 - COMMUNICATION INTERVIEW: HYBRID PRODUCT, AND INSTEAD MAKES REFERENCE TO KINKING IN A GORE TAG THORACIC STENT GRAFT. STENT THAT HAD KINK WAS BRIDGING THE PRIOR TO THORAFLEX TEVAR STENTS. 3331 - ANALYSIS OF PRODUCTION RECORDS: A FULL BATCH REVIEW WAS PERFORMED FOR TAG0183, AND NO ISSUES WERE IDENTIFIED DURING MANUFACTURING PROCESS FROM RAW MATERIALS TO FINISHED PRODUCTS. 4117 - DEVICE NOT ACCESSIBLE FOR TESTING: THE DEVICE WAS NOT EXPLANTED. 4112- ANALYSIS OF INFORMATION PROVIDED BY USER/THIRD PARTY: PRE OP IMAGING DEMONSTRATES PROTRUSION OF STENTS OF THE EXISTING TEVAR INTO THE AORTIC FLOW LUMEN. POST OP IMAGING SHOWS COMPRESSION OF THE 7TH AND 8TH STENT RING OF THE THORAFLEX HYBRID DEVICE AT THE LEVEL OF THE PROTRUDING STENT. INVESTIGATION FINDINGS: 213 - NO DEVICE PROBLEM FOUND- POST OP IMAGING SHOWS COMPRESSION OF THE 7TH AND 8TH STENT RING OF THE THORAFLEX HYBRID DEVICE AT THE LEVEL OF THE PROTRUDING STENT. THE LOCATION OF THE COMPRESSED RINGS CORRELATES WITH THE PREVIOUSLY DEMONSTRATED PROTRUDING STENT OF THE EXISTING TEVAR. INVESTIGATION CONCLUSION: 4323- DEVICE PROBLEM EXCLUDED- SITE CONFIRMED ON 03 APR 2025 AND 15 MAY 2025 THE DEVICE DEFICIENCY NO LONGER RELATES TO THE THORAFLEX HYBRID PRODUCT, AND INSTEAD MAKES REFERENCE TO KINKING IN A GORE TAG THORACIC STENT GRAFT. STENT THAT HAD KINK WAS BRIDGING THE PRIOR TO THORAFLEX TEVAR STENTS. Manufacturer Narrative: MANUFACTURER NARRATIVE: CLINICAL CODE: 4580 - INSUFFICIENT INFORMATION - DEVICE DEFICIENCY: GRAFT TWISTING/KINKING. IMPACT CODE: 1802- DEATH - THE SUBJECT IS REPORTED TO HAVE DIED ON (B)(6)2024. NO AUTOPSY WAS PERFORMED. THE PRIMARY CAUSE OF DEATH WAS CATEGORISED AS "DISEASE UNDER STUDY". 4625- ADDITIONAL SURGERY- THE EXTENSION PROCEDURE WAS REPORTED TO HAVE TAKEN PLACE ON THE (B)(6)2024 (DEVICE LISTED: 2 X GORE CTAG, 3 X GORE VIABAHN). DEVICE PROBLEM CODE: 2889- DEFORMATION DUE TO COMPRESSIVE STRESS - EVENT REPORTED AS THORAFLEX HYBRID GRAFT TWISTING/KINKING. COMPONENT CODE: 4755 - PART/COMPONENT/SUB-ASSEMBLY TERM NOT APPLICABLE. TYPE OF INVESTIGATION: 4110 - TREND ANALYSIS: A 4-YEAR REVIEW OF SIMILAR COMPLAINTS THORAFLEX HYBRID SALES JAN 21- JAN 25 VS KINK TYPE JAN 21 - MAR 25 AN OCCURRENCE RATE OF (B)(4). NO TREND REQUIRING ACTION HAS BEEN IDENTIFIED. 4111 - COMMUNICATION INTERVIEW: ADDITIONAL INFORMATION REQUESTED TO THE SITE 3331 - ANALYSIS OF PRODUCTION RECORDS: A REVIEW OF THE MANUFACTURING AND QUALITY CONTROL RECORDS FOR THIS BATCH CONFIRMED THAT THE PRODUCT WAS MANUFACTURED TO SPECIFICATION. 4117 - DEVICE NOT ACCESSIBLE FOR TESTING: THE DEVICE WAS NOT EXPLANTED. INVESTIGATION FINDINGS: 3233 - RESULTS PENDING COMPLETION OF INVESTIGATION. INVESTIGATION CONCLUSION: 11 - CONCLUSION NOT YET AVAILABLE AS INVESTIGATION IS ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21570995&pc=FRN | 2016493-2025-65186 | 2025-02-07 05:00 | Death | CAREFUSION SD | 2025-03-11 04:00 | FRN | ALARIS SYSTEM | Contamination; Pumping Stopped; Use of Device Problem | Insufficient Information | K211218 |  | 1 | Event Description: IT WAS INITIALLY REPORTED A REQUEST TO REVIEW THE DEVICE LOGS FOR (B)(6) 2025 BETWEEN THE HOURS OF 1320 AND 1640. NO FURTHER INFORMATION WAS PROVIDED AT THE TIME. BD RECEIVED ADDITIONAL INFORMATION THROUGH DUE DILIGENCE ATTEMPTS. IT WAS REPORTED THAT THE PATIENT WAS TRANSPORTED FROM EMERGENCY DEPARTMENT (ED) TO RADIOLOGY FOR APPROXIMATELY 15 MINUTES. LEVOPHED AND BICARB WAS INFUSING VIA PUMP AT THE TIME. UPON RETURN TO THE ED, THE PUMP WAS FOUND "NOT INFUSING." PER REPORT, THE PATIENT DIED AND HAD A DNR (DO NOT RESUSCITATE) CODE STATUS "WITH FAMILY DECISION OF NO FURTHER CARE." WHEN ASKED IF THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH, THE HOSPITAL'S PATIENT SAFETY MANAGER RN STATED THAT "CLARIFYING THE CONCERN IS IF THE PUMP SHUT DOWN DURING THE INFUSION POTENTIALLY DUE TO BATTERY LIFE AS IT WAS UNPLUGGED, WAS THE PUMP CONTINUOUSLY INFUSING AND THE DISPLAY WAS NOT VISIBLE. REALLY JUST SEEKING IF THE PUMP WAS CONTINUOUSLY INFUSING OR NOT AND IF NOT, WHY? THE PUMP NOR THE INFUSION(S) LED NOR CONTRIBUTED TO THE DEMISE OF THE PATIENT. WE ARE REVIEWING FOR PROCESS FAILURES INCLUDING TEAM MEMBER HANDOFF IN COMMUNICATION." THE CUSTOMER IS REQUESTING THE MANUFACTURER TO REVIEW THE LOGS, SPECIFICALLY IDENTIFYING "THE MEDICATION INFUSION DETAILS OF ML/HR. AND IF THERE WAS ANY TIME, THE MEDICATIONS WERE NOT INFUSING AS WE ARE REVIEWING FOR OUR INVESTIGATION OF PROCESSES. THE PATIENT WAS A DNR, WITH FAMILY PRESENTING THE DNR DOCUMENT UPON THEIR ARRIVAL TO THE ED. "Event Description: IT WAS INITIALLY REPORTED A REQUEST TO REVIEW THE DEVICE LOGS FOR (B)(6) 2025 BETWEEN THE HOURS OF 1320 AND 1640. NO FURTHER INFORMATION WAS PROVIDED AT THE TIME. BD RECEIVED ADDITIONAL INFORMATION THROUGH DUE DILIGENCE ATTEMPTS. IT WAS REPORTED THAT THE PATIENT WAS TRANSPORTED FROM EMERGENCY DEPARTMENT (ED) TO RADIOLOGY FOR APPROXIMATELY 15 MINUTES. LEVOPHED AND BICARB WAS INFUSING VIA PUMP AT THE TIME. UPON RETURN TO THE ED, THE PUMP WAS FOUND "NOT INFUSING." PER REPORT, THE PATIENT DIED AND HAD A DNR (DO NOT RESUSCITATE) CODE STATUS "WITH FAMILY DECISION OF NO FURTHER CARE." WHEN ASKED IF THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH, THE HOSPITAL'S PATIENT SAFETY MANAGER RN STATED THAT "CLARIFYING THE CONCERN IS IF THE PUMP SHUT DOWN DURING THE INFUSION POTENTIALLY DUE TO BATTERY LIFE AS IT WAS UNPLUGGED, WAS THE PUMP CONTINUOUSLY INFUSING AND THE DISPLAY WAS NOT VISIBLE. REALLY JUST SEEKING IF THE PUMP WAS CONTINUOUSLY INFUSING OR NOT AND IF NOT, WHY?...THE PUMP NOR THE INFUSION(S) LED NOR CONTRIBUTED TO THE DEMISE OF THE PATIENT. WE ARE REVIEWING FOR PROCESS FAILURES INCLUDING TEAM MEMBER HANDOFF IN COMMUNICATION." THE CUSTOMER IS REQUESTING THE MANUFACTURER TO REVIEW THE LOGS, SPECIFICALLY IDENTIFYING "THE MEDICATION INFUSION DETAILS OF ML/HR. AND IF THERE WAS ANY TIME, THE MEDICATIONS WERE NOT INFUSING AS WE ARE REVIEWING FOR OUR INVESTIGATION OF PROCESSES. THE PATIENT WAS A DNR, WITH FAMILY PRESENTING THE DNR DOCUMENT UPON THEIR ARRIVAL TO THE ED. " Manufacturer Narrative: THE ACTUAL DATE OF DEATH IS UNKNOWN. BD TECHNICAL SUPPORT TROUBLESHOOT WITH CUSTOMER OVER THE PHONE. THE EVENT LOGS HAVE BEEN RECEIVED AND THE EVALUATION IS PENDING. A FOLLOW UP REPORT WILL BE SUBMITTED WITH INVESTIGATION RESULTS ONCE THE EVALUATION HAS BEEN COMPLETED. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER. Manufacturer Narrative: OMIT : B17 - DEVICE NOT RETURNED, C20 - NO FINDINGS AVAILABLE, ADDITIONAL INFORMATION : DEVICE AVAILABLE FOR EVAL?, RETURNED TO MANUFACTURER ON, DEVICE EVAL BY MANUFACTURER?, IMDRF ANNEX A, G, B, C, D CODES AND MANUFACTURER NARRATIVE. A DEVICE HISTORY RECORD, COMPLAINT HISTORY REVIEW, AND RISK REVIEW ON FAILURE MODES ARE PERFORMED ON EACH DEVICE WITH REPORTABLE MALFUNCTION(S) ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN SECTION H6 OF THIS MDR REPORT. INVESTIGATION SUMMARY: THE CUSTOMER SOUGHT INFORMATION AROUND THE MEDICATION INFUSION DETAILS AND WHETHER THERE WAS ANY TIME WHERE THE MEDICATIONS WERE NOT INFUSING. REVIEW OF THE DEVICE LOGS IDENTIFIED INSTANCES OF OCCLUDED PATIENT SIDE ALARMS, BUT THERE WERE NO INDICATIONS OF DEVICE MALFUNCTIONS OR INTERRUPTIONS THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE PATIENT¿S CONDITION. INTERRUPTIONS IN INFUSION WERE NOT OBSERVED OR DURING DEVICE TESTING. INSPECTION AND TESTING OF THE SUSPECT PUMP MODULES AND PCU DID NOT FIND ANY EXISTING MALFUNCTIONS. THE SUSPECT PUMP MODULES WERE INFUSING WITHIN SPECIFICATIONS. THE PCU PASSED PREVENTATIVE MAINTENANCE TESTING AND WAS OPERATING AS INTENDED. REVIEW OF THE SUSPECT PCU BATTERY LOGS DID NOT SHOW ANY EVENTS OR MALFUNCTIONS THAT COULD HAVE CONTRIBUTED TO THE REPORTED INCIDENT. THERE WERE NO INSTANCES OF THE SYSTEM POWERING OFF DURING THE REPORTED INCIDENT TIME. THE INITIAL INFUSION PROGRAMMING OF NOREPINEPHRINE (DRUGID = 308), AT A CONCENTRATION OF 4 MG / 250 ML, WAS RECORDED ON 7FEB2025 AT 1:13 PM ON THE SUSPECT PUMP MODULE SN (B)(6). THE INFUSION RATE WAS SET AT 41.25 ML, THE VOLUME TO BE INFUSED (VTBI) WAS 100 ML. A BASIC INFUSION WAS PROGRAMMED ON (B)(6) 2025 AT 3:13 PM ON THE SUSPECT PUMP MODULE SN (B)(6) AT A RATE OF 100 ML/HR AND A VTBI OF 950 ML. A THIRD PUMP MODULE (SN (B)(6), NOT RECEIVED) WAS IDENTIFIED TO BE IN USE DURING THE REPORTED INCIDENT TIME. A BASIC INFUSION WAS PROGRAMMED AND STARTED ON (B)(6) 2025 AT 1:47 PM, WITH A RATE OF 999 ML/HR AND A VTBI OF 950 ML. AT 2:40 PM, THE CONCOMITANT PUMP MODULE SN (B)(6) WAS CHANNELED OFF. THEN AT 2:59 PM AN INFUSION FOR THE DRUG CEFTRIAXONE (DRUGID = 131), AT A CONCENTRATION OF 1000 MG / 50 ML, WAS PROGRAMMED AND STARTED. THE INFUSION RATE WAS 100 ML/HR AND THE VTBI WAS 50 ML. THE DOSE FOR THE INFUSION OF NOREPINEPHRINE ON THE PUMP MODULE SN (B)(6) WAS CHANGED TO 0.4 MCG/KG/HR AT 3:18 PM AND THEN CHANGED AGAIN TO 0.22 MCG/KG/HR AT 3:20 PM. THE DOSE WAS CHANGED TO 0.25 MCG/KG/HR AT 4:28 PM. AN AIR-IN-LINE ALARM WAS OBSERVED ON THE CONCOMITANT PUMP MODULE SN (B)(6), LASTING 2 MINUTES BEFORE THE UNIT WAS CHANNELED OFF. THE TOTAL VOLUME INFUSED (TVI) WAS 902.015 ML. FIVE (5) OCCLUDED PATIENT SIDE ALARMS WERE OBSERVED BETWEEN 1:13 PM AND 4:27 PM ON THE PUMP MODULE SN (B)(6). THE ALARM AT 2:00 PM LASTED 8 MINUTES BEFORE THE INFUSION WAS RESTARTED. THEN, AT 3:15 PM, THE ALARM LASTED 2 MINUTES AND 48 SECONDS BEFORE THE INFUSION WAS RESUMED. THE SUSPECT PUMP MODULE SN (B)(6) WAS CHANNELED OFF AT 3:46 PM, BUT WAS THEN POWERED ON AGAIN AT 4:24 PM AND AN INFUSION OF NOREPINEPHRINE AT A DOSE OF 0.22 MCG/KG/HR AND A VTBI OF 200 ML WAS PROGRAMMED AND STARTED AT 4:26 PM. THE SUSPECT PUMP MODULE SN (B)(6) WAS CHANNELED OFF AT 4:29 PM. THE TVI WAS RECORDED AT 101.491 ML. THE SUSPECT PUMP MODULE SN (B)(6) WAS CHANNELED OFF AT 4:30 PM. THE TVI WAS RECORDED AT 127.003 ML. IT IS NOT POSSIBLE TO DETERMINE WHAT THE ACTUAL VOLUME IS WITHIN AN IV CONTAINER AT THE START AND ENDING OF AN INFUSION FROM A REVIEW OF THE PCU EVENT LOG. IT IS ALSO NOT POSSIBLE TO DETERMINE WHAT THE FLUID IS THAT IS CONTAINED WITHIN THE IV CONTAINER. THIS IS NOT A FUNCTION OF THE PCU EVENT LOG; IT ONLY CAN REFLECT THE INPUTTED INFORMATION IT RECEIVES EITHER MANUALLY OR REMOTELY WITH RESPECT TO VOLUME TO BE DELIVERED AND FLUID SELECTED. THE DEVICE WAS REPORTEDLY IN USE FOR TREATMENT PURPOSES AS INTENDED PER 21 CFR 820.198(D)(2). ROOT CAUSE: THE ROOT CAUSE OF THE REPORTED INTERRUPTION OF INFUSIONS IS DUE TO THE USER POWERING OFF THE PUMP MODULES. IT WAS NOTED THAT THE SUSPECT PUMP MODULES SN (B)(6) AND SN (B)(6) WERE POWERED OFF ON (B)(6) 2025 AT 4:29 PM AND 4:30 PM, RESPECTIVELY. THE CONCOMITANT PUMP MODULE SN (B)(6) WAS POWERED OFF AT 3:28 PM. NO FUNCTIONALITY ISSUES OR ANOMALIES WERE OBSERVED WITH THE DEVICES THAT WOULD CONTRIBUTE TO THE REPORTED EVENT. THE RETURNED PUMP MODULES WERE FOUND TO BE INFUSING WITHIN SPECIFICATION. NOTE THAT THIS REPORT LISTS IMDRF ANNEX A1801, G04037, C0601 CODES NOT ASSOCIATED WITH THE REPORTED EVENT BUT IDENTIFIED AS REPORTABLE MALFUNCTIONS OBSERVED ON THE DEVICE DURING INVESTIGATION ARE UNRELATED TO THE REPORTED ISSUE. THESE OTHER FAILURES PRESUMPTIVELY DID NOT CAUSE OR CONTRIBUTE TO THE REPORTED ISSUE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21570996&pc=FRN | 2016493-2025-65187 | 2025-02-07 05:00 | Death | CAREFUSION SD | 2025-03-11 04:00 | FRN | ALARIS SYSTEM | Contamination; Pumping Stopped; Use of Device Problem | Insufficient Information | K211218 |  | 1 | Event Description: IT WAS INITIALLY REPORTED A REQUEST TO REVIEW THE DEVICE LOGS FOR (B)(6) 2025 BETWEEN THE HOURS OF 1320 AND 1640. NO FURTHER INFORMATION WAS PROVIDED AT THE TIME. BD RECEIVED ADDITIONAL INFORMATION THROUGH DUE DILIGENCE ATTEMPTS. IT WAS REPORTED THAT THE PATIENT WAS TRANSPORTED FROM EMERGENCY DEPARTMENT (ED) TO RADIOLOGY FOR APPROXIMATELY 15 MINUTES. LEVOPHED AND BICARB WAS INFUSING VIA PUMP AT THE TIME. UPON RETURN TO THE ED, THE PUMP WAS FOUND "NOT INFUSING." PER REPORT, THE PATIENT DIED AND HAD A DNR (DO NOT RESUSCITATE) CODE STATUS "WITH FAMILY DECISION OF NO FURTHER CARE." WHEN ASKED IF THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH, THE HOSPITAL'S PATIENT SAFETY MANAGER RN STATED THAT "CLARIFYING THE CONCERN IS IF THE PUMP SHUT DOWN DURING THE INFUSION POTENTIALLY DUE TO BATTERY LIFE AS IT WAS UNPLUGGED, WAS THE PUMP CONTINUOUSLY INFUSING AND THE DISPLAY WAS NOT VISIBLE. REALLY JUST SEEKING IF THE PUMP WAS CONTINUOUSLY INFUSING OR NOT AND IF NOT, WHY? THE PUMP NOR THE INFUSION(S) LED NOR CONTRIBUTED TO THE DEMISE OF THE PATIENT. WE ARE REVIEWING FOR PROCESS FAILURES INCLUDING TEAM MEMBER HANDOFF IN COMMUNICATION." THE CUSTOMER IS REQUESTING THE MANUFACTURER TO REVIEW THE LOGS, SPECIFICALLY IDENTIFYING "THE MEDICATION INFUSION DETAILS OF ML/HR. AND IF THERE WAS ANY TIME, THE MEDICATIONS WERE NOT INFUSING AS WE ARE REVIEWING FOR OUR INVESTIGATION OF PROCESSES. THE PATIENT WAS A DNR, WITH FAMILY PRESENTING THE DNR DOCUMENT UPON THEIR ARRIVAL TO THE ED. "Event Description: IT WAS INITIALLY REPORTED A REQUEST TO REVIEW THE DEVICE LOGS FOR (B)(6) 2025 BETWEEN THE HOURS OF 1320 AND 1640. NO FURTHER INFORMATION WAS PROVIDED AT THE TIME. BD RECEIVED ADDITIONAL INFORMATION THROUGH DUE DILIGENCE ATTEMPTS. IT WAS REPORTED THAT THE PATIENT WAS TRANSPORTED FROM EMERGENCY DEPARTMENT (ED) TO RADIOLOGY FOR APPROXIMATELY 15 MINUTES. LEVOPHED AND BICARB WAS INFUSING VIA PUMP AT THE TIME. UPON RETURN TO THE ED, THE PUMP WAS FOUND "NOT INFUSING." PER REPORT, THE PATIENT DIED AND HAD A DNR (DO NOT RESUSCITATE) CODE STATUS "WITH FAMILY DECISION OF NO FURTHER CARE." WHEN ASKED IF THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH, THE HOSPITAL'S PATIENT SAFETY MANAGER RN STATED THAT "CLARIFYING THE CONCERN IS IF THE PUMP SHUT DOWN DURING THE INFUSION POTENTIALLY DUE TO BATTERY LIFE AS IT WAS UNPLUGGED, WAS THE PUMP CONTINUOUSLY INFUSING AND THE DISPLAY WAS NOT VISIBLE. REALLY JUST SEEKING IF THE PUMP WAS CONTINUOUSLY INFUSING OR NOT AND IF NOT, WHY? THE PUMP NOR THE INFUSION(S) LED NOR CONTRIBUTED TO THE DEMISE OF THE PATIENT. WE ARE REVIEWING FOR PROCESS FAILURES INCLUDING TEAM MEMBER HANDOFF IN COMMUNICATION." THE CUSTOMER IS REQUESTING THE MANUFACTURER TO REVIEW THE LOGS, SPECIFICALLY IDENTIFYING "THE MEDICATION INFUSION DETAILS OF ML/HR. AND IF THERE WAS ANY TIME, THE MEDICATIONS WERE NOT INFUSING AS WE ARE REVIEWING FOR OUR INVESTIGATION OF PROCESSES. THE PATIENT WAS A DNR, WITH FAMILY PRESENTING THE DNR DOCUMENT UPON THEIR ARRIVAL TO THE ED. " Manufacturer Narrative: THE ACTUAL DATE OF DEATH IS UNKNOWN. BD TECHNICAL SUPPORT TROUBLESHOOT WITH CUSTOMER OVER THE PHONE. THE EVENT LOGS HAVE BEEN RECEIVED AND THE EVALUATION IS PENDING. A FOLLOW UP REPORT WILL BE SUBMITTED WITH INVESTIGATION RESULTS ONCE THE EVALUATION HAS BEEN COMPLETED. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER. Manufacturer Narrative: OMIT: B17 - DEVICE NOT RETURNED, C20 - NO FINDINGS AVAILABLE. ADDITIONAL INFORMATION: DEVICE AVAILABLE FOR EVAL? RETURNED TO MANUFACTURER ON, DEVICE EVAL BY MANUFACTURER? IMDRF ANNEX A,G,B,C AND D CODES AND MANUFACTURER NARRATIVE. A DEVICE HISTORY RECORD, COMPLAINT HISTORY REVIEW, AND RISK REVIEW ON FAILURE MODES ARE PERFORMED ON EACH DEVICE WITH REPORTABLE MALFUNCTION(S) ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN SECTION H6 OF THIS MDR REPORT. INVESTIGATON SUMMARY: THE CUSTOMER SOUGHT INFORMATION AROUND THE MEDICATION INFUSION DETAILS AND WHETHER THERE WAS ANY TIME WHERE THE MEDICATIONS WERE NOT INFUSING. REVIEW OF THE DEVICE LOGS IDENTIFIED INSTANCES OF OCCLUDED PATIENT SIDE ALARMS, BUT THERE WERE NO INDICATIONS OF DEVICE MALFUNCTIONS OR INTERRUPTIONS THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE PATIENT¿S CONDITION. INTERRUPTIONS IN INFUSION WERE NOT OBSERVED OR DURING DEVICE TESTING. INSPECTION AND TESTING OF THE SUSPECT PUMP MODULES AND PCU DID NOT FIND ANY EXISTING MALFUNCTIONS. THE SUSPECT PUMP MODULES WERE INFUSING WITHIN SPECIFICATIONS. THE PCU PASSED PREVENTATIVE MAINTENANCE TESTING AND WAS OPERATING AS INTENDED. REVIEW OF THE SUSPECT PCU BATTERY LOGS DID NOT SHOW ANY EVENTS OR MALFUNCTIONS THAT COULD HAVE CONTRIBUTED TO THE REPORTED INCIDENT. THERE WERE NO INSTANCES OF THE SYSTEM POWERING OFF DURING THE REPORTED INCIDENT TIME. THE INITIAL INFUSION PROGRAMMING OF NOREPINEPHRINE (DRUGID = 308), AT A CONCENTRATION OF 4 MG / 250 ML, WAS RECORDED ON (B)(6) 2025 AT 1:13 PM ON THE SUSPECT PUMP MODULE SN (B)(6). THE INFUSION RATE WAS SET AT 41.25 ML, THE VOLUME TO BE INFUSED (VTBI) WAS 100 ML. A BASIC INFUSION WAS PROGRAMMED ON (B)(6) 2025 AT 3:13 PM ON THE SUSPECT PUMP MODULE SN (B)(6) AT A RATE OF 100 ML/HR AND A VTBI OF 950 ML. A THIRD PUMP MODULE (SN (B)(6), NOT RECEIVED) WAS IDENTIFIED TO BE IN USE DURING THE REPORTED INCIDENT TIME. A BASIC INFUSION WAS PROGRAMMED AND STARTED ON (B)(6) 2025 AT 1:47 PM, WITH A RATE OF 999 ML/HR AND A VTBI OF 950 ML. AT 2:40 PM, THE CONCOMITANT PUMP MODULE SN (B)(6) WAS CHANNELED OFF. THEN AT 2:59 PM AN INFUSION FOR THE DRUG CEFTRIAXONE (DRUGID = 131), AT A CONCENTRATION OF 1000 MG / 50 ML, WAS PROGRAMMED AND STARTED. THE INFUSION RATE WAS 100 ML/HR AND THE VTBI WAS 50 ML. THE DOSE FOR THE INFUSION OF NOREPINEPHRINE ON THE PUMP MODULE SN (B)(6) WAS CHANGED TO 0.4 MCG/KG/HR AT 3:18 PM AND THEN CHANGED AGAIN TO 0.22 MCG/KG/HR AT 3:20 PM. THE DOSE WAS CHANGED TO 0.25 MCG/KG/HR AT 4:28 PM. AN AIR-IN-LINE ALARM WAS OBSERVED ON THE CONCOMITANT PUMP MODULE SN (B)(6), LASTING 2 MINUTES BEFORE THE UNIT WAS CHANNELED OFF. THE TOTAL VOLUME INFUSED (TVI) WAS 902.015 ML. FIVE (5) OCCLUDED PATIENT SIDE ALARMS WERE OBSERVED BETWEEN 1:13 PM AND 4:27 PM ON THE PUMP MODULE SN (B)(6). THE ALARM AT 2:00 PM LASTED 8 MINUTES BEFORE THE INFUSION WAS RESTARTED. THEN, AT 3:15 PM, THE ALARM LASTED 2 MINUTES AND 48 SECONDS BEFORE THE INFUSION WAS RESUMED. THE SUSPECT PUMP MODULE SN (B)(6) WAS CHANNELED OFF AT 3:46 PM, BUT WAS THEN POWERED ON AGAIN AT 4:24 PM AND AN INFUSION OF NOREPINEPHRINE AT A DOSE OF 0.22 MCG/KG/HR AND A VTBI OF 200 ML WAS PROGRAMMED AND STARTED AT 4:26 PM. THE SUSPECT PUMP MODULE SN (B)(6) WAS CHANNELED OFF AT 4:29 PM. THE TVI WAS RECORDED AT 101.491 ML. THE SUSPECT PUMP MODULE SN (B)(6) WAS CHANNELED OFF AT 4:30 PM. THE TVI WAS RECORDED AT 127.003 ML. IT IS NOT POSSIBLE TO DETERMINE WHAT THE ACTUAL VOLUME IS WITHIN AN IV CONTAINER AT THE START AND ENDING OF AN INFUSION FROM A REVIEW OF THE PCU EVENT LOG. IT IS ALSO NOT POSSIBLE TO DETERMINE WHAT THE FLUID IS THAT IS CONTAINED WITHIN THE IV CONTAINER. THIS IS NOT A FUNCTION OF THE PCU EVENT LOG; IT ONLY CAN REFLECT THE INPUTTED INFORMATION IT RECEIVES EITHER MANUALLY OR REMOTELY WITH RESPECT TO VOLUME TO BE DELIVERED AND FLUID SELECTED. ROOT CAUSE: THE ROOT CAUSE OF THE REPORTED INTERRUPTION OF INFUSIONS IS DUE TO THE USER POWERING OFF THE PUMP MODULES. IT WAS NOTED THAT THE SUSPECT PUMP MODULES SN (B)(6) AND SN (B)(6) WERE POWERED OFF ON (B)(6) 2025 AT 4:29 PM AND 4:30 PM, RESPECTIVELY. THE CONCOMITANT PUMP MODULE SN (B)(6) WAS POWERED OFF AT 3:28 PM. NO FUNCTIONALITY ISSUES OR ANOMALIES WERE OBSERVED WITH THE DEVICES THAT WOULD CONTRIBUTE TO THE REPORTED EVENT. THE RETURNED PUMP MODULES WERE FOUND TO BE INFUSING WITHIN SPECIFICATION. NOTE THAT THIS REPORT LISTS IMDRF ANNEX A1801, G04037, C0601 CODES NOT ASSOCIATED WITH THE REPORTED EVENT BUT IDENTIFIED AS REPORTABLE MALFUNCTIONS OBSERVED ON THE DEVICE DURING INVESTIGATION ARE UNRELATED TO THE REPORTED ISSUE. THESE OTHER FAILURES PRESUMPTIVELY DID NOT CAUSE OR CONTRIBUTE TO THE REPORTED ISSUE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21571772&pc=OZD | 1220648-2025-26615 | 2025-02-13 05:00 | Death | ABIOMED, INC. | 2025-03-11 04:00 | OZD | IMPELLA 5.5 | Increase in Pressure | Insufficient Information | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WAS ON AN IMPELLA 5.5 PUMP FOR SUPPORT. THE PATIENT WAS ADMITTED IN FOR HIGH RISK SURGERY. THE PUMP WAS PLACED BUT AFTER TWO DAYS WAS FOUND TO HAVE HIGH PURGE PRESSURE WHICH THE TEAM TROUBLESHOT WITH TISSUE PLASMINOGEN ACTIVATOR. THE PUMP REMAINED ON FOR SUPPORT FOR ALMOST 20 DAYS. THE PATIENT WAS NOT A CANDIDATE FOR TRANSPLANT AND SO THE CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE INVESTIGATION FOR THE HIGH PURGE PRESSURE HAS BEEN COMPLETED. NO PRODUCT WAS RETURNED FOR ANALYSIS. THE DATA LOGS DO NOT SHOW ANY MOTOR CURRENT SPIKES BUT THERE WAS A GRADUAL RISE IN PURGE PRESSURE BEFORE ALARM WAS TRIGGERED. THERE WERE CONTINUOUS SUCTION ALARMS AND AN EFFECT (DECREASE) IN PURGE FLOW SUGGESTING BIOMATERIAL BUILD UP. THE HIGH PURGE PRESSURE ALARMS PERSISTED FOR ABOUT AN HOUR BEFORE PURGE PRESSURE VALUES RETURNED TO NORMAL LEVELS. THIS ALIGNS WITH CLINICAL DETAIL OF TISSUE PLASMINOGEN ACTIVATOR (TPA) BEING ADDED TO THE PURGE AND ALARMS RESOLVING AFTER ADDITION. THERE WERE NO OTHER HIGH PURGE PRESSURE ALARMS FOR THE REMAINING DURATION OF THE CASE. THE ROOT CAUSE OF THE HIGH PURGE PRESSURE WAS MOST LIKELY BIOMATERIAL AS EVIDENCED BY LOG ANALYSIS AND TPA RESOLUTION. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST, SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA CATHETER: ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP. CLINICIANS SHOULD MONITOR MOTOR CURRENT AND CONSIDER REPLACING THE IMPELLA CATHETER WHENEVER A RISE IN MOTOR CURRENT IS SEEN.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21572024&pc=OZD | 1220648-2025-26618 | 2025-02-12 05:00 | Death | ABIOMED, INC. | 2025-03-11 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. AFTER THE IMPELLA WAS IMPLANTED, THE PATIENT WENT INTO VENTRICULAR TACHYCARDIA (VT). THE PATIENT WAS DEFIBRILLATED. THE PUMP WAS PULLED BACK FROM THE SEPTAL WALL AND THE VT RESOLVED. SUPPORT CONTINUED. ADDITIONAL INFORMATION NOTED THAT THE PATIENT'S FAMILY DECIDED TO WITHDRAW CARE. THE PATIENT SUBSEQUENTLY EXPIRED. THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA DEVICE AS AN ASSOCIATED FACTOR IN THE PATIENT'S DEMISE. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21572410&pc=OZD | 1220648-2025-26624 | 2025-02-14 05:00 | Death | ABIOMED, INC. | 2025-03-11 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT THE PATIENT'S LOWER EXTREMITIES WERE COLD TO THE TOUCH DURING IMPELLA CP SUPPORT. A DOPPLER STUDY SHOWED SIGNIFICANTLY DECREASED BLOOD FLOW TO THE BILATERAL LOWER EXTREMITIES. IT WAS DOCUMENTED THAT BOTH LEGS HAD SHUTDOWN AND A BILATERAL AMPUTATION WAS PLANNED FOR THE PATIENT. FASCIOTOMIES ON BOTH LEGS WERE PERFORMED, BUT THE DECISION WAS MADE BY THE FAMILY TO WITHDRAW CARE. AN AMPUTATION WAS NOT PERFORMED AND THE PATIENT PASSED AWAY. THERE IS NOT ENOUGH EVIDENCE TO DISASSOCIATE THE PATIENT'S PASSING WITH THE USE OF THE IMPELLA. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21572564&pc=OZD | 1220648-2025-26625 | 2025-02-14 05:00 | Death | ABIOMED, INC. | 2025-03-11 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hematoma | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS) AND EXPERIENCED A HEMATOMA REQUIRING BLOOD PRODUCTS. THE PATIENT WOULD SUBSEQUENTLY EXPIRE. ST-SEGMENT ELEVATION MYOCARDIAL INFARCTION WAS CONFIRMED, AND THE PATIENT HAD TWO STENTS PLACED. ON THE WAY TO THE UNIT, THE PATIENT DECOMPENSATED AND WAS BROUGHT BACK TO THE CATHETERIZATION LAB FOR IMPELLA CP DEVICE PLACEMENT. THE IMPELLA CP DEVICE WAS PLACED VIA THE RIGHT COMMON FEMORAL ARTERY WITHOUT INCIDENT. THE PATIENT WAS TRANSFERRED TO THE UNIT ON IMPELLA MCS. FOLLOWING THE IMPLANT, A HEMATOMA AT THE ACCESS SITE WAS NOTED. BLOOD PRODUCTS WERE REQUIRED AND MANUAL PRESSURE AND PRESSURE DRESSING WERE APPLIED. CARE GOALS WERE DISCUSSED WITH THE FAMILY AND THE PATIENT WAS MADE DO NOT RESUSCITATE WITH CARE WITHDRAWN. THE PATIENT WOULD EXPIRE, THEREAFTER. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM, SECTION: GENERAL PATIENT CARE CONSIDERATIONS: ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT, USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21572576&pc=LZG | 3004464228-2025-10453 | 2024-09-12 04:00 | Death | INSULET CORPORATION | 2025-03-11 04:00 | LZG | OMNIPOD DASH INSULIN MANAGEMENT SYSTEM | Insufficient Information | Hyperglycemia; Diabetic Ketoacidosis | K211575 |  | 1 | Event Description: IT HAS BEEN REPORTED THAT THE PATIENT HAS PASSED AWAY ON (B)(6) 2024. THE PATIENT REPORTEDLY DECEASED DUE TO DIABETIC KETOACIDOSIS (DKA). IT WAS ALSO REPORTED THAT INSULIN DELIVERY STOPPED ON (B)(6) 2024, WITH THE REASON FOR THIS BEING UNKNOWN. INFORMATION ON ANY PRODUCT ISSUES THE CUSTOMER EXPERIENCED WAS NOT PROVIDED. ACCORDING TO THE HOSPITAL, THE PATIENT'S BLOOD OXYGEN LEVELS WERE ACIDOTIC AND THERE WERE HIGH LEVELS OF TRAMADOL PRESENT. IT WAS ALSO REPORTED THAT THE PATIENT HAD A HISTORY OF DIABETES RELATED COMPLICATIONS THAT INCLUDED HIGH HBA1C (HEMOGLOBIN A1C), GASTROPARESIS, UVEITIS (TREATED WITH HIGH DOSE STEROIDS) AND RETINOPATHY. Manufacturer Narrative: ADDITIONAL INFORMATION, DEVICE INFORMATION, D4, H4. LOT RELEASE RECORDS WERE REVIEWED AND THE PRODUCT LOT MET ALL ACCEPTANCE CRITERIA. IN THE CASE DESCRIPTION, A DEATH DUE TO DIABETIC KETOACIDOSIS IS REPORTED. IT IS MENTIONED THAT THE USER'S GLOOKO DATA SHOWED INSULIN DELIVERY STOPPING ON (B)(6) 2024 WITH NO REASON KNOWN. REVIEW OF THE. IBF AND ANDROID LOG FILES DOWNLOADED FROM THE RETURNED PDM FOUND THAT THE SYSTEM WAS DELIVERING INSULIN AS EXPECTED. THE. IBF FILE SHOWED NO RECORDS OF BASAL INSULIN DELIVERY AFTER THE LAST POD ACTIVATION AT 13:19 ON (B)(6) 2024. THE AUDIT LOG FILES SHOW THAT THE TOTAL PULSE COUNT RECORDED BY THE POD STATUS WAS INCREASING WHEN NO BOLUSES WERE DELIVERED, INDICATING THAT THE POD WAS STILL DELIVERING THE USER'S BASAL PROGRAM. THE LAST POD STOPPED COMMUNICATING WITH THE PDM AT APPROXIMATELY 21:00 ON (B)(6) 2024. THE PDM KEPT ATTEMPTING TO ESTABLISH CONNECTION TO THE POD UNTIL IT WAS POWERED OFF ON (B)(6) 2024. THE LACK OF COMMUNICATION WITH THE LAST POD USED AND THE SUBSEQUENT DISCARD RESULTED IN THE BASAL RECORDS BEING INVALIDATED IN THE FILE. AS THE. IBF FILE POPULATES THE GLOOKO REPORT, THE INVALIDATED BASAL RECORDS RESULTED IN THE APPEARANCE OF NO BASAL INSULIN BEING DELIVERED AFTER (B)(6) 2025. DURING PHYSICAL TESTING, THE PDM WAS PAIRED WITH AN UNUSED POD. THE PDM WAS ABLE TO ACTIVATE THE POD, COMMAND BOLUSES, AND DEACTIVATE THE POD WITH NO ISSUES. ALL BOLUS, BASAL, AND BLOOD GLUCOSE RECORDS CREATED DURING TESTING WERE SUCCESSFULLY CAPTURED IN THE PDM HISTORY. NO EVIDENCE OF ISSUES WITH INSULIN DELIVERY WERE OBSERVED DURING THE INVESTIGATION. Manufacturer Narrative: ACCORDING TO THE COMPLAINANT THE DEVICE WILL NOT BE RETURNED FOR INVESTIGATION. WE ARE UNABLE TO DETERMINE IF ANY PRODUCT CONDITION COULD HAVE CONTRIBUTED TO THE REPORTED DEATH AND DIABETIC KETOACIDOSIS. NO LOT RELEASE RECORDS WERE REVIEWED, AS THE PRODUCT LOT NUMBER WAS NOT PROVIDED. HIGH BLOOD GLUCOSE IS A COMMON SYMPTOM FOR PEOPLE WITH DIABETES (GLUCOSE MONITORING DATA FROM PEOPLE WITH DIABETES INDICATE THAT ON AVERAGE, THEY CAN EXPERIENCE BLOOD GLUCOSE LEVELS ABOVE 250 MG/DL FOR 14-25% OF THE TIME[1][2][3]..), AND IT WOULD BE CHALLENGING TO SPECULATE ON A CAUSE FOR THE COMPLAINTS WITHOUT RECEIVING THE DEVICES BACK FOR AN ENGINEERING INVESTIGATION. [1] BECK RW, BERGENSTAL RM, CHENG P, KOLLMAN C, CARLSON AL, JOHNSON ML, RODBARD D. THE RELATIONSHIPS BETWEEN TIME IN RANGE, HYPERGLYCEMIA METRICS, AND HBA1C. J DIABETES SCI TECHNOL 2019;13:614-626. [1] WELSH JB, DERDZINSKI M, PARKER AS, PUHR S, JIMENEZ A, WALKER T. REAL-TIME SHARING AND FOLLOWING OF CONTINUOUS GLUCOSE MONITORING DATA IN YOUTH. DIABETES THER 2019; 10:751-755. [1] PUHR S, DERDZINSKI M, WELSH JB, PARKER AS, WALKER T, PRICE DA. REAL-WORLD HYPOGLYCEMIA AVOIDANCE WITH A CONTINUOUS GLUCOSE MONITORING SYSTEM'S PREDICTIVE LOW GLUCOSE ALERT. DIABETES TECHNOL THER 2019; 21:155-158.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21572780&pc=MCX | 2124215-2025-14604 | 2025-02-18 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-03-11 04:00 | MCX | ROTAPRO | Adverse Event Without Identified Device or Use Problem | Ischemia; Myocardial Infarction |  |  | 1 | Event Description: IT WAS REPORTED A PATIENT DEATH OCCURRED. DURING THE PROCEDURE IN THE RIGHT CORONARY ARTERY (RCA), OSTIUM SUBTOTAL, THE LESION WAS MODERATELY TORTUOUS, SEVERELY CALCIFIED AND 90% STENOSED. A 6F GUIDING CATHETER WAS SELECTED FOR TREATMENT , AND THEN REPLACED WITH A 7F CATHETER. INITIALLY A NON-BOSTON SCIENTIFIC WIRE WAS USED, THEN CHANGED TO A ROTAWIRE FLOPPY VIA A MICROCATHETER, ALONG WITH A 1.75 MM ROTAPRO. FOUR RUNS OF ROTATIONAL ATHERECTOMY WERE CONDUCTED, EACH ABOUT 20 SECONDS, WITH A MAXIMUM RPM DROP OF 15 K. HOWEVER, NO REFLOW OCCURRED, LEADING TO BALLOON ANGIOPLASTY AND IMPLANTATION OF TWO STENTS. THERE WAS NO ANTEGRADE FLOW OBSERVED IN THE CORONARY ARTERY (NO TIMI FLOW), CAUSING AV BLOCK, AND REQUIRING A TEMPORARY PACEMAKER. COMPLICATIONS THAT WERE ENCOUNTERED INCLUDED MYOCARDIAL INFARCTION/ISCHEMIA AND FURTHER INTERVENTIONS INCLUDED ADMINISTRATION OF ISOPREL, AGOSTAT, KATECHOMALIN, ATROPIN, AND USE OF A TEMPORARY PACEMAKER. FOLLOWING THE PROCEDURE, THE PATIENT WAS TRANSFERRED TO THE CRITICAL CARE UNIT (CCU). THOUGH NO DEVICE MALFUNCTIONS OCCURRED, THE PHYSICIAN IS UNSURE IF THE DEVICE OR PROCEDURE CONTRIBUTED TO THE COMPLICATIONS. HOWEVER, FURTHER INFORMATION CONFIRMED DEATH OF THE PATIENT ON AN UNKNOWN DATE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21572812&pc=OZD | 1220648-2025-26630 | 2025-02-13 05:00 | Death | ABIOMED, INC. | 2025-03-11 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. IT FOLLOWING 19 HOURS OF IMPELLA SUPPORT, THE PATIENT DECOMPENSATED OVERNIGHT AND WENT INTO VENTRICULAR TACHYCARDIA AND ARRESTED. TEAM WAS UNABLE TO RESUSCITATE THE PATIENT. THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21572839&pc=DRE | MW5167476 | 2025-02-26 05:00 | Death | COOK MEDICAL / COOK VANDERGRIFT INC. | 2025-03-11 04:00 | DRE | EVOLUTION RL CONTROLLED-ROTATION DILATOR SHEATH SET |  | Rupture |  |  | 1 | Event Description: THE SHEATH BROKE. WHEN IT WAS REMOVED IT SHEARED THE SUBCLAVIAN VEIN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21572851&pc=MCX | MW5167477 | 2025-02-07 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-03-11 04:00 | MCX | ROTAPRO PRE-CONNECTED BURR AND ADVANCING DEVICE |  | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: PROCEDURE COMPLICATED BY ROTATIONAL ATHERECTOMY BUR ENTRAPMENT WHICH COULD NOT BE REMOVED BY VARIOUS TECHNIQUES. INTRA-AORTIC BALLOON PUMP PLACED. THEREFORE REFERRED FOR EMERGENCY CABG TO REMOVE ROTA WIRE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21572883&pc=NIK | 9614453-2025-00981 | 2025-02-12 05:00 | Death | MEDTRONIC EUROPE SARL | 2025-03-11 04:00 | NIK | CROME¿ HF QUAD CRT-D MRI SURESCAN¿ | Use of Device Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P010031 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY LESS THAN A MONTH AFTER A DEVICE CHANGE OUT PROCEDURE. THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) WAS SENT TO THE CLINIC FOR REVIEW. IT WAS DETERMINED THAT THE IMPLANT DETECTION WAS STILL IN PROGRESS BECAUSE THE VENTRICULAR FIBRILLATION (VF) ZONE WAS NEVER PROGRAMMED ON, RESULTING IN NO DATA BEING COLLECTED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE DEVICE WAS NOT RETURNED FOR ANALYSIS, HOWEVER, PERFORMANCE DATA COLLECTED FROM THE DEVICE WAS RECEIVED AND ANALYZED. ANALYSIS OF THE DEVICE MEMORY HAD AN OBSERVATION RELATING TO IMPLANT DETECT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE DEVICE WAS RETURNED AND ANALYZED. RETURNED PRODUCT ANALYSIS WAS PERFORMED AND NO ANOMALIES WERE FOUND. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21573173&pc=NPT | 2135147-2025-01211 | 2025-02-19 05:00 | Death | ABBOTT MEDICAL | 2025-03-11 04:00 | NPT | NAVITOR VISION | Migration or Expulsion of Device | Cardiac Arrest; Obstruction/Occlusion; Vascular Dissection | P190023 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 27MM NAVITOR VALVE WAS SELECTED FOR IMPLANT. THE PATIENT HAD AORTIC STENOSIS, AORTIC INSUFFICIENCY AND LEFT BUNDLE BRANCH BLOCK (LBBB). WITH MILD CALCIUM WITH A VERY HORIZONTAL ROOT. THE PATIENT'S AORTIC VALVE ANGLE WAS NOTED AS 79 DEGREES. IT WAS NOTED THAT THERE WAS MILD CALCIUM. PRE-IMPLANTATION BALLOON AORTIC VALVULOPLASTY (BAV) WAS DONE WITH A 22MM NON-ABBOTT BALLOON. THE ANNULAR DIMENSIONS OF THE NATIVE VALVE WERE: ANNULUS MIN 22.6MM, MAX 24.3MM, AVG 23.5MM. AREA 429.2MM2, PERIMETER 74.2MM, LVOT 71.7MM PERIMETER MIN 20.7MM, MAX 24.1MM, AVG 22.4MM. DURING THE PROCEDURE, WHILE THE DEVICE WAS BEING DEPLOYED, THE VALVE POPPED UP TOWARDS THE AORTA. IT WAS NOTED THAT THE VIEW FOR CUSP OVERLAP WAS NOT GOOD AS THE HEART WAS ROTATED AND THERE WAS A STEEP ROOT ANGLE. WITH THE ANGLES AVAILABLE, IT WAS ASSESSED THAT THE IMPLANT DEPTH OF THE VALVE WAS 3MM ON THE NON-CORONARY CUSP (NCC) AND 5 ON THE LEFT CORONARY CUSP (LCC). ALL THREE TABS WERE CONFIRMED TO BE RELEASED IN TWO VIEWS. THE VALVE WAS NEVER RESHEATHED MORE THAN TWO TIMES. THE CAUSE OF THE MIGRATION WAS UNDETERMINED. THE PHYSICIAN PLANNED TO PERFORM A VALVE-IN-VALVE WITH A 23MM NON-ABBOTT VALVE. THE PHYSICIAN WAS ADVISED TO SNARE THE NAVITOR OUT OF THE WAY AND TO MOVE IT TO THE SINOTUBULAR JUNCTION (STJ), BELOW MAJOR VESSELS AS TO NOT RISK CORONARY OBSTRUCTION. THE PHYSICIAN PLACES THE NON-ABBOTT VALVE IN POSITION AT THE NATIVE ANNULUS AND THEN ATTEMPTED TO SNARE THE NAVITOR VALVE. ONLY ONE GOOSENECK SNARE WAS USED THE VALVE WAS ABLE TO BE MOVED UP SLIGHTLY, BUT THEN THE SNARE ON THE VALVE WAS LOST AND LEFT THE NAVITOR VALVE BLOCKING THE CORONARIES. THE PATIENT LOST HEMODYNAMIC STABILITY AND CHEST COMPRESSIONS WERE STARTED. ADDITIONAL ATTEMPTS WERE USED WITH THE GOOSENECK TO SNARE THE VALVE FOLLOWED BY A TULIP SNARE WHILE CHEST COMPRESSIONS WERE OCCURRING, BUT WAS NOT SUCCESSFUL. THE PHYSICIAN ASKED FOR PERFUSION TO PREP FOR ECMO (EXTRACORPOREAL MEMBRANE OXYGENATION) AND THE 23MM NON-ABBOTT VALVE WAS IMPLANTED WITHOUT ANY HEMODYNAMIC RECOVERY. THE PHYSICIAN THEN ATTEMPTED TO USE THE NON-ABBOTT VALVES BALLOON TO WALK UP THE NAVITOR SEVERAL TIMES UNSUCCESSFULLY. THIS WAS AGAIN ATTEMPTED WITH A 26MM NON-ABBOTT BALLOON AND THE NAVITOR WAS ABLE TO BE MOVED UP ENOUGH TO FREE THE LEFT MAIN, BUT NOT THE RIGHT. AS THE PATIENT WAS LEFT DOMINATE, THIS WASN'T A CONCERN. THE PHYSICIAN WANTED TO SNARE THE NAVITOR UP A LITTLE MORE AND WANTED TO PUT A WIRE THROUGH THE NAVITOR AND THEN SNARE THE WIRE. AS A RESULT OF THIS, THE TUBING FOR THE ECMO WAS PUNCTURED AND THE PERFUSIONIST WAS UNABLE TO HOLD HEMODYNAMIC STABILITY. THE PATIENT'S RHYTHM FLAT LINED AND THE PHYSICIANS DISCONTINUED LIFE SAVING MEASURES. THE PATIENT PASSED AWAY ON (B)(6) 2025.Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 27MM NAVITOR VALVE WAS SELECTED FOR IMPLANT. THE PATIENT HAD AORTIC STENOSIS, AORTIC INSUFFICIENCY AND LEFT BUNDLE BRANCH BLOCK (LBBB). WITH MILD CALCIUM WITH A VERY HORIZONTAL ROOT. THE PATIENT'S AORTIC VALVE ANGLE WAS NOTED AS 79 DEGREES. IT WAS NOTED THAT THERE WAS MILD CALCIUM. PRE-IMPLANTATION BALLOON AORTIC VALVULOPLASTY (BAV) WAS DONE WITH A 22MM NON-ABBOTT BALLOON. THE ANNULAR DIMENSIONS OF THE NATIVE VALVE WERE: ANNULUS MIN 22.6MM, MAX 24.3MM, AVG 23.5MM. AREA 429.2MM2, PERIMETER 74.2MM, LVOT 71.7MM PERIMETER MIN 20.7MM, MAX 24.1MM, AVG 22.4MM. DURING THE PROCEDURE, WHILE THE DEVICE WAS BEING DEPLOYED, THE VALVE POPPED UP TOWARDS THE AORTA. IT WAS NOTED THAT THE VIEW FOR CUSP OVERLAP WAS NOT GOOD AS THE HEART WAS ROTATED AND THERE WAS A STEEP ROOT ANGLE. WITH THE ANGLES AVAILABLE, IT WAS ASSESSED THAT THE IMPLANT DEPTH OF THE VALVE WAS 3MM ON THE NON-CORONARY CUSP (NCC) AND 5 ON THE LEFT CORONARY CUSP (LCC). ALL THREE TABS WERE CONFIRMED TO BE RELEASED IN TWO VIEWS. THE VALVE WAS NEVER RESHEATHED MORE THAN TWO TIMES. THE CAUSE OF THE MIGRATION WAS UNDETERMINED. THE PHYSICIAN PLANNED TO PERFORM A VALVE-IN-VALVE WITH A 23MM NON-ABBOTT VALVE. THE PHYSICIAN WAS ADVISED TO SNARE THE NAVITOR OUT OF THE WAY AND TO MOVE IT TO THE SINOTUBULAR JUNCTION (STJ), BELOW MAJOR VESSELS AS TO NOT RISK CORONARY OBSTRUCTION. THE PHYSICIAN PLACES THE NON-ABBOTT VALVE IN POSITION AT THE NATIVE ANNULUS AND THEN ATTEMPTED TO SNARE THE NAVITOR VALVE. ONLY ONE GOOSENECK SNARE WAS USED THE VALVE WAS ABLE TO BE MOVED UP SLIGHTLY, BUT THEN THE SNARE ON THE VALVE WAS LOST AND LEFT THE NAVITOR VALVE BLOCKING THE CORONARIES. THE PATIENT LOST HEMODYNAMIC STABILITY AND CHEST COMPRESSIONS WERE STARTED. ADDITIONAL ATTEMPTS WERE USED WITH THE GOOSENECK TO SNARE THE VALVE FOLLOWED BY A TULIP SNARE WHILE CHEST COMPRESSIONS WERE OCCURRING BUT WAS NOT SUCCESSFUL. THE USE OF THE SINGLE SNARE USED TO PULL BACK THE NAVITOR INTO THE ASCENDING AORTA CAUSED INJURY TO THE GREATER CURVATURE OF THE AORTA WITH THE CROWN BEING EMBEDDED IN THE AORTIC WALL AND BENT BACK WHICH RESULTED IN AN AORTIC DISSECTION. THE PHYSICIAN ASKED FOR PERFUSION TO PREP FOR ECMO (EXTRACORPOREAL MEMBRANE OXYGENATION) AND THE 23MM NON-ABBOTT VALVE WAS IMPLANTED WITHOUT ANY HEMODYNAMIC RECOVERY. THE PHYSICIAN THEN ATTEMPTED TO USE THE NON-ABBOTT VALVES BALLOON TO WALK UP THE NAVITOR SEVERAL TIMES UNSUCCESSFULLY. THIS WAS AGAIN ATTEMPTED WITH A 26MM NON-ABBOTT BALLOON AND THE NAVITOR WAS ABLE TO BE MOVED UP ENOUGH TO FREE THE LEFT MAIN, BUT NOT THE RIGHT. AS THE PATIENT WAS LEFT DOMINATE, THIS WASN'T A CONCERN. THE PHYSICIAN WANTED TO SNARE THE NAVITOR UP A LITTLE MORE AND WANTED TO PUT A WIRE THROUGH THE NAVITOR AND THEN SNARE THE WIRE. AS A RESULT OF THIS, THE TUBING FOR THE ECMO WAS PUNCTURED AND THE PERFUSIONIST WAS UNABLE TO HOLD HEMODYNAMIC STABILITY. THE PATIENT'S RHYTHM FLAT LINED AND THE PHYSICIANS DISCONTINUED LIFE SAVING MEASURES. THE PATIENT PASSED AWAY ON (B)(6) 2025. Manufacturer Narrative: IT WAS REPORTED THAT DURING NAVITOR IMPLANT PROCEDURE THE VALVE POPPED UP TOWARDS THE AORTA WHILE THE DEVICE WAS BEING DEPLOYED. A VALVE-IN-VALVE THE NON-ABBOTT VALVE WAS PLACED IN POSITION AT THE NATIVE ANNULUS AND THEN THE NAVITOR VALVE WAS ATTEMPTED TO BE SNARED. THE PATIENT LOST HEMODYNAMIC STABILITY AND CHEST COMPRESSIONS WERE STARTED. A SINGLE SNARE WAS USED TO PULL BACK THE NAVITOR INTO THE ASCENDING AORTA; THIS CAUSED INJURY TO THE GREATER CURVATURE OF THE AORTA WITH THE CROWN BEING EMBEDDED IN THE AORTIC WALL AND BENT BACK WHICH RESULTED IN AN AORTIC DISSECTION. A 23 MM NON-ABBOTT VALVE WAS IMPLANTED WITHOUT ANY HEMODYNAMIC RECOVERY. A SECOND SNARE WAS ATTEMPTED. HOWEVER, THE PATIENT WAS HEMODYNAMICALLY UNSTABLE AND CPR WAS REQUIRED. DESPITE ECMO, THE PATIENT DID NOT SURVIVE. THE DEVICE REMAINED IMPLANTED AND WILL NOT RETURN TO ABBOTT FOR ANALYSIS. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED, AND THE PRODUCT MET ALL SPECIFICATIONS. A REVIEW OF CINE REVIEW WAS PERFORMED AT ABBOTT. THE PATIENT HAD CHALLENGING ANATOMY WITH ~80 DEG HORIZONTAL AORTA. INITIAL NAVITOR IMPLANT POSITION APPEARED HIGH DURING PARTIAL DEPLOYMENT. IMPLANT WAS RELEASED WHILE AT ANNULUS OR SLIGHTLY ABOVE ON THE NCC SIDE AND 1MM DEPTH ON THE LCC SIDE. RELEASING AT HIGH IMPLANT POSITION INCREASES RISK OF VALVE MIGRATION/EMBOLIZATION. BASED ON INFORMATION AVAILABLE, THE CAUSE OF REPORTED DEVICE MIGRATION IS POSSIBLY RELATED TO THE CHALLENGING PATIENT ANATOMY WHICH MAY HAVE MADE IT DIFFICULT TO IMPLANT THE VALVE AT RECOMMENDED IMPLANT DEPTH OF 3 MM. THE EXACT CAUSE OF REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THE REPORTED SURGICAL INTERVENTION WAS A RESULT OF CASE-SPECIFIC CIRCUMSTANCES AS VALVE-IN-VALVE WAS PERFORMED. THE REPORTED COMPLICATIONS POST IMPLANT AND PATIENT DEATH APPEARS TO BE A CASCADING EFFECT. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO LABELING, DESIGN, OR MANUFACTURING OF THE DEVICE. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21573807&pc=NPT | 2025587-2025-01813 | 2025-01-29 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-11 04:00 | NPT | EVOLUT FX VALVE | Fluid/Blood Leak; Material Invagination; Patient Device Interaction Problem | Low Blood Pressure/ Hypotension; Aortic Valve Insufficiency/ Regurgitation | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC HEART VALVE PROCEDURE, USING TRANSFEMORAL ACCESS, THE 29 MM VALVE WAS ATTEMPTED TO BE IMPLANTED BUT INFOLDED. THE VALVE WAS RECAPTURED AND WITHDRAWN FROM THE PATIENT. THE PATIENT BECAME HYPOTENSIVE WITH SEVERE AORTIC INSUFFICIENCY. A SECOND 29 MM VALVE WAS PREPPED BUT WOULD NOT CROSS THE PATIENT¿S VALVE DURING CARDIOPULMONARY RESUSCITATION (CPR). CPR CONTINUED FOR APPROXIMATELY 20 MINUTES. A 23 MM NON-MEDTRONIC (EDWARDS SAPIEN 3 ULTRA) TRANSCATHETER VALVE WAS PREPPED AND IMPLANTED WITHOUT ISSUE OR LEAKS. THE NON-MEDTRONIC VALVE IMPLANT WAS CONSIDERED A BAILOUT IMPLANT AS IT TOOK PLACE DURING AN ACTIVE CODE EVENT WITH ACTIVE RESUSCITATION EFFORTS FROM THE PREVIOUS FAILED IMPLANT ATTEMPTS. THE PATIENT DID NOT RECOVER AND EXPIRED. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID D-EVOLUTFX-2329 (UNKNOWN SERIAL/LOT); PRODUCT TYPE: 0195-HEART VALVES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21573973&pc=MIH | 9612164-2025-01276 | 2025-03-06 05:00 | Death | MEDTRONIC IRELAND | 2025-03-11 04:00 | MIH | VALIANT CAPTIVIA - FF | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Hemorrhage/Bleeding; Perforation of Vessels | P100040 |  | 1 | Event Description: TWO VALIANT CAPTIVIA STENT GRAFTS WERE IMPLANTD DURING THE ENDOVASCULAR TREATMENT OF A 48MM TAA. IT WAS REPORTED THAT DURING THE INDEX PROCEDURE VAMF3834C150TJ WAS ATTEMPTED TO BE ADVANCED ONCE ACCESS WAS OBTAINED IN THE RIGHT FEMORAL ARTERY AND A CLOSURE DEVICE USED , BUT IT DID NOT ADVANCE NEAR THE EXTERNAL ILIAC ARTERY TO COMMON ILIAC ARTERY. IT WAS REMOVED AND A NON MDT SHEATH WAS ADVANCED FROM THE SAME RIGHT FEMORAL ARTERY. THE SHEATH WAS ADVANCED TO THE RENAL ARTERY LEVEL. THE VAMF3834C150TJ WAS INSERTED INTO THE SHEATH AND PLACED IN DESCENDING POSITION AS SCHEDULED. THE NEXT STENT GRAFT VAMF4040C150TJ WAS LANDED AT THE DISTAL ARCH, AND CONTRAST IMAGING WAS PERFORMED. DURING THE CONTRAST IMAGING, THE CONTRAST AGENT FLOW WAS POOR AND HAD RETAINED IN THE ARCH FOR A LONG TIME. VAMF4040C150TJ WAS PLACED AS SCHEDULED AND CONTRAST IMAGING WAS PERFORMED. THE TEVAR ITSELF WAS COMPLETED. THE ACCESS WAS CHECKED AND A CONTRAST AGENT LEAKED OUT OF THE BLOOD VESSEL AROUND THE RIGHT EIA-CIA, CONFIRMING DAMAGE TO THE ACCESS VESSEL . THE PATIENT SUDDENLY DEVELOPED SHOCK AND CARDIAC ARREST. CARDIAC MASSAGE WAS INITIATED. IT WAS DECIDED TO PLACE A NON-MDT IMMEDIATELY . THE FIRST NON MDT STENT 10MM X 100MM WAS PLACED IN CIA-EIA. AS THE DAMAGE REMAINED IN THE CONTRAST IMAGING AFTER PLACEMENT ANOTHER NON MDT STENT 9 MM X 100MM WAS ADDED PROXIMALLY SLIGHTLY FROM THE FEMORAL HEAD LEVEL. IT WAS NOTED THAT THE NATIVE BLOOD VESSEL DIAMETER IS ABOUT 8MM. THERE WAS A TYPE IIIA ENDOLEAK BETWEEN THE NON MDT STENTS, SO TOUCH UP WITH A NON-MDT BALLOON. THE IIIA ENDOLEAK DID NOT RESOLVE. WHEN THE 8FR SHEATH WAS ADVANCED TO PLACE A 3RD NON MDT STENT 11 MM STENT, THE JUNCTION OF THE ALREADY IMPLANTED NON MDT STENTS WAS DISCONNECTED. THE NON MDT, WHICH HAD BEEN PLACED PROXIMALLY, HAD BEEN COMPLETELY PUSHED OUT OF THE BLOOD VESSEL, AND THE GUIDEWIRE, WHICH HAD BEEN PASSING THROUGH THE TRUE LUMEN, HAD ALSO ESCAPED, MAKING CANNULATION IMPOSSIBLE. A RESCUE BALLOON WAS ADVANCED FROM THE OPPOSITE SIDE (LEFT) TO OCCLUDE. THE FOLLOWING PROCEDURES WERE BEING CONSIDERED, BUT THE CARDIAC MASSAGE COULD NOT BE STOPPED WHILE THE PATIENT REMAINED IN CARDIAC ARREST. BLEEDING FROM THE INTUBATION TUBE WAS ALSO CONFIRMED. TWO (2) HOURS HAD PASSED SINCE THE CARDIAC ARREST, AND THE PATIENT'S HEART RATE DID NOT RESUME EVEN ONCE; THE PHYSICIAN DECIDED TO STOP RESUSCITATION AND CONFIRMED THE PATIENT'S DEATH. PER THE PHYSICIAN THE CAUSE OF THE ACCESS DAMAGE WAS ANATOMY RELATED DUE TO THE PATIENT'S VASCULAR MORPHOLOGY AND THE ACCESS VESSEL WAS NARROW AT ABOUT 7.2MM -8MM. NO ADDITIONAL CLINICAL SEQUALAE WERE PROVIDED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: IMAGE ANALYSIS THE REPORTED ACCESS VESSEL DAMAGE COULD NOT BE CONFIRMED ON THE PRE-IMPLANT IMAGES PROVIDED; THEREFORE, THE CAUSE OF THE EVENT COULD NOT BE DETERMINED. PROCEDURAL ANGIOGRAMS SHOWING THE DELIVERY SYSTEM BEING INTRODUCED AND REMOVED FROM THE PATIENT WERE NOT AVAILABLE FOR ASSESSMENT OF THE EVENT. THE REPORTED VESSEL DIAMETER OF 7.2-8 MM COULD BE APPRECIATED IN THE IMAGES. THE REPORTED BLOOD LOSS, CARDIAC ARREST, AND PATIENT DEATH COULD NOT BE ASSESSED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21573983&pc=DSQ | 2916596-2025-01499 | 2025-02-23 05:00 | Death | THORATEC CORPORATION | 2025-03-11 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE PATIENT OUTCOME THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND THE REPORTED PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. THE OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. NO AUTOPSY WAS PERFORMED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT'S OUTCOME COULD NOT CONCLUSIVELY BE ESTABLISHED THROUGH THIS EVALUATION. HEARTMATE 3 LVAS INSTRUCTIONS FOR USE LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REVISION D, IS CURRENTLY AVAILABLE. SECTION 1 OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21574011&pc=DSQ | 2916596-2025-01535 | 2025-02-03 05:00 | Death | THORATEC CORPORATION | 2025-03-11 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21574059&pc=MIH | 9612164-2025-01277 | 2025-03-06 05:00 | Death | MEDTRONIC IRELAND | 2025-03-11 04:00 | MIH | VALIANT CAPTIVIA - FF | Positioning Problem | Cardiac Arrest; Hemorrhage/Bleeding; Perforation of Vessels | P100040 |  | 1 | Event Description: TWO VALIANT CAPTIVIA STENT GRAFTS WERE IMPLANTED DURING THE ENDOVASCULAR TREATMENT OF A 48MM TAA. IT WAS REPORTED THAT DURING THE INDEX PROCEDURE VAMF3834C150TJ WAS ATTEMPTED TO BE ADVANCED ONCE ACCESS WAS OBTAINED IN THE RIGHT FEMORAL ARTERY AND A CLOSURE DEVICE USED, BUT IT DID NOT ADVANCE NEAR THE EXTERNAL ILIAC ARTERY TO COMMON ILIAC ARTERY. IT WAS REMOVED AND A NON MDT SHEATH WAS ADVANCED FROM THE SAME RIGHT FEMORAL ARTERY. THE SHEATH WAS ADVANCED TO THE RENAL ARTERY LEVEL. THE VAMF3834C150TJ WAS INSERTED INTO THE SHEATH AND PLACED IN DESCENDING POSITION AS SCHEDULED. THE NEXT STENT GRAFT VAMF4040C150TJ WAS LANDED AT THE DISTAL ARCH, AND CONTRAST IMAGING WAS PERFORMED. DURING THE CONTRAST IMAGING, THE CONTRAST AGENT FLOW WAS POOR AND HAD RETAINED IN THE ARCH FOR A LONG TIME. VAMF4040C150TJ WAS PLACED AS SCHEDULED AND CONTRAST IMAGING WAS PERFORMED. THE TEVAR ITSELF WAS COMPLETED. THE ACCESS WAS CHECKED AND A CONTRAST AGENT LEAKED OUT OF THE BLOOD VESSEL AROUND THE RIGHT EIA-CIA, CONFIRMING DAMAGE TO THE ACCESS VESSEL. THE PATIENT SUDDENLY DEVELOPED SHOCK AND CARDIAC ARREST. CARDIAC MASSAGE WAS INITIATED. IT WAS DECIDED TO PLACE A NON-MDT IMMEDIATELY. THE FIRST NON MDT STENT 10MM X 100MM WAS PLACED IN CIA-EIA. AS THE DAMAGE REMAINED IN THE CONTRAST IMAGING AFTER PLACEMENT ANOTHER NON MDT STENT 9 MM X 100MM WAS ADDED PROXIMALLY SLIGHTLY FROM THE FEMORAL HEAD LEVEL. IT WAS NOTED THAT THE NATIVE BLOOD VESSEL DIAMETER IS ABOUT 8MM. THERE WAS A TYPE IIIA ENDOLEAK BETWEEN THE NON MDT STENTS, SO TOUCH UP WITH A NON-MDT BALLOON. THE IIIA ENDOLEAK DID NOT RESOLVE. WHEN THE 8FR SHEATH WAS ADVANCED TO PLACE A 3RD NON MDT STENT 11 MM STENT, THE JUNCTION OF THE ALREADY IMPLANTED NON MDT STENTS WAS DISCONNECTED. THE NON MDT, WHICH HAD BEEN PLACED PROXIMALLY, HAD BEEN COMPLETELY PUSHED OUT OF THE BLOOD VESSEL, AND THE GUIDEWIRE, WHICH HAD BEEN PASSING THROUGH THE TRUE LUMEN, HAD ALSO ESCAPED, MAKING CANNULATION IMPOSSIBLE. A RESCUE BALLOON WAS ADVANCED FROM THE OPPOSITE SIDE (LEFT) TO OCCLUDE. THE FOLLOWING PROCEDURES WERE BEING CONSIDERED, BUT THE CARDIAC MASSAGE COULD NOT BE STOPPED WHILE THE PATIENT REMAINED IN CARDIAC ARREST. BLEEDING FROM THE INTUBATION TUBE WAS ALSO CONFIRMED. TWO (2) HOURS HAD PASSED SINCE THE CARDIAC ARREST, AND THE PATIENT'S HEART RATE DID NOT RESUME EVEN ONCE; THE PHYSICIAN DECIDED TO STOP RESUSCITATION AND CONFIRMED THE PATIENT'S DEATH. PER THE PHYSICIAN THE CAUSE OF THE ACCESS DAMAGE WAS ANATOMY RELATED DUE TO THE PATIENT'S VASCULAR MORPHOLOGY AND THE ACCESS VESSEL WAS NARROW AT ABOUT 7.2MM -8MM. NO ADDITIONAL CLINICAL SEQUELAE WERE PROVIDED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: IMAGE ANALYSIS THE REPORTED POSITIONING DIFFICULTIES AND ACCESS VESSEL DAMAGE COULD NOT BE CONFIRMED ON THE PRE-IMPLANT IMAGES PROVIDED; THEREFORE, THE CAUSE OF THE EVENT COULD NOT BE DETERMINED. PROCEDURAL ANGIOGRAMS SHOWING THE DELIVERY SYSTEM BEING INTRODUCED AND REMOVED FROM THE PATIENT WERE NOT AVAILABLE FOR ASSESSMENT OF THE EVENT. THE REPORTED VESSEL DIAMETER OF 7.2-8 MM COULD BE APPRECIATED IN THE IMAGES. THE REPORTED BLOOD LOSS, CARDIAC ARREST, AND PATIENT DEATH COULD NOT BE ASSESSED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21574065&pc=DSQ | 2916596-2025-01513 | 2025-02-19 05:00 | Death | THORATEC CORPORATION | 2025-03-11 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND THE REPORTED PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. THE OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REV. C IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21574275&pc=MIH | 2017233-2025-05916 | 2025-03-04 05:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-03-11 04:00 | MIH | GORE® TAG® THORACIC BRANCH ENDOPROSTHESIS | Patient-Device Incompatibility | Ruptured Aneurysm | P210032 |  | 1 | Event Description: IT WAS REPORTED TO GORE THAT THE PATIENT UNDERWENT 2 PROCEDURES. THE PRIMARY ANEURYSM WAS ~89MM IN SIZE AND HAD ALREADY RUPTURED ONCE AT 2021, AT AN UNKNOWN DATE. THE PATIENT HAD 2 STENT GRAFTS (COOK® MEDICAL) IMPLANTED IN 2021. AND LATER, A GORE® TAG® THORACIC BRANCH ENDOPROSTHESIS [TBE], USING TSB081506E & TAC084515E, WAS IMPLANTED ON (B)(6) 2025, IN A PATIENT WITH A HISTORY OF HYPERTENSION. THIS PROCEDURE WAS COMPLETED SUCCESSFULLY. HOWEVER, IT HAS BEEN IMPLANTED DESPITE A VERY DIFFICULT ANATOMY. REPORTEDLY, THE PATIENT DIED 5 DAYS AFTER TBE IMPLANTATION BECAUSE OF AN RUPTURED THORACO-ABDOMINAL ANEURYSM. IT WAS ADDITIONALLY REPORTED: PATIENT WAS TREATED WITHIN THE IFU FOR THE TBE PROCEDURE, NO ENDOLEAK WAS PRESENT DURING OR POST-PROCEDURE OF THE TBE IMPLANT DATE, THE TBE WAS USED FOR THE SUPPLY OF THE THORACIC SECTION, THE ANEURYSM WAS IN THIS THORACO-ABDOMINAL AREA, THERE WAS NO ISSUE DURING THE IMPLANT PROCEDURE FOR THE SIDE BRANCH (TSB081506E) OBSERVED, OR PATENCY ISSUE OBSERVED AFTER IMPLANT, AND THE RESULT WAS VERY GOOD.Event Description: IT WAS REPORTED TO GORE THAT THE PATIENT UNDERWENT 2 PROCEDURES. THE PRIMARY ANEURYSM WAS ~89MM IN SIZE AND HAD ALREADY RUPTURED ONCE AT 2021, AT AN UNKNOWN DATE. THE PATIENT HAD 2 STENT GRAFTS (COOK® MEDICAL) IMPLANTED IN 2021. AND LATER, A GORE® TAG® THORACIC BRANCH ENDOPROSTHESIS [TBE], USING TSB081506E (THORACIC SIDE BRANCH COMPONENT) & TAC084515E (THORACIC STENT GRAFT), WAS IMPLANTED ON (B)(6) 2025, IN A PATIENT WITH A HISTORY OF HYPERTENSION. THIS PROCEDURE WAS COMPLETED SUCCESSFULLY. HOWEVER, IT HAS BEEN IMPLANTED DESPITE A VERY DIFFICULT ANATOMY. REPORTEDLY, THE PATIENT DIED 5 DAYS AFTER TBE IMPLANTATION BECAUSE OF AN RUPTURED THORACO-ABDOMINAL ANEURYSM. IT WAS ADDITIONALLY REPORTED: - PATIENT WAS TREATED WITHIN THE IFU FOR THE TBE PROCEDURE, - NO ENDOLEAK WAS PRESENT DURING OR POST-PROCEDURE OF THE TBE IMPLANT DATE, - THE TBE WAS USED FOR THE SUPPLY OF THE THORACIC SECTION, THE ANEURYSM WAS IN THE THORACO-ABDOMINAL AREA, - THERE WAS NO ISSUE DURING THE IMPLANT PROCEDURE FOR THE SIDE BRANCH (TSB081506E) OBSERVED. Manufacturer Narrative: SECTION C: CBAS® HEPARIN SURFACE INCORPORATES CARMEDA HEPARIN MANUFACTURED FROM HEPARIN SODIUM API, WHICH IS COVALENTLY BOUND TO THE DEVICE SURFACE AND IS ESSENTIALLY NON-ELUTING. D10: CONCOMITANT MEDICAL PRODUCTS: (RELEVANT ASSOCIATED DEVICES USED WITH THE DEVICE BEING REPORTED ON): GORE® TAG® THORACIC BRANCH ENDOPROSTHESIS (AORTIC COMPONENT). H3: PRELIMINARY RESULTS ARE NOT YET AVAILABLE. H6, B20: ENGINEERING EVALUATION COULD NOT BE PERFORMED AS THE DEVICE REMAINS IMPLANTED. H6, B15: IMAGING TASK STARTED FOR CLINICAL IMAGES. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. Manufacturer Narrative: UPDATED G3. H6: INVESTIGATION FINDINGS: CODE C19 WAS ADDED. THE C21 IS NO LONGER APPLICABLE. CODE D15 ADDED IN INVESTIGATION CONCLUSIONS. UPDATED B5, THE INCIDENT DESCRIPTION. GORE HAS FINISHED THE INVESTIGATION AND RESULTS STATED AS THE FOLLOWING: A PHR REVIEW FOR THIS DEVICE (SERIAL (B)(6)) WAS COMPLETED, THE REVIEW OF THE MANUFACTURING RECORDS FOR THIS DEVICE VERIFIED THAT THE LOT MET ALL PRE-RELEASE SPECIFICATIONS. ENGINEERING EVALUATION COULD NOT BE PERFORMED AS THE DEVICE REMAINS IMPLANTED. FURTHER, AN IMAGING TASK OF A NON-CLINICAL RADIOGRAPHIC FILE WAS EVALUATED, BUT NO DICOM IMAGES WERE PROVIDED: THE IMAGES RECEIVED CANNOT BE USED TO PERFORM A FULL IMAGING EVALUATION BECAUSE THEY DO NOT MEET THE DICOM STANDARD. THE EXTENT AND ACCURACY OF THE OBSERVATIONS AND FINDINGS MAY BE LIMITED DUE TO THE COMPLETENESS, FORMAT AND/OR QUALITY OF THE IMAGES PROVIDED FOR REVIEW. GORE CANNOT GUARANTEE THE IMAGES PROVIDED ARE COMPLETE, ACCURATE OR LACK ALTERATION. THE IMAGING SUMMARY WAS STUDIED BY MPDIMAGING TEAM, IT IS NOT CONCLUSIVE TO FIND A ROOT CAUSE RELATED TO THE DEVICES. NO RUPTURED AREA WAS SEEN IN THIS IMAGES EVALUATION. WITH THE INFORMATION GIVEN ABOVE, THE REPORTED ANEURYSM RUPTURE IS NOT ABLE TO CONCLUDE DUE TO MISSING RELEVANT INFORMATION AND NO DEVICE IS RETURNED. THE INITIAL REPORTER HAS BEEN CONTACTED IN ORDER TO RECEIVE MORE INFORMATION ON THE EVENT AND PATIENT DETAILS. PER THE GORE® TAG® THORACIC BRANCH ENDOPROSTHESIS INSTRUCTIONS FOR USE (IFU), POSSIBLE ADVERSE EVENTS AND COMPLICATIONS THAT MAY OCCUR WITH THE USE OF GORE® TAG® THORACIC BRANCH ENDOPROSTHESIS INCLUDE, BUT ARE NOT LIMITED TO: AORTIC RUPTURE, DEATH, DISSECTION, OR RUPTURE OF THE AORTIC VESSEL AND/OR SURROUNDING VASCULATURE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21574350&pc=FPD | 2245270-2025-00019 | 2024-12-09 05:00 | Death | VYGON | 2025-03-11 04:00 | FPD | FEEDING TUBE | Material Perforation | Perforation | K060944 |  | 1 | Event Description: GASTRIC PERFORATION WHEN USED FOR GASTRIC FEEDING. THE TUBE WAS PLACED IN A 27 WEEK PREMATURE BABY ON (B)(6) 2024, AND ON (B)(6) THEY DIAGNOSED NEONATAL GASTRIC PERFORATION (NGP). THE BABY WAS TRANSFERRED TO (B)(6) HOSPITAL, A SPECIALIST PAEDIATRIC HOSPITAL WHERE HE DIED. THE BABY WAS A TWIN, AND THE OTHER TWIN ALSO EXPERIENCED A NGP BUT HE SURVIVED. FOR ADMINISTRATION OF ENTERAL NUTRITION. Manufacturer Narrative: THE INVOLVED DEVICE IS NOT AVAILABLE. THE CUSTOMER DID NOT PROVIDE THE INVOLVED BATCH NUMBER. THEREFORE, VYGON IE HAS PROVIDED THE BATCHES SUPPLIED TO THIS CUSTOMER IN PREVIOUS 4 MONTHS: BATCH NUMBER 200324PA, 210224PA, 041223PA. THE BATCH REVIEW OF THESE 3 BATCHES CONFIRMS THAT THE FEEDING TUBES ARE COMPLIANT, NO ANOMALIES AND NO NON-CONFORMITIES WERE RECORDED. THESE FEEDING TUBE HAVE ALL A ROUNDED DISTAL END TO AVOID ANY TRAUMA TO PATIENT. THE MANUFACTURING PROCESS OF THESE NUTRISAFE FEEDING TUBES INCLUDES A 100% LEAK AND OBSTRUCTION TESTS. THERE ARE ALSO QUALITY CONTROL CHECKS ARE AS VISUAL ASPECT, DISTAL END ROUNDING CONTROL, DISTANCE BETWEEN THE FIRST MARKING AND THE DISTAL END CENTIMER'S MARKING, USEFUL LENGTH. IN THE INSTRUCTION OF USE THE INFORMATION ARE DESCRIBED AS FOLLOWS: PLACEMENT OF THE FEEDING TUBE: FOR ORAL OR NASAL PLACEMENT, MEASURE THE LENGTH OF TUBE TO BE INSERTED ACCORDING TO STANDARD PROTOCOLS. LUBRICATE THE FEEDING TUBE. INTRODUCE THE FEEDING TUBE IN ONE OF THE NOSTRILS AND MOVE THE TUBE CAREFULLY TOWARDS THE POSTERIOR NASOPHARYNX, ORIENTED TOWARDS THE OESOPHAGUS. PROCEED SLOWLY TO AVOID THE RISK OF THE PATIENT VOMITING. BEND THE HEAD FORWARD TO CLOSE THE TRACHEA AND OPEN THE ESOPHAGUS. ADVANCE THE TUBE INTO THE OESOPHAGUS. ONCE THE MARK IS REACHED, STOP THE ADVANCEMENT OF THE TUBE. FIXATION OF THE FEEDING TUBE: AFTER PREPARING THE SKIN, DRESS ADHESIVE STRIP AROUND THE TUBE AND FIX IT ON THE NOSE. VERIFICATION: CHECK FEEDING TUBE POSITIONING PREFERABLY BY X-RAY. IF USING A FEEDING TUBE WITH STYLET, REMOVE GENTLY THE STYLET FROM THE FEEDING TUBE. -PRECAUTION FOR USE: THE PLACEMENT OF THE TUBE IS VERIFIED ONCE A DAY AND SYSTEMATICALLY BEFORE EACH USE (IN ANY CASE FOLLOW THE PROTOCOLS OF YOUR INSTITUTION). NUTRISAFE2 AND NUTRIFIT FEEDING TUBES ARE CONTRAINDICATED FOR PATIENTS WITH CONGENITAL ABNORMALITIES OF THE GASTROINTESTINAL TRACT. THE INFORMATION WE RECEIVED ARE POOR. WE DO NOT KNOW IF THE PLACEMENT HAS BEEN CHECKED AFTER INSERTION, DAILY AND THEN BEFORE EACH USE, IF THE PATIENT HAD NO ABNORMAL ANATOMY. FURTHERMORE, CASES OF NEONATAL GASTRIC PERFORATION ARE VERY RARE, BUT THEY ARE NEVERTHELESS WELL REPORTED IN THE LITERATURE. IN THESE CASES, THE STOMACH WALLS ARE NOT MATURE AND ARE OFTEN VERY THIN AND WEAK AND NOT RESISTANT TO EVEN MINIMAL PRESSURE. PREMATURITY AND LOW BIRTH WEIGHT ARE THE MAIN FACTORS ASSOCIATED WITH NGP WITH HYPOTENSION, SEPSIS, HIGH GASTRIC ACIDITY AS SIGNIFICANT CONTRIBUTORY FACTORS. MECHANICAL VENTILATION, SPECIFICALLY NASAL VENTILATION ARE ALSO SUSPECTED OF BEING CONTRIBUTORY FACTORS. THE PRESENCE OF A NG TUBE IS OBVIOUSLY ANOTHER RISK FACTOR BUT THERE ARE EVEN CASES OF SPONTANEOUS NGP REPORTED WHEN THERE IS NO NG TUBE IN SITU. THE MOST LIKELY CAUSE IS A RARE COMPLICATION ON A PREMATURE BABY. IT SEEMS NOT TRACED TO THE DEVICE. FROM HISTORICAL DATA ANALYSIS OF THIS NUTRISAFE2 ENTERAL TUBES PVC RANGE CODES 36X.XXX OF SIMILAR INCIDENT (GASTRIC PERFORATION) THESE 3 LAST YEARS, WE CAN CONCLUDE THAT THIS IS THE VERY FIRST CASE OF PERFORATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21574520&pc=FRN | 2016493-2025-65206 | 2025-02-07 05:00 | Death | CAREFUSION SD | 2025-03-11 04:00 | FRN | ALARIS SYSTEM | Contamination; Pumping Stopped; Use of Device Problem | Insufficient Information | K211218 |  | 1 | Event Description: IT WAS INITIALLY REPORTED A REQUEST TO REVIEW THE DEVICE LOGS FOR 07 FEBRUARY 2025 BETWEEN THE HOURS OF 1320 AND 1640. NO FURTHER INFORMATION WAS PROVIDED AT THE TIME. BD RECEIVED ADDITIONAL INFORMATION THROUGH DUE DILIGENCE ATTEMPTS. IT WAS REPORTED THAT THE PATIENT WAS TRANSPORTED FROM EMERGENCY DEPARTMENT (ED) TO RADIOLOGY FOR APPROXIMATELY 15 MINUTES. LEVOPHED AND BICARB WAS INFUSING VIA PUMP AT THE TIME. UPON RETURN TO THE ED, THE PUMP WAS FOUND "NOT INFUSING." PER REPORT, THE PATIENT DIED AND HAD A DNR (DO NOT RESUSCITATE) CODE STATUS "WITH FAMILY DECISION OF NO FURTHER CARE." WHEN ASKED IF THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH, THE HOSPITAL'S PATIENT SAFETY MANAGER RN STATED THAT "CLARIFYING THE CONCERN IS IF THE PUMP SHUT DOWN DURING THE INFUSION POTENTIALLY DUE TO BATTERY LIFE AS IT WAS UNPLUGGED, WAS THE PUMP CONTINUOUSLY INFUSING AND THE DISPLAY WAS NOT VISIBLE. REALLY JUST SEEKING IF THE PUMP WAS CONTINUOUSLY INFUSING OR NOT AND IF NOT, WHY? THE PUMP NOR THE INFUSION(S) LED NOR CONTRIBUTED TO THE DEMISE OF THE PATIENT. WE ARE REVIEWING FOR PROCESS FAILURES INCLUDING TEAM MEMBER HANDOFF IN COMMUNICATION." THE CUSTOMER IS REQUESTING THE MANUFACTURER TO REVIEW THE LOGS, SPECIFICALLY IDENTIFYING "THE MEDICATION INFUSION DETAILS OF ML/HR. AND IF THERE WAS ANY TIME, THE MEDICATIONS WERE NOT INFUSING AS WE ARE REVIEWING FOR OUR INVESTIGATION OF PROCESSES. THE PATIENT WAS A DNR, WITH FAMILY PRESENTING THE DNR DOCUMENT UPON THEIR ARRIVAL TO THE ED."Event Description: IT WAS INITIALLY REPORTED A REQUEST TO REVIEW THE DEVICE LOGS FOR (B)(6) 2025 BETWEEN THE HOURS OF 1320 AND 1640. NO FURTHER INFORMATION WAS PROVIDED AT THE TIME. BD RECEIVED ADDITIONAL INFORMATION THROUGH DUE DILIGENCE ATTEMPTS. IT WAS REPORTED THAT THE PATIENT WAS TRANSPORTED FROM EMERGENCY DEPARTMENT (ED) TO RADIOLOGY FOR APPROXIMATELY 15 MINUTES. LEVOPHED AND BICARB WAS INFUSING VIA PUMP AT THE TIME. UPON RETURN TO THE ED, THE PUMP WAS FOUND "NOT INFUSING." PER REPORT, THE PATIENT DIED AND HAD A DNR (DO NOT RESUSCITATE) CODE STATUS "WITH FAMILY DECISION OF NO FURTHER CARE." WHEN ASKED IF THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH, THE HOSPITAL'S PATIENT SAFETY MANAGER RN STATED THAT "CLARIFYING THE CONCERN IS IF THE PUMP SHUT DOWN DURING THE INFUSION POTENTIALLY DUE TO BATTERY LIFE AS IT WAS UNPLUGGED, WAS THE PUMP CONTINUOUSLY INFUSING AND THE DISPLAY WAS NOT VISIBLE. REALLY JUST SEEKING IF THE PUMP WAS CONTINUOUSLY INFUSING OR NOT AND IF NOT, WHY? THE PUMP NOR THE INFUSION(S) LED NOR CONTRIBUTED TO THE DEMISE OF THE PATIENT. WE ARE REVIEWING FOR PROCESS FAILURES INCLUDING TEAM MEMBER HANDOFF IN COMMUNICATION." THE CUSTOMER IS REQUESTING THE MANUFACTURER TO REVIEW THE LOGS, SPECIFICALLY IDENTIFYING "THE MEDICATION INFUSION DETAILS OF ML/HR. AND IF THERE WAS ANY TIME, THE MEDICATIONS WERE NOT INFUSING AS WE ARE REVIEWING FOR OUR INVESTIGATION OF PROCESSES. THE PATIENT WAS A DNR, WITH FAMILY PRESENTING THE DNR DOCUMENT UPON THEIR ARRIVAL TO THE ED. " Manufacturer Narrative: OMIT : B17 - DEVICE NOT RETURNED, C20 - NO FINDINGS AVAILABLE. ADDITIONAL INFORMATION: DEVICE AVAILABLE FOR EVAL?, RETURNED TO MANUFACTURER ON, DEVICE EVAL BY MANUFACTURER?, IMDRF ANNEX A, B, C, D, G CODES AND MANUFACTURER NARRATIVE. A DEVICE HISTORY RECORD, COMPLAINT HISTORY REVIEW, AND RISK REVIEW ON FAILURE MODES ARE PERFORMED ON EACH DEVICE WITH REPORTABLE MALFUNCTION(S) ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN SECTION H6 OF THIS MDR REPORT. INVESTIGATON SUMMARY: THE CUSTOMER SOUGHT INFORMATION AROUND THE MEDICATION INFUSION DETAILS AND WHETHER THERE WAS ANY TIME WHERE THE MEDICATIONS WERE NOT INFUSING. REVIEW OF THE DEVICE LOGS IDENTIFIED INSTANCES OF OCCLUDED PATIENT SIDE ALARMS, BUT THERE WERE NO INDICATIONS OF DEVICE MALFUNCTIONS OR INTERRUPTIONS THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE PATIENT¿S CONDITION. INTERRUPTIONS IN INFUSION WERE NOT OBSERVED OR DURING DEVICE TESTING. INSPECTION AND TESTING OF THE SUSPECT PUMP MODULES AND PCU DID NOT FIND ANY EXISTING MALFUNCTIONS. THE SUSPECT PUMP MODULES WERE INFUSING WITHIN SPECIFICATIONS. THE PCU PASSED PREVENTATIVE MAINTENANCE TESTING AND WAS OPERATING AS INTENDED. REVIEW OF THE SUSPECT PCU BATTERY LOGS DID NOT SHOW ANY EVENTS OR MALFUNCTIONS THAT COULD HAVE CONTRIBUTED TO THE REPORTED INCIDENT. THERE WERE NO INSTANCES OF THE SYSTEM POWERING OFF DURING THE REPORTED INCIDENT TIME. THE INITIAL INFUSION PROGRAMMING OF NOREPINEPHRINE (DRUGID = 308), AT A CONCENTRATION OF 4 MG / 250 ML, WAS RECORDED ON (B)(6) 2025 AT 1:13 PM ON THE SUSPECT PUMP MODULE SN (B)(6). THE INFUSION RATE WAS SET AT 41.25 ML, THE VOLUME TO BE INFUSED (VTBI) WAS 100 ML. A BASIC INFUSION WAS PROGRAMMED ON (B)(6) 2025 AT 3:13 PM ON THE SUSPECT PUMP MODULE SN (B)(6) AT A RATE OF 100 ML/HR AND A VTBI OF 950 ML. A THIRD PUMP MODULE (SN (B)(6), NOT RECEIVED) WAS IDENTIFIED TO BE IN USE DURING THE REPORTED INCIDENT TIME. A BASIC INFUSION WAS PROGRAMMED AND STARTED ON (B)(6) 2025 AT 1:47 PM, WITH A RATE OF 999 ML/HR AND A VTBI OF 950 ML. AT 2:40 PM, THE CONCOMITANT PUMP MODULE SN (B)(6) WAS CHANNELED OFF. THEN AT 2:59 PM AN INFUSION FOR THE DRUG CEFTRIAXONE (DRUGID = 131), AT A CONCENTRATION OF 1000 MG / 50 ML, WAS PROGRAMMED AND STARTED. THE INFUSION RATE WAS 100 ML/HR AND THE VTBI WAS 50 ML. THE DOSE FOR THE INFUSION OF NOREPINEPHRINE ON THE PUMP MODULE SN (B)(6) WAS CHANGED TO 0.4 MCG/KG/HR AT 3:18 PM AND THEN CHANGED AGAIN TO 0.22 MCG/KG/HR AT 3:20 PM. THE DOSE WAS CHANGED TO 0.25 MCG/KG/HR AT 4:28 PM. AN AIR-IN-LINE ALARM WAS OBSERVED ON THE CONCOMITANT PUMP MODULE SN (B)(6), LASTING 2 MINUTES BEFORE THE UNIT WAS CHANNELED OFF. THE TOTAL VOLUME INFUSED (TVI) WAS 902.015 ML. FIVE (5) OCCLUDED PATIENT SIDE ALARMS WERE OBSERVED BETWEEN 1:13 PM AND 4:27 PM ON THE PUMP MODULE SN (B)(6). THE ALARM AT 2:00 PM LASTED 8 MINUTES BEFORE THE INFUSION WAS RESTARTED. THEN, AT 3:15 PM, THE ALARM LASTED 2 MINUTES AND 48 SECONDS BEFORE THE INFUSION WAS RESUMED. THE SUSPECT PUMP MODULE SN (B)(6) WAS CHANNELED OFF AT 3:46 PM, BUT WAS THEN POWERED ON AGAIN AT 4:24 PM AND AN INFUSION OF NOREPINEPHRINE AT A DOSE OF 0.22 MCG/KG/HR AND A VTBI OF 200 ML WAS PROGRAMMED AND STARTED AT 4:26 PM. THE SUSPECT PUMP MODULE SN (B)(6) WAS CHANNELED OFF AT 4:29 PM. THE TVI WAS RECORDED AT 101.491 ML. THE SUSPECT PUMP MODULE SN (B)(6) WAS CHANNELED OFF AT 4:30 PM. THE TVI WAS RECORDED AT 127.003 ML. IT IS NOT POSSIBLE TO DETERMINE WHAT THE ACTUAL VOLUME IS WITHIN AN IV CONTAINER AT THE START AND ENDING OF AN INFUSION FROM A REVIEW OF THE PCU EVENT LOG. IT IS ALSO NOT POSSIBLE TO DETERMINE WHAT THE FLUID IS THAT IS CONTAINED WITHIN THE IV CONTAINER. THIS IS NOT A FUNCTION OF THE PCU EVENT LOG; IT ONLY CAN REFLECT THE INPUTTED INFORMATION IT RECEIVES EITHER MANUALLY OR REMOTELY WITH RESPECT TO VOLUME TO BE DELIVERED AND FLUID SELECTED. ROOT CAUSE: THE ROOT CAUSE OF THE REPORTED INTERRUPTION OF INFUSIONS IS DUE TO THE USER POWERING OFF THE PUMP MODULES. IT WAS NOTED THAT THE SUSPECT PUMP MODULES SN (B)(6) WERE POWERED OFF ON (B)(6) 2025 AT 4:29 PM AND 4:30 PM, RESPECTIVELY. THE CONCOMITANT PUMP MODULE SN (B)(6) WAS POWERED OFF AT 3:28 PM. NO FUNCTIONALITY ISSUES OR ANOMALIES WERE OBSERVED WITH THE DEVICES THAT WOULD CONTRIBUTE TO THE REPORTED EVENT. THE RETURNED PUMP MODULES WERE FOUND TO BE INFUSING WITHIN SPECIFICATION. NOTE THAT THIS REPORT LISTS IMDRF ANNEX A1801, G04067, C0601 CODES NOT ASSOCIATED WITH THE REPORTED EVENT BUT IDENTIFIED AS REPORTABLE MALFUNCTIONS OBSERVED ON THE DEVICE DURING INVESTIGATION ARE UNRELATED TO THE REPORTED ISSUE. THESE OTHER FAILURES PRESUMPTIVELY DID NOT CAUSE OR CONTRIBUTE TO THE REPORTED ISSUE. Manufacturer Narrative: THE ACTUAL DATE OF DEATH IS UNKNOWN. BD TECHNICAL SUPPORT TROUBLESHOOT WITH CUSTOMER OVER THE PHONE. THE EVENT LOGS HAVE BEEN RECEIVED AND THE EVALUATION IS PENDING. A FOLLOW UP REPORT WILL BE SUBMITTED WITH INVESTIGATION RESULTS ONCE THE EVALUATION HAS BEEN COMPLETED. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21574589&pc=DSQ | 2916596-2025-01436 | 2025-02-12 05:00 | Death | THORATEC CORPORATION | 2025-03-11 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Sepsis; Fungal Infection; Multiple Organ Failure; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO CANDIDA SEPSIS. THE PATIENT SUFFERED FROM A SEVERE CANDIDA SEPSIS FOLLOWED BY MULTIPLE ORGAN FAILURE (MOF) WHICH WAS RELATED TO THE OUTCOME. NO DEVICE AND/OR DRIVELINE INFECTION WAS IN PLACE AT THE TIME OF THE OUTCOME. THE ROOT CAUSE FOR INFECTION WAS NOT IDENTIFIED BUT PULMONARY INFECTION WAS SUSPECTED. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. Manufacturer Narrative: H6-3261 - MULTIPLE ORGAN DYSFUNCTION SYNDROME. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REV. C AND THE HEARTMATE 3 PATIENT HANDBOOK, REV. B ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, INFECTION, SEPSIS, AND MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RESPIRATORY FAILURE, RIGHT HEART FAILURE, RENAL FAILURE, AND HEPATIC DYSFUNCTION), THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, LISTS INFECTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. SEVERAL SECTIONS OF THE IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION, AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21574702&pc=NPT | 2025587-2025-01819 | 2025-01-22 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-11 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Non specific EKG/ECG Changes | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 3 MONTHS AND 3 WEEKS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS ARRHYTHMIA. PER THE PHYSICIAN, NEITHER THE DEVICE NOR THE PROCEDURE CONTRIBUTED TO THE CAUSE OF DEATH. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21574934&pc=NIQ | 2124215-2025-11704 | 2024-05-01 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-03-11 04:00 | NIQ | SYNERGY | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. IN (B)(6) 2023, THE SUBJECT PRESENTED WITH STABLE ANGINA AND WAS REFERRED FOR CARDIAC CATHETERIZATION. THE FIRST TARGET LESION WAS LOCATED IN THE DISTAL LEFT ANTERIOR DESCENDING ARTERY (LAD) WITH 90% STENOSIS AND WAS 21 MM LONG, WITH A REFERENCE VESSEL DIAMETER OF 2.25 MM. THE FIRST TARGET LESION WAS TREATED WITH PRE-DILATATION AND FOLLOWED BY THE PLACEMENT OF A 2.25 MM X 24 SYNERGY STENT SYSTEM. FOLLOWING POST-DILATATION, THE RESIDUAL STENOSIS WAS NOTED TO BE 0%. THE SECOND TARGET LESION WAS LOCATED IN THE FIRST DIAGONAL WITH 90% STENOSIS AND WAS 21 MM LONG, WITH A REFERENCE VESSEL DIAMETER OF 2.25 MM. THE SECOND TARGET LESION WAS TREATED WITH PRE-DILATATION AND FOLLOWED BY THE PLACEMENT OF A 2.25 MM X 24 SYNERGY STENT SYSTEM. FOLLOWING POST-DILATATION, THE RESIDUAL STENOSIS WAS NOTED TO BE 0%. NINE DAYS LATER, STAGED PROCEDURE WAS PERFORMED. THE FIRST TARGET LESION WAS LOCATED IN THE RIGHT POSTERIOR DESCENDING ARTERY (R-PDA) WITH 90% STENOSIS AND WAS 13 MM LONG, WITH A REFERENCE VESSEL DIAMETER OF 2.25 MM. THE FIRST TARGET LESION WAS TREATED WITH PRE-DILATATION AND FOLLOWED BY THE PLACEMENT OF A 2.25 MM X 16 SYNERGY STENT SYSTEM. THE RESIDUAL STENOSIS WAS NOTED TO BE 0%. POST-DILATATION WAS NOT PERFORMED. THE SECOND TARGET LESION WAS LOCATED IN THE DISTAL RIGHT CORONARY ARTERY (RCA) WITH 80% STENOSIS AND WAS 21 MM LONG, WITH A REFERENCE VESSEL DIAMETER OF 3.5 MM. THE SECOND TARGET LESION WAS TREATED WITH PRE-DILATATION AND FOLLOWED BY THE PLACEMENT OF A 3.50 MM X 24 SYNERGY STENT SYSTEM. FOLLOWING POST-DILATATION, THE RESIDUAL STENOSIS WAS NOTED TO BE 0%. THREE DAYS LATER, THE SUBJECT WAS DISCHARGED ON ASPIRIN AND CLOPIDOGREL. IN (B)(6) 2024, THE SUBJECT DIED. AT THE TIME OF EVENT, THE SUBJECT WAS ON ASPIRIN AND CLOPIDOGREL. NO ACTION WAS TAKEN TO TREAT THE EVENT. IT IS UNKNOWN IF AUTOPSY WAS PERFORMED OR NOT. Manufacturer Narrative: B3: DATE OF EVENT: USED MAY 01, 2024, AS THE EXACT DEATH DATE WAS UNKNOWN. E1: INITIAL REPORTER ADDRESS 1: (B)(6).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21574937&pc=NIQ | 2124215-2025-11703 | 2024-05-01 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-03-11 04:00 | NIQ | SYNERGY | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. IN (B)(6) 2023, THE SUBJECT PRESENTED WITH STABLE ANGINA AND WAS REFERRED FOR CARDIAC CATHETERIZATION. THE FIRST TARGET LESION WAS LOCATED IN THE DISTAL LEFT ANTERIOR DESCENDING ARTERY (LAD) WITH 90% STENOSIS AND WAS 21 MM LONG, WITH A REFERENCE VESSEL DIAMETER OF 2.25 MM. THE FIRST TARGET LESION WAS TREATED WITH PRE-DILATATION AND FOLLOWED BY THE PLACEMENT OF A 2.25 MM X 24 SYNERGY STENT SYSTEM. FOLLOWING POST-DILATATION, THE RESIDUAL STENOSIS WAS NOTED TO BE 0%. THE SECOND TARGET LESION WAS LOCATED IN THE FIRST DIAGONAL WITH 90% STENOSIS AND WAS 21 MM LONG, WITH A REFERENCE VESSEL DIAMETER OF 2.25 MM. THE SECOND TARGET LESION WAS TREATED WITH PRE-DILATATION AND FOLLOWED BY THE PLACEMENT OF A 2.25 MM X 24 SYNERGY STENT SYSTEM. FOLLOWING POST-DILATATION, THE RESIDUAL STENOSIS WAS NOTED TO BE 0%. NINE DAYS LATER, STAGED PROCEDURE WAS PERFORMED. THE FIRST TARGET LESION WAS LOCATED IN THE RIGHT POSTERIOR DESCENDING ARTERY (R-PDA) WITH 90% STENOSIS AND WAS 13 MM LONG, WITH A REFERENCE VESSEL DIAMETER OF 2.25 MM. THE FIRST TARGET LESION WAS TREATED WITH PRE-DILATATION AND FOLLOWED BY THE PLACEMENT OF A 2.25 MM X 16 SYNERGY STENT SYSTEM. THE RESIDUAL STENOSIS WAS NOTED TO BE 0%. POST-DILATATION WAS NOT PERFORMED. THE SECOND TARGET LESION WAS LOCATED IN THE DISTAL RIGHT CORONARY ARTERY (RCA) WITH 80% STENOSIS AND WAS 21 MM LONG, WITH A REFERENCE VESSEL DIAMETER OF 3.5 MM. THE SECOND TARGET LESION WAS TREATED WITH PRE-DILATATION AND FOLLOWED BY THE PLACEMENT OF A 3.50 MM X 24 SYNERGY STENT SYSTEM. FOLLOWING POST-DILATATION, THE RESIDUAL STENOSIS WAS NOTED TO BE 0%. THREE DAYS LATER, THE SUBJECT WAS DISCHARGED ON ASPIRIN AND CLOPIDOGREL. IN (B)(6) 2024, THE SUBJECT DIED. AT THE TIME OF EVENT, THE SUBJECT WAS ON ASPIRIN AND CLOPIDOGREL. NO ACTION WAS TAKEN TO TREAT THE EVENT. IT IS UNKNOWN IF AUTOPSY WAS PERFORMED OR NOT. Manufacturer Narrative: B3: DATE OF EVENT: USED MAY 01, 2024, AS THE EXACT DEATH DATE WAS UNKNOWN. E1: INITIAL REPORTER ADDRESS 1: (B)(6).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21574938&pc=NIQ | 2124215-2025-11705 | 2024-05-01 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-03-11 04:00 | NIQ | SYNERGY | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. IN (B)(6) 2023, THE SUBJECT PRESENTED WITH STABLE ANGINA AND WAS REFERRED FOR CARDIAC CATHETERIZATION. THE FIRST TARGET LESION WAS LOCATED IN THE DISTAL LEFT ANTERIOR DESCENDING ARTERY (LAD) WITH 90% STENOSIS AND WAS 21 MM LONG, WITH A REFERENCE VESSEL DIAMETER OF 2.25 MM. THE FIRST TARGET LESION WAS TREATED WITH PRE-DILATATION AND FOLLOWED BY THE PLACEMENT OF A 2.25 MM X 24 SYNERGY STENT SYSTEM. FOLLOWING POST-DILATATION, THE RESIDUAL STENOSIS WAS NOTED TO BE 0%. THE SECOND TARGET LESION WAS LOCATED IN THE FIRST DIAGONAL WITH 90% STENOSIS AND WAS 21 MM LONG, WITH A REFERENCE VESSEL DIAMETER OF 2.25 MM. THE SECOND TARGET LESION WAS TREATED WITH PRE-DILATATION AND FOLLOWED BY THE PLACEMENT OF A 2.25 MM X 24 SYNERGY STENT SYSTEM. FOLLOWING POST-DILATATION, THE RESIDUAL STENOSIS WAS NOTED TO BE 0%. NINE DAYS LATER, STAGED PROCEDURE WAS PERFORMED. THE FIRST TARGET LESION WAS LOCATED IN THE RIGHT POSTERIOR DESCENDING ARTERY (R-PDA) WITH 90% STENOSIS AND WAS 13 MM LONG, WITH A REFERENCE VESSEL DIAMETER OF 2.25 MM. THE FIRST TARGET LESION WAS TREATED WITH PRE-DILATATION AND FOLLOWED BY THE PLACEMENT OF A 2.25 MM X 16 SYNERGY STENT SYSTEM. THE RESIDUAL STENOSIS WAS NOTED TO BE 0%. POST-DILATATION WAS NOT PERFORMED. THE SECOND TARGET LESION WAS LOCATED IN THE DISTAL RIGHT CORONARY ARTERY (RCA) WITH 80% STENOSIS AND WAS 21 MM LONG, WITH A REFERENCE VESSEL DIAMETER OF 3.5 MM. THE SECOND TARGET LESION WAS TREATED WITH PRE-DILATATION AND FOLLOWED BY THE PLACEMENT OF A 3.50 MM X 24 SYNERGY STENT SYSTEM. FOLLOWING POST-DILATATION, THE RESIDUAL STENOSIS WAS NOTED TO BE 0%. THREE DAYS LATER, THE SUBJECT WAS DISCHARGED ON ASPIRIN AND CLOPIDOGREL. IN (B)(6) 2024, THE SUBJECT DIED. AT THE TIME OF EVENT, THE SUBJECT WAS ON ASPIRIN AND CLOPIDOGREL. NO ACTION WAS TAKEN TO TREAT THE EVENT. IT IS UNKNOWN IF AUTOPSY WAS PERFORMED OR NOT. Manufacturer Narrative: B3: DATE OF EVENT: USED MAY 01, 2024, AS THE EXACT DEATH DATE WAS UNKNOWN. E1: INITIAL REPORTER ADDRESS 1: (B)(6).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21575209&pc=NIQ | 2124215-2025-12922 | 2024-05-01 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-03-11 04:00 | NIQ | SYNERGY | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. IN (B)(6) 2023, THE SUBJECT PRESENTED WITH STABLE ANGINA AND WAS REFERRED FOR CARDIAC CATHETERIZATION. THE FIRST TARGET LESION WAS LOCATED IN THE DISTAL LEFT ANTERIOR DESCENDING ARTERY (LAD) WITH 90% STENOSIS AND WAS 21 MM LONG, WITH A REFERENCE VESSEL DIAMETER OF 2.25 MM. THE FIRST TARGET LESION WAS TREATED WITH PRE-DILATATION AND FOLLOWED BY THE PLACEMENT OF A 2.25 MM X 24 SYNERGY STENT SYSTEM. FOLLOWING POST-DILATATION, THE RESIDUAL STENOSIS WAS NOTED TO BE 0%. THE SECOND TARGET LESION WAS LOCATED IN THE FIRST DIAGONAL WITH 90% STENOSIS AND WAS 21 MM LONG, WITH A REFERENCE VESSEL DIAMETER OF 2.25 MM. THE SECOND TARGET LESION WAS TREATED WITH PRE-DILATATION AND FOLLOWED BY THE PLACEMENT OF A 2.25 MM X 24 SYNERGY STENT SYSTEM. FOLLOWING POST-DILATATION, THE RESIDUAL STENOSIS WAS NOTED TO BE 0%. NINE DAYS LATER, STAGED PROCEDURE WAS PERFORMED. THE FIRST TARGET LESION WAS LOCATED IN THE RIGHT POSTERIOR DESCENDING ARTERY (R-PDA) WITH 90% STENOSIS AND WAS 13 MM LONG, WITH A REFERENCE VESSEL DIAMETER OF 2.25 MM. THE FIRST TARGET LESION WAS TREATED WITH PRE-DILATATION AND FOLLOWED BY THE PLACEMENT OF A 2.25 MM X 16 SYNERGY STENT SYSTEM. THE RESIDUAL STENOSIS WAS NOTED TO BE 0%. POST-DILATATION WAS NOT PERFORMED. THE SECOND TARGET LESION WAS LOCATED IN THE DISTAL RIGHT CORONARY ARTERY (RCA) WITH 80% STENOSIS AND WAS 21 MM LONG, WITH A REFERENCE VESSEL DIAMETER OF 3.5 MM. THE SECOND TARGET LESION WAS TREATED WITH PRE-DILATATION AND FOLLOWED BY THE PLACEMENT OF A 3.50 MM X 24 SYNERGY STENT SYSTEM. FOLLOWING POST-DILATATION, THE RESIDUAL STENOSIS WAS NOTED TO BE 0%. THREE DAYS LATER, THE SUBJECT WAS DISCHARGED ON ASPIRIN AND CLOPIDOGREL. IN (B)(6) 2024, THE SUBJECT DIED. AT THE TIME OF EVENT, THE SUBJECT WAS ON ASPIRIN AND CLOPIDOGREL. NO ACTION WAS TAKEN TO TREAT THE EVENT. IT IS UNKNOWN IF AUTOPSY WAS PERFORMED OR NOT. Manufacturer Narrative: B3: DATE OF EVENT: USED MAY 01, 2024, AS THE EXACT DEATH DATE WAS UNKNOWN. E1: INITIAL REPORTER ADDRESS 1: (B)(6). E1: INITIAL REPORTER PHONE: (B)(6).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21575374&pc=DSQ | 2916596-2025-01522 | 2023-11-13 05:00 | Death | THORATEC CORPORATION | 2025-03-11 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. Manufacturer Narrative: SECTION D4: PRIMARY UDI CORRECTED MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21575396&pc=DSQ | 2916596-2025-01426 | 2025-02-22 05:00 | Death | THORATEC CORPORATION | 2025-03-11 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. PER THE VAD TEAM, THE PATIENT WAS ADMITTED TO A HOSPICE FACILITY, HAD HIS DEVICE DEACTIVATED, AND PASSED AWAY ON (B)(6) 2025. THE DEVICE FUNCTIONED AS EXPECTED. IT WAS UNKNOWN IF AN AUTOPSY WAS PERFORMED AND WOULD BE PROVIDED. IT WAS UNKNOWN IF THE PUMP WAS EXPLANTED. THE PUMP WOULD NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE (HM) 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE (HM) 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HM 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21575466&pc=NAY | 2955842-2025-06857 | 2025-02-07 05:00 | Death | INTUITIVE SURGICAL, INC | 2025-03-11 04:00 | NAY | DAVINCI 5 | Adverse Event Without Identified Device or Use Problem | Coagulation Disorder; Hemorrhage/Bleeding; Thromboembolism; Insufficient Information | K232610 |  | 1 | Event Description: IT WAS REPORTED THAT AFTER A DA VINCI-ASSISTED BENIGN HYSTERECTOMY PROCEDURE, THE PATIENT EXPIRED ON POST-OPERATIVE DAY ONE. THE INTUITIVE SURGICAL CLINICAL SALES REPRESENTATIVE (CSR) WAS INFORMED THAT THE SURGEON OF THE PROCEDURE COULD NOT DETERMINE IF THE DEATH WAS RELATED TO THE PROCEDURE. THE CSR WAS NOT PROVIDED ANY DETAILS REGARDING THE COURSE OF EVENTS, EITHER DURING THE SURGICAL PROCEDURE OR POST-PROCEDURE LEADING TO THE PATIENT DEATH. THE PATIENT WAS STATED TO HAVE UNDERLYING HEALTH CONDITIONS, BUT SPECIFICS WERE NOT AVAILABLE. MULTIPLE REQUESTS FOR ADDITIONAL INFORMATION FROM THE SURGEON HAVE BEEN MADE; HOWEVER, NO RESPONSE HAS BEEN RECEIVED.Event Description: REFER TO H11 FOR FOLLOW-UP INFORMATION. Manufacturer Narrative: A REVIEW OF THE SYSTEM LOGS FOUND THAT NO RELEVANT ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: 0 DEGREE ENDOSCOPE, MONOPOLAR CURVED SCISSORS, FENESTRATED BIPOLAR FORCEPS, PROGRASP FORCEPS, MEGA SUTURECUT NEEDLE DRIVER. A SITE HISTORY REVIEW FOUND NO COMPLAINTS WERE MADE FOR THE INSTRUMENTS USED DURING THIS PROCEDURE. A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED FOR THE DEVICE(S) USED DURING THE PROCEDURE AND NO NON-CONFORMANCES WERE IDENTIFIED TO BE RELATED TO THIS EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT DIED ONE DAY AFTER A HYSTERECTOMY. NO DETAILS OF THE CASE ARE PROVIDED NOR ARE ANY OF THE EVENTS WHICH LED TO THE DEATH. BASED ON THE INFORMATION IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT. Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED FROM THE SITE RISK MANAGER. SECTION B5 INFORMATION: IT WAS REPORTED THAT THE PROCEDURE WAS A HYSTERECTOMY FOR A RIGHT OVARIAN MASS. THERE WERE NO INTRA-OPERATIVE COMPLICATIONS DURING THE DA VINCI-ASSISTED PROCEDURE. POST-OPERATIVELY, THE PATIENT WAS INITIALLY DOING WELL IN POST-ANESTHESIA CARE (PACU);HOWEVER, WAS LATER FOUND TO BE UNRESPONSIVE. THE PATIENT WAS RETURNED TO THE OR WHERE TWO AREAS OF RUPTURED ABDOMINAL WALL VARICES WERE FOUND AND LIGATED. THE ESTIMATED BLOOD LOSS WAS 6.7 LITERS. THE PATENT WAS TRANSFUSED WITH MULTIPLE BLOOD PRODUCTS. THE ABDOMEN WAS PACKED, ABTHERA WOUND DRESSING WAS PLACED, AND THE PATIENT WAS TRANSFERRED TO THE SURGICAL INTENSIVE CARE UNIT. THE PATIENT WAS LATER RETURNED TO THE OR AGAIN FOR BLEEDING. NO LARGE VESSELS WERE VISIBLE, THE BLEEDING APPEARED TO BE COAGULOPATHIC AND WAS NOTED FROM THE PERIHEPATIC REGION, AND FROM THE EDGES OF THE MIDLINE LAPAROTOMY. THE ABDOMEN WAS PACKED WITH LAP PADS AND THE BLEEDING SLOWED. DURING THE PROCEDURE, THE PATIENT'S PULSE WAS LOST; RESUSCITATION WAS SUCCESSFUL WITH RETURN OF SPONTANEOUS CIRCULATION. THE VASCULAR TEAM PERFORMED SUCTION EMBOLECTOMY OF CLOT FROM THE RIGHT ATRIUM AND RIGHT VENTRICLE; HOWEVER, SIGNIFICANT CLOT BURDEN REFORMED AS THE CLOTS WERE MECHANICALLY EXTRACTED. THE PATIENT'S FAMILY ELECTED TO WITHDRAW CARE AND THE PATIENT EXPIRED. THE CUSTOMER STATED THAT NO DA VINCI PRODUCTS CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE CUSTOMER ATTRIBUTED THE PATIENT OUTCOME TO THEIR HISTORY OF MELANOMA, ESOPHAGEAL VARICES, GASTROESOPHAGEAL REFLUX DISEASE (GERD), CIRRHOSIS, SEIZURES, AND HYPERTENSION. RE-REVIEW OF THE EVENT BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT HAD SIGNIFICANT COMORBIDITIES PRIOR TO THE OPERATION, SPECIFICALLY ALCOHOLIC LIVER DISEASE. THIS DIAGNOSIS CAN LEAD TO A PATIENT HAVING SIGNIFICANT ABNORMAL VENOUS COLLATERALIZATION BECAUSE OF PORTAL HYPERTENSION INCLUDING ABDOMINAL WALL AND GASTRIC VARICES AND COAGULOPATHY AS RESULT OF POOR HEPATIC FUNCTION. THIS PATIENT UNDERWENT A HYSTERECTOMY BUT LATER BLED FROM THE VARICES AND EVENTUALLY DIED AS A RESULT OF A CASCADE OF COMPLICATIONS FROM THESE VARICES. NO ALLEGATIONS HAVE BEEN MADE AGAINST ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS. THE POSTOPERATIVE COMPLICATIONS SUGGEST THIS DEATH WAS MOST LIKELY A RESULT OF THE PATIENT PRE-EXISTING DISEASE CONTRIBUTION. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, NO INTUITIVE SURGICAL PRODUCT OR INSTRUMENT CONTRIBUTED TO THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21575547&pc=OZD | 1220648-2025-26647 | 2025-02-12 05:00 | Death | ABIOMED, INC. | 2025-03-11 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Perforation of Vessels; Cardiac Perforation | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED AN IMPELLA CP PUMP WAS BEING PLACED ON A PATIENT FOR THE SUPPORT OF AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WHEN THERE WAS A GUIDEWIRE PERFORATION OF THE LEFT VENTRICLE. A THORACOTOMY HEMOSTASIS WAS PERFORMED AND CIRCULATION CONTINUED. THE PATIENT CONTINUED ON SUPPORT. CIRCULATION BECAME DIFFICULT TO MAINTAIN, AND THE DEVICE WAS REMOVED DUE TO DEATH. THE CAUSE OF DEATH WAS STATED BY THE FACILITY TO BE ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK AND THE LEFT VENTRICLE PERFORATION. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM, SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS & CAUTIONS: WARNINGS, SECTION: PRE-SUPPORT EVALUATION, SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER, SECTION: DIRECT AORTIC INSERTION, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION. [ADDITION FOR 5.0, 5.5 ONLY] IN INSTANCES WHERE THE IMPELLA PUMP HAS BEEN PLACED PRIOR TO PERFORMING CARDIAC SURGERY WITH AORTIC CROSS CLAMPING AND CARDIOPLEGIC ARREST, CARE SHOULD BE TAKEN WHEN MANIPULATING THE HEART WHEN THE PUMP POSITION IS FIXED WITH APPLICATION OF THE AORTIC CROSS CLAMP ACROSS THE PUMP CATHETER.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE. ¿TO REDUCE THE RISK OF CARDIAC INJURY (INCLUDING VENTRICULAR PERFORATION), PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY IN THE SETTINGS OF ACUTE INFARCTION WITH TISSUE NECROSIS.¿ ¿TO REDUCE THE RISK OF VASCULAR INJURY, PHYSICIANS SHOULD EXERCISE CAUTION WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX PERIPHERAL VASCULAR ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: UNREPAIRED ABDOMINAL AORTIC ANEURYSM, SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM, DISSECTION OF THE ASCENDING/ TRANSVERSE/DESCENDING AORTA, CHRONIC ANATOMICAL CHANGES IN THE RELATIONSHIP OF THE AORTA/AORTIC VALVE/VENTRICULAR ALIGNMENT, SIGNIFICANT MOBILE ATHEROMATOUS DISEASE IN THE THORACIC OR ABDOMINAL AORTA OR PERIPHERAL VESSELS.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING A RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY IN THE LEFT VENTRICLE AFTER CPR WITH ECHOCARDIOGRAPHY GUIDANCE. SECTION: WARNINGS & CAUTIONS: CAUTIONS: ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ Manufacturer Narrative: THE INVESTIGATION OF VENTRICLE PERFORATION HAS BEEN COMPLETED. NO PRODUCT WAS RETURNED FOR EVALUATION. CLINICAL DETAILS NOTED THAT PERFORATION OCCURRED DURING USE OF THE GUIDEWIRE DURING IMPELLA INSERTION STEPS. ADDITIONALLY, PHYSICIAN NOTED THAT THE PERFORATION OF THE VENTRICLE WAS SHARP AND RESEMBLED A PUNCTURE FROM A GUIDEWIRE AND NOT FROM FRAGILE TISSUE AND PIGTAIL PERFORATION. THE ROOT CAUSE OF THE VENTRICLE PERFORATION WAS MOST LIKELY DUE TO A GUIDEWIRE PUNCTURE DURING IMPELLA INSERTION. B2 DATE OF DEATH WAS ERRONEOUSLY ENTERED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26647. D6B EXPLANT DATE WAS ERRONEOUSLY ENTERED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26647. H6 CODE 2135 WAS ERRONEOUSLY ENTERED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26647. NEW CODE WAS ADDED TO HEALTH EFFECT - CLINICAL CODE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21575711&pc=DSQ | 2916596-2025-01429 | 2025-02-18 05:00 | Death | THORATEC CORPORATION | 2025-03-11 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Loss of consciousness; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS IN THE CARDIOLOGY UNIT AFTER THE INTENSIVE CARE UNIT (ICU), RECOVERING FROM SURGERY AND SUDDENLY UNDERWENT A RIGHT HEART FAILURE EPISODE. THE PATIENT WAS RETURNED TO THE ICU AND NEEDED TO BE INTUBATED AND GIVEN INOTROPES BEFORE THE PATIENT STABILIZED. THE PATIENT HAD STILL NOT AWOKEN. IT WAS NOTED THERE WAS A LOW FLOW ALARM. THE RIGHT HEART FAILURE (RHF) SEEMED TO HAVE COME FROM A BLOCKAGE OF ONE OF THE BYPASSES THAT DROVE BLOOD TO THE RIGHT VENTRICLE, AS THE PATIENT HAD BYPASSES FROM PREVIOUS SURGERIES. THE RHF WAS SUDDEN AND UNEXPECTED. THE RHF HAD BEEN CONTROLLED, AND THE PATIENT WAS BEING MONITORED. THE PATIENT WAS STABLE HEMODYNAMICALLY. THE PATIENTS BRAIN STATE WAS A CONCERN.Event Description: IT WAS ADDITIONALLY REPORTED THAT THE HEALTHCARE PROVIDERS DISCONTINUED THERAPY AND THEREFORE THE PATIENT EXPIRED FROM BRAIN DEATH DUE TO RIGHT HEART FAILURE FROM RIGHT VENTRICULAR INFARCTION. IT WAS NOTED THAT THE DEATH WAS NOT DEVICE RELATED. THE DEVICE OPERATED AS EXPECTED AND NO PRODUCT WOULD BE RETURNING. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT EXPIRATION CANNOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), WAS NOT RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), IS CURRENTLY AVAILABLE. SECTION 1 OF THIS DOCUMENT LISTS ADVERSE EVENTS, INCLUDING RIGHT HEART FAILURE, RESPIRATORY FAILURE, NEUROLOGICAL EVENTS (NOT STROKE-RELATED), AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, LISTS NEUROLOGIC DYSFUNCTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. SECTION 6 "PATIENT CARE AND MANAGEMENT" (UNDER "RIGHT HEART FAILURE") DISCUSSES THE POTENTIAL DEVELOPMENT OF RIGHT HEART FAILURE FOLLOWING IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS, INCLUDING INOTROPES AND RIGHT VENTRICULAR ASSIST DEVICE (RVAD) PLACEMENT. SECTION 6 (UNDER CAUTION!) STATES: "RIGHT HEART FAILURE CAN OCCUR FOLLOWING IMPLANTATION OF THE PUMP. RIGHT VENTRICULAR DYSFUNCTION, ESPECIALLY WHEN COMBINED WITH ELEVATED PULMONARY VASCULAR RESISTANCE, MAY LIMIT THE EFFECTIVENESS OF THE LVAS DUE TO REDUCED FILLING OF THE PUMP." NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: E1: REPORTER CONTACT INFORMATION WAS NOT AVAILABLE. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21575749&pc=DSQ | 2916596-2025-01523 | 2023-12-17 05:00 | Death | THORATEC CORPORATION | 2025-03-11 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE DEVICE WAS NOT RETURNED FOR EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED BY THE ACCOUNT. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM INSTRUCTIONS FOR USE (IFU), IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21575983&pc=OZD | 1220648-2025-26649 | 2025-02-13 05:00 | Death | ABIOMED, INC. | 2025-03-11 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Cardiovascular Insufficiency | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT PRESENTED FOR SHORTNESS OF BREATH AND WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. DURING SUPPORT, IT WAS REPORTED THAT HEPARIN WAS STARTED. THE PATIENT COUGHED AND EXPERIENCED BLEEDING AT THE ACCESS SITE. DRESSING WAS CHANGED AND PRESSURE WAS HELD. HEPARIN WAS TURNED OFF TO OBTAIN HEMOSTASIS. NO HEMATOMA WAS PRESENT. SUPPORT CONTINUED. SUBSEQUENTLY IT WAS NOTED THAT THE PATIENT DECOMPENSATED AFTER ROLLING FOR CLEANING. HE WENT INTO A WIDE COMPLEX RHYTHM, THEN INTO VENTRICULAR TACHYCARDIA AND THEN VENTRICULAR FIBRILLATION. ADVANCED CARDIOVASCULAR LIFE SUPPORT (ACLS) PROTOCOL WAS INITIATED AND PATIENT CODED FOR 20 MINUTES, UNTIL IT WAS CALLED AND HE WAS PRONOUNCED EXPIRED. THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA DEVICE AS AN ASSOCIATED FACTOR IN THE PATIENT'S DEMISE. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM, SECTION: GENERAL PATIENT CARE CONSIDERATIONS: ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT, USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21576730&pc=DSQ | 2916596-2025-01372 | 2025-02-07 05:00 | Death | THORATEC CORPORATION | 2025-03-11 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Pulmonary Embolism; Cardiac Arrest | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD CARDIAC ARREST PULSELESS ELECTRICAL ACTIVITY (PEA), LIKELY SECONDARY TO MASSIVE PULMONARY EMBOLISM PER TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE). CHEST COMPRESSIONS WERE INITIATED. THE PATIENT REQUIRED CARDIOPULMONARY RESUSCITATION (CPR), INOTROPES, VASOPRESSORS, TISSUE PLASMINOGEN ACTIVATOR (TPA), AND INHALED NITRIC OXIDE. THE PATIENT WAS ON ELIQUIS 2.5 MG TWICE A DAY AND ACETYLSALICYLIC ACID (ASA) 325 MG DAILY PER THE HOSPITAL ANTICOAGULATION PROTOCOL. A LOW FLOW ALARM OCCURRED DUE TO CARDIAC ARREST. THERE WERE NO CHANGES TO PUMP PARAMETERS. Manufacturer Narrative: SECTION B2: CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6) AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C, IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, LISTS VENOUS THROMBOEMBOLISM, ARTERIAL NON-CENTRAL NERVOUS SYSTEM (CNS) THROMBOEMBOLISM, AND DEATH AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS. ADDITIONALLY, SECTION 6 ¿PATIENT CARE AND MANAGEMENT¿ LISTS THROMBOEMBOLISM AS A POTENTIAL LATE POSTIMPLANT COMPLICATION THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, AND OUTLINES INDICATIONS OF THROMBUS AS WELL AS HOW TO RESPOND TO SUCH EVENTS. THIS SECTION ALSO PROVIDES INFORMATION REGARDING THE RECOMMENDED ANTICOAGULATION THERAPY AND INR (INTERNATIONAL NORMALIZED RATIO) RANGE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21576741&pc=NPT | 2025587-2025-01827 | 2025-03-06 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-11 04:00 | NPT | EVOLUT FX PLUS VALVE | Gradient Increase; Insufficient Information; Activation Failure; Patient Device Interaction Problem | Rupture | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING THE IMPLANT OF A TRANSCATHETER VALVE, A POST-IMPLANT BALLOON AORTIC VALVULOPLASTY (BAV) WAS PERFORMED USING A 22 MILLIMETER (MM) NON-MEDTRONIC BALLOON. SUBSEQUENTLY, AN ANNULAR RUPTURE WAS OBSERVED AND SURGICAL REPAIR WAS PERFORMED TO EXPLANT THE TRANSCATHETER VALVE AND ADDRESS THE TEAR. IT WAS REPORTED THAT THE POST-IMPLANT BAV AND CALCIFICATION CONTRIBUTED TO THE ANNULAR RUPTURE.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE POST-IMPLANT BAV WAS PERFORMED DUE TO A CONSTRAINED VALVE ON ECHOCARDIOGRAM, AND AN ELEVATED GRADIENT OF 14 MILLIMETERS OF MERCURY (MM HG). A NON-COMPLIANT BALLOON WAS USED AND 1 INFLATION WAS PERFORMED FOR 5 SECONDS USING A 30 CUBIC CENTIMETER (CC) SYRINGE. A PRESSURE GAUGE WAS NOT USED. SUBSEQUENTLY, INFORMATION WAS RECEIVED THAT THE PATIENT DIED.Event Description: ADDITIONAL INFORMATION WAS RECEIVED WHICH CONFIRMED THAT PER THE PHYSICIAN, THE CAUSE OF DEATH WAS THE ANNULAR RUPTURE RELATED TO THE POST-IMPLANT BAV. Manufacturer Narrative: UPDATED DATA: B2. DATE OF DEATH ADDED B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B1. B2. B5. H1. H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID: {D-EVOLUTFX-2329}; PRODUCT TYPE: {0195-HEARTVALVES}; PRODUCT ID: [22 MM DILATATION BALLOON}; PRODUCT TYPE: {BALLOON AORTIC VALVULOPLASTY}. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21576962&pc=DSQ | 2916596-2025-01343 | 2025-02-18 05:00 | Death | THORATEC CORPORATION | 2025-03-11 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Cardiogenic Shock; Respiratory Failure; Unspecified Nervous System Problem | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED A TEAR IN THE POSTERIOR WALL OF THE LEFT VENTRICLE IMMEDIATELY POST OP WITH INITIAL LVAD IMPLANT SURGERY. ACCORDING TO THE APP, THE FLOWS WENT TO 0.7 AND THE PATIENT EXPERIENCED IMMENSE BLEEDING FROM CHEST TUBES. PATIENT LOST A PULSE AND BLOOD PRESSURE. THE SURGEON EMERGENTLY REOPENED THE PATIENT'S CHEST AND CANNULATED FOR CARDIOPULMONARY BYPASS (CPB). PATIENT'S DOWN TIME WAS 20 MINS BEFORE INITIATION OF CPB AND RESTORATION OF BLOOD FLOW. THE PATIENT WAS CRITICALLY ILL AND INTUBATED WITH NO SEDATION WITH PARTIAL LOSS OF CRANIAL NERVE REFLEXES. PATIENT WAS TERMINALLY EXTUBATED AND PASSED AWAY ON (B)(6) 2025 DUE TO CARDIOGENIC SHOCK. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED; IT WAS DUE TO TERMINAL EXTUBATION. AN AUTOPSY WOULD NOT BE PERFORMED, AND THE PUMP WOULD NOT BE EXPLANTED OR RETURNED.Event Description: SUBMITTED LOG FILES CAPTURED THE PUMP STARTUP AND LOW FLOW EVENTS ON (B)(6) 2025. THE CALCULATED FLOW APPEARS TO HAVE FLUCTUATED BELOW THE ALARM THRESHOLD OF 2.0 LPM DURING THE EVENTS. THE LOG FILE CONTAINED LOW FLOW ESTIMATES AS WELL AS SUSTAINED LOW FLOW HAZARD EVENTS. THE PUMP APPEARED TO STABILIZE AS THE SPEED WAS ADJUSTED TO 4800 RPMS. STARTING ON 19FEB2025 AT 4:13:25, THE FLOW APPEARED TO HAVE STABILIZED. Manufacturer Narrative: B2 - OUTCOMES ATTRIBUTED TO ADVERSE: CORRECTED MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS INVESTIGATION. REVIEW OF THE SUBMITTED LOG FILES CONFIRMED LOW FLOW ALARMS; HOWEVER, A SPECIFIC CAUSE FOR THE LOW FLOW ALARMS COULD NOT CONCLUSIVELY BE DETERMINED THROUGH THIS INVESTIGATION. THE CONTROLLER EVENT LOG FILE CONTAINED EVENTS FROM (B)(6) 2025, PER THE TIMESTAMPS. MULTIPLE LOW FLOW ALARMS WERE CAPTURED ON (B)(6) 2025, AS WELL AS NUMEROUS LOW FLOW FAULT FLAGS THAT DID NOT PERSIST LONG ENOUGH TO ACTIVATE A LOW FLOW ALARM. FLOW APPEARED TO IMPROVE WHEN THE PUMP SPEED WAS INCREASED TO 4800 RPM. NO OTHER NOTABLE ALARMS WERE CAPTURED, AND THE PUMP APPEARED TO FUNCTION AS INTENDED THROUGHOUT THE LOG FILE. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C, IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, "INTRODUCTION," LISTS ADVERSE EVENTS, INCLUDING BLEEDING, NEUROLOGICAL EVENTS, RESPIRATORY FAILURE AND DEATH, WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS. THIS DOCUMENT ALSO PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) RANGE, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. SECTION 1 ALSO PROVIDES AN EXPLANATION OF EACH OF THE PUMP PARAMETERS, INCLUDING PUMP FLOW. SECTION 4, ¿SYSTEM MONITOR,¿ EXPLAINS THAT THE LOW FLOW HAZARD ALARM OCCURS WHEN PUMP FLOW IS LESS THAN 2.5 LPM. A 10-SECOND DELAY IS IMPOSED BETWEEN THE DETECTION OF THE LOW FLOW STATUS AND THE ACTIVATION OF THE ASSOCIATED AUDIO AND VISUAL INDICATORS ON THE SYSTEM CONTROLLER. CHANGES IN PATIENT CONDITIONS CAN RESULT IN LOW FLOW. SECTION 7, ¿ALARMS AND TROUBLESHOOTING,¿ EXPLAINS SYSTEM ALARMS AND THE RECOMMENDED ACTIONS ASSOCIATED WITH THEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21576972&pc=LWS | 9614453-2025-00987 | 2025-03-04 05:00 | Death | MEDTRONIC EUROPE SARL | 2025-03-11 04:00 | LWS | EVERA MRI S VR SURESCAN | Device Sensing Problem | Ventricular Fibrillation | P980016 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS TRANSPORTED TO HOSPITAL VIA AMBULANCE. ELECTROCARDIOGRAM (ECG) STRIPS SHOWED VENTRICULAR FIBRILLATION (VF) SO THE PATIENT WAS TREATED WITH EXTERNAL DEFIBRILLATION AND PASSED AWAY SHORTLY THEREAFTER. UNDERSENSING OF THE RIGHT VENTRICULAR (RV) LEAD AND A POTENTIAL IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) DETECTION ISSUE WERE SUSPECTED.Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED THAT, PER THE AUTOPSY REPORT- THE CAUSE OF DEATH WAS ACUTE CARDIOVASCULAR FAILURE SECONDARY TO DILATED MYOCARDIAL ISCHEMIA. IT WAS FURTHER NOTED FROM THE AUTOPSY THAT THERE WAS MARKED CORONARY SCLEROSIS, SEVERE MYOCARDIOSCLEROSIS WITH LARGE AREAS OF MYOCARDIAL SCARRING IN THE ANTERIOR AND INFERIOR LEFT VENTRICULAR REGION, PULMONARY EDEMA, AND DIABETIC GLOMERULOSCLEROSIS. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID: 6935M62, IMPLANT DATE: (B)(6) 2016. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE DEVICE WAS NOT RETURNED FOR ANALYSIS, HOWEVER, PERFORMANCE DATA COLLECTED FROM THE DEVICE WAS RECEIVED AND ANALYZED. ANALYSIS OF THE DEVICE MEMORY WAS PERFORMED AND NO ANOMALIES WERE FOUND. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE DEVICE WAS NOT RETURNED FOR ANALYSIS; HOWEVER, PERFORMANCE DATA COLLECTED FROM THE DEVICE WAS RECEIVED AND ANALYZED. ANALYSIS OF THE DEVICE MEMORY WAS PERFORMED AND NO ANOMALIES WERE FOUND. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21576992&pc=LWS | 3008973940-2025-01830 | 2025-03-04 05:00 | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-03-11 04:00 | LWS | SPRINT QUATTRO SECURE S MRI SURESCAN | Under-Sensing | Ventricular Fibrillation | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS TRANSPORTED TO HOSPITAL VIA AMBULANCE. ELECTROCARDIOGRAM (ECG) STRIPS SHOWED VENTRICULAR FIBRIL LATION (VF) SO THE PATIENT WAS TREATED WITH EXTERNAL DEFIBRILLATION, AND PASSED AWAY SHORTLY THEREAFTER. UNDERSENSING OF THE RIGHT VENTRICULAR (RV) LEAD AND A POTENTIAL IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) DETECTION ISSUE WERE SUSPECTED.Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED THAT, PER THE AUTOPSY REPORT- THE CAUSE OF DEATH WAS ACUTE CARDIOVASCULAR FAILURE SECONDARY TO DILATED MYOCARDIAL ISCHEMIA. IT WAS FURTHER NOTED FROM THE AUTOPSY THAT THERE WAS MARKED CORONARY SCLEROSIS, SEVERE MYOCARDIOSCLEROSIS WITH LARGE AREAS OF MYOCARDIAL SCARRING IN THE ANTERIOR AND INFERIOR LEFT VENTRICULAR REGION, PULMONARY EDEMA, AND DIABETIC GLOMERULOSCLEROSIS. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID: DVMC3D4, IMPLANT DATE (B)(6) 2016. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE LEAD WAS NOT RETURNED FOR ANALYSIS, HOWEVER, PERFORMANCE DATA COLLECTED FROM THE DEVICE WAS RECEIVED AND ANALYZED. ANALYSIS OF THE DEVICE MEMORY WAS PERFORMED AND NO ANOMALIES WERE FOUND. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE PROXIMAL PORTION OF THE LEAD WAS RETURNED, ANALYZED, AND NO ANOMALIES WERE FOUND. ANALYSIS OF THE DEVICE MEMORY WAS PERFORMED AND NO ANOMALIES WERE FOUND. VISUAL ANALYSIS OF THE LEAD INDICATED APPARENT EXPLANT DAMAGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21576994&pc=CBK | 2518422-2025-031958 | 2023-07-10 04:00 | Death | RESPIRONICS, INC. | 2025-03-11 04:00 | CBK | TRILOGY EVO | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K181166 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT PASSED AWAY WHILE USING A VENTILATOR. THE INFORMATION INDICATES EMPORIUM IS UNABLE TO RECOVER THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21577017&pc=DSQ | 2916596-2025-01449 | 2024-11-20 05:00 | Death | THORATEC CORPORATION | 2025-03-11 04:00 | DSQ | THORATEC® HEARTMATE II®, LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2024 AND WAS FOUND DEAD IN THEIR HOME. THE CAUSE WAS UNKNOWN. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. THE PUMP WOULD NOT BE EXPLANTED OR RETURNED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE SUBMITTED LOG FILE CAPTURED THE PUMP OPERATING AS INTENDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE II LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21577130&pc=DSQ | 2916596-2025-01379 | 2025-01-22 05:00 | Death | THORATEC CORPORATION | 2025-03-11 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT WITH WORSENING RIGHT VENTRICULAR FAILURE WAS ADMITTED ON (B)(6) 2025. DISCUSSION WAS MADE TO HAVE COMFORT MEASURES AND TRANSITION TO PALLIATIVE CARE. THE PATIENT WAS ADMITTED TO HOSPICE (B)(6) 2025. THEY DID NOT HAVE RIGHT HEART FAILURE EXIST PRE-LEFT VENTRICULAR ASSIST DEVICE (LVAD) IMPLANT. RIGHT HEART WAS NOT CAUSED OR CONTRIBUTED BY ANY DEVICE RELATED ISSUES. THE OUTCOME WAS NOT CONSIDERED TO BE DEVICE RELATED. THE DEVICE WAS FUNCTIONING AS INTENDED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND THE PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING RIGHT HEART FAILURE AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6, UNDER ¿RIGHT HEART FAILURE¿, STATES THAT PATIENTS MAY DEVELOP RIGHT VENTRICULAR FAILURE DURING OR SHORTLY AFTER IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21577725&pc=FKX | 0008030665-2025-00572 | 2025-02-27 05:00 | Death | ERIKA DE REYNOSA, S.A. DE C.V. | 2025-03-11 04:00 | FKX | LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Dyspnea | K181108 |  | 1 | Event Description: A PERITONEAL DIALYSIS REGISTERED NURSE (PDRN) REPORTED THAT PERITONEAL DIALYSIS (PD) PATIENT PASSED AWAY ON (B)(6) 2025. ADDITIONAL INFORMATION WAS OBTAINED THROUGH FOLLOW-UP WITH THE PDRN. THE PATIENT WAS IN PD TREATMENT CONNECTED TO THE LIBERTY SELECT CYCLER ON (B)(6)2025. THE PATIENT EXPERIENCED SHORTNESS OF BREATH. THE PATIENT¿S PARTNER CALLED 911 AND DISCONNECTED THE PATIENT FROM THE TREATMENT. EMERGENCY MEDICAL SERVICES (EMS) ARRIVED TO FIND THE PATIENT IN CARDIAC ARREST. CARDIOPULMONARY RESUSCITATION (CPR) WAS ADMINISTERED. RESUSCITATIVE EFFORTS WERE UNSUCCESSFUL, AND THE PATIENT WAS PRONOUNCED DECEASED IN THE HOME. THE PATIENT WAS TRANSPORTED TO THE FUNERAL HOME. NO AUTOPSY WAS CONDUCTED. THE CAUSE OF DEATH IS DOCUMENTED AS CARDIAC ARREST, CAUSE UNKNOWN. NO ISSUES WITH THE LIBERTY SELECT CYCLER OR THE PATIENT¿S TREATMENT WERE REPORTED PRIOR TO THE PATIENT¿S DEATH. NO ADDITIONAL INFORMATION WAS AVAILABLE.Event Description: A PERITONEAL DIALYSIS REGISTERED NURSE (PDRN) REPORTED THAT PERITONEAL DIALYSIS (PD) PATIENT PASSED AWAY ON (B)(6) 2025. ADDITIONAL INFORMATION WAS OBTAINED THROUGH FOLLOW-UP WITH THE PDRN. THE PATIENT WAS IN PD TREATMENT CONNECTED TO THE LIBERTY SELECT CYCLER ON (B)(6) 2025. THE PATIENT EXPERIENCED SHORTNESS OF BREATH. THE PATIENT¿S PARTNER CALLED 911 AND DISCONNECTED THE PATIENT FROM THE TREATMENT. EMERGENCY MEDICAL SERVICES (EMS) ARRIVED TO FIND THE PATIENT IN CARDIAC ARREST. CARDIOPULMONARY RESUSCITATION (CPR) WAS ADMINISTERED. RESUSCITATIVE EFFORTS WERE UNSUCCESSFUL, AND THE PATIENT WAS PRONOUNCED DECEASED IN THE HOME. THE PATIENT WAS TRANSPORTED TO THE FUNERAL HOME. NO AUTOPSY WAS CONDUCTED. THE CAUSE OF DEATH IS DOCUMENTED AS CARDIAC ARREST, CAUSE UNKNOWN. NO ISSUES WITH THE LIBERTY SELECT CYCLER OR THE PATIENT¿S TREATMENT WERE REPORTED PRIOR TO THE PATIENT¿S DEATH. NO ADDITIONAL INFORMATION WAS AVAILABLE. Manufacturer Narrative: CLINICAL REVIEW: THERE IS A TEMPORAL RELATIONSHIP BETWEEN PERITONEAL DIALYSIS THERAPY UTILIZING THE LIBERTY SELECT CYCLER AND THE PATIENT EVENT OF SHORTNESS OF BREATH WITH SUBSEQUENT CARDIAC ARREST AND DEATH. HOWEVER, THERE IS NO DOCUMENTATION TO SHOW A CAUSAL RELATIONSHIP BETWEEN THE ADVERSE EVENT AND USE OF THE LIBERTY SELECT CYCLER. ADDITIONALLY, THERE IS NO ALLEGATION OF A DEVICE MALFUNCTION OR DEFICIENCY REPORTED FOR THIS EVENT. SUDDEN CARDIAC DEATH (SCD) ACCOUNTS FOR A LARGE PROPORTION OF CARDIOVASCULAR DEATHS AND ALMOST ONE FOURTH OF THE OVERALL MORTALITY IN DIALYSIS PATIENTS. SCD IS PREVALENT IN BOTH PATIENTS WITH A HISTORY OF CARDIAC DISEASE AS WELL AS THOSE WITHOUT. BASED ON THE AVAILABLE INFORMATION AND NO ALLEGATION OR EVIDENCE OF A MALFUNCTION OR DEFICIENCY, THE LIBERTY SELECT CYCLER CAN BE EXCLUDED AS THE CAUSE OF THE PATIENT¿S SHORTNESS OF BREATH FOLLOWED BY CARDIAC ARREST AND DEATH. THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. Manufacturer Narrative: PLANT INVESTIGATION: THE ACTUAL DEVICE WAS RETURNED. A VISUAL INSPECTION OF THE RETURNED CYCLER EXTERIOR SHOWED NO SIGNS OF PHYSICAL DAMAGE. THERE WERE NO VISUAL INDICATIONS OF DRIED FLUID WITHIN THE CASSETTE COMPARTMENT ON THE PUMP. THERE WERE NO VISUAL INDICATIONS OF PARTICULATES WITHIN THE CASSETTE AREA. THERE WERE NO BURRS OR SHARP EDGES IN CASSETTE AREA THAT MAY HAVE PUNCTURED A CASSETTE MEMBRANE. A (AS-RECEIVED WITH REDUCED DWELL) SIMULATED TREATMENT WAS PERFORMED AND COMPLETED. VALVE ACTUATION TEST ¿ PASSED. SYSTEM AIR LEAK TEST ¿ PASSED. THE DEVICE HISTORY RECORD DID NOT REVEAL ANY ISSUES OR PROBLEMS RELATED TO THE REPORTED SYMPTOM CODE(S).UPON COMPLETION OF THE EVALUATION, THERE WERE NO MALFUNCTIONS THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. AN ASSOCIATED CAUSE COULD NOT BE DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21578691&pc=LWS | 3004209178-2025-04417 | 2021-04-06 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-03-11 04:00 | LWS | EVERA¿ XT DR | Unauthorized Access to Computer System | Myocardial Infarction | P980016 |  | 1 | Event Description: IT WAS REPORTED BY THE PATIENT'S REPRESENTATIVE THAT "SOMEONE WITH A TOWER PROGRAMMER HACKED INTO THE STREAM AND TURNED OFF PEOPLE'S PACEMAKERS, CAUSING THEM TO HAVE HEART ATTACKS." THE PATIENT PASSED AWAY. INFORMATION REGARDING THE PATIENT'S CAUSE OF DEATH COULD NOT BE OBTAINED. Manufacturer Narrative: CONTINUATION OF D10: 407652 LEAD IMPLANTED: (B)(6) 2010. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21579917&pc=MIH | 2017233-2025-05919 | 2025-02-17 05:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-03-11 04:00 | MIH | GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM | Adverse Event Without Identified Device or Use Problem | Ruptured Aneurysm | P040043 |  | 1 | Event Description: THE FOLLOWING INFORMATION WAS REPORTED TO GORE: ON (B)(6) 2025, THE PATIENT UNDERWENT AN EMERGENCY ENDOVASCULAR TREATMENT FOR A RUPTURED THORACIC ANEURYSM USING GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM. THE PATIENT COMPLAINED OF BACK PAIN SEVERAL DAYS AGO, AND IT WAS REPORTEDLY DUE TO A SEALED ANEURYSM RUPTURE. AFTER A 2-DEBRANCH BYPASS, THE FIRST DEVICE (TGM343420J) WAS IMPLANTED IN THE DESCENDING AORTA. THE SECOND DEVICE (TGM454520J) WAS DEPLOYED FROM A POSITION WHERE ITS PARTIAL UNCOVER STENT COVERED THE BRACHIOCEPHALIC ARTERY. FOLLOWING IMPLANTATION, THE BLOOD PRESSURE (MEASURED AT THE RIGHT HAND) DECREASED, AND OCCLUSION OF THE BRACHIOCEPHALIC ARTERY WAS SUSPECTED. ANGIOGRAPHY REVEALED THAT THE BRACHIOCEPHALIC ARTERY WAS COVERED BY THE SECOND DEVICE, AND BLOOD FLOW WAS REDUCED. ADDITIONALLY, AORTIC REGURGITATION (INFLOW OF CONTRAST MEDIUM INTO THE HEART) WAS ALSO CONFIRMED. THERE WAS NO ST ELEVATION OBSERVED ON THE MONITOR. A COVERED STENT (NON-GORE DEVICE) WAS INITIALLY PLACED IN THE BRACHIOCEPHALIC ARTERY AS A TREATMENT. BLOOD FLOW TO THE BRACHIOCEPHALIC ARTERY WAS IMPROVED; HOWEVER, THE HEARTBEAT REMAINED WEAK. CARDIAC MASSAGE AND ADRENALINE ADMINISTRATION WERE PERFORMED, AND THE PATIENT INITIALLY RECOVERED, BUT A FEW MINUTES LATER, VENTRICULAR FIBRILLATION OCCURRED, AND THE PATIENT WENT INTO CARDIAC ARREST. PERCUTANEOUS CARDIOPULMONARY SUPPORT (PCPS) WAS INITIATED, BUT THE PATIENT DID NOT RECOVER. ECHOCARDIOGRAPHY SHOWED BLOOD LEAKAGE INTO THE LEFT THORACIC CAVITY, AND FRESH BLOOD WAS CONFIRMED BY PUNCTURE. IT COULD NOT BE DETERMINED WHETHER THE BLOOD LEAKAGE INTO THE LEFT THORACIC CAVITY WAS DUE TO A RE-RUPTURED ANEURYSM DURING THE PROCEDURE OR FROM BLEEDING CAUSED BY THE INITIAL RUPTURE PRIOR TO THE PROCEDURE. CARDIAC MASSAGE WAS CONTINUED, BUT THE CARDIAC RHYTHM COULD NOT BE RESTORED, AND THE PATIENT EXPIRED. NO AUTOPSY INFORMATION IS AVAILABLE. THE PHYSICIAN'S COMMENTS: THE DEFINITE CAUSE OF DEATH IS UNKNOWN. THERE WERE NO CHOICES IN LIFE-SAVING MEASURES OTHER THAN TEVAR, AND THE INTERVENTION ITSELF WAS NOT AN INCORRECT CHOICE. SEVERAL POSSIBILITIES ARE AS FOLLOWS: INSUFFICIENT BLOOD FLOW TO THE BRACHIOCEPHALIC ARTERY DUE TO ITS COVERAGE BY THE CTAGAC. HOWEVER, THIS EVENT ALONE IS NOT BELIEVED TO HAVE AFFECTED THE CARDIAC RHYTHM. GUIDEWIRE MANIPULATION RESULTING IN DAMAGE TO THE VALVE LEAFLETS AND AORTIC REGURGITATION. THE GUIDEWIRE POTENTIALLY INTERFERING WITH THE ANASTOMOSIS OF THE ASCENDING AORTA FOR CORONARY ARTERY BYPASS GRAFTING. ANEURYSM RE-RUPTURE DUE TO ELEVATED PRESSURE CAUSED BY CARDIAC MASSAGE AND ADRENALINE ADMINISTRATION. Manufacturer Narrative: H6: A REVIEW OF THE MANUFACTURING RECORDS FOR THE DEVICE VERIFIED THAT THE LOT MET ALL PRE-RELEASE SPECIFICATIONS. H6: CODE D12: ACCORDING TO THE GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM INSTRUCTIONS FOR USE (IFU), ADVERSE EVENTS THAT MAY OCCUR AND/OR REQUIRE INTERVENTION INCLUDE, BUT ARE NOT LIMITED TO : DEATH, AORTIC RUPTURE, CARDIAC (E.G., ARRHYTHMIA, MYOCARDIAL INFARCTION, CONGESTIVE HEART FAILURE, HYPOTENSION OR HYPERTENSION) W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21579947&pc=DYB | 3020347218-2025-00016 | 2025-02-11 05:00 | Death | INARI MEDICAL, INC. | 2025-03-11 04:00 | DYB | INTRI24 INTRODUCER SHEATH | Improper or Incorrect Procedure or Method | Hemorrhage/Bleeding | K233646 |  | 1 | Event Description: AN 85-YEAR-OLD FEMALE ARRIVED AT THE HOSPITAL WITH SHORTNESS OF BREATH AND CHEST PAIN. THE PATIENT HAD A HISTORY OF CARDIAC ISSUES AND WAS LISTED AS "DO NOT RESUSCITATE". THE PATIENT WAS DIAGNOSED WITH MASSIVE PULMONARY EMBOLISM AND A THROMBECTOMY WAS PERFORMED WITH INARI DEVICES. THE PATIENT WAS INTUBATED AND ON TWO PRESSORS AT THE BEGINNING OF THE THROMBECTOMY. THE PROCEDURE BEGAN WITHOUT ISSUES AND A LARGE AMOUNT OF CLOT WAS REMOVED FROM THE PATIENT'S RIGHT SIDE THROUGH FIVE ASPIRATIONS WITH THE TRIEVER20 (T20). THE PATIENT'S OXYGEN (O2) SATURATION ROSE, AND IMAGING REVEALED THAT THE RIGHT SIDE WAS CLEAR. THE PHYSICIAN RE-POSITIONED TO THE LEFT AND WAS IN PROCESS OF OBTAINING WIRE POSITIONING WHEN A NURSE NOTICED A LARGE VOLUME OF BLOOD ON THE GROUND. UPON FURTHER INVESTIGATION, THE STOPCOCK OF THE LARGE BORE SYRINGE SIDE PORT WAS NOT CLOSED AFTER IMAGING WAS PERFORMED. IT WAS ESTIMATED THAT THERE WAS 500 CC OF BLOOD LOST FROM THE SIDE PORT BEING LEFT OPEN. THIS WAS IN ADDITION TO 300 CC OF BLOOD LOSS FROM ASPIRATIONS LEADING TO A TOTAL OF APPROXIMATELY 800 CC OF BLOOD LOSS. THE PATIENT'S BLOOD PRESSURE AND O2 DECLINED. THE PHYSICIAN MANAGED TO RE-POSITION TO THE LEFT SIDE WHILE AN ORDER OF BLOOD WAS PROCESSING. TWO FURTHER ASPIRATIONS WERE PERFORMED WHICH CLEARED SOME CLOT. FLOWSAVER WAS USED TO RETURN BLOOD. FLUID AND BLOOD WERE ADMINISTERED TO THE PATIENT. THE PHYSICIAN PERFORMED TWO ADDITIONAL ASPIRATIONS BUT WAS UNABLE TO FULLY CLEAR THE LEFT SIDE. AT THE TIME THE PROCEDURE WAS CONCLUDED, AND THE PATIENT WAS TRANSFERRED TO THE INTENSIVE CARE UNIT. THE PATIENT PASSED AWAY IN THE MIDDLE OF THE NIGHT. Manufacturer Narrative: THE INARI DEVICES WERE DISCARDED BY THE USER FACILITY AND ARE NOT AVAILABLE FOR EVALUATION. THE DEVICE IDENTIFIERS WERE NOT PROVIDED, THE LOT HISTORY RECORDS OF ALL POTENTIAL LOTS SHIPPED TO THE FACILITY WERE REVIEWED. THERE WERE NO DISCREPANCIES OR UNUSUAL FINDINGS. THERE WAS NO REPORT OF ANY DEVICE MALFUNCTION. BASED ON THE INFORMATION REVIEWED, THERE IS NO EVIDENCE TO INDICATE THE PRESENCE OF A POTENTIAL QUALITY ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. THE CASE WAS REVIEWED BY INARI'S CHIEF MEDICAL OFFICER, WHO CONCLUDED THAT THE PATIENT'S DEATH WAS THE RESULT OF SEVERE DISEASE STATE AND USER ERROR (BLOOD LOSS DUE TO FAILURE TO CLOSE THE SIDE PORT STOPCOCK). BLOOD LOSS, BLEEDING, AND DEATH ARE LISTED AS POTENTIAL COMPLICATIONS/ADVERSE EVENTS. MANUFACTURER REFERENCE #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21580132&pc=NPT | 2025587-2025-01835 | 2025-02-11 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-11 04:00 | NPT | EVOLUT FX PLUS VALVE | Patient Device Interaction Problem | Dysphagia/ Odynophagia; Ecchymosis; Hypoxia; Muscle Weakness; Visual Impairment; Burning Sensation; Dizziness; Sore Throat; Respiratory Failure; Speech Disorder; Ischemia Stroke; Heart Block | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT THE SUBJECT WITH A HISTORY OF FIRST DEGREE ATRIO-VENTRICULAR BLOCK AND BIFASCICULAR BLOCK WITH A RIGHT BUNDLE AND LEFT ANTERIOR FASCICULAR BLOCK HAD DISCUSSION OF A PERMANENT PACEMAKER IMPLANT PRIOR TO THE TRANSCATHETER BIOPROSTHETIC AORTIC VALVE IMPLANT. DURING THE VALVE IMPLANT PROCEDURE, THE SUBJECT DEVELOPED COMPLETE HEART BLOCK AND BECAME PACEMAKER DEPENDENT. A TEMPORARY PACEMAKER WAS PLACED. THE COMPLETE HEART BLOCK WAS REPORTED AS CAUSAL TO THE IMPLANT PROCEDURE AND VALVE. ALSO DURING THE IMPLANT PROCEDURE, ACUTE ONSET DYSARTHRIA AND EYE MOVEMENT ABNORMALITIES DEVELOPED. THE INITIAL NEUROLOGIC ASSESSMENT SUSPECTED AN EMBOLIC STROKE BUT SUGGESTED FURTHER WORK-UP. THE VALVE WAS IMPLANTED. NEUROLOGY CONTINUED TO ASSESS THE SUBJECT AND SUSPECTED A STROKE. PATIENT WAS UNDER OBSERVATION AND HAD NOT YET BEEN GIVEN TREATMENT FOR STROKE-LIKE SYMPTOMS. THE PHYSICIAN REPORTED THE SUSPECT STROKE WAS CAUSAL RELATED TO THE DELIVERY CATHETER SYSTEM (DCS) AND IMPLANT PROCEDURE. THE DAY FOLLOWING THE VALVE IMPLANT A PERMANENT PACEMAKER WAS IMPLANTED. HOWEVER, IT WAS ALSO REPORTED THAT THE PERMANENT PACEMAKER HAD NOT BEEN IMPLANTED. THE PRIMARY INDICATION FOR THE PERMANENT PACEMAKER WAS A CHRONIC BIFASCICULAR BLOCK. HOSPITALIZATION WAS PROLONGED AS RESULT OF THE EVENT. TWO DAYS LATER THE WOUND NURSE WAS CONSULTED TO ASSESS GROIN SKIN CONCERNS. UPON ASSESSMENT, A ¿LARGE¿ AREA OF ECCHYMOSIS WAS PRESENT ON THE BILATERAL GROIN AND THE LOWER ABDOMEN DUE TO THE RECENT PROCEDURES. HOWEVER, NO SIGNIFICANT OPEN ULCERATIONS WERE NOTED. ALL AREAS WERE TO BE LEFT OPEN AND TOPICAL DRESSINGS WERE NOT REQUIRED. THE PHYSICIAN INDICATED THE ECCHYMOSIS WAS CAUSAL RELATED TO THE PROCEDURE. ULTIMATELY, AN ISCHEMIC STROKE WAS REPORTED. THE PATIENT MOVED TO COMFORT CARE AND DIED 5 DAYS FOLLOWING THE VALVE IMPLANT DUE TO THE STROKE. THE DEATH WAS REPORTED AS A NON-SUDDEN CARDIAC DEATH.Event Description: ADDITIONALLY IT WAS REPORTED THAT DURING THE IMPLANT PROCEDURE THE VALVE AND DELIVERY CATHETER SYSTEM HAD BEN ADVANCED TO THE ASCENDING AORTA BEGAN BEFORE ACHIEVING A THERAPEUTIC ACTIVATED CLOTTING TIME (ACT) LEVEL. THIS WAS NOTED WHEN TWO READINGS IN THE MID TO LOW 100 SECOND LEVEL WERE OBSERVED. A LEAK IN THE INTRAVENOUS TUBING PREVENTED THE DELIVERY OF ANTICOAGULANT AND OTHER ANESTHETIC DRUGS. AS A RESULT, THE PATIENT EXPERIENCED A FACIAL BURNING SENSATION, DIZZINESS, AND THROAT PAIN, INITIALLY THOUGHT TO BE A CONTRAST REACTION. STEROIDS AND ANTIHISTAMINES WERE ADMINISTERED, BUT IT WAS LATER DISCOVERED THAT NO DRUGS HAD REACHED THE PATIENT DUE TO THE DEFECTIVE TUBING. HEPARIN WAS SUBSEQUENTLY ADMINISTERED VIA A STERILE CENTRAL VENOUS LINE, AND THE PROCEDURE CONTINUED. THE FINAL ACTIVATED CLOTTING TIME MEASUREMENT WAS IN THE LOW 300 SECOND RANGE, WHICH WAS TAKEN FOLLOWING VALVE DEPLOYMENT. MAGNETIC RESONANCE IMAGING LATER CONFIRMED THE PATIENT SUFFERED A STROKE, DESPITE THREE COMPUTED TOMOGRAPHY SCANS SHOWING NO EVIDENCE OF STROKE. DUAL ANTIPLATELET THERAPY (DAPT) WAS STARTED FOR THE STROKE AND SYMPTOMS. AS REPORTED, THE PATIENT EXPERIENCED ONGOING DISABILITY AS A R ESULT OF THE STROKE. DUE TO ¿SEVERE¿ DYSARTHRIA AND BILATERAL LOWER FACIAL WEAKNESS DYSPHAGIA, A NASOGASTRIC (NG) TUBE WAS PLACED THE DAY FOLLOWING THE VALVE IMPLANT TO FACILITATE MEDICATION ADMINISTRATION. THE PHYSICIAN INDICATED THE PATIENT¿S MAIN RISK FACTORS FOR STROKE WERE BILATERAL CAROTID ARTERY PLAQUE WITH 50-60% STENOSIS OF THE RIGHT INTERNAL CAROTID ARTERY AND <(><<)> 50% STENOSIS IN THE LEFT INTERNAL CAROTID ARTERY. AN AUTOPSY WAS NOT PERFORMED. THE ISCHEMIC EMBOLIC STROKE AND SUBSEQUENT DEATH WERE REPORTED AS CAUSAL TO THE PROCEDURE AND THE DELIVERY CATHETER SYSTEM (DCS).Event Description: IT WAS REPORTED THAT THE SUBJECT WITH A HISTORY OF FIRST DEGREE ATRIO-VENTRICULAR BLOCK AND BIFASCICULAR BLOCK WITH A RIGHT BUNDLE AND LEFT ANTERIOR FASCICULAR BLOCK HAD DISCUSSION OF A PERMANENT PACEMAKER IMPLANT PRIOR TO THE TRANSCATHETER BIOPROSTHETIC AORTIC VALVE IMPLANT. DURING THE VALVE IMPLANT PROCEDURE, THE SUBJECT DEVELOPED COMPLETE HEART BLOCK AND BECAME PACEMAKER DEPENDENT. A TEMPORARY PACEMAKER WAS PLACED. THE COMPLETE HEART BLOCK WAS REPORTED AS CAUSAL TO THE IMPLANT PROCEDURE AND VALVE. ALSO DURING THE IMPLANT PROCEDURE, ACUTE ONSET DYSARTHRIA AND EYE MOVEMENT ABNORMALITIES DEVELOPED. THE INITIAL NEUROLOGIC ASSESSMENT SUSPECTED AN EMBOLIC STROKE BUT SUGGESTED FURTHER WORK-UP. THE VALVE WAS IMPLANTED. NEUROLOGY CONTINUED TO ASSESS THE SUBJECT AND SUSPECTED A STROKE. PATIENT WAS UNDER OBSERVATION AND HAD NOT YET BEEN GIVEN TREATMENT FOR STROKE-LIKE SYMPTOMS. THE PHYSICIAN REPORTED THE SUSPECT STROKE WAS CAUSAL RELATED TO THE DELIVERY CATHETER SYSTEM (DCS) AND IMPLANT PROCEDURE. THE DAY FOLLOWING THE VALVE IMPLANT A PERMANENT PACEMAKER WAS IMPLANTED. HOWEVER, IT WAS ALSO REPORTED THAT THE PERMANENT PACEMAKER HAD NOT BEEN IMPLANTED. THE PRIMARY INDICATION FOR THE PERMANENT PACEMAKER WAS A CHRONIC BIFASCICULAR BLOCK. HOSPITALIZATION WAS PROLONGED AS RESULT OF THE EVENT. TWO DAYS LATER THE WOUND NURSE WAS CONSULTED TO ASSESS GROIN SKIN CONCERNS. UPON ASSESSMENT, A ¿LARGE¿ AREA OF ECCHYMOSIS WAS PRESENT ON THE BILATERAL GROIN AND THE LOWER ABDOMEN DUE TO THE RECENT PROCEDURES. HOWEVER, NO SIGNIFICANT OPEN ULCERATIONS WERE NOTED. ALL AREAS WERE TO BE LEFT OPEN AND TOPICAL DRESSINGS WERE NOT REQUIRED. THE PHYSICIAN INDICATED THE ECCHYMOSIS WAS CAUSAL RELATED TO THE PROCEDURE. ULTIMATELY, AN ISCHEMIC STROKE WAS REPORTED. THE PATIENT MOVED TO COMFORT CARE AND DIED 5 DAYS FOLLOWING THE VALVE IMPLANT DUE TO THE STROKE. THE DEATH WAS REPORTED AS A NON-SUDDEN CARDIAC DEATH. ADDITIONALLY IT WAS REPORTED THAT DURING THE IMPLANT PROCEDURE THE VALVE AND DELIVERY CATHETER SYSTEM HAD BEN ADVANCED TO THE ASCENDING AORTA BEGAN BEFORE ACHIEVING A THERAPEUTIC ACTIVATED CLOTTING TIME (ACT) LEVEL. THIS WAS NOTED WHEN TWO READINGS IN THE MID TO LOW 100 SECOND LEVEL WERE OBSERVED. A LEAK IN THE INTRAVENOUS TUBING PREVENTED THE DELIVERY OF ANTICOAGULANT AND OTHER ANESTHETIC DRUGS. AS A RESULT, THE PATIENT EXPERIENCED A FACIAL BURNING SENSATION, DIZZINESS, AND THROAT PAIN, INITIALLY THOUGHT TO BE A CONTRAST REACTION. STEROIDS AND ANTIHISTAMINES WERE ADMINISTERED, BUT IT WAS LATER DISCOVERED THAT NO DRUGS HAD REACHED THE PATIENT DUE TO THE DEFECTIVE TUBING. HEPARIN WAS SUBSEQUENTLY ADMINISTERED VIA A STERILE CENTRAL VENOUS LINE, AND THE PROCEDURE CONTINUED. THE FINAL ACTIVATED CLOTTING TIME MEASUREMENT WAS IN THE LOW 300 SECOND RANGE, WHICH WAS TAKEN FOLLOWING VALVE DEPLOYMENT. MAGNETIC RESONANCE IMAGING LATER CONFIRMED THE PATIENT SUFFERED A STROKE, DESPITE THREE COMPUTED TOMOGRAPHY SCANS SHOWING NO EVIDENCE OF STROKE. DUAL ANTIPLATELET THERAPY (DAPT) WAS STARTED FOR THE STROKE AND SYMPTOMS. AS REPORTED, THE PATIENT EXPERIENCED ONGOING DISABILITY AS A RESULT OF THE STROKE. DUE TO ¿SEVERE¿ DYSARTHRIA AND BILATERAL LOWER FACIAL WEAKNESS DYSPHAGIA, A NASOGASTRIC (NG) TUBE WAS PLACED THE DAY FOLLOWING THE VALVE IMPLANT TO FACILITATE MEDICATION ADMINISTRATION. THE PHYSICIAN INDICATED THE PATIENT¿S MAIN RISK FACTORS FOR STROKE WERE BILATERAL CAROTID ARTERY PLAQUE WITH 50-60% STENOSIS OF THE RIGHT INTERNAL CAROTID ARTERY AND 50% STENOSIS IN THE LEFT INTERNAL CAROTID ARTERY. AN AUTOPSY WAS NOT PERFORMED. THE ISCHEMIC EMBOLIC STROKE AND SUBSEQUENT DEATH WERE REPORTED AS CAUSAL TO THE PROCEDURE AND THE DELIVERY CATHETER SYSTEM (DCS).Event Description: IT WAS REPORTED THAT THE SUBJECT WITH A HISTORY OF FIRST DEGREE ATRIO-VENTRICULAR BLOCK AND BIFASCICULAR BLOCK WITH A RIGHT BUNDLE AND LEFT ANTERIOR FASCICULAR BLOCK HAD DISCUSSION OF A PERMANENT PACEMAKER IMPLANT PRIOR TO THE TRANSCATHETER BIOPROSTHETIC AORTIC VALVE IMPLANT. DURING THE VALVE IMPLANT PROCEDURE, THE SUBJECT DEVELOPED COMPLETE HEART BLOCK AND BECAME PACEMAKER DEPENDENT. A TEMPORARY PACEMAKER WAS PLACED. THE COMPLETE HEART BLOCK WAS REPORTED AS CAUSAL TO THE IMPLANT PROCEDURE AND VALVE. ALSO DURING THE IMPLANT PROCEDURE, ACUTE ONSET DYSARTHRIA AND EYE MOVEMENT ABNORMALITIES DEVELOPED. THE INITIAL NEUROLOGIC ASSESSMENT SUSPECTED AN EMBOLIC STROKE BUT SUGGESTED FURTHER WORK-UP. THE VALVE WAS IMPLANTED. NEUROLOGY CONTINUED TO ASSESS THE SUBJECT AND SUSPECTED A STROKE. PATIENT WAS UNDER OBSERVATION AND HAD NOT YET BEEN GIVEN TREATMENT FOR STROKE-LIKE SYMPTOMS. THE PHYSICIAN REPORTED THE SUSPECT STROKE WAS CAUSAL RELATED TO THE DELIVERY CATHETER SYSTEM (DCS) AND IMPLANT PROCEDURE. THE DAY FOLLOWING THE VALVE IMPLANT A PERMANENT PACEMAKER WAS IMPLANTED. HOWEVER, IT WAS ALSO REPORTED THAT THE PERMANENT PACEMAKER HAD NOT BEEN IMPLANTED. THE PRIMARY INDICATION FOR THE PERMANENT PACEMAKER WAS A CHRONIC BIFASCICULAR BLOCK. HOSPITALIZATION WAS PROLONGED AS RESULT OF THE EVENT. TWO DAYS LATER THE WOUND NURSE WAS CONSULTED TO ASSESS GROIN SKIN CONCERNS. UPON ASSESSMENT, A ¿LARGE¿ AREA OF ECCHYMOSIS WAS PRESENT ON THE BILATERAL GROIN AND THE LOWER ABDOMEN DUE TO THE RECENT PROCEDURES. HOWEVER, NO SIGNIFICANT OPEN ULCERATIONS WERE NOTED. ALL AREAS WERE TO BE LEFT OPEN AND TOPICAL DRESSINGS WERE NOT REQUIRED. THE PHYSICIAN INDICATED THE ECCHYMOSIS WAS CAUSAL RELATED TO THE PROCEDURE. ULTIMATELY, AN ISCHEMIC STROKE WAS REPORTED. THE PATIENT MOVED TO COMFORT CARE AND DIED 5 DAYS FOLLOWING THE VALVE IMPLANT DUE TO THE STROKE. THE DEATH WAS REPORTED AS A NON-SUDDEN CARDIAC DEATH. ADDITIONALLY IT WAS REPORTED THAT DURING THE IMPLANT PROCEDURE THE VALVE AND DELIVERY CATHETER SYSTEM HAD BEN ADVANCED TO THE ASCENDING AORTA BEGAN BEFORE ACHIEVING A THERAPEUTIC ACTIVATED CLOTTING TIME (ACT) LEVEL. THIS WAS NOTED WHEN TWO READINGS IN THE MID TO LOW 100 SECOND LEVEL WERE OBSERVED. A LEAK IN THE INTRAVENOUS TUBING PREVENTED THE DELIVERY OF ANTICOAGULANT AND OTHER ANESTHETIC DRUGS. AS A RESULT, THE PATIENT EXPERIENCED A FACIAL BURNING SENSATION, DIZZINESS, AND THROAT PAIN, INITIALLY THOUGHT TO BE A CONTRAST REACTION. STEROIDS AND ANTIHISTAMINES WERE ADMINISTERED, BUT IT WAS LATER DISCOVERED THAT NO DRUGS HAD REACHED THE PATIENT DUE TO THE DEFECTIVE TUBING. HEPARIN WAS SUBSEQUENTLY ADMINISTERED VIA A STERILE CENTRAL VENOUS LINE, AND THE PROCEDURE CONTINUED. THE FINAL ACTIVATED CLOTTING TIME MEASUREMENT WAS IN THE LOW 300 SECOND RANGE, WHICH WAS TAKEN FOLLOWING VALVE DEPLOYMENT. MAGNETIC RESONANCE IMAGING LATER CONFIRMED THE PATIENT SUFFERED A STROKE, DESPITE THREE COMPUTED TOMOGRAPHY SCANS SHOWING NO EVIDENCE OF STROKE. DUAL ANTIPLATELET THERAPY (DAPT) WAS STARTED FOR THE STROKE AND SYMPTOMS. AS REPORTED, THE PATIENT EXPERIENCED ONGOING DISABILITY AS A RESULT OF THE STROKE. DUE TO ¿SEVERE¿ DYSARTHRIA AND BILATERAL LOWER FACIAL WEAKNESS DYSPHAGIA, A NASOGASTRIC (NG) TUBE WAS PLACED THE DAY FOLLOWING THE VALVE IMPLANT TO FACILITATE MEDICATION ADMINISTRATION. THE PHYSICIAN INDICATED THE PATIENT¿S MAIN RISK FACTORS FOR STROKE WERE BILATERAL CAROTID ARTERY PLAQUE WITH 50-60% STENOSIS OF THE RIGHT INTERNAL CAROTID ARTERY AND 50% STENOSIS IN THE LEFT INTERNAL CAROTID ARTERY. AN AUTOPSY WAS NOT PERFORMED. THE ISCHEMIC EMBOLIC STROKE AND SUBSEQUENT DEATH WERE REPORTED AS CAUSAL TO THE PROCEDURE AND THE DELIVERY CATHETER SYSTEM (DCS).Event Description: ADDITIONAL INFORMATION RECEIVED INDICATED A NEUROLOGIC CONSULT IDENTIFIED A NATIONAL INSTITUTES OF HEALTH STROKE SCALE (NIHSS) OF 4. A SECOND STROKE ALERT ACTIVATED OVERNIGHT WITH AN NIHSS OF 6. THE DAY FOLLOWING THE VALVE IMPLANT FURTHER PATIENT EVALUATION NOTED THE NIHSS HAD WORSENED TO 14. THE NEUROLOGIC EVALUATION WAS POSITIVE FOR LIMB ATAXIA, PARTIAL HEMIANOPIA AND PARTIAL GAZE PALSY. ADDITIONAL HEAD COMPUTED TOMOGRAPHIES (CT) GAVE CONCERN FOR WORSENING SYMPTOMS. ALSO ON THE DAY FOLLOWING THE VALVE IMPLANT A BRAIN MAGNETIC RESONANCE IMAGING (MRI) DEMONSTRATED ACUTE INFARCTS AT BILATERAL CENTRAL PONS. DUAL ANTI-PLATELET THERAPY (DAPT) WAS STARTED. THE PHYSICIAN FURTHER INDICATED THE ISCHEMIA STROKE WAS CAUSAL RELATED TO THE TRANSCATHETER AORTIC VALVE AS WELL AS THE DELIVERY CATHETER SYSTEM (DCS) AND PROCEDURE. THREE DAYS FOLLOWING THE VALVE IMPLANT THE PATIENT WAS ADMITTED TO THE INTENSIVE CARE UNIT (ICU) AFTER DEVELOPING ACUTE HYPOXIA. HYPOXIC RESPIRATORY FAILURE DEVELOPED, AND THE RESPIRATORY STATUS CONTINUED TO DECOMPENSATE. THIS CONDITION REQUIRED 15 LITERS OF OXYGEN VIA A NON-REBREATHER MASK. THE RESPIRATORY FAILURE AND COMPLETE HEART BLOCK HAD ALSO CONTRIBUTED TO THE PROLONGED HOSPITALIZATION. AS REPORTED, THE RESPIRATORY FAILURE WAS A RESULT OF THE STROKE. THE PATIENT WAS TRANSITIONED TO THE INPATIENT PALLIATIVE CARE UNIT. THE RESPIRATORY FAILURE WAS REPORTED AS POSSIBLY RELATED TO THE VALVE IMPLANT PROCEDURE. Manufacturer Narrative: UPDATED DATA: B5, D4, H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA RELATED TO THE SYSTEM REPORTED DELIVERY CATHETER SYSTEM (DCS): H6. EVAL CODE RESULTS AND CONCLUSION WERE UPDATED SYSTEM REPORTED DEVICE CONCLUSION: THE SUBJECT DCS WAS DISCARDED BY THE CUSTOMER AND AS SUCH NO ANALYSIS COULD BE PERFORMED. PROCEDURAL IMAGES WERE NOT PROVIDED FOR REVIEW. THE FOLLOWING TECHNICAL ASSESSMENTS WERE NOT REQUIRED AS THE PRODUCT EVENT DOES NOT MEET THE CRITERIA TO INITIATE; DHR REVIEW, MANUFACTURING ASSESSMENT, SUPPLIER MANUFACTURING ASSESSMENT, DESIGN ASSESSMENT AND RISK DOCUMENTATION ASSESSMENT. STROKE IS A KNOWN POTENTIAL ADVERSE EFFECT PER THE DEVICE INSTRUCTIONS FOR USE (IFU). ISCHEMIC STROKES HAVE MANY ETIOLOGIES, INCLUDING EMBOLIZATION OF CALCIFIC DEBRIS, AND ARE HIGHLY DEPENDENT ON PATIENT MEDICAL HISTORY AND PROCEDURAL FACTORS. DEATH IS A KNOWN POTENTIAL ADVERSE EFFECT PER THE DEVICE IFU. A MEDICAL SAFETY ASSESSMENT WAS PERFORMED AND BASED ON THE UPDATED INFORMATION PRE-EXISTING CONDITIONS SUCH AS BILATERAL CAROTID ARTERY PLAQUE WITH 50-60% STENOSIS OF THE RIGHT INTERNAL CAROTID ARTERY AND <(><<)>50% STENOSIS IN THE LEFT INTERNAL CAROTID ARTERY LIKELY CONTRIBUTED TO THE STROKE AND SUBSEQUENT DEATH. IT IS ALSO LIKELY THE HOLE IN THE INTRAVENOUS TUBING CONTRIBUTED TO THE STROKE AND DEATH AS THE PATIENT WAS NOT ADEQUATELY ANTICOAGULATED DURING A PORTION OF THE PROCEDURE. IT IS STILL LIKELY THE STROKE AND SUBSEQUENT DEATH WERE RELATED TO THE PROCEDURE AND POSSIBLY THE DELIVERY CATHETER SYSTEM AS THE DELIVERY SYSTEM COULD HAVE INTERACTED WITH PLAQUE DURING ADVANCEMENT. THERE WAS NO INFORMATION TO SUGGEST A DEVICE QUALITY DEFICIENCY OR A FAILURE TO MEET MANUFACTURING SPECIFICATIONS. THERE WAS NO NEW OR UNEXPECTED DEVICE OR THERAPY ISSUE IDENTIFIED. THE EVENT IS FORESEEN, WITHIN EXPECTED SEVERITY, AND DOCUMENTED IN THE RISK MANAGEMENT FILE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: H6. PRODUCT ANALYSIS: THE SUSPECT DEVICE REMAINED IN THE PATIENT; THEREFORE, ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED. CONCLUSION: THERE WAS NO INFORMATION TO INDICATE THE EVENT WAS POTENTIALLY RELATED TO A MANUFACTURING ISSUE. A DEVICE HISTORY RECORD REVIEW WAS NOT REQUIRED. HOWEVER, MANUFACTURING CONTROLS ARE IN PLACE TO ENSURE THAT THE DEVICE MET SPECIFICATION REQUIREMENTS PRIOR TO RELEASE FROM THE MANUFACTURING FACILITY. THE REPORTED EVENT IS UNCONFIRMED AS THE DEVICE MET SPECIFICATIONS FOR THE REPORTED ISSUE. THE COMPLAINT INVESTIGATION DETERMINED THIS EVENT IS FORESEEN IN RISK MANAGEMENT AND IS INCLUDED IN MONITORING/TRENDING. THE DEVICE FAILURE MODE WAS UNIDENTIFIED; IT IS NOT POSSIBLE TO DETERMINE A CAUSE OF THE EVENT FROM THE INFORMATION PROVIDED. THIS EVENT WILL BE INCLUDED IN MONITORING AND TRENDING. MEDICAL SAFETY ASSESSMENT: A MEDICAL SAFETY ASSESSMENT WAS PERFORMED FOR THIS EVENT. THE ADVERSE EVENTS/HARMS, PATIENT OUTCOME, AND TREATMENTS FOR THIS EVENT INCLUDE COMPLETE HEART BLOCK, MOST LIKELY TREATED WITH A PERMANENT PACEMAKER; HEMATOMA; DYSARTHRIA; EYE MOVEMENT ABNORMALITIES; ISCHEMIC STROKE; AND NON-SUDDEN CARDIAC DEATH. POTENTIAL CONTRIBUTING FACTORS INCLUDED PRE-EXISTING CONDUCTION DISTURBANCES. THE RELATEDNESS ASSESSMENT AND RATIONALE NOTED, BASED ON MEDICAL EXPERTISE, THE PRE-EXISTING CONDUCTION DISTURBANCES LIKELY CONTRIBUTED TO THE NEW CONDUCTION DISTURBANCES. THEY WERE ALSO LIKELY RELATED TO THE VALVE AND INTERACTION WITH THE NATIVE CONDUCTION PATHWAY. BASED ON THE LOCATION, THE HEMATOMAS WERE LIKELY RELATED TO THE PROCEDURE, INCLUDING THE ANTICOAGULATION MEDICATION USED DURING THE IMPLANT PROCEDURE AND DEVICES INSERTED AT THE GROIN, INCLUDING THE DELIVERY CATHETER SYSTEM. BASED ON THE TIMING OF ONSET, THE DYSARTHRIA AND EYE MOVEMENT ABNORMALITIES AND STROKE WERE LIKELY RELATED TO THE PROCEDURE AND POSSIBLY THE DELIVERY CATHETER SYSTEM. THE NON-SUDDEN CARDIAC DEATH WAS LIKELY RELATED TO THE PROCEDURE AS IT OCCURRED WITHIN DAYS OF THE VALVE IMPLANT PROCEDURE. PRODUCT LABELING WAS REVIEWED AND THE PRIMARY HARMS APPEAR TO BE DOCUMENTED IN THE PRODUCT LABELING. A RISK MANAGEMENT FILE WAS REVIEWED AND CONFIRMED THE PRIMARY HARMS AND SEVERITY APPEAR TO BE SUFFICIENTLY DOCUMENTED. BASED ON THE MEDICAL SAFETY ASSESSMENT OF THIS EVENT NO FURTHER ACTIONS ARE RECOMMENDED THE DEVICE INSTRUCTIONS FOR USE (IFU) LISTS CONDUCTION DISTURBANCES AS A POTENTIAL RISK ASSOCIATED WITH THE TREATMENT PROCEDURES. THE DEVICE IFU WAS DEVELOPED FROM THE PRODUCT RISK DOCUMENTATION AS IS THE PRODUCT LABELLED ADVERSE EVENTS DOCUMENT. THERE WAS NO IDENTIFIED INADEQUACY WITH THE DEVICE IFU IN RELATION TO THIS EVENT. NEITHER THE PRODUCT DESIGN NOR MANUFACTURING PROCESSES OF THE MEDTRONIC DEVICE HAVE BEEN IDENTIFIED AS THE CAUSE OF THE EVENT. THEREFORE, THE CAUSE OF THE EVENT IS UNDETERMINED. CONDUCTION DISTURBANCES, SUCH AS COMPLETE HEART BLOCK, ARE KNOWN POTENTIAL ADVERSE EFFECTS PER THE DEVICE IFU AND CAN BE RESOLVED WITH THE IMPLANT OF A PERMANENT PACEMAKER WITH THE RISK-BENEFIT RATIO IN FAVOR OF THE TRANSCATHETER AORTIC VALVE REPLACEMENT. FACTORS THAT MAY IMPACT THE DEVELOPMENT OF CONDUCTION DISTURBANCES INCLUDE DEPTH OF IMPLANT, BASELINE CONDUCTION DEFECTS, AND ANATOMICAL CONSIDERATIONS. BASED ON THE MEDICAL SAFETY ASSESSMENT, PRE-EXISTING CONDUCTION DISTURBANCES LIKELY CONTRIBUTED TO THE NEW CONDUCTION DISTURBANCES. ADDITIONALLY, THE CONDUCTION DISTURBANCE WAS ALSO LIKELY RELATED TO THE VALVE AND INTERACTION WITH THE NATIVE CONDUCTION PATHWAY; HOWEVER, A CONCLUSIVE CAUSE COULD NOT BE DETERMINED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA; A1, A2, B5, D1, D4, H4, H6 A DUPLICATE EVENT REPORT, # 2025587-2025-01792, WAS IDENTIFIED. GOING FORWARD NEW REPORTABLE INFORMATION WILL CONTINUED TO BE SUBMITTED WITHIN THIS CURRENT REGULATORY REPORT # 2025587-2025-01835. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SECTION D REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: BRAND NAME EVOLUT FX DCS; PRODUCT ID D-EVOLUTFX-2329 (LOT: 0012508008); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE NA; EXPLANT DATE NA; PRODUCT ID L-EVOLUTFX-2329; PRODUCT LOT/SERIAL NUMBER (B)(6); PRODUCT TYPE: COMPRESSION LOADING SYSTEM; IMPLANT DATE NA; EXPLANT DATE NA. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21580417&pc=PNJ | 2017865-2025-17247 | 2024-04-27 04:00 | Death | ABBOTT | 2025-03-11 04:00 | PNJ | AVEIR¿ LEADLESS PACEMAKER, RIGHT VENTRICLE | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P150035 |  | 1 | Event Description: A PATIENT WAS IMPLANTED WITH A LEADLESS PACEMAKER SYSTEM. ONE DAY POST-OPERATION, THE PATIENT DIED DUE TO HYPERCAPNIC COMA WITH BILATERAL PLEURAL EFFUSIONS AND CARDIAC DECOMPENSATION. THE PHYSICIAN DOES NOT BELIEVE THE LEADLESS PACEMAKER CAUSE OR CONTRIBUTED TO THE DEATH OF THE PATIENT. THE LEADLESS PACEMAKER WAS IMPLANTED WITHOUT COMPLICATION, HOWEVER, IT WAS ALLEGED THAT THE DEATH WAS DUE TO THE ANESTHESIA. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21557199&pc=PRC | 3006513822-2025-00013 | 2024-09-16 04:00 | Death | LUTONIX, INC | 2025-03-10 04:00 | PRC | LUTONIX 035AV DRUG COATED PTA DILATION CATHETER | Adverse Event Without Identified Device or Use Problem; Insufficient Information | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P170003 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESULTS OF A CLINICAL TRIAL THAT APPROXIMATELY ONE MONTH AND TWENTY-TWO DAYS POST AN INDEX PROCEDURE COMPLETION USING A DRUG-COATED BALLOON CATHETER ON THE LEFT UPPER ARM BASILIC VEIN STENOSIS, THE SUBJECT WAS EXPIRED WHILE ATTEMPTING SIX MONTHS FOLLOW-UP OFFICE VISIT. THE CAUSE OF DEATH AND THE CAUSAL RELATIONSHIP TO DEVICE AND PROCEDURE DETAILS WERE NOT PROVIDED.Event Description: IT WAS REPORTED THROUGH THE RESULTS OF A CLINICAL TRIAL THAT APPROXIMATELY ONE MONTH AND TWENTY TWO DAYS POST AN INDEX PROCEDURE COMPLETION USING A DRUG-COATED BALLOON CATHETER ON THE LEFT UPPER ARM BASILIC VEIN STENOSIS, THE SUBJECT WAS EXPIRED WHILE ATTEMPTING SIX MONTH FOLLOW-UP OFFICE VISIT. REPORTEDLY, THE ADVERSE EVENT WAS NOT RELATED TO STUDY DEVICE, PROCEDURE AND AV ACCESS CIRCUIT. HOWEVER, THE CAUSE OF DEATH WAS NOT PROVIDED. Manufacturer Narrative: H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. Manufacturer Narrative: H10: THE FILE WAS REASSESSED FOR REPORTABILITY AND DETERMINED TO BE NO LONGER REPORTABLE. SINCE AN INITIAL MDR WAS SUBMITTED, THEREFORE, THE FILE WILL REMAIN ASSESSED AS A DEATH. H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. G10: G3. G11: B5, H6(PATIENT, DEVICE). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21557292&pc=NAY | 2955842-2025-05793 | 2025-01-25 05:00 | Death | INTUITIVE SURGICAL, INC | 2025-03-10 04:00 | NAY | DAVINCI X | Adverse Event Without Identified Device or Use Problem | Pulmonary Embolism; Respiratory Arrest | K171294 |  | 1 | Event Description: IT WAS REPORTED THAT AFTER A DA VINCI-ASSISTED BYPASS GASTROPLASTY PROCEDURE, THE PATIENT SUFFERED A CARDIAC ARREST AND EXPIRED. THE DA VINCI PROCEDURE WAS COMPLETED WITHOUT ANY INTRA-OPERATIVE COMPLICATIONS. THE PATIENT EXPERIENCED A PULMONARY THROMBOEMBOLISM, DEVELOPED DYSPNEA, AND WAS ADMITTED TO THE INTENSIVE CARE UNIT (ICU) 72 HOURS POST-OPERATION. THE NEXT DAY, THE PATIENT WENT INTO CARDIAC ARREST AGAIN. RESUSCITATION EFFORTS WERE PERFORMED, BUT THE PATIENT COULD NOT BE REVIVED AND EXPIRED. THE HOSPITAL DOCUMENTED THE OFFICIAL CAUSE OF DEATH WAS PULMONARY THROMBOEMBOLISM. IT WAS REPORTED THAT NO DA VINCI PRODUCT CAUSED OR CONTRIBUTED TO THE PATIENT EVENTS. THE CUSTOMER REPORTED THAT THE DEATH WAS ATTRIBUTED TO PULMONARY EMBOLISM, PNEUMONIA, OBESITY, AND SEQUELAE FROM A HISTORY OF SEVERE COVID-19 WITH PULMONARY INVOLVEMENT.Event Description: REFER TO H11 FOR FOLLOW-UP INFORMATION.Event Description: REFER TO H11 FOR FOLLOW-UP INFORMATION. Manufacturer Narrative: A REVIEW OF THE DA VINCI SYSTEM LOGS FOUND THAT NO ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM ALSO FOUND THAT NO ERRORS OCCURRED DURING THE PROCEDURES. Manufacturer Narrative: THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THE PROCEDURE: 30 DEGREE ENDOSCOPE PLUS, LARGE NEEDLE DRIVER, TWO CADIERE FORCEPS, SYNCHROSEAL. A SITE HISTORY REVIEW FOUND NO COMPLAINTS WERE MADE FOR THE INSTRUMENTS USED DURING THIS PROCEDURE. A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED FOR THE DEVICE(S) USED DURING THE PROCEDURE AND NO NON-CONFORMANCES WERE IDENTIFIED TO BE RELATED TO THIS EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT DIED OF A PULMONARY EMBOLUS A FEW DAYS AFTER A ROBOTIC BARIATRIC PROCEDURE THAT WAS UNEVENTFUL WITH NO INTRAOPERATIVE COMPLICATIONS OR ISSUES REPORTED. THE PATIENT HAD SIGNIFICANT COMORBIDITIES WHICH MAY HAVE CONTRIBUTED TO THIS EVENT. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, NO INTUITIVE SURGICAL PRODUCT OR INSTRUMENT CONTRIBUTED TO THIS EVENT. Manufacturer Narrative: THE FOLLOWING FIELDS WERE UPDATED: B2,B3, H11. ALL MULTI-USE INSTRUMENTS USED IN THE CASE HAVE BEEN USED IN SUBSEQUENT PROCEDURES WITH THE EXCEPTION OF THE SINGLE-USE INSTRUMENT. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: 30 DEGREE ENDOSCOPE PLUS, LARGE NEEDLE DRIVER, TWO CADIERE FORCEPS INSTRUMENTS, AND A SYNCHROSEAL.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21558018&pc=DSP | 2248146-2025-0000129 | 2025-01-26 05:00 | Death | DATASCOPE CORP. | 2025-03-10 04:00 | DSP | LINEAR 7.5FR. 40CC IAB WITH ACCESSORIES, US ONLY | Fluid/Blood Leak; Material Rupture | Insufficient Information | K041281 |  | 1 | Event Description: IT WAS REPORTED THAT THE INTRA-AORTIC BALLOON (IAB) EXPERIENCED A CATHETER RUPTURE. UNFORTUNATELY, THE PATIENT DID NOT SURVIVE.Event Description: IT WAS REPORTED THAT THE INTRA-AORTIC BALLOON (IAB) CATHETER EXPERIENCED A RUPTURE. UNFORTUNATELY, THE PATIENT DID NOT SURVIVE. Manufacturer Narrative: UPDATED FIELDS: B4, B5, D8, D9, G3, G6, H2, H3, H6 (MEDICAL DEVICE ¿ PROBLEM CODE, TYPE OF INVESTIGATION, INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS), H8, H11 CORRECTED FIELDS: D7A. THE IAB WAS RETURNED WITH THE MEMBRANE COMPLETELY UNFOLDED AND BLOOD ON THE INTERIOR AND EXTERIOR OF THE CATHETER. THE SHEATH HUB WAS OVER THE MEMBRANE. THE SHEATH TUBING APPEARS TO BE BROKEN OFF FROM THE SHEATH AND NOT RETURNED. THE EXTENDER TUBING WAS ALSO RETURNED. AN UNDERWATER LEAK TEST OF THE BALLOON, CATHETER, Y-FITTING, EXTRACORPOREAL AND EXTENDER TUBING WAS PERFORMED AND A LEAK WAS DETECTED ON THE MEMBRANE APPROXIMATELY 1.3CM FROM THE REAR SEAL MEASURING 0.013CM IN LENGTH. THE REPORTED PROBLEM WAS MOST LIKELY TRIGGERED BY A LEAK WHICH WAS FOUND ON THE MEMBRANE. UNDER MAGNIFICATION, A WHITISH PATCH WAS OBSERVED AROUND THE LEAK. THIS WHITISH PATCH IS THE TYPICAL APPEARANCE OF AN ABRASION MARK WHICH IS CAUSED BY CALCIFIED PLAQUE IN THE AORTA. A LOT HISTORY RECORD REVIEW WAS COMPLETED FOR THE REPORTED PRODUCT. NO NONCONFORMANCES WERE FOUND THAT ARE CONSIDERED TO BE RELATED TO THE EVENT. COMPLAINT RECORD ID #(B)(4). Manufacturer Narrative: DUE TO CHARACTER LIMIT IN BLOCK E1 EVENT SITE NAME IS (B)(6). THE DEVICE IS NOT AVAILABLE TO BE RETURNED TO THE MANUFACTURER SO WE ARE UNABLE TO COMPLETE AN EVALUATION. IF PROVIDED WE WILL SEND A SUPPLEMENTAL REPORT WITH OUR ADDITIONAL FINDINGS. COMPLAINT RECORD ID #(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21558319&pc=BZD | 2518422-2025-102794 | 2025-02-12 05:00 | Death | RESPIRONICS, INC. | 2025-03-10 04:00 | BZD | PHILIPS CPAP DEVICE | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: THE MANUFACTURER HAS BEEN NOTIFIED OF THE PASSING OF A PATIENT WHO WAS USING A PHILIPS CPAP DEVICE. ACCORDING TO THE PATIENT'S WIFE, HER HUSBAND PASSED AWAY IN 2023, AND SHE IS NOW IN POSSESSION OF ONE OF HIS CPAP DEVICES. HOWEVER, SHE IS UNSURE OF WHAT STEPS TO TAKE WITH IT. THE WIFE MENTIONED RECEIVING A LETTER REGARDING THE DEVICE, BUT SHE IS UNABLE TO READ IT DUE TO RECENT EYE SURGERY. SHE ALSO INDICATED THAT SHE WAS UNABLE TO VERIFY THE DEVICE'S SERIAL NUMBER. NO ADDITIONAL DETAILS WERE PROVIDED, AND THERE ARE NO ALLEGATIONS SUGGESTING THAT THE DEVICE CONTRIBUTED TO OR CAUSED THE PATIENT'S DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21558402&pc=OZD | 1220648-2025-26587 | 2025-02-08 05:00 | Death | ABIOMED, INC. | 2025-03-10 04:00 | OZD | IMPELLA CP | Increase in Pressure | Insufficient Information | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS) AND DEVELOPED HIGH PURGE PRESSURE AND WOULD EXPIRE THE SAME DAY NOTED TO BE FROM MULTI-ORGAN FAILURE. THE IMPELLA WAS IMPLANTED VIA A FEMORAL ARTERY PRIOR TO PERCUTANEOUS CORONARY INTERVENTION. VESSEL TREATMENT INFORMATION WAS NOTED AS UNKNOWN. THE PATIENT WAS EVENTUALLY SENT TO THE UNIT ON IMPELLA MCS. IT WAS NOTED THERE WERE NO OTHER THERAPIES PRIOR TO IMPELLA MCS. APPROXIMATELY 11 DAYS AFTER THE START OF THE SUPPORT, THE IMPELLA CP PUMP EXPERIENCED A RISE IN PURGE PRESSURE COINCIDING WITH A DECREASE IN PURGE FLOW. TROUBLESHOOTING WAS COMPLETED WITHOUT SUCCESS AS THE PURGE CASSETTE WAS REPLACED, BUT THERE WAS NO IMPROVEMENT WITH THE HIGH PRESSURES. NO FURTHER DETAILS WERE PROVIDED REGARDING THE PURGE PRESSURE ISSUE WERE MADE AVAILABLE. HOWEVER, IT WAS NOTED THE PATIENT WOULD EXPIRE THIS SAME DAY FROM MULTI-ORGAN FAILURE AS THE CAUSE OF DEATH. NO DETAILS SURROUNDING AND CAUSE OF THE MULTI-ORGAN FAILURE WERE PROVIDED OTHER THAN THIS CLINICAL ADVERSE EVENT IS UNRELATED TO THE IMPELLA DEVICE. ALTHOUGH THE REPORTER INDICATES THE CAUSE OF DEATH AS MULTIPLE ORGAN FAILURE AND UNRELATED TO THE IMPELLA DEVICE USED, IT IS UNKNOWN IF THE UNRESOLVED HIGH PURGE PRESSURE DEVICE MALFUNCTION MAY HAVE IMPACTED THE PATIENT'S SUPPORT. THE HIGH PURGE PRESSURE ISSUE WAS UNRESOLVED WITH THE CASSETTE CHANGE AND OTHER DEVICE FUNCTION FACTORS CANNOT BE RULED OUT WITH ONLY "NOT RELATED." THUS, THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN OVERALL OUTCOME. Manufacturer Narrative: THE IMPELLA CP PUMP AND PURGE CASSETTE WERE RETURNED, AND AN EVALUATION IS UNDERWAY. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. Manufacturer Narrative: THE INVESTIGATION OF HIGH PURGE PRESSURE HAS BEEN COMPLETED. IMPELLA CP WAS RETURNED FOR EVALUATION. UPON VISUAL INSPECTION, BIOMATERIAL WAS PRESENT IN PURGE GAP BUT NOT FULLY COVERING ALL OF THE PURGE GAP. THE RETURNED PUMP AND RETURNED PURGE CASSETTE WERE CONNECTED TO THE CONSOLE AND PURGED IN THE LAB FOR APPROXIMATELY 30 MINUTES AND NO HIGH PURGE PRESSURE ALARMS REPRODUCED. THE PUMP WAS TURNED UP TO P-9 AND AGAIN NO HIGH PURGE PRESSURE ALARMS WERE TRIGGERED. DATA LOGS WERE REVIEWED AND LONG-TERM LOG OF FULL CASE SHOWED TWO INSTANCES OF HIGH PURGE PRESSURE OCCURRING DURING THE CASE. NO ¿HIGH PURGE PRESSURE¿ ALARMS WERE TRIGGERED DURING EITHER INSTANCE OF PURGE PRESSURE INCREASES. BOTH INSTANCES OF INCREASES IN PURGE PRESSURE OCCURRED OVER A PERIOD OF APPROXIMATELY 12 HOURS FOR EACH INSTANCE AND NO MOTOR CURRENT SPIKES WERE PRESENT PRIOR TO PURGE PRESSURE RISE. FIRST INSTANCE OF HIGH PURGE PRESSURE BEGINS TO DECREASE PRIOR TO A PURGE CASSETTE CHANGE. CLINICAL DETAILS NOTED THAT HIGH PURGE PRESSURE ALARMS OCCURRED ON DAY 12 OF SUPPORT. THE CLINICAL TEAM ATTEMPTED TO EXCHANGE THE PURGE CASSETTE BUT DID NOT RESOLVE THE ISSUE. THE PURGE COMPOSITION AT TIME OF ALARMS WAS NOT NOTED. THE ROOT CAUSE OF THE HIGH PURGE PRESSURE COULD NOT BE DEFINITIVELY DETERMINED AS HIGH PURGE PRESSURE WAS UNABLE TO BE REPRODUCED IN THE LAB. B.2 REQUIRED INTERVENTION WAS REPORTED INCORRECTLY ON INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26587.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21558770&pc=NVZ | 3008973940-2025-01780 |  | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-03-10 04:00 | NVZ | SENSIA IPG | Premature Discharge of Battery | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P980035 |  | 1 | Event Description: IT WAS REPORTED THAT THE IMPLANTABLE PULSE GENERATOR (IPG) EXHIBITED PREMATURE PREMATURE BATTERY DEPLETION AND THE LEAD EXHIBITED FRACTURE. THE PATIENT IS NOW DECEASED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE DEVICE WAS RETURNED AND ANALYZED. ANALYSIS OF THE DEVICE REVEALED NORMAL BATTERY DEPLETION. THE BATTERY INDICATOR SIGNIFYING THAT IT IS TIME FOR DEVICE REPLACEMENT OCCURRED AFTER EXPLANT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21558801&pc=DTB | 2649622-2025-06721 |  | Death | MPRI | 2025-03-10 04:00 | DTB | CAPSURE EPICARDIAL LEAD | Fracture | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P950024 |  | 1 | Event Description: IT WAS REPORTED THAT THE IMPLANTABLE PULSE GENERATOR (IPG) EXHIBITED PREMATURE PREMATURE BATTERY DEPLETION AND THE LEAD EXHIBITED FRACTURE. THE PATIENT IS NOW DECEASED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE PROXIMAL PORTION OF THE LEAD WAS RETURNED, ANALYZED, AND NO ANOMALIES WERE FOUND. VISUAL ANALYSIS OF THE LEAD INDICATED APPARENT EXPLANT DAMAGE. THE ANALYST NOTED A PROXIMAL SEGMENT OF THE LEAD 5.4 CM LONG WAS RETURNED, MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTED: H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21558973&pc=OZD | 1220648-2025-26588 | 2025-02-12 05:00 | Death | ABIOMED, INC. | 2025-03-10 04:00 | OZD | IMPELLA CP | Restricted Flow rate | Perforation; Cardiac Tamponade | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT (UNKNOWN RACE AND ETHNICITY) UNDERGOING A HIGH-RISK PERCUTANEOUS CORONARY INTERVENTION (PCI) AND IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS) EXPERIENCED LEFT VENTRICLE PERFORATION WITH A DEMISE OUTCOME. THE PATIENT PRESENTED TO THE HOSPITAL WITH CHEST PAIN. A CATHETERIZATION REVEALED MULTI-VESSEL DISEASE INCLUDING A HEAVILY CALCIFIED LEFT MAIN AND AN OCCLUSION OF THE RIGHT CORONARY ARTERY. CARDIOTHORACIC SURGERY WAS CONSULTED. AFTER FURTHER WORK UP, PATIENT WAS TURNED DOWN FOR SURGERY DUE TO HAVING A PORCELAIN AORTA. THE PLAN WAS NOTED FOR CARDIOLOGY TO PERFORM PROTECTIVE PCI WITH IMPELLA SUPPORT TO ROTABLATION TO THE LEFT MAIN/LEFT ANTERIOR DESCENDING ARTERIES AND STENT THE RIGHT CORONARY ARTERY. THE IMPELLA CP DEVICE WAS INSERTED RIGHT FEMORAL ARTERY, UTILIZING 14FR. PEEL-AWAY SHEATH. JR4 AND 0.035 J GUIDEWIRE UTILIZED FOR VENTRICULAR ACCESS. THE GUIDEWIRE EXCHANGED FOR ABIOMED 0.018, OF NOTE 0.018 WIRE WAS CURLED PRIOR TO INSERTION. IMPELLA PRIMED AND BACKLOADED ONTO 0.018 AND INSERTED INTO LEFT VENTRICLE. THE GUIDEWIRE REMOVED AND THE IMPELLA CP STARTED IN AUTO MODE AND ENCOUNTERED REDUCED FLOWS. PATIENT BECAME MORE HEMODYNAMICALLY UNSTABLE, CORRELATING WITH POOR PUMP PERFORMANCE. PATIENT REQUIRED CARDIOPULMONARY RESUSCITATION, AND AN EMERGENT ECHO WAS COMPLETED REVEALING A LARGE CARDIAC TAMPONADE. EMERGENT PERICARDIOCENTESIS AND CARDIOTHORACIC SURGERY WAS SUMMONED TO THE BEDSIDE. PATIENT CONTINUED TO REQUIRE RESUSCITATION INCLUDING TWO UNITS OF PACKED RED BLOOD CELLS. THE DECISION MADE TO EMERGENTLY TRANSPORT THE PATIENT TO THE OPERATING ROOM FOR PERICARDIAL WINDOW. IN OPERATING ROOM, THE PATIENT HAD MEDIAN STERNOTOMY, CARDIAC TAMPONADE WAS RELIEVED, 500 CC OF BLOOD EVACUATED, AND THE PATIENT WAS FOUND TO HAVE LEFT VENTRICLE PERFORATION. AT THIS POINT IMPELLA WAS WEANED TO PERFORMANCE LEVEL P-2 AND WITHDRAWN INTO THE DESCENDING AORTA TO ALLOW FOR LEFT VENTRICLE REPAIR. THE DECISION WAS MADE TO REESTABLISH IMPELLA MCS WITH A NEW IMPELLA CP. THE FIRST IMPELLA CP WAS REMOVED THROUGH THE PEEL-AWAY SHEATH, AND NEW IMPELLA CP WAS INSERTED WITHOUT ISSUE. THE REPLACEMENT IMPELLA CP WAS STARTED IN AUTO MODE. HOWEVER, THE PUMP AND PATIENT WERE DEPENDENT ON CARDIAC MASSAGE. THE PATIENT HAD NO SIGNS OF NATIVE ACTIVITY; THEREFORE, THE DECISION WAS MADE BY CARDIOTHORACIC SURGERY AND CARDIOLOGY THAT PATIENT WAS DECLARED EXPIRED. FOLLOWING, THE PHYSICIAN COMMENTED A SECOND IMPELLA CP WAS UTILIZED DUE TO CONCERNS THAT THE LEFT VENTRICLE PERFORATION WAS CAUSED BY THE INITIAL IMPELLA CP. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: PRECAUTION IMPELLA CP WITH SMARTASSIST SYSTEM. SECTION: POTENTIAL ADVERSE EVENTS: ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS & CAUTIONS: WARNINGS, SECTION: PRE-SUPPORT EVALUATION, SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER, SECTION: DIRECT AORTIC INSERTION, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION. ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE. SECTION: WARNINGS & CAUTIONS: WARNINGS, SECTION: PRE-SUPPORT EVALUATION, SECTION: DIRECT AORTIC INSERTION, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: ¿TO REDUCE THE RISK OF CARDIAC INJURY (INCLUDING VENTRICULAR PERFORATION), PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY IN THE SETTINGS OF ACUTE INFARCTION WITH TISSUE NECROSIS.¿ ¿TO REDUCE THE RISK OF VASCULAR INJURY, PHYSICIANS SHOULD EXERCISE CAUTION WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX PERIPHERAL VASCULAR ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: UNREPAIRED ABDOMINAL AORTIC ANEURYSM, SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM, DISSECTION OF THE ASCENDING/ TRANSVERSE/DESCENDING AORTA, CHRONIC ANATOMICAL CHANGES IN THE RELATIONSHIP OF THE AORTA/AORTIC VALVE/VENTRICULAR ALIGNMENT, SIGNIFICANT MOBILE ATHEROMATOUS DISEASE IN THE THORACIC OR ABDOMINAL AORTA OR PERIPHERAL VESSELS.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING A RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY IN THE LEFT VENTRICLE AFTER CPR WITH ECHOCARDIOGRAPHY GUIDANCE. SECTION: WARNINGS & CAUTIONS: CAUTIONS: ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ Manufacturer Narrative: THE INVESTIGATION FOR THE VENTRICLE PERFORATION AND LOW PUMP FLOW HAS BEEN COMPLETED. THE PRODUCT WAS RETURNED AND EVALUATED. EVIDENCE OF CANNULA KINK WAS FOUND. THERE WAS NO BIOMATERIAL FOUND NEAR THE IMPELLER OR IN THE CANNULA. THE PUMP WAS RUN IN A BUCKET AT P-6, P-9, AND P-2. LOW FLOW WAS NOT REPRODUCED. IN THE BEGINNING OF THE CASE, REDUCED FLOWS WERE REPORTED WITH THE PATIENT BECOMING MORE HEMODYNAMICALLY UNSTABLE. DATA LOG REVIEW CONFIRMS REDUCED FLOW IN THE FIRST 45 MINUTES OF THE CASE. DURING THIS TIME, THERE ARE SEVERAL SUCTION, IMPELLA POSITION UNKNOWN, AND IMPELLA IN VENTRICLE ALARMS. THE PUMP IS THEN TURNED UP TO P-6 FROM P-2 AND EVENTUALLY REACHES ADEQUATE FLOWS (~2.7 L/MIN). THE PUMP IS THEN TURNED DOWN TO P-2 AND THE CALCULATED PUMP FLOW STAYS HIGH AROUND 2.7 L/MIN UNTIL THE END OF IMPELLA USE. THE ROOT CAUSE OF THE VENTRICLE PERFORATION WAS NOT DETERMINED SINCE INSUFFICIENT CLINICAL DETAILS WERE PROVIDED RELATED TO TIMING/CAUSE OF THE PERFORATION. THE ROOT CAUSE OF THE LOW PUMP FLOW WAS MOST LIKELY A CANNULA KINK SINCE THE RETURNED PRODUCT SHOWED EVIDENCE OF A KINK AND DATA LOGS CONFIRM POSITIONING ISSUES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21559148&pc=NGV | 2124215-2025-14108 | 2025-02-18 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-03-10 04:00 | NGV | WATCHMAN FLX PRO | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension | P130013 |  | 1 | Event Description: IT WAS REPORTED THAT DEATH OCCURRED. IT WAS NOTED THE PATIENT HAD A PLEURAL EFFUSION, AND RENAL FAILURE WITH A CREATININE LEVEL OF ABOVE 4.0 PRIOR TO THE PROCEDURE AND OVERALL CONDITION WAS NOT OPTIMAL PER THE PHYSICIAN. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED. SINCE THE START OF THE INDEX PROCEDURE, THE PATIENT'S HEART RATE WAS NOTED TO BE HIGH AT 120BPM, OCCASIONALLY RISING TO 150BPM. AS THE HEART RATE INCREASED, THE BLOOD PRESSURE DECREASED, BUT NO PERICARDIAL FLUID ACCUMULATION WAS OBSERVED SO THE CONDITION WAS TREATED WITH VASOPRESSORS, FLUID LOADING, AND AN ANTIARRHYTHMIC DRUG. THE INTRAPROCEDURAL LEFT ATRIAL PRESSURE WAS ALSO HIGH, SUGGESTING THIS WAS DUE TO WORSENING HEART FAILURE. A 35MM WATCHMAN FLX PRO DELIVERY SYSTEM AND CLOSURE DEVICE (WDS) WAS INSERTED AND DEPLOYED BUT DID NOT MEET RELEASE CRITERIA SO IT WAS RECAPTURED AND REMOVED. A 40MM WDS WAS INSERTED, DEPLOYED, AND IMPLANTED AFTER MEETING RELEASE CRITERIA. HEART RATE WAS NOTED TO DECREASE AT THE CONCLUSION OF THE PROCEDURE, SO A DECISION WAS MADE TO INTRODUCE AN INTRA-AORTIC BALLOON PUMP, AND THE PATIENT WAS TRANSFERRED TO THE INTENSIVE CARE UNIT (ICU). THE PATIENT DIED APPROXIMATELY FIVE (5) DAYS POST INDEX PROCEDURE. THE PRESUMED CAUSE OF DEATH ARE EITHER NON-OCCLUSIVE MESENTERIC ISCHEMIA (NOMI) OR SEPSIS. THE PHYSICIAN CONCLUDED THAT IT WAS NOT DUE TO CARDIOGENIC SHOCK OR DUE TO THE WATCHMAN PROCEDURE OR DEVICE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21559737&pc=DSQ | 2916596-2025-01439 | 2024-03-21 04:00 | Death | THORATEC CORPORATION | 2025-03-10 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6) , AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. IT WAS REPORTED THAT THE PATIENT EXPIRED. THE DEVICE WAS NOT RETURNED FOR EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED BY THE ACCOUNT. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21559833&pc=DQY | 2124215-2025-14203 | 2025-02-27 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-03-10 04:00 | DQY | WATCHMAN TRUSTEER ACCESS SYSTEM | Material Integrity Problem | Air Embolism; Non specific EKG/ECG Changes; Heart Block | K240018 |  | 1 | Event Description: IT WAS REPORTED THAT AN AIR EMBOLISM OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED. A WATCHMAN TRUSTEER ACCESS SYSTEM (WAS) AND A 31MM WATCHMAN FLX PRO LAA CLOSURE DEVICE & DELIVERY SYSTEM (WDS) WERE SELECTED TO BE USED. THE PHYSICIAN POSITIONED THE WAS IN THE LEFT ATRIUM (LA) OF THE PATIENT. WHILE PREPARING TO INSERT THE PIGTAIL CATHETER THE PATIENT COUGHED, AND AIR WAS SUCTIONED INTO THE LA THROUGH THE WAS. AN AIR EMBOLISM WAS SEEN IN THE LA AND THE PHYSICIAN ASPIRATED THE AIR BUBBLES OUT OF THE PATIENT. THE PHYSICIAN CONTINUED THE PROCEDURE WITH THE HOPE THAT ANY REMAINING AIR COULD BE TRAPPED BY IMPLANTING THE CLOSURE DEVICE. THE WDS WAS INSERTED, AND THE CLOSURE DEVICE WAS SUCCESSFULLY IMPLANTED IN THE LAA OF THE PATIENT. SHORTLY AFTER THE CLOSURE DEVICE HAD BEEN RELEASED THE PATIENT BEGAN TO EXPERIENCE ST SEGMENT ELEVATION WHICH LED TO COMPLETE HEART BLOCK. THE PATIENT RECEIVED CARDIOPULMONARY RESUSCITATION AND IV MEDICATION BEFORE BECOMING STABLE. THE PATIENT WAS STILL RECOVERING FROM THIS EVENT THE NEXT DAY IN THE INTENSIVE CARE UNIT. THE PHYSICIAN BELIEVES THAT THE AIR EMBOLISM REACHED THE PATIENT'S BRAIN.Event Description: IT WAS REPORTED THAT AN AIR EMBOLISM OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED. A WATCHMAN TRUSTEER ACCESS SYSTEM (WAS) AND A 31MM WATCHMAN FLX PRO LAA CLOSURE DEVICE & DELIVERY SYSTEM (WDS) WERE SELECTED TO BE USED. THE PHYSICIAN POSITIONED THE WAS IN THE LEFT ATRIUM (LA) OF THE PATIENT. WHILE PREPARING TO INSERT THE PIGTAIL CATHETER THE PATIENT COUGHED, AND AIR WAS SUCTIONED INTO THE LA THROUGH THE WAS. AN AIR EMBOLISM WAS SEEN IN THE LA AND THE PHYSICIAN ASPIRATED THE AIR BUBBLES OUT OF THE PATIENT. THE PHYSICIAN CONTINUED THE PROCEDURE WITH THE HOPE THAT ANY REMAINING AIR COULD BE TRAPPED BY IMPLANTING THE CLOSURE DEVICE. THE WDS WAS INSERTED, AND THE CLOSURE DEVICE WAS SUCCESSFULLY IMPLANTED IN THE LAA OF THE PATIENT. SHORTLY AFTER THE CLOSURE DEVICE HAD BEEN RELEASED THE PATIENT BEGAN TO EXPERIENCE ST SEGMENT ELEVATION WHICH LED TO COMPLETE HEART BLOCK. THE PATIENT RECEIVED CARDIOPULMONARY RESUSCITATION AND IV MEDICATION BEFORE BECOMING STABLE. THE PATIENT WAS STILL RECOVERING FROM THIS EVENT THE NEXT DAY IN THE INTENSIVE CARE UNIT. THE PHYSICIAN BELIEVES THAT THE AIR EMBOLISM REACHED THE PATIENT'S BRAIN. IT WAS FURTHER REPORTED THAT THE PATIENT DID NOT RECOVER FROM THIS EVENT AND THE PATIENT DIED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21559848&pc=LWR | 9612164-2025-01242 | 2024-03-02 05:00 | Death | MEDTRONIC MEXICO | 2025-03-10 04:00 | LWR | AVALUS AORTIC TISSUE VALVE | Insufficient Information | Insufficient Information | P170006 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT HAD A TISSUE VALVE IMPLANTED IN THE AORTIC POSITION ON (B)(6) 2024. THE PATIENT WAS DECEASED ON (B)(6) 2024. THE CAUSE OF DEATH WAS NOT RECEIVED. THEREFORE, RELATEDNESS OF THE DEATH TO THE PRODUCT COULD NOT BE EXCLUDED. IT IS UNKNOWN WHETHER AN AUTOPSY WAS PERFORMED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21560208&pc=DQY | 2124215-2025-14239 | 2025-03-04 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-03-10 04:00 | DQY | WATCHMAN FXD CURVE ACCESS SYSTEM | Adverse Event Without Identified Device or Use Problem | Air Embolism; Arrhythmia; Stroke/CVA; Non specific EKG/ECG Changes; Low Blood Pressure/ Hypotension | K212228 |  | 1 | Event Description: IT WAS REPORTED THAT AN AIR EMBOLISM OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED. A WATCHMAN FXD CURVE ACCESS SYSTEM (WAS) AND A WATCHMAN FLX PRO LAA CLOSURE DEVICE & DELIVERY SYSTEM (WDS) WERE SELECTED TO BE USED. THE PHYSICIAN POSITIONED THE WAS AND THE DILATOR IN THE LEFT ATRIUM (LA) OF THE PATIENT. WHEN THE PHYSICIAN REMOVED THE DILATOR, THE PATIENT COUGHED DEEPLY RESULTING IN AIR BEING PULLED IN THROUGH THE WAS INTO THE LA OF THE PATIENT. THE PHYSICIAN ATTEMPTED TO ASPIRATE THE AIR OUT OF THE PATIENT. THE PATIENT BECAME HEMODYNAMICALLY UNSTABLE WITH A DROP IN BLOOD PRESSURE, OXYGEN LEVELS, ST SEGMENT ELEVATIONS AND BRIEF EPISODES OF VENTRICULAR TACHYCARDIA. THE PATIENT WAS PACED AND INTUBATED AT THIS TIME. ALL THE DEVICES WERE REMOVED FROM THE PATIENT, AND THEY STABILIZED AFTER ALL THE INTERVENTIONS. THE PROCEDURE WAS ENDED WITH NO CLOSURE DEVICE ATTEMPTED TO BE IMPLANTED. IMAGING SHOWED NO FURTHER AIR IN THE LEFT ATRIUM OR ANY OTHER COMPLICATIONS. THE PATIENT WAS EXTUBATED AND HEMODYNAMICALLY STABLE. FOLLOWING MONITORING OF THE PATIENT THE PHYSICIAN BELIEVED THE PATIENT MAY HAVE EXPERIENCED A CEREBRAL VASCULAR ACCIDENT DUE TO THE PATIENT'S SYMPTOMS OF FACIAL DROOP AND SLURRED SPEECH FOLLOWING EXTUBATION. THE PATIENT UNDERWENT COMPUTED TOMOGRAPHY (CT) IMAGING FOR EVALUATION. THE CT OF THE HEAD AND CT ANGIOGRAM WERE ALL NEGATIVE. THE PATIENT WAS KEPT OVERNIGHT FOR OBSERVATION AND TO UNDERGO A MAGNETIC RESONANCE IMAGING SCAN.Event Description: IT WAS REPORTED THAT AN AIR EMBOLISM OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED. A WATCHMAN FXD CURVE ACCESS SYSTEM (WAS) AND A WATCHMAN FLX PRO LAA CLOSURE DEVICE & DELIVERY SYSTEM (WDS) WERE SELECTED TO BE USED. THE PHYSICIAN POSITIONED THE WAS AND THE DILATOR IN THE LEFT ATRIUM (LA) OF THE PATIENT. WHEN THE PHYSICIAN REMOVED THE DILATOR, THE PATIENT COUGHED DEEPLY RESULTING IN AIR BEING PULLED IN THROUGH THE WAS INTO THE LA OF THE PATIENT. THE PHYSICIAN ATTEMPTED TO ASPIRATE THE AIR OUT OF THE PATIENT. THE PATIENT BECAME HEMODYNAMICALLY UNSTABLE WITH A DROP IN BLOOD PRESSURE, OXYGEN LEVELS, ST SEGMENT ELEVATIONS AND BRIEF EPISODES OF VENTRICULAR TACHYCARDIA. THE PATIENT WAS PACED AND INTUBATED AT THIS TIME. ALL THE DEVICES WERE REMOVED FROM THE PATIENT, AND THEY STABILIZED AFTER ALL THE INTERVENTIONS. THE PROCEDURE WAS ENDED WITH NO CLOSURE DEVICE ATTEMPTED TO BE IMPLANTED. IMAGING SHOWED NO FURTHER AIR IN THE LEFT ATRIUM OR ANY OTHER COMPLICATIONS. THE PATIENT WAS EXTUBATED AND HEMODYNAMICALLY STABLE. FOLLOWING MONITORING OF THE PATIENT THE PHYSICIAN BELIEVED THE PATIENT MAY HAVE EXPERIENCED A CEREBRAL VASCULAR ACCIDENT DUE TO THE PATIENT'S SYMPTOMS OF FACIAL DROOP AND SLURRED SPEECH FOLLOWING EXTUBATION. THE PATIENT UNDERWENT COMPUTED TOMOGRAPHY (CT) IMAGING FOR EVALUATION. THE CT OF THE HEAD AND CT ANGIOGRAM WERE ALL NEGATIVE. THE PATIENT WAS KEPT OVERNIGHT FOR OBSERVATION AND TO UNDERGO A MAGNETIC RESONANCE IMAGING SCAN. IT WAS FURTHER REPORTED THAT NO CEREBRAL VASCULAR ACCIDENT WAS CONFIRMED. CT AND MRI IMAGING WERE ALL NEGATIVE. THE PATIENTS SYMPTOMS RESOLVED FROM THIS EVENT. OTHER CHALLENGES AROSE DURING THE HOSPITALIZATION AND THE PATIENT REMAINED IN THE HOSPITAL. TWO (2) WEEKS POST PROCEDURE THE PATIENT DIED DUE TO THE COMPLICATIONS THAT RESULTED DURING HOSPITALIZATION. THE PATIENTS FAMILY CHOSE TO PROVIDE COMFORT CARE TO THE PATIENT BEFORE THE PATIENT DIED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21560227&pc=NIK | 9614453-2025-00969 | 2024-12-19 05:00 | Death | MEDTRONIC EUROPE SARL | 2025-03-10 04:00 | NIK | CLARIA MRI QUAD CRT-D SURESCAN | Adverse Event Without Identified Device or Use Problem | Cardiomyopathy; Septic Shock; Multiple Organ Failure | P010031 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IS DECEASED AND THE CAUSE OF DEATH WAS PROVIDED AS NATURAL WITH THE PRIMARY CAUSE LISTED AS MULTIORGAN FAILURE AND THE SECONDARY CAUSES LISTED AS SEPTIC SHOCK, EXTENDED-SPECTRUM BETA-LACTAMASE (ESBL) BACTEREMIA, AND END STAGE CARDIOMYOPATHY. Manufacturer Narrative: CONTINUATION OF D10: 6935M62 LEAD, IMPLANTED: (B)(6) 2015. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE DEVICE WAS RETURNED AND ANALYZED. RETURNED PRODUCT ANALYSIS WAS PERFORMED AND NO ANOMALIES WERE FOUND. THE DEVICE WAS RETURNED AND SCREENING ANALYSIS WAS PERFORMED, BUT NO ISSUE WAS IDENTIFIED REQUIRING FULL ANALYSIS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21560244&pc=LWS | 2649622-2025-06761 | 2024-12-19 05:00 | Death | MPRI | 2025-03-10 04:00 | LWS | SPRINT QUATTRO SECURE S MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Cardiomyopathy; Septic Shock; Multiple Organ Failure | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IS DECEASED AND THE CAUSE OF DEATH WAS PROVIDED AS NATURAL WITH THE PRIMARY CAUSE LISTED AS MULTIORGAN FAILURE AND THE SECONDARY CAUSES LISTED AS SEPTIC SHOCK, EXTENDED-SPECTRUM BETA-LACTAMASE (ESBL) BACTEREMIA, AND END STAGE CARDIOMYOPATHY. Manufacturer Narrative: CONTINUATION OF D10: 6935M62 LEAD, IMPLANTED: (B)(6) 2015. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE PROXIMAL PORTION OF THE LEAD WAS RETURNED, ANALYZED, AND NO ANOMALIES WERE FOUND. VISUAL ANALYSIS OF THE LEAD INDICATED APPARENT EXPLANT DAMAGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21560256&pc=DTB | 2649622-2025-06760 | 2024-12-19 05:00 | Death | MPRI | 2025-03-10 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Cardiomyopathy; Septic Shock; Multiple Organ Failure | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IS DECEASED AND THE CAUSE OF DEATH WAS PROVIDED AS NATURAL WITH THE PRIMARY CAUSE LISTED AS MULTIORGAN FAILURE AND THE SECONDARY CAUSES LISTED AS SEPTIC SHOCK, EXTENDED-SPECTRUM BETA-LACTAMASE (ESBL) BACTEREMIA, AND END STAGE CARDIOMYOPATHY. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE PROXIMAL PORTION OF THE LEAD WAS RETURNED, ANALYZED, AND NO ANOMALIES WERE FOUND. VISUAL ANALYSIS OF THE LEAD INDICATED APPARENT EXPLANT DAMAGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: 6935M62 LEAD, IMPLANTED (B)(6) 2015. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21560259&pc=OJX | 2649622-2025-06762 | 2024-12-19 05:00 | Death | MPRI | 2025-03-10 04:00 | OJX | ATTAIN PERFORMA STRAIGHT MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Cardiomyopathy; Septic Shock; Multiple Organ Failure | P080006 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IS DECEASED AND THE CAUSE OF DEATH WAS PROVIDED AS NATURAL WITH THE PRIMARY CAUSE LISTED AS MULTIORGAN FAILURE AND THE SECONDARY CAUSES LISTED AS SEPTIC SHOCK, EXTENDED-SPECTRUM BETA-LACTAMASE (ESBL) BACTEREMIA, AND END STAGE CARDIOMYOPATHY. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE PROXIMAL PORTION OF THE LEAD WAS RETURNED, ANALYZED, AND NO ANOMALIES WERE FOUND. VISUAL ANALYSIS OF THE LEAD INDICATED APPARENT EXPLANT DAMAGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: 6935M62 LEAD, IMPLANTED: (B)(6) 2015. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21560453&pc=DSQ | 3007042319-2025-00635 | 2025-02-01 05:00 | Death | HEARTWARE, INC. | 2025-03-10 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage; Confusion/ Disorientation; Heart Failure/Congestive Heart Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WITH A HISTORY OF CARDIOMYOPATHY EXPERIENCED A SERIES OF MEDICAL ISSUES, INCLUDING A SUBDURAL HEMATOMA AND WORSENING RIGHT VENTRICULAR (RV) HEART FAILURE. THE PATIENT PRESENTED WITH PROGRESSIVE CONFUSION AND AGITATION, WHICH NECESSITATED INTUBATION AND SEDATION TO MANAGE THE AGITATION. A HEAD COMPUTED TOMOGRAPHY CONFIRMED THE PRESENCE OF A SUBDURAL HEMATOMA. DESPITE THESE FINDINGS, NO INTERVENTION WAS TAKEN FOR THE HEMATOMA AS PER THE FAMILY'S WISHES. A TRANSTHORACIC ECHOCARDIOGRAM INDICATED WORSENING RV HEART FAILURE. ULTIMATELY, THE FAMILY DECIDED TO PURSUE HOSPICE CARE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN USE.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT DIED IN HOSPICE CARE. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: THE VENTRICULAR ASSIST DEVICE (VAD)(B)(6) WAS NOT RETURNED FOR EVALUATION AS IT REMAINED IMPLANTED. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORDS WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, RIGHT VENTRICULAR FAILURE, WORSENING HEART FAILURE, NEUROLOGICAL DYSFUNCTION AND DEATH ARE KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. THERE IS NO EVIDENCE THAT THE PATIENT HAD A HISTORY OF SIMILAR ADVERSE EVENTS. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: (B)(6) WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORDS WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, RIGHT VENTRICULAR FAILURE, WORSENING HEART FAILURE, AND NEUROLOGICAL DYSFUNCTION ARE KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. THERE IS NO EVIDENCE THAT THE PATIENT HAD A HISTORY OF SIMILAR ADVERSE EVENTS. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED AS ADDITIONAL INFORMATION HAS BEING RECEIVED FOR THIS EVENT AND SECTIONS HAVE BEEN UPDATED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21560571&pc=DSQ | 2916596-2025-01633 | 2025-02-23 05:00 | Death | THORATEC CORPORATION | 2025-03-10 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. DUE TO PATIENT PRIVACY LAWS THE MANAGING HOSPITAL WOULD NOT COMMUNICATE PATIENT OUTCOME INFORMATION. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. THE OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. NO AUTOPSY WAS PERFORMED. THE DEVICE WAS NOT EXPLANTED. THE DEVICE WOULD NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: SECTION A1-A4: PATIENT INFORMATION WAS NOT PROVIDED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. DUE TO HOSPITAL PRIVACY POLICY, NO ADDITIONAL INFORMATION WILL BE PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, REV. G IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21561028&pc=NPT | 2025587-2025-01782 | 2025-02-12 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-10 04:00 | NPT | EVOLUT FX VALVE | Difficult to Open or Close; Device Dislodged or Dislocated; Patient Device Interaction Problem | Ventricular Fibrillation; Rupture; Cardiac Tamponade; Cardiogenic Shock; Heart Block | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT PRIOR TO THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE IN A PATIENT WITH AN ANNULUS PERIMETER OF 90.6 MM, AN ELECTROCARDIOGRAM REVEALED BOTH FIRST DEGREE ATRIO-VENTRICULAR BLOCK AND COMPLETE RIGHT BUNDLE BRANCH BLOCK. DURING IMPLANT, WHEN THE VALVE WAS PLACED AT A DEPTH OF 2 MM ON THE NON-CORONARY CUSP (NCC) AND 8 MM ON THE LEFT CORONARY CUSP (LCC), COMPLETE HEART BLOCK OCCURRED. THE TRANSVENOUS PACEMAKER THAT WAS PLACED DURING THE PROCEDURE WAS LEFT IN PLACE POST-OPERATIVELY.Event Description: IT WAS REPORTED THAT PRIOR TO THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE IN A PATIENT WITH AN ANNULUS PERIMETER OF 90.6 MM, AN ELECTROCARDIOGRAM REVEALED BOTH FIRST DEGREE ATRIO-VENTRICULAR BLOCK AND COMPLETE RIGHT BUNDLE BRANCH BLOCK. DURING IMPLANT, WHEN THE VALVE WAS PLACED AT A DEPTH OF 2 MM ON THE NON-CORONARY CUSP (NCC) AND 8 MM ON THE LEFT CORONARY CUSP (LCC), COMPLETE HEART BLOCK OCCURRED. THE TRANSVENOUS PACEMAKER THAT WAS PLACED DURING THE PROCEDURE WAS LEFT IN PLACE POST-OPERATIVELY. ADDITIONAL INFORMATION WAS RECEIVED INDICATING THAT 2 DAYS FOLLOWING THE PROCEDURE, VENTRICULAR FIBRILLATION CARDIOPULMONARY ARREST OCCURRED AND AN ECHOCARDIOGRAM REVEALED CARDIAC TAMPONADE. SUBSEQUENTLY, INTUBATION AND A THORACOTOMY WERE PERFORMED AND A RUPTURE OF THE HEART BASE WAS OBSERVED. AN AORTIC DISSECTION HAD OCCURRED; HOWEVER, PERICARDIAL GAIN WAS NOT INITIALLY OBSERVED IN THE POSTOPERATIVE ECHOCARDIOGRAM. DESPITE THIS, IT WAS BELIEVED THAT THE DISSECTION HAD PROGRESSED CAUSING THE HEART BASE TO RUPTURE LEADING TO CARDIAC TAMPONADE. BLOOD WAS TRANSFUSED. INTERVENTION WAS UNSUCCESSFUL AND THE PATIENT PASSED AWAY THE SAME DAY.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT PER THE PHYSICIAN, THE VALVE AND DCS CONTRIBUTED TO THE VASCULAR INJURY. THE CAUSE OF THE VASCULAR INJURY WAS DUE TO THE VALVE NOT BEING STABLE WHICH REQUIRED 3 RECAPTURES. THE VALVE WAS ABLE TO BE PASSED THROUGH THE PATIENT'S NATIVE VALVE WITHOUT BEING COMPLETELY RECAPTURED WITHIN THE CAPSULE.Event Description: IT WAS REPORTED THAT PRIOR TO THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE IN A PATIENT WITH AN ANNULUS PERIMETER OF 90.6 MM, AN ELECTROCARDIOGRAM REVEALED BOTH FIRST DEGREE ATRIO-VENTRICULAR BLOCK AND COMPLETE RIGHT BUNDLE BRANCH BLOCK. DURING IMPLANT, WHEN THE VALVE WAS PLACED AT A DEPTH OF 2 MM ON THE NON-CORONARY CUSP (NCC) AND 8 MM ON THE LEFT CORONARY CUSP (LCC), COMPLETE HEART BLOCK OCCURRED. THE TRANSVENOUS PACEMAKER THAT WAS PLACED DURING THE PROCEDURE WAS LEFT IN PLACE POST-OPERATIVELY. ADDITIONAL INFORMATION WAS RECEIVED INDICATING THAT 2 DAYS FOLLOWING THE PROCEDURE, VENTRICULAR FIBRILLATION CARDIOPULMONARY ARREST OCCURRED AND AN ECHOCARDIOGRAM REVEALED CARDIAC TAMPONADE. SUBSEQUENTLY, INTUBATION AND A THORACOTOMY WERE PERFORMED AND A RUPTURE OF THE HEART BASE WAS OBSERVED. AN AORTIC DISSECTION HAD OCCURRED; HOWEVER, PERICARDIAL GAIN WAS NOT INITIALLY OBSERVED IN THE POSTOPERATIVE ECHOCARDIOGRAM. DESPITE THIS, IT WAS BELIEVED THAT THE DISSECTION HAD PROGRESSED CAUSING THE HEART BASE TO RUPTURE LEADING TO CARDIAC TAMPONADE. BLOOD WAS TRANSFUSED. INTERVENTION WAS UNSUCCESSFUL AND THE PATIENT PASSED AWAY THE SAME DAY. ADDITIONAL INFORMATION WAS RECEIVED THAT PER THE PHYSICIAN, THE VALVE AND DCS CONTRIBUTED TO THE VASCULAR INJURY. THE CAUSE OF THE VASCULAR INJURY WAS DUE TO THE VALVE NOT BEING STABLE WHICH REQUIRED 3 RECAPTURES. THE VALVE WAS ABLE TO BE PASSED THROUGH THE PATIENT'S NATIVE VALVE WITHOUT BEING COMPLETELY RECAPTURED WITHIN THE CAPSULE. ADDITIONAL INFORMATION WAS RECEIVED INDICATING THAT A PRE-IMPLANT BALLOON DILATION WAS PERFORMED USING A 22 MM BALLOON. THE VALVE WAS DEPLOYED AT A DEPTH OF 1-2 MM ON THE NON-CORONARY CUSP (NCC) AND 7-8 MM ON THE LEFT CORONARY CUSP (LCC) USING A MID PIGTAIL STARTING POINT. THE VALVE WAS NOTED TO BE "UNSTABLE" DUE TO THE PLACEMENT DEPTH NOT BEING COAXIAL. AT 80% DEPLOYMENT, THE VALVE DISLODGED IN THE AORTIC DIRECTION TO A DEPTH OF 0 MM ON THE NCC AND 7 MM ON THE LCC. AS A RESULT, THE VALVE WAS RECAPTURED AND RE-POSITIONED. THE VALVE WAS RECAPTURED A TOTAL OF 3 TIMES DUE TO DISLODGEMENTS. PER THE PHYSICIAN, THE PATIENT'S HORIZONTAL AORTA CONTRIBUTED TO THE DISLODGEMENTS. DURING THE RECAPTURES, THE VALVE COULD NOT BE FULLY RETRACTED WITHIN THE CAPSULE DUE TO THE INFLUENCE OF THE HORIZONAL AORTA AND VASCULAR TORTUOSITY. THERE WAS DIFFICULTY RECAPTURING THE VALVE, DESPITE THE USE OF TORQUE AND FULLY CLOSING THE DCS HANDLE. OF NOTE, A NON-MEDTRONIC GUIDEWIRE WAS USED. Manufacturer Narrative: UPDATED DATA: B5. ADDED FOURTH PARAGRAPH. H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID D-EVOLUTFX-34 (LOT: 0012402703); PRODUCT TYPE: 0195-HEART VALVES; PRODUCT ID L-EVOLUTFX-34 (LOT: 0012210199); PRODUCT TYPE: 0195-HEART VALVES; SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. H6. CORRECTED DATA: D4. THE CODES PRESENT IN SECTION H6 CORRESPOND TO MULTIPLE COMPONENTS/PRODUCTS THAT COMPRISE THIS REPORTED EVENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME EVOLUT FX DCS; PRODUCT ID D-EVOLUTFX-2329 (LOT: 0012402703); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: NON-IMPLANTABLE DEVICE UPDATED DATA: B2. B5. ADDED. SECOND PARAGRAPH. H1. H6. H11. SYSTEM REPORTED D-EVOLUTFX-2329. ADDITIONAL CODES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21561050&pc=LDD | 3003832357-2025-000181 | 2025-03-04 05:00 | Death | SCHILLER AG | 2025-03-10 04:00 | LDD | TEMPUS LS | Failure to Deliver Shock/Stimulation | Cardiac Arrest | K200849 |  | 1 | Event Description: THIS REPORT IS BASED ON INFORMATION PROVIDED BY REMOTE SERVICE ENGINEER RSE AND IS CURRENTLY UNDER INVESTIGATION BY THE PHILIPS COMPLAINT HANDLING TEAM. PHILIPS RECEIVED A COMPLAINT ON THE TEMPUS LS MANUAL INDICATING THE DEVICE, DURING A CARDIAC ARREST ON THE MORNING OF (B)(6) 2025, ENCOUNTERED A "CHECK DEFIBRILLATION PADS" ERROR MESSAGE. THE CREWS PLACED DEFIBRILLATOR PADS ON THE PATIENT (35YO MALE), THEN PLUGGED THE PADS INTO THE TEMPUS LS. THE MONITOR SHOWED A RHYTHM DURING COMPRESSIONS FOR 2-3 SECONDS, THEN NO WAVEFORM AND THE SCREEN DISPLAYED THE MESSAGE "CHECK DEFIBRILLATION PADS." PADS WERE CONFIRMED TO BE IN PLACE AND APPROPRIATELY ADHERED TO THE PATIENT. WHEN THE CORDS WERE UNPLUGGED THEN PLUGGED BACK IN, THE MONITOR RECOGNIZED THAT SOMETHING HAD BEEN ATTACHED TO IT, THEN WENT BACK TO THE ERROR MESSAGE. ONCE OR TWICE, THERE WERE 1-2 SECONDS OF A RHYTHM SHOWING ON THE MONITOR, BUT FOR THE MAJORITY OF THE TIME THE CREW SAW THE "CHECK DEFIBRILLATION PADS" ERROR MESSAGE WITH NO RHYTHM SHOWN. THE ISSUE WAS RESOLVED WHEN WE REMOVED THE FIRST SET OF PADS AND REPLACED THEM WITH ANOTHER SET OF PADS. Manufacturer Narrative: UPDATED HEALTH IMPACT GRID, ADVERSE EVENT/PRODUCT PROBLEM AND DEATH DATE ORIGINALLY PROVIDED ON MFR REPORT NUMBER 3003832357-2025-000181. Manufacturer Narrative: EVENT DATE AND DATE OF DEATH UPDATED AFTER RECEIVING GFE RESPONSE THAT WAS ORIGINALLY PROVIDED ON MFR REPORT NUMBER 3003832357-2025-000181.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21561948&pc=CBK | 9617566-2025-00002 | 2024-10-11 04:00 | Death | BREAS MEDICAL AB | 2025-03-10 04:00 | CBK | VIVO 45LS | Insufficient Information | Insufficient Information | K193586 |  | 1 | Event Description: JANUARY 9, 2025: BREAS MEDICAL RECEIVED A CUSTOMER REPORT THAT A PATIENT HAD EXPIRED WHILE USING A BREAS VIVO 45LS. THIS WAS PROVIDED BY THE CUSTOMER ROTECH INC.: 'THIS PT DAUGHTER SENT US A TEXT STATING THE MOM PASSED AWAY AND STATED THE MACHINE PLAYED A FACTOR IN HER DEATH I HAVE BEEN TRYING TO GET THIS SENT OFF AND I HAVE NOT HEARD ANYTHING FROM THE FAMILY SINCE THIS DATE ' BREAS RECEIVED THE DEVICE ON JANUARY 16, 2025. THE INVESTIGATION IS ONGOING. LOG FILES HAVE BEEN REVIEWED BUT THERE IS NO DATA IN THE FILES FROM THE REPORTED DATE OF THE EVENT. NOTHING IN THE LOG FILES INDICATES THAT THE VENTILATOR STOPPED WORKING. BREAS REACHED OUT TO THE CUSTOMER VIA EMAIL ON JANUARY 27, 2025 TO REQUEST MORE INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21562049&pc=DQY | 2029214-2025-00604 | 2024-08-30 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-03-10 04:00 | DQY | REFLEX CATHETER | Adverse Event Without Identified Device or Use Problem | Renal Failure | K110055 |  | 1 | Event Description: YAO, J., ZHENG, B., SUN, Q., ZHANG, F., JI, Z., WANG, C., WU, P., & SHI, H. (2024). HYBRID SURGERY FOR SYMPTOMATIC CHRONIC INTERNAL CAROTID ARTERY OCCLUSION: A SINGLE-CENTER EXPERIENCE. ACTA NEUROCHIRURGICA: THE EUROPEAN JOURNAL OF NEUROSURGERY, 166(1). HTTP S://DOI.ORG/10.1007/S00701-024-06250-4. MEDTRONIC REVIEW OF THE LITERATURE ARTICLE FOUND A RETROSPECTIVE CASE SERIES WAS CONDUCTED AT A SINGLE CENTER REVIEWING CASES OF 27 PATIENTS WHO UNDERWENT HYBRID SURGERY FOR THE TREATMENT OF SYMPTOMATIC CHRONIC INTERNAL CAROTID ARTERY OCCLUSION (ICAO). IT WAS NOTED THAT PATIENT WITH ICAO ARE AT HIGH RISK FOR REOCCLUSION. THE HYBRID SURGICAL TECHNIQUE COMBINES CAROTID ENDARTERECTOMY (CEA) AND ENDOVASCULAR INTERVENTION (EI). THIS PROCEDURE INVOLVES PERFORMING CEA AT THE PROXIMAL END OF THE CAROTID ARTERY TO REMOVE ATHEROSCLEROTIC PLAQUES AND THROMBI. SUBSEQUENTLY, EI TARGETS DISTAL STENOSIS OR OCCLUSION IN THE INTERNAL CAROTID ARTERY THAT CEA CANNOT REACH, THUS SUCCESSFULLY RECANALIZING THE CHRONIC ICAO. MULTIPLE MANUFACTURERS' DEVICES WERE USED IN THE TREATMENT PROCEDURES. IN SOME CASES, THE INTERMEDIATE CATHETER WAS A NAVIEN 6F 115CM CATHETER. HOWEVER, OTHER CATHETERS WERE USED AND IT WAS NOT SPECIFIED WHICH DEVICES WERE USED IN EACH PROCEDURE. NO DEVICE MALFUNCTION OR DEFICIENCY OF THE CATHETER WAS REPORTED. IT WAS NOTED THAT SUCCESSFUL REVASCULARIZATION WAS ACHIEVED IN 100% OF CASES. THE FOLLOWING PERI-PROCEDURE ADVERSE EVENTS WERE REPORTED: ONE PATIENT DEATH OCCURRED. THE PATIENT DEVELOPED ACUTE RENAL FAILURE POST-SURGERY MARKED BY ANURIA, AND UNFORTUNATELY PASSED AWAY ON THE THIRD DAY AFTER THE OPERATION. TWO PATIENTS EXPERIENCED ISCHEMIC COMPLICATIONS. ONE OF THESE PATIENTS HAD INTRAOPERATIVE STENT THROMBOSIS, AS CONFIRMED BY DIGITAL SUBTRACTION ANGIOGRAPHY (DSA) SHOWING THROMBI IN THE DISTAL ICA AND M2 SEGMENT OF THE MIDDLE CEREBRAL ARTERY. THIS PATIENT UNDERWENT A SUCCESSFUL MECHANICAL THROMBECTOMY AND HAD A FAVORABLE RECOVERY, EXHIBITING NO NEUROLOGICAL ISCHEMIC SYMPTOMS AND WAS DISCHARGED WITH AN MRS SCORE OF 0. THE OTHER PATIENT PRESENTED WITH LIMB WEAKNESS IN THE VASCULAR SUPPLY AREA OF THE SURGERY, CONFIRMED BY POSTOPERATIVE DIFFUSION-WEIGHTED IMAGING (DWI) AS A SMALL FOCAL ACUTE CEREBRAL INFARCTION, LIKELY DUE TO EMBOLIZATION OF DEBRIS. SYMPTOMS PARTIALLY ALLEVIATED FOLLOWING IMPROVED CIRCULATION AND REHABILITATION INTERVENTIONS. ONE PATIENT EXPERIENCED A HEMORRHAGIC COMPLICATION THAT WAS ATTRIBUTED TO THE ANTIPLATELET THERAPY. AS THERE WAS NO DEVICE MALFUNCTION AND THE EVENT WAS NOT ASSOCIATED WITH THE DEVICE OR PROCEDURE, THE HEMORRHAGIC EVENT IS NOT CONSIDERED A COMPLAINT. Manufacturer Narrative: **SERIOUS ADVERSE EVENTS/INJURIES REPORTED IN THE ARTICLE WILL BE REPORTED SEPARATELY. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21562449&pc=DSQ | 2916596-2025-01440 | 2023-08-24 04:00 | Death | THORATEC CORPORATION | 2025-03-10 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT'S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. IT WAS REPORTED THAT THE PATIENT PASSED AWAY. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED BY THE ACCOUNT. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21562450&pc=OZD | 1220648-2025-26605 | 2025-02-05 05:00 | Death | ABIOMED, INC. | 2025-03-10 04:00 | OZD | IMPELLA CP | Pumping Stopped | Insufficient Information | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WITH SYSTEMIC CIRCULATORY FAILURE WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. IT WAS REPORTED THAT THE PUMP STOPPED ON THE 10TH DAY OF SUPPORT. THE PUMP WAS EXPLANTED AND THERE WERE NO PLANS TO ESCALATE CARE. THE PATIENT HAD AN IMPELLA DEPENDENT CONDITION AND EXPIRED SHORTLY AFTER THE PUMP WAS EXPLANTED. Manufacturer Narrative: THE IMPELLA PUMP AND PURGE CASSETTE WERE RETURNED, AND AN EVALUATION IS UNDERWAY. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH RISK PERCUTANEOUS CORONARY INTERVENTION. SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA CATHETER: ¿ACHIEVING AN ACT = 250 SECONDS PRIOR TO REMOVING THE DILATOR WILL HELP PREVENT A THROMBUS FROM ENTERING THE CATHETER AND CAUSING A SUDDEN STOP ON STARTUP.¿ SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA RP WITH SMARTASSIST SYSTEM CATHETER. ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA RP WITH SMARTASSIST SYSTEM CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP. MONITOR MOTOR CURRENT AND CONSIDER REPLACING IMPELLA RP WITH SMARTASSIST SYSTEM CATHETER WHENEVER A RISE IN MOTOR CURRENT IS SEEN.¿ IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK. SECTION: TROUBLESHOOTING THE PURGE SYSTEM ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP.¿ SECTION: IMPELLA STOPPED ¿IF THE IMPELLA CATHETER HAS STOPPED SUDDENLY: 1. TRY TO RESTART THE CATHETER AT PREVIOUS P-LEVEL. 2. IF THE IMPELLA DOES NOT RESTART, TRY TO RESTART AT P-2. 3. IF THE IMPELLA DOES NOT RESTART OR STOPS AGAIN, WAIT 1 MINUTE AND TRY TO RESTART AGAIN. 4. IF THE IMPELLA RESTARTS, WEAN DOWN TO P-2 AS THE PATIENT CAN TOLERATE. UNDER THESE CIRCUMSTANCES, CATHETER FUNCTION IS NOT RELIABLE AND THE IMPELLA MAY STOP AGAIN. 5. IF THE IMPELLA DOES NOT RESTART, REMOVE THE IMPELLA FROM THE VENTRICLE AS SOON AS POSSIBLE TO AVOID AORTIC INSUFFICIENCY.¿ IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH RISK PERCUTANEOUS CORONARY INTERVENTION SECTION: WARNINGS, CAUTIONS, AND PRECAUTIONS ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿ Manufacturer Narrative: THE INVESTIGATION OF PUMP STOP HAS BEEN COMPLETED. THE IMPELLA CP WAS RETURNED FOR EVALUATION. NO PHYSICAL ABNORMALITIES WERE FOUND ON THE RETURNED PRODUCT, AND THE PUMP PASSED ELECTRICAL TESTING. UPON FURTHER INSPECTION, A LARGE PURGE GAP WAS FOUND WHICH IS DUE TO BEARING WEAR. DATA LOGS WERE NOT PROVIDED FOR THIS CASE RUN. AS PER THE CLINICAL INFORMATION, THE PUMP STOP HAPPED ON THE TENTH DAY OF USE. THE ROOT CAUSE OF THE PUMP STOP WAS DUE TO BEARING WEAR AS THE PUMP WAS RUN PAST THE INDICATED DESIGN LIFE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21562462&pc=DSQ | 2916596-2025-01384 | 2024-07-11 04:00 | Death | THORATEC CORPORATION | 2025-03-10 04:00 | DSQ | THORATEC® HEARTMATE II®, LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. ADDITIONAL INFORMATION WAS NOT PROVIDED. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE II LVAS, SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C, IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21562487&pc=NPT | 2025587-2025-01794 | 2024-12-20 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-10 04:00 | NPT | EVOLUT R TRANSCATHETER AORTIC VALVE | Perivalvular Leak; Insufficient Information; Patient Device Interaction Problem; Central Regurgitation | Anemia; Hemorrhage/Bleeding; Low Blood Pressure/ Hypotension; Pleural Effusion; Renal Failure; Shock; Cardiogenic Shock; Heart Failure/Congestive Heart Failure; Valvular Insufficiency/ Regurgitation; Aortic Valve Insufficiency/ Regurgitation | P130021 |  | 1 | Event Description: APPROXIMATELY 3 MONTHS LATER THE PATIENT UNDERWENT A SURGICAL PROCEDURE FOR THE EXPLANT OF THE TRANSCATHETER AORTIC VALVE. HOSPITALIZATION WAS REQUIRED. A SURGICAL VALVE WAS IMPLANTED AND A CORONARY ARTERY BYPASS WAS ALSO PERFORMED. AN INTRA-AORTIC BALLOON PUMP WAS PLACED. BLOOD PRODUCTS WERE GIVEN FOR THE EXPECTED POST-OPERATIVE ANEMIA FROM THE SURGICAL VALVE IMPLANT. TWO UNITS OF FRESH FROZEN PLASMA (FFP), 2 UNITS OF PLATELETS, AND 2 UNITS OF CRYO WERE GIVEN INTRA-OPERATIVELY. THREE UNITS OF PACKED RED BLOOD CELLS, 1 CRYO, AND 1 UNIT OF PLATELETS WERE GIVEN POST-OPERATIVELY. PRESSOR SUPPORT WAS REQUIRED. A SWAN CATHETER PLACED. THE EJECTION FRACTION MEASURED 20%. THE DAY FOLLOWING THE VALVE IMPLANT AN TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) WAS PERFORMED AND BIVENTRICULAR HEART FAILURE WAS IDENTIFIED WITH AN EJECTION FRACTION LESS THAN 15%. CARDIOGENIC SHOCK WAS NOTED AND A NEED FOR VASOPRESSOR MEDICATION. AS RESULT, AN IMPELLA DEVICE REPLACED THE BALLOON PUMP AND VENOARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION (VA ECMO) WAS PLACED. THE BIVENTRICULAR HEART FAILURE WAS REPORTED AS RELATED TO THE SURGICAL VALVE IMPLANT PROCEDURE. TWO DAYS FOLLOWING THE VALVE IMPLANT THE EXPECTED BLOOD LOSS FROM THE VALVE INTERVENTION WAS GREATER THAN EXPECTED. MULTIPLE HEMOGLOBIN DRAWS RANGED FROM 7.1 TO 11.2. ON THIS DATE, THE HEMOGLOBIN MEASURED 8. AS RESULT, 1 UNIT OF PACKED RED BLOOD CELLS WAS TRANSFUSED. CONTINUOUS RENAL REPLACEMENT THERAPY (CRRT) WAS ALSO REQUIRED. THREE DAYS FOLLOWING THE VALVE REPLACEMENT A CHEST ULTRASOUND IDENTIFIED A LEFT PLEURAL EFFUSION. TREATMENT WAS NOT REQUIRED FOR THE EFFUSION. THIS SAME DAY AN ECHOCARDIOGRAM SHOWED A DECREASED EJECTION FRACTION. ECMO SUPPORT REQUIRED. THIS WAS REPORTED AS PROBABLY RELATED TO THE SURGICAL VALVE IMPLANT PROCEDURE. THE SEVERE PARAVALVULAR LEAK (PVL) EVENT WAS REPORTED RESOLVED WITHOUT SEQUELAE 8 DAYS FOLLOWING THE VALVE REPLACEMENT PROCEDURE. HOWEVER, IT WAS ALSO REPORTED THAT TRANSAMINITIS LEVELS TRENDED DOWN. LIVER SHOCK OCCURRED SECONDARY TO THE CARDIOGENIC SHOCK. THE CARDIOGENIC SHOCK WAS REPORTED AS RELATED TO THE SURGICAL VALVE AND ITS IMPLANT PROCEDURE. ULTIMATELY, THE PATIENT DIED 8 DAYS FOLLOWING THE TRANSCATHETER VALVE REPLACEMENT. THE DEATH TYPE WAS REPORTED AS CARDIAC.Event Description: IT WAS REPORTED THAT APPROXIMATELY 7 YEARS AND 10 MONTHS FOLLOWING THE IMPLANT OF THE TRANSCATHETER BIOPROSTHETIC VALVE AN ECHOCARDIOGRAM IDENTIFIED SEVERE TRANSVALVULAR REGURGITATION. INTRAVENOUS DIURETIC WAS PROVIDED. SURGICAL WORK-UP WHICH INCLUDED A CORONARY ANGIOGRAPHY WAS REPORTED. PATIENT COMPLICATIONS HAVE NOT BEEN REPORTED AS A RESULT OF THIS EVENT. APPROXIMATELY 3 MONTHS LATER THE PATIENT UNDERWENT A SURGICAL PROCEDURE FOR THE EXPLANT OF THE TRANSCATHETER AORTIC VALVE. HOSPITALIZATION WAS REQUIRED. A SURGICAL VALVE WAS IMPLANTED, AND A CORONARY ARTERY BYPASS WAS ALSO PERFORMED. AN INTRA-AORTIC BALLOON PUMP WAS PLACED. BLOOD PRODUCTS WERE GIVEN FOR THE EXPECTED POST-OPERATIVE ANEMIA FROM THE SURGICAL VALVE IMPLANT. TWO UNITS OF FRESH FROZEN PLASMA (FFP), 2 UNITS OF PLATELETS, AND 2 UNITS OF CRYO WERE GIVEN INTRA-OPERATIVELY. THREE UNITS OF PACKED RED BLOOD CELLS, 1 CRYO, AND 1 UNIT OF PLATELETS WERE GIVEN POST-OPERATIVELY. PRESSOR SUPPORT WAS REQUIRED. A SWAN CATHETER PLACED. THE EJECTION FRACTION MEASURED 20%. THE DAY FOLLOWING THE VALVE IMPLANT A TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) WAS PERFORMED AND BIVENTRICULAR HEART FAILURE WAS IDENTIFIED WITH AN EJECTION FRACTION LESS THAN 15%. CARDIOGENIC SHOCK WAS NOTED AND A NEED FOR VASOPRESSOR MEDICATION. AS RESULT, AN IMPELLA DEVICE REPLACED THE BALLOON PUMP AND VENOARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION (VA ECMO) WAS PLACED. THE BIVENTRICULAR HEART FAILURE WAS REPORTED AS RELATED TO THE SURGICAL VALVE IMPLANT PROCEDURE. TWO DAYS FOLLOWING THE VALVE IMPLANT THE EXPECTED BLOOD LOSS FROM THE VALVE INTERVENTION WAS GREATER THAN EXPECTED. MULTIPLE HEMOGLOBIN DRAWS RANGED FROM 7.1 TO 11.2. ON THIS DATE, THE HEMOGLOBIN MEASURED 8. AS RESULT, 1 UNIT OF PACKED RED BLOOD CELLS WAS TRANSFUSED. CONTINUOUS RENAL REPLACEMENT THERAPY (CRRT) WAS ALSO REQUIRED. THREEDAYS FOLLOWING THE VALVE REPLACEMENT A CHEST ULTRASOUND IDENTIFIED A LEFT PLEURAL EFFUSION. TREATMENT WAS NOT REQUIRED FOR THE EFFUSION. THIS SAME DAY AN ECHOCARDIOGRAM SHOWED A DECREASED EJECTION FRACTION. ECMO SUPPORT REQUIRED. THIS WAS REPORTED AS PROBABLY RELATED TO THE SURGICAL VALVE IMPLANT PROCEDURE. THE SEVERE TRANSVALVULAR REGURGITATION EVENT WAS REPORTED RESOLVED WITHOUT SEQUELAE 8 DAYS FOLLOWING THE VALVE REPLACEMENT PROCEDURE. HOWEVER, IT WAS ALSO REPORTED THAT TRANSAMINITIS LEVELS TRENDED DOWN. LIVER SHOCK OCCURRED SECONDARY TO THE CARDIOGENIC SHOCK. THE CARDIOGENIC SHOCK WAS REPORTED AS RELATED TO THE SURGICAL VALVE AND ITS IMPLANT PROCEDURE. ULTIMATELY, THE PATIENT DIED 8 DAYS FOLLOWING THE TRANSCATHETER VALVE REPLACEMENT. THE DEATH TYPE WAS REPORTED AS CARDIAC. ADDITIONAL INFORMATION RECEIVED INDICATED THAT THE DEATH WAS CAUSED BY THE CARDIOGENIC SHOCK THAT OCCURRED FOLLOWING THE IMPLANT OF THE SURGICAL VALVE AND ITS IMPLANT PROCEDURE TO REPLACE THE TRANSCATHETER AORTIC VALVE (TAV) DUE TO SEVERE AORTIC INSUFFICIENCY.Event Description: IT WAS REPORTED THAT APPROXIMATELY 7 YEARS AND 10 MONTHS FOLLOWING THE IMPLANT OF THE TRANSCATHETER BIOPROSTHETIC VALVE AN ECHOCARDIOGRAM IDENTIFIED SEVERE PARAVALVULAR LEAK (PVL). INTRAVENOUS DIURETIC WAS PROVIDED. SURGICAL WORK-UP WHICH INCLUDED A CORONARY ANGIOGRAPHY WAS REPORTED. PATIENT COMPLICATIONS HAVE NOT BEEN REPORTED AS A RESULT OF THIS EVENT. Manufacturer Narrative: DEVICE ANALYSIS: THE PRODUCT HAS BEEN RETURNED AND THE RESULTS OF THE ANALYSIS ARE PENDING. UPDATED DATA: B2, B5, B6B, B9, H1, H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTED DATA: B5 (FIRST AND SECOND PARAGRAPH) - SEVERE TRANSVALVULAR REGURGITATION WAS REPORTED AND NOT PARAVALVULAR LEAK (PVL). H6 - ANNEX E, ANNEX A, ANNEX G - SEVERE TRANSVALVULAR REGURGITATION WAS REPORTED AND NOT PARAVALVULAR LEAK (PVL). UPDATED DATA: B5 (THIRD PARAGRAPH). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT ANALYSIS: THE VALVE WAS RECEIVED INSIDE A SPECIMEN JAR WITHOUT SOLUTION. VISUAL EVALUATION OF THE VALVE NOTED LEAFLET 1 (L1) AND LEAFLET 3 (L3) APPEARED TO BE IN A CLOSED POSITION WHILE LEAFLET 2 (L2) WAS IN AN OPEN POSITION TOWARDS THE FRAME WALL. LEAFLET DEHISCENCE ON L1 AND L3 APPEARED TO BE ASSOCIATED WITH INCOMPLETE LEAFLET COAPTATION. ALL LEAFLETS WERE TAN AND FLEXIBLE. LEAFLET DEHISCENCE (APPROXIMATELY 10 MM) WAS OBSERVED ON L1, ALONG THE SUPERIOR COAPTIVE JUNCTION OF COMMISSURE 3-1. MANUFACTURING SUTURES ON THE MARGIN OF ATTACHMENT APPEARED INTACT. A LARGE FOLD WAS NOTED ON THE BELLY OF THE LEAFLET. A SMALL DEHISCENCE (APPROXIMATELY 2 MILLIMETERS (MM)) WAS NOTED ON THE FREE MARGIN OF L1, APPROACHING COMMISSURE 1-2. LEAFLET DEHISCENCE (APPROXIMATELY 7MM) WAS OBSERVED ON L2, ALONG THE INFERIOR COAPTIVE AREA BETWEEN LEAFLET 2 AND LEAFLET 1. THE LEAFLET APPEARED TO REMAIN ATTACHED TO THE COMMISSURE. THICK PANNUS EXTENDED ONTO THE LEAFLET OF THE SHOULDER. LEAFLET DEHISCENCE (APPROXIMATELY 12MM) WAS NOTED ON L3, ALONG THE SUPERIOR COAPTIVE JUNCTION OF COMMISSURE 2-3. TISSUE ALONG THE DEHISCENCE APPEARED TEXTURED. MANUFACTURING SUTURES ON THE MARGIN OF ATTACHMENT APPEARED INTACT. A SMALL FOLD WAS NOTED ON THE BELLY OF THE LEAFLET. TAN PANNUS COVERED COMMISSURE 3-1 AND THERE WAS THE INABILITY TO ASSESS THE CONDITION DUE TO THE HOST TISSUE GROWTH. THICK PANNUS ENCAPSULATED COMMISSURE 1-2 AND COMMISSURE 2-3 AND THERE WAS THE INABILITY TO ASSESS THE CONDITION DUE TO HOST TISSUE GROWTH. MULTIPLE BROKEN SUTURES WERE NOTED ALONG THE OUTER VALVE SKIRT WHERE THE HOST TISSUE GROWTH WAS PRESENT. NO DAMAGE SUCH AS KINKS OR BENDS WERE OBSERVED ON THE FRAME. THICK, TAN PANNUS WAS ADHERED TO THE OUTFLOW FRAME OF L2 AND L3, EXTENDING AND COVERING THE FRONT AND BACK OF ALL COMMISSURES. GLISTENING TAN PANNUS REMAINED ATTACHED TO THE OUTSIDE OF THE FRAME EXTENDING TO THE OUTFLOW MARGIN OF ATTACHMENTS OF ALL LEAFLETS. THICK, TAN PANNUS LINED THE INFLOW CROWN EXTENDING INTO THE INNER LUMEN. AN UNKNOWN AMOUNT OF PANNUS MAY HAVE BEEN REMOVED DURING EXPLANT. RADIOGRAPHY DID NOT REVEAL CALCIFICATION OR FRAME FRACTURES ON THE RETURNED VALVE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: D4: UDI H6: ANNEX D. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21562537&pc=DSQ | 2916596-2025-01441 | 2025-02-16 05:00 | Death | THORATEC CORPORATION | 2025-03-10 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Electrical Power Problem; Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT ARRESTED AND WAS PRONOUNCED DECEASED ON THE SCENE. LOG FILES WERE SUBMITTED FOR REVIEW AND CAPTURED SEVERAL LOW FLOW EVENTS STARTING ON (B)(6) 2025 FROM 1618 THROUGH 1901, WITH FLOW NOTED AS LOW AS 1.5 LITERS PER MINUTE (LPM). MORE LOW FLOW EVENTS WERE NOTED LATER ON FROM 1938 THROUGH 2309 THEN THE DRIVELINE WAS DISCONNECTED AT 2309. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: SECTION B1: TYPE OF REPORT CORRECTED. SECTION B3: DATE OF EVENT AND DATE OF DEATH CORRECTED. SECTION H6: MEDICAL DEVICE PROBLEM CODE CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: REVIEW OF THE SUBMITTED LOG FILES CONFIRMED THE LOW FLOW ALARMS AND THE DRIVELINE DISCONNECTIONS. A SPECIFIC CAUSE FOR THESE EVENTS COULD NOT CONCLUSIVELY BE DETERMINED THROUGH THIS EVALUATION. ADDITIONALLY, A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE CONTROLLER EVENT LOG FILE CONTAINED EVENTS FROM (B)(6) 2025 THROUGH (B)(6) 2025. THE FILE CAPTURED INTERMITTENT LOW FLOW ALARMS ON (B)(6) 2025. A SUSTAINED LOW FLOW ALARM WAS CAPTURED ON (B)(6) 2025 UNTIL THE DRIVELINE WAS DISCONNECTED ON THAT DAY, CAUSING THE PUMP TO STOP. THE DRIVELINE WAS SUBSEQUENTLY RECONNECTED ON (B)(6) 2025 AND THE PUMP RAMPED BACK UP TO THE FIXED SPEED WITHOUT ISSUE. FOLLOWING THE RECONNECTION, A SUSTAINED LOW FLOW ALARM WAS CAPTURED UNTIL THE DRIVELINE WAS DISCONNECTED AGAIN AND THE CONTROLLER WAS SHUTDOWN ON (B)(6) 2025. THE CONTROLLER WAS TURNED ON BRIEFLY ON (B)(6) 2025 LIKELY TO DOWNLOAD LOG FILES FOR REVIEW. THE DRIVELINE WAS NOT CONNECTED TO THE CONTROLLER AT THIS TIME. NO OTHER NOTABLE EVENTS OR ALARMS WERE CAPTURED. THE DEVICE WAS NOT RETURNED FOR EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) AND THE PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, LISTS DEATH AS POTENTIAL ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 2 OF THE IFU AND SECTION 2 OF THE PATIENT HANDBOOK INSTRUCT THE USER TO CHECK THE SYSTEM CONTROLLER DRIVELINE CONNECTOR TO CONFIRM THAT THE DRIVELINE IS SECURELY INSERTED IN THE SOCKET. IF THE DRIVELINE DISCONNECTS FROM THE SYSTEM CONTROLLER, THE PUMP STOPS. IF THE DRIVELINE DISCONNECTS FROM THE SYSTEM CONTROLLER, PROMPTLY RECONNECT IT TO RESUME PUMP OPERATION. THE PATIENT HANDBOOK ALSO EXPLAINS THAT THE PUMP CANNOT RUN WITHOUT POWER. ADDITIONALLY, SECTION 2 OF THE IFU, UNDER "CONNECTING THE DRIVELINE TO THE SYSTEM CONTROLLER", PROVIDES INSTRUCTIONS ON CONNECTING/DISCONNECTING THE DRIVELINE TO/FROM THE SYSTEM CONTROLLER AND MAKING SURE THAT IT IS FULLY AND PROPERLY INSERTED INTO THE SYSTEM CONTROLLER SOCKET. THE IFU AND PATIENT HANDBOOK CONTAIN INFORMATION ON ALL SYSTEM CONTROLLER ALARMS, INCLUDING DRIVELINE DISCONNECT ALARMS, AS WELL AS THE PROPER ACTIONS ASSOCIATED WITH THEM. THESE DOCUMENTS ALSO WARN OF EVENTS THAT MAY CAUSE THE PUMP TO STOP AND HOW TO PREVENT PUMP STOPS FROM OCCURRING. ADDITIONALLY, THE ¿HANDLING EMERGENCIES¿ SECTION OF THE PATIENT HANDBOOK LISTS EXAMPLES OF EMERGENCIES AND WHAT ACTIONS TO TAKE. FURTHERMORE, THE PATIENT HANDBOOK CAUTIONS THE USER TO CALL THEIR HOSPITAL CONTACTS IF THEY THINK, FOR ANY REASON, ANY PORTION OF THEIR EQUIPMENT IS NOT FUNCTIONING AS USUAL, IS BROKEN, OR THEY ARE UNCOMFORTABLE WITH THE OPERATION OF THE EQUIPMENT. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21563406&pc=DSQ | 2916596-2025-01595 | 2024-03-01 05:00 | Death | THORATEC CORPORATION | 2025-03-10 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Sepsis; No Clinical Signs, Symptoms or Conditions | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE "OUTCOMES OF HEARTMATE 3 IN PEDIATRIC PATIENTS WITH END-STAGE HEART FAILURE: A SINGLE-CENTER PRELIMINARY EXPERIENCE FROM TURKEY" THAT HEARTMATE 3 (HM3) MAY BE ASSOCIATED WITH PLEURAL EFFUSION, CARDIAC TAMPONADE, POLYURIA, POSITIVE BLOOD CULTURES, THROMBUS, INFECTION, RIGHT VENTRICULAR FAILURE AND DEATH. PATIENT #3 EXPERIENCED AN ICU STAY OF 28 DAYS, PLEURAL EFFUSION, RIGHT HEART FAILURE, LIVER FAILURE, KIDNEY FAILURE, MYOCYTE HYPERTROPHY, AND DEATH DUE TO SEPSIS.Event Description: THIS EVENT WILL BE REPORTED UNDER MFR#: 2916596-2025-01398. Manufacturer Narrative: CORRECTED DATA: THIS EVENT WILL BE REPORTED UNDER MFR#: 2916596-2025-01398. MANUFACTURER'S INVESTIGATION CONCLUSION: DEVICE RELATED ISSUES ASSOCIATED WITH HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), ARE COVERED UNDER (B)(6). IT WAS DETERMINED THAT THE REPORTED INFORMATION DID NOT MEET THE REQUIREMENTS OF A COMPLAINT; HOWEVER, THIS RECORD WILL REMAIN A COMPLAINT AS AN INITIAL MDR (MEDICAL DEVICE REPORT) HAS ALREADY BEEN SUBMITTED PRIOR TO THIS ADDITIONAL INFORMATION BEING PROVIDED. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. G IS CURRENTLY AVAILABLE. ALTHOUGH NO DEVICE RELATED ISSUES WERE REPORTED BY THE ACCOUNT, SECTION 1 OF THE IFU ¿INTRODUCTION¿ LISTS POTENTIAL ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. THE SERIAL NUMBER OR OTHER IDENTIFYING INFORMATION OF THE PRODUCT WAS NOT REPORTED AND WAS NOT ABLE TO BE DETERMINED DURING THE INVESTIGATION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SECTION A, D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. SECTION D4: DEVICE SERIAL NUMBER AND LOT NUMBER WERE NOT PROVIDED, AND EXPIRATION DATE AND MANUFACTURE DATE (H4) CANNOT BE DETERMINED. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION. DOGAN, E., ULGER TUTAR, Z., TUNCER, O. N., LEVENT, R. E., ENGIN, Ç., YAGDI, T., ATAY, Y., & ÖZBARAN, M. (2024). OUTCOMES OF HEARTMATE 3 IN PEDIATRIC PATIENTS WITH END-STAGE HEART FAILURE: A SINGLE-CENTER PRELIMINARY EXPERIENCE FROM TURKEY. FRONTIERS IN CARDIOVASCULAR MEDICINE, 11. HTTPS://DOI.ORG/10.3389/FCVM.2024.1472663. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21563593&pc=MHY | 3006630150-2025-01386 | 2025-01-16 05:00 | Death | BOSTON SCIENTIFIC NEUROMODULATION CORPORATION | 2025-03-10 04:00 | MHY | VERCISE GENUS | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P150031 |  | 1 | Event Description: IT WAS REPORTED THAT THE DEEP BRAIN STIMULATION (DBS) PATIENT PASSED AWAY FOR AN UNKNOWN REASON. NO ADDITIONAL INFORMATION WAS AVAILABLE. Manufacturer Narrative: GOOD FAITH EFFORT ATTEMPTS WERE MADE TO TRY AND RETRIEVE ADDITIONAL DETAILS REGARDING THE REPORTED EVENT, BUT FURTHER INFORMATION WAS UNABLE TO BE OBTAINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21563886&pc=DSQ | 2916596-2025-01188 | 2024-09-25 04:00 | Death | THORATEC CORPORATION | 2025-03-10 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE ¿THE IMPACT OF BLOOD TYPE AND BODY SIZE ON SUCCESSFUL BRIDGING TO HEART TRANSPLANTATION USING INTRA-AORTIC BALLOON PUMP OR SURGICALLY IMPLANTED MICROAXIAL LEFT VENTRICULAR ASSIST DEVICE¿ THAT HEARTMATE 3 (HM3) MAY BE ASSOCIATED WITH DEATH. THE ARTICLE COMPARED A BRIDGING STRATEGY USING EITHER THE IMPELLA 5.0/5.5 (¿IMPELLA¿) OR INTRA-AORTIC BALLOON PUMP WITH A DURABLE LEFT VENTRICULAR ASSIST DEVICE (LVAD) IN PATIENTS WITH BLOOD TYPE O STRATIFIED BY BODY HABITUS. THE UNITED NETWORK FOR ORGAN SHARING REGISTRY WAS QUERIED FOR ADULTS LISTED FOR DE NOVO HEART TRANSPLANTATION (WITHOUT DIALYSIS) SUPPORTED BY THE IMPELLA OR THE HEARTMATE 3 (HM3) LVAD. THERE WERE 1402 LVAD PATIENTS HAD AN AVERAGE AGE OF 56 AND AVERAGE BMI OF 29.6. 102 PATIENTS HAD HISTORY OF STROKE, 719 WERE PRIOR SMOKERS, 457 HAD DIABETES, 1019 WERE IMPLANTED WITH AN IMPLANTABLE CARDIOVERTER-DEFIBRILLATOR (ICD). THE PRIMARY OUTCOME WAS 1-YEAR POSTLISTING SURVIVAL, DEFINED AS THE SUM OF WAITLIST TIME AND POST¿HEART TRANSPLANT SURVIVAL TIME IF THE CANDIDATE UNDERWENT HEART TRANSPLANTATION. 2942 CANDIDATES WERE INCLUDED (IMPELLA: 214: 1326; HEARTMATE 3: 1402). LISTING WITH THE IMPELLA OR INTRA-AORTIC BALLOON PUMP WAS ASSOCIATED WITH WORSE POSTLISTING SURVIVAL COMPARED WITH THE HEARTMATE 3 LVAD IN TYPE O CANDIDATES (IMPELLA: HAZARD RATIO, 2.90 [95% CI, 1.48-5.67], P ¼ .002; INTRA-AORTIC BALLOON PUMP: HAZARD RATIO, 2.42 [95% CI, 1.59-3.68], P < .001) BUT LESS SO IN NON¿TYPE O CANDIDATES. FURTHER ANALYSIS OF TYPE O CANDIDATES DEMONSTRATED THAT THE IMPELLA AND INTRA-AORTIC BALLOON PUMP WERE ASSOCIATED WITH A LOWER RISK OF POSTLISTING MORTALITY AMONG THOSE WITH NORMAL HEIGHT (25-75TH PERCENTILE) AND NONOBESE (BODY MASS INDEX<.001). LVAD PATIENTS HAD A SURVIVABILITY OF 86.1% (95% CI, 80.1-90.5), 88.5% (95% CI, 86.6-90.2), AND 91.2% (95% CI, 89.5-92.7), RESPECTIVELY (LOG-RANK P ¼ .007). BLOOD TYPE O CANDIDATES LISTED WITH THE IMPELLA OR INTRA-AORTIC BALLOON PUMP ARE AT INCREASED RISK OF POST LISTING MORTALITY COMPARED WITH A DURABLE LVAD IN THE CURRENT ORGAN ALLOCATION ERA. THESE EFFECTS ARE DIMINISHED AMONG THOSE WITH NORMAL HEIGHT AND BODY HABITUS. Manufacturer Narrative: SECTIONS A AND D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. SECTION B2: DATE OF DEATH WAS ESTIMATED AS DATE OF EVENT. SECTION B3: DATE OF EVENT HAS BEEN ENTERED AS THE DATE OF PUBLICATION (25SEP2024) AS THE DATE OF DATA COLLECTION WAS NOT PROVIDED. SECTION D4: DEVICE SERIAL NUMBER AND LOT NUMBER WERE NOT PROVIDED, AND EXPIRATION DATE AND MANUFACTURE DATE (H4) CANNOT BE DETERMINED. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION. SECTION E1: REPORTER EMAIL WAS NOT AVAILABLE. AUTHOR INFORMATION: NORDAN, T., KESHK, K., HACKMANN, A., GIVERTZ, M. M., MEHRA, M. R., CUNNINGHAM, M. J., ITOH, A. THE IMPACT OF BLOOD TYPE AND BODY SIZE ON SUCCESSFUL BRIDGING TO HEART TRANSPLANTATION USING INTRA-AORTIC BALLOON PUMP OR SURGICALLY IMPLANTED MICROAXIAL LEFT VENTRICULAR ASSIST DEVICE. (2024). JOURNAL OF THORACIC AND CARDIOVASCULAR SURGERY. HTTPS://DOI.ORG/10.1016/J.JTCVS.2024.10.013. (B)(6). MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED EVENTS COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE SERIAL NUMBER OR OTHER IDENTIFYING INFORMATION OF THE PRODUCT WAS NOT REPORTED AND WAS NOT ABLE TO BE DETERMINED DURING THE INVESTIGATION. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU) AND THE HEARTMATE 3 PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21563892&pc=OAE | 2029046-2025-00733 | 2025-02-10 05:00 | Death | BIOSENSE WEBSTER INC | 2025-03-10 04:00 | OAE | QDOT MICRO¿ CATHETER | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | P210027 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN ATRIAL FIBRILLATION ABLATION WITH A QDOT MICRO¿ CATHETER AND THE PATIENT EXPERIENCED CARDIAC ARREST THAT RESULTED IN DEATH. THE PATIENT DIED OF CARDIAC ARREST DUE TO "AN ELECTROLYTE IMBALANCE". THE PROCEDURE TOOK PLACE THE DAY BEFORE. THE PATIENT WAS STABLE DURING THE PROCEDURE AND ON TRANSFER POST PROCEDURE. IT WAS REPORTED THAT THE PROCEDURE WENT SMOOTHLY. THE ONLY EVENT DURING THE PROCEDURE WAS A QUESTION OF ACT (ACTIVATED CLOTTING TIME) VALUES THAT INVOLVED CHECKING THE ACT ON A SEPARATE MACHINE. IT IS NOT KNOWN IF THE ACT WAS TOO HIGH OR TOO LOW. THE FOLLOWING BWI CATHETERS WERE USED DURING THE PROCEDURE: QDOT MICRO ABLATION CATHETER, SOUNDSTAR IMAGING CATHETER, DECANAV CS CATHETER, PENTARAY MAPPING CATHETER AND A VIZIGO SHEATH. Manufacturer Narrative: D4: UDI: AS THE LOT NUMBER FOR THE DEVICE INVOLVED IN THE EVENT WAS NOT PROVIDED, THE FULL UDI IS CURRENTLY NOT AVAILABLE. THE DEVICE HAS BEEN REPORTED AS DISCARDED; THEREFORE NO PRODUCT INVESTIGATION CAN BE PERFORMED, AND THE CUSTOMER COMPLAINT CANNOT BE CONFIRMED. AN EVALUATION OF THE MANUFACTURING RECORD COULD NOT BE PERFORMED AS THE REQUIRED PRODUCT IDENTIFICATION NUMBER WAS NOT PROVIDED TO COMPLETE THE EVALUATION. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, BIOSENSE WEBSTER, INC. OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. MANUFACTURER'S REFERENCE NUMBER: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21565409&pc=QAD | 3006845464-2025-00002 | 2025-02-10 05:00 | Death | ILLUMINOSS MEDICAL INC | 2025-03-10 04:00 | QAD | PHOTODYNAMIC BONE STABILIZATION SYSTEM | Adverse Event Without Identified Device or Use Problem | Embolism/Embolus | K181228 |  | 1 | Event Description: A PATIENT PASSED AWAY DURING THE CLOSING OF AN ILLUMINOSS PROCEDURE, AFTER THE CURING PROCESS WAS COMPLETED. THEY WERE NOT ABLE TO REVIVE THE PATIENT. THE ANATOMY TREATED WAS THE PELVIS, BOTH SIDES BEING TREATED. THE PATIENT WAS "THOUGHT TO HAVE HAD AN EMBOLIC EVENT" PER THE USER. THE PATIENT HAD COPD.Event Description: A PATIENT PASSED AWAY DURING THE CLOSING OF AN ILLUMINOSS PROCEDURE, AFTER THE CURING PROCESS WAS COMPLETED. THEY WERE NOT ABLE TO REVIVE THE PATIENT. THE ANATOMY TREATED WAS THE PELVIS, BOTH SIDES BEING TREATED. THE PATIENT WAS "THOUGHT TO HAVE HAD AN EMBOLIC EVENT" PER THE USER. THE PATIENT HAD COPD. Manufacturer Narrative: AT THE TIME OF THIS MDR SUBMISSION, THE INVESTIGATION IS STILL UNDERWAY. A FOLLOW-UP MDR WILL BE SUBMITTED WITH THE CONCLUSIONS OF THE INVESTIGATION. DHR REVIEW: THE DHRS FOR THE TWO LOTS OF IMPLANTS USED WERE REVIEWED AND FOUND TO BE IN SPECIFICATION AT THE TIME OF MANUFACTURE AND RELEASE. IFU 900356_X STATES THAT RISKS INCLUDE THROMBOEMBOLIC EVENT OR FAT EMBOLISM (BLOOD CLOT, FAT, OR OTHER MATERIAL THAT COULD RESULT IN ORGAN DAMAGE OR FAILURE). A CALL WAS HELD WITH THE TREATING PHYSICIAN TO DISCUSS THE EVENT. THE 73-YEAR-OLD PATIENT HAD A FRAGILITY FRACTURE IN PELVIS WITH A HISTORY OF SMOKING WITH COMORBIDITIES WHICH INCLUDED COPD, EMPHYSEMA, OSTEOPOROSIS, AND RESPIRATORY DISEASES. THE PHYSICIAN TREATED THE SACRUM AND BOTH ILIA, USED 4 BALLOONS BILATERALLY. AT THE COMPLETION OF THE PROCEDURE DURING CLOSING, THE PATIENT BECAME HYPOXIA AND HYPOTENSIVE. THEY COVERED THE WOUNDS AND RAN A CODE FOR A PRESUMED DIAGNOSIS OF A FAT EMBOLISM. THERE WERE DIFFICULTIES DUE TO THE PATIENT'S RESPIRATORY COMORBIDITIES AND THEY WERE NOT ABLE TO STABILIZE THE PATIENT. AFTER 40 MINUTES OF THE CODE, THE CARDIAC PHYSICIAN DETERMINED THAT GIVEN THE LENGTH OF TIME OF HYPOXIA THE PATIENT'S PROGNOSIS WAS NOT GOOD, AND THE PATIENT WAS PRONOUNCED DEAD. THE USER ATTRIBUTED THE CAUSE OF THE DECLINE AND DEATH TO FAT EMBOLISM WHICH WAS SEEN IN THE ECHOCARDIOGRAM. THE PHYSICIAN IN THIS CASE, WAS A VERY EXPERIENCED ILLUMINOSS USER SO IT IS UNLIKELY THAT ANY USER ERROR CONTRIBUTED TO THE EVENT. Manufacturer Narrative: AT THE TIME OF THE INITIAL MDR SUBMISSION, THE INVESTIGATION WAS STILL UNDERWAY. THIS UPDATE IS BEING PROVIDED WITH THE CONCLUSION FROM THE INVESTIGATION. INVESTIGATION: A CALL WAS HELD WITH THE TREATING PHYSICIAN TO DISCUSS THE EVENT. THE 73 YEAR OLD PATIENT HAD A FRAGILITY FRACTURE IN PELVIS WITH A HISTORY OF SMOKING WITH COMORBIDITIES WHICH INCLUDED COPD, EMPHYSEMA, OSTEOPOROSIS, AND RESPIRATORY DISEASES. THE PHYSICIAN TREATED THE SACRUM AND BOTH ILIA, USED 4 BALLOONS BILATERALLY. AT THE COMPLETION OF THE PROCEDURE DURING CLOSING, THE PATIENT BECAME HYPOXIA AND HYPOTENSIVE. THEY COVERED THE WOUNDS AND RAN A CODE FOR A PRESUMED DIAGNOSIS OF A FAT EMBOLISM. THERE WERE DIFFICULTIES DUE TO THE PATIENT'S RESPIRATORY COMORBIDITIES AND THEY WERE NOT ABLE TO STABILIZE THE PATIENT. AFTER 40 MINUTES OF THE CODE, THE CARDIAC PHYSICIAN DETERMINED THAT GIVEN THE LENGTH OF TIME OF HYPOXIA THE PATIENT'S PROGNOSIS WAS NOT GOOD, AND THE PATIENT WAS PRONOUNCED DEAD. THE USER ATTRIBUTED THE CAUSE OF THE DECLINE AND DEATH TO FAT EMBOLISM WHICH WAS SEEN IN THE ECHOCARDIOGRAM. THE PHYSICIAN IN THIS CASE, WAS A VERY EXPERIENCED ILLUMINOSS USER SO IT IS UNLIKELY THAT ANY USER ERROR CONTRIBUTED TO THE EVENT. A MEDICAL OVERSIGHT REVIEW MEETING WAS HELD WITH A MEDICAL REVIEWER, REGULATORY, QUALITY, AND ENGINEERING. THE INFORMATION FROM THE COMPLAINT WAS REVIEWED AND THE MEDICAL REVIEWER STATED THAT GIVEN THE TIMING AND REPORT FROM THE ECHOCARDIOGRAM OF THE FAT EMBOLISM, AND THAT FROM THE LITERATURE A FAT EMBOLISM IS A KNOWN RISK OF THE ILLUMINOSS PROCEDURE, THEY AGREE WITH THE USES CONCLUSIONS. THE MEDICAL REVIEWER ALSO OBSERVED THAT THE PATIENT HAD MANY RESPIRATORY COMPLICATIONS WHICH WOULD INCREASE THE RISK OF EXPERIENCING A FAT EMBOLISM. DHR REVIEW: THE DHRS FOR THE TWO LOTS OF IMPLANTS USED WERE REVIEWED AND FOUND TO BE IN SPECIFICATION AT THE TIME OF MANUFACTURE AND RELEASE. IFU REVIEW: IFU 900356_X STATES THAT RISKS INCLUDE THROMBOEMBOLIC EVENT OR FAT EMBOLISM (BLOOD CLOT, FAT, OR OTHER MATERIAL THAT COULD RESULT IN ORGAN DAMAGE OR FAILURE). POTENTIAL FOR USER ERROR: THE USER IS AN EXPERIENCED ILLUMINOSS USER, SO IT IS UNLIKELY THAT ANY USER ERROR CONTRIBUTED TO THE EVENT. THE USER PERFORMED REAMING AS INSTRUCTED IN THE SURGICAL TECHNIQUE GUIDE 900598_A FOR THE PELVIS. CONCLUSION: THE CAUSE OF PATIENT DEATH WAS A FAT EMBOLISM. THIS PATIENT HAD SEVERAL COMORBIDITIES WHICH INCREASE THE RISK OF FAT EMBOLISM AND THE RISK OF FAT EMBOLISM IS INHERENT TO AN ORTHOPEDIC SURGERY INVOLVING ANESTHESIA.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21566832&pc=OZD | 1220648-2025-26609 | 2025-02-12 05:00 | Death | ABIOMED, INC. | 2025-03-10 04:00 | OZD | IMPELLA 5.5 | Increase in Pressure | Hemorrhage/Bleeding; Perforation of Vessels | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT DURING PATIENT TRANSPORT IMPELLA 5.5 PUMP, A PURGE PRESSURE HIGH / PURGE SYSTEM BLOCKED ALARM APPEARED ON THE CONSOLE (AIC). THE PATIENT WAS TAKEN TO THE OPERATING ROOM WHERE A VENTRICLE RUPTURE WAS OBSERVED. THE PHYSICIAN ATTEMPTED TO CHANGE THE PUMP, BUT THE PATIENT DIED DUE TO SEVERE BLEEDING. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA 5.5 WITH SMARTASSIST STATES THE FOLLOWING: SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA CATHETER ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP. CLINICIANS SHOULD MONITOR MOTOR CURRENT AND CONSIDER REPLACING THE IMPELLA CATHETER WHENEVER A RISE IN MOTOR CURRENT IS SEEN.¿ SECTION: TROUBLESHOOTING THE PURGE SYSTEM ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿ Manufacturer Narrative: A.2 ADDITIONAL INFORMATION WAS RECEIVED CONFIRMING THAT THE PATIENT WAS AN ADULT OVER THE AGE OF 50. A.3 REVISED AS SELECTION WAS INADVERTENTLY OMITTED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26609. B.3 REVISED DATE OF EVENT AS IT WAS PREVIOUSLY REPORTED INCORRECTLY ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26609. E.1 REVISED NAME PREFIX/TITLE, FIRST NAME AND LAST NAME AS THEY WERE INADVERTENTLY OMITTED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26609. E.4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26609 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. G.1 REVISED REPORTING CONTACT FACILITY NAME, ADDRESS, AND FAX NUMBER AS THEY WERE INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26609. Manufacturer Narrative: THE INVESTIGATION OF HIGH PURGE PRESSURE AND VENTRICLE PERFORATION HAS BEEN COMPLETED. THE PUMP WAS NOT RETURNED FOR INVESTIGATION. THE DATA LOG SHOWS A SUDDEN RISE IN PURGE PRESSURE, ACCOMPANIED BY A MOTOR CURRENT SPIKE AND SATURATED PUMP FLOW. THE PURGE PRESSURE AND PUMP FLOW VALUES RETURNED TO NORMAL AFTER 1 HOUR. ACCORDING TO THE CLINICAL REPORT, A HIGH PURGE PRESSURE ALARM WAS OBSERVED DURING TRANSPORT. ADDITIONALLY, A VENTRICULAR RUPTURE WAS OBSERVED IN THE OPERATING ROOM. THE CAUSE OF THE HIGH PURGE PRESSURE ISSUE WAS MOST LIKELY BIOMATERIAL, BASED ON DATA LOG REVIEW. THE CAUSE OF THE VENTRICLE PERFORATION ISSUE WAS NOT DETERMINED SINCE PRODUCT WAS NOT RETURNED AND SUFFICIENT CLINICAL INFORMATION WAS NOT PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21566876&pc=MCX | 2124215-2025-14610 | 2025-02-18 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-03-10 04:00 | MCX | ROTAWIRE AND WIRECLIP TORQUER | Adverse Event Without Identified Device or Use Problem | Ischemia; Myocardial Infarction |  |  | 1 | Event Description: IT WAS REPORTED A PATIENT DEATH OCCURRED. DURING THE PROCEDURE IN THE RIGHT CORONARY ARTERY (RCA), OSTIUM SUBTOTAL, THE LESION WAS MODERATELY TORTUOUS, SEVERELY CALCIFIED AND 90% STENOSED. A 6F GUIDING CATHETER WAS SELECTED FOR TREATMENT , AND THEN REPLACED WITH A 7F CATHETER. INITIALLY A NON-BOSTON SCIENTIFIC WIRE WAS USED, THEN CHANGED TO A ROTAWIRE FLOPPY VIA A MICROCATHETER, ALONG WITH A 1.75 MM ROTAPRO. FOUR RUNS OF ROTATIONAL ATHERECTOMY WERE CONDUCTED, EACH ABOUT 20 SECONDS, WITH A MAXIMUM RPM DROP OF 15 K. HOWEVER, NO REFLOW OCCURRED, LEADING TO BALLOON ANGIOPLASTY AND IMPLANTATION OF TWO STENTS. THERE WAS NO ANTEGRADE FLOW OBSERVED IN THE CORONARY ARTERY (NO TIMI FLOW), CAUSING AV BLOCK, AND REQUIRING A TEMPORARY PACEMAKER. COMPLICATIONS THAT WERE ENCOUNTERED INCLUDED MYOCARDIAL INFARCTION/ISCHEMIA AND FURTHER INTERVENTIONS INCLUDED ADMINISTRATION OF ISOPREL, AGOSTAT, KATECHOMALIN, ATROPIN, AND USE OF A TEMPORARY PACEMAKER. FOLLOWING THE PROCEDURE, THE PATIENT WAS TRANSFERRED TO THE CRITICAL CARE UNIT (CCU). THOUGH NO DEVICE MALFUNCTIONS OCCURRED, THE PHYSICIAN IS UNSURE IF THE DEVICE OR PROCEDURE CONTRIBUTED TO THE COMPLICATIONS. HOWEVER, FURTHER INFORMATION CONFIRMED DEATH OF THE PATIENT ON AN UNKNOWN DATE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21566975&pc=DQY | 2029214-2025-00613 | 2024-09-05 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-03-10 04:00 | DQY | NAVIEN CATHETER | Adverse Event Without Identified Device or Use Problem | Hemorrhagic Stroke |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING: THE USE OF FLOW DIVERTERS (FDS) FOR TREATING ANEURYSMS BEYOND THE CIRCLE OF WILLIS, EMPHASIZING A LACK OF EVIDENCE ON THEIR SAFETY AND EFFECTIVENESS IN SUCH CASES. THE SILK VISTA BABY (SVB) FD HAS BEEN USED FOR UNRUPTURED, LARGE, AND COMPLEX ANEURYSMS, BUT DATA ON LONG-TERM OUTCOMES FOR BIFURCATION AND BEYOND THE CIRCLE OF WILLIS ANEURYSMS IS LIMITED. (PAGES 846-847) THE TIME FRAME OF THIS STUDY WAS: OCTOBER 2018 TO DECEMBER 2022. (PAGE 848) MULTIPLE MANUFACTURER¿S DEVICES WERE USED IN THE STUDY POPULATION.: DEVICES FROM MULTIPLE MANUFACTURERS WERE UTILIZED, INCLUDING THE SILK VISTA BABY (SVB) FD FROM BALT, AND OTHER DEVICES FROM PENUMBRA INC. AND MICROVENTION-TERUMO. (PAGE 848) THE FOLLOWING MEDTRONIC DEVICES WERE USED: NAVIEN 058 CATHETER. IT WAS NOT STATED HOW MANY PROCEDURES USED A NAVIEN CATHETER. (PAGE 848) DEATHS OCCURRED IN THE STUDY POPULATION.: THERE WAS ONE DEATH DUE TO COMPLICATIONS FROM SUBARACHNOID HEMORRHAGE (SAH). (PAGE 849) AMONG PATIENT ADVERSE EVENTS INCLUDED: - INTRAPROCEDURAL COMPLICATIONS ACCOUNTED FOR 3 (5.3%) WHERE ONE PATIENT HAD A DISSECTING HEMATOMA OF THE RECTUS ABDOMINAL MUSCLE, AND ONE PATIENT HAD A RETROPERITONEAL HEMATOMA, WHICH LED TO CLINICAL INSTABILITY, ULTIMATELY WITH GOOD PATIENT OUTCOME AND DISCHARGE MRS OF 2. THE THIRD PATIENT HAD A DISSECTION OF THE ORIGIN OF THE POSTERIOR INFERIOR CEREBELLAR ANEURYSM (PICA) ARTERY WITH AN ISCHEMIC EVENT ON THE FOLLOW-UP MRI, HOWEVER WITH NO CLINICAL CORRELATION. - THERE WAS 1 COMPLICATION OF ISCHEMIA ARISING FROM IN-STENT THROMBOSIS (1.8%) ON THE THIRD DAY AFTER THE PROCEDURE. THE PATIENT WAS DISCHARGED AND RETURNED WITH HEADACHES, VISUAL SCOTOMAS, AND VOMITING, WITH NO HEMORRHAGE OF ADMISSION PLAIN CT. AN IMMEDIATE CEREBRAL ANGIOGRAM SHOWED AN OCCLUDED PERICALLOSAL STENT; HOWEVER, REVASCULARIZATION WAS NOT ATTEMPTED BECAUSE OF EXCELLENT COLLATERALS AND THE ABSENCE OF MAJOR NEUROLOGICAL SYMPTOMS. THERE WERE ASYMPTOMATIC PUNCTATE DWI-RESTRICTION LESIONS OVER THE RIGHT FRONTAL LOBE, RIGHT CINGULATE GYRUS, AND CORPUS CALLOSUM. THE EPISODE WAS CONSIDERED A SEIZURE BY THE NEUROLOGY TEAM AND THE PATIENT WAS DISCHARGED WITH SEIZURE PROPHYLAXIS. - MINOR COMPLICATIONS ACCOUNTED FOR 3.5% (2/57) (ONE CASE OF ASYMPTOMATIC PICA OCCLUSION WITH GOOD COLLATERALS AND CONSERVATIVE MANAGEMENT), WITH NO HEMORRHAGIC COMPLICATION, AND ONE PRESENTED WITH A LATE TRANSITORY SYMPTOMATIC ISCHEMIA. - ALL PATIENTS DISCHARGED WITH MRS > 2 (3/57, 5.3%) PRESENTED WITH HIGH-GRADE SAH, AND IN-HOSPITAL COMPLICATIONS (DELAYED CEREBRAL ISCHEMIA, AND INFECTIONS). NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID UNK-NV-RFX (UNKNOWN); PRODUCT TYPE: IMPLANT DATE N/A; EXPLANT DATE N/A G2: CITATION: AUTHORS: VASCONCELLOS DE OLIVEIRA SOUZA, N., BENALIA, V. H., ORTEGA MORENO, D. A., LIU, E., CHAN, V., BHARATHA, A., MAROTTA, T. R., SPEARS, J., PEREIRA, V. M. SILK VISTA BABY FLOW DIVERSION BEYOND THE CIRCLE OF WILLIS: A SINGLE-CENTER EXPERIENCE WITH LONG-TERM OUTCOMES. INTERVENTIONAL NEURORADIOLOGY¿: JOURNAL OF PERITHERAPEUTIC N 30(6):846-853 2024. DOI:10.1177/15910199241285504 EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21567018&pc=DQO | 3010909876-2025-01379 | 2025-02-10 05:00 | Death | TRISALUS LIFE SCIENCES | 2025-03-10 04:00 | DQO | TRINAV INFUSION SYSTEM | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | K180677 |  | 1 | Event Description: IT WAS REPORTED THAT A TRINAV CATHETER WAS USED IN A ROUTINE THYROID GOITER EMBOLIZATION PROCEDURE EMPLOYING RADIAL APPROACH VIA THE PERIPHERAL VASCULATURE/INFERIOR THYROID ARTERY. EMBOLIZATION WAS SUCCESSFUL AND CARRIED TO STASIS. POST-EMBOLIZATION ANGIOGRAPHY WAS PERFORMED VIA THE TRINAV CATHETER INSTEAD OF THE THROUGH THE BASE CATHETER IN THE SUBCLAVIAN ARTERY. POST-EMBOLIZATION CONTRAST INJECTION VOLUMES AND FORCES WERE HIGH AND THERE WAS NO FORWARD FLOW/ROOM, DUE TO PRESENCE OF SPHERES, RESULTING IN REFLUX. ARTERY WAS SUPERIMPOSED OVER RIBS AND SPINE MAKING VISUALIZATION OF REFLUX DIFFICULT. ORIGIN OF THE LEFT VERTEBRAL ARTERY WAS DOMINANT (LARGER THAN THE RIGHT) AND IMMEDIATELY PROXIMAL TO THE INFERIOR THYROID ARTERY, RESULTING IN HIGH FLOW TO THE POSTERIOR PORTION OF THE BRAIN. PATIENT EXPERIENCED MULTIPLE SMALL STROKES IDENTIFIED VIA POST PROCEDURE MRI, STARTED ON DUAL-ANTIPLATELET MEDICATIONS (ASPIRIN AND PLAVIX, LOADING DOSES), AND MONITORED FOR 24 HOURS. CT SCAN SHOWED A LARGE AMOUNT OF HEMORRHAGE INTO THE LEFT LATERAL VENTRICLE THAT SPREAD INTO ADJACENT VENTRICLES AND RESULTED IN LOSS OF DIFFERENTIATION OF GREY AND WHITE MATTER. THE PATIENT WAS TRANSITIONED TO HOSPICE CARE THAT DAY AND DIED SHORTLY THEREAFTER. Manufacturer Narrative: NO DEVICE MALFUNCTION OR PERFORMANCE ISSUE ASSERTED OR SUSPECTED BASED ON USER FACILITY REPORT. SUBJECT DEVICE NOT AVAILABLE FOR INVESTIGATION, SCRAPPED BY USER FACILITY AT TIME OF PROCEDURE. DEVICE VARIANT (MODEL NUMBER) AND/OR LOT NUMBER NOT REPORTED BY USER FACILITY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21567623&pc=DSQ | 2916596-2025-01157 | 2022-10-01 04:00 | Death | THORATEC CORPORATION | 2025-03-10 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE "EARLY VENTRICULAR ARRHYTHMIAS AFTER LEFT VENTRICULAR ASSIST DEVICE IMPLANTATION" THAT EARLY VENTRICULAR ARRHYTHMIAS (EVAS) WERE A SIGNIFICANT CONCERN FOLLOWING LEFT VENTRICULAR ASSIST DEVICE (LVAD) IMPLANTATION. A SINGLE CENTER RETROSPECTIVE ANALYSIS WAS CONDUCTED ON PATIENTS WHO UNDERWENT LVAD IMPLANTATION FROM 01OCT2006, TO 01OCT2022. EVA WAS DEFINED AS AN EPISODE OF SUSTAINED VENTRICULAR ARRHYTHMIAS (VA) IDENTIFIED 30 DAYS AFTER LVAD IMPLANTATION. A TOTAL OF 789 PATIENTS UNDERWENT LVAD IMPLANTATION (MEAN AGE OF 62.9 YEARS, HEARTMATE 3 41.4%, DESTINATION THERAPY 43.3%). EVAS HAD OCCURRED IN 100 PATIENTS (12.7%). PATIENTS HAD A HISTORY OF END-STAGE RENAL DISEASE (ODDS RATIO [OR] 5.6, 95% CONFIDENCE INTERVAL [CI] 1.45-21.70), PREOPERATIVE ELECTRICAL STORM (OR 2.82, 95% CI 1.11-7.16), APPROPRIATE IMPLANTABLE CARDIAC DEFIBRILLATOR THERAPY BEFORE IMPLANTATION (OR 2.8, 95% CI 1.26-6.19), AND WERE INDEPENDENTLY ASSOCIATED WITH EVAS. EVA WAS ASSOCIATED WITH DECREASED 30-DAY SURVIVAL (HAZARD RATIO 3.02, 95% CI 1.1-8.3, P = .032). THE MEDIAN NUMBER OF DAYS OF FOLLOW-UP WHILE ON LVAD SUPPORT WAS 1594 DAYS. 55 PATIENTS HAD A 30-DAY MORTALITY, 31 PATIENTS HAD PASSED AWAY DUE TO CARDIOVASCULAR RELATED DEATH, 24 HAD PASSED AWAY DUE TO NON-CARDIOVASCULAR RELATED DEATHS, 173 WERE TRANSPLANTED. THE MEDIAN DAYS TO TRANSPLANT FOR PATIENTS WAS 534 DAYS. CARDIOVASCULAR DEATH HAD OCCURRED MORE FREQUENTLY IN PATIENTS WITH EVAS (57.1% VS 38.3%), INCLUDING A NUMERICALLY GREATER RATE OF DEATH FROM RIGHT VENTRICULAR FAILURE AMONG PATIENTS WITH EVAS (33.3% VS 18.6%). THERE WAS NO DIFFERENCE IN TRANSPLANT-FREE SURVIVAL TIME BETWEEN PATIENTS WITH AND WITHOUT EVAS (HAZARD RATIO 0.82, 95% CI 0.5-1.4, P = .454). IN CONCLUSION, IDENTIFYING PATIENTS AT HIGHER RISK FOR EVAS ALLOWED FOR TARGETED INTERVENTIONS, WHICH MAY HAVE IMPROVED OUTCOMES IN THIS POPULATION. Manufacturer Narrative: SECTIONS A AND D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. SECTION B3: DATE OF EVENT HAS BEEN ENTERED AS (01OCT2022) AS DATA WAS COLLECTED BETWEEN 01OCT2006 AND 01OCT2022. SECTION D4: DEVICE SERIAL NUMBER AND LOT NUMBER WERE NOT PROVIDED, AND EXPIRATION DATE AND MANUFACTURE DATE (H4) CANNOT BE DETERMINED. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION. AUTHOR INFORMATION: OATES, C. P., LAM, P. H., LAWRENCE, L., BIGHAM, G., MEDA, N. S., BINAYA BASYAL, HADADI, C. A., RAO, S. D., HOCKSTEIN, M., SHAH, M., & SHEIKH, F. H. (2023). EARLY VENTRICULAR ARRHYTHMIAS AFTER LEFT VENTRICULAR ASSIST DEVICE IMPLANTATION. JOURNAL OF CARDIAC FAILURE. HTTPS://DOI.ORG/10.1016/J.CARDFAIL.2023.11.018 GEORGETOWN UNIVERSITY SCHOOL OF MEDICINE, WASHINGTON, DC. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: SECTION D1: BRAND NAME CORRECTED. SECTION D4: MODEL NUMBER AND CATALOG NUMBER CORRECTED. SECTION E: REPORTER INFORMATION CORRECTED. SECTION G2: REPORT SOURCE CORRECTED. SECTION H6: HEALTH EFFECT - CLINICAL CODE CORRECTED. MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 DEVICES AND THE PATIENT OUTCOMES COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RESEARCH ARTICLE TITLED ¿EARLY VENTRICULAR ARRHYTHMIAS AFTER LEFT VENTRICULAR ASSIST DEVICE IMPLANTATION¿ WAS RECEIVED. THE ARTICLE SOUGHT TO IDENTIFY THE INCIDENCE OF EARLY VENTRICULAR ARRHYTHMIAS (EVAS), RISK FACTORS ASSOCIATED WITH THE DEVELOPMENT OF EVA, AND THE PROGNOSTIC SIGNIFICANCE OF EVA IN DURABLE LEFT VENTRICULAR ASSIST DEVICE (LVAD) PATIENTS. A SINGLE-CENTER RETROSPECTIVE ANALYSIS WAS PERFORMED OF PATIENTS OLDER THAN 18 YEARS OF AGE WHO UNDERWENT LVAD IMPLANTATION FROM (B)(6) 2006 THROUGH (B)(6) 2022. PATIENTS WERE EITHER IMPLANTED WITH HEARTMATE II, HEARTMATE 3, OR HEARTWARE DEVICES. THE STUDY INCLUDED A TOTAL OF 789 PATIENTS WITH 188 PATIENTS (23.8%) IMPLANTED WITH HEARTMATE II, 327 PATIENTS (41.4%) IMPLANTED WITH HEARTMATE 3, AND 273 PATIENTS (34.6%) IMPLANTED WITH HEARTWARE. EVAS WERE DEFINED AS SUSTAINED VENTRICULAR ARRHYTHMIAS OCCURRING LESS THAN 30 DAYS AFTER LVAD IMPLANTATION. OF THE 789 LVAD PATIENTS, 43 PATIENTS (6.2%) WHO HAD NO EVA AFTER LVAD IMPLANT EXPIRED WITHIN 30-DAYS AND 12 PATIENTS (12.0%) WHO HAD EVA AFTER LVAD IMPLANT EXPIRED WITHIN 30 DAYS. THE ARTICLE SHOWED THAT, OF THE PATIENTS THAT EXPIRED WITHIN 30 DAYS, 31 PATIENTS EXPIRED DUE TO CARDIOVASCULAR REASONS AND 24 PATIENTS EXPIRED DUE TO NON-CARDIOVASCULAR REASONS. ADDITIONALLY, THE ARTICLE SHOWED THAT EXPIRATIONS DUE TO CARDIOVASCULAR REASONS WERE MORE FREQUENT IN PATIENTS WITH EVAS (57.1% VERSUS 38.3%), INCLUDING A NUMERICALLY GREATER RATE OF DEATH FROM RIGHT VENTRICULAR FAILURE AMONG PATIENTS WITH EVAS (33.3% VERSUS 18.6%). THE STUDY CONCLUDED THAT EVAS ARE A FREQUENT COMPLICATION OCCURRING IN MORE THAN 1 IN 10 LVAD PATIENTS, EVAS ARE ASSOCIATED WITH AN INCREASED 30-DAY MORTALITY, THE HEARTMATE 3 DEVICE SERIAL NUMBERS, AS WELL AS OTHER SPECIFIC CASE/PATIENT INFORMATION, ARE NOT AVAILABLE. THE SERIAL NUMBER OR OTHER IDENTIFYING INFORMATION OF THE PRODUCT WAS NOT REPORTED AND WAS NOT ABLE TO BE DETERMINED DURING THE INVESTIGATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C, IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING RIGHT VENTRICULAR FAILURE AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6, UNDER ¿RIGHT HEART FAILURE¿, ALSO STATES THAT PATIENTS MAY DEVELOP RIGHT VENTRICULAR FAILURE DURING OR SHORTLY AFTER IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21568042&pc=QZI | 2124215-2025-14718 | 2025-02-12 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-03-10 04:00 | QZI | FARAWAVE | Adverse Event Without Identified Device or Use Problem | Bradycardia; Cardiac Arrest; Ventricular Fibrillation; Respiratory Failure |  |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING A PULSE FIELD ABLATION (PFA) PROCEDURE USING A FARAWAVE PFA CATHETER THE PATIENT PASSED AWAY DUE TO CARDIAC ARREST. AFTER THE PATIENT WAS EXTUBATED, THEY WERE BEING PREPARED TO BE MOVED FROM THE PROCEDURE ROOM. AT THIS TIME, THEY STARTED HAVING BREATHING ISSUES AND THE ANESTHESIOLOGIST STARTED TO MANUALLY BREATH FOR THEM BEFORE REINTUBATING THE PATIENT. THE EKG SHOWED THE PATIENT WAS BRADYCARDIC AND THEIR HEART RATE WAS 20 - 30 BEATS PER MINUTE. A CODE WAS CALLED BUT ULTIMATELY THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS DETERMINED TO BE CARDIAC ARREST, BUT THE PHYSICIAN STATED THERE WAS NO EVIDENCE OF AN EFFUSION OR CLOT AND DOES NOT BELIEVE THE DEATH WAS RELATED TO THE PFA PROCEDURE OR THE CATHETER. THE CATHETER WAS DISPOSED BY THE FACILITY AND WILL NOT BE RETURNED AS THE DEVICE PERFORMED AS EXPECTED DURING THE PROCEDURE AND THERE WERE NO REPORTS OF ANY PERFORMANCE CONCERNS.Event Description: IT WAS REPORTED THAT FOLLOWING A PULSE FIELD ABLATION (PFA) PROCEDURE USING A FARAWAVE PFA CATHETER THE PATIENT PASSED AWAY DUE TO CARDIAC ARREST. AFTER THE PATIENT WAS EXTUBATED, THEY WERE BEING PREPARED TO BE MOVED FROM THE PROCEDURE ROOM. AT THIS TIME, THEY STARTED HAVING BREATHING ISSUES AND THE ANESTHESIOLOGIST STARTED TO MANUALLY BREATH FOR THEM BEFORE REINTUBATING THE PATIENT. THE EKG SHOWED THE PATIENT WAS BRADYCARDIC AND THEIR HEART RATE WAS 20 - 30 BEATS PER MINUTE. A CODE WAS CALLED BUT ULTIMATELY THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS DETERMINED TO BE CARDIAC ARREST, BUT THE PHYSICIAN STATED THERE WAS NO EVIDENCE OF AN EFFUSION OR CLOT AND DOES NOT BELIEVE THE DEATH WAS RELATED TO THE PFA PROCEDURE OR THE CATHETER. THE CATHETER WAS DISPOSED BY THE FACILITY AND WILL NOT BE RETURNED AS THE DEVICE PERFORMED AS EXPECTED DURING THE PROCEDURE AND THERE WERE NO REPORTS OF ANY PERFORMANCE CONCERNS. IT WAS FURTHER REPORTED THAT PHYSICIANS ATTEMPTED TO PLACE PACER WIRES INTO THE PATIENT, BUT THE PATIENT WENT INTO VENTRICULAR FIBRILLATION BEFORE PASSING AWAY. Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21568838&pc=FKX | 0008030665-2025-00561 | 2025-02-22 05:00 | Death | ERIKA DE REYNOSA, S.A. DE C.V. | 2025-03-10 04:00 | FKX | LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | K181108 |  | 1 | Event Description: A PERITONEAL DIALYSIS (PD) PATIENT'S DAUGHTER REPORTED THAT THIS PATIENT ON PD EXPIRED WHILE STILL CONNECTED TO THE FRESENIUS CYCLER. THE CAUSE OF DEATH WAS UNKNOWN. HOWEVER, THERE WERE NO REPORTED ALLEGATIONS THAT THE DEATH WAS RELATED TO ANY PRODUCT RELATED ISSUES. IN ADDITIONAL FOLLOW-UP, THE PATIENT¿S PD NURSE STATED THAT THE PATIENT HAD PRE-EXISTING COMORBID CONDITIONS OF END STAGE RENAL DISEASE (ESRD) WITH ADDITIONAL CARDIAC COMORBIDITIES INCLUDED AORTIC VALVE STENOSIS AND HYPERTENSION. PER THE NURSE, THE PATIENT EXPIRED ON (B)(6) 2025. THE NURSE STATED THAT THE PATIENT¿S DAUGHTER REPORTED THAT THE PATIENT WAS ATTACHED TO THE FRESENIUS CYCLER. HOWEVER, NO TREATMENT DATA WAS AVAILABLE AND IT WAS UNKNOWN IF THE PATIENT WAS IN ACTIVE TREATMENT WHEN SHE EXPIRED. THE NURSE STATED THE PATIENT WAS AWAKE AND WELL AT 8AM. HOWEVER, SUBSEQUENTLY, AT APPROXIMATELY 10AM THE PATIENT WAS FOUND UNRESPONSIVE (BY DAUGHTER). THE NURSE STATED THAT EMERGENCY MEDICAL SERVICES WAS CALLED. THE PATIENT WAS A DNR AND NOT BROUGHT TO THE HOSPITAL; PRONOUNCED DECEASED IN THE HOME. THE NURSE STATED THERE WERE NO REPORTED ISSUES WITH THE PATIENT¿S PD TREATMENT OR PRODUCT IN RELATION TO THE DEATH. THE NURSE STATED THAT THE PATIENT¿S NEPHROLOGIST SUSPECTS THE CAUSE OF DEATH WAS FROM CARDIAC DUE TO NATURAL CAUSES FROM PRE-EXISTING COMORBID CONDITIONS. NO FURTHER INFORMATION ABOUT THE DEATH WAS AVAILABLE. THE FRESENIUS CYCLER WAS PENDING RETURN TO MANUFACTURER AT THE TIME FOLLOW-UP WAS COMPLETED.Event Description: A PERITONEAL DIALYSIS (PD) PATIENT'S DAUGHTER REPORTED THAT THIS PATIENT ON PD EXPIRED WHILE STILL CONNECTED TO THE FRESENIUS CYCLER. THE CAUSE OF DEATH WAS UNKNOWN. HOWEVER, THERE WERE NO REPORTED ALLEGATIONS THAT THE DEATH WAS RELATED TO ANY PRODUCT RELATED ISSUES. IN ADDITIONAL FOLLOW-UP, THE PATIENT¿S PD NURSE STATED THAT THE PATIENT HAD PRE-EXISTING COMORBID CONDITIONS OF END STAGE RENAL DISEASE (ESRD) WITH ADDITIONAL CARDIAC COMORBIDITIES INCLUDED AORTIC VALVE STENOSIS AND HYPERTENSION. PER THE NURSE, THE PATIENT EXPIRED ON (B)(6) 2025. THE NURSE STATED THAT THE PATIENT¿S DAUGHTER REPORTED THAT THE PATIENT WAS ATTACHED TO THE FRESENIUS CYCLER. HOWEVER, NO TREATMENT DATA WAS AVAILABLE AND IT WAS UNKNOWN IF THE PATIENT WAS IN ACTIVE TREATMENT WHEN SHE EXPIRED. THE NURSE STATED THE PATIENT WAS AWAKE AND WELL AT 8AM. HOWEVER, SUBSEQUENTLY, AT APPROXIMATELY 10AM THE PATIENT WAS FOUND UNRESPONSIVE (BY DAUGHTER). THE NURSE STATED THAT EMERGENCY MEDICAL SERVICES WAS CALLED. THE PATIENT WAS A DNR AND NOT BROUGHT TO THE HOSPITAL; PRONOUNCED DECEASED IN THE HOME. THE NURSE STATED THERE WERE NO REPORTED ISSUES WITH THE PATIENT¿S PD TREATMENT OR PRODUCT IN RELATION TO THE DEATH. THE NURSE STATED THAT THE PATIENT¿S NEPHROLOGIST SUSPECTS THE CAUSE OF DEATH WAS FROM CARDIAC DUE TO NATURAL CAUSES FROM PRE-EXISTING COMORBID CONDITIONS. NO FURTHER INFORMATION ABOUT THE DEATH WAS AVAILABLE. THE FRESENIUS CYCLER WAS PENDING RETURN TO MANUFACTURER AT THE TIME FOLLOW-UP WAS COMPLETED. Manufacturer Narrative: CLINICAL REVIEW: TEMPORAL RELATIONSHIP EXISTS BETWEEN PD THERAPY WITH THE LIBERTY SELECT CYCLER AND THE PATIENT¿S DEATH. HOWEVER, THERE IS NO ALLEGATION OR OBJECTIVE EVIDENCE THAT THE PATIENT¿S DEATH WAS RELATED TO A PRODUCT ISSUE OR MALFUNCTION WITH THE FRESENIUS CYCLER OR ANY ISSUES WITH PD TREATMENT. IT WAS REPORTED THAT THE SUSPECTED CAUSE OF DEATH WAS CARDIAC ARREST. THE PATIENT HAD A MEDICAL HISTORY THAT INCLUDED ESRD WITH PRE-EXISTING CARDIAC COMORBIDITIES THAT INCLUDED AORTIC VALVE STENOSIS AND HYPERTENSION. PATIENTS WITH END STAGE RENAL DISEASE (ESRD) ON DIALYSIS HAVE A SIGNIFICANTLY HIGHER MORTALITY RISK THAN THE GENERAL POPULATION. FURTHERMORE, THE PATIENT¿S PRE-EXISTING CARDIAC COMORBID CONDITIONS INCREASE RISK OF MORTALITY. THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. Manufacturer Narrative: ADDITIONAL INFORMATION: D.9., H.3. PLANT INVESTIGATION: THE ACTUAL DEVICE WAS RETURNED. A VISUAL INSPECTION OF THE RETURNED CYCLER EXTERIOR SHOWED SIGNS OF PHYSICAL DAMAGE: HEATER TRAY DISCOLORED. THERE WERE VISUAL INDICATIONS OF DRIED FLUID WITHIN THE CASSETTE COMPARTMENT ON THE PUMP. THERE WERE NO VISUAL INDICATIONS OF PARTICULATES WITHIN THE CASSETTE AREA. THERE WERE NO BURRS OR SHARP EDGES IN CASSETTE AREA THAT MAY HAVE PUNCTURED A CASSETTE MEMBRANE. A (AS-RECEIVED WITH REDUCED DWELL) SIMULATED TREATMENT WAS PERFORMED AND COMPLETED. VALVE ACTUATION TEST ¿ PASSED. SYSTEM AIR LEAK TEST ¿ PASSED. AN INTERNAL VISUAL INSPECTION OF THE RETURNED CYCLER WAS PERFORMED. THERE WERE VISUAL INDICATIONS OF DRIED FLUID ON THE BOTTOM COVER UNDERNEATH THE PUMP ASSEMBLY. THE CAUSE OF THE OBSERVED DRIED FLUID COULD NOT BE DETERMINED. THERE WERE NO DISCREPANCIES ENCOUNTERED WITH THE MUSHROOM HEADS. THE DEVICE HISTORY RECORD DID NOT REVEAL ANY ISSUES OR PROBLEMS RELATED TO THE REPORTED SYMPTOM CODE(S).UPON COMPLETION OF THE EVALUATION, THERE WERE NO MALFUNCTIONS THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. AN ASSOCIATED CAUSE COULD NOT BE DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21568962&pc=MNS | 2518422-2025-102831 | 2025-03-05 05:00 | Death | RESPIRONICS, INC. | 2025-03-10 04:00 | MNS | RENTAL, DREAMSTATION ST30, GB | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K102465 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT A PATIENT WITH A RENTAL DREAMSTATION ST30 GB DEVICE HAD PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE CLINIC REACHED OUT TO THE MANUFACTURER TO INQUIRE ABOUT RETURNING THE DEVICE. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. A FOLLOW-UP REPORT WILL BE FILED IF ANY FURTHER RELEVANT INFORMATION BECOMES AVAILABLE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. Manufacturer Narrative: THE MANUFACTURER PREVIOUSLY REPOTED THAT THEY RECEIVED INFORMATION ALLEGING THAT A PATIENT WITH A RENTAL DREAMSTATION ST30 GB DEVICE HAD PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE CLINIC REACHED OUT TO THE MANUFACTURER TO INQUIRE ABOUT RETURNING THE DEVICE. THE DEVICE UNDERWENT A REPAIR EVALUATION BY A THIRD-PARTY SERVICE CENTER, WHERE IT WAS VISIBLY INSPECTED, TESTED, AND SUBSEQUENTLY PASSED, WITH NO FAULTS FOUND. THE UNIT WAS THEN RETURNED TO LOAN STOCK . NO FURTHER INVESTIGATION IS NECESSARY. IF WE RECEIVE ANY RELEVANT INFORMATION, A FOLLOW UP REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21568974&pc=MIH | 2017233-2025-05908 | 2025-02-25 05:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-03-10 04:00 | MIH | GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM | Adverse Event Without Identified Device or Use Problem | Paralysis; Vascular Dissection | P040043 |  | 1 | Event Description: THE FOLLOWING INFORMATION WAS REPORTED TO GORE: ON (B)(6) 2025, THE PATIENT UNDERWENT AN EMERGENCY ENDOVASCULAR TREATMENT FOR A TYPE B ACUTE AORTIC DISSECTION WITH MALPERFUSION USING GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM. THE PROCEDURE WAS COMPLETED WITH IMPLANTATION OF ONE CTAGAC FROM ZONE 2, ONE STENT GRAFT AND ONE AORTIC STENT FROM ANOTHER MANUFACTURER (ZENITH DISSECTION ENDOVASCULAR SYSTEM). AFTER ENTERING TO THE ICU, THE PATIENT DEVELOPED SYMPTOMS OF ABDOMINAL PAIN AND PARALYSIS OF THE LOWER LIMB. A CONTRAST CT SCAN CONFIRMED THAT RETROGRADE AORTIC DISSECTION HAD OCCURRED FROM THE PROXIMAL END OF THE CTAGAC. AFTER THE ONSET OF CHEST PAIN AND HYPOTENSION, THE PATIENT WENT INTO CARDIAC ARREST, AND CARDIAC MASSAGE WAS PERFORMED. AN ASCENDING AORTIC ARCH REPLACEMENT WAS PERFORMED WITH ON-PUMP, BUT THE PATIENT'S CARDIAC FUNCTION DID NOT RECOVER, AND THE PATIENT EXPIRED. NO AUTOPSY INFORMATION WAS AVAILABLE. THE PHYSICIAN'S COMMENT: THE IMPLANTATION OF CTAGAC LIKELY CAUSED THE RETROGRADE AORTIC DISSECTION. THE DIAMETER OF THE AORTIC VESSEL UNDER THE LEFT COMMON CAROTID ARTERY WAS ABOUT 37 MM, SO THE 40 MM CTAGAC DEVICE WAS IMPLANTED. HOWEVER, IT WOULD HAVE BEEN BETTER TO SELECT THE 37 MM DIAMETER DEVICE. AFTER THE CTAGAC WAS DEPLOYED TO THE INTERMEDIATE, THE ANGULATION CONTROL WAS APPLIED 30 TIMES TO ADJUST THE ANGLE, AND THEN THE DEVICE WAS FULLY DEPLOYED, BUT A "BIRD BEAK" REMAINED. IT WOULD HAVE BEEN BETTER TO APPLY FURTHER ANGULATION CONTROL AFTER FULL DEPLOYMENT. Manufacturer Narrative: H6: A REVIEW OF THE MANUFACTURING RECORDS FOR THE DEVICE VERIFIED THAT THE LOT MET ALL PRE-RELEASE SPECIFICATIONS. H6: CODE D12: ACCORDING TO THE GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM INSTRUCTIONS FOR USE (IFU), ADVERSE EVENTS THAT MAY OCCUR AND/OR REQUIRE INTERVENTION INCLUDE, BUT ARE NOT LIMITED TO : DEATH, DISSECTION, PERFORATION, OR RUPTURE OF THE AORTIC VESSEL & SURROUNDING VASCULATURE, NEUROLOGIC DAMAGE, LOCAL OR SYSTEMIC (E.G., STROKE, PARAPLEGIA, PARAPARESIS) W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21568982&pc=NPT | 2015691-2025-01636 |  | Death | EDWARDS LIFESCIENCES | 2025-03-10 04:00 | NPT | SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Cardiac Tamponade | P140031 |  | 1 | Event Description: AS REPORTED FROM PATIENT SUPPORT IN REGARD TO A CALL FROM A CARE ADVOCATE WHO RECEIVED ID CARDS FOR THE PATIENT, POST TRANSFEMORAL TAVR PROCEDURE WITH A 23MM SAPIEN 3 ULTRA, THE PATIENT HAS PASSED AWAY. PER CARE ADVOCATE, THE PATIENT PASSED DUE TO CARDIAC TAMPONADE AND ACUTE HEART FAILURE. THE CALLER DID NOT ASK TO SPEAK WITH AN EDWARDS PHYSICIAN AND DID NOT MAKE ANY ALLEGATIONS AGAINST THE DEVICE. THE DATE OF DEATH IS UNKNOWN Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. ACCORDING TO THE DEVICE TRAINING MANUALS, RISK FACTORS FOR AORTIC DISSECTION, HEMATOMA, OR ANNULAR RUPTURE DURING THE TVR PROCEDURE INCLUDE SIGNIFICANT TRANSCATHETER HEART VALVE OVERSIZING, SEVERELY OBLITERATED SINUSES OF VALSALVA, PORCELAIN AORTA, AND/OR PRESENCE OF BULKY CALCIFICATION AND NARROW CALCIFIED SINOTUBULAR JUNCTION. IN ADDITION, ADVANCED AGE, FEMALE GENDER, SMALL BODY WEIGHT, AND STEROID DEPENDENCY CAN ALSO BE CONTRIBUTING FACTORS. THE SAPIEN VALVE (ALL MODELS) RELIES ON VALVE CALCIUM TO SECURELY ANCHOR TO THE LANDING ZONE. DESPITE THIS BENEFICIAL ASPECT OF CALCIUM, BULKY CALCIUM CAN INCREASE THE RISK OF CALCIFIC NODULE DISPLACEMENT INTO THE VASCULATURE, WHICH CAN LEAD TO CARDIOVASCULAR INJURY. AT TIMES, THE EXTENT AND DISTRIBUTION OF CALCIUM CAN IMPAIR EASE OF DELIVERY OF THE VALVE, CORRECT POSITIONING OF THE VALVE, DEPLOYMENT OF THE VALVE, AND PROCEDURAL SUCCESS. PERICARDIAL EFFUSION IS AN ABNORMAL ACCUMULATION OF FLUID IN THE PERICARDIAL CAVITY. BECAUSE OF THE LIMITED AMOUNT OF SPACE IN THE PERICARDIAL CAVITY, FLUID ACCUMULATION WILL LEAD TO AN INCREASED INTRA PERICARDIAL PRESSURE WHICH CAN NEGATIVELY AFFECT HEART FUNCTION. WHEN THERE IS A PERICARDIAL EFFUSION WITH ENOUGH PRESSURE TO ADVERSELY AFFECT HEART FUNCTION, THIS IS CALLED CARDIAC TAMPONADE. IT CAN BE CAUSED BY A VARIETY OF LOCAL AND SYSTEMIC DISORDERS, OR IT MAY BE IDIOPATHIC. PERICARDIAL EFFUSIONS CAN BE ACUTE OR CHRONIC, AND THE TIME COURSE OF DEVELOPMENT HAS A SIGNIFICANT IMPACT ON THE PATIENT'S SYMPTOMS. IN TVR PATIENTS, IT MAY BE CAUSED BY GUIDEWIRE PERFORATIONS OR ANNULAR RUPTURES. IN TRANSAPICAL PATIENTS, POSTOPERATIVE PERICARDIAL EFFUSIONS ARE COMMON, MOST WILL RESOLVE SPONTANEOUSLY, AND SOME MAY REQUIRE INTERVENTION. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. ALTHOUGH THE EXACT CAUSE AND SOURCE OF THE CARDIAC TAMPONADE CANNOT BE DETERMINED, IT IS POSSIBLE THE MECHANISMS OF THE TAVR PROCEDURE DESCRIBED ABOVE OR PATIENT FACTORS NOT PROVIDED MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21569061&pc=DQX | 3003775027-2025-00086 |  | Death | ASAHI INTECC CO., LTD. | 2025-03-10 04:00 | DQX | PROWATERFLEX | Material Separation | Foreign Body In Patient | K163426 |  | 1 | Event Description: IT WAS REPORTED THAT AN ASAHI PROWATERFLEX GUIDE WIRE WAS FRACTURED IN THE LEFT MAIN TRUNK (LMT) DURING A PERCUTANEOUS CORONARY INTERVENTION (PCI) TO TREAT THE SEGMENT FROM THE LMT TO THE LEFT ANTERIOR DESCENDING ARTERY (LAD). THE WIRE FRAGMENT WAS LEFT IN SITU. IT WAS INFORMED THAT THE PATIENT EVENTUALLY DIED.Event Description: IT WAS REPORTED THAT THE PATIENT WITH A 100%-OCCLUDED LEFT MAIN TRUNK (LMT) WAS IN CARDIOGENIC SHOCK WHEN HE WAS BROUGHT IN THE CATHETERIZATION LABORATORY. PERCUTANEOUS CORONARY INTERVENTION (PCI) WAS PERFORMED TO TREAT A 100%-OCCLUDED, HEAVILY CALCIFIED LESION OF 30MM IN LENGTH IN THE LMT. AN ASAHI PROWATERFLEX GUIDE WIRE WAS USED WITH AN UNSPECIFIED BALLOON CATHETER TO DEPLOY A 3.5MM-IN-DIAMETER RESOLUTE ONYX STENT (MEDTRONIC) TO SUCCESSFULLY RE-CANALIZE THE LMT. AT THE END OF THE PROCEDURE, THE SUBJECT PROWATERFLEX WAS BEING REMOVED BUT SLIGHT RESISTANCE WAS MET. AS THE GUIDE WIRE WAS CONTINUED TO BE PULLED BACK, THE DISTAL SEGMENT OF THE GUIDE WIRE WAS FOUND BROKEN OFF. THE DEPLOYED STENT WAS IMMEDIATELY CRUSHED IN THE LMT, RESULTING IN UNSUCCESSFUL RE-CANALIZATION. IT WAS INFORMED THAT THE PATIENT EVENTUALLY DIED. THE PHYSICIAN COMMENTED THAT THE SUBJECT PROWATERFLEX MET RESISTANCE PROBABLY BECAUSE THE GUIDE WIRE MIGHT HAVE BEEN TRAPPED BY A PART OF THE STENT STRUT. AS HE KEPT PULLING BACK THE GUIDE WIRE AGAINST RESISTANCE FOR REMOVAL, THE STENT WENT CRUSHED. CONSEQUENTLY, THE LMT WAS COMPLETELY CLOSED OFF, PREVENTING THE PHYSICIAN FROM ENTERING ANY DEVICE IN THE LMT AND CAUSING CARDIAC ARREST. NOTHING WAS ABLE TO RE-CANALIZE THE LESION AT THAT POINT. Manufacturer Narrative: MANUFACTURING SITE: ASAHI INTECC (THAILAND) CO., LTD., PATHUM THANI, THAILAND, REGISTRATION NUMBER: 3003780911. DEVICE EVALUATION COULD NOT BE PERFORMED BECAUSE THE AFFECTED DEVICE HAD NOT BEEN RETURNED YET. LOT HISTORY REVIEW REVEALED NO ANOMALY RELATING TO THE REPORTED EVENT. NO OTHER SIMILAR PRODUCT EXPERIENCE REPORT WAS RECEIVED FROM THIS LOT. BASED ON THE OBTAINED INFORMATION AND REVIEWING THE KNOWN SIMILAR EVENTS, IT WAS PRESUMED THAT STRESS GENERATED WITH GUIDE WIRE MANIPULATION WHILE THE SUBJECT PROWATERFLEX GUIDE WIRE WAS CAUGHT DUE TO ANATOMICAL AND CONDITIONS. CONSEQUENTLY, THE GUIDE WIRE WAS FRACTURED. IT WAS CONCLUDED THAT THIS EVENT WAS NOT ATTRIBUTED TO PRODUCT QUALITY. CAPA: NO CAPA WILL BE TAKEN. THE INSTRUCTIONS FOR USE (IFU) STATES: [WARNINGS]: OBSERVE MOVEMENT OF THIS GUIDE WIRE IN THE VESSELS. BEFORE THIS GUIDE WIRE IS MOVED OR TORQUED, THE TIP MOVEMENT SHOULD BE EXAMINED AND MONITORED UNDER FLUOROSCOPY. DO NOT MOVE OR TORQUE THE GUIDE WIRE WITHOUT OBSERVING CORRESPONDING MOVEMENT OF THE TIP; OTHERWISE, THE GUIDE WIRE MAY BE DAMAGED AND/OR TRAUMA MAY OCCUR. IN ADDITION, ENSURE THAT THE DISTAL TIP OF THIS GUIDE WIRE AND ITS LOCATION IN THE VESSEL ARE VISIBLE DURING MANIPULATIONS OF THE GUIDE WIRE. NEVER PUSH, AUGER, WITHDRAW, OR TORQUE THIS GUIDE WIRE THAT MEETS RESISTANCE. TORQUING OR PUSHING THIS GUIDE WIRE AGAINST RESISTANCE MAY CAUSE DAMAGE AND/OR TIP SEPARATION OF THIS GUIDE WIRE OR DIRECT DAMAGE TO A VESSEL. RESISTANCE MAY BE FELT AND/OR OBSERVED UNDER FLUOROSCOPY BY NOTING ANY BUCKLING OF THE GUIDE WIRE. IF THE PROLAPSE OF THE GUIDE WIRE TIP IS OBSERVED, DO NOT ALLOW THE TIP TO REMAIN IN A PROLAPSED POSITION; OTHERWISE, DAMAGE TO THE GUIDE WIRE MAY OCCUR. DETERMINE THE CAUSE OF RESISTANCE UNDER FLUOROSCOPY AND TAKE ANY NECESSARY REMEDIAL ACTION. IF RESISTANCE IS FELT DUE TO SPASM, BENDING OF THE GUIDE WIRE, OR DUE TO TRAP WHILE OPERATING THIS GUIDE WIRE IN THE BLOOD VESSEL OR REMOVING IT, DO NOT TORQUE AND/OR PULL THE GUIDE WIRE ITSELF. STOP THE PROCEDURE. DETERMINE THE CAUSE OF RESISTANCE UNDER FLUOROSCOPY AND TAKE APPROPRIATE REMEDIAL ACTION. IF THE GUIDE WIRE IS MOVED EXCESSIVELY, IT MAY BREAK OR BECOME DAMAGED, WHICH MAY CAUSE BLOOD VESSEL INJURY OR RESULT IN FRAGMENTS BEING LEFT INSIDE THE VESSEL. [MALFUNCTION AND ADVERSE EFFECTS]: SEPARATION OF THE GUIDE WIRE. Manufacturer Narrative: MANUFACTURING SITE: ASAHI INTECC (THAILAND) CO., LTD., PATHUM THANI, THAILAND, REGISTRATION NUMBER: (B)(4). THE REPORTED PROWATERFLEX GUIDE WIRE WAS RETURNED FOR INVESTIGATION. THE RETURNED PROWATERLFLEX GUIDE WIRE WAS FOUND THREE-DIMENSIONALLY DEFORMED FOR APPROXIMATELY 30MM FROM THE DISTAL MID SOLDER (SET AT 25MM FROM THE WIRE TIP TO FIX THE OUTER COIL ONTO THE CORE WIRE). THE OUTER COIL OF THE GUIDE WIRE WAS FOUND ELONGATED FOR APPROXIMATELY 45MM FROM THE DISTAL MID SOLDER AND FRACTURED. THE BALL TIP WAS NOT FOUND AT THE VERY DISTAL END OF THE OUTER COIL. THE CORE WIRE WAS COMING OUT OF THE DISTAL MID SOLDER FOR APPROXIMATELY 10MM. THE OUTER COIL WAS FOUND DISARRANGED AT APPROXIMATELY 10MM AND 14MM PROXIMAL TO THE DISTAL MID SOLDER. OBSERVATION BY MICROSCOPE AND SCANNING ELECTRON MICROSCOPE (SEM) OF THE CORE WIRE FOUND NECKING OF THE FRACTURE END AND DIMPLES ON A RELATIVELY FLAT FRACTURE SURFACE, INDICATING THAT TENSIONAL STRESS HAD CONTRIBUTED TO THE CORE WIRE FRACTURE. OBSERVATION BY MICROSCOPE AND SEM OF THE OUTER COIL FRACTURE END FOUND A RELATIVELY FLAT FRACTURE SURFACE, INDICATING THAT TORSION WAS GENERATED AS THE COIL STRUCTURE HAD BEEN STRAIGHTENED BY TENSION. TWISTING OF THE OUTER COIL WAS ALSO OBSERVED AT THE FRACTURE END. MEASUREMENT OF THE RETURNED PROWATERFLEX GUIDE WIRE SUGGESTED THAT THE CORE WIRE WAS FRACTURED AT APPROXIMATELY 15MM FROM THE WIRE TIP AND THE OUTER COIL WAS ELONGATED AND DETACHED TOGETHER WITH THE BALL TIP AT APPROXIMATELY 20MM FROM THE WIRE TIP. LOT HISTORY REVIEW REVEALED NO ANOMALY RELATING TO THE REPORTED EVENT. NO OTHER SIMILAR PRODUCT EXPERIENCE REPORT WAS RECEIVED FROM THIS LOT. BASED ON THE OBTAINED INFORMATION AND THE INVESTIGATION OUTCOME, IT WAS PRESUMED THAT TENSIONAL STRESS GENERATED WITH GUIDE WIRE MANIPULATION DURING WITHDRAWAL ATTEMPTS MIGHT HAVE BEEN LOCALLY ACCUMULATED ON THE DISTAL SEGMENT OF THE SUBJECT PROWATERFLEX GUIDE WIRE WHILE ITS DISTAL SEGMENT HAD BEEN CAUGHT BY THE DEPLOYED STENT STRUT. CONSEQUENTLY, THE GUIDE WIRE WAS FRACTURED. IT WAS CONCLUDED THAT THIS EVENT WAS NOT ATTRIBUTED TO PRODUCT QUALITY. CAPA: NO CAPA WILL BE TAKEN. THE INSTRUCTIONS FOR USE (IFU) STATES: [WARNINGS] OBSERVE MOVEMENT OF THIS GUIDE WIRE IN THE VESSELS. BEFORE THIS GUIDE WIRE IS MOVED OR TORQUED, THE TIP MOVEMENT SHOULD BE EXAMINED AND MONITORED UNDER FLUOROSCOPY. DO NOT MOVE OR TORQUE THE GUIDE WIRE WITHOUT OBSERVING CORRESPONDING MOVEMENT OF THE TIP; OTHERWISE, THE GUIDE WIRE MAY BE DAMAGED AND/OR TRAUMA MAY OCCUR. IN ADDITION, ENSURE THAT THE DISTAL TIP OF THIS GUIDE WIRE AND ITS LOCATION IN THE VESSEL ARE VISIBLE DURING MANIPULATIONS OF THE GUIDE WIRE. NEVER PUSH, AUGER, WITHDRAW, OR TORQUE THIS GUIDE WIRE THAT MEETS RESISTANCE. TORQUING OR PUSHING THIS GUIDE WIRE AGAINST RESISTANCE MAY CAUSE DAMAGE AND/OR TIP SEPARATION OF THIS GUIDE WIRE OR DIRECT DAMAGE TO A VESSEL. RESISTANCE MAY BE FELT AND/OR OBSERVED UNDER FLUOROSCOPY BY NOTING ANY BUCKLING OF THE GUIDE WIRE. IF THE PROLAPSE OF THE GUIDE WIRE TIP IS OBSERVED, DO NOT ALLOW THE TIP TO REMAIN IN A PROLAPSED POSITION; OTHERWISE DAMAGE TO THE GUIDE WIRE MAY OCCUR. DETERMINE THE CAUSE OF RESISTANCE UNDER FLUOROSCOPY AND TAKE ANY NECESSARY REMEDIAL ACTION. IF RESISTANCE IS FELT DUE TO SPASM, BENDING OF THE GUIDE WIRE, OR DUE TO TRAP WHILE OPERATING THIS GUIDE WIRE IN THE BLOOD VESSEL OR REMOVING IT, DO NOT TORQUE AND/OR PULL THE GUIDE WIRE ITSELF. STOP THE PROCEDURE. DETERMINE THE CAUSE OF RESISTANCE UNDER FLUOROSCOPY AND TAKE APPROPRIATE REMEDIAL ACTION. IF THE GUIDE WIRE IS MOVED EXCESSIVELY, IT MAY BREAK OR BECOME DAMAGED, WHICH MAY CAUSE BLOOD VESSEL INJURY OR RESULT IN FRAGMENTS BEING LEFT INSIDE THE VESSEL. [MALFUNCTION AND ADVERSE EFFECTS] SEPARATION OF THE GUIDE WIRE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21569181&pc=NPT | 2015691-2025-01638 | 2024-09-26 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-10 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Endocarditis | P140031 |  | 1 | Event Description: AS REPORTED THROUGH THE JAPANESE TAVI REGISTRY REPORTING SYSTEM, THE PATIENT UNDERWENT A TRANSAXILLARY/SUBCLAVIAN TAVR AND RECEIVED A 29 MM SAPIEN 3 ULTRA RESILIA VALVE IN THE NATIVE AORTIC POSITION. ON POSTOPERATIVE DAY (POD) 6, THE PATIENT DEVELOPED SEVERE CONDUCTION DISTURBANCE AND A PERMANENT PACEMAKER WAS IMPLANTED. ON POSTOPERATIVE DAY (POD) 22, THE PATIENT DEVELOPED INFECTIVE ENDOCARDITIS (IE) AND CONTINUED TO RECEIVE ANTIBIOTIC TREATMENT AT ANOTHER HOSPITAL. ALTHOUGH THE IE IMPROVED WITH ANTIBIOTICS, THE PATIENT WAS UNABLE TO TAKE ORAL FOOD AND PASSED AWAY ON POD 42. PER THE MEDICAL OPINION, THE EVENT RELATIONSHIP TO THE VALVE WAS AFFIRMED. ON POD 6 ((B)(6) 2024), THE PATIENT DEVELOPED SEVERE CONDUCTION DISTURBANCE AND A PERMANENT PACEMAKER WAS IMPLANTED. ON POD 22 ((B)(6) 2024), THE PATIENT DEVELOPED INFECTIVE ENDOCARDITIS (IE) AND CONTINUED TO RECEIVE ANTIBIOTIC TREATMENT AT ANOTHER HOSPITAL. ON POD 42 ((B)(6) 2024), THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS NOT REPORTED. INVESTIGATION FOLLOW UP WAS ATTEMPTED, NO FURTHER INFORMATION WAS AVAILABLE. IN THIS CASE, ALL ADVERSE EVENT (AE) S OF THIS CASE REPORTED TO BE RELATED TO EW DEVICES THOUGH THE JAPANESE TAVI REGISTRY. THE FACILITY REPORTED THAT THERE WAS A CAUSAL RELATIONSHIP BETWEEN THE CONDUCTION DISTURBANCE, IE, AND DEATH AND THE EW DEVICE, BUT NOT RELATED TO THE TAVI PROCEDURE. Manufacturer Narrative: THIS IS 2 OF 2 REPORTS. PER THE INSTRUCTIONS FOR USE (IFU), INFECTIONS SUCH AS SEPTICEMIA AND ENDOCARDITIS ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. ENDOCARDITIS IS AN INFECTION OF A NATIVE OR PROSTHETIC VALVE, IS TREATED WITH ANTIBIOTICS, AND MAY REQUIRE VALVE REPLACEMENT IF ANTIBIOTIC THERAPY IS NOT EFFECTIVE. CAUSES OF PROSTHETIC VALVE ENDOCARDITIS ARE WELL DOCUMENTED IN THE LITERATURE AND ARE TYPICALLY CLASSIFIED AS EARLY (<60 DAYS) OR LATE (>60 DAYS). THE TWO CATEGORIES SHOW DEFINITE DIFFERENCES IN CLINICAL FEATURES, MICROBIAL PATTERNS, AND MORTALITY. ACCORDING TO THE LITERATURE REVIEW, AND AS DOCUMENTED IN A TECHNICAL SUMMARY WRITTEN BY EW, EARLY PROSTHETIC VALVE ENDOCARDITIS IS USUALLY CAUSED BY PERIOPERATIVE BACTERIAL CONTAMINATION OF THE VALVE. IN EARLY CASES OF PROSTHETIC ENDOCARDITIS, SUBSEQUENT INFECTIONS MAY BE RELATED TO CONTAMINATION AT THE TIME OF SURGERY (POOR STERILE TECHNIQUE). ADDITIONAL CAUSES OF EARLY PVE CAN OCCUR FROM OTHER SOURCES OF INFECTION AND NOT FROM THE VALVE (E.G., DENTAL TREATMENT OR SURGICAL PROCEDURES INVOLVING THE UPPER RESPIRATORY TRACT, UTIS, PNEUMONIA). ADDITIONALLY, A NON-CONFORMANCE IN THE STERILITY OR PACKAGING PROCESSES OF THE VALVE MAY ALSO MANIFEST IN THE EARLY POST-OPERATIVE PERIOD CAUSING EARLY PVE. HOWEVER, THERE ARE SEVERAL TYPES OF BACTERIA THAT WOULD NOT SURVIVE THE EDWARDS STERILIZATION PROCESS. EDWARDS LIFESCIENCES PRODUCES AND PROVIDES STERILE TISSUE BIOPROSTHESIS TO ITS CUSTOMERS BY FOLLOWING CAREFULLY DESIGNED ROBUST STERILIZATION PROCESSES. THESE MANUFACTURING PROCESSES HAVE BEEN VALIDATED AND DEMONSTRATED TO CONSISTENTLY PROVIDE A SIGNIFICANT SAFETY FACTOR FROM WHICH MICROORGANISMS COULD NOT SURVIVE. MICROBIOLOGY AND PROCESS MONITORING IS ROUTINELY REVIEWED WITHIN QUALITY SYSTEMS TO MAINTAIN STERILITY CONTROL. VALIDATED TESTING HAS DEMONSTRATED THAT MICROORGANISMS COULD NOT SURVIVE EDWARDS' MULTI-STAGE PROCESSING WITH ENHANCED STERILANT OR HEATED GLUTARALDEHYDE TERMINAL STERILANT SOLUTION. THESE MULTIPLE, REDUNDANT MANUFACTURING CONTROLS ENSURE THE STERILITY OF EDWARDS' VALVES AS PROVIDED TO CUSTOMERS. THEREFORE, THE PROBABILITY OF ENDOCARDITIS RELATED TO EDWARDS' BIOPROSTHESIS IS REMOTE. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. WITH THE LIMITED INFORMATION PROVIDED, THE CAUSE OF THE ENDOCARDITIS IS UNKNOWN BUT MAY POSSIBLY BE RELATED TO BACTERIA THAT WAS INTRODUCED DURING THE IMPLANTATION OF THE PERMANENT PACEMAKER. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21569182&pc=NPT | 2015691-2025-01639 | 2024-09-10 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-10 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Heart Block | P140031 |  | 1 | Event Description: AS REPORTED THROUGH THE JAPANESE TAVI REGISTRY REPORTING SYSTEM, THE PATIENT UNDERWENT A TRANSAXILLARY/SUBCLAVIAN TAVR AND RECEIVED A 29 MM SAPIEN 3 ULTRA RESILIA VALVE IN THE NATIVE AORTIC POSITION. ON POSTOPERATIVE DAY (POD) 6, THE PATIENT DEVELOPED SEVERE CONDUCTION DISTURBANCE AND A PERMANENT PACEMAKER WAS IMPLANTED. IN THIS CASE, ALL ADVERSE EVENT (AE) S OF THIS CASE REPORTED TO BE RELATED TO EW DEVICES THOUGH THE JAPANESE TAVI REGISTRY. THE FACILITY REPORTED THAT THERE WAS A CAUSAL RELATIONSHIP BETWEEN THE CONDUCTION DISTURBANCE, IE, AND DEATH AND THE EW DEVICE, BUT NOT RELATED TO THE TAVI PROCEDURE. Manufacturer Narrative: THIS IS 1 OF 2 REPORTS. PER THE INSTRUCTIONS FOR USE (IFU), CONDUCTION SYSTEM DEFECTS (HEART BLOCK), ARRHYTHMIAS AND CONDUCTION SYSTEM DEFECTS THAT MAY REQUIRE A PERMANENT PACEMAKER ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH BALLOON AORTIC VALVULOPLASTY, BIOPROSTHETIC HEART VALVES, AND THE THV PROCEDURE. ACCORDING TO THE VALVE ACADEMIC RESEARCH CONSORTIUM (VARC) GUIDELINES, THE CLOSE ANATOMICAL RELATIONSHIP BETWEEN THE AORTIC VALVE COMPLEX AND THE BRANCHING ATRIOVENTRICULAR BUNDLE MAY EXPLAIN THESE COMPLICATIONS OF THE THV PROCEDURE. ACCORDING TO THE LITERATURE REVIEW, AND AS DOCUMENTED IN EW CLINICAL TECHNICAL SUMMARY FOR COMPLAINTS-CONDUCTION DISTURBANCES/ HEART BLOCK, ATRIOVENTRICULAR CONDUCTION DISTURBANCES AFTER THV ARE ASSOCIATED WITH MANY PATIENT-RELATED AND PROCEDURAL RELATED FACTORS, INCLUDING PRE-OPERATIVE CO-MORBID STATUS, THE DEGREE, AND BULKINESS OF AORTIC VALVE AND ANNULAR CALCIFICATION, INTER-VENTRICULAR SEPTAL THICKNESS, PRE-EXISTING ELECTROCARDIOGRAM ABNORMALITIES, THE DEPTH OF PROSTHESIS IMPLANTATION, AND THE PROFILE OF THE IMPLANTED PROSTHESIS. UNLIKE CONVENTIONAL AVR, WHERE THERE MAY BE LOCALIZED TRAUMA DUE TO DECALCIFICATION OF THE ANNULUS AND/OR SUTURE PLACEMENT IN THE PROXIMITY OF THE AV NODE OR THE BUNDLES, THV MAY CAUSE CONDUCTION ABNORMALITIES THROUGH MECHANICAL IMPINGEMENT OF THE CONDUCTION SYSTEM BY THE PROSTHESIS. THE MECHANISMS OF THE DEVELOPMENT OF HEART BLOCK AFTER THV ARE WELL DOCUMENTED AND DESCRIBED IN THE LITERATURE. IT IS ALSO DOCUMENTED THAT PRE-EXISTING HEART BLOCK IS COMMON IN PATIENTS UNDERGOING THV OR SURGICAL AVR AND ANOTHER 4-6 % WILL DEVELOP POSTOPERATIVE HEART BLOCK, POTENTIALLY REQUIRING A PERMANENT PACEMAKER. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. WITH THE LIMITED INFORMATION PROVIDED, THE CAUSE OF THE HEART BLOCK IS UNKNOWN BUT MAY BE RELATED TO THE MECHANISMS DESCRIBED ABOVE. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21556288&pc=QZK | 2017233-2025-05903 | 2025-02-07 05:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-03-09 05:00 | QZK | GORE® EXCLUDER® THORACOABDOMINAL BRANCH ENDOPROSTHESIS | Patient-Device Incompatibility | Abdominal Pain; Ischemia; Thromboembolism | P230023 |  | 1 | Event Description: ON (B)(6) 2025, THE PATIENT UNDERWENT ENDOVASCULAR TREATMENT OF A THORACOABDOMINAL AORTIC ANEURYSM USING A GORE® EXCLUDER® THORACOABDOMINAL BRANCH ENDOPROSTHESIS. IT WAS REPORTED THAT THE DEVICES WERE SUCCESSFULLY IMPLANTED AND WERE PATENT AT THE CLOSE OF THE PROCEDURE. IT WAS NOTED THAT MURAL THROMBUS WAS PRESENT IN THE THORACIC AORTA. ON (B)(6) 2025, THE PATIENT BEGAN TO EXPERIENCE STOMACH AND INTESTINAL PAIN. IT WAS SUSPECTED THAT WIRE MANIPULATION DURING THE PROCEDURE MAY HAVE CAUSED SHOWER EMBOLI FROM THE THORACIC AORTA TO SPRAY INTO THE CELIAC OR SUPERIOR MESENTERIC ARTERY. A REINTERVENTION WAS PERFORMED ON THE PATIENT'S COLON. THE PROCEDURE WAS SUCCESSFULLY COMPLETED. HOWEVER, THE PATIENT CONTINUED TO DETERIORATE AND STOMACH ISCHEMIA WAS REPORTED. DETERIORATION CONTINUED AND CTA OF THE PATIENT'S ABDOMEN AND PELVIS CONFIRMED THAT ALL IMPLANTED DEVICES REMAINED PATENT. ON (B)(6) 2025, THE PATIENT EXPIRED. Manufacturer Narrative: B.2. OUTCOMES ATTRIBUTED TO ADVERSE EVENT: ADDED. D.2A./D.2B COMMON DEVICE NAME/PRODUCT CODE: CORRECTED. D.5. OPERATOR OF DEVICE: ADDED. H.6. COMPONENT CODE: CODE G04122 UPDATED TO CODE G07001. H.6. TYPE OF INVESTIGATION: CODE B15 ADDED. H.6. TYPE OF INVESTIGATION: CODE B15 - IMAGES WERE PROVIDED, AND AN IMAGING EVALUATION WAS PERFORMED. H.6. INVESTIGATION FINDINGS FOR IMAGING EVALUATION: CODE C19 ADDED. H.6. INVESTIGATION FINDINGS FOR IMAGING EVALUATION: CODE C19 - TWO SETS OF PRE-IMPLANTATION CTA IMAGES DATED (B)(6) 2024, AND (B)(6) 2024, WERE AVAILABLE FOR EVALUATION. THE IMAGING EVALUATION, PERFORMED BY A CLINICAL IMAGING SPECIALIST, SHOWED THAT THERE APPEARED TO BE IRREGULAR THROMBUS IN THE DESCENDING THORACIC AORTA (DTA) AND ABDOMINAL AORTA ON BOTH TIME-POINTS. THE ABDOMINAL/PELVIS CTA SCAN ONLY SHOWED HALF OF THE AORTA: POSSIBLE PROXIMAL LANDING ZONE (45MM ¿ 65MM) IN THE DTA. DIAMETERS APPEARED TO RANGE FROM 35.3MM ¿ 37.4MM, WITH THROMBUS. THE PARTIAL 3D IMAGES SHOWED AORTIC TORTUOSITY JUST DISTAL TO THE RENAL ARTERIES. IMAGES SHOWED THROMBUS AT THE LEVEL OF THE RENAL ARTERIES AND DISTALLY IN THE TORTUOUS AORTA. AORTIC ANEURYSM DIAMETER ON THE (B)(6) 2024, STUDY: 48.9MM X 48.7MM. AORTIC ANEURYSM DIAMETER ON THE (B)(6) 2024, STUDY: 52.4MM X 50.2MM. THERE APPEARED TO BE MINIMAL CHANGE IN THROMBUS BETWEEN THE 2 TIME-POINTS. H.6. INVESTIGATION CONCLUSIONS: CODES D1001 AND D12 REMAIN UNCHANGED. ACCORDING TO THE GORE® EXCLUDER® THORACOABDOMINAL BRANCH ENDOPROSTHESIS INSTRUCTIONS FOR USE (IFU), POTENTIAL DEVICE AND PROCEDURE-RELATED ADVERSE EVENTS AND COMPLICATIONS THAT MAY OCCUR WITH THE USE OF THE GORE® EXCLUDER® THORACOABDOMINAL BRANCH ENDOPROSTHESIS, OR IN ANY ENDOVASCULAR REPAIR PROCEDURE, WHICH MAY OR MAY NOT REQUIRE INTERVENTION INCLUDE, BUT ARE NOT LIMITED TO, BOWEL COMPLICATIONS (E.G., ILEUS, TRANSIENT ISCHEMIA, MESENTERIC ISCHEMIA, INFARCTION, NECROSIS), AND DEATH. THE IFU ALSO NOTES THAT WIRE POSITIONING WITHIN THE VISCERAL ARTERIES AND AORTA SHOULD BE MONITORED AND MAINTAINED THROUGHOUT THE PROCEDURE TO MINIMIZE THE RISK OF IATROGENIC COMPLICATIONS DUE TO WIRE MANIPULATION. Manufacturer Narrative: A.4. PATIENT WEIGHT: ASKED BUT UNAVAILABLE. B.7. OTHER RELEVANT HISTORY, INCLUDING PREEXISTING MEDICAL CONDITIONS: ASKED BUT UNAVAILABLE. D.10. CONCOMITANT MEDICAL PRODUCTS AND THERAPY DATES: ASKED BUT UNAVAILABLE. H.6. INVESTIGATION FINDINGS FOR ANALYSIS OF PRODUCTION RECORDS: CODE C19 - THE REVIEW OF THE MANUFACTURING PAPERWORK VERIFIED THAT THE LOTS INVOLVED IN THIS EVENT MET ALL PRE-RELEASE SPECIFICATIONS. ACCORDING TO THE GORE® EXCLUDER® THORACOABDOMINAL BRANCH ENDOPROSTHESIS INSTRUCTIONS FOR USE, POTENTIAL DEVICE AND PROCEDURE-RELATED ADVERSE EVENTS AND COMPLICATIONS THAT MAY OCCUR WITH THE USE OF THE GORE® EXCLUDER® THORACOABDOMINAL BRANCH ENDOPROSTHESIS, OR IN ANY ENDOVASCULAR REPAIR PROCEDURE, WHICH MAY OR MAY NOT REQUIRE INTERVENTION INCLUDE, BUT ARE NOT LIMITED TO, ENDOLEAK. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21544479&pc=QQG | 3014265054-2025-00001 | 2025-01-28 05:00 | Death | PHAGENESIS LTD | 2025-03-07 05:00 | QQG | PHAGENYX | Adverse Event Without Identified Device or Use Problem | Aspiration/Inhalation | DEN220025 |  | 1 | Event Description: DEATH OF A PATIENT THAT HAD A PHAGENYX PNX-1000 CATHETER IN PLACE, WHILST RECEIVING TUBE FEEDS VIA THE CATHETER. THE HOSPITAL FOLLOWED CORRECT PROCEDURES IN CONFIRMING CORRECT POSITION OF THE TUBE IN THE STOMACH BY X-RAY PRIOR TO INITIATING FEEDING. THE CATHETER HAD NOT BEEN USED TO DELIVER ELECTRICAL STIMULATION. PT WAS IN DIALYSIS, WAS INCONTINENT OF BOWEL AND THEY CLEANED HIM UP. THIS WAS AT THE END OF HIS DIALYSIS RUN. THE PATIENT WAS FULLY ALERT AND ORIENTED WHEN HE WENT TO DIALYSIS, AND THEY NOTED HE BECAME LETHARGIC AFTER CLEANING HIM UP. WHEN THEY TOOK HIM BACK TO PCU THEY IMMEDIATELY CALLED A RAPID RESPONSE AND THEN A CODE BLUE. TUBE FEEDS STOPPED PRIOR TO STARTING COMPRESSIONS. WHEN THEY WERE INTUBATING HIM, THEY FOUND TUBE FEEDING IN HIS LUNGS, PHYSICIANS REPORTED THAT THE PATIENT MAY HAVE ASPIRATED TUBE FEEDS. THERE IF NO DIRECT EVIDENCE THAT THE TUBE ENTERED THE AIRWAYS ONLY THAT FEED WAS FOUND IN THE AIRWAYS. CAUSE OF DEATH HAS NOT YET BEEN COMMUNICATED BY THE HOSPITAL TO THE MANUFACTURER. THE CATHETER WAS DICARDED. THERE IS CURRENTLY NO EVIDENCE OF ANY DEFECT OR FAULT WITH THE DEVICE USED, AND NO CASUAL RELATIONSHIP BETWEEN THE DEATH AND DEVICE HAS CURRENTLY BEEN ESTABLISHED. FURTHER INFORMATION HAS BEEN REQUESTED FROM THE HOSPITAL TO CONFIRM THE CAUSE OF DEATH AND ANY CONTRIBUTORY EFFECT THAT THE DEVICE MAY HAVE HAD.Event Description: DEATH OF A PATIENT THAT HAD A PHAGENYX PNX-1000 CATHETER IN PLACE, WHILST RECEIVING TUBE FEEDS VIA THE CATHETER. THE HOSPITAL FOLLOWED CORRECT PROCEDURES IN CONFIRMING CORRECT POSITION OF THE TUBE IN THE STOMACH BY X-RAY PRIOR TO INITIATING FEEDING. THE CATHETER HAD NOT BEEN USED TO DELIVER ELECTRICAL STIMULATION. PT WAS IN DIALYSIS, WAS INCONTINENT OF BOWEL AND THEY CLEANED HIM UP. THIS WAS AT THE END OF HIS DIALYSIS RUN. THE PATIENT WAS FULLY ALERT AND ORIENTED WHEN HE WENT TO DIALYSIS, AND THEY NOTED HE BECAME LETHARGIC AFTER CLEANING HIM UP. WHEN THEY TOOK HIM BACK TO PCU THEY IMMEDIATELY CALLED A RAPID RESPONSE AND THEN A CODE BLUE. TUBE FEEDS STOPPED PRIOR TO STARTING COMPRESSIONS. WHEN THEY WERE INTUBATING HIM, THEY FOUND TUBE FEEDING IN HIS LUNGS, PHYSICIANS REPORTED THAT THE PATIENT MAY HAVE ASPIRATED TUBE FEEDS. THERE IF NO DIRECT EVIDENCE THAT THE TUBE ENTERED THE AIRWAYS ONLY THAT FEED WAS FOUND IN THE AIRWAYS. CAUSE OF DEATH HAS NOT YET BEEN COMMUNICATED BY THE HOSPITAL TO THE MANUFACTURER. THE CATHETER WAS DISCARDED. THERE IS CURRENTLY NO EVIDENCE OF ANY DEFECT OR FAULT WITH THE DEVICE USED, AND NO CASUAL RELATIONSHIP BETWEEN THE DEATH AND DEVICE HAS CURRENTLY BEEN ESTABLISHED. FURTHER INFORMATION HAS BEEN REQUESTED FROM THE HOSPITAL TO CONFIRM THE CAUSE OF DEATH AND ANY CONTRIBUTORY EFFECT THAT THE DEVICE MAY HAVE HAD. UPDATED INFORMATION 21 MAR 2025. UPDATED INFORMATION WAS RECEIVED VIA EMAIL FROM THE REPORTING HOSPITAL, STATING THAT THE TUBE FEED WAS NOT LOCATED IN THE PATIENT LUNGS BUT WAS OBSERVED IN THE UPPER AIRWAYS AND THAT WITHOUT AN AUTOPSY THERE IS NO WAY TO KNOW IF IT WAS IN THE LOWER AIRWAYS. THE UPDATED INFORMATION DOES NOT CHANGE THE OUTCOME OF THIS INVESTIGATION AND IS STILL DEEMED REPORTABLE AS THE DEVICE COULD STILL HAVE LED TO THE PATIENT DEATH. Manufacturer Narrative: MANUFACTURER COMPLIANT REFERENCE (B)(4). THE DEVICE WAS DISPOSED OF BY THE HOSPITAL AND IS NOT AVAILABLE FOR EVALUATION. SUBMISSION DELAYED DUE TO ESG SET UP. FORM FDA 3500A PDF WAS UPLOADED ONTO WEBTRADER TEST ACCOUNT 27 FEB 2025. Manufacturer Narrative: MANUFACTURER COMPLIANT REFERENCE (B)(4). THE DEVICE WAS DISPOSED OF BY THE HOSPITAL AND IS NOT AVAILABLE FOR EVALUATION. SUBMISSION DELAYED DUE TO ESG SET UP. FORM FDA 3500A PDF WAS UPLOADED ONTO WEBTRADER TEST ACCOUNT 27 FEB 2025. ADDITION OF SUPPLEMENTAL FOLLOW UP INFORMATION 21 MAR 2025 UPDATED INFORMATION WAS RECEIVED VIA EMAIL FROM THE REPORTING HOSPITAL, STATING THAT THE TUBE FEED WAS NOT LOCATED IN THE PATIENT LUNGS BUT WAS OBSERVED IN THE UPPER AIRWAYS AND THAT WITHOUT AN AUTOPSY THERE IS NO WAY TO KNOW IF IT WAS IN THE LOWER AIRWAYS. THE UPDATED INFORMATION DOES NOT CHANGE THE OUTCOME OF THIS INVESTIGATION AND IS STILL DEEMED REPORTABLE AS THE DEVICE COULD STILL HAVE LED TO THE PATIENT DEATH. IF FURTHER INFORMATION IS MADE AVAILABLE, A SUBSEQUENT FOLLOW UP REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21544922&pc=LZG | 3004464228-2025-09749 | 2025-02-20 05:00 | Death | INSULET CORPORATION | 2025-03-07 05:00 | LZG | OMNIPOD DASH INSULIN MANAGEMENT SYSTEM | Insufficient Flow or Under Infusion; Insufficient Information | Diabetic Ketoacidosis | K211575 |  | 1 | Event Description: AN MHRA USER REPORT HAS BEEN RECEIVED, REPORTED BY A HCP "PATIENT HAD USED THE OMNIPOD PUMP TO DELIVER MULTIPLE INSULIN BOLUSES FOR HYPERGLYCEMIA CORRECTION, HOWEVER IT DID NOT APPEAR THAT THE INSULIN BOLUSES WORKED TO REDUCE GLUCOSE LEVELS". PATIENT HEALTH IMPACT IS DEATH. THE REPORTING HCP HAS BEEN IN CONTACT WITH INSULET ON (B)(6)2025 TO ADVISE THAT THE PATIENT WAS FOUND DEAD IN HIS RESIDENCE IN (B)(6) 2024 AND THE CORONER HAS ADVISED THAT THE CAUSE OF DEATH IS DIABETIC KETOACIDOSIS. THE HCP RECEIVED THE DEVICE ON (B)(6)2025. THE PATIENT'S BLOOD GLUCOSE LEVELS WERE REPORTED TO BE ABOVE 27.8 MMOL/L (500.4 MG/DL) BEFORE THEY DECEASED. Manufacturer Narrative: THE DEVICE HAS NOT BEEN RETURNED TO INSULET. THE DEVICE IS WITH THE REPORTING HEALTHCARE PROFESSIONAL. IF THE DEVICE IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED WITH THE INVESTIGATION RESULTS. WE ARE UNABLE TO DETERMINE IF ANY PRODUCT CONDITION COULD HAVE CONTRIBUTED TO THE REPORTED DEATH. NO LOT RELEASE RECORDS WERE REVIEWED, AS THE PRODUCT LOT NUMBER WAS NOT PROVIDED. THE INVESTIGATION IS ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21545803&pc=MIH | 3007284313-2025-03815 | 2025-02-26 05:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-03-07 05:00 | MIH | GORE® EXCLUDER® AAA ENDOPROSTHESIS | Insufficient Information | Insufficient Information | P020004 |  | 1 | Event Description: THE FOLLOWING CONFERENCE ABSTRACT WAS REVIEWED: RAKEMAA, L., BJÖRKMAN, P., LAUKONTAUS, S., AHO, P., TULAMO, R., LAINE, M., TURKKI, M., VENERMO M. (PUBLISHING YEAR UNKNOW). "ABSTRACT P-062 LONG TERM DEVICE PERFORMANCE DIFFERENCES AFTER ENDOVASCULAR AORTIC REPAIR." THE AIM OF THE STUDY WAS TO COMPARE WAS TO COMPARE THE REAL-LIFE PERFORMANCE OF THE THREE MOST USED ENDOGRAFTS WITH OVER TWO DECADES OF SURVEILLANCE. 461 CONSECUTIVE PATIENTS WHO UNDERWENT EVAR FOR INTACT INFRARENAL AAA IN SINGLE CENTRE BETWEEN 2000 AND 2016. DATA WAS COLLECTED UNTIL THE END OF 2023 AND INCLUDED GORE®¿EXCLUDER®¿AAA ENDOPROSTHESIS [N=127] AND OTHER DEVICES. IN THE CLINICAL FOLLOW UP: ONE-, THREE- AND FIVE-YEAR FREEDOM FROM FREEDOM FROM ANEURYSM RELATED DEATH FOR GORE EXCLUDER WAS 99.2%, 97.9%, AND 97.9%. Manufacturer Narrative: A1: NO PATIENT SPECIFIC DETAILS HAVE BEEN PROVIDED. THEREFORE, THE PATIENT IDENTIFIER REFLECT THE W. L. GORE INTERNAL CASE NUMBER. G2: LITERATURE, RAKEMAA, L., BJÖRKMAN, P., LAUKONTAUS, S., AHO, P., TULAMO, R., LAINE, M., TURKKI, M., VENERMO M. (PUBLISHING YEAR UNKNOW). "ABSTRACT P-062 LONG TERM DEVICE PERFORMANCE DIFFERENCES AFTER ENDOVASCULAR AORTIC REPAIR." H3: REVIEW OF THE MANUFACTURING AND STERILIZATION RECORDS COULD NOT BE PERFORMED AS A VALID LOT NUMBER WAS NOT PROVIDED. H6, CODE C21: REVIEW OF THE MANUFACTURING RECORDS COULD NOT BE PERFORMED AS A VALID LOT NUMBER WAS NOT PROVIDED. ENGINEERING EVALUATION COULD NOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED. THE INITIAL REPORTER HAS BEEN CONTACTED IN ORDER TO RECEIVE MORE INFORMATION ON THE EVENT AND PATIENT DETAILS. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. Manufacturer Narrative: H5: UPDATED THE DATE TO APRIL 24, 2025 AS ADDITIONAL INFORMATION WAS PROVIDED IN THE INVESTIGATION FROM GORE. H6, UPDATED CODES: D15: THERE IS INSUFFICIENT INFORMATION AVAILABLE FOR GORE TO REASONABLY DRAW CONCLUSIONS RELATED TO ASPECTS OF THE EVENT, THEREFORE CONCLUSION CODE ¿D15: CAUSE NOT ESTABLISHED¿ IS BEING USED. INSUFFICIENT INFORMATION MAY INCLUDE LIMITED OR MISSING RELEVANT MEDICAL RECORDS. CODE C19: REVIEW OF THE MANUFACTURING AND STERILIZATION RECORDS COULD NOT BE PERFORMED AS A VALID LOT NUMBER WAS NOT PROVIDED. ENGINEERING EVALUATION COULD NOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED. THE INITIAL REPORTER HAS BEEN CONTACTED IN ORDER TO RECEIVE MORE INFORMATION ON THE EVENT AND PATIENT DETAILS. CODE B13: QUERY FOR THE INFORMATION WITH DEVICE DETAILS WAS REQUESTED. CODE B11: THE HISTORICAL DATA ANALYSIS IS STARTED, WAITING FOR THE CONCLUSIVE RESULTS AND WILL BE DOCUMENTED. Manufacturer Narrative: UPDATED G3. H6: ADDED INVESTIGATION FINDINGS CODE C19, CODE C21 IS NO LONGER APPLICABLE. ADDED CODE D15, CODE D16 IS NO LONGER APPLICABLE. ADDED TYPE OF INVESTIGATION CODE B11 AND B20. ADDED HEALTH IMPACT CODE: F02. DUE TO AN UNKNOWN LOT/SERIAL NUMBER AND NO DEVICE RETURN, AN INVESTIGATION COULD NOT BE PERFORMED. THE REPORTED EVENT IN THIS ARTICLE HAS NO CLINICAL IMAGES NOR EVIDENCE TO PERFORM INVESTIGATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION ABOUT THIS EVENT. NO ADDITIONAL INFORMATION WAS PROVIDED. ACCORDING TO THE GORE® EXCLUDER® AAA ENDOPROSTHESIS INSTRUCTIONS FOR USE (IFU), IT STATES THE FOLLOWING: POTENTIAL DEVICE OR PROCEDURE-RELATED ADVERSE EVENTS ADVERSE EVENTS THAT MAY OCCUR AND / OR REQUIRE INTERVENTION OR ADDITIONAL INTRAOPERATIVE PROCEDURE TIME INCLUDE BUT ARE NOT LIMITED TO: ANEURYSM RUPTURE AND DEATH NO EVENT DATES WERE FOUND FOR THE PATIENTS IN THIS ARTICLE (ABSTRACT). B3: THE DATE INCIDENT WILL BE SELECTED FOR THE END OF STUDY OR THE PUBLICATION DATE. IN THIS ABSTRACT THE DATE OF PUBLISHING IS MISSING, SO THE DATE EVENT COULD NOT BE DETERMINED AND SUBSEQUENTLY IT WAS SELECTED THE DATE OF THE INFORMATION CAME TO GORE. Manufacturer Narrative: UPDATED G3. H6 SECTION: INVESTIGATION CONCLUSIONS, CODE ADDED D12. ATTACHMENTS SECTION: ADDED THE RELATED LITERATURE TO THIS REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21546159&pc=OZD | 1220648-2025-26565 | 2025-02-12 05:00 | Death | ABIOMED, INC. | 2025-03-07 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT EXPERIENCED BLEEDING FROM THEIR IMPELLA CP ACCESS SITE DURING SUPPORT. MANUAL PRESSURE WAS APPLIED AND THE SHEATH WAS MANIPULATED TO ACHIEVE HEMOSTASIS. THREE UNITS PACKED RED BLOOD CELLS, TWO UNITS' FRESH FROZEN PLASMA, AND TWO UNIT'S PLATELETS WERE TRANSFUSED TO TREAT THE BLEED. SUPPORT WAS SET TO CONTINUE, HOWEVER, DUE TO THE SEVERITY OF THE PATIENT'S DISEASE, THE FAMILY OPTED TO MOVE TO COMFORT CARE. THE PUMP WAS REMOVED AND THE PATIENT LATER PASSED AWAY. A RELATIONSHIP BETWEEN THE BLEED AND THE PATIENT'S PASSING WAS NOT DEFINED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21546561&pc=OZP | 2032227-2025-132567 | 2025-02-01 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-03-07 05:00 | OZP | MINIMED 780G US SYSTEM BLE CONNECT 3.0 MG/DL | Adverse Event Without Identified Device or Use Problem | High Blood Pressure/ Hypertension; Pneumonia; Renal Failure; Diabetic Ketoacidosis | P160017 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2025. CAUSE OF DEATH WAS DIABETIC KETOACIDOSIS DURING HOSPITAL STAY. OTHER RELEVANT HISTORY, INCLUDING PREEXISTING MEDICAL CONDITIONS: PNEUMONIA, ALSO TRIGGERED HYPERTENSION AND RENAL FAILURE. THE PUMP WAS WORN AT THE TIME OF PASSING. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: MMT-1884, MMT-396A, MMT-332A. THE CUSTOMER WILL DISCONTINUE USE OF THE MMT-1884. MMT-1884 WAS REQUESTED AND CUSTOMER RESPONSE WAS THE DEVICE WILL BE RETURNED. NO PRODUCT RETURN IS REQUIRED FOR MMT-396A. NO PRODUCT RETURN IS REQUIRED FOR MMT-332A.Event Description: UPDATED SUMMARY: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2025 AND THE CAUSE OF DEATH WAS DIABETIC KETOACIDOSIS DURING THE HOSPITAL STAY. OTHER RELEVANT HISTORY, INCLUDING PREEXISTING MEDICAL CONDITIONS SUCH AS PNEUMONIA, ALSO TRIGGERED HYPERTENSION AND RENAL FAILURE. THE PUMP WAS WORN AT THE TIME OF PASSING. IT WAS UNKNOWN WHETHER THE CUSTOMER WAS USING THE AUTO MODE/SMARTGUARD FEATURE AT THE TIME OF THE EVENT. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: MMT-1884, MMT-396A, MMT-332A. THE INSULIN PUMP MMT-1884 WILL BE RETURNED FOR ANALYSIS. NO PRODUCT RETURN IS REQUIRED FOR MMT-396A. NO PRODUCT RETURN IS REQUIRED FOR MMT-332A. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. THE DEVICE WILL BE RETURNED FOR ANALYSIS AND FURTHER INFORMATION WILL FOLLOW ONCE THE ANALYSIS HAS BEEN COMPLETED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21546804&pc=OZD | 1220648-2025-26572 | 2025-02-09 05:00 | Death | ABIOMED, INC. | 2025-03-07 05:00 | OZD | IMPELLA CP | Malposition of Device; Deformation Due to Compressive Stress | Hemolysis | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS). HOWEVER, DURING INSERTION, A KINK IN THE CATHETER WAS NOTED TO HAVE DAMAGED THE CABLE AND AORTIC MONITORING WAS DYSFUNCTIONAL. THE IMPELLA CP DEVICE CONTINUED TO SUPPORT THE PATIENT. TWO DAYS LATER, HEMOLYSIS WAS NOTED WHEN BLOOD DRAWS FROM A-LINE, BUT NO HEMOLYSIS WHEN BLOOD DRAWN FROM V-LINE. ECHOCARDIOGRAM WAS COMPLETED AND SHOWED THE IMPELLA CATHETER IN TOO FAR. THE PHYSICIAN DECIDED NOT TO REPOSITION THE PUMP AT SUCH TIME. TWO DAYS FOLLOWING, THE PATIENT WAS NOTED ON A VERY HIGH DOSE OF LEVOPHED AND VASOPRESSIN. LABS ARE ALL HEMOLYZING. THEREFORE, A NEW SET WAS INITIATED. GOOD FLOW ON THE IMPELLA PLACEMENT WAS OBSERVED ON ECHOCARDIOGRAM. THE OUTCOME OF THE PATIENT AS A RESULT OF THE EVENT IS UNKNOWN. HOWEVER, THE PATIENT CONTINUES IMPELLA MCS.Event Description: ADDITIONAL INFORMATION RECEIVED THE PATIENT EXPIRED WHILE ON SUPPORT. IT IS UNKNOWN IF THE USE OF THE IMPELLA WAS RELATED TO PATIENT DEATH, THUS THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: B2 REVISED TO REFLECT THE ADDITIONAL INFORMATION RECEIVED REGARDING THE PATIENT OUTCOME DEATH. B5 REVISED TO REFLECT THE ADDITIONAL INFORMATION RECEIVED REGARDING THE PATIENT OUTCOME DEATH. D6B EXPLANT DATE ADDED. H1 REVISED TO REFLECT THE ADDITIONAL INFORMATION RECEIVED REGARDING THE PATIENT OUTCOME DEATH. H6 HEALTH EFFECT - IMPACT CODE 1802 ADDED REFLECTING PATIENT OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER. ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE. ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS. ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION. ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿ Manufacturer Narrative: THE INVESTIGATION OF PRODUCT DAMAGE (CANNULA KINK), POSITIONING ISSUES, AND HEMOLYSIS HAS BEEN COMPLETED. THE ROOT CAUSE OF THE HEMOLYSIS WAS MOST LIKELY BIOMATERIAL AS EVIDENCED BY INGESTION PATTERN IN THE LOGS. THE ROOT CAUSE OF THE POSITIONING ISSUES AND PRODUCT DAMAGE (CANNULA KINK) WAS NOT DETERMINED AS NO PRODUCT WAS RETURNED FOR ANALYSIS AND LIMITED INFORMATION RELATED TO TIME OF FAILURE OCCURRENCE WAS PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21546837&pc=FRO | 3005920706-2025-00007 |  | Death | INTEGRA - ACELL(COLUMBIA) | 2025-03-07 05:00 | FRO | MEDIHONEY GEL IN TUBE, .5 FL OZ TWISTOFF - STERILE | Lack of Effect | Pain | K101793 |  | 1 | Event Description: IT WAS REPORTED: "MY MOTHER BOUGHT THIS (MEDIHONEY) FOR MY FATHER'S WOUNDED TOE. IT WAS ONE OF HIS LAST THREE TOES REMAINING. NOT ONLY DID IT NOT WORK, BUT IT ALSO MADE IT WORSE. AMPUTATION WAS IMMINENT. MY FATHER COULDN'T HANDLE ANYMORE. THE PAIN, THE EXHAUSTION, THE STRESS HE WAS PUTTING ON MY MOTHER. ON (B)(6) 2024, HE SHOT HIMSELF IN THE FACE."Event Description: N/A. Manufacturer Narrative: AN INVESTIGATION HAS BEEN INITIATED BASED ON THE REPORTED INFORMATION. UPON COMPLETION OF THE INVESTIGATION, A FOLLOW-UP REPORT WILL BE SUBMITTED. Manufacturer Narrative: FAILURE ANALYSIS/ROOT CAUSE - THIS COMPLAINT DESCRIBES USE OF MEDIHONEY FOR TREATMENT OF A WOUNDED TOE AND STATES THAT IT DID NOT WORK AND "MADE IT WORSE". IN A FOLLOW UP COMMUNICATION, THE CUSTOMER PROVIDED THE SKU 31805, HOWEVER, THEY DID NOT PROVIDE ANY CLARIFICATION REGARDING "MADE IT WORSE." AS THE REPORTER DESCRIBED THE PATIENT'S CONDITION IN THE INITIAL COMPLAINT AS "AMPUTATION WAS IMMINENT". A MEDICAL ASSESSMENT WAS REQUESTED TO INTEGRA'S MEDICAL AFFAIRS DIRECTOR AND THE MEDICAL ASSESSMENT DETERMINED THE ADVERSE EVENT WAS POSSIBLY INFECTION. THEY CONCLUDED THAT A CAUSAL RELATIONSHIP BETWEEN THE DEVICE AND AN ADVERSE EVENT OF INFECTION IS HIGHLY UNLIKELY. FURTHERMORE, THE REPORTER DID NOT PROVIDE ANY INFORMATION RELATED TO CHANGES IN TEXTURE, COLOR, SMELL, OR CONSISTENCY OF THE PRODUCT, NOR PRESENCE OF FUZZY PATCHES OR FOREIGN PARTICULATE, WHICH, IF PRESENT, COULD INDICATE MICROBIAL PROLIFERATION. IN AN ADDITIONAL COMMUNICATION FROM THE CUSTOMER, THEY STATE THEY DO NOT HAVE ACCESS TO ADDITIONAL MEDICAL RECORDS AND THAT THE PRODUCT HAS BEEN DISCARDED. DUE TO THE LACK OF FURTHER INFORMATION REGARDING THE PRODUCT, SUCH AS PHOTOS OR RETURNED PRODUCT, AND FURTHER INFORMATION REGARDING THE PATIENT, SUCH AS MEDICAL HISTORY, THIS COMPLAINT IS UNCONFIRMED. A COMPLAINT ADJUDICATION REPORT WAS INITIATED TO PROVIDE 3RD PARTY MEDICAL ASSESSMENT FROM BOARD-CERTIFIED INFECTION DISEASE SPECIALIST (INITIALS (B)(6)). THIS REVIEW INDICATED THAT BASED ON THE INFORMATION PROVIDED, AND IN THE ABSENCE OF DETAILED CLINICAL DOCUMENTATION, IT IS THE PROFESSIONAL OPINION OF THE OUTSIDE EXPERT THAT EVEN IF THE MEDIHONEY WAS APPLIED FROM A POTENTIALLY COMPROMISED TUBE, IT DID NOT CONTRIBUTE TO ANY OF THE ADVERSE OUTCOME REPORTED IN THE COMPLAINT AS THE PATIENT PRESENTED WITH SIGNIFICANT COMORBIDITIES THAT OFFER MORE PLAUSIBLE EXPLANATIONS FOR THEIR CLINICAL DETERIORATION OR DEATH. POST THE COMPLETION OF THIS ASSESSMENT, FURTHER INFORMATION PROVIDED BY THE CUSTOMER INDICATES THAT THIS PRODUCT COULD NOT BE PART OF THE RECALL, FOR WHICH THE EXPERT MADE THIS ASSESSMENT. THE CUSTOMER HAD PROVIDED ADDITIONAL INFORMATION ON THIS COMPLAINT WITH THE ORIGINAL ORDER AND SHIPPING INFORMATION FOR THE PRODUCT, PURCHASED ON AMAZON IN MARCH 2023. THE CUSTOMER, UPON THE INITIATION OF THE COMPLAINT, HAD THE UNDERSTANDING THAT THE PRODUCT IN QUESTION WAS PART OF THE RECALL FOR SKU 31815 (LOTS 2328, 2332 AND 2333). THE LOT INFORMATION WAS NOT PROVIDED FOR THIS COMPLAINT HOWEVER, THE EARLIEST LOT OF THE THREE IN THE SCOPE OF THE RECALL, LOT 2328, WAS RELEASED ON 18AUG2023. THIS IS 5 MONTHS AFTER THE PRODUCT IN QUESTION WAS PURCHASED, MEANING THE COMPLAINT PRODUCT COULD NOT HAVE BEEN IN SCOPE OF THE RECALL LOTS, EVEN WITHOUT CONFIRMATION OF WHICH LOT NUMBER THE PRODUCT IS. THE PHOTOGRAPH OF THE PRODUCT PURCHASED ALSO INDICATES THAT IT IS SKU 31805, WHICH IS ALSO OUTSIDE THE SCOPE OF THE RECALL. ADDITIONALLY, THE CUSTOMER PROVIDED PURCHASE DETAILS SHOWING THE PRODUCT WAS PURCHASED FROM AMAZON, SELLER PREFERRED PHARMACY PLUS. FEDERAL LAW REQUIRES THAT THE PRODUCT BE SOLD AND USED ONLY UNDER THE SUPERVISION OF A PHYSICIAN PER THE IFU. THERE ARE NO LEGAL SALES OF ANY MEDIHONEY PRODUCTS ON AMAZON AND INTEGRA HAS CONFIRMED THAT AMAZON IS SELLING COUNTERFEIT PRODUCTS UNDER THE PRODUCT NAME OF MEDIHONEY. WITHOUT THE PRODUCT RETURNED FOR EVALUATION OR PHOTOGRAPHS PROVIDED, THERE IS NO WAY TO CONFIRM IF THE PRODUCT IS A GENUINE INTEGRA DEVICE. INTEGRA REGULATORY HAS BEEN IN COMMUNICATION WITH THE FDA ABOUT THE INAPPROPRIATE SALES OF MEDIHONEY. THERE IS NO CONFIRMATION IF THE CUSTOMER FOLLOWED THE INSTRUCTIONS FOR USE ON THE INDICATIONS, PRECAUTIONS, OR INSTRUCTIONS. IF THE CUSTOMER DID NOT FOLLOW GUIDANCE FROM THEIR DOCTOR, OR INSTRUCTIONS FOR USE OF THE PRODUCT, INFECTION MAY BE CAUSED BY ANY NUMBER OF SOURCES UNRELATED TO THE USE OF A MEDIHONEY PRODUCT. A CORRECTIVE AND PREVENTATIVE ACTION (CAPA) HAS BEEN CREATED FOR PRESCRIPTION/OTC DISTRIBUTION ISSUE TO ADDRESS HOW MEDIHONEY WAS DISTRIBUTED FOR SALE WITHOUT A PRESCRIPTION. ADDITIONAL INFORMATION RECEIVED: THE PATIENT'S DAUGHTER STATED: "MY FATHER WAS 82 YEARS OLD. I DON'T HAVE ANY OF MY FATHER'S MEDICAL INFORMATION AVAILABLE. HE WAS A DIABETIC AND HAD ASTHMA. NO OTHER PROBLEMS THAT I KNOW OF. I CAN'T PROVIDE MEDICAL RECORDS OR ANY OF THAT INFORMATION."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21547226&pc=OZD | 1220648-2025-26574 | 2025-02-05 05:00 | Death | ABIOMED, INC. | 2025-03-07 05:00 | OZD | AUTOMATED IMPELLA CONTROLLER | Failure to Read Input Signal | Cardiac Arrest | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT TURNED DOWN FOR CORONARY ARTERY BYPASS GRAFT SURGERY PRESENTED FOR A HIGH-RISK PERCUTANEOUS CORONARY INTERVENTION (PCI) WITH IMPELLA CP MECHANICAL CIRCULATORY SUPPORT (MCS). DURING INITIAL PURGE SETUP, THE AUTOMATED IMPELLA CONTROLLER DISPLAYED A PURGE DISC NOT DETECTED MESSAGE. THE AIC WAS EXCHANGED AND NO FURTHER ISSUES WITH THE PURGE SYSTEM WERE ENCOUNTERED. THE PATIENT, HOWEVER, WOULD EVENTUALLY PASS AWAY. THE IMPELLA CP HAD BEEN PLACED AND STARTED IN AUTO WITH DECISION TO PLACE PERFORMANCE LEVEL P-5 AND EVENTUALLY INCREASING TO P-7 - 3.1L FOR THE PROCEDURE. CONTRALATERAL ACCESS WAS OBTAINED FOR PCI AND EXTENSIVE EFFORT AND TIME FOR TEAM TO CONFIDENTLY WIRE THE LEFT ANTERIOR DESCENDING (LAD) AND LEFT CIRCUMFLEX (LCX) ARTERIES; ULTIMATELY ACCESS WAS SUCCESSFUL. THE LEFT MAIN AND LAD WERE THEN BALLOONED AND STENTED FROM MID-LAD TO PROXIMAL LEFT MAIN WITH "GREAT" RESULT. AFTER THE PCI WAS COMPLETED, REPOSITIONING SHEATH WAS THEN TIPPED INTO THE 16FR. COOK SHEATH AND PER THE PLAN IT WAS LEFT IN FOR THE FOLLOWING 24-48 HOURS FOR THE PATIENT TO REST ON SUPPORT. THE FOLLOWING DAY AFTER START OF THE IMPELLA MCS, CONTINUOUS RENAL REPLACEMENT THERAPY (CRRT) WAS STARTED TO HELP WITH FLUID REMOVAL AND FUNCTION. THE PATIENT CONTINUED ON IMPELLA MCS AND CRRT FOR TWO ADDITIONAL DAYS. LAST DAY, OVERNIGHT, THE PATIENT WAS INTUBATED, MAXED OUT ON PRESSORS AND LATER THIS DAY THE PATIENT WOULD EXPIRE. FAMILY WITHDREW CARE. THE CAUSE OF DEATH IS UNKNOWN AT THIS TIME, AND THE TEMPORARY PERIOD TO ADDRESS THE PURGE DISC ISSUE COMPROMISING SUPPORT IS AN UNCLEAR ASSOCIATION WITH THE EVENT OUTCOME AND THEREFORE CANNOT BE DISSOCIATED. AFTER EXCHANGE OF THE AIC, THE PATIENT WOULD CONTINUE WITH IMPELLA MCS FOR AN ADDITIONAL THREE DAY UNTIL HIS/HER DEMISE. MEDICAL SAFETY REVIEWED THE EVENT NOTED THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE AIC AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE AUTOMATED IMPELLA CONTROLLER (AIC) DATA LOGS AND CONSOLE WERE RETURNED FOR EVALUATION AND ANALYSIS. DATA ANALYSIS OF THE CONSOLE LOGS FROM THE REPORTED DAY OF THE EVENT REVEALED THAT THE CONSOLE ISSUED PURGE DISC NOT DETECTED ALARM SEVERAL TIMES. THE CONSOLE WAS EXAMINED, AND A BROKEN PURGE DISC FLAG WAS IDENTIFIED. A PRODUCT HISTORY REVIEW WAS COMPLETED, AND NO ACTION WAS REQUIRED AS A RESULT OF THE REVIEW. IN CONCLUSION, THE PURGE DISC FLAG WAS OBSERVED TO BE BROKEN AND WAS THE ROOT CAUSE OF THE REPORTED PURGE DISC NOT DETECTED ALARM ISSUE. APPLICABLE REPAIRS AND FULL FUNCTIONAL CHECK ON CONSOLE AND OPTICAL SYSTEM WERE COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21547311&pc=NIQ | 2124215-2025-09963 | 2021-04-01 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-03-07 05:00 | NIQ | PROMUS PREMIER | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. IN (B)(6) 2019, THE SUBJECT WAS REFERRED FOR CARDIAC CATHETERIZATION. THE FIRST TARGET LESION WAS LOCATED IN THE PROXIMAL LEFT ANTERIOR DESCENDING ARTERY (LAD) EXTENDING UP TO DISTAL LAD WITH 99% STENOSIS AND WAS 64 MM LONG AND A REFERENCE VESSEL DIAMETER OF 3.00 MM. THE FIRST TARGET LESION WAS TREATED WITH PRE-DILATATION AND PLACEMENT OF 2.50 MM X 24 MM, OVERLAPPED WITH ANOTHER 2.75 MM X 24 MM AND 3.00 MM X 28 MM PROMUS PREMIER STENTS. POST-DILATATION WAS PERFORMED WITH 0% RESIDUAL STENOSIS. TWENTY-ONE DAYS LATER, IN (B)(6) 2019, THE SUBJECT WAS DISCHARGED ON ASPIRIN AND CLOPIDOGREL. ON AN UNKNOWN DATE, THE SUBJECT DIED. AT THE TIME OF THE EVENT, THE SUBJECT WAS ON ASPIRIN AND CLOPIDOGREL. THE SUBJECT WAS HOSPITALIZED AND TREATED MEDICALLY. THE PRIMARY CAUSE OF DEATH WAS UNKNOWN; HOWEVER, IT WAS CLASSIFIED AS CARDIAC. IT IS UNKNOWN WHETHER AN AUTOPSY WAS PERFORMED, AND A DEATH CERTIFICATE IS NOT AVAILABLE FOR REVIEW. Manufacturer Narrative: B3. DATE OF EVENT- USED THE FIRST DATE OF THE MONTH OF THE AWARE DATE AS NO EVENT DATE WAS PROVIDED. E1. INITIAL REPORTER PHONE: (B)(6).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21547313&pc=NIQ | 2124215-2025-09970 | 2021-04-01 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-03-07 05:00 | NIQ | PROMUS PREMIER | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. IN (B)(6) 2019, THE SUBJECT WAS REFERRED FOR CARDIAC CATHETERIZATION. THE FIRST TARGET LESION WAS LOCATED IN THE PROXIMAL LEFT ANTERIOR DESCENDING ARTERY (LAD) EXTENDING UP TO DISTAL LAD WITH 99% STENOSIS AND WAS 64 MM LONG AND A REFERENCE VESSEL DIAMETER OF 3.00 MM. THE FIRST TARGET LESION WAS TREATED WITH PRE-DILATATION AND PLACEMENT OF 2.50 MM X 24 MM, OVERLAPPED WITH ANOTHER 2.75 MM X 24 MM AND 3.00 MM X 28 MM PROMUS PREMIER STENTS. POST-DILATATION WAS PERFORMED WITH 0% RESIDUAL STENOSIS. TWENTY-ONE DAYS LATER, IN (B)(6) 2019, THE SUBJECT WAS DISCHARGED ON ASPIRIN AND CLOPIDOGREL. ON AN UNKNOWN DATE, THE SUBJECT DIED. AT THE TIME OF THE EVENT, THE SUBJECT WAS ON ASPIRIN AND CLOPIDOGREL. THE SUBJECT WAS HOSPITALIZED AND TREATED MEDICALLY. THE PRIMARY CAUSE OF DEATH WAS UNKNOWN; HOWEVER, IT WAS CLASSIFIED AS CARDIAC. IT IS UNKNOWN WHETHER AN AUTOPSY WAS PERFORMED, AND A DEATH CERTIFICATE IS NOT AVAILABLE FOR REVIEW. Manufacturer Narrative: B3. DATE OF EVENT- USED THE FIRST DATE OF THE MONTH OF THE AWARE DATE AS NO EVENT DATE WAS PROVIDED. E1. INITIAL REPORTER PHONE: (B)(6).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21547314&pc=NIQ | 2124215-2025-09975 | 2021-04-01 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-03-07 05:00 | NIQ | PROMUS PREMIER | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. IN (B)(6) 2019, THE SUBJECT WAS REFERRED FOR CARDIAC CATHETERIZATION. THE FIRST TARGET LESION WAS LOCATED IN THE PROXIMAL LEFT ANTERIOR DESCENDING ARTERY (LAD) EXTENDING UP TO DISTAL LAD WITH 99% STENOSIS AND WAS 64 MM LONG AND A REFERENCE VESSEL DIAMETER OF 3.00 MM. THE FIRST TARGET LESION WAS TREATED WITH PRE-DILATATION AND PLACEMENT OF 2.50 MM X 24 MM, OVERLAPPED WITH ANOTHER 2.75 MM X 24 MM AND 3.00 MM X 28 MM PROMUS PREMIER STENTS. POST-DILATATION WAS PERFORMED WITH 0% RESIDUAL STENOSIS. TWENTY-ONE DAYS LATER, IN (B)(6) 2019, THE SUBJECT WAS DISCHARGED ON ASPIRIN AND CLOPIDOGREL. ON AN UNKNOWN DATE, THE SUBJECT DIED. AT THE TIME OF THE EVENT, THE SUBJECT WAS ON ASPIRIN AND CLOPIDOGREL. THE SUBJECT WAS HOSPITALIZED AND TREATED MEDICALLY. THE PRIMARY CAUSE OF DEATH WAS UNKNOWN; HOWEVER, IT WAS CLASSIFIED AS CARDIAC. IT IS UNKNOWN WHETHER AN AUTOPSY WAS PERFORMED, AND A DEATH CERTIFICATE IS NOT AVAILABLE FOR REVIEW. Manufacturer Narrative: B3. DATE OF EVENT- USED THE FIRST DATE OF THE MONTH OF THE AWARE DATE AS NO EVENT DATE WAS PROVIDED. E1. INITIAL REPORTER PHONE: (B)(6).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21547374&pc=BTL | 3012307300-2025-02588 | 2025-01-17 05:00 | Death | SMITHS MEDICAL CZECH REPUBLIC A.S. | 2025-03-07 05:00 | BTL | PARAPAC PLUS | Defective Alarm; Contamination of Device Ingredient or Reagent | Respiratory Failure | K123957 |  | 1 | Event Description: IT WAS REPORTED THAT THE FILTER ON THE DEVICE WHERE THE CIRCUIT WAS LINKED TO WAS NOT WORKING AND NOT ALARMING. THE PATIENT WAS REPORTED TO HAVE CODED AND EXPIRED DUE TO ACUTE RESPIRATORY FAILURE. THE DATE OF DEATH WAS NOT PROVIDED. ADDITIONAL INFORMATION WAS REQUESTED ON THE CAUSE OF DEATH. TO DATE NO FURTHER INFORMATION HAS BEEN PROVIDED. A CLINICAL REVIEW IS IN PROGRESS. Manufacturer Narrative: D3: CORRECTED. H6: EVALUATION CODES UPDATED. DEVICE EVALUATION: ONE DEVICE WAS RECEIVED FOR EVALUATION. THE DEVICE WAS VISUALLY AND FUNCTIONALLY TESTED BUT THE CUSTOMER REPORTED PROBLEM WAS NOT ABLE TO BE VERIFIED DURING TESTING BOTH WITH AND WITHOUT THE BACTERIAL FILTER. NO ACTION WILL BE TAKEN TO ADDRESS THE REPORTED ISSUE. SERVICE HISTORY REVIEW IDENTIFIED THERE WAS NO INDICATION THAT THE COMPLAINT WAS RELATED TO A SERVICE OF THE DEVICE WITHIN THE REVIEW PERIOD. Manufacturer Narrative: B2 DATE OF DEATH IS UNKNOWN. INVESTIGATION INCLUDING ROOT CAUSE ANALYSIS IS IN PROGRESS. A SUPPLEMENTAL MDR WILL BE FILED AS NECESSARY IN ACCORDANCE WITH 21 CFR 803.56 WHEN ADDITIONAL REPORTABLE INFORMATION BECOMES AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21547422&pc=DTB | 2649622-2025-06624 | 2024-12-03 05:00 | Death | MPRI | 2025-03-07 05:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Pacing Problem; Under-Sensing | Insufficient Information; No Clinical Signs, Symptoms or Conditions | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WAS RIGHT ATRIAL (RA) LEAD UNDERSENSING TRIGGERING DEVICE-CLASSIFIED FALSE TERMINATION OF ATRIAL TACHYCARDIA/ATRIAL FIBRILLATION EVENTS AND RESULTED IN INAPPROPRIATE ATRIAL PACING. THE LEAD REMAINS IN USE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT SUBSEQUENTLY EXPIRED THREE MONTHS LATER. THE CAUSE OF DEATH WAS NOT PROVIDED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: 429678 LEAD IMPLANTED: (B)(6) 2014 AND 6947M62 LEAD IMPLANTED: (B)(6) 2013. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21547444&pc=NKM | 2135147-2025-01154 | 2016-01-01 05:00 | Death | ABBOTT MEDICAL | 2025-03-07 05:00 | NKM | MITRACLIP® SYSTEM CLIP DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P100009 |  | 1 | Event Description: THE ARTICLE "OUTCOME IMPROVEMENT WITH LAST-GENERATION DEVICES IN MITRAL TRANSCATHETER EDGE-TO-EDGE REPAIR: INSIGHTS FROM THE REAL-WORLD MITRACLIP FLORENCE REGISTRY" WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE SINGLE CENTER STUDY, TO EVALUATE THE SAFETY AND EFFICACY OF THE LATEST GENERATION OF MITRACLIP COMPARED TO EARLIER MODELS IN THE REAL-WORLD MITRACLIP FLORENCE REGISTRY. THE PRIMARY EFFICACY ENDPOINT WAS A COMPARISON IN TERMS OF THE RATE OF SUCCESSFUL PROCEDURES, THE TIME TO DEVICE DEPLOYMENT AND THE DURATION OF THE HOSPITAL STAY. THE SECONDARY SAFETY ENDPOINT REGARDED LONG-TERM ALL-CAUSE MORTALITY AND HOSPITALIZATION FOR HEART FAILURE. DEVICES MENTIONED INCLUDE MITRACLIP. THE ARTICLE CONCLUDED THAT THE MITRACLIP G4 SYSTEM IN THE REAL WORLD FOR THE TREATMENT OF SEVERE MR IS SAFE AND EFFECTIVE, ACHIEVING IMMEDIATE AND DURABLE PROCEDURAL SUCCESS, ACCOMPANIED BY AN IMPROVED NYHA FUNCTIONAL CLASS. MOREOVER, A BETTER LONG-TERM SURVIVAL RATE WAS OBSERVED, ALONG WITH A COMPARABLE HIGH RATE OF RECURRENT HF HOSPITALIZATION, REFLECTING A HIGH COMORBIDITY BURDEN IN THIS FRAIL POPULATION. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS NAZARIO CARRABBA AT CAREGGI HOSPITAL INSTITUTION WITH CORRESPONDING EMAIL N.CARRABBA@VIRGILIO.IT]. THE TIME FRAME OF THE STUDY WAS JANUARY 2016 TO JUNE 2022. A TOTAL OF 131 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH 131 RECEIVED AN ABBOTT DEVICE. AS THIS EVENT IS FROM A LITERATURE REVIEW, THERE IS NO RELEVANT PATIENT INFORMATION (DATE OF BIRTH, AGE, GENDER, WEIGHT, AND MEDICAL HISTORY) TO REPORT. COMORBIDITIES INCLUDED DYSLIPIDEMIA, DIABETES, HYPERTENSION, ATRIAL FIBRILLATION. (B)(6), UNK MITRACLIP: PERI- AND POST-PROCEDURAL COMPLICATIONS INCLUDED ARRHYTHMIA, CARDIOGENIC SHOCK, DEATH, HEART FAILURE, PROLONGED HOSPITALIZATION. Manufacturer Narrative: B2: DATE OF DEATH ESTIMATED. B3: DATE OF EVENT ESTIMATED. D4: THE UDI IS UNKNOWN DUE TO THE PART/LOT NUMBER WAS NOT PROVIDED. SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF MITRACLIP DEVICE WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING COMORBIDITIES INCLUDED DYSLIPIDEMIA, DIABETES, HYPERTENSION, ATRIAL FIBRILLATION. COMPLICATIONS REPORTED INCLUDED ARRHYTHMIA, CARDIOGENIC SHOCK, DEATH, HEART FAILURE, PROLONGED HOSPITALIZATION; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21547569&pc=NVN | 2182208-2025-01176 | 2025-01-01 05:00 | Death | MEDTRONIC, INC. | 2025-03-07 05:00 | NVN | SELECTSECURE MRI SURESCAN | Appropriate Term/Code Not Available | Insufficient Information | P030036 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING LEFT BUNDLE BRANCH AREA PACING (LBBAP). THE AUTHORS DESCRIBED PATIENT DEATHS; HOWEVER, THE CAUSES OF DEATH WERE UNKNOWN AND DESCRIBED AS ALL-CAUSE. THERE WERE PATIENTS WHO EXPERIENCED SEPTAL PERFORATIONS, HEMO/PNEUMOTHORAX, INFECTIONS, DEEP VEIN THROMBOSIS, AND HEART FAILURE RELATED HOSPITALIZATIONS. THERE WERE UNKNOWN ACUTE COMPLICATIONS POST IMPLANT AND OTHER PROCEDURE RELATED COMPLICATIONS IN THE FOLLOW UP PERIOD. THERE WERE LEADS WHICH EXHIBITED DISLODGMENTS WHICH INCLUDED BOTH MACRO DISLODGEMENTS IN WHICH VISIBLE DISPLACEMENT OF THE LEAD WAS OBSERVED ON FLUOROSCOPY WITH LOSS OF CAPTURE; AND MICRO DISLODGEMENTS IN WHICH NO LEAD DISPLACEMENT WAS OBSERVED ON FLUOROSCOPY BUT LOSS OF LBBAP OCCURRED. SOME OF THE PATIENTS UNDERWENT LEAD REPOSITIONING. THE STATUS OF THE LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/69 YEARS OLD. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: LEFT BUNDLE BRANCH AREA PACING VERSUS ENDOCARDIAL RESYNCHRONIZATION IN PATIENTS WITH HEART FAILURE. JOURNAL OF C ARDIOVASCULAR ELECTROPHYSIOLOGY. 2025; 36:135¿143. DOI: 10.1111/JCE.16479. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21547864&pc=NKM | 2135147-2025-01165 | 2018-04-01 04:00 | Death | ABBOTT MEDICAL | 2025-03-07 05:00 | NKM | MITRACLIP® SYSTEM CLIP DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P100009 |  | 1 | Event Description: THE ARTICLE "INCREASED AFTERLOAD IN PATIENTS WITH ACUTE REDUCTION IN LEFT VENTRICULAR EJECTION FRACTION FOLLOWING MITRAL VALVE TRANSCATHETER EDGE-TO-EDGE REPAIR" WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE SINGLE CENTER STUDY, TO EVALUATE THE CARDIOMECHANICAL PARAMETERS, INCLUDING END-SYSTOLIC ELASTANCE (EES) AND ARTERIAL ELASTANCE (EA), TO EXPLORE THEIR ASSOCIATION WITH ACUTE REDUCTION (ARD) FOLLOWING MITRAL TRANSCATHETER EDGE-TO-EDGE REPAIR (TEER). DEVICES MENTIONED INCLUDE MITRACLIP. THE ARTICLE CONCLUDED THAT HEMODYNAMIC AND CARDIO-MECHANICAL PARAMETERS ASSOCIATED WITH ARD IN LEFT VENTRICULAR EJECTION FRACTION (LVEF) FOLLOWING SUCCESSFUL M- TEER WERE ASSESSED. A SUBSTANTIAL PROPORTION OF PATIENTS EXHIBITED ARD IN LVEF AFTER TEER, WHICH MAY HAVE PROGNOSTIC IMPLICATIONS. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS TOMOO NAGAI AT TOKAI UNIVERSITY SCHOOL OF MEDICINE, ISEHARA, KANAGAWA, JAPAN WITH CORRESPONDING EMAIL NAGAIT@TOKAI.AC.JP]. THE TIME FRAME OF THE STUDY WAS APRIL 2018 AND JULY 2024. A TOTAL OF 49 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH ALL RECEIVED AN ABBOTT DEVICE. THE AVERAGE AGE WAS 81 AND MAJORITY GENDER WAS MALE. COMORBIDITIES INCLUDED HYPERTENSION, DIABETES, DYSLIPIDEMIA, ATRIAL FIBRILLATION, PACEMAKER, CORONARY ARTERY DISEASE, AORTIC VALVE PROSTHESIS, CHRONIC PULMONARY DISEASE, CHRONIC KIDNEY DISEASE, CHRONIC HEART FAILURE, PERIPHERAL ARTERY DISEASE, CEREBRAL VASCULAR DISEASE, HEART FAILURE MEDICATION, DIURETICS DEGENERATIVE MITRAL REGURGITATION, FUNCTIONAL MITRAL REGURGITATION. PERI AND POST-PROCEDURAL COMPLICATIONS INCLUDED DEATH, HEART FAILURE, AND HOSPITALIZATION. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF MITRACLIP WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING HYPERTENSION, DIABETES, DYSLIPIDEMIA, ATRIAL FIBRILLATION, PACEMAKER, CORONARY ARTERY DISEASE, AORTIC VALVE PROSTHESIS, CHRONIC PULMONARY DISEASE, CHRONIC KIDNEY DISEASE, CHRONIC HEART FAILURE, PERIPHERAL ARTERY DISEASE, CEREBRAL VASCULAR DISEASE, HEART FAILURE MEDICATION, DIURETICS DEGENERATIVE MITRAL REGURGITATION, FUNCTIONAL MITRAL REGURGITATION. COMPLICATIONS REPORTED INCLUDED DEATH, HEART FAILURE, HOSPITALIZATION; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. B2 DEATH DATE IS ESTIMATED. B3 EVENT DATE IS ESTIMATED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21548569&pc=OZD | 1220648-2025-26577 | 2025-02-04 05:00 | Death | ABIOMED, INC. | 2025-03-07 05:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Thrombosis/Thrombus | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT THE DAY AFTER UPGRADING TO IMPELLA 5.5, THE PATIENT EXPIRED WHILE STILL RECEIVING ECPELLA SUPPORT. A COMPUTED TOMOGRAPHY SCAN SHOWED SUPERIOR MESENTERIC ARTERY (SMA) OCCLUSION DUE TO A THROMBUS. IT IS UNKNOWN WHAT CAUSE THIS OCCLUSION. AS A RESULT OF THE SMA OCCLUSION THE PATIENT SUBSEQUENTLY DIED. THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER SECTION: DIRECT AORTIC INSERTION SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES "IN PATIENTS WITH TRANSCATHETER AORTIC VALVES POSITION THE IMPELLA SYSTEM CAREFULLY TO AVOID INTERACTION WITH THE TRANSCATHETER AORTIC VALVE PROSTHESIS. UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH." . IN THIS SITUATION, AVOID REPOSITIONING WHILE THE DEVICE IS RUNNING; TURN THE DEVICE TO P0 DURING REPOSITIONING OR ANY MOVEMENT THAT COULD BRING THE OUTLET WINDOWS INTO PROXIMITY TO THE VALVE STENT STRUCTURES. IF THERE IS LOW FLOW OBSERVED IN A PATIENT IMPLANTED WITH A TRANSCATHETER AORTIC VALVE PROSTHESIS, CONSIDER DAMAGE OF THE IMPELLER AND REPLACE THE IMPELLA AS SOON AS POSSIBLE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21548612&pc=DYB | 1220648-2025-26578 | 2025-02-12 05:00 | Death | ABIOMED, INC. | 2025-03-07 05:00 | DYB | KIT, 14FR INTRODUCER, 13&25 CM, ABMD | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT DURING THE IMPELLA CP PROCEDURE THE PATIENT EXPERIENCED BLEEDING. THE SHEATH WAS INSERTED OVER THE .035 WIRE ON A 14FR DILATOR AND WHEN THE DILATOR WAS UNSCREWED FROM THE 13CM SHEATH, BLOOD BEGAN TO EMERGE. MANUAL PRESSURE WAS HELD WHILE THE 13CM SHEATH WAS REMOVED AND REPLACED WITH THE 25CM SHEATH. NO BLEEDING WAS OBSERVED WITH THE NEW SHEATH. HOWEVER, TWO UNITS OF BLOOD WERE GIVEN TO THE PATIENT BECAUSE OF THIS EVENT. THE PATIENT WAS STABLE AND REMAINS ON SUPPORT.Event Description: THE DECISION WAS MADE TO WITHDRAW CARE AND THE PATIENT EXPIRED. IT IS UNKNOWN IF THE USE OF THE IMPELLA WAS RELATED TO PATIENT DEATH, THUS THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM, SECTION: GENERAL PATIENT CARE CONSIDERATIONS: ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT, USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. BASED ON THE CLINICAL DETAILS THE ROOT CAUSE OF THE ACCESS SITE BLEEDING ISSUE WAS MOST LIKELY USE RELATED AS THE 13CM SHEATH AND REPLACED IT WITH THE 25CM SHEATH, NO BLEEDING WAS OBSERVED WITH THE NEW SHEATH. B2 REVISED TO REFLECT THE PATIENT OUTCOME DEATH. B5 REVISED TO REFLECT THE PATIENT OUTCOME DEATH. H1 TYPE OF REPORTABLE EVENT REVISED TO REFLECT THE PATIENT OUTCOME DEATH. H6 REVISED TO REFLECT THE PATIENT OUTCOME DEATH D1 REVISED BRAND NAME AND PRODUCT CODE SINCE THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26578 WAS SUBMITTED IN ACCORDANCE WITH UPDATED PROCEDURES. D2 REVISED COMMON DEVICE NAME SINCE THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26578 WAS SUBMITTED IN ACCORDANCE WITH UPDATED PROCEDURES. D4 REVISED MODEL NUMBER, CATALOG NUMBER, AND LOT NUMBER SINCE THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26578 WAS SUBMITTED IN ACCORDANCE WITH UPDATED PROCEDURES. SERIAL NUMBER, EXPIRATION AND UNIQUE IDENTIFIER (UDI) # SHOULD HAVE BEEN REFLECTED AS BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER. D6B WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26578. D10 ADDED PUMP INFORMATION SINCE THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26578 WAS SUBMITTED IN ACCORDANCE WITH UPDATED PROCEDURES. E4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26578 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. H6 REVISED COMPONENT CODE SINCE THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26578 WAS SUBMITTED IN ACCORDANCE WITH UPDATED PROCEDURES. SINCE THE INITIAL MANUFACTURER DEVICE REPORT NUMBER WAS SUBMITTED IN ACCORDANCE WITH UPDATED PROCEDURES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21548956&pc=KNS | 1528319-2025-00013 | 2024-12-17 05:00 | Death | HOBBS MEDICAL, INC. | 2025-03-07 05:00 | KNS | HOBBS MEDICAL TOUCHSOFT COAGULATOR | Adverse Event Without Identified Device or Use Problem | Abdominal Pain; Unspecified Infection; Perforation | K093079 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT A PATIENT REPORTED ABDOMINAL PAIN AND RETURNED TO THE HOSPITAL ONE DAY AFTER A PROCEDURE WHICH INCLUDED USE OF THE TOUCHSOFT COAGULATOR FOR TREATMENT OF TWO SMALL ARTERIOVENOUS MALFORMATIONS (AVMS) IN THE SECOND PORTION OF THE DUODENUM. UPON RETURN TO THE HOSPITAL, THE PATIENT WAS FOUND TO HAVE AN INFECTION IN THE DUODENUM AND A PERFORATION IN THE THIRD PORTION OF THE DUODENUM WHICH REQUIRED SURGERY. AFTER THE SURGERY, THE PATIENT REMAINED IN THE ICU FOR APPROXIMATELY SIX WEEKS AND THEN PASSED AWAY. Manufacturer Narrative: US ENDOSCOPY IS THE COMPLAINT FILE ESTABLISHMENT AND REPACKAGER/RELABELER FOR THE DEVICE SUBJECT OF THIS EVENT. THE DEVICE IS MANUFACTURED BY HOBBS MEDICAL, INC. AND DISTRIBUTED BY US ENDOSCOPY. ON FOLLOW-UP WITH THE USER FACILITY, US ENDOSCOPY LEARNED THAT THE PATIENT'S MEDICAL HISTORY AND COMORBIDITIES INCLUDED GASTRIC CARCINOID TUMOR, COLONIC POLYPS, RENAL CELL CANCER, HYPERTENSION, AND HYPERLIPIDEMIA. UNRELATED TO USE OF THE TOUCHSOFT COAGULATOR FOR AVM TREATMENT, THE PROCEDURE SUBJECT OF THIS REPORTED EVENT INCLUDED BIOPSY OF A 10MM NODULE IN THE DUODENAL BULB. TWO SMALL NONBLEEDING ANGIODYSPLASIAS WERE ALSO NOTED DURING THE PROCEDURE. THE USER FACILITY STATED THAT THEY DO NOT KNOW EXACTLY WHAT CAUSED THIS EVENT TO HAPPEN; THE DOCTOR STATED HE DID NOT THINK THAT THE PROBE CONTRIBUTED TO THE PATIENT'S PASSING BUT COULD NOT RULE IT OUT. ADDITIONALLY, US ENDOSCOPY LEARNED THAT THE USER FACILITY USED A THIRD PARTY ERBE GENERATOR SET AT 30-40 WATTS WITH THE TOUCHSOFT COAGULATOR DURING THE PROCEDURE. THE SETTING WAS CONTRARY TO THE TOUCHSOFT COAGULATOR INSTRUCTIONS FOR USE (IFU) WHICH STATES, "THE TOUCHSOFT COAGULATOR CAN BE USED WITH...GENERATOR MODEL ERBE, RANGE IN WATTS 45-60." THE IFU STATES, "CAUTION: DO NOT APPLY EXCESSIVE HEAT, FORCE OR PROLONGED APPLICATION TIME, AS THIS MAY CAUSE DAMAGE TO THE DEVICE, STICKING OF THE DEVICE TO THE TISSUE OR OTHER UNDESIRABLE EFFECTS, INCLUDING POSSIBLE PERFORATION." THE IFU FURTHER STATES, "ALWAYS CHOOSE THE LOWEST POWER SETTING AND SHORTEST TIME DURATION (3 SECONDS OR LESS) THAT WILL ACHIEVE THE DESIRED CLINICAL RESULT." THE DEVICE SUBJECT OF THE EVENT WAS DISCARDED BY USER FACILITY PERSONNEL AND IS NOT AVAILABLE FOR EVALUATION. THE USER FACILITY WAS UNABLE TO PROVIDE A LOT NUMBER FOR THE DEVICE SUBJECT OF THE EVENT. A REVIEW OF DISTRIBUTION RECORDS INDICATES THAT THE USER FACILITY HAS ORDERED 5 LOTS OF THIS PRODUCT SINCE 2022. THERE HAVE BEEN NO OTHER COMPLAINTS ASSOCIATED WITH ANY OF THESE 5 LOTS. THE IFU INCLUDES THE FOLLOWING WARNINGS AND CAUTIONS, "WARNING: POTENTIAL ADVERSE EFFECTS INCLUDE, BUT ARE NOT LIMITED TO: RISK OF PERFORATION, HEMORRHAGE, INFECTION, THERMAL BURNS, AND EXPLOSION. CAUTION: ONLY THE USER CAN EVALUATE THE CLINICAL FACTORS INVOLVED WITH EACH PATIENT TO DETERMINE IF THE USE OF THIS DEVICE IS INDICATED. THE USER MUST THEN DECIDE ON THE SPECIFIC TECHNIQUE AND PROCEDURE THAT WILL BEST ACCOMPLISH THE DESIRED CLINICAL EFFECT. DO NOT USE THIS DEVICE FOR ANY PURPOSE OTHER THAN THE STATED INDICATION FOR USE. WARNING: THE USER MUST THOROUGHLY UNDERSTAND THE TECHNICAL PRINCIPLES, CLINICAL TECHNIQUES AND RISKS ASSOCIATED WITH THE USE OF HIGH FREQUENCY ELECTROSURGICAL ENERGY IN GENERAL AND THE CHARACTERISTICS OF THEIR CHOSEN ELECTROSURGICAL GENERATOR IN PARTICULAR BEFORE USING THIS PRODUCT." A US ENDOSCOPY ACCOUNT MANAGER OFFERED THE USER FACILITY IN-SERVICE TRAINING ON THE PROPER USE AND OPERATION OF THE TOUCHSOFT COAGULATOR; HOWEVER, THE USER FACILITY DECLINED. NO ADDITIONAL ISSUES HAVE BEEN REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21549392&pc=OZD | 1220648-2025-26582 | 2025-02-09 05:00 | Death | ABIOMED, INC. | 2025-03-07 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Ischemia; Hemorrhagic Stroke | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK (AMI/CGS) IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS) LIMB ISCHEMIA, ACCESS SITE BLEEDING AND INTRACRANIAL HEMORRHAGE WITH A DEMISED OUTCOME. THE PATIENT WAS TAKEN FOR EMERGENT PERCUTANEOUS CORONARY INTERVENTION (PCI) TO LEFT ANTERIOR DESCENDING (LAD) ARTERY. DURING THE STAY, PATIENT BEGAN TO EXPERIENCE RESPIRATORY DISTRESS, PROGRESSIVELY BECAME WORSE THE FOLLOWING MORNING AND WAS EMERGENTLY INTUBATED AND REQUIRED INOTROPE/PRESSOR SUPPORT. TROPONIN LEVEL WAS GREATER THAN 200, AND THE PHYSICIAN WAS CONCERNED THAT LAD STENT WAS OCCLUDED. THE DECISION WAS MADE TO PLACE AN IMPELLA CP IMPELLA IN THE PRESENCE OF AMI/CGS, WHICH WAS SUCCESSFULLY IMPLANTED WITHOUT INCIDENT. FOLLOWING THE IMPLANT, HOWEVER, BOTH OF THE PATIENT'S LEGS WERE COLD TO THE TOUCH WITH NO PALPABLE PULSES. DISTAL FLOWS WERE OPTIMIZED AND THE ISSUE RESOLVED BY THE FOLLOWING DAY. A COMPUTED TOMOGRAPHY (CT) SCAN THEN SHOWED A SMALL SUBARACHNOID HEMORRHAGE WITH NO RETROPERITONEAL BLEED OR ANY OTHER SOURCE OF BLEEDING. THE CLINICAL CONSULTANT WENT BACK TO THE CT SCANNER TO CHECK AGAIN, AND IT NOTED INTRAPARENCHYMAL HEMORRHAGE. THE PATIENT'S NEUROLOGICAL RESPONSES DECLINED. THE FOLLOWING DAY ANOTHER CT SCAN WAS COMPLETED, AND IT SHOWED STABLE HEMORRHAGE. THE DECISION WAS MADE TO HOLD PLAVIX AND START HEPARIN. A BLEED FROM THE IMPELLA SITE DEVELOPED, AND ONE UNIT OF FRESH FROZEN PLASMA AND ONE UNIT OF PACKED RED BLOOD CELLS WERE GIVEN AS TREATMENT. SUPPORT CONTINUED FOR APPROXIMATELY ANOTHER FOUR DAYS WITH THE IMPLANTED IMPELLA CP PUMP DESPITE THE NOTED CONDITIONS. IT WAS NOTED THE PATIENT ALSO REMAINED ON MECHANICAL VENTILATOR SUPPORT, VENO-ARTERIAL-VENOUS EXTRACORPOREAL MEMBRANE OXYGENATIONS, AND CONTINUOUS RENAL REPLACEMENT THERAPY. AFTER FIVE DAYS OF TOTAL SUPPORT, THE PATIENT'S FAMILY MADE THE DECISION TO WITHDRAW CARE DUE TO THE PATIENT'S LACK OF NEUROLOGICAL PROGRESS. THE PATIENT EXPIRED, AND THE IMPELLA RELATED COMPLICATIONS MAY HAVE POSSIBLY ADDED TO OR COMPOUNDED AN ALREADY COMPLEX SITUATION. MEDICAL SAFETY REVIEW OF THE EVENT NOTED THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: PRECAUTION IMPELLA CP WITH SMARTASSIST SYSTEM, SECTION: GENERAL PATIENT CARE CONSIDERATIONS: ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT, USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS: ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY¿ (INCLUDING VENTRICULAR PERFORATION). SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21549413&pc=DSQ | 3007042319-2025-00628 | 2025-03-03 05:00 | Death | HEARTWARE, INC. | 2025-03-07 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Pumping Problem | Viral Infection; Decreased Appetite | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT A VENTRICULAR ASSIST DEVICE EXHIBITED 100 LOW FLOW ALARMS SINCE (B)(6) 2025. THE PATIENT HAD INFLUENZA A IN THE LAST FEW DAYS, WHICH COINCIDED WITH A DRAMATIC DROP IN FLOW ON (B)(6) 2025. THE HEALTHCARE PROFESSIONAL SUGGESTED THAT THE PATIENT MIGHT NOT BE CONSUMING ENOUGH FLUIDS. FLOW AND PULSATILITY WERE OBSERVED TO BE BELOW NORMAL VALUES. ALSO, IT WAS REPORTED THAT REVIEW OF LOG FILE DATA REVEALED HISTORY OF VARIOUS MOTOR START EVENTS WITH PARAMETERS OUTSIDE OF THE TYPICAL RANGE. THE VAD REMAINS IN USE.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT WAS SEND HOME AFTER THE ROUTINE CHECK.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT WAS SEND HOME AFTER THE ROUTINE CHECK BUT IT WAS HOSPITALIZED LAST WEEK AND KEPT UNDER OBSERVATION. NO TREATMENT WAS INDICATED. LOGFILES WERE SENT AGAIN AND IT APPEARS THAT THE FLOW HAS DROPPED AGAIN. THE PATIENT IS THOUGHT TO BE IN THE TERMINAL PHASE OF DEATH. DECREASED MOVEMENT AND DIFFICULTY EATING HAVE BEEN OBSERVED IN THE PAST FEW DAYS. THERE IS NO PLAN TO SHUT DOWN THE SYSTEM. INSTEAD, THEY PLAN TO CHANGE THE VAD ROTATION SPEED TO 2500 AND THE ALARM LIMIT SETTINGS (LOW FLOW AND HIGH WATT) TO CREATE A CALM ENVIRONMENT FOR THE PATIENT AND RELATIVES.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT PASSED AWAY. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR ADDITIONAL INFORMATION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR DEVICE EVALUATION AND INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: THE VENTRICULAR ASSIST DEVICE (VAD) WAS NOT RETURNED FOR EVALUATION. A REVIEW OF THE MANUFACTURING DOCUMENTATION CONFIRMED THAT THE ASSOCIATED DEVICE MET ALL REQUIREMENTS FOR RELEASE. LOG FILE ANALYSIS REVEALED A MOTOR START EVENT RECORDED ON (B)(6) 2024 AT 15:10:32 IN WHICH THE MOTOR START PARAMETERS WERE ATYPICAL. LOG FILES ALSO REVEALED SEVERAL INTERMITTENT DECREASES IN POWER CONSUMPTION AND ESTIMATED FLOWS, WITH ASSOCIATED DECREASES IN PULSATILITY, STARTING ON (B)(6) 2025. IN ADDITION, ONE-HUNDRED FORTY-SIX (146) LOW FLOW ALARMS LOGGED SINCE (B)(6) 2025. AS A RESULT, THE REPORTED EVENTS WERE CONFIRMED. BASED ON THE INFORMATION AVAILABLE, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. BASED ON RISK DOCUMENTATION, MULTIPLE FACTORS MAY HAVE CONTRIBUTED TO THE LOW FLOW EVENT INCLUDING BUT NOT LIMITED TO THROMBUS AT THE INFLOW CANNULA/OUTFLOW GRAFT, CONSTRICTION AT THE OUTFLOW GRAFT, POOR VAD FILLING, AND/OR INAPPROPRIATE PUMP ROTATIONAL SPEED. BASED ON RISK DOCUMENTATION, THE MOST LIKELY ROOT CAUSE OF ATYPICAL MOTOR START PARAMETERS MAY BE ATTRIBUTED, BUT NOT LIMITED, TO OUTER SHROUD CONTACT THAT CREATES MORE FRICTION AT THE HOUSING TO IMPELLER INTERFACE DURING PUMP STARTUP. PER THE INSTRUCTIONS FOR USE, POSSIBLE ROOT CAUSES OF CHANGES IN PULSATILITY CAN BE ATTRIBUTED TO PATIENT CONDITIONS INCLUDING LEFT VENTRICULAR CONTRACTILITY, RIGHT HEART FUNCTION AND LEFT VENTRICULAR AFTERLOAD, WHICH CAN BE AFFECTED BY PUMP ROTATIONAL SPEED. PER THE INSTRUCTIONS FOR USE, INFECTION AND DEATH ARE KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. THERE WAS NO EVIDENCE THAT THE PATIENT HAD A HISTORY OF SIMILAR ADVERSE EVENTS. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: A CORRECTION SUPPLEMENTAL REPORT IS BEING SUBMITTED AS SECTION B5 WAS REVISED THE ADDITIONAL SECTIONS WERE UPDATED ACCORDINGLY. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE PRODUCT WAS DISTRIBUTED OUTSIDE OF THE UNITED STATES AND D4 UDI IS THEREFORE BLANK INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED AS ADDITIONAL INFORMATION HAS BEING RECEIVED FOR THIS EVENT AND SECTIONS HAVE BEEN UPDATED.MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21549438&pc=NPT | 2025587-2025-01747 | 2024-08-17 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-07 05:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Cardiomyopathy; Non specific EKG/ECG Changes; Myocardial Infarction; Pneumonia; Ventricular Fibrillation; Loss of consciousness | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 1 YEAR, 2 MONTHS, AND 18 DAYS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, DURING A HOSPITALIZATION WHERE THE PATIENT WAS BEING TREATED FOR PNEUMONIA, THE PATIENT'S CONSCIOUSNESS LEVEL DECREASED. AN ELECTROCARDIOGRAM REVEALED AN ST INCREASE. THERE WAS A NOTED INCREASE IN CREATINE KINASE, LACTATE DEHYDROGENASE, AND AMINOTRANSFERASE LEVELS. A MYOCARDIAL INFARCTION WAS SUSPECTED DUE TO A POSITIVE TROPONIN TEST. THE PATIENT WAS TRANSPORTED TO A DIFFERENT HOSPITAL WHERE THEY WERE ADMITTED DUE TO STRESS CARDIOMYOPATHY. UPON ADMISSION, VITALS WERE STABLE. APPROXIMATELY 6 DAYS LATER, TORSADES DE POINTES WAS NOTED. SUBSEQUENTLY, ONE TIME DEFIBRILLATION WAS PERFORMED DUE TO PULSELESS ELECTRICAL ACTIVITY. DUE TO A "DO NOT ATTEMPT RESUSCITATION" ORDER, NO FURTHER CARDIOPULMONARY RESUSCITATION WAS PERFORMED. THE PATIENT'S DEATH WAS CONFIRMED THE NEXT DAY. THE CAUSE OF DEATH WAS REPORTED TO BE VENTRICULAR FIBRILLATION. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21549439&pc=NIK | 2182208-2025-01181 | 2025-02-21 05:00 | Death | MEDTRONIC, INC. | 2025-03-07 05:00 | NIK | MEDTRONIC CRT | Appropriate Term/Code Not Available | Insufficient Information |  |  | 1 | Event Description: IT WAS REPORTED BY A PATIENT'S DEVICE EXHIBITED AN UNSPECIFIED ISSUE AND THE PATIENT HAD PASSED AWAY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21549525&pc=DSQ | 3007042319-2025-00630 | 2025-03-03 05:00 | Death | HEARTWARE, INC. | 2025-03-07 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Sepsis; Lethargy | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WITH A HISTORY OF CARDIOMYOPATHY AND MSSA BACTEREMIA PRESENTED TO THE EMERGENCY ROOM WITH LETHARGY AND SUSPECTED SEPSIS. THE PATIENT HAD ONGOING INFECTION ISSUES, AND THE FAMILY OPTED FOR COMFORT CARE. THE PATIENT PASSED AWAY. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: THE VENTRICULAR ASSIST DEVICE (VAD) (B)(6) WAS NOT RETURNED FOR EVALUATION AS IT REMAINED IMPLANTED. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORDS WAS NOT PERFORMED AS THE REPORTED EVENT AND ANALYSIS ARE NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS INFECTION AND SEPSIS. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE INFECTION AND DEATH ARE KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. BASED ON REVIEW OF PAST ADVERSE EVENTS, THIS PATIENT HAD A HISTORY OF INFECTION. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21549575&pc=NPT | 2025587-2025-01750 | 2023-06-15 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-07 05:00 | NPT | EVOLUT PRO PLUS VALVE | Insufficient Information | Dehydration | P130021 |  | 1 | Event Description: APPROXIMATELY 4 MONTHS AND 1 WEEK AFTER THE PROCEDURE, THE PATIENT WAS TRANSFERRED TO REHABILITATION WHERE THEY EXPERIENCED A LOSS OF APPETITE AND DEHYDRATION. THE PATIENT SUBSEQUENTLY PASSED AWAY. THE CAUSE OF DEATH IS UNKNOWN. PER THE PHYSICIAN, THERE IS NO RELATIONSHIP BETWEEN THE DEATH AND EITHER THE DEVICE OR THE PROCEDURE. Manufacturer Narrative: B3: DATE IS APPROXIMATE. MONTH AND YEAR ARE CONFIRMED VALID. SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21549849&pc=MAX | 3004142400-2025-00069 | 2025-01-09 05:00 | Death | GLOBUS MEDICAL, INC. | 2025-03-07 05:00 | MAX | SABLE | Insufficient Information | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K192115 |  | 1 | Event Description: IT WAS REPORTED A WAS SABLE SPACER WAS REMOVED FROM A PATIENT WHO THEN LATER DIED.Event Description: IT WAS REPORTED A WAS SABLE SPACER WAS REMOVED FROM A PATIENT WHO THEN LATER DIED. Manufacturer Narrative: THE DEVICE WAS UNAVAILABLE FOR EVALUATION. NO DETERMINATIONS COULD BE MADE AS TO THE CAUSE OF THE REPORTED ISSUE. Manufacturer Narrative: THE DEVICE WAS NOT AVAILABLE FOR EVALUATION. IN THE INITIAL SURGERY, THE PATIENT HAD A PRIOR KYPHOPLASTY WITH CEMENT AT T12, WHICH WAS THE SUPERIOR VERTEBRAL BODY OF THE SITE OF INTENDED INTERBODY FUSION. IT IS UNKNOWN WHAT SCREW SYSTEM WAS PRESENT IN THE PATIENT DUE TO THE KYPHOPLASTY. THE INITIAL PROCEDURE WAS ON THURSDAY, (B)(6) 2025. A SABLE SPACER WAS PLACED AT T12-L1 WITH SUPPLEMENTAL FIXATION USING THE REVERE STABILIZATION SYSTEM. IN THE DAYS AFTER SURGERY, THE PATIENT COMPLAINED OF PAIN. A REVISION CASE WAS PERFORMED THE FOLLOWING MONDAY, (B)(6) 2025, TO REMOVE THE IMPLANT. THE IMPLANT WAS REMOVED WITHOUT ISSUE, HOWEVER IN THE FOLLOWING DAYS AFTER SURGERY THE PATIENT PASSED AWAY. IMAGING OF THE SURGICAL SITE WAS PROVIDED. IN THE IMAGING, THE SABLE SPACER APPEARS TO BE IN AN EXPANDED STATE. THE IMPLANT APPEARS TO BE SUBSIDED INTO THE T12 VERTEBRA WITH THE DISTAL TIP OF THE IMPLANT ANGLED CAUDALLY, WITH IMPLANT ENDPLATES NOT PARALLEL TO THE T12 OR L1 VERTEBRAL ENDPLATES. IT APPEARS THAT THE T12 VERTEBRA MAY HAVE FRACTURED, BUT THIS CANNOT BE DEFINITIVELY DETERMINED BASED ON THE PROVIDED IMAGING. NO DETERMINATIONS COULD BE MADE AS TO THE CAUSE OF THE REPORTED ISSUE. THE FOLLOWING SECTIONS HAVE BEEN UPDATED: B4; D6B; G6; H2; H3; H6; H10.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21550535&pc=DSQ | 2916596-2025-01315 | 2019-11-16 05:00 | Death | THORATEC CORPORATION | 2025-03-07 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Unspecified Infection | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE "ORAL AMIODARONE AND PROPRANOLOL IN MAINTENANCE THERAPY OF POST IMPLANTATION TACHYCARDIA: AN OBSERVATIONAL STUDY" THAT A 57-YEAR-OLD PATIENT WAS IMPLANTED WITH HEARTMATE 3 AND PASSED AWAY 12 DAYS AFTER IMPLANTATION. THE PATIENT HAD A HISTORY OF ARRHYTHMIA, WAS NON-ISCHEMIC AND HAD MODERATELY DEPRESSED RIGHT VENTRICULAR FUNCTION. THE PATIENT DID NOT RECEIVE AMIODARONE TREATMENT. CANNULA-RELATED VENTRICULAR TACHYCARDIAS (VTS) WERE OBSERVED IN >50% OF PATIENTS IN THE FIRST MONTH AFTER LEFT VENTRICULAR ASSIST DEVICE (LVAD) IMPLANTATION. IN ADDITION TO CHRONIC VENTRICULAR SCARS, DE NOVO SCARS AT LVAD ENTRY CANNULATION SITES PROVIDED POTENTIAL SUBSTRATES FOR REENTRY ARRHYTHMIA. THUS, VT ORIGINATING FROM THE LATERAL WALL OF THE LEFT VENTRICLE (LV) IN THE EVENT OF NON-MYOCARDITIS, FROM THE LV SEPTUM DUE TO SARCOIDOSIS, AND FROM THE APICAL ANEURYSM-INDUCED SCAR TISSUE IN THE EVENT OF ANTERIOR MYOCARDIAL INFARCTION CAN OCCUR, AND NOT ALL VTS WERE CAUSED BY DE NOVO SCAR TISSUE. THE STUDY CONCLUDED THAT PROPRANOLOL WAS WELL-TOLERATED BY PATIENTS WITH LVAD IMPLANTATION AND COULD BE ADMINISTERED IN HIGHER DOSES. THUS, PROPRANOLOL WAS THE PREFERRED BETA-BLOCKER FOR THIS PATIENT GROUP. THIS MEDICATION CAN BE USED IN HIGH DOSES IN PATIENTS WITH LVAD IMPLANTATION, ACT ON THE CENTRAL NERVOUS SYSTEM RECEPTORS THROUGH ITS LIPOSOLUBLE CHARACTERISTICS, AND PROVIDE SIGNIFICANTLY REDUCED CATECHOLAMINE DISCHARGE BY NON SELECTIVE BETA-BLOCKADE. Manufacturer Narrative: SECTION B3: DATE OF DEATH AND DATE OF EVENT CORRECTED. SECTION E: INITIAL REPORTER INFORMATION CORRECTED. SECTION G2: REPORT SOURCE CORRECTED. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21550811&pc=DSQ | 2916596-2025-01352 | 2025-02-15 05:00 | Death | THORATEC CORPORATION | 2025-03-07 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS FOUND DEAD.Event Description: IT WAS REPORTED THAT THE CAUSE OF DEATH WAS UNKNOWN. IT WAS KNOWN IF THE DEATH CONSIDERED TO BE DEVICE RELATED OR IS THE DEVICE OPERATE AS EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: EVALUATION OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), DID NOT REVEAL ANY DEVICE-RELATED ISSUES. A SPECIFIC CAUSE FOR THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. (B)(6) WAS RETURNED ASSEMBLED WITH THE PUMP CABLE SEVERED APPROXIMATELY 2.0¿ FROM THE PUMP HEADER. THE DISTAL PORTION OF THE PUMP CABLE AND THE MODULAR CABLE WERE NOT RETURNED. APPROXIMATELY 0.5¿ OF THE SEALED OUTFLOW GRAFT WAS RETURNED ATTACHED TO THE PUMP COVER OUTLET PORT. THE OUTFLOW GRAFT BEND RELIEF WAS NOT RETURNED. AN OUTFLOW GRAFT CLIP WAS RETURNED ATTACHED TO THE GRAFT HARDWARE. THE APICAL CUFF WAS RETURNED WITH THE CUFF LOCK FULLY ENGAGED. UPON DISASSEMBLY OF THE PUMP BODY, EXAMINATION OF THE BLOOD-CONTACTING SURFACES REVEALED NO EVIDENCE OF DEVELOPED DEPOSITIONS OR THROMBUS FORMATIONS WITHIN THE DEVICE THAT WOULD HAVE CONTRIBUTED TO A FUNCTIONAL ISSUE DURING SUPPORT. VISUAL INSPECTION OF THE PUMP ROTOR AND ROTOR WELL DID NOT REVEAL ANY OBVIOUS SURFACE SCRATCHES OR DEFECTS. THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) EVENT AND PERIODIC LOG FILES RETRIEVED FROM THE RETURNED DEVICE APPEARED TO CAPTURE THE DEVICE FUNCTIONING AS INTENDED. (B)(6) WAS CLEANED, REBUILT, AND FUNCTIONALLY TESTED ON A MOCK CIRCULATORY LOOP. THE DEVICE OPERATED AS INTENDED AND IN ACCORDANCE WITH MANUFACTURING SPECIFICATIONS. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. G IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21550877&pc=NPS | 2135147-2025-01178 | 2010-01-01 05:00 | Death | ABBOTT MEDICAL | 2025-03-07 05:00 | NPS | TRICLIP DELIVERY SYSTEM (TCDS) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P230007 |  | 1 | Event Description: THE ARTICLE "TRICUSPID REGURGITATION DISEASE STAGES AND TREATMENT OUTCOMES AFTER TRANSCATHETER TRICUSPID VALVE REPAIR" WAS REVIEWED. THE ARTICLE PRESENTED A PROSPECTIVE MULTI CENTER STUDY, TO EVALUATE THE ASSOCIATION BETWEEN DISEASE STAGE AND OUTCOMES FOLLOWING TRICUSPID TRANSCATHETER EDGE-TO-EDGE REPAIR (T-TEER). DEVICES MENTIONED INCLUDE TRICLIP AND PASCAL. THE ARTICLE CONCLUDED THAT PATIENTS PRESENTING IN AN INTERMEDIATE STAGE MAY RESPOND FAVORABLY TO T-TEER IN TERMS OF SURVIVAL AS OPPOSED TO CONSERVATIVE THERAPY. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS DR PHILIPP LURZ AT GUTENBERG-UNIVERSITY, MAINZ, GERMANY WITH CORRESPONDING EMAIL ALURZPHIL@UNI-MAINZ.DE.]. THE TIME FRAME OF THE STUDY WAS 2010 TO 2017. A TOTAL OF 1885 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH 1300 UNDERWENT T-TEER WHERE AN UNKNOWN AMOUNT RECEIVED AN ABBOTT DEVICE. THE AVERAGE AGE WAS 79 AND MAJORITY GENDER WAS FEMALE. COMORBIDITIES INCLUDED ATRIAL FIBRILLATION, HEART FAILURE, TRICUSPID REGURGITATION, PACEMAKER, CORONARY ARTERY DISEASE. PERI AND POST-PROCEDURAL COMPLICATIONS INCLUDED DEATH. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF TRICLIP WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING ATRIAL FIBRILLATION, HEART FAILURE, TRICUSPID REGURGITATION, PACEMAKER, CORONARY ARTERY DISEASE. COMPLICATIONS REPORTED INCLUDED DEATH; THIS COMPLICATION IS ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. B2 DEATH DATE IS ESTIMATED. B3 EVENT DATE IS ESTIMATED. LITERATURE ATTACHMENT: TRICUSPID REGURGITATION DISEASE STAGES AND TREATMENT OUTCOMES AFTER TRANSCATHETER TRICUSPID VALVE REPAIR
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21550949&pc=NPT | 2025587-2025-01753 | 2025-02-26 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-07 05:00 | NPT | EVOLUT FX DCS | Failure to Advance; Insufficient Information; Patient Device Interaction Problem | Cardiac Arrest | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE IMPLANT PROCEDURE, A PRE-IMPLANT BALLOON DILATION WAS PERFORMED ON THE NATIVE BICUSPID VALVE. FOLLOWING THE VALVE IMPLANT, A BALLOON DILATION WAS PERFORMED. DURING THE POST-IMPLANT DILATION, THERE WAS A NON-CAPTURED BEAT DURING RAPID PACING, WHICH CAUSED THE VALVE TO MIGRATE TO THE AORTA. A SECOND VALVE IMPLANT WAS ATTEMPTED, BUT IT WAS REPORTED THAT THE PATIENT CRASHED AND WAS TAKEN TO THE OPERATING ROOM FOR OPEN HEART SURGERY, WHERE THE PATIENT WAS LATER FOUND DECEASED.Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE IMPLANT PROCEDURE, A PRE-IMPLANT BALLOON DILATION WAS PERFORMED ON THE NATIVE BICUSPID VALVE. FOLLOWING THE VALVE IMPLANT, A BALLOON DILATION WAS PERFORMED. DURING THE POST-IMPLANT DILATION, THERE WAS A NON-CAPTURED BEAT DURING RAPID PACING, WHICH CAUSED THE VALVE TO MIGRATE TO THE AORTA. A SECOND VALVE IMPLANT WAS ATTEMPTED, BUT IT WAS REPORTED THAT THE PATIENT CRASHED AND WAS TAKEN TO THE OPERATING ROOM FOR OPEN HEART SURGERY, WHERE THE PATIENT WAS LATER FOUND DECEASED. ADDITIONAL INFORMATION WAS RECEIVED THAT THE POST-IMPLANT BALLOON DILATION WAS PERFORMED BECAUSE THE VALVE DID NOT FULLY EXPAND. UPON ATTEMPTING TO IMPLANT THE SECOND VALVE, THE DELIVERY CATHETER SYSTEM (DCS) HAD DIFFICULTY CROSSING THE FIRST VALVE, AND SUBSEQUENTLY, AN AORTIC DISSECTION OCCURRED DUE TO THE VALVE MOVING DOWN FROM BEING PUSHED BY THE DCS. THE PATIENT WAS TAKEN TO SURGERY TO TREAT THE DISSECTION, AND THE VALVE WAS EXPLANTED. PER THE PHYSICIAN, THE CAUSE OF THE PATIENT'S ARREST AND SUBSEQUENT DEATH WAS DUE TO THE DISSECTION. Manufacturer Narrative: UPDATED DATA: B2. B5. SECOND PARAGRAPH ADDED H6. ADDITIONAL CODES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED H.6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SECTION D REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: PRODUCT ID L-EVOLUTFX-34 (LOT: 0012318213); PRODUCT TYPE: 0195-HEART VALVES; PRODUCT ID EVFXPLUS-34 (SERIAL: (B)(6); PRODUCT TYPE: 0195-HEART VALVES; PRODUCT ID EVFXPLUS-34 ((B)(6); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE (B)(6) 2025. SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21551031&pc=OUT | 2029214-2025-00594 | 2021-07-15 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-03-07 05:00 | OUT | PIPELINE | Adverse Event Without Identified Device or Use Problem | Ruptured Aneurysm |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING: THE STUDY AIMS TO EVALUATE CLINICAL AND RADIOLOGICAL OUTCOMES AS WELL AS TREATMENT-RELATED COMPLICATIONS OF PATIENTS TREATED WITH FLOW DIVERTERS (FDS) FOR ACUTELY RUPTURED INTRACRANIAL ANEURYSMS (IAS). THE TIME FRAME OF THIS STUDY WAS: JANUARY 2012 TO DECEMBER 2019. MULTIPLE MANUFACTURER¿S DEVICES WERE USED IN THE STUDY POPULATION. THE FOLLOWING MEDTRONIC DEVICES WERE USED: PIPELINE EMBOLIZATION DEVICE (PED), PED WITH SHIELD TECHNOLOGY (62 DEVICES), AND OTHER UNSPECIFIED MEDTRONIC DEVICES. DEATHS OCCURRED IN THE STUDY POPULATION. THE OVERALL MORTALITY RATE WAS 18% (18 OUT OF 102 PATIENTS), WITH A TREATMENT-RELATED MORTALITY RATE OF 5% (5 OUT OF 102 PATIENTS). ONE PATIENT DIED DUE TO REBLEEDING. THE CAUSES OF DEATH WERE: REBLEEDING (1 DEATH) AMONG PATIENT ADVERSE EVENTS INCLUDED: TREATMENT-RELATED COMPLICATIONS IN 45% OF PATIENTS (49 PATIENTS). POSTOPERATIVE INTRACRANIAL HEMORRHAGE (ICH) IN 17% OF PATIENTS (19 PATIENTS). NEW ICH OR EXPANSION OF EXISTING ICH IN 17% OF PATIENTS (19 PATIENTS). REBLEEDING RATE OF 3% (3 PATIENTS). OTHER HEMORRHAGIC COMPLICATIONS: GROIN HEMORRHAGE (2%), SERIOUS EPISTAXIS (1%), UPPER GASTROINTESTINAL BLEEDING (1%). ISCHEMIC COMPLICATIONS IN 29% OF PATIENTS (32 PATIENTS). STENT THROMBOSIS IN 8% OF PATIENTS (9 PATIENTS). POSTOPERATIVE INTRAVENTRICULAR HEMORRHAGE (IVH) IN 23% OF PATIENTS (25 PATIENTS). UNFAVORABLE FUNCTIONAL OUTCOME (MRS 3¿6) IN 27% OF PATIENTS (28 OUT OF 102). NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS.Event Description: 2025-FEB-28 N/A (REP, HCP, FOR, LIT): LITERATURE WAS REVIEWED REGARDING: THE STUDY AIMS TO EVALUATE CLINICAL AND RADIOLOGICAL OUTCOMES AS WELL AS TREATMENT-RELATED COMPLICATIONS OF PATIENTS TREATED WITH FLOW DIVERTERS (FDS) FOR ACUTELY RUPTURED INTRACRANIAL ANEURYSMS (IAS). THE TIME FRAME OF THIS STUDY WAS: JANUARY 2012 TO DECEMBER 2019. MULTIPLE MANUFACTURER¿S DEVICES WERE USED IN THE STUDY POPULATION. THE FOLLOWING MEDTRONIC DEVICES WERE USED: PIPELINE EMBOLIZATION DEVICE (PED), PED WITH SHIELD TECHNOLOGY (62 DEVICES), AND OTHER UNSPECIFIED MEDTRONIC DEVICES. DEATHS OCCURRED IN THE STUDY POPULATION. THE OVERALL MORTALITY RATE WAS 18% (18 OUT OF 102 PATIENTS), WITH A TREATMENT-RELATED MORTALITY RATE OF 5% (5 OUT OF 102 PATIENTS). ONE PATIENT DIED DUE TO REBLEEDING. THE CAUSES OF DEATH WERE: REBLEEDING (1 DEATH) AMONG PATIENT ADVERSE EVENTS INCLUDED: TREATMENT-RELATED COMPLICATIONS IN 45% OF PATIENTS (49 PATIENTS). POSTOPERATIVE INTRACRANIAL HEMORRHAGE (ICH) IN 17% OF PATIENTS (19 PATIENTS). NEW ICH OR EXPANSION OF EXISTING ICH IN 17% OF PATIENTS (19 PATIENTS). REBLEEDING RATE OF 3% (3 PATIENTS). OTHER HEMORRHAGIC COMPLICATIONS: GROIN HEMORRHAGE (2%), SERIOUS EPISTAXIS (1%), UPPER GASTROINTESTINAL BLEEDING (1%). ISCHEMIC COMPLICATIONS IN 29% OF PATIENTS (32 PATIENTS). STENT THROMBOSIS IN 8% OF PATIENTS (9 PATIENTS). POSTOPERATIVE INTRAVENTRICULAR HEMORRHAGE (IVH) IN 23% OF PATIENTS (25 PATIENTS). UNFAVORABLE FUNCTIONAL OUTCOME (MRS 3¿6) IN 27% OF PATIENTS (28 OUT OF 102). NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID NV UNK PIPELINE (LOT: UNKNOWN); PRODUCT TYPE: ; IMPLANT DATE N/A; EXPLANT DATE N/A. PRODUCT ID NV UNK PIPELINE (UNKNOWN); PRODUCT TYPE: ; IMPLANT DATE N/A; EXPLANT DATE N/A. G2: CITATION: AUTHORS: KEMAL ALPAY, TERO HINKKA, ANTTI E LINDGREN, JUHA-MATTI ISOKANGAS, RAHUL RAJ, RIITTA PARKKOLA, MATIAS SINISALO, JUSSI NUMMINEN, JUHA-PEKKA PIENIMÄKI, PETRI SAARI, JANNE SEPPÄNEN, K. FINNISH FLOW DIVERTER STUDY: 8 YEARS OF EXPERIENCE IN THE TREATMENT OF ACUTELY RUPTURED INTRACRANIAL ANEURYSMS. J NEURO INTERVENT SURG 14:699¿703 2022. DOI: 10.1136/NEURINTSURG-2021-017641 EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21551160&pc=NPT | 2025587-2025-01754 | 2025-02-26 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-07 05:00 | NPT | EVOLUT FX PLUS VALVE | Device Dislodged or Dislocated; Insufficient Information; Activation Failure; Patient Device Interaction Problem; Migration | Cardiac Arrest | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE IMPLANT PROCEDURE, A PRE-IMPLANT BALLOON DILATION WAS PERFORMED ON THE NATIVE BICUSPID VALVE. FOLLOWING THE VALVE IMPLANT, A BALLOON DILATION WAS PERFORMED. DURING THE POST-IMPLANT DILATION, THERE WAS A NON-CAPTURED BEAT DURING RAPID PACING, WHICH CAUSED THE VALVE TO MIGRATE TO THE AORTA. A SECOND VALVE IMPLANT WAS ATTEMPTED, BUT IT WAS REPORTED THAT THE PATIENT CRASHED AND WAS TAKEN TO THE OPERATING ROOM FOR OPEN HEART SURGERY, WHERE THE PATIENT WAS LATER FOUND DECEASED.Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE IMPLANT PROCEDURE, A PRE-IMPLANT BALLOON DILATION WAS PERFORMED ON THE NATIVE BICUSPID VALVE. FOLLOWING THE VALVE IMPLANT, A BALLOON DILATION WAS PERFORMED. DURING THE POST-IMPLANT DILATION, THERE WAS A NON-CAPTURED BEAT DURING RAPID PACING, WHICH CAUSED THE VALVE TO MIGRATE TO THE AORTA. A SECOND VALVE IMPLANT WAS ATTEMPTED, BUT IT WAS REPORTED THAT THE PATIENT CRASHED AND WAS TAKEN TO THE OPERATING ROOM FOR OPEN HEART SURGERY, WHERE THE PATIENT WAS LATER FOUND DECEASED. ADDITIONAL INFORMATION WAS RECEIVED THAT THE POST-IMPLANT BALLOON DILATION WAS PERFORMED BECAUSE THE VALVE DID NOT FULLY EXPAND. UPON ATTEMPTING TO IMPLANT THE SECOND VALVE, THE DELIVERY CATHETER SYSTEM (DCS) HAD DIFFICULTY CROSSING THE FIRST VALVE, AND SUBSEQUENTLY, AN AORTIC DISSECTION OCCURRED DUE TO THE VALVE MOVING DOWN FROM BEING PUSHED BY THE DCS. THE PATIENT WAS TAKEN TO SURGERY TO TREAT THE DISSECTION, AND THE VALVE WAS EXPLANTED. PER THE PHYSICIAN, THE CAUSE OF THE PATIENT'S ARREST AND SUBSEQUENT DEATH WAS DUE TO THE DISSECTION. Manufacturer Narrative: UPDATED H.6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SECTION D REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: PRODUCT ID L-EVOLUTFX-34 (LOT: 0012318213); PRODUCT TYPE: 0195-HEART VALVES; PRODUCT ID EVFXPLUS-34 (SERIAL: (B)(6)); PRODUCT TYPE: 0195-HEART VALVES; PRODUCT ID D-EVOLUTFX-34 (0012399256); PRODUCT TYPE: 0195-HEART VALVES. SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B2. B5. SECOND PARAGRAPH ADDED D6B. H6. ADDITIONAL CODES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21551442&pc=CBK | 2518422-2025-029621 | 2023-03-08 05:00 | Death | RESPIRONICS, INC. | 2025-03-07 05:00 | CBK | TRILOGY EVO | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K181166 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A TRILOGY EVO DEVICE IS RETURNING BECAUSE THE PATIENT HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. LIMITED INFORMATION WAS PROVIDED BY THE CUSTOMER; HOWEVER THE CUSTOMER DID INDICATE THAT THE PATIENT EXPIRED DUE TO A NON-SPECIFIC NATURAL DISEASE PROGRESSION. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR EVALUATION AND WAS FOUND TO PASS TESTING AND TO BE IN OVERALL GOOD CONDITION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21551629&pc=OUT | 2029214-2025-00595 | 2021-07-15 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-03-07 05:00 | OUT | PIPELINE | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Intracranial Hemorrhage; Thrombosis/Thrombus; Gastrointestinal Hemorrhage |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING: THE STUDY AIMS TO EVALUATE CLINICAL AND RADIOLOGICAL OUTCOMES AS WELL AS TREATMENT-RELATED COMPLICATIONS OF PATIENTS TREATED WITH FLOW DIVERTERS (FDS) FOR ACUTELY RUPTURED INTRACRANIAL ANEURYSMS (IAS). THE TIME FRAME OF THIS STUDY WAS: JANUARY 2012 TO DECEMBER 2019. MULTIPLE MANUFACTURER¿S DEVICES WERE USED IN THE STUDY POPULATION. THE FOLLOWING MEDTRONIC DEVICES WERE USED: PIPELINE EMBOLIZATION DEVICE (PED), PED WITH SHIELD TECHNOLOGY (62 DEVICES), AND OTHER UNSPECIFIED MEDTRONIC DEVICES. DEATHS OCCURRED IN THE STUDY POPULATION. THE OVERALL MORTALITY RATE WAS 18% (18 OUT OF 102 PATIENTS), WITH A TREATMENT-RELATED MORTALITY RATE OF 5% (5 OUT OF 102 PATIENTS). ONE PATIENT DIED DUE TO REBLEEDING. THE CAUSES OF DEATH WERE: REBLEEDING (1 DEATH) AMONG PATIENT ADVERSE EVENTS INCLUDED: TREATMENT-RELATED COMPLICATIONS IN 45% OF PATIENTS (49 PATIENTS). POSTOPERATIVE INTRACRANIAL HEMORRHAGE (ICH) IN 17% OF PATIENTS (19 PATIENTS). NEW ICH OR EXPANSION OF EXISTING ICH IN 17% OF PATIENTS (19 PATIENTS). REBLEEDING RATE OF 3% (3 PATIENTS). OTHER HEMORRHAGIC COMPLICATIONS: GROIN HEMORRHAGE (2%), SERIOUS EPISTAXIS (1%), UPPER GASTROINTESTINAL BLEEDING (1%). ISCHEMIC COMPLICATIONS IN 29% OF PATIENTS (32 PATIENTS). STENT THROMBOSIS IN 8% OF PATIENTS (9 PATIENTS). POSTOPERATIVE INTRAVENTRICULAR HEMORRHAGE (IVH) IN 23% OF PATIENTS (25 PATIENTS). UNFAVORABLE FUNCTIONAL OUTCOME (MRS 3¿6) IN 27% OF PATIENTS (28 OUT OF 102). NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID NV UNK PIPELINE (LOT: UNKNOWN); PRODUCT TYPE: ; IMPLANT DATE N/A; EXPLANT DATE N/A. PRODUCT ID NV UNK PIPELINE (UNKNOWN); PRODUCT TYPE: ; IMPLANT DATE N/A; EXPLANT DATE N/A. G2: CITATION: AUTHORS: KEMAL ALPAY, TERO HINKKA, ANTTI E LINDGREN, JUHA-MATTI ISOKANGAS, RAHUL RAJ, RIITTA PARKKOLA, MATIAS SINISALO, JUSSI NUMMINEN, JUHA-PEKKA PIENIMÄKI, PETRI SAARI, JANNE SEPPÄNEN, K. FINNISH FLOW DIVERTER STUDY: 8 YEARS OF EXPERIENCE IN THE TREATMENT OF ACUTELY RUPTURED INTRACRANIAL ANEURYSMS. J NEURO INTERVENT SURG 14:699¿703 2022. DOI: 10.1136/NEURINTSURG-2021-017641 EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21551731&pc=DRE | 2522007-2025-00003 | 2025-02-27 05:00 | Death | COOK VANDERGRIFT INC. | 2025-03-07 05:00 | DRE | LEAD EXTRACTION EVOLUTION RL CONTROLLED-ROTATION DILATOR SHEATH SET | Adverse Event Without Identified Device or Use Problem | Great Vessel Perforation | K141148 |  | 1 | Event Description: AS REPORTED TO CUSTOMER RELATIONS: CUSTOMER REPORTS DURING THE USE OF A 13 FR EVOLUTION RL (G23747), THE EVOLUTION CHEWED THROUGH THE OUTER SHEATH WHILE BEING ACTIVATED INSIDE OF SAID SHEATH. THIS CAUSED THE SYSTEM TO NOT BE COAXIAL WITH THE LEAD AND CAUSED AN SUPERIOR VENA CAVA (SVC) TEAR WHICH RESULTED IN A PATIENT DEATH.Event Description: AS REPORTED TO CUSTOMER RELATIONS: CUSTOMER REPORTS DURING THE USE OF A 13 FR EVOLUTION RL (G23747), THE EVOLUTION CHEWED THROUGH THE OUTER SHEATH WHILE BEING ACTIVATED INSIDE OF SAID SHEATH. THIS CAUSED THE SYSTEM TO NOT BE COAXIAL WITH THE LEAD AND CAUSED AN SUPERIOR VENA CAVA (SVC) TEAR WHICH RESULTED IN A PATIENT DEATH.Event Description: AS REPORTED TO CUSTOMER RELATIONS: CUSTOMER REPORTS DURING THE USE OF A LEAD EXTRACTION EVOLUTION RL CONTROLLED-RL CONTROLLED-ROTATION DILATOR SHEATH SET (G23747), THE EVOLUTION CHEWED THROUGH THE OUTER SHEATH WHILE BEING ACTIVATED INSIDE OF SAID SHEATH. THIS CAUSED THE SYSTEM TO NOT BE COAXIAL WITH THE LEAD AND CAUSED A SUPERIOR VENA CAVA (SVC) TEAR WHICH RESULTED IN A PATIENT DEATH. Manufacturer Narrative: BLANK FIELDS ON THIS FORM INDICATED THE INFORMATION IS UNKNOWN, UNCHANGED, OR UNAVAILABLE. THE UDI INFORMATION PER 21CFR 803.52(C)(4) AND 21 CFR 803.50(B) THAT IS REQUIRED IS REASONABLY UNKNOWN DUE TO THE DEVICE ITSELF WAS DISCARDED/NOT RETURNED FOR INVESTIGATION AND THE LOT OF THE DEVICE WAS UNKNOWN. D4 - LOT #: UNKNOWN. E3 - OCCUPATION: UNKNOWN. THE EVENT IS CURRENTLY UNDER INVESTIGATION. THIS REPORT INCLUDES INFORMATION KNOWN AT THIS TIME. A FOLLOW UP REPORT WILL BE SUBMITTED SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE OR UPON COMPLETION OF INVESTIGATION. THIS REPORT IS REQUIRED BY THE FDA UNDER 21 CFR PART 803. THIS REPORT IS BASED ON UNCONFIRMED INFORMATION SUBMITTED BY OTHERS. NEITHER THE SUBMISSION OF THIS REPORT NOR ANY STATEMENT MADE IN IT IS INTENDED TO BE AN ADMISSION THAT ANY COOK DEVICE IS DEFECTIVE OR MALFUNCTIONED; THAT A DEATH OR SERIOUS INJURY OCCURRED; OR THAT ANY COOK DEVICE CAUSED OR CONTRIBUTED TO; OR IS LIKELY TO CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY IF A MALFUNCTION OCCURRED. Manufacturer Narrative: BLANK FIELDS ON THIS FORM INDICATED THE INFORMATION IS UNKNOWN, UNCHANGED, OR UNAVAILABLE. THE UDI INFORMATION PER 21CFR 803.52(C)(4) AND 21 CFR 803.50(B) THAT IS REQUIRED IS REASONABLY UNKNOWN DUE TO THE DEVICE ITSELF WAS DISCARDED/NOT RETURNED FOR INVESTIGATION AFTER MANUFACTURER REQUESTED FOR THE RETURN AND THE LOT OF THE DEVICE WAS UNKNOWN. . THE DEVICE WAS NOT RETURNED FOR THE COMPLAINT; THEREFORE, A PHYSICAL INVESTIGATION COULD NOT BE PERFORMED, AND THE CUSTOMER'S COMPLAINT IS ACKNOWLEDGED, BUT COULD NOT BE CONFIRMED, OTHER THAN BY THE CUSTOMER'S TESTIMONY. THE COMPLAINT/EVENT THAT WAS ENTERED AND REPORTED WITHIN TRACKWISE: "THE EVOLUTION CHEWED THROUGH THE OUTER SHEATH WHILE BEING ACTIVATED INSIDE OF SAID SHEATH. THIS CAUSED THE SYSTEM TO NOT BE COAXIAL WITH THE LEAD AND CAUSED AN SVC TEAR WHICH RESULTED IN A DEATH." THE DEVICE HISTORY RECORD (DHR) WAS UNABLE TO BE VIEWED DUE TO THE LOT NUMBER SPECIFIC TO THIS COMPLAINT WAS UNKNOWN. THIS COMPLAINT MODE IS BEING MONITORED, TRACKED AND TRENDED PER THE CVI POST MARKET SURVEILLANCE AND COMPLAINT HANDLING PROCESSES. A RISK ASSESSMENT WILL BE PERFORMED AND DOCUMENTED IN TRACKWISE. THIS REPORT IS REQUIRED BY THE FDA UNDER 21 CFR PART 803. THIS REPORT IS BASED ON UNCONFIRMED INFORMATION SUBMITTED BY OTHERS. NEITHER THE SUBMISSION OF THIS REPORT NOR ANY STATEMENT MADE IN IT IS INTENDED TO BE AN ADMISSION THAT ANY COOK DEVICE IS DEFECTIVE OR MALFUNCTIONED; THAT A DEATH OR SERIOUS INJURY OCCURRED; OR THAT ANY COOK DEVICE CAUSED OR CONTRIBUTED TO; OR IS LIKELY TO CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY IF A MALFUNCTION OCCURRED. Manufacturer Narrative: THE UDI INFORMATION PER 21CFR 803.52(C)(4) AND 21 CFR 803.50(B) THAT IS REQUIRED IS REASONABLY UNKNOWN DUE TO THE DEVICE ITSELF WAS DISCARDED/NOT RETURNED FOR INVESTIGATION AFTER MANUFACTURER REQUESTED FOR THE RETURN AND THE LOT OF THE DEVICE WAS UNKNOWN. THE EVENT IS CURRENTLY UNDER INVESTIGATION AND AWAITING AN AD-HOC RISK REVIEW. THIS REPORT INCLUDES INFORMATION KNOWN AT THIS TIME. AN ADDITIONAL FOLLOW UP REPORT WILL BE SUBMITTED SHOULD ADDITIONAL RELEVANT INFORMATION WERE TO COME FROM THE AD-HOC MEETING AND/OR HAS BECOME AVAILABLE UPON COMPLETION OF INVESTIGATION. THIS REPORT IS REQUIRED BY THE FDA UNDER 21 CFR PART 803. THIS REPORT IS BASED ON UNCONFIRMED INFORMATION SUBMITTED BY OTHERS. NEITHER THE SUBMISSION OF THIS REPORT NOR ANY STATEMENT MADE IN IT IS INTENDED TO BE AN ADMISSION THAT ANY COOK DEVICE IS DEFECTIVE OR MALFUNCTIONED; THAT A DEATH OR SERIOUS INJURY OCCURRED; OR THAT ANY COOK DEVICE CAUSED OR CONTRIBUTED TO; OR IS LIKELY TO CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY IF A MALFUNCTION OCCURRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21551811&pc=MVK | 3008642652-2025-02429 | 2025-01-25 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-03-07 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Failure to Convert Rhythm | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED FIVE APPROPRIATE TREATMENTS AND A NON-LIFEVEST DEFIBRILLATION. THE DEVICE WAS STARTED UP AT 20:32:45 ON (B)(6) 2025. AT 02:35:40 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS SINUS RHYTHM @ 80 BPM WITH PVCS TRANSITIONING TO VT @ 260 BPM. AT 02:36:17, THE PATIENT RECEIVED THE FIRST APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 270 BPM. POST SHOCK RHYTHM WAS SINUS BRADYCARDIA @ 40 BPM. AT 05:47:05, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VT @ 280 BPM. AT 05:47:35, THE PATIENT RECEIVED THE SECOND APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS SINUS BRADYCARDIA @ 40 BPM. AT 06:12:53, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VT @ 280 BPM. AT 06:13:23, THE PATIENT RECEIVED THE THIRD APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 270 BPM. POST SHOCK RHYTHM WAS SINUS BRADYCARDIA @ 40 BPM. AT 10:24:01, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VT @ 260 BPM. AT 10:24:31, THE PATIENT RECEIVED THE FOURTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 290 BPM WITH FUSION BEAT. POST SHOCK RHYTHM WAS SINUS BRADYCARDIA @ 50 BPM. AT 10:32:57, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VT @ 290 BPM. AT 10:33:14, THE PATIENT RECEIVED THE NON-LIFEVEST DEFIBRILLATION. THE RHYTHM AT THE TIME OF THE NON-LIFEVEST DEFIBRILLATIONS WAS VT @ 290 BPM. POST SHOCK RHYTHM WAS SINUS TACHYCARDIA @ 120 BPM WITH PVCS TRANSITIONING BACK TO VT @ 290 BPM. THE PATIENT RECEIVED A NON-LIFEVEST DEFIBRILLATION 17 SECONDS INTO THE DETECTION SEQUENCE. AT 10:33:32, THE PATIENT RECEIVED THE FIFTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS VT @ 280 BPM DEGRADING TO VF. THE ELECTRODE BELT WAS DISCONNECTED AT 10:33:43 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR HAS BEEN COMPLETED. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION. THE ELECTRODE BELT HAS NOT BEEN RECOVERED. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21552195&pc=MRN | 3012796317-2025-00002 | 2025-02-07 05:00 | Death | NOXBOX LTD | 2025-03-07 05:00 | MRN | NOXBOXI NITRIC OXIDE DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem; Insufficient Information | Insufficient Information | K231823 |  | 1 | Event Description: SUMMERLIN HOSPITAL REPORTED NITRIC OXIDE (NO) DOSE FLUCTUATIONS ON A NOXBOXI DEVICE DURING INHALED NITRIC OXIDE (INO) THERAPY. AS REPORTED, THE DEVICE WAS USED OFF-LABEL ON A <34-WEEK NEONATE PATIENT AND WAS ISSUING HIGH AND LOW DOSE ALARMS. HOSPITAL STAFF ATTEMPTED TO RESOLVE THE ISSUE BY REPLACING THE DEVICE WITH A DIFFERENT NOXBOXI UNIT; HOWEVER, THE DOSE FLUCTUATIONS PERSISTED. THREE HOURS AFTER THE INITIAL REPORT WAS MADE TO LINDE, THE HOSPITAL NOTIFIED LINDE THAT THE PATIENT HAD EXPIRED. THERE WAS NO REPORTED RELATIONSHIP BETWEEN THE DEVICE AND PATIENT'S OUTCOME. THE PATIENT WAS ON A BUNNELL JET LIFEPULSE VENTILATOR USING HIGH FREQUENCY JET VENTILATION. Manufacturer Narrative: THIS REPORT CONCERNS TWO NOXBOXI DEVICES (SERIAL NUMBERS (B)(6)) AND A REPORTED ADVERSE EVENT ON FEBRUARY 7, 2025. DESPITE A REQUEST FOR DEVICE RETURN, THE HOPITAL DECLINED TO PROVIDE THE DEVICES FOR PHYSICAL EVALUATION. LOG FILE DATA, PROVIDED BY THE HOSPITAL, INDICATES NO RECORDED DEVICE ACTIVITY: (B)(6) FROM (B)(6) 2025, AND (B)(6) FROM (B)(6) 2024 TO (B)(6) 2025. THEREFORE, NEITHER DEVICE WAS POWERED ON FEBRUARY 7, 2025, THE DATE OF THE REPORTED EVENT. A FORMAL REQUEST FOR DEVICE RETURN AND MORE INFORMATION ABOUT THE EVENT HAS BEEN SUBMITTED. HOWEVER, WE HAVE NOT RECEIVED FURTHER INFORMATION REGARDING THE PATIENT'S CONDITION OR CAUSE OF DEATH. CONSEQUENTLY, BASED ON THE AVAILABLE DATA, A CAUSAL RELATIONSHIP BETWEEN DEVICE FUNCTION AND THE REPORTED EVENT CANNOT BE ESTABLISHED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21552354&pc=DSQ | 3007042319-2025-00632 | 2024-01-01 05:00 | Death | HEARTWARE, INC. | 2025-03-07 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P100047 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING THE CARDIAC FUNCTIONAL AND STRUCTURAL CHANGES FOLLOWING MECHANICAL CIRCULATORY SUPPORT. THE AUTHORS DESCRIBED PATIENT DEATHS; HOWEVER, THE CAUSES OF DEATH WERE UNKNOWN. THERE WERE PATIENTS WHO EXPERIENCED RIGHT VENTRICULAR (RV) FAILURE POST VENTRICULAR ASSIST DEVICE (VAD) IMPLANT. THE STATUS OF THE VADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/56 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: CARDIAC REVERSE REMODELING MEDIATED BY HEARTMATE 3 LEFT VENTRICULAR ASSIST DEVICE: COMPARISON TO OLDER GENERATION DEVICES. ASAIO JOURNAL. 2024. 70:1060¿1066. DOI: 10.1097/MAT.0000000000002245. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21552381&pc=DSQ | 2916596-2025-01559 | 2025-02-06 05:00 | Death | THORATEC CORPORATION | 2025-03-07 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Mechanical Problem; Infusion or Flow Problem | Stroke/CVA; Tachycardia; Thromboembolism; Hemorrhagic Stroke; Heart Failure/Congestive Heart Failure; Valvular Insufficiency/ Regurgitation | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT WAS STENTED FOR EXTERNAL OUTFLOW GRAFT OBSTRUCTION (EOGO). THE PATIENT'S FLOW PROCEEDED TO DROP TO MID 1L/MIN AND THEY APPEARED TO BE IN A TACHYCARDIA ARRYTHMIA. THEY RECEIVED DEFIBRILLATOR SHOCK AND THEIR MEAN ARTERIAL PRESSURE IMPROVED BUT THE FLOWS DID NOT IMPROVE. LOG FILES WERE SUBMITTED FOR FURTHER EVALUATION. THE PERIODIC LOG FILE CAPTURED AN ELEVATION IN FLOW AND POWER ON (B)(6)2025 AT 16:44, THE EVENT LOG FILE CAPTURED LOW FLOW EVENTS BEGINNING ON 14FEB2025 AT 8:52:17. AT THE SAME TIME, THE PUMP EVENT FILE SHOWED ELEVATIONS IN ROTOR NOISE AND DISPLACEMENT WHICH RESULTED IN HARMONIC COMPENSATION EVENTS. IT WAS LATER REPORTED THAT THE PATIENT PASSED AWAY DUE TO STROKE POST OUTFLOW GRAFT STENTS. IT WAS SAID THE DEVICE OPERATED AS EXPECTED AFTER THE INITIAL ISSUE. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER.Event Description: IT WAS LATER REPORTED THAT THE DEVICE OPERATED AS INTENDED AFTER THE INITIAL EVENT RESOLVED UNTIL THE PATIENT'S DEATH. THE DEATH WAS CONSIDERED TO BE DEVICE RELATED.Event Description: IT WAS LATER REPORTED THAT ON (B)(6) 2025 A COMPUTED TOMOGRAPHY ANGIOGRAPHY (CTA) OF THE ABDOMEN AND PELVIS WAS PERFORMED WHICH SHOWED PARTIAL THROMBOSIS OF THE OUTFLOW TRACT. THERE WERE ALSO WEDGE-SHAPED HYPODENSITIES IN THE KIDNEY AND SPLEEN SUGGESTIVE OF INFARCTS. THROMBOEMBOLIC PHENOMENON WAS SUSPECTED. FINDINGS FROM A CTA CARDIAC TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) ON (B)(6) 2025 ALSO SHOWED A NEAR OCCLUSIVE THROMBUS IN THE OUTFLOW CANNULA AND WEDGE-SHAPED HYPODENSITIES IN THE POSTERIOR SPLEEN AND LEFT KIDNEY THAT EXTENDED TO THE RENAL CORTEX WHICH WERE SUSPICIOUS FOR INFARCTS. THE ABDOMINAL AORTA, CELIAC TRUNK, SUPERIOR MESENTERIC ARTERY, AND ILIAC BIFURCATION WERE ALL VISUALIZED AND PATENT. ECHOCARDIOGRAM ON (B)(6) 2025 SHOWED SEVERE RIGHT VENTRICULAR ENLARGEMENT WITH MODERATELY REDUCED SYSTOLIC FUNCTION, SEVERE BIATRIAL ENLARGEMENT, AND MODERATE MITRAL REGURGITATION AND MILD TO MODERATE TRICUSPID REGURGITATION. THE PATIENT WAS NOTED TO HAVE AN ALFIERI STICH IN THE MITRAL VALVE WHICH WAS PLACED AT THE TIME OF IMPLANT. THERE WAS ALSO NOTED ELEVATED VENOUS PRESSURE IN THE INFERIOR VENA CAVA AND SUPERIOR VENA CAVA. THE PATIENT EXPERIENCED INCREASED ACUTE SYMPTOMS OF ACUTE DECOMPENSATED HEART FAILURE, DIAPHORESIS, NAUSEA, DIZZINESS, INCREASED LACTATE DEHYDROGENASE TO THE 1200 UNITS PER LITER RANGE, INCREASED LIVER FUNCTION TESTS TO THE 900 UNITS PER LITER RANGE, AND DECREASED SODIUM TO 117 MILLIEQUIVALENTS PER LITER. THERE WERE NO CHANGES TO PATIENT CONDITION OR ANTICOAGULATION STATUS THAT MAY HAVE CONTRIBUTED TO THE EVENT AND THERE WERE NO CHANGES MADE TO PUMP PARAMETERS. THE PATIENT'S STROKE WAS CONFIRMED TO BE HEMORRHAGIC ON CT IMAGING. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: SECTION B2: OUTCOMES ATTRIBUTED TO ADVERSE CORRECTED. MANUFACTURER¿S INVESTIGATION CONCLUSION: THE REPORTED EXTRINSIC OUTFLOW GRAFT OBSTRUCTION (EOGO) WAS CONFIRMED THROUGH ANALYSIS OF THE SUBMITTED COMPUTED TOMOGRAPHY (CT) IMAGES. REVIEW OF THE SUBMITTED LOG FILES CONFIRMED LOW FLOW ALARMS; HOWEVER, A SPECIFIC CAUSE FOR THESE EVENTS COULD NOT CONCLUSIVELY BE DETERMINED. THE CONTROLLER EVENT LOG FILE CONTAINED EVENTS FROM (B)(6) 2025, PER THE TIMESTAMPS. THE LOG FILE CAPTURED LOW FLOW ALARMS STARTING ON (B)(6) 2025 AND PERSISTING FOR THE REMAINDER OF THE LOG FILE. THE LOG FILE ALSO CAPTURED NUMEROUS PULSATILITY INDEX (PI) EVENTS. NO OTHER NOTABLE EVENTS WERE RECORDED. THE PUMP APPEARED TO FUNCTION AS INTENDED THROUGHOUT THE DURATION OF THE LOG FILE. THE CUSTOMER PROVIDED 3 CT IMAGES FOR ANALYSIS. THE FIRST IMAGE WAS A CROSS-SECTION OF THE OUTFLOW GRAFT, SHOWING SIGNIFICANT OCCLUSION. THE REMAINDER OF THE IMAGES SHOWED THE LENGTH OF THE OUTFLOW GRAFT. EOGO WAS PRESENT ALONG THE LENGTH OF THE BEND RELIEF, AND THE OUTFLOW GRAFT WAS SIGNIFICANTLY OBSTRUCTED AT THE NARROWEST POINT. THE PUMP WAS NOT RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(4) WERE REVIEWED AND SHOWED NO DEVIATION FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE ELECTRONIC IFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) IFU, REV. D, AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK, REV. A ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU LISTS STROKE, RIGHT HEART FAILURE, AORTIC INSUFFICIENCY, THROMBOEMBOLISM, ARRHYTHMIA, AND DEATH AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 1 ALSO PROVIDES AN EXPLANATION OF EACH OF THE PUMP PARAMETERS, INCLUDING PUMP FLOW. SECTION 4, ¿SYSTEM MONITOR,¿ EXPLAINS THAT THE LOW FLOW HAZARD ALARM OCCURS WHEN PUMP FLOW IS LESS THAN 2.5 LPM. A 10-SECOND DELAY IS IMPOSED BETWEEN THE DETECTION OF THE LOW FLOW STATUS AND THE ACTIVATION OF THE ASSOCIATED AUDIO AND VISUAL INDICATORS ON THE SYSTEM CONTROLLER. CHANGES IN PATIENT CONDITIONS CAN RESULT IN LOW FLOW. SECTION 5, "SURGICAL PROCEDURES", OUTLINES CONSIDERATIONS FOR PUMP PLACEMENT AND PROVIDES INSTRUCTIONS REGARDING THE PREPARATION, INSTALLATION, AND ORIENTATION OF THE SEALED OUTFLOW GRAFT. SECTION 5, UNDER "ATTACHING THE SEALED OUTFLOW GRAFT TO THE PUMP," INSTRUCTS THE USER TO VERIFY THAT THE GRAFT IS NOT TWISTED OR KINKED BY CHECKING THE POSITION OF THE BLACK LINE ON THE GRAFT ABOVE AND BELOW THE BEND RELIEF AND ENSURING THAT THE LINE IS STRAIGHT. SECTION 5, UNDER "PREIMPLANT PROCEDURES" AND "IMPLANT PROCEDURES," CAUTIONS THE USER: "THE SEALED OUTFLOW GRAFT MUST NOT BE KINKED OR POSITIONED WHERE IT COULD ABRADE AGAINST A PUMP COMPONENT OR BODY STRUCTURE" AND "STRETCH THE SEALED OUTFLOW GRAFT COMPLETELY PRIOR TO MEASURING AND CUTTING THE GRAFT TO THE APPROPRIATE LENGTH." SECTION 7, ¿ALARMS AND TROUBLESHOOTING,¿ EXPLAINS SYSTEM ALARMS AND THE RECOMMENDED ACTIONS ASSOCIATED WITH THEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21552415&pc=DSQ | 2916596-2025-01316 | 2025-01-03 05:00 | Death | THORATEC CORPORATION | 2025-03-07 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Sepsis; Ascites; Hemorrhagic Stroke; Heart Failure/Congestive Heart Failure; Peripheral Edema | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO SEPSIS.Event Description: IT WAS COMMUNICATED ON 27FEB2025 THAT THE DEVICE WAS FUNCTIONING NORMALLY, AND THERE WAS NO OPERATION ASSOCIATED WITH DEVICE MALFUNCTION. THE DEVICE WAS STATED TO HAVE BEEN EXPLANTED POSTHUMOUSLY. IT WAS UNKNOWN IF THE DEVICE WAS RETURNED TO THE MANUFACTURER. NO AUTOPSY WAS PERFORMED. IT WAS STATED THAT THE PATIENT'S IMPROVED CONSCIOUSNESS BECAME SEVERE DUE TO CEREBRAL HEMORRHAGE AFTER VENTRICULAR ASSIST DEVICE (VAD) IMPLANTATION. ALTHOUGH THERE WAS NO DIRECT CAUSAL RELATIONSHIP, ANTICOAGULANT WITH VAD WAS OPTIMAL, BUT THE EFFECT COULD NOT BE RULED OUT. ALTHOUGH RIGHT HEART FAILURE COULD NOT BE CONTROLLED AND THE PATIENT DIED, THE REASON FOR THE LACK OF ACTIVE INTERVENTIONS FOR RIGHT HEART FAILURE SUCH AS RIGHT VENTRICULAR ASSIST DEVICE (RVAD) WAS THE FAMILY'S WISH AFTER CEREBRAL HEMORRHAGE.Event Description: IT WAS FURTHER COMMUNICATED ON 01APR2025 THAT THE PATIENT HAD RIGHT HEART FAILURE (RHF) ON (B)(6) 2025. THEIR CENTRAL PULSE PRESSURE WAS GREATER THAN 18 MM HG, AND THE PATIENT HAD ASCITES AND PERIPHERAL EDEMA. IT WAS STATED THAT THE RHF WAS NOT DEVICE RELATED. DUE TO THE CEREBRAL HEMORRHAGE AFTER VAD IMPLANTATION, IT WAS DIFFICULT TO IMPROVE THE PATIENT'S LEVEL OF CONSCIOUSNESS. THE DEVICE WAS FUNCTIONING PROPERLY AT THE TIME OF DEATH. THE PRIMARY CAUSE OF DEATH WAS STATED TO BE RHF. Manufacturer Narrative: SECTION A4 CORRECTION, SECTION E3 CORRECTION, SECTION G2 CORRECTION. MANUFACTURER'S INVESTIGATION CONCLUSION: EVALUATION OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), DID NOT REVEAL ANY DEVICE-RELATED ISSUES. THE HEARTMATE 3 LVAS WAS RETURNED ASSEMBLED WITH THE PUMP CABLE SEVERED APPROXIMATELY 2.5¿ FROM THE PUMP HEADER. THE REMAINDER OF THE PUMP CABLE WAS NOT RETURNED. APPROXIMATELY 1.5¿ OF THE SEALED OUTFLOW GRAFT WAS RETURNED ATTACHED TO THE PUMP COVER OUTLET PORT WITH THE OUTFLOW GRAFT BEND RELIEF HARDWARE ENGAGED. THE APICAL CUFF WAS RETURNED SECURED WITH THE CUFF LOCK ENGAGED. UPON DISASSEMBLY OF THE RETURNED PUMP, EXAMINATION OF THE BLOOD-CONTACTING SURFACES REVEALED NO DEVELOPED DEPOSITIONS OR THROMBUS FORMATIONS THAT WOULD HAVE CONTRIBUTED TO A FLOW OR FUNCTIONAL ISSUE. VISUAL INSPECTION OF THE PUMP ROTOR AND ROTOR WELL DID NOT REVEAL ANY OBVIOUS SURFACE SCRATCHES OR DEFECTS. REVIEW OF THE CONTROLLER AND LEFT VENTRICULAR ASSIST DEVICE (LVAD) EVENT AND PERIODIC LOG FILES RETRIEVED FROM THE RETURNED DEVICES APPEARED TO CAPTURE THE DEVICE FUNCTIONING AS INTENDED. (B)(6) WAS CLEANED, REBUILT, AND FUNCTIONALLY TESTED ON A MOCK CIRCULATORY LOOP. THE DEVICE OPERATED AS INTENDED AND IN ACCORDANCE WITH MANUFACTURING SPECIFICATIONS. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C IS CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ OF THIS DOCUMENT LISTS SEPSIS, LOCALIZED INFECTION, RIGHT HEART FAILURE, STROKE, BLEEDING, AND DEATH AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. ADDITIONALLY, SECTION 6 ¿PATIENT CARE AND MANAGEMENT¿ LISTS INFECTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS. THIS SECTION ALSO INCLUDES INFORMATION REGARDING HOW TO PREVENT AND CONTROL INFECTION. SECTION 6 ALSO LISTS RIGHT HEART FAILURE AS A POTENTIAL RISK/ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS. THIS SECTION (UNDER ¿RIGHT HEART FAILURE¿) ALSO OUTLINES INDICATIONS OF RIGHT HEART FAILURE AS WELL AS POSSIBLE TREATMENTS. ADDITIONALLY, THIS SECTION PROVIDES INFORMATION REGARDING ANTICOAGULATION, INCLUDING RECOMMENDED INR (INTERNATIONAL NORMALIZED RATIO) RANGE, AS WELL AS CONSIDERATIONS FOR WHEN THERE IS A RISK OF BLEEDING. THE HEARTMATE 3 LVAS PATIENT HANDBOOK REV. D, CONTAINS SEVERAL SECTIONS WITH INFORMATION ABOUT HOW TO PREVENT AND CONTROL INFECTION. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: D9 UPDATED. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: SECTION B5 CORRECTION. SECTION H6 CODING CORRECTION. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21552725&pc=CBK | 3016723884-2025-00005 | 2025-02-25 05:00 | Death | HAMILTON MEDICAL AG | 2025-03-07 05:00 | CBK | HAMILTON G5 CUSTOMIZED | Insufficient Information | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  |  | 1 | Event Description: HAMILTON G5 SN (B)(6) WAS ON A PATIENT WHO EXPIRED, UNKNOWN OF ANY PROBLEMS OR ISSUES THAT LED TO THIS. QUESTIONER SENT. TECH DID INSPECT MACHINE AFTER INCIDENT AND MENTIONED THAT EXP. VALVE MAY BE STICKY. SINCE PATIENT INCIDENT THEY DID WANT FIELD SERVICE TO CERTIFY THE MACHINE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21552766&pc=NPT | 2015691-2025-01600 | 2025-02-10 05:00 | Death | EDWARDS LIFESCIENCES | 2025-03-07 05:00 | NPT | EDWARDS LIFESCIENCES COMMANDER DELIVERY SYSTEM | Difficult or Delayed Positioning | Great Vessel Perforation | P140031 |  | 1 | Event Description: AS REPORTED BY THE EDWARDS' AFFILIATE IN JAPAN, APPROXIMATELY 2 YEARS AND 8 MONTHS AFTER A TRANSFEMORAL (TF) TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) WITH A 26 MM SAPIEN 3 VALVE, AN INCREASE OF AORTIC REGURGITATION DUE TO REDUCED LEAFLET MOTION WAS OBSERVED. IT WAS NOTED THE VALVE WAS DEPLOYED WITH 3.5 ML LESS THAN NOMINAL VOLUME AND THAT THE VALVE WAS DEPLOYED TOO LOW, WHICH MAY BE A CAUSE OF THE REDUCED LEAFLET MOTION. A VALVE IN VALVE WAS PERFORMED WITH A 20 MM SAPIEN 3 ULTRA RESILIA (S3UR) VALVE. DURING THE PROCEDURE, WHEN THE DOCTOR PUSHED A COMMANDER DELIVERY SYSTEM (DS) STRONGLY SINCE THE DS COULD NOT CROSS THE AORTIC VALVE, BLOOD PRESSURE (BP) DECREASED. A PERCUTANEOUS CARDIOPULMONARY SUPPORT (PCPS) WAS TRIED TO BE INTRODUCED, BUT IT WAS DIFFICULT TO RUN THE PCPS SINCE THE VENOUS LINE KINKED. AFTER THAT, IT WAS REVEALED THAT THE CAUSE OF DECREASED BP WAS A RUPTURE OF AORTIC ARCH. EMERGENCY MEDICAL CARE WAS PERFORMED, BUT THE PATIENT EXPIRED. Manufacturer Narrative: THIS IS ONE OF TWO EVENTS BEING REPORTED FOR THIS CASE. INVESTIGATION IS ONGOING. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED DUE TO ENGINEERING EVALUATION FINDINGS. SECTIONS G6, H2, H6 (TYPE OF INVESTIGATION, INVESTIGATION FINDINGS, INVESTIGATION CONCLUSIONS), AND CONCLUSIONS IN THIS SECTION HAVE BEEN UPDATED. SECTION H10 WAS UPDATED WITH REFERENCE TO THE OTHER REPORT SUBMITTED FOR THIS CASE. THE EVENTS REPORTED ARE ANTICIPATED IN THE RISK MANAGEMENT DOCUMENTATION FOR TRANSCATHETER HEART VALVE PROCEDURES. A PREVIOUS INVESTIGATION INTO THIS TYPE OF EVENT IS CAPTURED IN AN EDWARDS LIFESCIENCES TECHNICAL SUMMARY AND APPLIES TO THIS COMPLAINT. ADDITIONAL ASSESSMENT OF THE FAILURE MODE IS NOT REQUIRED AT THIS TIME. THE DEVICE WAS NOT RETURNED FOR EVALUATION. DUE TO THE UNAVAILABILITY OF THE COMPLAINT DEVICE, ENGINEERING WAS UNABLE TO PERFORM ANY VISUAL INSPECTION, FUNCTIONAL TESTING, OR DIMENSIONAL ANALYSIS. PER THE TECHNICAL SUMMARY, THE INSTRUCTIONS FOR USE (IFU), CURRENT RISK MITIGATIONS INCLUDING DESIGN AND MANUFACTURING CONTROLS, AND TRAINING MANUALS HAVE BEEN REVIEWED, AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. THROUGH EXTENSIVE COMPLAINT INVESTIGATIONS, DIFFICULTY OR INABILITY IN TRACKING THE DELIVERY SYSTEM TO THE LANDING ZONE, INCLUDING TRAVERSING THE AORTIC ARCH AND CROSSING THE NATIVE ANNULUS/BIOPROSTHESIS HAS NOT BEEN ASSOCIATED WITH DEVICE MALFUNCTIONS OR MANUFACTURING NONCONFORMANCES. RATHER, THE ROOT CAUSE FOR THESE EVENTS HAVE HISTORICALLY BEEN DUE TO PATIENT FACTORS, POOR GUIDEWIRE MANAGEMENT, DAMAGED GUIDEWIRE, AND/OR EXCESSIVE MANIPULATION OF THE FLEX WHEEL. THE COMPLAINT FOR DIFFICULTY CROSSING THE AORTIC VALVE WAS UNABLE TO BE CONFIRMED AS THE DEVICE WAS NOT RETURNED AND NO APPLICABLE IMAGERY WAS PROVIDED FOR EVALUATION. IN ADDITION, THE PATIENT ANATOMY INFORMATION SUCH AS DEGREE OF CALCIFICATION, TORTUOSITY, MINIMUM LUMINAL DIAMETER WAS NOT PROVIDED. DUE TO A LACK OF INFORMATION, A DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED, ALTHOUGH IT IS POSSIBLY RELATED TO THE ROOT CAUSES MENTIONED ABOVE. ADDITIONALLY, IT WAS REPORTED THAT THE BLOOD PRESSURE (BP) DROPPED AFTER THE DOCTOR APPLIED SIGNIFICANT FORCE WITH A COMMANDER DS TO OVERCOME CROSSING DIFFICULTIES. IT WAS LATER REVEALED THAT THE DROP IN BP WAS DUE TO A RUPTURE OF AORTIC ARCH. IN THIS CASE, THE HIGH PUSH FORCE LIKELY CAUSED THE DEVICE TO INTERACT AND PERFORATE THE AORTIC ARCH, RESULTING IN INJURY. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. SINCE NO EDWARDS PRODUCT DEFECTS OR LABELING DEFICIENCIES WERE IDENTIFIED, NO CORRECTIVE OR PREVENTATIVE ACTION IS REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21552804&pc=OZD | 1220648-2025-26583 | 2025-02-17 05:00 | Death | ABIOMED, INC. | 2025-03-07 05:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Cerebral Edema | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A 64 YEAR OLD FEMALE PATIENT ON AN IMPELLA 5.5 PUMP FOR ESCALATION OF THERAPY AND THE REQUIREMENT FOR LONG TERM THERAPY. THE PATIENT'S HEART FUNCTION IMPROVED SOME, HOWEVER, THE PATIENT DIED OF CEREBRAL EDEMA. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK STATES THE FOLLOWING: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿ THIS COMPLAINT IS FOR ESCALATION TO THE IMPELLA 5.5 DEVICE FROM IMPELLA CP DEVICE USED IN THE 1220648-2025-26536. Manufacturer Narrative: THE INVESTIGATION OF STROKE/CVA HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. B2 SELECTION REQUIRING INTERVENTION WAS ERRONEOUSLY SELECTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26583 B3 DATE OF EVENT WAS ERRONEOUSLY ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26583. E4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26583 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. G1 REPORTING CONTACT FACILITY NAME AND FAX NUMBER WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26583. G2 REPORT SOURCE; FOREIGN WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26583.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21554263&pc=DQX | 3004718255-2025-00375 | 2025-03-05 05:00 | Death | ASAHI INTECC | 2025-03-07 05:00 | DQX | PROWATERFLEX | Material Fragmentation | Foreign Body In Patient |  |  | 1 | Event Description: DETAILS TO FOLLOW AS OUR SALES REP WILL BE ONSITE ON (B)(6) 2025. AT THIS MOMENT, ALL HE KNOWS IS THAT THE TIP OF THE PROWATER BROKE OFF IN THE LEFT MAIN AND PATIENT ENDED UP DYING. HE WILL GATHER MORE INFORMATION ABOUT LESION AND OTHER DEVICES USED WHEN HE PICKS UP THE WIRE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21534537&pc=DYB | 9612164-2025-01187 | 2025-02-10 05:00 | Death | MEDTRONIC IRELAND | 2025-03-06 05:00 | DYB | MICRA | Leak/Splash; Use of Device Problem | Air Embolism; Cardiac Arrest | K132030 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE IMPLANT PROCEDURE OF LEADLESS IMPLANTABLE PULSE GENERATOR (IPG), THE INTRODUCER SHEATH WAS ADVANCED TO THE MID RIGHT ATRIUM AND APPARENT AIR EMBOLISM WAS OBSERVED DURING SHEATH FLUSHING. CARDIAC ARREST WAS REPORTED. THE PATIENT DECEASED AFTER CARDIOPULMONARY RESUSCITATION AND A CODE.Event Description: IT WAS FURTHER REPORTED THAT THE AIR EMBOLISM RELATED TO THE SHEATH INSERTION WAS A CONTRIBUTING FACTOR TO THE DEATH. ACCORDING TO THE PHYSICIAN NO BLOOD RETURN WAS NOTED WHEN ASPIRATION WAS ATTEMPTED. A FLUOROSCOPY X-RAY IMAGE THEN DEMONSTRATED APPARENT AIR IN THE RIGHT ATRIUM AND RIGHT VENTRICLE. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21534648&pc=MVK | 3008642652-2025-02368 | 2025-02-03 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-03-06 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED THREE INAPPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 16:36:12 ON (B)(6) 2024. AT 19:36:38, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH CPR/MOTION ARTIFACT. AT 19:37:43, THE PATIENT RECEIVED THE FIRST INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON LIFE-SUSTAINING RHYTHM, WITH CPR/MOTION ARTIFACT. AT 19:38:10, THE PATIENT RECEIVED THE SECOND INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON LIFE-SUSTAINING RHYTHM, WITH CPR/MOTION ARTIFACT. AT 19:38:38, THE PATIENT RECEIVED THE THIRD INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON LIFE-SUSTAINING RHYTHM, WITH CPR/MOTION ARTIFACT. THE ELECTRODE BELT WAS DISCONNECTED AT 20:23:48 ON (B)(6) 2025. Manufacturer Narrative: INCOMING TESTING OF THE ELECTRODE BELT HAS BEEN COMPLETED. THE ELECTRODE BELT ECG ACQUISITION AND PULSE DELIVERY CIRCUITRY WAS TESTED AND FOUND TO BE FULLY FUNCTIONAL. THE MONITOR HAS BEEN RETURNED AND THE EVALUATION IS UNDERWAY. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21534737&pc=DTB | 2649622-2025-06456 | 2025-02-16 05:00 | Death | MPRI | 2025-03-06 05:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Failure to Capture; Electrical /Electronic Property Problem; Low impedance | Insufficient Information; No Clinical Signs, Symptoms or Conditions | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WAS A SUSPECT OF NON CAPTURE IN THE CURRENT ELECTROGRAM AND IMPEDANCE WARNINGS ON THE RIGHT ATRIAL (RA) LEAD FOR BOTH UNIPOLAR AND BIPOLAR CIRCUITS. A POLARITY SWITCH OCCURRED, WITH PREVIOUS BIPOLAR READINGS IN THE 300S AND UNIPOLAR READINGS IN THE UPPER 200S. THE IMPLANTABLE PULSE GENERATOR (IPG) ALSO EXHIBITED A FALSE LOW IMPEDANCE WARNING. THE LEAD AND IPG REMAIN IN USE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS NOTED THAT THE PATIENT PASSED AWAY. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE LEAD WAS NOT RETURNED FOR ANALYSIS, HOWEVER, PERFORMANCE DATA COLLECTED FROM THE DEVICE WAS RECEIVED AND ANALYZED. ANALYSIS OF THE DEVICE MEMORY INDICATED THE IMPEDANCE TREND OF THE ATRIAL PACING LEAD WAS DECREASING. ANALYSIS OF THE DEVICE MEMORY INDICATED A FALSE ALERT FOR LEAD IMPEDANCE OUT OF RANGE. ANALYSIS OF THE DEVICE MEMORY INDICATED A POLARITY SWITCH. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21534742&pc=NVZ | 3008973940-2025-01704 | 2025-02-16 05:00 | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-03-06 05:00 | NVZ | AZURE XT DR MRI SURESCAN | False Alarm | Insufficient Information; No Clinical Signs, Symptoms or Conditions | P980035 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WAS A SUSPECT OF NON CAPTURE IN THE CURRENT ELECTROGRAM AND IMPEDANCE WARNINGS ON THE RIGHT ATRIAL (RA) LEAD FOR BOTH UNIPOLAR AND BIPOLAR CIRCUITS. A POLARITY SWITCH OCCURRED, WITH PREVIOUS BIPOLAR READINGS IN THE 300S AND UNIPOLAR READINGS IN THE UPPER 200S. THE IMPLANTABLE PULSE GENERATOR (IPG) ALSO EXHIBITED A FALSE LOW IMPEDANCE WARNING. THE LEAD AND IPG REMAIN IN USE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS NOTED THAT THE PATIENT PASSED AWAY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE DEVICE WAS NOT RETURNED FOR ANALYSIS, HOWEVER, PERFORMANCE DATA COLLECTED FROM THE DEVICE WAS RECEIVED AND ANALYZED. ANALYSIS OF THE DEVICE MEMORY INDICATED A FALSE ALERT FOR LEAD IMPEDANCE OUT OF RANGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21535301&pc=PYX | 1220648-2025-26534 | 2025-02-10 05:00 | Death | ABIOMED, INC. | 2025-03-06 05:00 | PYX | IMPELLA RP FLEX | Pumping Stopped | Cardiac Arrest | P170011 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT IN POSTCARDIOTOMY CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA RP FLEX DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS). FOLLOWING A DIFFICULT DELIVERY, (THE IMPELLA WAS UNABLE TO ADVANCE THROUGH THE SHEATH INITIALLY), A CRITICAL ALARM APPEARED ON THE AUTOMATED IMPELLA CONTROLLER WHEN THE IMPELLA RP FLEX DEVICE WAS STARTED FOR PATIENT SUPPORT. AFTER SEVERAL FAILED TROUBLESHOOTING ATTEMPTS, THE PUMP ULTIMATELY WOULD NOT START AND HAD TO BE REPLACED, WHICH WAS SUCCESSFULLY COMPLETED FROM A DIFFERENT ACCESS POINT. THE FOLLOWING DAY, THE PATIENT PASSED AWAY NOTED FROM THE ("VERY POOR PROGNOSIS") CONDITION. IMPELLA AS A POSSIBLE ASSOCIATED FACTOR CANNOT BE DISCOUNTED; IT IS UNCLEAR HOW THE IMPELLA RELATED EVENT MAY HAVE AFFECTED THE OUTCOME. Manufacturer Narrative: THE INVESTIGATION OF PUMP DID NOT START HAS BEEN COMPLETED. THE DATA LOGS ARE UNAVAILABLE. THE PRODUCT WAS RETURNED AND EVALUATED. TWO LARGE KINKS WERE FOUND IN THE CATHETER. NO OTHER ABNORMALITIES WERE FOUND. THE PUMP PASSED ALL ELECTRICAL TESTING AND WAS RUN IN A BUCKET OF WATER FOR 15 MINUTES WITHOUT ISSUE. THE ROOT CAUSE OF THE PUMP NOT STARTING WAS NOT DETERMINED SINCE THE ISSUE WAS NOT REPRODUCED AND DATA LOGS ARE UNAVAILABLE. G1 MANUFACTURING SITE NAME AND ADDRESS WAS ERRONEOUSLY ENTERED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26534. Manufacturer Narrative: THE IMPELLA DEVICE WAS RETURNED, AND AN EVALUATION IS UNDERWAY. UPON COMPLETION OF THE ANALYSIS, A SUPPLEMENTAL REPORT WILL BE FILED. SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA CATHETER ¿ACHIEVING AN ACT = 250 SECONDS PRIOR TO REMOVING THE DILATOR WILL HELP PREVENT A THROMBUS FROM ENTERING THE CATHETER AND CAUSING A SUDDEN STOP ON STARTUP.¿ SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA RP WITH SMARTASSIST SYSTEM CATHETER: ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA RP WITH SMARTASSIST SYSTEM CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP. MONITOR MOTOR CURRENT AND CONSIDER REPLACING THE IMPELLA RP WITH SMARTASSIST SYSTEM CATHETER WHENEVER A RISE IN MOTOR CURRENT IS SEEN.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21535904&pc=DSQ | 2916596-2025-01178 | 2025-02-13 05:00 | Death | THORATEC CORPORATION | 2025-03-06 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTI-SYSTEM ORGAN FAILURE. ADDITIONAL INFORMATION WAS NOT PROVIDED.Event Description: AN AUTOPSY WAS NOT PERFORMED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. DUE TO PATIENT PRIVACY LAWS, THE MANAGING HOSPITAL WILL NOT COMMUNICATE ANY ADDITIONAL INFORMATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) REV. G, IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RESPIRATORY FAILURE, RIGHT HEART FAILURE, RENAL FAILURE, HEPATIC DYSFUNCTION), AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21536055&pc=OZD | 1220648-2025-26544 | 2025-02-08 05:00 | Death | ABIOMED, INC. | 2025-03-06 05:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. THE PATIENT EXPERIENCED EPISODES OF VENTRICULAR TACHYCARDIA DURING IMPELLA 5.5 SUPPORT. THE PATIENT WAS SHOCKED OVER 20 TIMES IN THE PERIOD OF A TWENTY-FOUR HOURS TO MANAGE THE CONDITION. THE PUMP WAS REPOSITIONED IN RESPONSE TO THE EPISODES AND A GANGLION BLOCK WAS UTILIZED TO RESOLVE THE ISSUE. SUPPORT CONTINUED WITH THE IMPLANTED PUMP. TWENTY-FOUR DAYS INTO SUPPORT, CARE WAS WITHDRAWN AND THE PATIENT PASSED AWAY. THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA DEVICE AS AN ASSOCIATED FACTOR IN THE PATIENT'S DEMISE. Manufacturer Narrative: THE INVESTIGATION OF VF/VT/AF/AT HAS BEEN COMPLETED. AS THE NECESSARY CLINICAL INFORMATION WAS NOT PROVIDED AND THE DEVICE WAS NOT RETURNED, THE ROOT CAUSE OF THE VT/VF WAS NOT DETERMINED. E4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26544 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21536072&pc=OZD | 1220648-2025-26545 | 2025-02-09 05:00 | Death | ABIOMED, INC. | 2025-03-06 05:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Atrial Fibrillation; Hematoma; Thrombosis/Thrombus | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK IMPLANTED WITH AN IMPELLA 5.5 DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS) EXPERIENCED A CLOT ON THE CATHETER INLET AND AN ACCESS SITE HEMATOMA REQUIRING ONE UNIT OF RED BLOOD CELLS. THE PATIENT WOULD SUBSEQUENTLY EXPIRE AFTER CHOOSING PALLIATIVE CARE. THE PATIENT UNDERWENT AN EMERGENT LEFT HEART CATHETERIZATION FOR ST-SEGMENT ELEVATION MYOCARDIAL INFARCTION; THROMBECTOMY TO THE LEFT ANTERIOR DESCENDING ARTERY WAS COMPLETED, AND AN IATROGENIC LAD PERFORATION OCCURRED. WHILE ON THE UNIT, THE PATIENT DECOMPENSATED AND THE DECISION WAS MADE TO PLACE AN IMPELLA 5.5, WHICH WAS SUCCESSFULLY COMPLETED VIA THE RIGHT AXILLARY ARTERY. THE PATIENT WAS TRANSFERRED BACK TO THE UNIT IN STABLE CONDITION. APPROXIMATELY EIGHT DAYS AFTER STARTING IMPELLA MCS, AN ECHOCARDIOGRAM WAS COMPLETED AND A CLOT ON THE IMPELLA INLET WAS OBSERVED. THE PATIENT WAS ON SYSTEMIC HEPARIN; PLANS WERE TO MONITOR MOTOR TRENDS. THE FOLLOWING DAY, PERFORMANCE LEVEL WAS REDUCED DUE TO A SUSPECTED CLOT ON THE LEFT VENTRICLE. TWO DAYS THEREAFTER, A HEMATOMA AT THE ACCESS SITE WAS NOTED REQUIRING ONE UNIT OF PACKED RED BLOOD CELLS. THE PATIENT HAD INCREASING PRESSOR AND DOBUTAMINE REQUIREMENTS WITH CONTINUOUS RENAL REPLACEMENT THERAPY FLUID REMOVAL OVERNIGHT. CENTRAL VENOUS PRESSURE WAS ON THE HIGHER END DESPITE AGGRESSIVE FLUID REMOVAL. THE PLAN WAS DISCUSSED WITH STAFF TO CONTINUE WITH CURRENT THERAPIES; HOWEVER, POTENTIALLY SWITCH TO COMFORT CARE. THE FOLLOWING DAY THE PATIENT WENT INTO ATRIAL FIBRILLATION AND AN AMIODARONE DRIP WAS STARTED. THE PATIENT STATED BEING TIRED AND NO LONGER WANTED TO BE ON SUPPORT. THERAPIES WOULD CONTINUE UNTIL THE NEXT DAY. THE PATIENT WITH FAMILY DECIDED ON PALLIATIVE CARE. PERFORMANCE LEVEL WAS DECREASED TO P-2 (FROM P-4) TO PREPARE FOR IMPELLA EXPLANT BY THE SURGICAL TEAM, WHICH WAS SUCCESSFULLY COMPLETED. THE PATIENT WAS HEMODYNAMICALLY STABLE UNTIL THE DEMISE. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: IMPELLA 5.5 WITH SMARTASSIST SYSTEM. SECTION: GENERAL PATIENT CARE CONSIDERATIONS: ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: POTENTIAL ADVERSE EVENTS: ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY¿ (INCLUDING VENTRICULAR PERFORATION).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21536079&pc=OZD | 1220648-2025-26546 | 2025-02-11 05:00 | Death | ABIOMED, INC. | 2025-03-06 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Perforation of Vessels | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A 69-YEAR-OLD MALE PATIENT HAD AN IMPELLA CP PLACED DUE TO A STEMI. AFTER THE CP WAS PLACED, THERE WAS SUCTION THAT WAS TREATED BY RED BLOOD CELLS (RBCS) AND OTHER BLOOD PRODUCTS/VOLUME. THE TEAM THEN OBSERVED THERE WAS A HEMATOMA AT THE ACCESS SITE. PATIENT CONDITION CONTINUED TO DETERIORATE AND THE TEAM OBSERVED IN THE OR THAT THE LV WAS PERFORATED. THE PERFORATION WAS REPAIRED AND THE PUMP REMAINS ON WHILE PATIENT BEING SENT TO TERTIARY CENTER. FURTHER DETAILS STATE THAT THE PATIENT SUBSEQUENTLY EXPIRED WHILE ON SUPPORT AS THE BLEEDING COULD NOT BE CONTROLLED. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMART ASSIST SYSTEM STATES THE FOLLOWING: SECTION: GENERAL PATIENT CARE CONSIDERATIONS. ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT: USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS & CAUTIONS: WARNINGS, SECTION: PRE-SUPPORT EVALUATION, SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER, SECTION: DIRECT AORTIC INSERTION, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION. [ADDITION FOR 5.0, 5.5 ONLY] IN INSTANCES WHERE THE IMPELLA PUMP HAS BEEN PLACED PRIOR TO PERFORMING CARDIAC SURGERY WITH AORTIC CROSS CLAMPING AND CARDIOPLEGIC ARREST, CARE SHOULD BE TAKEN WHEN MANIPULATING THE HEART WHEN THE PUMP POSITION IS FIXED WITH APPLICATION OF THE AORTIC CROSS CLAMP ACROSS THE PUMP CATHETER.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE. SECTION: WARNINGS & CAUTIONS: WARNINGS, SECTION: PRE-SUPPORT EVALUATION, SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER, SECTION: DIRECT AORTIC INSERTION, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: ¿TO REDUCE THE RISK OF CARDIAC INJURY (INCLUDING VENTRICULAR PERFORATION), PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY IN THE SETTINGS OF ACUTE INFARCTION WITH TISSUE NECROSIS.¿ ¿TO REDUCE THE RISK OF VASCULAR INJURY, PHYSICIANS SHOULD EXERCISE CAUTION WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX PERIPHERAL VASCULAR ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: UNREPAIRED ABDOMINAL AORTIC ANEURYSM, SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM, DISSECTION OF THE ASCENDING/ TRANSVERSE/DESCENDING AORTA, CHRONIC ANATOMICAL CHANGES IN THE RELATIONSHIP OF THE AORTA/AORTIC VALVE/VENTRICULAR ALIGNMENT, SIGNIFICANT MOBILE ATHEROMATOUS DISEASE IN THE THORACIC OR ABDOMINAL AORTA OR PERIPHERAL VESSELS.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING A RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY IN THE LEFT VENTRICLE AFTER CPR WITH ECHOCARDIOGRAPHY GUIDANCE. SECTION: WARNINGS & CAUTIONS: CAUTIONS ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21536171&pc=MIH | 9612164-2025-01202 | 2022-07-12 04:00 | Death | MEDTRONIC IRELAND | 2025-03-06 05:00 | MIH | VALIANT CAPTIVIA UNK | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P100040 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING ¿ OUTCOMES OF ZONE 1 THORACIC ENDOVASCULAR AORTIC REPAIR WITH FENESTRATED SURGEON-MODIFIED STENT-GRAFT FOR AORTIC ARCH PATHOLOGIES¿. THE TIME FRAME OF THIS STUDY WAS OVER A FOUR-YEAR PERIOD. VALIANT CAPTIVIA STENT GRAFTS WERE FENESTRATED AND IMPLANTED IN 34 PATIENTS. THIS STUDY EVALUATED THE FEASIBILITY AND SAFETY OF ZONE 1 THORACIC ENDOVASCULAR AORTIC REPAIR (TEVAR) WITH FENESTRATED SURGEON-MODIFIED STENT-GRAFT (SMSG) FOR AORTIC ARCH PATHOLOGIES. PATIENT MORTALITY WAS REPORTED INCLUDING AORTIC RELATED DEATHS. NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. Manufacturer Narrative: MEDTRONIC RECEIVED THE FOLLOWING INFORMATION FROM A JOURNAL ARTICLE ENTITLED; ¿OUTCOMES OF ZONE 1 THORACIC ENDOVASCULAR AORTIC REPAIR WITH FENESTRATED SURGEON MODIFIED STENT-GRAFT FOR AORTIC ARCH PATHOLOGIES¿. LI X, ZHANG L, SONG C, ZHANG H, XIA S, ZHU L, GUO W, LI H, JING Z, LU Q. JOURNAL OF ENDOVASCULAR THERAPY. 2024 FEB;31(1):62-68. DOI: 10.1177/15266028221108903. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATE TO H6; CODING. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21536199&pc=QNR | 3003306248-2025-00062 | 2025-02-04 05:00 | Death | THORATEC SWITZERLAND GMBH | 2025-03-06 05:00 | QNR | CENTRIMAG BLOOD PUMP | Fracture; Gas/Air Leak; Adverse Event Without Identified Device or Use Problem | Air Embolism; Exsanguination | K020271 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS ON VENO-VENOUS EXTRACORPOREAL MEMBRANE OXYGENATION (VV ECMO). ON (B)(6) 2025 AT APPROXIMATELY 23:32, IT WAS IDENTIFIED THAT THE VENOUS CANNULA CRACKED AND DISCONNECTED FROM CIRCUIT. A CIRCUIT EXCHANGE WAS PERFORMED IN ATTEMPT TO STOP EXSANGUINOUS DRAINAGE. DURING TIME OF CIRCUIT, THE PATIENT CODED AND WAS PLACED BACK ON VV ECMO BUT BUBBLES WERE IDENTIFIED IN THE CIRCUIT WITH INABILITY TO GET OUT WITHOUT CAUSING ADDITIONAL HARM TO PATIENT. THE PATIENT PASSED AWAY AT 00:05 ON (B)(6) 2025 WHEN THE CENTRIMAG BLOOD PUMP WAS "BUMPED BY STAFF" WHICH CAUSED AIR TO GET IN THE CIRCUIT. THE PATIENT WAS NON-PULSATILE ALREADY AND THE AIR CAUSED THE PATIENT TO PASS. PER THE SITE, GIVEN THE CIRCUMSTANCES, IT WAS BELIEVED THAT THE HARDWARE WAS NOT AT FAULT AND WOULD NOT BE RETURNED.Event Description: THE STAFF IN THE ROOM "KNOCKED INTO THE PUMP" ALLEGEDLY CAUSING A CRACK WHICH REQUIRED SWITCHING THE CIRCUIT DUE TO THE PATIENT EXSANGUINATING. A NEW PRE-PRIMED CIRCUIT WAS INITIATED, AND AIR WAS IDENTIFIED POST INITIATION. THE PATIENT DECEASED BECAUSE THE AIR WAS UNABLE TO BE SAFELY REMOVED. THE PUMP WAS A PRE-PRIMED CIRCUIT THAT WAS READY FOR EMERGENCIES PRIOR TO THE EVENT. PER TWO SEPARATE PHONE CALLS, AIR WAS IDENTIFIED IN THE CIRCUIT BEFORE INITIATING SUPPORT AND LIKELY ENTERED THE SYSTEM DURING PRIMING. AFTER FOLLOW-UP THE ECMO COORDINATOR COMMUNICATED "THE CHANGE OUT WAS ATTEMPTED BUT WAS UNSUCCESSFUL BECAUSE THE PATIENT [HAD PASSED AWAY]¿.Event Description: THE ECMO SPECIALIST AND REGISTERED NURSE (RN) WERE BATHING THE PATIENT, AFTER THE BATH, THE RN WALKED PAST THE MACHINE AND BUMPED IT WITH THEIR HIP. THE MACHINE THEN BEGAN ALARMING LOW FLOW AND THE ECMO SPECIALIST RESPONDED AND NOTICED EXSANGUINATION COMING FROM THE TUBING. REGARDING THE SECOND CIRCUIT THAT COULD NOT BE PLACED ON THE PATIENT, IT WAS UNABLE TO BE IDENTIFIED HOW THERE WAS AIR IN IT. IT BEGAN ALARMING WHEN THERE WERE BUBBLES AFTER PERFORMING THE CIRCUIT EXCHANGE. Manufacturer Narrative: CORRECTED DATA SECTION H10: MEDWATCH NUMBER ADDED. NO FURTHER INFO WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE REPORTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: D4: DEVICE EXPIRATION DATE CANNOT BE DETERMINED. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: SECTION D4: LOT NUMBER PREVIOUSLY PROVIDED WAS INVALID. SERIAL NUMBER WAS NOT PROVIDED. EXPIRATION DATE AND MANUFACTURE DATE (H4) CANNOT BE DETERMINED. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION. SECTION D5: OPERATOR OF DEVICE CORRECTED. SECTION H6: HEALTH EFFECT - CLINICAL CODE CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A SPECIFIC CAUSE FOR THE REPORTED EVENTS COULD NOT BE CONCLUSIVELY DETERMINED, AND A DIRECT CORRELATION WITH THE CENTRIMAG BLOOD PUMP COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE PATIENT WAS PLACED ON VENO-VENOUS EXTRACORPOREAL MEMBRANE OXYGENATION (VV ECMO) ON (B)(6) 2025. AT APPROXIMATELY 23:32 ON (B)(6) 2025, THE VENOUS CANNULA CRACKED AND DISCONNECTED FROM THE CIRCUIT. THE ACCOUNT ATTEMPTED TO STOP EXSANGUINOUS DRAINAGE AND PERFORMED A CIRCUIT EXCHANGE. AS THE CIRCUIT WAS BEING EXCHANGED, THE PATIENT CODED AND WAS PLACED BACK ON VV ECMO. HOWEVER, BUBBLES WERE IDENTIFIED IN THE CIRCUIT THAT WERE UNABLE TO BE REMOVED WITHOUT CAUSING ADDITIONAL HARM TO THE PATIENT. DESPITE EFFORTS, THE PATIENT EXPIRED AT 00:05 ON (B)(6) 2025. IT WAS REPORTED THAT THE CENTRIMAG PUMP HAD BEEN BUMPED BY THE STAFF, WHICH CAUSED AIR TO GET INTO THE CIRCUIT. THE CENTRIMAG EQUIPMENT WAS NOT CONSIDERED TO BE AT FAULT, SO IT WAS REPORTED THAT IT WOULD NOT BE RETURNED FOR EVALUATION. THE PROVIDED LOT NUMBER WAS NOT VALID, AND THE CONTACTED CUSTOMER DID NOT HAVE IDENTIFYING INFORMATION FOR THE PRODUCT. THE CENTRIMAG BLOOD PUMP INSTRUCTIONS FOR USE (IFU) LISTS BLEEDING AND DEATH AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE CENTRIMAG CIRCULATORY SUPPORT SYSTEM UNDER ¿ADVERSE EVENTS¿. THIS IFU ALSO PROVIDES THE FOLLOWING WARNINGS AND CAUTIONS: IFU WARNING #2: ENSURE THAT THE PUMP AND CIRCUIT HAVE BEEN DEBUBBLED AND PRIMED PROPERLY PRIOR TO USE TO MINIMIZE THE RISK OF AIR ENTRY TO THE PATIENT. IFU WARNING #3: MASSIVE AIR ENTRY INTO THE PUMP WILL CAUSE THE PUMP TO DEPRIME AND BLOOD FLOW TO STOP. CLAMP THE OUTLET TUBING, STOP THE PUMP, AND REMOVE AIR PRIOR TO RESUMING CIRCULATION. IFU WARNING #7: IT IS INTENDED THAT SYSTEMIC ANTICOAGULATION BE UTILIZED WHILE THIS DEVICE IS IN USE. ANTICOAGULATION LEVELS SHOULD BE DETERMINED BY THE PHYSICIAN BASED ON RISKS AND BENEFITS TO THE PATIENT. IFU WARNING #10: FREQUENT PATIENT AND DEVICE MONITORING IS RECOMMENDED. IFU WARNING #14: DO NOT OPERATE THE PUMP WITH ITS INLET TUBING CLAMPED AS A NEGATIVE PRESSURE WILL BE GENERATED IN THE PUMP AND AIR BUBBLES MAY BE FORMED IN THE PRIMING FLUID OR BLOOD. IFU WARNING #16: AS WITH ALL CONTINUOUS FLOW PUMPS, OPERATING AT TOO HIGH A SPEED CAN RESULT IN NEGATIVE PRESSURE AT THE INLET WHICH CAN LEAD TO COLLAPSE OF THE VENTRICLE OR BLOOD VESSELS, INLET CANNULA OBSTRUCTION, INSPIRATION OF AIR, OUTGASSING, CAVITATION AND INCREASED RISK OF EMBOLISM. ALWAYS OPERATE THE SYSTEM AT THE LOWEST SPEED CONSISTENT WITH THE VOLUME OF BLOOD AVAILABLE TO BE PUMPED AND CLINICALLY ACCEPTABLE CIRCULATORY SUPPORT. IFU CAUTION #2: THIS DEVICE SHOULD ONLY BE USED BY PERSONS THOROUGHLY TRAINED IN EXTRACORPOREAL CIRCULATION PROCEDURES. IFU CAUTION #15: ALWAYS HAVE A BACKUP CENTRIMAG PUMP, CONSOLE, MOTOR, AND ACCESSORIES AVAILABLE FOR USE. THE SECTION TITLED ¿PUMP SETUP AND OPERATION¿ PROVIDES INSTRUCTIONS FOR INSERTING THE CANNULA AND DE-AIRING THE SYSTEM. THIS SECTION WARNS ¿DO NOT OVER TIGHTEN SUTURES WHEN SECURING CANNULAE TO TISSUES AND VESSELS. OVER TIGHTENED SUTURES MAY RESULT IN OBSTRUCTION AND INTERRUPTION OF BLOOD FLOW THROUGH THE CANNULAE. SUTURING USED TO SECURE CANNULA MUST BE MADE WITH SUFFICIENT TENSION TO HOLD THE CANNULA IN PLACE OVER THE FULL RANGE OF PATIENT ACTIVITY. FAILURE TO EFFECTIVELY SECURE CANNULAE IN PLACE POSES RISK OF DECANNULATION, BLEEDING, OR AIR EMBOLUS.¿ THIS SECTION ALSO WARNS ¿IF LEAKS OR OTHER ANOMALIES ARE FOUND IN THE CIRCUIT, REMOVE THE PUMP AND REPLACE WITH A NEW, STERILE PUMP, REPEATING THE ABOVE STEPS TO PRIME.¿ NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: D2B: CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: IT WAS REPORTED THAT A CIRCUIT CRACK AND DISCONNECTION WAS DUE TO A STAFF MEMBER ACCIDENTALLY BUMPING INTO THE CENTRIMAG PUMP; HOWEVER, A SPECIFIC CAUSE FOR THESE EVENTS COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. NO DAMAGE OR DEVICE-RELATED ISSUES WERE IDENTIFIED THROUGH THE EVALUATION OF THE RETURNED CENTRIMAG BLOOD PUMP. THE PATIENT WAS PLACED ON VENO-VENOUS EXTRACORPOREAL MEMBRANE OXYGENATION (VV ECMO) ON (B)(6) 2025. AT APPROXIMATELY 23:32 ON (B)(6) 2025, IT WAS REPORTED THAT A STAFF MEMBER ACCIDENTALLY BUMPED INTO THE CENTRIMAG PUMP, CAUSING A CRACK AND CIRCUIT DISCONNECTION. THE ACCOUNT ATTEMPTED TO STOP EXSANGUINOUS DRAINAGE AND PERFORMED A CIRCUIT EXCHANGE. THE CENTRIMAG BLOOD PUMP, LOT # L08279-LA4, WAS RETURNED WITHOUT ANY CONNECTIONS TO TUBING. SECTIONS OF TUBING WERE ALSO RETURNED; HOWEVER, IT WAS UNABLE TO BE DETERMINED WHICH OF THE RETURNED TUBING, IF ANY, HAD BEEN CONNECTED TO THE RETURNED PUMP. THERE WERE NO REMARKABLE FINDINGS REGARDING THE RETURNED TUBING. THE INLET AND OUTLET PORTS SHOWED NO EVIDENCE OF CRACKING OR OTHER DAMAGE. EXAMINATION OF THE PUMP'S BLOOD-CONTACTING SURFACES AND TUBING REVEALED NO EVIDENCE OF DEPOSITIONS OR THROMBUS FORMATIONS. VISUAL INSPECTION OF THE BLOOD PUMP REVEALED NO EVIDENCE OF DAMAGE TO THE IMPELLER BLADES, ROTOR BODY, ROTOR WELL, OR PUMP HOUSING. FOLLOWING CLEANING, VISUAL AND MICROSCOPIC EXAMINATION OF THE BLOOD PUMP REVEALED NO DAMAGE OR ABNORMALITIES. THE BLOOD PUMP WAS FUNCTIONALLY TESTED ON A MOCK CIRCULATORY LOOP WITH A TEST MOTOR AND EQUIPMENT. THE BLOOD PUMP FUNCTIONED AS INTENDED IN ACCORDANCE WITH MANUFACTURING SPECIFICATIONS. THE CENTRIMAG BLOOD PUMP INSTRUCTIONS FOR USE (IFU) CAUTIONS TO NOT HIT OR STRIKE THE PUMP WITH HANDS, OBJECTS, OR INSTRUMENTS. DO NOT STRIKE THE PUMP AGAINST ANY SURFACE OR OBJECT. THE IFU ALSO CAUTIONS THAT SHOCK MAY CAUSE DAMAGE TO THE DEVICE, WHICH MAY CAUSE DEVICE MALFUNCTION. BLEEDING IS LISTED IN THE IFU AS AN ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE CENTRIMAG CIRCULATORY SUPPORT SYSTEM. THE IFU STATES THAT IT IS INTENDED THAT SYSTEMIC ANTICOAGULATION BE UTILIZED WHILE THE DEVICE IS IN USE AND ANTICOAGULATION LEVELS SHOULD BE DETERMINED BY THE PHYSICIAN BASED ON RISKS AND BENEFITS TO THE PATIENT. THE IFU ALSO WARNS THAT FREQUENT PATIENT AND DEVICE MONITORING IS RECOMMENDED. REVIEW OF THE DEVICE HISTORY RECORD (DHR) FOR THE CENTRIMAG BLOOD PUMP, LOT #L08279-LA4, REVEALED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CENTRIMAG BLOOD PUMP INSTRUCTIONS FOR USE (IFU) LISTS BLEEDING AS AN ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE CENTRIMAG CIRCULATORY SUPPORT SYSTEM UNDER ¿ADVERSE EVENTS¿. THIS IFU ALSO PROVIDES THE FOLLOWING WARNINGS AND CAUTIONS: IFU WARNING #7: IT IS INTENDED THAT SYSTEMIC ANTICOAGULATION BE UTILIZED WHILE THIS DEVICE IS IN USE. ANTICOAGULATION LEVELS SHOULD BE DETERMINED BY THE PHYSICIAN BASED ON RISKS AND BENEFITS TO THE PATIENT. IFU WARNING #10: FREQUENT PATIENT AND DEVICE MONITORING IS RECOMMENDED. IFU CAUTION #2: THIS DEVICE SHOULD ONLY BE USED BY PERSONS THOROUGHLY TRAINED IN EXTRACORPOREAL CIRCULATION PROCEDURES. IFU CAUTION #3: DO NOT HIT OR STRIKE THE PUMP WITH HANDS, OBJECTS, OR INSTRUMENTS. DO NOT STRIKE THE PUMP AGAINST ANY SURFACE OR OBJECT. SHOCK MAY CAUSE DAMAGE TO THE DEVICE, WHICH MAY CAUSE DEVICE MALFUNCTION. IFU CAUTION #12: DO NOT USE THE PUMP IF IT HAS BEEN DROPPED. DROPPING OR OTHER SEVERE SHOCK MAY CAUSE DAMAGE WHICH COULD LEAD TO DEVICE MALFUNCTION. IFU CAUTION #15: ALWAYS HAVE A BACKUP CENTRIMAG PUMP, CONSOLE, MOTOR, AND ACCESSORIES AVAILABLE FOR USE. THE SECTION TITLED ¿PUMP SETUP AND OPERATION¿ WARNS ¿IF LEAKS OR OTHER ANOMALIES ARE FOUND IN THE CIRCUIT, REMOVE THE PUMP AND REPLACE WITH A NEW, STERILE PUMP, REPEATING THE ABOVE STEPS TO PRIME.¿ NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21536365&pc=MCX | 2124215-2025-08855 | 2025-02-03 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-03-06 05:00 | MCX | ROTAPRO | Adverse Event Without Identified Device or Use Problem | Cardiac Tamponade | P900056 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. AFTER A PROCEDURE INVOLVING A 1.50MM ROTAPRO WAS FINALIZED, THE PATIENT PRESENTED WITH CARDIAC TAMPONADE. THE PATIENT WAS SEDATED, AND CARDIAC SURGERY WAS PERFORMED. HOWEVER, THE PATIENT DIED DUE TO HEART FAILURE. Manufacturer Narrative: E1: INITIAL REPORTER FIRST NAME: (B)(6). GOOD FAITH EFFORT ATTEMPTS WERE MADE TO TRY AND RETRIEVE ADDITIONAL DETAILS REGARDING THE REPORTED EVENT, BUT FURTHER INFORMATION WAS UNABLE TO BE OBTAINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21536485&pc=HIF | 1320894-2025-00067 | 2025-01-20 05:00 | Death | W.O.M. WORLD OF MEDICINE GMBH | 2025-03-06 05:00 | HIF | AIRSEAL | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | K190303 |  | 1 | Event Description: THE SALES REPRESENTATIVE REPORTED ON BEHALF OF THE CUSTOMER THAT THE AS-IFS1, AIRSEAL IFS, 120V WAS BEING USED ON 20JAN25 AND ¿WHILE THE CLAMP WAS PLACED, THEY ENCOUNTERED A SIGNIFICANT AMOUNT OF BLEEDING. BLOOD ENTERED THE AIRSEAL TUBE SET, AND THE FLUID INGRESS WARNING APPEARED. THE TUBE SET WAS NOT CHANGED AND THE UNIT SHUT DOWN. SURGEON CONVERTED TO OPEN.¿. IT WAS FURTHER REPORTED THAT THE BLOOD ENTERED THE AIRSEAL TUBE SET AS "...THEY GOT INTO HEAVY BLEEDING AND WERE USING EXCESSIVE SUCTION.". WHEN THE AIRSEAL FLUID INGRESS WARNING APPEARED THE TUBESET WAS NOT CHANGED OUT "...AS THEY STARTED RECEIVING ERROR CODES, THEY WERE ON CLAMP AND TRYING TO MANAGE BLEEDING SO POSSIBLY DUE TO MULTIPLE THINGS GOING ON. THEY ULTIMATELY CONVERTED TO OPEN INSTEAD OF SWAPPING THE TUBE SET.". THE IAS12-120LPI AND ASM-EVAC1 WILL BE LISTED AS CONCOMITANT DEVICES AND THERE WERE NO ALLEGATIONS AGAINST THEM. THIS REPORT IS BEING RAISED ON THE BASIS OF DEATH RELATED TO A DEVICE BEING USED DURING A PROCEDURE WITHOUT ANY REPORT OF A MALFUNCTION.Event Description: THE SALES REPRESENTATIVE REPORTED ON BEHALF OF THE CUSTOMER THAT THE AS-IFS1, AIRSEAL IFS, 120V WAS BEING USED ON (B)(6) 2025 AND ¿WHILE THE CLAMP WAS PLACED, THEY ENCOUNTERED A SIGNIFICANT AMOUNT OF BLEEDING. BLOOD ENTERED THE AIRSEAL TUBE SET, AND THE FLUID INGRESS WARNING APPEARED. THE TUBE SET WAS NOT CHANGED AND THE UNIT SHUT DOWN. SURGEON CONVERTED TO OPEN.¿ IT WAS FURTHER REPORTED THAT THE BLOOD ENTERED THE AIRSEAL TUBE SET AS "...THEY GOT INTO HEAVY BLEEDING AND WERE USING EXCESSIVE SUCTION." WHEN THE AIRSEAL FLUID INGRESS WARNING APPEARED THE TUBESET WAS NOT CHANGED OUT "...AS THEY STARTED RECEIVING ERROR CODES, THEY WERE ON CLAMP AND TRYING TO MANAGE BLEEDING SO POSSIBLY DUE TO MULTIPLE THINGS GOING ON. THEY ULTIMATELY CONVERTED TO OPEN INSTEAD OF SWAPPING THE TUBE SET." THE IAS12-120LPI AND ASM-EVAC1 WILL BE LISTED AS CONCOMITANT DEVICES AND THERE WERE NO ALLEGATIONS AGAINST THEM. THIS REPORT IS BEING RAISED ON THE BASIS OF DEATH RELATED TO A DEVICE BEING USED DURING A PROCEDURE, WITHOUT ANY REPORT OF A MALFUNCTION, WHERE THE PATIENT EXPERIENCED HEAVY BLEEDING AND EXPIRED. Manufacturer Narrative: THE REPORTED DEVICE HAS BEEN RETURNED TO CONMED; HOWEVER, THE INVESTIGATION OF THE DEVICE HAS NOT BEEN COMPLETED. A SUPPLEMENTAL AND FINAL REPORT WILL BE FILED FOLLOWING THE COMPLETION OF THE DEVICE COMPLAINT INVESTIGATION. WE WILL CONTINUE TO MONITOR PER REGULATORY REQUIREMENTS TO HELP ENSURE PATIENT SAFETY. Manufacturer Narrative: THE IAS12-120LPI AND ASM-EVAC1 DEVICES WERE NOT RETURNED FOR EVALUATION. AS-IFS1: THE PREVENTATIVE MAINTENANCE WAS OVERDUE AND ADDITIONAL PROBLEMS WERE FOUND. THE COMPRESSOR VALVE ERRED. THE HIGH-PRESSURE SENSOR FAILED. THE COMPRESSOR WAS REFURBISHED. THE UNIT ALSO PRESENTED PHYSICAL DAMAGE WITH A CRACKED BEZEL, AND HOUSING DAMAGE. THE SERVICE HISTORY WAS REVIEWED AND NO RELATIONSHIP TO THIS COMPLAINT WAS FOUND. A DEVICE HISTORY RECORD (DHR) REVIEW WAS REQUESTED FROM THE MANUFACTURER, AND NO INDICATION OF ABNORMALITIES WAS COMMUNICATED. (B)(4). PER THE INSTRUCTIONS FOR USE, THE USER IS ADVISED THE FOLLOWING: WHEN WORKING IN THE AIRSEAL MODE, IT IS POSSIBLE THAT FLUIDS CAN BE EXTRACTED FROM THE SURGICAL FIELD OVER THE EVACUATION LINE AND INTO THE FILTER HOUSING. IN ORDER TO PREVENT CONTAMINATION OF THE DEVICE, BODILY FLUID IS RETAINED BY THE FLUID TRAP OF THE HOUSING COMPONENT. THE CAPACITY OF THE FLUID TRAP IS LIMITED. A WARNING MESSAGE AND AN AUDIBLE SIGNAL WILL BE EMITTED IF THE FILL LEVEL LOW IS REACHED. AIRSEAL CANNULA AND TUBING PLACEMENT SHOULD BE CHECKED TO MAKE SURE NO MORE FLUID ENTERS THE FLUID TRAP. AT THIS POINT, INSERT AIRSEAL OBTURATOR AND CHANGE THE FILTERED TUBE SET. INSUFFLATION CAN BE CONTINUED WITH FULL FUNCTIONALITY. IF FLUID LEVEL HIGH IS REACHED, THE AIRSEAL FUNCTION WILL SHUT DOWN. INSERT OBTURATOR IN AIRSEAL CANNULA TO MAINTAIN PRESSURE WITH NORMAL INSUFFLATION. CHANGE THE TUBE SET TO FINISH THE PROCEDURE. WE WILL CONTINUE TO MONITOR PER REGULATORY REQUIREMENTS TO HELP ENSURE PATIENT SAFETY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21536636&pc=DTZ | 9681834-2025-00035 | 2025-02-11 05:00 | Death | TERUMO CORPORATION, ASHITAKA | 2025-03-06 05:00 | DTZ | CAPIOX | Failure to Deliver | Insufficient Information | K022115 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT THE INVOLVED OXYGENATOR DID NOT FUNCTION PROPERLY. THE EVENT OCCURRED DURING CARDIOPULMONARY BYPASS. THE EVENT RESULTED IN DELAY IN BEGINNING OR CONTINUING THE SURGICAL PROCEDURE. AFTER 3 HOURS AND 40 MINUTES THE DEVICE'S ABILITY TO OXYGENATE WAS NOT FUNCTIONING. SWEEP WAS AT 1.2 LPM AND 60% FIO2, PO2 WAS 295 THEN NEXT ENTRY FELL TO 95 AND CONTINUED FALLING DESPITE INCREASES IN FIO2 AND SWEEP. A SECOND OXYGENATOR WAS ADDED TO THE RECIRCULATION LINE TO FACILITATE OXYGENATION. THEY DID NOT CHANGE OUT THE OXYGENATOR THEY ADDED ANOTHER OXYGENATOR TO THE CIRCUIT TO FINISH THE CASE. THE SURGERY WAS COMPLETED SUCCESSFULLY. NO BLOOD LOSS WAS REPORTED. A DEATH OCCURRED. THE PATIENT DID NOT DIE IN SURGERY, THEY DIED IN INTENSE CARE UNIT (ICU). Manufacturer Narrative: D6A: IMPLANTED DATE: DEVICE WAS NOT IMPLANTED. D6B: EXPLANTED DATE: DEVICE WAS NOT EXPLANTED. E3: OCCUPATION: PERFUSION SUPERVISOR. THE ACTUAL SAMPLE WAS NOT RETURNED. THE MANUFACTURING RECORD AND THE SHIPPING INSPECTION RECORD OF THE ACTUAL PRODUCT FOUND NO ANOMALY. NO OTHER SIMILAR REPORT OF THE PRODUCT WITH THE INVOLVED PRODUCT CODE/LOT NUMBER WAS FOUND. BASED ON THE INVESTIGATION RESULT, NO ANOMALY WAS FOUND IN THE MANUFACTURING RECORDS. AS A CAUSE OF DECREASING IN PAO2, BASED ON OUR PAST EXPERIENCE, FOLLOWING FACTORS WERE INFERRED. HOWEVER, SINCE THE ACTUAL SAMPLE COULD NOT BE CONFIRMED, IT WAS NOT POSSIBLE TO CLARIFY THE CAUSE OF OCCURRENCE. REGARDING FIO2 DURING CIRCULATION, THE AMOUNT OF OXYGEN SUPPLY WAS INSUFFICIENT COMPARED TO THE PATIENT'S OXYGEN CONSUMPTION. GAS EXCHANGE WAS HINDERED BY SOME FACTOR (E.G., BLOOD CLOT, WET LUNG). AS THE PATIENT'S METABOLISM BECAME MORE ACTIVE DUE TO REWARMING, THE OXYGEN SUPPLY WAS INSUFFICIENT COMPARED TO THE PATIENT'S OXYGEN CONSUMPTION, CAUSING A DECREASE IN SVO2. THE INSTRUCTIONS FOR USE (IFU) (CAPIOX RX05) HAS THE FOLLOWING WARNINGS, METHOD OF OPERATION AND PRECAUTION REGARDING GAS EXCHANGE FAILURE: "START GAS SUPPLY WITH V/Q=1, AND FIO2=100%, THEN MAKE ADJUSTMENTS BASED ON BLOOD GAS MEASUREMENTS." "MEASURE BLOOD GASES AND MAKE NECESSARY ADJUSTMENTS AS FOLLOWS. A. CONTROL PAO2 BY CHANGING CONCENTRATION OF OXYGEN IN VENTILATING GAS USING GAS BLENDER. TO DECREASE PAO2, DECREASE FIO2. TO INCREASE PAO2, INCREASE FIO2. B. CONTROL PACO2 BY CHANGING THE TOTAL GAS FLOW. TO DECREASE PACO2, INCREASE TOTAL GAS FLOW. TO INCREASE PACO2, DECREASE TOTAL GAS FLOW." "UPON PATIENT REWARMING, ADJUST O2 CONCENTRATION, GAS FLOW RATE AND BLOOD FLOW RATE BY INCREASING THEM AS NEEDED BASED ON AN INCREASE IN PATIENTS METABOLISM. FAILURE TO ADJUST THE GAS SUPPLY AND THE BLOOD FLOW RATE APPROPRIATELY MAY CAUSE INSUFFICIENT O2 SUPPLY NEEDED OR THE AMOUNT OF THE PATIENT'S GASEOUS METABOLISM." "A PHENOMENON CALLED WET LUNG MAY OCCUR WHEN WATER CONDENSATION OCCURS INSIDE FIBERS OF MICROPOROUS MEMBRANE OXYGENATORS WITH BLOOD FLOWING EXTERIOR TO THE FIBERS. THIS MAY OCCUR WHEN OXYGENATORS ARE USED FOR A LONGER PERIOD OF TIME. IF WATER CONDENSATION AND/OR A DECREASE IN PAO2 AND/OR AN INCREASE IN PACO2 IS NOTED DURING EXTENDED OXYGENATOR USE, BRIEFLY INCREASING THE GAS FLOW RATE MAY IMPROVE THE PERFORMANCE. INCREASE GAS FLOW RATE, TO 5 L/MIN FOR 10 SECONDS. DO NOT REPEAT THIS FLUSHING TECHNIQUE, EVEN IF OXYGENATOR PERFORMANCE IS NOT IMPROVED." TERUMO MEDICAL CORPORATION (TMC) (IMPORTER) REGISTRATION NO. (B)(4) IS SUBMITTING THIS REPORT ON BEHALF OF ASHITAKA FACTORY OF TERUMO CORPORATION (MANUFACTURER) REGISTRATION NO. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21536708&pc=OZD | 1220648-2025-26555 | 2025-02-07 05:00 | Death | ABIOMED, INC. | 2025-03-06 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT AND APPROXIMATELY FOUR DAYS AFTER START OF SUPPORT, THE PATIENT DEVELOPED LIMB ISCHEMIA IN THE IMPELLA LEG. THE STAFF WAS UNSURE IF THE ISCHEMIA WAS DUE TO THE REPOSITIONING SHEATH OR IF THE PATIENT HAD A THROMBUS IN THE AREA. ACTIVATED CLOTTING TIME WAS BETWEEN 150-160 SECONDS DURING IMPELLA THERAPY. THE IMPELLA CP DEVICE WAS EXPLANTED, AND THE PATIENT EXPIRED THE SAME DAY AS THE ISCHEMIC EVENT. THE PUMP WORKED WELL SO THE USER FACILITY THREW THE PUMP IN THE TRASH. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21536806&pc=LWS | 2649622-2025-06499 | 2025-02-07 05:00 | Death | MPRI | 2025-03-06 05:00 | LWS | SPRINT QUATTRO SECURE S MRI SURESCAN | Defibrillation/Stimulation Problem; Under-Sensing | Ventricular Fibrillation | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED TO THE EMERGENCY DEPARTMENT AFTER RECEIVING HIGH-VOLTAGE THERAPY. AN INTERROGATION SHOWED THE RIGHT VENTRICULAR (RV) LEAD EXHIBITED OCCASIONAL UNDERSENSING, CAUSING A POSSIBLE DELAY TO THERAPY. ADDITIONAL INFORMATION RECEIVED INDICATED THE PATIENT EXPIRED LESS THAN FIVE DAYS LATER. A CAUSE OF DEATH WAS REQUESTED AND WAS NOT PROVIDED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21536898&pc=NPT | 2025587-2025-01684 | 2025-03-03 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-06 05:00 | NPT | EVOLUT FX DCS | Difficult or Delayed Positioning; Patient Device Interaction Problem; Migration | Cardiac Arrest; Obstruction/Occlusion | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE IMPLANT PROCEDURE, RIGHT FEMORAL ACCESS, AN 18 FRENCH INTRODUCER (SENTRANT), AND NON-MEDTRONIC CLOSURE DEVICES (PROGLIDE) WERE USED. THE PRE-SHAPED GUIDEWIRE (CONFIDA) WAS ADVANCED INTO THE LEFT VENTRICLE, AND ARTERIAL PRESSURE CHANGES (HYPOTENSION) WERE OBSERVED. AN URGENT CARDIOLOGY CONSULTATION AND CONTROL ECHOCARDIOGRAM WERE PERFORMED, RULING OUT PERICARDIAL EFFUSION. CONTRAST INJECTIONS AT THE FEMORAL ACCESS SITES AND ABDOMINAL AORTA TO ASSESS FOR BLEEDING OR DISSECTIONS WERE NEGATIVE, AND AN AORTOGRAM SHOWED NO EVIDENCE OF BLEEDING OR ABNORMALITIES. DUE TO THE PATIENT'S DETERIORATING CONDITION, CARDIAC ARREST OCCURRED. RESUSCITATION MANEUVERS WERE INITIATED. ONCE STABILIZED, A PRE-IMPLANT BALLOON DILATION WAS PERFORMED WITH A 18X40 MM SEMI-COMPLIANT BALLOON. URGENT VALVE DEPLOYMENT PROCEEDED UNDER FLUOROSCOPIC GUIDANCE BY ADVANCING THE DELIVERY CATHETER SYSTEM (DCS) WITH THE PORT POSITIONED AT 3 O¿CLOCK. DEPLOYMENT OF THE VALVE WAS INITIATED IN THE IMPLANTATION PROJECTION. THE POSITIONING WAS ADJUSTED, NOTING THAT THE THREE CUSPS WERE VISIBLE UNDER CONTRAST INJECTION. UPON REACHING THE THIRD NODE, PACING WAS INITIATED VIA THE GUIDEWIRE, AND THE VALVE WAS OPENED TO 80%. THE PACEMAKER WAS THEN DISCONNECTED, AND CONTRAST INJECTIONS CONFIRMED VALVE POSITIONING, WITH A 2 MM DEPTH IN BOTH THE NON-CORONARY AND LEFT CORONARY SINUSES. HOWEVER, THE PATIENT BECAME HEMODYNAMICALLY UNSTABLE, LEADING THE SPECIALIST TO PROCEED WITH FULL VALVE DEPLOYMENT. THE PATIENT SUBSEQUENTLY EXPERIENCED ANOTHER CARDIAC ARREST, REQUIRING FURTHER RESUSCITATION EFFORTS. CARDIOPULMONARY RESUSCITATION WAS ADMINISTERED FOR 37 MINUTES. A CONTROL INJECTION REVEALED THAT THE VALVE HAD MIGRATED TOWARD THE NON-CORONARY SIDE, WITH EVIDENCE OF RIGHT CORONARY ARTERY OCCLUSION. AN ATTEMPT WAS MADE TO CANNULATE THE RIGHT CORONARY ARTERY THROUGH THE VALVE, BUT ALL ATTEMPTS WERE UNSUCCESSFUL. THE PATIENT WAS PRONOUNCED DECEASED.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT PER THE PHYSICIAN, THE PATIENT WAS VERY UNSTABLE, AND THE MEDTRONIC PRODUCTS INCLUDING THE VALVE, DCS, AND GUIDEWIRE DID NOT CAUSE OR CONTRIBUTE TO THE PATIENT'S DEATH. Manufacturer Narrative: UPDATED DATA: H2 H3 H6 IMAGE REVIEW: INTRAPROCEDURAL FLUOROSCOPIC IMAGES WERE PROVIDED FOR REVIEW OF THE EVENT. PATIENT¿S EXECUTIVE SUMMARY WAS PROVIDED FOR ANATOMICAL REVIEW. PATIENT¿S ANNULUS PERIMETER MEASURED 63.7 MILLIMETER (MM) WITH A PERIMETER DERIVED DIAMETER OF 21.3MM SUGGESTING A 26MM EVOLUT. SINUSES OF VALSALVA DIAMETERS MEASURED LESS THAN THE RECOMMENDED 27MM FOR THE 26MM EVOLUT. EVEN THOUGH A 23MM EVOLUT WAS SELECTED FOR THE PROCEDURE, THE SINUSES OF VALSALVA DIAMETERS ALSO DID NOT MEET THE MINIMUM RECOMMENDED CRITERIA OF 25MM FOR THE 23MM EVOLUT. IT WAS REPORTED THAT AN 18F INTRODUCER SHEATH WAS USED FOR THE PROCEDURE. BASED ON THE PATIENT¿S SUMMARY, THE ILIOFEMORAL ARTERIES MEASURED GREATER THAN 6MM WHICH IS THE MINIMUM REQUIRED TO ACCOMMODATE AN 18F INTRODUCER SHEATH. FLUOROSCOPIC VALVE LOAD INSPECTION WAS PERFORMED AND CONFIRMED A GOOD LOAD. IT WAS REPORTED THAT AFTER INSERTION OF THE GUIDEWIRE INTO THE LEFT VENTRICLE, THE PATIENT BECAME HYPOTENSIVE. AN ABDOMINAL AORTA ANGIOGRAM REVEALED NO SIGNS OF DISSECTION OR PERFORATION, AND AN AORTOGRAM REVEALED NO ABNORMALITIES OF THE GUIDEWIRE, GOOD CORONARY PERFUSION, AND NO EVIDENCE OF A LEFT VENTRICULAR PERFORATION. IT WAS REPORTED THAT CARDIOPULMONARY RESUSCITATION EFFORTS WERE REQUIRED TO STABILIZE THE PATIENT. ONCE STABILIZED, A PRE BALLOON AORTIC VALVULOPLA STY WAS PERFORMED. THE DELIVERY CATHETER SYSTEM WAS INSERTED AND ADVANCED ACROSS THE AORTIC VALVE AND SUBSEQUENT ANGIOGRAM REVEALS ADEQUATE PERFUSION TO THE RIGHT AND LEFT CORONARY ARTERIES AND NO SIGNS OF VENTRICULAR PERFORATION. THE VALVE WAS DEPLOYED JUST PRIOR TO THE POINT OF NO RECAPTURE AND DEPTH ASSESSMENT WAS PERFORMED IN THE LEFT ANTERIOR OBLIQUE (LAO) VIEW ONLY. HOWEVER, THERE APPEARS TO BE SIGNIFICANT PARALLAX IN THE EVOLUT THEREFORE ACCURATE DEPTH CANNOT BE CONFIRMED. PER MEDTRONIC BEST PRACTICES, DEPTH ASSESSMENT AT THE NON-CORONARY CUSP (NCC) IS PERFORMED IN A CUSP OVERLAP VIEW, AND IF ACCEPTABLE, NEXT STEP IS TO MOVE TO THE 3-CUSP COPLANAR VIEW AND REMOVE PARALLAX FROM THE INFLOW PORTION OF THE VALVE (ROLL LAO NO GREATER THAN 25°) TO VERIFY DEPTH AT THE LEFT CORONARY CUSP (LCC). THE VALVE MAY BE RELEASED ONCE THESE STEPS ARE COMPLETED. THE RECOMMENDED TARGET DEPTH IS 3 MM RELATIVE TO THE VALVE ANNULUS. IF THE IMPLANT DEPTH IS =1 MM OR >5 MM, CONSIDER RECAPTURE. IN THIS CASE, DEPTH IS UNDETERMINED. IT WAS REPORTED THAT THE PATIENT BECAME HEMODYNAMICALLY UNSTABLE, AND THE ANGIOGRAM SHOWED ABSENCE OF PERFUSION TO THE RIGHT CORONARY ARTERY. DUE TO THE PATIENT INSTABILITY, THE VALVE WAS RELEASED, AND SUBSEQUENT ANGIOGRAMS REVEALED THAT THE FINAL DEPTH OF THE VALVE APPEARED TO BE SHALLOW AND CONFIRMED THE RIGHT CORONARY OCCLUSION/OBSTRUCTION. THERE IS EVIDENCE OF A FAILED ATTEMPT TO RE-ACCESS THE RIGHT CORONARY WITH A CATHETER. IT WAS REPORTED THAT THE PATIENT SUBSEQUENTLY DIED. EVIDENCE SUPPORTS THAT RIGHT CORONARY OCCLUSION/OBSTRUCTION MIGHT HAVE CONTRIBUTED TO THE PATIENT¿S DEATH. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: PRODUCT ID: EVOLUTFX-23, SERIAL#: (B)(6), UBD: 16-SEP-2026, UDI#: (B)(4). CONTINUATION OF D10: SECTION D INFORMATION REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER RELEVANT DEVICE(S) ARE: PRODUCT ID L-EVOLUTFX-2329 (LOT: 0012526083); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE N/A; EXPLANT DATE N/A PRODUCT ID GWBC30 (LOT: UNKNOWN); PRODUCT TYPE: 0193-GUIDEWIRE; IMPLANT DATE N/A; EXPLANT DATE N/A PRODUCT ID SENTRANT SHEATH (LOT: UNKNOWN); PRODUCT TYPE: 0199-INTRODUCER; IMPLANT DATE N/A; EXPLANT DATE N/A MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21536906&pc=NRY | 2029214-2025-00572 | 2024-06-26 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-03-06 05:00 | NRY | UNKNOWN SOLITAIRE | Adverse Event Without Identified Device or Use Problem | Hypoxia; Ischemia |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING: 'EXTRACORPOREAL MEMBRANE OXYGENATION STATES BASILAR ARTERY THROMBECTOMY AND LEFT POSTERIOR CEREBRAL ARTERY STENT THROMBECTOMY: A CASE REPORT.' THE OBJECTIVE IS TO REPORT THE FEASIBILITY OF CEREBRAL ANGIOGRAPHY AND THROMBECTOMY IN PATIENTS WITH IMPLANTED ECMO IN AN INTUBATED STATE. THE TIME FRAME OF THIS STUDY WAS: THE CASE REPORT DOES NOT SPECIFY A DETAILED TIME FRAME FOR THE STUDY. HOWEVER, IT MENTIONS THAT THE PATIENT WAS ADMITTED TO THE HOSPITAL AND UNDERWENT PROCEDURES SHORTLY AFTER, INDICATING A TIME FRAME OF DAYS TO WEEKS FOR THE CASE PRESENTED. MULTIPLE MANUFACTURER¿S DEVICES WERE USED IN THE STUDY POPULATION: THE CASE REPORT DOES NOT EXPLICITLY MENTION THE USE OF DEVICES FROM MULTIPLE MANUFACTURERS. IT PRIMARILY DISCUSSES THE USE OF ECMO AND VARIOUS CATHETERS AND GUIDEWIRES WITHOUT SPECIFYING MANUFACTURERS. THE FOLLOWING MEDTRONIC DEVICES WERE USED: 6F INTERMEDIATE CATHETER (NAVIEN), MICROCATHETER (REBAR-18), STENT (SOLITAIRE 4 MM × 20 MM) DEATHS OCCURRED IN THE STUDY POPULATION: THE PATIENT EVENTUALLY DIED OF CIRCULATORY FAILURE. DESPITE THE SURGICAL INTERVENTIONS, THE CARDIAC FUNCTION NEVER RECOVERED, LEADING TO THE PATIENT'S DEATH. THE CAUSES OF DEATH WERE: CIRCULATORY FAILURE DUE TO LACK OF RECOVERY IN CARDIAC FUNCTION. AMONG PATIENT ADVERSE EVENTS INCLUDED: SEVERE ISCHEMIA AND HYPOXIA DUE TO CARDIAC ARREST. NO OBVIOUS COMPLICATIONS OF BLEEDING, INFECTION, AND ISCHEMIA REPORTED DURING THE POSTOPERATIVE PERIOD. THE FAMILY ULTIMATELY DECIDED TO GIVE UP TREATMENT DUE TO THE LACK OF RECOVERY IN CARDIAC FUNCTION. NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. Manufacturer Narrative: G2: CITATION: AUTHORS: DING, L. S., LIANG, H., ZHENG, M., SHEN, M., LI, Z. J., SONG, R. P., CHEN, Q. L.. EXTRACORPOREAL MEMBRANE OXYGENATION STATES BASILAR ARTERY THROMBECTOMY AND LEFT POSTERIOR CEREBRAL ARTERY STENT THROMBECTOMY: A CASE REPORT. . WORLD J CLIN CASES 12(18): 3589-3595 2024. DOI: 10.12998/WJCC.V12.I18.3589. EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21536979&pc=DQX | 2025587-2025-01686 | 2025-03-03 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-06 05:00 | DQX | CONFIDA BRECKER CURVE GUIDEWIRE | Patient Device Interaction Problem; Migration | Cardiac Arrest; Low Blood Pressure/ Hypotension; Obstruction/Occlusion | K132623 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE IMPLANT PROCEDURE, RIGHT FEMORAL ACCESS, AN 18 FRENCH INTRODUCER (SENTRANT), AND NON-MEDTRONIC CLOSURE DEVICES (PROGLIDE) WERE USED. THE PRE-SHAPED GUIDEWIRE (CONFIDA) WAS ADVANCED INTO THE LEFT VENTRICLE, AND ARTERIAL PRESSURE CHANGES (HYPOTENSION) WERE OBSERVED. AN URGENT CARDIOLOGY CONSULTATION AND CONTROL ECHOCARDIOGRAM WERE PERFORMED, RULING OUT PERICARDIAL EFFUSION. CONTRAST INJECTIONS AT THE FEMORAL ACCESS SITES AND ABDOMINAL AORTA TO ASSESS FOR BLEEDING OR DISSECTIONS WERE NEGATIVE, AND AN AORTOGRAM SHOWED NO EVIDENCE OF BLEEDING OR ABNORMALITIES. DUE TO THE PATIENT'S DETERIORATING CONDITION, CARDIAC ARREST OCCURRED. RESUSCITATION MANEUVERS WERE INITIATED. ONCE STABILIZED, A PRE-IMPLANT BALLOON DILATION WAS PERFORMED WITH A 18X40 MM SEMI-COMPLIANT BALLOON. URGENT VALVE DEPLOYMENT PROCEEDED UNDER FLUOROSCOPIC GUIDANCE BY ADVANCING THE DELIVERY CATHETER SYSTEM (DCS) WITH THE PORT POSITIONED AT 3 O¿CLOCK. DEPLOYMENT OF THE VALVE WAS INITIATED IN THE IMPLANTATION PROJECTION. THE POSITIONING WAS ADJUSTED, NOTING THAT THE THREE CUSPS WERE VISIBLE UNDER CONTRAST INJECTION. UPON REACHING THE THIRD NODE, PACING WAS INITIATED VIA THE GUIDEWIRE, AND THE VALVE WAS OPENED TO 80%. THE PACEMAKER WAS THEN DISCONNECTED, AND CONTRAST INJECTIONS CONFIRMED VALVE POSITIONING, WITH A 2 MM DEPTH IN BOTH THE NON-CORONARY AND LEFT CORONARY SINUSES. HOWEVER, THE PATIENT BECAME HEMODYNAMICALLY UNSTABLE, LEADING THE SPECIALIST TO PROCEED WITH FULL VALVE DEPLOYMENT. THE PATIENT SUBSEQUENTLY EXPERIENCED ANOTHER CARDIAC ARREST, REQUIRING FURTHER RESUSCITATION EFFORTS. CARDIOPULMONARY RESUSCITATION WAS ADMINISTERED FOR 37 MINUTES. A CONTROL INJECTION REVEALED THAT THE VALVE HAD MIGRATED TOWARD THE NON-CORONARY SIDE, WITH EVIDENCE OF RIGHT CORONARY ARTERY OCCLUSION. AN ATTEMPT WAS MADE TO CANNULATE THE RIGHT CORONARY ARTERY THROUGH THE VALVE, BUT ALL ATTEMPTS WERE UNSUCCESSFUL. THE PATIENT WAS PRONOUNCED DECEASED.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT PER THE PHYSICIAN, THE PATIENT WAS VERY UNSTABLE, AND THE MEDTRONIC PRODUCTS INCLUDING THE VALVE, DCS, AND GUIDEWIRE DID NOT CAUSE OR CONTRIBUTE TO THE PATIENT'S DEATH. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: H2, H3, H6. IMAGE REVIEW: INTRAPROCEDURAL FLUOROSCOPIC IMAGES WERE PROVIDED FOR REVIEW OF THE EVENT. PATIENT¿S EXECUTIVE SUMMARY WAS PROVIDED FOR ANATOMICAL REVIEW. PATIENT¿S ANNULUS PERIMETER MEASURED 63.7 MILLIMETER (MM) WITH A PERIMETER DERIVED DIAMETER OF 21.3MM SUGGESTING A 26MM EVOLUT. SINUSES OF VALSALVA DIAMETERS MEASURED LESS THAN THE RECOMMENDED 27MM FOR THE 26MM EVOLUT. EVEN THOUGH A 23MM EVOLUT WAS SELECTED FOR THE PROCEDURE, THE SINUSES OF VALSALVA DIAMETERS ALSO DID NOT MEET THE MINIMUM RECOMMENDED CRITERIA OF 25MM FOR THE 23MM EVOLUT. IT WAS REPORTED THAT AN 18F INTRODUCER SHEATH WAS USED FOR THE PROCEDURE. BASED ON THE PATIENT¿S SUMMARY, THE ILIOFEMORAL ARTERIES MEASURED GREATER THAN 6MM WHICH IS THE MINIMUM REQUIRED TO ACCOMMODATE AN 18F INTRODUCER SHEATH. FLUOROSCOPIC VALVE LOAD INSPECTION WAS PERFORMED AND CONFIRMED A GOOD LOAD. IT WAS REPORTED THAT AFTER INSERTION OF THE GUIDEWIRE INTO THE LEFT VENTRICLE, THE PATIENT BECAME HYPOTENSIVE. AN ABDOMINAL AORTA ANGIOGRAM REVEALED NO SIGNS OF DISSECTION OR PERFORATION, AND AN AORTOGRAM REVEALED NO ABNORMALITIES OF THE GUIDEWIRE, GOOD CORONARY PERFUSION, AND NO EVIDENCE OF A LEFT VENTRICULAR PERFORATION. IT WAS REPORTED THAT CARDIOPULMONARY RESUSCITATION EFFORTS WERE REQUIRED TO STABILIZE THE PATIENT. ONCE STABILIZED, A PRE BALLOON AORTIC VALVULOPLASTY WAS PERFORMED. THE DELIVERY CATHETER SYSTEM WAS INSERTED AND ADVANCED ACROSS THE AORTIC VALVE AND SUBSEQUENT ANGIOGRAM REVEALS ADEQUATE PERFUSION TO THE RIGHT AND LEFT CORONARY ARTERIES AND NO SIGNS OF VENTRICULAR PERFORATION. THE VALVE WAS DEPLOYED JUST PRIOR TO THE POINT OF NO RECAPTURE AND DEPTH ASSESSMENT WAS PERFORMED IN THE LEFT ANTERIOR OBLIQUE (LAO) VIEW ONLY. HOWEVER, THERE APPEARS TO BE SIGNIFICANT PARALLAX IN THE EVOLUT THEREFORE ACCURATE DEPTH CANNOT BE CONFIRMED. PER MEDTRONIC BEST PRACTICES, DEPTH ASSESSMENT AT THE NON-CORONARY CUSP (NCC) IS PERFORMED IN A CUSP OVERLAP VIEW, AND IF ACCEPTABLE, NEXT STEP IS TO MOVE TO THE 3-CUSP COPLANAR VIEW AND REMOVE PARALLAX FROM THE INFLOW PORTION OF THE VALVE (ROLL LAO NO GREATER THAN 25°) TO VERIFY DEPTH AT THE LEFT CORONARY CUSP (LCC). THE VALVE MAY BE RELEASED ONCE THESE STEPS ARE COMPLETED. THE RECOMMENDED TARGET DEPTH IS 3 MM RELATIVE TO THE VALVE ANNULUS. IF THE IMPLANT DEPTH IS =1 MM OR >5 MM, CONSIDER RECAPTURE. IN THIS CASE, DEPTH IS UNDETERMINED. IT WAS REPORTED THAT THE PATIENT BECAME HEMODYNAMICALLY UNSTABLE, AND THE ANGIOGRAM SHOWED ABSENCE OF PERFUSION TO THE RIGHT CORONARY ARTERY. DUE TO THE PATIENT INSTABILITY, THE VALVE WAS RELEASED, AND SUBSEQUENT ANGIOGRAMS REVEALED THAT THE FINAL DEPTH OF THE VALVE APPEARED TO BE SHALLOW AND CONFIRMED THE RIGHT CORONARY OCCLUSION/OBSTRUCTION. THERE IS EVIDENCE OF A FAILED ATTEMPT TO RE-ACCESS THE RIGHT CORONARY WITH A C ATHETER. IT WAS REPORTED THAT THE PATIENT SUBSEQUENTLY DIED. EVIDENCE SUPPORTS THAT RIGHT CORONARY OCCLUSION/OBSTRUCTION MIGHT HAVE CONTRIBUTED TO THE PATIENT¿S DEATH. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: SECTION D INFORMATION REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER RELEVANT DEVICE(S) ARE: PRODUCT ID EVOLUTFX-23 (SERIAL: (B)(6)); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE (B)(6) 2025; EXPLANT DATE N/A PRODUCT ID D-EVOLUTFX-2329 (LOT: 0012483145); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE N/A; EXPLANT DATE N/A PRODUCT ID L-EVOLUTFX-2329 (LOT: 0012526083); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE N/A; EXPLANT DATE N/A PRODUCT ID SENTRANT SHEATH (LOT: UNKNOWN); PRODUCT TYPE: 0199-INTRODUCER; IMPLANT DATE N/A; EXPLANT DATE N/A MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21537136&pc=DSQ | 2916596-2025-01217 | 2025-02-06 05:00 | Death | THORATEC CORPORATION | 2025-03-06 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Renal Failure; Fluid Discharge; Hemorrhagic Stroke; Skin Infection; Swelling/ Edema | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2025, DUE TO A HEMORRHAGIC STROKE AND RENAL FAILURE. NO OBVIOUS NEUROLOGICAL DEFICITS WERE NOTED UPON ADMISSION, THOUGHT THE PATIENT WAS BLEEDING FROM THEIR MOUTH, AND SO A HEAD COMPUTE TOMOGRAPHY (CT) SCAN WAS ORDERED, AND FOUND SMALL AREAS OF ACUTE SUBARACHNOID HEMORRHAGE. THE PATIENT WAS NOTED AS HAVING RENAL FAILURE KNOWN SINCE 2023 (MANUFACTURER REPORT 2916596-2023-02122), AND CHRONIC DRIVELINE INFECTIONS, WITH THE ONSET BEING BLOOD CULTURES POSITIVE METHICILLIN-RESISTANT STAPHYLOCOCCUS AUREUS (MRSA) ON (B)(6) 2021. (MANUFACTURER REPORT 2916596-2025-01357), AS WELL AS ANOTHER POSITIVE MRSA CULTURE ON (B)(6) 2023. (MANUFACTURER REPORT 2916596-2023-02125) AND PSEUDOMONAS ON (B)(6) 2025.THE PATIENT HAD BEEN ON ANTIBIOTICS, CEFAZOLIN AND BACTRIM (TRIMETHOPRIM/SULFAMETHOXAZOLE). IT WAS NOTED THE PATIENT HAD DRAINAGE, REDNESS, AND SWELLING, AS WELL AS ELEVATED ERYTHROCYTE SEDIMENTATION RATE (ESR) AND C-REACTIVE PROTEIN (CRP), DUE TO THE PSEUDOMONAS INFECTION, WHICH DID NOT RESOLVE. THE PATIENT WAS TRANSFERRED TO COMFORT CARE, AND LEFT-VENTRICULAR ASSIST DEVICE (LVAD) SUPPORT WAS TURNED OFF ON (B)(6) 2025. THE DEATH WAS NOT CONSIDERED DEVICE RELATED, AND THE DEVICE WAS NOTED TO FUNCTION AS EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A SPECIFIC CAUSE FOR THE PATIENT¿S INFECTION COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. ADDITIONALLY, A DIRECT CORRELATION BETWEEN HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS ADVERSE EVENTS, INCLUDING RENAL DYSFUNCTION, INFECTION, BLEEDING, AND STROKE, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21537253&pc=NPT | 2025587-2025-01688 | 2024-09-14 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-06 05:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Non specific EKG/ECG Changes; Heart Failure/Congestive Heart Failure | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 8 DAYS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT WAS HOSPITALIZED DUE TO CONGESTIVE HEART FAILURE. THREE DAYS LATER, A PERMANENT PACEMAKER WAS IMPLANTED DUE TO A SEVERE CONDUCTION DISTURBANCE. APPROXIMATELY 2 MONTHS AND 24 DAYS LATER, THE PATIENT PASSED AWAY DUE TO WORSENING HEART FAILURE. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21537283&pc=NIM | 2032493-2025-00202 | 2024-07-29 04:00 | Death | MICROVENTION, INC. | 2025-03-06 05:00 | NIM | CASPER X | Adverse Event Without Identified Device or Use Problem | Obstruction/Occlusion; Cerebral Edema |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PROCEDURE WAS PERFORMED IN EMERGENCY, UNDER GENERAL ANESTHESIA. ANGIOGRAPHY REVEALED OCCLUSION OF THE RIGHT ICA AT THE OSTIUM. THE STENOSIS WAS CROSSED WITH A MICROGUIDEWIRE, PREDILATION WAS PERFORMED WITH A BALLOON, TIROFIBAN WAS ADMINISTERED, AND THE CASPER X STENT WAS IMPLANTED. IT WAS REPORTED THAT ¿THE CALIBER RECOVERY¿ WAS DEEMED ¿NOT SUFFICIENT ¿ BY THE TREATING PHYSICIAN, SO AN INTRA-STENT ANGIOPLASTY WAS PERFORMED WITH A BALLOON, OBTAINING ¿GOOD RECOVERY OF THE VESSEL LUMEN¿. A VASOSPASM REPORTEDLY OCCURRED AT THE CERVICAL TRACT OF THE ICA AND 1 MG OF NIMODIPINE WAS ADMINISTERED. AT THE FINAL CHECK, THE INTRACRANIAL CIRCULATION WAS REPORTED TO APPEAR ¿REGULARLY REPRESENTED¿. MAINTENANCE THERAPY WAS CONTINUED WITH TIROFIBAN. IT WAS REPORTED THAT THE STENT OCCLUDED 40 HOURS AFTER IMPLANTATION. THE PATIENT WAS SAID TO HAVE EXPERIENCED SWELLING IN THE BRAIN ON "LARGE LESION IN RIGHT ICA ISCHEMIC LESION¿ AND WAS STATED TO HAVE PASSED AWAY 28 DAYS AFTER THE STENT OCCLUSION. Manufacturer Narrative: THE DEVICE WAS IMPLANTED IN THE PATIENT AND IS NOT AVAILABLE FOR EVALUATION. MEDICAL NOTES AND IMAGING WERE NOT PROVIDED. THE EVENT AS DESCRIBED COULD NOT BE VERIFIED. IF MORE INFORMATION IS RECEIVED AT A LATER DATE, AN INVESTIGATION WILL BE PERFORMED AND A SUPPLEMENTAL MDR WILL BE SUBMITTED. VISUAL ANALYSIS: A VISUAL INSPECTION OF THE DEVICE CAPTURED IN THIS FILE COULD NOT BE PERFORMED AS A PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION, NOR WERE ANY IMAGES OF THE DEVICE PROVIDED IN PLACE OF A DEVICE RETURN. PROCEDURE AND MEDICAL IMAGING WERE NOT PROVIDED FOR THIS INVESTIGATION. INVESTIGATION FINDINGS: WITHOUT THE RETURN AND PHYSICAL EVALUATION OF THE DEVICE, THE INVESTIGATION CANNOT FURTHER EXAMINE IF A CONDITION EXISTED THAT WOULD HAVE CONTRIBUTED TO THE REPORTED EVENT. WITHOUT IMAGING, THE INVESTIGATION CANNOT FURTHER VERIFY IF THE EVENT OCCURRED AS DESCRIBED, NOR COULD THE INVESTIGATION FURTHER EVALUATE THE CAUSE OF THE REPORTED EVENT. BASED ON A REVIEW OF THE DEVICE¿S RISK DOCUMENTATION, THE REPORTED EVENT DID NOT INDICATE THERE WERE THE PRESENCE OF ANY POTENTIAL OR NEW MANUFACTURING, DESIGN, QUALITY, OR OTHER SYSTEMIC ISSUES, OR NON-CONFORMANCES. THE COMPLAINT CODE IS MONITORED THROUGH THE TRENDING PROCESS; CORRECTIVE ACTION IS DETERMINED, AS NEEDED, THROUGH THIS PROCESS. INVESTIGATIONS OF HISTORIC COMPLAINT FILES WITH SIMILAR COMPLAINT CATEGORY CODING ARE RECORDED IN THE COMPLAINT HANDLING SYSTEM; WITHOUT THE ABILITY TO PERFORM AN ANALYSIS OF THE DEVICE, THIS INVESTIGATION CANNOT FURTHER EXAMINE ANY POTENTIAL ROOT CAUSES. BATCH REVIEW: A SEARCH FOR NON-CONFORMANCES ASSOCIATED WITH THIS PART/LOT NUMBER COMBINATION DID NOT REVEAL ANY PRODUCTION-RELATED ISSUES RELEVANT TO THE COMPLAINT THAT OCCURRED DURING MANUFACTURING OF THE DEVICE. COMPLAINT SYSTEM REVIEW: THERE ARE NO SIMILAR COMPLAINTS BASED ON THE COMPLAINT CATEGORY REGARDING THIS BATCH NUMBER FROM THE LAST TWO YEARS RECORDED IN THE COMPLAINT SYSTEM AT THE TIME OF THIS INVESTIGATION. IFU REVIEW (ADDITIONAL INFORMATION CAN BE FOUND IN THE IFU): ¿POTENTIAL COMPLICATIONS POSSIBLE COMPLICATIONS INCLUDE BUT ARE NOT LIMITED TO THE FOLLOWING: ARRHYTHMIA, ANEURYSM AND PSEUDOANEURYSM FORMATION, ABRUPT VESSEL CLOSURE, ALLERGIC REACTIONS (INCLUDING TO ANTIPLATELET AGENTS, CONTRAST MEDIUM OR STENT MATERIALS), ARTERIOVENOUS FISTULA, BLEEDING FROM ANTICOAGULATION/ANTIPLATELET MEDICATION, BRADYCARDIA AND HYPOTENSION, CAROTID ARTERY SPASM, CORONARY ISCHEMIA, DEATH, DISSEMINATED INTRAVASCULAR COAGULATION, EMBOLI (AIR, TISSUE, PLAQUE, THROMBUS, DEVICE OR OTHER), EMERGENCY ARTERY BYPASS GRAFT SURGERY, HEMATOMA, HEMORRHAGIC OR EMBOLIC STROKE/TIA, INFECTION AND/OR PAIN AT INSERTION SITE/SEPSIS, INTIMAL TEAR/DISSECTION, MYOCARDIAL INFARCTION (MI), NEW OR WORSE ENCEPHALOPATHY, RENAL FAILURE/INSUFFICIENCY, RESPIRATORY ARREST, STENT MISPLACEMENT, TISSUE NECROSIS, VESSEL INJURY/DISSECTION/PERFORATION/RUPTURE/TRAUMA, VESSEL OCCLUSION OR THROMBOSIS, VESSEL SPASM OR RECOIL. WARNINGS: SHOULD UNUSUAL RESISTANCE BE FELT AT ANY TIME DURING ACCESS OR REMOVAL, THE SHEATH INTRODUCER/GUIDE CATHETER AND CASPER X SYSTEM SHOULD BE REMOVED AS A SINGLE UNIT. APPLYING EXCESSIVE FORCE DURING DELIVERY OR RETRIEVAL OF THE CASPER X SYSTEM CAN POTENTIALLY RESULT IN LOSS OR DAMAGE TO THE DEVICE AND DELIVERY COMPONENTS. IF AN EPD DEVICE IS USED, ALLOW FOR AND MAINTAIN ADEQUATE DISTANCE BETWEEN THE FILTER, THE STENT DELIVERY SYSTEM OR DEPLOYED STENT TO AVOID POTENTIAL ENTANGLEMENT. DO NOT REPOSITION THE CASPER X SYSTEM WITHOUT FULLY RETRIEVING THE DEVICE. THE CASPER X STENT MUST BE RETRIEVED INTO THE DELIVERY SYSTEM AND REDEPLOYED AT THE DESIRED TARGET LOCATION OR REMOVED COMPLETELY FROM THE PATIENT. THE USE OF A GUIDING SHEATH OR GUIDING CATHETER WITH A FIXED HEMOSTASIS VALVE MAY CAUSE THE DISTAL TIP OF THE DELIVERY CATHETER TO DETACH AT THE HEMOSTASIS VALVE UPON REMOVAL IF THE VALVE IS NOT ADEQUATELY OPENED. PRECAUTIONS: EXERCISE CAUTION WHEN CROSSING THE DEPLOYED/DETACHED CASPER X SYSTEM WITH ADJUNCTIVE DEVICES SUCH AS GUIDEWIRES, CATHETERS OR BALLOON CATHETERS TO AVOID DISRUPTING THE DEVICE GEOMETRY AND DEVICE PLACEMENT. DIRECTIONS FOR USE: INTRODUCTION OF THE STENT DELIVERY SYSTEM: ACCESS THE TREATMENT SITE USING THE APPROPRIATE ACCESSORY EQUIPMENT AND INSERT AN EPD USABLE AS A GUIDEWIRE. IF AN EPD IS NOT PREFERRED, USE A SUFFICIENTLY LONG .014¿ GUIDEWIRE AND PLACE IT ACROSS THE LESION. WARNING: ALWAYS USE A GUIDEWIRE OR EPD WHEN ADVANCING OR WITHDRAWING THE DELIVERY SYSTEM. IF NECESSARY, PERFORM A PRE-DILATATION AT THE LESION TARGET SITE USING A STANDARD PTA TECHNIQUE USING THE RECOMMENDED GUIDE SHEATH OR GUIDING CATHETER ON THE OUTER BOX LABEL. THE GUIDE WIRE OR EPD DEVICE SHOULD BE ADVANCED PAST THE TARGET LESION SITE. THEN THE DELIVERY SYSTEM SHOULD BE ADVANCED OVER THE GUIDE WIRE OR EPD TO THE LESION SITE USING FLUOROSCOPY. CAUTION: DO NOT ADVANCE THE DELIVERY SYSTEM AGAINST SIGNIFICANT RESISTANCE. THE CAUSE OF THE RESISTANCE SHOULD BE DETERMINED BY FLUOROSCOPY AND REMEDIAL ACTION TAKEN. WITHDRAW THE SYSTEM AND USE A NEW ONE. SLACK REMOVAL: ENSURE THAT THE DELIVERY SYSTEM CATHETER OUTSIDE THE PATIENT REMAINS FLAT AND STRAIGHT. WARNING: SLACK IN THE DELIVERY SYSTEM CATHETER EITHER OUTSIDE OR WITHIN THE PATIENT MAY RESULT IN MISPLACEMENT OF THE STENT BEYOND THE LESION. STENT DEPLOYMENT: NOTE: THE CASPER X STENT SHOULD BE SIZED TO THE VESSEL ACCORDING TO THE REFERENCE VESSEL DIAMETER RANGE LISTED ON THE PRODUCT LABEL. THE CASPER X STENT FORESHORTENS AS IT EXPANDS TO THE DIAMETER OF THE NATIVE VESSEL. THE STENT FORESHORTENING SHOULD BE TAKEN INTO ACCOUNT WHEN SIZING AND DEPLOYING THE CASPER X CAROTID STENT SYSTEM. IF THE CASPER X STENT POSITIONING IS NOT SATISFACTORY ACROSS THE TARGET LESION, THE STENT MAY BE RECAPTURED AND REPOSITIONED IF IT IS NOT FULLY DEPLOYED. THE IMPLANT MAY BE RECAPTURED UP TO 50% OF ITS DEPLOYMENT FROM THE CATHETER DELIVERY SYSTEM. RECAPTURING OF STENT CAN BE ACCOMPLISHED BY REMOVING THE SLACK IN THE DELIVERY CATHETER SHAFT AND THEN MAINTAINING THE INNER CATHETER (HANDLE) POSITION AND ADVANCING OR PUSHING FORWARD THE OUTER CATHETER. AFTER RECAPTURING THE STENT INTO THE DELIVERY SYSTEM THEN YOU CAN REDEPLOY THE STENT ACROSS THE LESION, TAKING INTO CONSIDERATION THE FORESHORTENING OF THE STENT. ENSURE THE RHV IS OPEN. THE STENT IS DEPLOYED BY OUTER CATHETER RETRACTION. DEPLOYMENT IS COMPLETE BY MAINTAINING INNER CATHETER POSITION (HOLDING HANDLE) WHILE RETRACTING THE OUTER CATHETER AND ALLOWING THE STENT TO EXPAND. WITHDRAWING STENT DELIVERY SYSTEM: AFTER THE STENT HAS BEEN COMPLETELY DEPLOYED, USING FLUOROSCOPY, WITHDRAW THE DELIVERY SYSTEM CAREFULLY, LEAVING THE GUIDEWIRE OR EPD IN PLACE. PERFORM STANDARD POST-PROCEDURAL ANGIOGRAPHY. POST-DEPLOYMENT STENT DILATATION: IF THE STENT IS NOT COMPLETELY EXPANDED WITHIN THE LENGTH OF THE LESION, POST-DILATION WITH THE BALLOON CATHETER INSIDE THE STENT MAY BE DONE WITH STANDARD PTA TECHNIQUE. THE INFLATED NOMINAL DIAMETER OF THE PTA BALLOON USED FOR POST DILATION SHOULD BE APPROXIMATELY THE SAME AS THE DIAMETER OF THE REFERENCED NATIVE VESSEL.¿ THE PRODUCT REPORTED IS AN EXPORTED DEVICE NOT CLEARED OR APPROVED FOR MARKETING IN THE US. THE ADVERSE EVENT IS BEING REPORTED BASED ON THIS PRODUCT MEETING SIMILAR PRODUCT CRITERIA (SIMILAR TO ROADSAVER/CASPER CAROTID STENT DEVICE ¿ PMA # P210030). MICROVENTION IS SUBMITTING THIS REPORT TO COMPLY WITH FDA REPORTING REGULATIONS UNDER 21 CFR PARTS 4 AND 803. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MICROVENTION, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MICROVENTION HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION AND HAS PROVIDED AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MICROVENTION, OR ITS EMPLOYEES THAT THE DEVICE, MICROVENTION, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. A GOOD FAITH EFFORT WILL BE MADE TO OBTAIN THE APPLICABLE INFORMATION RELEVANT TO THE REPORT. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21537876&pc=NIK | 2017865-2025-16869 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-03-06 05:00 | NIK | GALLANT VR | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P030054 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBER:2017865-2025-16869. IT WAS REPORTED THAT THE PATIENT HAS DECEASED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE RELATED. THE CAUSE OF DEATH WAS HEART FAILURE, DIABETES AND FEMUR FRACTURE. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: THE DEVICE WAS RETURNED DUE TO PATIENT DEATH. INTERROGATION OF THE DEVICE REVEALED IT WAS ABOVE ELECTIVE REPLACEMENT INDICATOR (ERI) WHEN RECEIVED. THE DEVICE¿S SENSING, PACING, LEAD IMPEDANCE, HIGH VOLTAGE CHARGING, AND HIGH VOLTAGE SHOCK IMPEDANCE WERE TESTED, AND THE DEVICE¿S FUNCTION WAS NORMAL. NO ANOMALIES WERE FOUND. Manufacturer Narrative: D6A SHOULD BE 10 NOV 2022 RATHER THAN BLANK.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21537885&pc=NVY | 2017865-2025-16870 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-03-06 05:00 | NVY | OPTISURE ACTIVE FIXATION, DF-4 CONNECTOR | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P950022 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBER:2017865-2025-16869. IT WAS REPORTED THAT THE PATIENT HAS DECEASED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE RELATED. THE CAUSE OF DEATH WAS HEART FAILURE, DIABETES AND FEMUR FRACTURE. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: THE LEAD WAS RETURNED FOR PATIENT DEATH. AS RECEIVED A COMPLETE LEAD WAS RETURNED IN ONE PIECE. VISUAL ANALYSIS DID NOT FIND ANY ANOMALIES. ELECTRICAL TESTING DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS. NO ANOMALIES FOUND,
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21537897&pc=DSQ | 2916596-2025-01374 | 2023-03-01 05:00 | Death | THORATEC CORPORATION | 2025-03-06 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Mechanical Problem; Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Unspecified Infection; Sepsis; Multiple Organ Failure; Thrombosis/Thrombus; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THE RESEARCH ARTICLE "FIVE-YEAR OUTCOMES IN PATIENTS IMPLANTED WITH A HEARTMATE 3 LEFT VENTRICULAR ASSIST DEVICE AT A HIGH-VOLUME CENTER" THAT HEARTMATE 3 (HM3) PATIENTS EXPERIENCED STROKE, GASTROINTESTINAL BLEEDING, DRIVELINE INFECTION, DE NOVO RENAL REPLACEMENT THERAPY, TRANSPLANT, AND DEATH. THE STUDY EXAMINED THE LONG-TERM OUTCOMES OF 280 CONSECUTIVE HM3 PATIENTS AT A SINGLE, HIGH-VOLUME LEFT VENTRICULAR ASSIST DEVICE (LVAD) CENTER, INCLUDING: 5-YEAR SURVIVAL, 5-YEAR FREEDOM FROM STROKE, GASTROINTESTINAL BLEEDING (GIB), AND DRIVELINE INFECTION (DLI), AND INCIDENCE OF DE NOVO RENAL REPLACEMENT THERAPY (RRT) AT DISCHARGE FOLLOWING DEVICE IMPLANTATION, CONSIDERING THEIR IMPACT ON PATIENTS¿ PROGNOSIS AND QUALITY OF LIFE. PATIENTS IMPLANTED WITH THEIR FIRST HM3 LVAD BETWEEN 30NOV2014 AND 01JUL2022 WERE STUDIED. FOLLOW-UP CLINICAL DATA INFORMATION WAS RECORDED UP TO 5 YEARS AND UNTIL 01MAR2023. TWO HUNDRED-EIGHTY HM3 PATIENTS WERE STUDIED. THE MEDIAN AGE WITH INTERQUARTILE RANGE (IQR) WAS 61.0 (52.9¿68.9) YEARS; MAJORITY WERE MALE (84.0%), WHITE (61.4%) WITH NON-ISCHEMIC ETIOLOGY (59.3%); 62.9% WERE INTERMACS PROFILE 1 OR 2 AND 85.0% RECEIVED THE HM3 AS DESTINATION THERAPY. MEDIAN FOLLOW-UP WAS 694 (379¿1,235) DAYS. AT THE END OF FOLLOW-UP, 143 (51.1%) PATIENTS REMAINED ON LVAD SUPPORT, 76 (27.1%) HAD UNDERGONE HEART TRANSPLANT, 2 (0.7%) WERE EXPLANTED, 10 (3.6%) TRANSFERRED CARE, AND 49 (17.5%) DIED. CAUSES OF DEATH WERE AS FOLLOWS: 11 DEATHS DUE TO INFECTION/SEPSIS (22.4%), 9 DUE TO RIGHT VENTRICULAR FAILURE, 5 DUE TO MULTISYSTEM ORGAN FAILURE (10.2%), 5 DUE TO FAILURE TO THRIVE (10.2%), 8 DUE TO STROKE (16.3%), 2 DUE TO DEVICE/OUTFLOW GRAFT THROMBOSIS (4.1%), AND 9 DUE TO OTHER CAUSES (18.4%). THERE WERE 20 STROKES IN 20 (7.1%) PATIENTS THAT OCCURRED AT A MEDIAN OF 9 (IQR: 3¿86) DAYS FOLLOWING LVAD. THERE WERE 92 GIB EVENTS IN 66 (23.6%) PATIENTS. THE FIRST GIB EVENT OCCURRED AT A MEDIAN OF 171 (IQR: 29¿515) DAYS FOLLOWING LVAD. THERE WERE 56 DLIS IN 56 (20.0%) PATIENTS AT A MEDIAN OF 459 (IQR: 257¿844) DAYS FOLLOWING LVAD. AT INDEX DISCHARGE, DE NOVO RRT WAS REQUIRED IN TWO (0.7%) PATIENTS. Manufacturer Narrative: SECTIONS A AND D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. SECTION B3: DATE OF EVENT HAS BEEN ENTERED AS 01MAR2023 AS DATA WAS COLLECTED UNTIL 01MAR2023. AUTHOR INFORMATION: MONDELLINI, G. M., VINOGRADSKY, A. V., KIRSCHNER, M., PINSINO, A., LADANYI, A., KURLANSKY, P. A., SUN, J. Y., WANG, C., HASTIE, J. M., CLERKIN, K. J., NAKA, Y., KAKU, Y., SAYER, G. T., URIEL, N., YUZEFPOLSKAYA, M., TAKEDA, K., & COLOMBO, P. C. (2024). FIVE-YEAR OUTCOMES IN PATIENTS IMPLANTED WITH A HEARTMATE 3 LEFT VENTRICULAR ASSIST DEVICE AT A HIGH-VOLUME CENTER. ASAIO JOURNAL. HTTPS://DOI.ORG/10.1097/MAT.0000000000002286. COLUMBIA UNIVERSITY IRVING MEDICAL CENTER, NEW YORK, NEW YORK. MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED EVENTS COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. ADDITIONALLY, THE REPORTS OF PUMP THROMBOSIS COULD NOT BE CONFIRMED BASED ON THE PROVIDED INFORMATION. A SPECIFIC CAUSE FOR THESE EVENTS COULD NOT BE CONCLUSIVELY DETERMINED. THE SERIAL NUMBERS OR OTHER IDENTIFYING INFORMATION OF THE PRODUCTS WERE NOT REPORTED AND WERE NOT ABLE TO BE DETERMINED DURING THE INVESTIGATION. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU), REV. C, AND THE HEARTMATE 3 PATIENT HANDBOOK, REV. D, ARE CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, LISTS INFECTION (DRIVELINE, LOCALIZED, AND PUMP POCKET OR PSEUDO POCKET INFECTION), SEPSIS, RIGHT HEART FAILURE, STROKE, MULTIPLE ORGAN FAILURES/DYSFUNCTIONS, PUMP THROMBOSIS, AND DEATH AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 5 OF THE IFU, "SURGICAL PROCEDURES", INSTRUCTS TO INSPECT THE VENTRICULAR CHAMBER FOR MURAL THROMBI AND CROSSING TRABECULAE FOLLOWING REMOVAL OF THE CORE AND TO ADDRESS ONE OR BOTH, AS NEEDED. THIS SECTION ALSO INSTRUCTS TO INSPECT THE VENTRICLE AND REMOVE ANY PREVIOUSLY FORMED CLOTS THAT MAY CAUSE EMBOLISM OR ANY TRABECULAE THAT MAY IMPEDE FLOW. SECTION 6, ¿PATIENT CARE AND MANAGEMENT¿ PROVIDES INFORMATION REGARDING ANTICOAGULATION, INCLUDING RECOMMENDED ANTICOAGULATION THERAPY AND INR (INTERNATIONAL NORMALIZED RATIO) RANGE. THIS SECTION ALSO LISTS INFECTION AND THROMBOEMBOLISM AS POTENTIAL LATE POSTIMPLANT COMPLICATIONS. FURTHERMORE, SEVERAL SECTIONS OF THE IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21538156&pc=NPT | 2025587-2025-01700 | 2024-05-26 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-06 05:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Bradycardia; Unspecified Infection; Heart Failure/Congestive Heart Failure | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT TWO DAYS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, FEVER AND IMPAIRED CONSCIOUSNESS WERE OBSERVED. A BLOOD CULTURE TEST REVEALED. METHICILLIN-RESISTANT STAPHYLOCOCCUS AUREUS. SUBSEQUENTLY, ANTIBIOTIC TREATMENT WAS INITIATED. REPEAT BLOOD CULTURE APPROXIMATELY 24 DAYS LATER WAS NEGATIVE, HOWEVER DISTURBANCE OF CONSCIOUSNESS HAD WORSENED. THE PATIENT'S GENERAL CONDITION WAS NOTED TO HAVE WORSENED GRADUALLY WITH THE PROLONGED HOSPITALIZATION. APPROXIMATELY 1 MONTH AND 12 DAYS AFTER THE NEGATIVE BLOOD CULTURE, THE PATIENT DEVELOPED BRADYCARDIA AND PASSED AWAY. PER THE PHYSICIAN, THE CAUSE OF DEATH WAS CHRONIC HEART FAILURE.Event Description: IT WAS REPORTED THAT TWO DAYS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, FEVER AND IMPAIRED CONSCIOUSNESS WERE OBSERVED. A BLOOD CULTURE TEST REVEALED METHICILLIN-RESISTANT STAPHYLOCOCCUS AUREUS. SUBSEQUENTLY, ANTIBIOTIC TREATMENT WAS INITIATED. REPEAT BLOOD CULTURE APPROXIMATELY 24 DAYS LATER WAS NEGATIVE, HOWEVER DISTURBANCE OF CONSCIOUSNESS HAD WORSENED. THE PATIENT'S GENERAL CONDITION WAS NOTED TO HAVE WORSENED GRADUALLY WITH THE PROLONGED HOSPITALIZATION. APPROXIMATELY 1 MONTH AND 12 DAYS AFTER THE NEGATIVE BLOOD CULTURE, THE PATIENT DEVELOPED BRADYCARDIA AND PASSED AWAY. PER THE PHYSICIAN, THE CAUSE OF DEATH WAS CHRONIC HEART FAILURE. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTED DATA: B5. UPDATED DATA: D4. H4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21538170&pc=NPT | 2135147-2025-01134 | 2025-02-10 05:00 | Death | ABBOTT MEDICAL | 2025-03-06 05:00 | NPT | NAVITOR VISION | Perivalvular Leak | Mitral Valve Insufficiency/ Regurgitation | P190023 |  | 1 | Event Description: IT WAS REPORTED THAT ON 10 FEBRUARY 2025, A 25MM NAVITOR VALVE WAS SELECTED FOR IMPLANT USING A SMALL FLEXNAV DELIVERY SYSTEM. THE PATIENT HAS A PERTINENT PAST MEDICAL HISTORY OF MILD MITRAL REGURGITATION. PRIOR TO PROCEDURE, THE PATIENT'S CONDITION INCLUDED SEVERE AORTIC STENOSIS AND MODERATE MITRAL REGURGITATION. PRE-DILATATION WAS PERFORMED WITH A 20MM BALTON; THE BALLOON DID NOT EXCEED THE MIN ANNULUS DIAMETER. DURING IMPLANT, THE VALVE WAS DEPLOYED IN OPTIMAL POSITION USING THE REFINE TECHNIQUE. NON-UNIFORM EXPANSION WAS CHECKED. THEN, THE PATIENT WENT INTO SEVERE MITRAL REGURGITATION (MR) WITH SIGNIFICANT HEART FAILURE SYMPTOMS. THE PATIENT WAS STABILIZED AND THE NAVITOR WAS IMPLANTED AT AN IMPLANT DEPTH OF 3-4MM. POST-DILATATION WAS ALSO PERFORMED WITH A 20MM BALTON AND THE RESULT WAS TRACE PARAVALVULAR LEAK. OVERNIGHT WHILST IN THE INTENSIVE CARE UNIT (ICU), SEVERE MR RE-OCCURRED AND THE PATIENT RECEIVED MEDICATION THERAPIES; PALLIATIVE CARE WAS ULTIMATELY DECIDED. THE USER DOES NOT ALLEGE THAT THE NAVITOR VALVE CAUSED OR CONTRIBUTED TO THE SEVERE MR. ON 11 FEBRUARY 2025, THE PATIENT PASSED AWAY DUE TO A "SUICIDE VENTRICLE." THE IMPLANTER DID NOT CLASSIFY THE PATIENT'S PASSING AS RELATED TO THE DEVICE. Manufacturer Narrative: AN EVENT OF PERIVALVULAR LEAK WAS REPORTED. A RETURNED DEVICE ASSESSMENT COULD NOT BE PERFORMED AS THE DEVICE REMAINS IMPLANTED AND WAS NOT RETURNED FOR ANALYSIS. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED, AND THE PRODUCT MET ALL SPECIFICATIONS, INCLUDING SPECIFICATION FOR RADIAL FORCE. THE PROVIDED INFORMATION WAS REVIEWED BY MEDICAL AFFAIRS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED PERIVALVULAR LEAK COULD NOT BE CONCLUSIVELY DETERMINED. THE REPORTED PATIENT EFFECTS OF MITRAL REGURGITATION COULD NOT BE DETERMINED. DEATH WAS DUE TO SUICIDE VENTRICLE. THE REPORTED HOSPITALIZATION WAS A RESULT OF CASE-SPECIFIC CIRCUMSTANCES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH REGARDS TO MANUFACTURE, DESIGN, OR LABELING. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21538183&pc=OUT | 2029214-2025-00577 | 2022-11-03 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-03-06 05:00 | OUT | PIPELINE | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING: THE STUDY DISCUSSES THE TREATMENT OF BLOOD BLISTER-LIKE ANEURYSMS (BBA) AND COMPARES THE OUTCOMES OF MICROSURGERY AND ENDOVASCULAR TREATMENT. THE TIME FRAME OF THIS STUDY WAS: THE STUDY COLLECTED DATA FROM PATIENTS ADMITTED FROM JANUARY 2012 TO DECEMBER 2021. THE FOLLOWING MEDTRONIC DEVICES WERE USED: IT WAS REPORTED THAT PIPELINE EMBOLIZATION DEVICES WERE USED FOR FLOW DIVERSION. THERE WERE 95 TOTAL PATIENTS IN THE ENDOVASCULAR GROUP, 7 OF WHICH UNDERWENT FLOW DIVERSION (5 FLOW DIVERSION ONLY, 2 FLOW DIVERSION + COILING). THE MANUFACTURER WAS NOT SPECIFIED FOR OTHER DEVICES USED IN THE STUDY. DEATHS OCCURRED IN THE STUDY POPULATION. THERE WERE 4 DEATHS REPORTED IN THE ENDOVASCULAR GROUP. THE TEXT DOES NOT SPECIFY THE CAUSES OF DEATH OR WHAT TREATMENT THE PATIENT RECEIVED. AMONG PATIENT ADVERSE EVENTS INCLUDED: THE FOLLOWING ADVERSE EVENTS AND OUTCOMES WERE REPORTED IN THE ENDOVASCULAR GROUP. THE TEXT DOES NOT SPECIFY WHAT TREATMENT THE PATI ENTS RECEIVED: 26 CASES OF VASOSPASM, 1 CASE OF INTRAOPERATIVE LACERATION 2 CASES OF POSTOPERATIVE BLEEDING, 9 CASES OF LARGE-AREA BRAIN INFARCTION. 6 CASES OF RELAPSE OR RESIDUAL ANEURYSM, 1 CASE OF TERSON¿S SYNDROME (VISUAL IMPAIRMENT). 7 CASES OF MODERATE DISABILITY WITH NO LIVING INDEPENDENCY. 3 CASES OF SEVERE DISABILITY WITH NO LIVING INDEPENDENCY NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID NV UNK PIPELINE (UNKNOWN); PRODUCT TYPE: ; IMPLANT DATE N/A; EXPLANT DATE N/A G2: CITATION: AUTHORS: WEN, D., WANG, X., CHEN, R., LI, H., ZHENG, J., FU, W., LIU, D., XIE, X., YOU, C., ZHANG, C., <(>&<)> MA, L.. A PROPENSITY SCORE¿MATCHED STUDY ON THE SHORT-TERM OUTCOME OF RUPTURED BLOOD BLISTER-LIKE ANEURYSM TREATED BY MICROSURGERY OR ENDOVASCULAR SURGERY: A SINGLE-CENTER STUDY OF 155 CASES. NEUROSURGICAL REVIEW 45(6):3789-3800 2022. DOI:10.1007/S10143-022-01887-0. EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21538217&pc=NGV | 2135147-2025-01135 | 2025-02-11 05:00 | Death | ABBOTT MEDICAL | 2025-03-06 05:00 | NGV | AMPLATZER AMULET | Migration or Expulsion of Device | Cardiac Arrest | P200049 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 22MM AMPLATZER AMULET WAS SUCCESSFULLY IMPLANTED. THE DEVICE WAS ABLE TO BE SUCCESSFULLY IMPLANTED. THE LANDING ZONE MEASUREMENTS FOUND A MAX DIAMETER OF 17MM AS FOUND BY BOTH 2D-TEE AND ANGIOGRAPHY. IT WAS NOTED THAT CLOSE CRITERIA WERE SATISFIED AS WELL AS A TENSION TEST WAS PERFORMED. DURING THE POST-PROCEDURE ASSESSMENT, A TTE REVEALED THAT THE PROSTHESIS HAD EMBOLIZED INTO THE LEFT FLOW TRACT. ON (B)(6) 2025, PERCUTANEOUS RETRIEVAL WAS ATTEMPTED, BUT WAS UNSUCCESSFUL. THE PATIENT ENDED UP DECOMPENSATING DURING THE ONE THE RECAPTURE ATTEMPTS. IT WAS THEN DECIDED TO PERFORM SURGICAL RETRIEVAL OF THE DEVICE. BUT THE PATIENT PASSED AWAY THE SAME DAY. Manufacturer Narrative: IT WAS REPORTED THAT DURING THE POST-PROCEDURE ASSESSMENT AFTER AN SUCCESSFUL IMPLANT OF AMULET DEVICE, A TTE REVEALED THAT THE PROSTHESIS HAD EMBOLIZED INTO THE LEFT FLOW TRACT. A PERCUTANEOUS RETRIEVAL WAS ATTEMPTED TWO DAYS AFTER IMPLANT, BUT WAS UNSUCCESSFUL. THE PATIENT ENDED UP DECOMPENSATING DURING RECAPTURE ATTEMPTS. IT WAS THEN DECIDED TO PERFORM SURGICAL RETRIEVAL OF THE DEVICE. BUT THE PATIENT EXPIRED THE SAME DAY. INFORMATION FROM FIELD INDICATED THAT THE DEVICE SIZE WAS DETERMINED THROUGH 2D-TEE AND ANGIOGRAPHY. FIELD ALSO INDICATED THAT CLOSE CRITERIA WERE SATISFIED AS WELL AS A TENSION TEST WAS COMPLETED. A RETURNED DEVICE ASSESSMENT COULD NOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED FOR ANALYSIS. REQUESTS WERE MADE FOR ADDITIONAL PROCEDURAL DETAILS SURROUNDING THE CASE (E.G., OPERATIVE NOTES, PROCEDURE IMAGING), HOWEVER THIS INFORMATION WAS NOT SUPPLIED BY THE SITE. BASED ON THE LIMITED INFORMATION AVAILABLE, THE CAUSE OF REPORTED DEVICE EMBOLIZATION, CARDIAC ARREST AND PATIENT DEATH COULD NOT BE CONCLUSIVELY DETERMINED. THE REPORTED SURGICAL INTERVENTION WAS A RESULT OF CASE-SPECIFIC CIRCUMSTANCES AS SURGICAL RETRIEVAL OF THE DEVICE WAS PERFORMED AFTER UNSUCCESSFUL ATTEMPT OF PERCUTANEOUS RETRIEVAL. THERE WERE NO COMPLAINTS ASSOCIATED WITH ANY OTHER DEVICES FROM THE LOT. BASED ON THE INFORMATION RECEIVED, THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO LABELING DESIGN OR MANUFACTURING OF THE DEVICE. Manufacturer Narrative: IT IS UNKNOWN IF THE DEVICE IS RETURNING FOR ANALYSIS. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. THE DATE OF DEATH IS ESTIMATED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21538243&pc=DSQ | 2916596-2025-01503 | 2025-02-23 05:00 | Death | THORATEC CORPORATION | 2025-03-06 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THERE WERE NO DEVICE ISSUES, AND THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. Manufacturer Narrative: A1-A4: PATIENT INFORMATION CANNOT BE PROVIDED DUE TO PERSONAL DATA PRIVACY LEGISLATION/POLICY. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. AN AUTOPSY WAS NOT PERFORMED, AND THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. DUE TO PATIENT PRIVACY LAWS, NO ADDITIONAL INFORMATION WILL BE PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21538504&pc=NPT | 2025587-2025-01704 | 2024-09-29 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-06 05:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Hemorrhage/Bleeding; Heart Failure/Congestive Heart Failure | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 1 MONTH AND 9 DAYS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, A MASSIVE HEMORRHAGE WAS NOTED. THE HEMORRHAGE WAS A 3 OR HIGHER ON THE BLEEDING ACADEMIC RESEARCH CONSORTIUM SCALE. THAT SAME DAY THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS REPORTED TO BE HEART FAILURE.Event Description: IT WAS REPORTED THAT APPROXIMATELY 1 MONTH AND 9 DAYS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, A MASSIVE HEMORRHAGE WAS NOTED. THE HEMORRHAGE WAS A 3 OR HIGHER ON THE BLEEDING ACADEMIC RESEARCH CONSORTIUM SCALE. THAT SAME DAY THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS REPORTED TO BE HEART FAILURE. ADDITIONAL INFORMATION WAS RECEIVED THAT PER THE PHYSICIAN, THE CAUSE OF THE INJURY WAS NOT KNOWN BUT IT WAS NOT ATTRIBUTED TO THE VALVE OR DELIVERY CATHETER SYSTEM (DCS). THE EXACT LOCATION OF THE INJURY WAS NOT KNOWN AND IT WAS UNKNOWN WHETHER TREATMENT WAS PRESCRIBED. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED: B5, D4, H4 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21539069&pc=NIK | 2017865-2025-16808 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-03-06 05:00 | NIK | GALLANT DR | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P030054 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-16809. RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-16810. IT WAS REPORTED THAT THE PATIENT HAS DECEASED. THE CAUSE OF DEATH WAS VENTRICULAR TACHYCARDIA (VT). THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) RIGHT VENTRICULAR (RV) AND RIGHT ATRIAL (RA) LEADS WERE EXPLANTED. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: INTERROGATION OF THE DEVICE REVEALED IT WAS ABOVE ELECTIVE REPLACEMENT INDICATOR (ERI) WHEN RECEIVED. THE DEVICE¿S SENSING, PACING, LEAD IMPEDANCE, HIGH VOLTAGE CHARGING, AND HIGH VOLTAGE SHOCK IMPEDANCE WERE TESTED, AND THE DEVICE¿S FUNCTION WAS NORMAL.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21539164&pc=NVY | 2017865-2025-16809 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-03-06 05:00 | NVY | DURATA STS OPTIM ACTIVE FIXATION, DF-4 CONNECTOR | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P950022 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-16808. RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-16810. IT WAS REPORTED THAT THE PATIENT HAS DECEASED. THE CAUSE OF DEATH WAS VENTRICULAR TACHYCARDIA (VT). THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) RIGHT VENTRICULAR (RV) AND RIGHT ATRIAL (RA) LEADS WERE EXPLANTED. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: THE REPORTED EVENT WAS PATIENT DECEASED. AS RECEIVED, A COMPLETE LEAD WAS RETURNED IN ONE PIECE. ELECTRICAL TESTING WAS NORMAL. VISUAL INSPECTION OF THE LEAD DID NOT FIND ANY ANOMALIES EXCEPT FOR PROCEDURAL DAMAGE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21539213&pc=NVN | 2017865-2025-16810 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-03-06 05:00 | NVN | ULTIPACE¿ LEAD, 52CM MRI | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P960013 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-16808. RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-16809. IT WAS REPORTED THAT THE PATIENT HAS DECEASED. THE CAUSE OF DEATH WAS VENTRICULAR TACHYCARDIA (VT). THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) RIGHT VENTRICULAR (RV) AND RIGHT ATRIAL (RA) LEADS WERE EXPLANTED. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: THE LEAD WAS RETURNED DUE TO PATIENT DEATH. AS RECEIVED, A COMPLETE LEAD WAS RETURNED IN ONE PIECE. ELECTRICAL TESTING WAS NORMAL. VISUAL AND X-RAY EXAMINATIONS DID NOT FIND ANY ANOMALIES EXCEPT FOR PROCEDURAL DAMAGE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21539638&pc=BZD | 2518422-2025-102718 | 2025-02-05 05:00 | Death | RESPIRONICS, INC. | 2025-03-06 05:00 | BZD | REP DREAMSTATION AUTO BIPAP, DOM | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION FROM THE PATIENT'S REPRESENTATIVE THAT THE PATIENT HAD PASSED AWAY AND THEY WOULD LIKE TO DONATE THE PATIENT'S DEVICE. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH OF THE PATIENT. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21539644&pc=MHX | 1218950-2025-000127 | 2024-11-23 05:00 | Death | PHILIPS MEDICAL SYSTEMS | 2025-03-06 05:00 | MHX | PATIENT INFORMATION CENTER IX | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | K183387 |  | 1 | Event Description: PHILIPS RECEIVED A COMPLAINT ON THE PATIENT INFORMATION CENTER IX INDICATING THAT THE CUSTOMER REQUESTED INFORMATION ON WHETHER THE DEVICE HAD AN 'OPTIONAL FUNCTION' FOR EGC ANALYSIS AND ALARMS DURING POSTOPERATIVE CARE AND CARDIAC ARREST. PRIOR TO THE ARREST, THE PATIENT WAS PACED AT 80 BPM WITH AN EXTERNAL PACEMAKER VIA THE JUGULAR VEIN WITH A PLANNED INSERTION OF AN IMPLANTABLE PACEMAKER. THE PATIENT HAD AN UNWITNESSED CARDIAC ARREST, WAS FOUND UNCONSCIOUS, AND CPR WAS INITIATED. THE DEFIBRILLATOR DISPLAYED THE PACED RHYTHM BUT THERE WAS NO PULSE. AN ULTRASOUND WAS PERFORMED, WHICH REVEALED CARDIAC STANDSTILL, AND THE PATIENT WAS PRONOUNCED DEAD. IT WAS CLARIFIED THAT THE DEVICE FUNCTIONED AS INTENDED AND THAT THERE WAS AN ASYSTOLE ALARM GENERATED ON THE PIC IX WHICH NOTIFIED THE NURSING STAFF. THE ALARM WAS SILENCED AT THE BEDSIDE 2 MINUTES LATER. THE RESUSCITATION ATTEMPTS WERE UNSUCCESSFUL, AND THE PATIENT PASSED AWAY. THE DEATH WAS UNRELATED TO THE DEVICE AND THERE WAS NO DEVICE MALFUNCTION. Manufacturer Narrative: THIS COMPLAINT IS A SUPPLEMENTAL TO (9610816-2024-001005) DUE TO INCORRECT REGISTRATION NUMBER USED. THE COMPLAINT WAS ESCALATED FOR TECHNICAL INVESTIGATION, AND A CLINICAL APPLICATION SPECIALIST REVIEWED THE EVENT, REVEALING THE ECG RHYTHM AT 1826 IS PRESENT BUT THEN BECOMES VERY CHAOTIC OVER THE NEXT FEW SECONDS, AND TRANSITIONS TO ECG LEADS OFF WITH NO DATA STORED AFTER THAT TIME. THE LOGS ALSO INDICATE THE PATIENT'S DATA WAS ACCESSED BY A USER FIVE TIMES BETWEEN 1830 AND 1900. THE DEVICE WAS FOUND TO BE OPERATING PER SPECIFICATIONS, INCLUDING ALARMS AND CAREEVENT NOTIFICATIONS. BASED ON THE INFORMATION AVAILABLE AND THE TESTING CONDUCTED, THE DEVICE WAS FUNCTIONING AS INTENDED, AND THERE WAS NO TROUBLE FOUND WITH THE DEVICE. THE REPORTED PROBLEM WAS NOT CONFIRMED. THE DEVICE WAS CONFIRMED TO BE OPERATING PER SPECIFICATIONS AND NO FAILURE WAS IDENTIFIED. SECTION E REPORTING INSTITUTION PHONE #: (B)(6), SECTION E REPORTER PHONE #: (B)(6).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21539652&pc=DSQ | 2916596-2025-01348 | 2025-02-19 05:00 | Death | THORATEC CORPORATION | 2025-03-06 05:00 | DSQ | HEARTMATE II LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED AND THE DEVICE FUNCTIONED AS INTENDED. AN AUTOPSY WAS NOT PERFORMED, AND THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: A2 - PATIENT AGE: CORRECTION. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. IT WAS COMMUNICATED THAT NO FURTHER INFORMATION ABOUT THE PATIENT¿S OUTCOME IS ABLE TO BE PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE II LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: A1-A4: PATIENT INFORMATION CANNOT BE PROVIDED DUE TO PERSONAL DATA PRIVACY LEGISLATION/POLICY. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21539826&pc=DSQ | MW5167344 | 2025-02-22 05:00 | Death | BERLIN HEART, INC. | 2025-03-06 05:00 | DSQ | EXCOR PEDIATRIC ATRIAL CANNULA |  | Air Embolism |  |  | 1 | Event Description: PATIENT WITH SINGLE VENTRICLE ANATOMY/PHYSIOLOGY WAS ON A BERLIN HEART-VENTRICULAR ASSIST DEVICE FOR 17 MONTHS WITH 9 MM ATRIAL CANNULA ATTACHED TO THE COMMON ATRIUM AND 9 MM COMPOSITE ARTERIAL CANNULA TO THE AORTA. PATIENT TO OPERATING ROOM FOR HEART TRANSPLANT. IN OPERATING ROOM AND UPON ENTERING THE CHEST CAVITY, AIR WAS NOTED IN THE ATRIAL CANNULA AND EVENTUALLY WITHIN THE ENTIRE VENTRICULAR ASSIST DEVICE. THE VENOUS CANNULAS WERE ALSO FILLED WITH AIR. PATIENT EXPIRED. POST OPERATIVELY IT WAS DISCOVERED THAT THE BERLIN CANNULA HAD SEPARATED FROM THE CANNULA'S PLASTIC SUTURED SEWING RING. THE SUTURE RING WAS STILL SUTURED IN PLACE IN THE ATRIUM. CANNULA ALSO HAD A SMALL CRACK AT THE MEDIAL PORTION OF CANNULA.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21539907&pc=DSQ | 2916596-2025-01589 | 2025-02-12 05:00 | Death | THORATEC CORPORATION | 2025-03-06 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Dyspnea; Shock; Heart Failure/Congestive Heart Failure; Liver Failure; Lactate Dehydrogenase Increased | P160054 |  | 1 | Event Description: IT WAS LATER REPORTED THAT THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO THE PATIENTS' RIGHT HEART FAILURE OR RENAL DYSFUNCTION. THE ONSET OF THE PATIENT'S RENAL FUNCTION WAS SAID TO HAVE BEEN ON (B)(6) 2023. PATIENT WAS DISCHARGED HOME ON MILRINONE DRIP.Event Description: NO OUTFLOW GRAFT OBSTRUCTION OR INFLOW CANNULA OBSTRUCTION WAS CONFIRMED. IT WAS LATER REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2025 DUE TO RIGHT VENTRICULAR (RV) FAILURE. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED.Event Description: IT WAS REPORTED THAT PATIENT PRESENTED TO CLINIC WITH SHORTNESS OF BREATH AND FATIGUE. THE PATIENT WAS ALSO FOUND TO BE IN SHOCK WITH ACUTE KIDNEY INJURY, ACUTE LIVER INJURY, ELEVATED N-TERMINAL PRO B-TYPE NATRIURETIC PEPTIDE, AND A LACTATE VALUE OF 3.4 MMOL/L. IT WAS NOTED THAT PATIENT HAD AN LDH VALUE OF 2531 UNITS PER LITER (U/L) WITHOUT INFECTIOUS PROCESS IN THE MORNING AND THEN HAD A VALUE OF 777U/L LATER THAT DAY. TRANSTHORACIC ECHOCARDIOGRAM AND RIGHT HEART CATHETERIZATION SHOWED EVIDENCE OF RIGHT VENTRICULAR FAILURE AND PATIENT WAS PUT ON A DOBUTAMINE INFUSION AT 5MCG/KG/MIN WITH A LEAVE-IN SWAN-GANZ CATHETER FOR OPTIMIZATION. LOG FILES WERE SUBMITTED FOR REVIEW. THE LOG FILE EVENT HISTORY CAPTURED NO UNUSUAL EVENTS. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: SECTION E1: REPORTER EMAIL WAS NOT AVAILABLE. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: SECTION H3: CORRECTION. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED. ADDITIONALLY, THE REPORTED INFLOW CANNULA AND/OR EXTRINSIC OUTFLOW GRAFT OBSTRUCTIONS COULD NOT BE CONFIRMED THROUGH THIS EVALUATION. THE SUBMITTED CONTROLLER EVENT LOG FILE CAPTURED EVENTS FROM 03FEB2025 THROUGH 13FEB2025. NO NOTABLE EVENTS OR ALARMS WERE CAPTURED, AND THE PUMP APPEARED TO FUNCTION AS INTENDED AT THE FIXED SPEED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE ELECTRONIC IFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) IFU, REV. D, IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (INCLUDING RIGHT HEART FAILURE, HEPATIC DYSFUNCTION, AND RENAL DYSFUNCTION,), HEMOLYSIS, DEVICE THROMBOSIS, AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6, ¿PATIENT CARE AND MANAGEMENT¿, OUTLINES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING THE INTERNATIONAL NORMALIZED RATIO RANGE FOR PATIENTS USING THE HEARTMATE 3 LVAS, AS WELL AS THE SUGGESTED ANTICOAGULATION MODIFICATIONS. SECTION 6, UNDER ¿RIGHT HEART FAILURE¿, ALSO STATES THAT PATIENTS MAY DEVELOP RIGHT VENTRICULAR FAILURE DURING OR SHORTLY AFTER IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS. SECTION 5 OF THE IFU, "SURGICAL PROCEDURES", PROVIDES INSTRUCTIONS ON HOW TO INSERT THE PUMP IN THE VENTRICLE AND INSTRUCTS THE USER TO TAKE CARE TO AVOID ORIENTING THE INLET TOWARDS THE INTERVENTRICULAR SEPTUM AS PUMP FUNCTION WILL BE COMPROMISED IN THE PRESENCE OF INLET OBSTRUCTION. THE STEPS TO ADJUST THE ORIENTATION OF THE PUMP ARE ALSO OUTLINED IN THIS SECTION. SECTION 5, UNDER ¿PREPARING THE VENTRICULAR APEX SITE¿, INSTRUCTS TO INSPECT THE VENTRICULAR CHAMBER FOR MURAL THROMBI AND CROSSING TRABECULAE FOLLOWING REMOVAL OF THE CORE AND TO ADDRESS ONE OR BOTH, AS NEEDED. FURTHERMORE, SECTION 5, UNDER ¿IMPLANT PROCEDURES¿, WARNS TO INSPECT THE VENTRICLE AND REMOVE ANY PREVIOUSLY FORMED CLOTS THAT MAY CAUSE EMBOLISM OR ANY TRABECULAE THAT MAY IMPEDE FLOW. SECTION 5 ALSO OUTLINES CONSIDERATIONS FOR PUMP PLACEMENT AND PROVIDES INSTRUCTIONS REGARDING THE PREPARATION, INSTALLATION, AND ORIENTATION OF THE SEALED OUTFLOW GRAFT. THIS SECTION UNDER "ATTACHING THE SEALED OUTFLOW GRAFT TO THE PUMP," INSTRUCTS THE USER TO VERIFY THAT THE GRAFT IS NOT TWISTED OR KINKED BY CHECKING THE POSITION OF THE BLACK LINE ON THE GRAFT ABOVE AND BELOW THE BEND RELIEF AND ENSURING THAT THE LINE IS STRAIGHT. SECTION 5, UNDER "PREIMPLANT PROCEDURES" AND "IMPLANT PROCEDURES" CAUTIONS THE USER: "THE SEALED OUTFLOW GRAFT MUST NOT BE KINKED OR POSITIONED WHERE IT COULD ABRADE AGAINST A PUMP COMPONENT OR BODY STRUCTURE" AND "STRETCH THE SEALED OUTFLOW GRAFT COMPLETELY PRIOR TO MEASURING AND CUTTING THE GRAFT TO THE APPROPRIATE LENGTH." NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21540007&pc=DSQ | 2916596-2025-01590 | 2025-02-12 05:00 | Death | THORATEC CORPORATION | 2025-03-06 05:00 | DSQ | HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM | Adverse Event Without Identified Device or Use Problem | Bacterial Infection | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY DUE TO PSEUDOMONAS INFECTION AND FAILURE TO THRIVE. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: A2 - PATIENT AGE: CORRECTION A4 - PATIENT WEIGHT WAS INADVERTENTLY REPORTED AS 0 KG IN THE INITIAL REPORT. PATIENT WEIGHT WAS REQUESTED BUT NOT PROVIDED. MANUFACTURER¿S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF INFECTION COULD NOT BE CONCLUSIVELY DETERMINED. INFECTION IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. INFECTION HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21540077&pc=NPT | 2025587-2025-01710 | 2025-02-14 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-06 05:00 | NPT | EVOLUT FX PLUS VALVE | Patient Device Interaction Problem | Tachycardia | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING THE IMPLANT OF A TRANSCATHETER VALVE, THE PATIENT DIED DUE TO VENTRICULAR TACHYCARDIA. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID {D-EVOLUTFX-2329}; PRODUCT LOT/SERIAL NUMBER {UNKNOWN}; PRODUCT TYPE: {0195-HEARTVALVES}; IMPLANT DATE {N/A}; EXPLANT DATE {N/A} MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21540090&pc=DSQ | 2916596-2025-01417 | 2025-02-11 05:00 | Death | THORATEC CORPORATION | 2025-03-06 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REV. C IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21540135&pc=LWS | 2649622-2025-06527 | 2025-03-04 05:00 | Death | MPRI | 2025-03-06 05:00 | LWS | SPRINT QUATTRO SECURE S MRI SURESCAN | Defibrillation/Stimulation Problem; Under-Sensing | Insufficient Information | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IS DECEASED. THE RIGHT VENTRICULAR (RV) LEAD EXHIBITED UNDERSENSING OF THE VENTRICULAR FIBRILLATION (VF) THAT RESULTED IN INAPPROPRIATE WITHHOLDING OF HIGH VOLTAGE THERAPY. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE LEAD WAS NOT RETURNED FOR ANALYSIS; HOWEVER, PERFORMANCE DATA COLLECTED FROM THE DEVICE WAS RECEIVED AND ANALYZED. ANALYSIS OF THE DEVICE MEMORY INDICATED RIGHT VENTRICULAR UNDERSENSING. ANALYSIS OF THE DEVICE MEMORY INDICATED THE CRITERIA FOR THE RIGHT VENTRICULAR LEAD INTEGRITY ALERT WERE MET. ANALYSIS OF THE DEVICE MEMORY INDICATED OVERSENSING DUE TO NON-PHYSIOLOGIC SIGNALS/SENSING INTEGRITY COUNTER. ANALYSIS OF THE DEVICE MEMORY INDICATED OVERSENSING ASSOCIATED WITH THE RIGHT VENTRICULAR LEAD. ANALYSIS OF THE DEVICE MEMORY INDICATED DIMINISHED RIGHT VENTRICULAR SENSING. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21540594&pc=OZD | 1220648-2025-26559 | 2025-02-04 05:00 | Death | ABIOMED, INC. | 2025-03-06 05:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Atrial Fibrillation; Hemorrhage/Bleeding; Perforation; Cardiovascular Insufficiency | P140003 |  | 1 | Event Description: THE COMPLAINANT HAD PLACED A IMPELLA 5.5 PUMP TO SUPPORT A HIGH-RISK PERCUTANEOUS CORONARY INTERVENTION PATIENT. THE TEAM USED IMPELLA AND NON-IMPELLA PRODUCT. THERE IS AN ALLEGATION THAT THE NON IMPELLA WIRE CAUSED A LEFT VENTRICLE PERFORATION THAT WAS OBSERVED AND WAS REPAIRED. THE PATIENT SURVIVED. LATER IN THE MONTH OF THE ONGOING 5.5 PUMP SUPPORT THE LOQI DATABASE REPORTED UPON 3 ADVERSE EVENTS, 2 OF WHICH WERE OF UNKNOWN RELATIONSHIP TO THE PUMP, THESE BEING RIGHT VENTRICLE FAILURE AND ATRIAL ARRYTHMIA/ATRIAL FIBRILLATION WITH RAPID VENTRICULAR RATE, AND THE 3RD WAS DEFINITE RELATIONSHIP, THIS BEING BLEEDING. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING, SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS & CAUTIONS: WARNINGS, SECTION: PRE-SUPPORT EVALUATION, SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER, SECTION: DIRECT AORTIC INSERTION, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION. [ADDITION FOR 5.0, 5.5 ONLY] IN INSTANCES WHERE THE IMPELLA PUMP HAS BEEN PLACED PRIOR TO PERFORMING CARDIAC SURGERY WITH AORTIC CROSS CLAMPING AND CARDIOPLEGIC ARREST, CARE SHOULD BE TAKEN WHEN MANIPULATING THE HEART WHEN THE PUMP POSITION IS FIXED WITH APPLICATION OF THE AORTIC CROSS CLAMP ACROSS THE PUMP CATHETER.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE. SECTION: WARNINGS & CAUTIONS: WARNINGS, SECTION: PRE-SUPPORT EVALUATION, SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER, SECTION: DIRECT AORTIC INSERTION, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: ¿TO REDUCE THE RISK OF CARDIAC INJURY (INCLUDING VENTRICULAR PERFORATION), PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY IN THE SETTINGS OF ACUTE INFARCTION WITH TISSUE NECROSIS.¿ ¿TO REDUCE THE RISK OF VASCULAR INJURY, PHYSICIANS SHOULD EXERCISE CAUTION WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX PERIPHERAL VASCULAR ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: UNREPAIRED ABDOMINAL AORTIC ANEURYSM, SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM, DISSECTION OF THE ASCENDING/ TRANSVERSE/DESCENDING AORTA, CHRONIC ANATOMICAL CHANGES IN THE RELATIONSHIP OF THE AORTA/AORTIC VALVE/VENTRICULAR ALIGNMENT, SIGNIFICANT MOBILE ATHEROMATOUS DISEASE IN THE THORACIC OR ABDOMINAL AORTA OR PERIPHERAL VESSELS.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING A RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY IN THE LEFT VENTRICLE AFTER CPR WITH ECHOCARDIOGRAPHY GUIDANCE. SECTION: WARNINGS & CAUTIONS: CAUTIONS: ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21540733&pc=DSQ | 2916596-2025-01331 | 2025-02-15 05:00 | Death | THORATEC CORPORATION | 2025-03-06 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Fever; Low Blood Pressure/ Hypotension; Sepsis; Vasodilatation; Malaise; Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO SEPSIS, MULTISYSTEM ORGAN FAILURE. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED DUE TO SEPSIS, MULTISYSTEM ORGAN FAILURE, POSITIVE BLOOD CULTURES, NO DRIVELINE EXIT SITE INFECTION. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED AND WILL NOT BE RETURNED FOR EVALUATION.
ADDITIONAL INFORMATION PROVIDED THAT 2 TO 3 DAYS PRIOR TO HOSPITALIZATION, THE PATIENT CALLED WITH REPORTS OF FEVER AND GENERAL MALAISE. THE PATIENT WAS INSTRUCTED TO SEEK CARE LOCALLY BUT REFUSED. A BLOOD CULTURE AND GRAM STAIN ON (B)(6) 2025 DETERMINED THE ONSET OF THE INFECTION AT AN OUTSIDE HOSPITAL. THE SOURCE OF SEPSIS WAS UNKNOWN. THE PATIENT WAS HYPOTENSIVE AND HAD SEVERE DILATORY SHOCK. THE INFECTION WAS TREATED WITH VANCOMYCIN. Manufacturer Narrative: SECTION B3 - DATE OF EVENT: CORRECTED. SECTION E3 - OCCUPATION: CORRECTED. MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C, AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK, REV. D, ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING SEPSIS, INFECTION, MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RESPIRATORY FAILURE, RIGHT HEART FAILURE, RENAL FAILURE, HEPATIC DYSFUNCTION), AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, LISTS INFECTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. SEVERAL SECTIONS OF THE IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING INFECTION PREVENTION, AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21540806&pc=DSQ | 2916596-2025-01327 | 2025-02-12 05:00 | Death | THORATEC CORPORATION | 2025-03-06 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO CARDIAC RHYTHM DISORDERS. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED, AND THE PUMP WAS NOT EXPLANTED.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE PATIENT HAD HISTORY OF ARRHYTHMIA BEFORE IMPLANT. THE ARRHYTHMIA IN THIS EVENT WAS NOT THOUGHT TO BE DEVICE OR THERAPY RELATED. THE ARRHYTHMIA RESULTED IN CLINICAL COMPROMISE. AN IMPLANTABLE CARDIOVERTER DEFIBRILLATOR WAS IMPLANTED PRIOR TO THE VENTRICULAR ASSIST DEVICE. DETAILS LEADING UP TO DEATH WAS NOT AVAILABLE AND THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED THROUGH THE MANUFACTURING EXECUTION SYSTEM (MES) AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. G, IS CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ OF THIS DOCUMENT LISTS CARDIAC ARRHYTHMIA AND DEATH AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. ADDITIONALLY, SECTION 6 ¿PATIENT CARE AND MANAGEMENT¿ LISTS ARRHYTHMIA AS A POTENTIAL LATE POSTIMPLANT COMPLICATION THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21540880&pc=NPT | 2025587-2025-01714 | 2025-02-28 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-06 05:00 | NPT | EVOLUT PRO TRANSCATHETER AORTIC VALVE | Degraded; Patient Device Interaction Problem; Central Regurgitation | Cardiac Arrest; Aortic Valve Insufficiency/ Regurgitation | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT FOUR YEARS AND THREE MONTHS AFTER THE IMPLANT OF THE 29 MM EVOLUT PRO TRANSCATHETER AORTIC VALVE, THE PATIENT PRESENTED WITH ACUTE AORTIC REGURGITATION, AND THE EVOLUT PRO APPEARED DEGENERATED. CONSEQUENTLY, A 23 MM NON-MEDTRONIC TRANSCATHETER VALVE (SAPIEN S3) WAS IMPLANTED VALVE-IN-VALVE IN THE EVOLUT PRO. HOWEVER, AFTER THE IMPLANT OF THE SAPIEN S3 VALVE, AND AS THE IMPLANTING PHYSICIANS WERE CLOSING UP THE VASCULAR ACCESS, THE PATIENT ARRESTED AND WENT INTO PULSELESS ELECTRICAL ACTIVITY. CARDIOPULMONARY RESUSCITATION WAS PERFORMED, AND THE LUCAS (LUND UNIVERSITY CARDIOPULMONARY ASSIST SYSTEM) CHEST COMPRESSION DEVICE WAS USED. FLUOROSCOPY REVEALED THAT THE SAPIEN S3 VALVE HAD MIGRATED LOWER INTO THE LEFT VENTRICLE. THE PATIENT DIED IN THE CATHETER LAB A SHORT TIME LATER.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT PER THE PHYSICIAN, THE CAUSE OF DEATH WAS SEVERE AORTIC REGURGITATION AND THE VALVE INDIRECTLY CONTRIBUTED TO THE PATIENT DEATH AS IT FAILED WITHIN 5 YEARS OF IMPLANT. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID: UNKNOWN 23 MM EDWARDS SAPIEN S3 VALVE; PRODUCT LOT/SERIAL NUMBER: UNKNOWN; PRODUCT TYPE: TRANSCATHETER AORTIC VALVE; IMPLANT DATE: (B)(6) 2025 EXPLANT DATE: N/A. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5 H2 H3 H6. IMAGE REVIEW: MEDIA FILES WERE PROVIDED FOR REVIEW OF THE EVENT DESCRIPTION. PATIENT¿S EXECUTIVE SUMMARY WAS NOT PROVIDED FOR ANATOMICAL REVIEW. THE MEDIA FILES PROVIDED SHOW EVIDENCE OF A NON-MEDTRONIC TRANSCATHETER AORTIC VALVE IN A PREVIOUSLY IMPLANTED EVOLUT. AS LISTED IN THE INSTRUCTIONS FOR USE (IFU), POTENTIAL RISKS ASSOCIATED WITH THE IMPLANTATION OF THE COREVALVE¿ EVOLUT¿ PRO BIOPROSTHESIS MAY INCLUDE, BUT ARE NOT LIMITED TO, THE FOLLOWING: PROSTHETIC VALVE DYSFUNCTION (REGURGITATION OR STENOSIS) DUE TO FRACTURE; BENDING (OUT-OF-ROUND CONFIGURATION) OF THE VALVE FRAME; UNDEREXPANSION OF THE VALVE FRAME; CALCIFICATION; PANNUS; LEAFLET WEAR, TEAR, PROLAPSE, OR RETRACTION; POOR VALVE COAPTATION; SUTURE BREAKS OR DISRUPTION; LEAKS; MAL-SIZING (PROSTHESIS-PATIENT MISMATCH); MALPOSITION (EITHER TOO HIGH OR TOO LOW)/MALPLACEMENT. THERE IS EVIDENCE THAT CARDIOPULMONARY RESUSCITATION WAS INITIATED SOMETIME AFTER THE NON-MEDTRONIC VALVE WAS IMPLANTED, AND FLUOROSCOPY SHOWS THAT THE NON-MEDTRONIC VALVE APPEARED TO HAVE MIGRATED INTO A DEEPER POSITION TOWARDS THE LEFT VENTRICLE. IT WAS REPORTED THAT THE PATIENT SUBSEQUENTLY DIED. ADDITIONAL IMAGES WOULD BE NEEDED TO VALIDATE THE CAUSE OF DEATH. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21541081&pc=KDJ | 1416980-2025-01010 |  | Death | VANTIVE US HEALTHCARE LLC | 2025-03-06 05:00 | KDJ | NI | Adverse Event Without Identified Device or Use Problem | Abdominal Pain; Cardiac Arrest; Peritonitis; Discomfort | NI |  | 1 | Event Description: A PERITONEAL DIALYSIS (PD) PATIENT EXPERIENCED PERITONITIS MANIFESTED BY CLOUDY PD EFFLUENT, ABDOMINAL PAIN AND DISCOMFORT. THE CAUSE OF PERITONITIS WAS UNKNOWN. THE PATIENT WAS HOSPITALIZED AND TREATED WITH VANCOMYCIN INJECTION (1GM, INTRAPERITONEAL, EVERY 5TH DAY, ONGOING) AND CEFTAZIDIME INJECTION (1GM, INTRAPERITONEAL, ONCE DAILY, ONGOING) FOR PERITONITIS. THE PATIENT WAS RECOVERING FROM THE EVENTS. ON AN UNKNOWN DATE, THE PATIENT SUBSEQUENTLY PASSED AWAY. THE CAUSE OF DEATH WAS REPORTED TO BE CARDIAC ARREST. THE PD THERAPY WAS ONGOING TILL DEATH. IT WAS NOT REPORTED IF AN AUTOPSY WAS PERFORMED. NO ADDITIONAL INFORMATION IS AVAILABLE. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED AND THE LOT NUMBER IS UNKNOWN; THEREFORE, A DEVICE ANALYSIS COULD NOT BE COMPLETED. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21542296&pc=NIQ | 2124215-2025-12160 | 2025-02-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-03-06 05:00 | NIQ | SYNERGY XD | Adverse Event Without Identified Device or Use Problem | Thrombosis/Thrombus |  |  | 1 | Event Description: THE PATIENT EXPERIENCED THROMBOSIS AND DIED. A SYNERGY XD WAS IMPLANTED. THE PATIENT PRESENTED WITH STENT THROMBOSIS AND DIED. Manufacturer Narrative: E1 INITIAL REPORTER ADDRESS:(B)(6). B2 DATE OF DEATH, B3 DATE OF EVENT, D6A IMPLANT DATE: USED THE FIRST DAY OF THE MONTH BSC BECAME AWARE OF THIS EVENT AS SPECIFIC DATES WERE NOT PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21542299&pc=MIH | 2017233-2025-05895 | 2025-02-17 05:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-03-06 05:00 | MIH | GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM | Adverse Event Without Identified Device or Use Problem | Obstruction/Occlusion; Ruptured Aneurysm | P040043 |  | 1 | Event Description: THE FOLLOWING INFORMATION WAS REPORTED TO GORE: ON (B)(6) 2025, THE PATIENT UNDERWENT AN EMERGENCY ENDOVASCULAR TREATMENT FOR A RUPTURED THORACIC ANEURYSM USING GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM. THE PATIENT COMPLAINED OF BACK PAIN SEVERAL DAYS AGO, AND IT WAS REPORTEDLY DUE TO A SEALED ANEURYSM RUPTURE. AFTER A 2-DEBRANCH BYPASS, THE FIRST DEVICE (TGM343420J) WAS IMPLANTED IN THE DESCENDING AORTA. THE SECOND DEVICE (TGM454520J) WAS DEPLOYED FROM A POSITION WHERE ITS PARTIAL UNCOVER STENT COVERED THE BRACHIOCEPHALIC ARTERY. FOLLOWING IMPLANTATION, THE BLOOD PRESSURE (MEASURED AT THE RIGHT HAND) DECREASED, AND OCCLUSION OF THE BRACHIOCEPHALIC ARTERY WAS SUSPECTED. ANGIOGRAPHY REVEALED THAT THE BRACHIOCEPHALIC ARTERY WAS COVERED BY THE SECOND DEVICE, AND BLOOD FLOW WAS REDUCED. ADDITIONALLY, AORTIC REGURGITATION (INFLOW OF CONTRAST MEDIUM INTO THE HEART) WAS ALSO CONFIRMED. THERE WAS NO ST ELEVATION OBSERVED ON THE MONITOR. A COVERED STENT (NON-GORE DEVICE) WAS INITIALLY PLACED IN THE BRACHIOCEPHALIC ARTERY AS A TREATMENT. BLOOD FLOW TO THE BRACHIOCEPHALIC ARTERY WAS IMPROVED; HOWEVER, THE HEARTBEAT REMAINED WEAK. CARDIAC MASSAGE AND ADRENALINE ADMINISTRATION WERE PERFORMED, AND THE PATIENT INITIALLY RECOVERED, BUT A FEW MINUTES LATER, VENTRICULAR FIBRILLATION OCCURRED, AND THE PATIENT WENT INTO CARDIAC ARREST. PERCUTANEOUS CARDIOPULMONARY SUPPORT (PCPS) WAS INITIATED, BUT THE PATIENT DID NOT RECOVER. ECHOCARDIOGRAPHY SHOWED BLOOD LEAKAGE INTO THE LEFT THORACIC CAVITY, AND FRESH BLOOD WAS CONFIRMED BY PUNCTURE. IT COULD NOT BE DETERMINED WHETHER THE BLOOD LEAKAGE INTO THE LEFT THORACIC CAVITY WAS DUE TO A RE-RUPTURED ANEURYSM DURING THE PROCEDURE OR FROM BLEEDING CAUSED BY THE INITIAL RUPTURE PRIOR TO THE PROCEDURE. CARDIAC MASSAGE WAS CONTINUED, BUT THE CARDIAC RHYTHM COULD NOT BE RESTORED, AND THE PATIENT EXPIRED. NO AUTOPSY INFORMATION IS AVAILABLE. THE PHYSICIAN'S COMMENTS: THE DEFINITE CAUSE OF DEATH IS UNKNOWN. THERE WERE NO CHOICES IN LIFE-SAVING MEASURES OTHER THAN TEVAR, AND THE INTERVENTION ITSELF WAS NOT AN INCORRECT CHOICE. SEVERAL POSSIBILITIES ARE AS FOLLOWS: INSUFFICIENT BLOOD FLOW TO THE BRACHIOCEPHALIC ARTERY DUE TO ITS COVERAGE BY THE CTAGAC. HOWEVER, THIS EVENT ALONE IS NOT BELIEVED TO HAVE AFFECTED THE CARDIAC RHYTHM. GUIDEWIRE MANIPULATION RESULTING IN DAMAGE TO THE VALVE LEAFLETS AND AORTIC REGURGITATION. THE GUIDEWIRE POTENTIALLY INTERFERING WITH THE ANASTOMOSIS OF THE ASCENDING AORTA FOR CORONARY ARTERY BYPASS GRAFTING. ANEURYSM RE-RUPTURE DUE TO ELEVATED PRESSURE CAUSED BY CARDIAC MASSAGE AND ADRENALINE ADMINISTRATION. Manufacturer Narrative: H6: A REVIEW OF THE MANUFACTURING RECORDS FOR THE DEVICE VERIFIED THAT THE LOT MET ALL PRE-RELEASE SPECIFICATIONS. H6: CODE D12: ACCORDING TO THE GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM INSTRUCTIONS FOR USE (IFU), ADVERSE EVENTS THAT MAY OCCUR AND/OR REQUIRE INTERVENTION INCLUDE, BUT ARE NOT LIMITED TO : DEATH, AORTIC RUPTURE, BRANCH VESSEL OCCLUSION OR OBSTRUCTION, CARDIAC (E.G., ARRHYTHMIA, MYOCARDIAL INFARCTION, CONGESTIVE HEART FAILURE, HYPOTENSION OR HYPERTENSION) W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21542354&pc=NIQ | 2124215-2025-14098 | 2025-02-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-03-06 05:00 | NIQ | SYNERGY MEGATRON | Adverse Event Without Identified Device or Use Problem | Thrombosis/Thrombus |  |  | 1 | Event Description: THE PATIENT EXPERIENCED THROMBOSIS AND DIED. A SYNERGY MEGATRON WAS IMPLANTED. THE PATIENT PRESENTED WITH STENT THROMBOSIS AND DIED. Manufacturer Narrative: E1 INITIAL REPORTER ADDRESS: (B)(6). B2 DATE OF DEATH, B3 DATE OF EVENT, D6A IMPLANT DATE: USED THE FIRST DAY OF THE MONTH BSC BECAME AWARE OF THIS EVENT AS SPECIFIC DATES WERE NOT PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21542363&pc=PYX | 1220648-2025-26560 | 2025-02-11 05:00 | Death | ABIOMED, INC. | 2025-03-06 05:00 | PYX | IMPELLA RP FLEX | Positioning Failure | Insufficient Information | P170011 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT AN IMPELLA RP FLEX PUMP WAS INADVERTENTLY PULLED BACK DURING PATIENT SUPPORT. IT WAS NOTED THAT THE SUTURE HUB HAD BROKEN OFF THE REPOSITIONING SHEATH BUT WAS STILL SUTURED INTACT TO THE NECK. THE PATIENT PASSED AWAY LATER THAT DAY. IT IS UNKNOWN IF THE COMPLICATION CONTRIBUTED TO THE PATIENT'S PASSING. Manufacturer Narrative: THE INVESTIGATION OF POSITIONING ISSUES HAS BEEN COMPLETED. THE PRODUCT WAS NOT RETURNED FOR REVIEW. IT WAS NOTED IN THE COMPLAINT DESCRIPTION THAT THE SUTURE PAD WAS BROKEN OFF THE REPO AT THE SAME TIME THE IMPELLA RP FLEX WAS FOUND TO BE DISLODGED INTO THE IVC, INDICATING FORCE WAS LIKELY A FACTOR. OTHER DETAILS SURROUNDING WAT WAS HAPPENING AT THE TIME OF THE DISLODGMENT ARE UNKNOWN. THE ROOT CAUSE OF POSITIONING ISSUE WAS UNABLE TO BE DETERMINED AS LIMITED CLINICAL DETAILS WERE PROVIDED E4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26560 IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21542376&pc=OZD | 1220648-2025-26563 | 2025-02-12 05:00 | Death | ABIOMED, INC. | 2025-03-06 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. THE PATIENT EXPERIENCED A CEREBRAL HEMORRHAGE AND WAS NOTED TO BE IN CRITICAL CONDITION. THE PATIENTS BLOOD PRESSURE WAS CONTROLLED BY ADMINISTERING NICORANDIL AND AFTER CEREBRAL HEMORRHAGE WAS CONFIRMED, THE ADDITION OF HEPARIN TO THE PURGE FLUID WAS DISCONTINUED. THE EXACT CAUSE OF THE CVA IS UNKNOWN, AND THE CAUSAL RELATIONSHIP WITH IMPELLA IS ALSO UNKNOWN. THE PATIENT CONTINUED ON SUPPORT. ADDITIONAL INFORMATION NOTED CARE WAS THE PATIENT WAS MADE DO NOT RESUSCITATE AND THE PATIENT EXPIRED. THE CAUSE OF DEATH WAS ACUTE MYOCARDIAL INFARCTION AND CEREBRAL HEMORRHAGE. NO ISSUES WERE IDENTIFIED WITH IMPELLA OPERATION. THE CAUSAL RELATIONSHIP BETWEEN CEREBRAL HEMORRHAGE, ONE OF THE CAUSES OF DEATH, AND IMPELLA IS UNKNOWN. THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA DEVICE AS AN ASSOCIATED FACTOR IN THE PATIENT'S DEMISE. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA IFU: IMPELLA CP WITH SMARTASSIST SYSTEM: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿ Manufacturer Narrative: THE INVESTIGATION OF STROKE/CEREBROVASCULAR ACCIDENT (CVA) HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. AS THIS CLINICAL INFORMATION WAS NOT PROVIDED, THE TRUE ROOT CAUSE OF THE STROKE/CVA COULD NOT BE DETERMINED. E.1 ADDED FACILITY NAME AS IT WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26563. E.4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26563 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. G.2 REPORT SOURCE; FOREIGN WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26563.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21543016&pc=NPT | 2015691-2025-01580 | 2025-02-13 05:00 | Death | EDWARDS LIFESCIENCES | 2025-03-06 05:00 | NPT | EDWARDS COMMANDER DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | Cardiac Perforation | P140031 |  | 1 | Event Description: AS REPORTED BY AN EDWARDS LIFESCIENCES AFFILIATE IN FRANCE, REGARDING AN IMPLANT OF A 29MM SAPIEN 3 IN MITRAL POSITION BY VENOUS TRANSFEMORAL APPROACH WITHIN A SURGICAL DEGENERATED BIOPROSTHESIS. THE PROCEDURE WENT WITHOUT COMPLICATIONS, THE VALVE WAS CORRECTLY DEPLOYED, AND THE DELIVERY SYSTEM RETRIEVED SMOOTHLY. HOWEVER, IMMEDIATE ECHO ASSESSMENT REVELED AN ABNORMAL FLOW PATTERN, LEADING TO THE IDENTIFICATION OF A PERICARDIAL LACERATION WITH SIGNIFICANT BLEEDING. CARDIAC SURGEON ASSESSED THE SITUATION AND DETERMINED THAT NO SURGICAL INTERVENTION WAS FEASIBLE GIVEN PATIENT'S INOPERABLE STATUS. TO CONTROL THE BLEEDING, THE TEAM PROCEED WITH A PERCUTANEOUS CLOSURE USING A 1.8CM VASCULAR OCCLUDER. IT FAILED SINCE THE DEVICE PASSED THROUGH THE LACERATION. A SECOND 2.6CM VASCULAR OCCLUDER WAS USED APPARENTLY ACHIEVING HOMEOSTASIS. HOWEVER, FIVE MINUTES LATER ECHO RELEVELED THE OCCLUDER HAD DISLODGED AND WAS FREELY MOBILE WITHIN THE PERICARDIAL SPACE. WITH NO OTHER VIABLE INTERVENTIONAL OPTIONS, THE DECISION WAS MADE TO WITHDRAW ACTIVE THERAPEUTIC MEASURES. PATIENT REMAINED WITH 10MG/H OF NOREPINEFRINE FOR HEMODYNAMIC SUPPORT AND WAS TRANSFERRED TO THE ICU WITH AN EXTREMELY POOR PROGNOSIS. AS PER MEDICAL OPINION, THE ROOT CAUSE OF THE EVENT WAS THE SAFARI GUIDEWIRE WHICH CAUSED THE LACERATION AT THE APEX. ALTHOUGH SEPTAL PASSAGE APPEARED UNEVENTFUL AND THE PERICARDIUM WAS NOTABLY THICK, THE TEAM CONCLUDED THAT THE WIRE LIKELY EXERTED EXCESSIVE FORCE AT THE APEX, LEADING TO THE INJURY. THEY ACKNOWLEDGED THAT THIS TYPE OF COMPLICATION REMAINS DIFFICULT TO ANTICIPATE AND MANAGE.Event Description: ADDITIONAL INFORMATION RECEIVED: THE PATIENT PASSED AWAY IN THE RECOVERY ROOM. Manufacturer Narrative: E1 PHONE NUMBER (B)(6). PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. THERE ARE SEVERAL POTENTIAL ETIOLOGIES FOR VENTRICULAR PERFORATION DURING A TVR PROCEDURE, INCLUDING PERFORATION BY THE GUIDEWIRE, THE DELIVERY SYSTEM, OR THE TRANSVENOUS PACER (TVP) LEAD. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVE (THV). TRAINING INCLUDES PROPER GUIDEWIRE POSITIONING, FIXATION OF THE TVP TO PREVENT VENTRICLE PERFORATION, AND CAREFUL MANIPULATION OF DEVICES. PER THE PROCEDURE DIDACTIC, PATIENTS WITH SMALL VENTRICLES ARE AT PARTICULARLY HIGH RISK FOR VENTRICULAR PERFORATION. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST/INDICATE PROCEDURAL FACTORS (THE ROOT CAUSE OF THE EVENT WAS THE SAFARI GUIDEWIRE WHICH CAUSED THE LACERATION AT THE APEX) CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED DUE TO ADDITIONAL INFORMATION. SECTIONS: B2, B5, G3, G6, H1, H2, H6 HEALTH EFFECT - IMPACT CODE HAVE BEEN UPDATED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21543129&pc=DSQ | 2916596-2025-01268 | 2025-02-12 05:00 | Death | THORATEC CORPORATION | 2025-03-06 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD TROUBLE OXYGENATING AND DID NOT DO WELL AFTER CARE WAS DISCONTINUED. THE PATIENT PASSED AWAY ON (B)(6) 2025 DUE TO RIGHT HEART FAILURE THAT DID NOT EXIST PRE-LEFT VENTRICULAR ASSIST DEVICE (LVAD). THE CAUSE OF DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED. A DEVICE RELATED ISSUE DID NOT CAUSE OR CONTRIBUTE TO RIGHT HEART FAILURE. THE DEVICE FUNCTIONED AS INTENDED AND NO EQUIPMENT WOULD BE RETURNED. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. IT WAS REPORTED THAT HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6) WAS USED AS A RIGHT VENTRICULAR ASSIST DEVICE (RVAD) WHICH IS NOT AN APPROVED INDICATION. THE HEARTMATE 3 IS INDICATED FOR LVAD SUPPORT, AND ALL LABELING, PHYSICIAN TRAINING AND CUSTOMER MARKETING MATERIALS ARE ALIGNED WITH THIS INDICATION. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING RIGHT HEART FAILURE AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 OF THE IFU (UNDER "RIGHT HEART FAILURE") DISCUSSES THE POTENTIAL DEVELOPMENT OF RIGHT HEART FAILURE FOLLOWING IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS, INCLUDING INOTROPES AND RVAD PLACEMENT. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21543133&pc=DSQ | 2916596-2025-01265 | 2025-02-17 05:00 | Death | THORATEC CORPORATION | 2025-03-06 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Insufficient Information; No Clinical Signs, Symptoms or Conditions | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THIS EVENT IS SUPPLEMENTAL AND THAT THE INVESTIGATION FINDINGS WILL BE SUBMITTED UNDER MFR # 2916596-2024-01768. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21543262&pc=NAY | 2955842-2025-05720 | 2025-02-03 05:00 | Death | INTUITIVE SURGICAL, INC | 2025-03-06 05:00 | NAY | DA VINCI ENERGY | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Perforation of Vessels | K191280 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED MEDIASTINAL MASS RESECTION, THE BRACHIOCEPHALIC ARTERY WAS INJURED AND THE PROCEDURE WAS CONVERTED TO OPEN. THE PATIENT EXPIRED ON POST-OPERATIVE DAY ONE. THE SURGEON REPORTED THAT THE ARTERY WAS INJURED WHEN USING THE SYNCHROSEAL INSTRUMENT TO DISSECT LYMPH NODES. THE ARTERY WAS ADHERED TO TISSUE AND TO THE LYMPH NODES, RESULTING IN THE ARTERY NOT BEING VISIBLE TO THE SURGEON. THE ARTERY BEGAN TO BLEED AFTER IT WAS GRASPED WITH THE SYNCHROSEAL INSTRUMENT. THE PROCEDURE WAS CONVERTED TO THORACOTOMY. THERE WAS NO MALFUNCTION OF THE SYNCHROSEAL INSTRUMENT OR ANY OTHER DA VINCI PRODUCTS USED DURING THIS PROCEDURE. NO ADDITIONAL INFORMATION WAS PROVIDED.Event Description: REFER TO H11 FOR FOLLOW-UP INFORMATION. Manufacturer Narrative: THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: 30 DEGREE ENDOSCOPE PLUS, CURVED BIPOLAR DISSECTOR, TIP-UP FENESTRATED GRASPER, FENESTRATED BIPOLAR FORCEPS, MEDIUM-LARGE CLIP APPLIER, SYNCHROSEAL. A SITE HISTORY REVIEW FOUND NO RELATED COMPLAINTS WERE RECEIVED FOR THE INSTRUMENTS USED DURING THIS PROCEDURE. A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED FOR THE DEVICE(S) USED DURING THE PROCEDURE AND NO NON-CONFORMANCES WERE IDENTIFIED TO BE RELATED TO THIS EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT HAD A CATASTROPHIC INJURY TO THE INNOMINATE ARTERY WHILE ATTEMPTING TO REMOVE A MEDIASTINAL MASS AND ADHERENT LYMPH NODES IN THE VICINITY OF THIS LARGE ARTERY. THE SURGEON COMMENTED THEY DID NOT RECOGNIZE THE ARTERY AS THEY WERE ATTEMPTING TO DISSECT THE NEARBY LYMPH NODES AND INADVERTENTLY GRASPED THIS LARGE VESSEL. THIS SYNCHROSEAL IS NOT INTENDED FOR MAJOR ARTERIES AND THIS IS A SURGICAL TECHNIQUE ERROR. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, NO INTUITIVE SURGICAL PRODUCT OR INSTRUMENT CONTRIBUTED TO THIS EVENT. Manufacturer Narrative: A REVIEW OF THE SYSTEM LOGS FOUND THAT NO RELEVANT ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21543490&pc=NAY | 2955842-2025-05723 | 2022-07-07 04:00 | Death | INTUITIVE SURGICAL, INC | 2025-03-06 05:00 | NAY | DAVINCI X | Insufficient Information | Sepsis; Peritonitis; Unspecified Tissue Injury | K171294 |  | 1 | Event Description: IT WAS REPORTED IN A LOCAL NEWSPAPER ARTICLE THAT A PATIENT EXPIRED AFTER A DA VINCI-ASSISTED HYSTERECTOMY FOR METASTATIC UTERINE CANCER. THE PROCEDURE WAS CONVERTED TO OPEN DUE TO UNSPECIFIED COMPLICATIONS. DURING THE DA VINCI-ASSISTED PORTION OF THE PROCEDURE, LYMPH NODES WERE REMOVED FROM THE PATIENT'S LEFT PELVIS. THE DA VINCI SYSTEM WAS REPOSITIONED TO REMOVE LYMPH NODES ALONG THE ABDOMINAL AORTA, WHICH REQUIRED PLACEMENT OF THREE ADDITIONAL TROCARS THAT WERE INSERTED UNDER VISUAL GUIDANCE. AFTER THE DA VINCI SYSTEM WAS REPOSITIONED, INSUFFLATION WAS LOST AND THE ABDOMINAL CAVITY COLLAPSED. EFFORTS TO RE-INSUFFLATE THE ABDOMEN FAILED. THE PROCEDURE WAS CONVERTED TO OPEN AND LASTED FOR 11 HOURS AND 21 MINUTES. THE ARTICLE STATES THAT THE PATIENT'S SMALL INTESTINE WAS INJURED AND THE LEFT URETER WAS SEVERED. DETAILS OF HOW OR WHEN THESE INJURIES OCCURRED IS NOT PROVIDED. THE ARTICLE STATED INJURIES REPORTEDLY CAUSED LEAKS LEADING TO PERITONITIS AND SEVERE SEPSIS, RESULTING IN A TOTAL OF 20 ADDITIONAL UNSPECIFIED PROCEDURES PERFORMED. THE PATIENT EXPIRED ON POST-OPERATIVE DAY 47. THE NEWSPAPER ARTICLE DID NOT INDICATE WHETHER THE COMPLICATIONS OCCURRED DURING THE ROBOTIC PORTION OR THE OPEN PORTION OF THE PROCEDURE, AS THE INJURIES WERE NOT DISCOVERED UNTIL OVER A DAY LATER. THE SURGEON IDENTIFIED IN THE ARTICLE WAS CONTACTED BUT DECLINED TO PROVIDE ANY INFORMATION. Manufacturer Narrative: REVIEW OF THE SYSTEM LOGS FOUND NO RELEVANT ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THE PROCEDURE: 30 DEGREE ENDOSCOPE PLUS, MONOPOLAR CURVED SCISSORS, SMALL GRASPING RETRACTOR, FENESTRATED BIPOLAR FORCEPS. A SITE HISTORY REVIEW FOUND NO RELATED COMPLAINTS WERE RECEIVED FOR THE INSTRUMENTS USED DURING THE PROCEDURE. A DEVICE HISTORY RECORD (DHR) REVIEW FOR THE DEVICE(S) USED DURING THE PROCEDURE FOUND NO NON-CONFORMANCES WERE IDENTIFIED TO BE RELATED TO THIS EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT UNDERWENT A ROBOTIC HYSTERECTOMY FOR UTERINE CANCER AND THE PROCEDURE WAS EVENTUALLY CONVERTED TO OPEN. INJURIES TO THE BOWEL OCCURRED BUT THE DETAILS OF HOW OR WHEN THESE INJURIES OCCURRED IS NOT PROVIDED. THE PATIENT SUBSEQUENTLY UNDERWENT MANY ADDITIONAL PROCEDURES AND ULTIMATELY EXPIRED 47 DAYS AFTER THE INDEX OPERATION. THE REASONS FOR THESE ADDITIONAL PROCEDURES, FINDINGS, AND OTHER DETAILS ARE NOT PROVIDED. HOW THE PATIENT DEVELOPED SEPSIS, HOW THEY DIED, AND THE DETAILS OF EVENTS THAT LED TO THEIR DEATH IS NOT PROVIDED. NO ALLEGATION WAS MADE REGARDING ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21522101&pc=BZD | 2518422-2025-026756 | 2024-12-10 05:00 | Death | RESPIRONICS, INC. | 2025-03-05 05:00 | BZD | DS2ADV AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K200480 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH OF THE PATIENT. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS INSPECTED AND THE UNIT PASSED THE FUNCTIONAL TEST. THE DEVICE'S EVENT LOG WAS REVIEWED BY THE SERVICE CENTER AND ERROR CODE FOUND. ADDITIONALLY, CUSTOMER COMPLAINT WAS NOT REPRODUCED, AND THE DEVICE WAS SCRAPPED. THERE WAS AN ALLEGATION OF SERIOUS OR PERMANENT HARM OR INJURY. NO MEDICAL INTERVENTION WAS SPECIFIED. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21523575&pc=BZD | 2518422-2025-026601 | 2024-06-05 04:00 | Death | RESPIRONICS, INC. | 2025-03-05 05:00 | BZD | DS2ADV AUTO CPAP | Insufficient Information | Insufficient Information | K200480 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION REGARDING A DS2ADV AUTO CPAP. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT PASSED AWAY. AT THIS TIME MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21523741&pc=NIP | 2017233-2025-05890 | 2025-02-07 05:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-03-05 05:00 | NIP | GORE® VIABAHN® ENDOPROSTHESIS WITH HEPARIN BIOACTIVE SURFACE | Patient-Device Incompatibility | Necrosis | P040037 |  | 1 | Event Description: THE FOLLOWING LITERATURE WAS REVIEWED. "CLINICAL OUTCOMES AND RISK FACTORS FOR VIABAHN STENT GRAFT OCCLUSION: IN THE TREATMENT OF VISCERAL ARTERIAL INJURIES IN CANCER PATIENTS" YUJI KORETSUNE, ET AL. INTERVENTIONAL RADIOLOGY, 9(3):172-179. SUMMARY: THIS IS A MULTI-CENTER RETROSPECTIVE STUDY TO EVALUATE TECHNICAL AND CLINICAL SUCCESS RATES, PERIPROCEDURAL COMPLICATIONS AND STENT GRAFT [SG] PATENCY OF 26 [21 MEN AND 5 WOMEN, AGES 59.5-75 YEARS OLD] PATIENTS THAT UNDERWENT GORE®VIABAHN®ENDOPROSTHESIS [VSX] PLACEMENT FOR VISCERAL ARTERIAL INJURIES BETWEEN DECEMBER 2017 AND NOVEMBER 2022. IN THE STUDY, 29 VSX DEVICES WERE PLACE. THE DIAGNOSES FOR THE PATIENTS INCLUDED PANCREATIC CANCER [N=10], DISTAL CHOLANGIOCARCINOMA [N=7], GASTRIC CANCER [N=3], DUODENAL PAPILLARY CARCINOMA [N=3], AND TESTICULAR TUMORS [N=1]. INDICATIONS FOR THE 26 CASES, 22 HAD PRIOR SURGERY, WITH A MEDIAN INTERVAL OF 15 DAYS BETWEEN SURGERY AND ONSET OF HEMORRHAGE. PSEUDOANEURYSM WAS CAUSED BY VESSEL EROSION RESULTING FROM PANCREATIC CANCER IN THREE CASES, WHILE THE ETIOLOGY OF BLEEDING WAS UNKNOWN IN THE REMAINING ONE CASE. DURING THE PROCEDURE, VSX WAS PLACED ACCORDING TO INSTRUCTIONS FOR USE AND IN THE EVENT OF A BRANCH ARTERY AROUND THE BLEEDING POINT, COIL EMBOLIZATION OF THE BRANCH ARTERY WAS PERFORMED BEFORE VSX PLACEMENT PREVENT A TYPE 2 ENDOLEAK AND BALLOON DILATION OR GRAFT EXTENSION WAS CHOSEN. THE TECHNICAL AND CLINICAL SUCCESS RATES WERE 96.6% (28 OF 29) AND 86.2% (25 OF 29), RESPECTIVELY. ¿ONE PATIENT EXPERIENCED TECHNICAL FAILURE DUE TO CELIAC ARTERY DISSECTION DURING THE PROCEDURE. ¿FOUR CASES RESULTED IN CLINICAL FAILURE; TWO HAD REBLEEDING FROM THE PROXIMAL SIDE OF THE VIABAHN SG, AND TWO EXPERIENCED A BLOODY FLUID COLLECTION FROM SURGICAL DRAINS; NEVERTHELESS, NO SOURCE OF BLEEDING COULD BE IDENTIFIED. THE PERIPROCEDURAL COMPLICATION RATE WAS 17.2% (5 OF 29), INCLUDING FOUR MODERATE ADVERSE EVENTS AND ONE PATIENT DEATH. ¿PARTICULARLY, MINOR THROMBOSES IN THE HEPATIC ARTERY WERE OBSERVED IN TWO INSTANCES ON POST-SG IMPLANTATION ANGIOGRAPHY; HOWEVER, THEY RESOLVED WITHOUT FURTHER CARE AND CAUSED NO HEPATIC INFARCTION. ¿IN ONE CASE, CELIAC ARTERY DISSECTION WAS FOUND; HOWEVER, IT DID NOT CAUSE ISCHEMIC SEQUELAE. ¿IN ANOTHER INSTANCE, A SPLENIC ARTERY PERFORATION OCCURRED IMMEDIATELY DISTAL TO THE BLEEDING SPOT WITH A 0.035- INCH HYDROPHILIC GUIDEWIRE, WHICH WAS FIXED BY COVERING BOTH WOUNDED SITES WITH A SINGLE VIABAHN SG. ¿ONE PATIENT, A 76-YEAR-OLD WOMAN WITH A HISTORY OF PANCREATICODUODENECTOMY FOR DUODENAL PAPILLA CARCINOMA 3 YEARS EARLIER AND DISTAL GASTRECTOMY FOR GASTRIC CANCER 13 YEARS PRIOR. 3.0-CM PSEUDOANEURYSM IN THE CELIAC ARTERY AND TWO VSX WERE PLACED [6MM AND 5MM] WERE PLACED WERE PLACED FROM THE COMMON HEPATIC ARTERY TO THE CELIAC ARTERY, FOLLOWING 8-MM AVP-4 INSERTION IN THE SPLENIC ARTERY TO PREVENT TYPE 2 ENDOLEAK. CELIAC ARTERIOGRAPHY REVEALED THE DISAPPEARANCE OF THE PSEUDOANEURYSM AND THROMBUS IN THE PROPER HEPATIC ARTERY, RESULTING IN REDUCED BLOOD FLOW IN THE LEFT HEPATIC ARTERY. THE CLEARANCE OF THIS THROMBUS WAS VERIFIED THROUGH POSTOPERATIVE CT WITH NO FURTHER THERAPY. CONTRAST-ENHANCED CT 2 DAYS FOLLOWING SG INSTALLATION SHOWED NECROSIS OF THE STOMACH, SPLEEN, AND PANCREAS, WHICH WERE THOUGHT TO BE RELATED WITH SG PLACEMENT. THE PATIENT DIED 6 DAYS AFTER SG INSTALLATION. ¿BRANCH ARTERY EMBOLIZATION WAS DONE IN FIVE CASES TO PREVENT TYPE 2 ENDOLEAKS. ¿ANGIOGRAMS FOLLOWING SG INSTALLATION SHOWED A TYPE 1 ENDOLEAK IN TWO CASES; ONE WAS RESOLVED FOLLOWING SUBSEQUENT BALLOON PERCUTANEOUS TRANSLUMINAL ANGIOPLASTY, AND THE OTHER WAS ADDRESSED BY ADDITIONAL SG IMPLANTATION ON THE PROXIMAL SIDE. ¿THE 30-DAY MORTALITY RATE WAS 15.4% (4 OF 26 CASES), INCLUDING 3 CASES OF MULTIORGAN FAILURE AND 1 CASE OF PNEUMONIA-RELATED RESPIRATORY FAILURE. FOLLOW UP CT PATENCY [206-584.25 DAYS] THE SG PATENCY RATES AT 1, 3, 6, 12, AND 24 MONTHS WERE 85%, 84.2%, 77.8%, 66.7%, AND 50%, RESPECTIVELY FOLLOW-UP CT WAS PERFORMED IN 20 OF THE 26 PATIENTS (4 PATIENTS DIED, 1 PATIENT EXPERIENCED TECHNICAL ISSUE, AND 1 PATIENT WAS LOST TO FOLLOW-UP). AMONG THE 20 PATIENTS, 15 UNDERWENT ANTIPLATELET THERAPY AFTER SG INSERTION. ¿SG OCCLUSION OCCURRED IN SEVEN INDIVIDUALS (AFTER 5, 20, 21, 176, 191, 218, AND 492 DAYS, RESPECTIVELY). ¿IN FOUR CASES, OCCLUDED SGS WERE IMPLANTED IN THE SPLENIC ARTERY, TWO IN THE COMMON HEPATIC ARTERY TO THE RIGHT HEPATIC ARTERY, AND ONE IN THE PROPER HEPATIC ARTERY TO THE RIGHT HEPATIC ARTERY. Manufacturer Narrative: HEPARIN SURFACE INCORPORATES CARMEDA HEPARIN MANUFACTURED FROM HEPARIN SODIUM API, WHICH IS COVALENTLY BOUND TO THE DEVICE SURFACE AND IS ESSENTIALLY NON-ELUTING. A1: NO PATIENT SPECIFIC DETAILS HAVE BEEN PROVIDED. THEREFORE, THE PATIENT INITIALS REFLECT THE W. L. GORE INTERNAL CASE NUMBER. B2: THE DATE OF PATIENT DEATH IS UNKNOWN. H6: A REVIEW OF THE MANUFACTURING RECORDS FOR THE DEVICE COULD NOT BE PERFORMED AS A VALID LOT NUMBER WAS NOT PROVIDED. H6: CODE D12: ACCORDING TO THE GORE® VIABAHN® ENDOPROSTHESIS WITH HEPARIN BIOACTIVE SURFACE INSTRUCTIONS FOR USE (IFU), ADVERSE EVENTS THAT MAY OCCUR AND/OR REQUIRE INTERVENTION INCLUDE, BUT ARE NOT LIMITED TO: DEATH, TISSUE NECROSIS. LITERATURE TITLE: CLINICAL OUTCOMES AND RISK FACTORS FOR VIABAHN STENT GRAFT OCCLUSION : IN THE TREATMENT OF VISCERAL ARTERIAL INJURIES IN CANCER PATIENTS SOURCE: INTERVENTIONAL RADIOLOGY, 9(3):172-179. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21524132&pc=PNJ | 9612164-2025-01156 | 2025-02-14 05:00 | Death | MEDTRONIC IRELAND | 2025-03-05 05:00 | PNJ | MICRA¿ AV2 | Device Dislodged or Dislocated | Insufficient Information | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IS DECEASED. IT WAS FURTHER REPORTED THAT THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) DELIVERY SYSTEM EXHIBITED DISLODGEMENT DUE TO TETHER CUTTING. IT WAS NOTED THAT THE LEADLESS IPG EXHIBITED DISLODGEMENT/DISLOCATION/MIGRATION. THE LEADLESS IPG WAS EXPLANTED AND REPLACED.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT'S DEATH WAS NOT DEVICE RELATED. Manufacturer Narrative: CONTINUATION OF D10: 3058 STIM IPF; THERAPY DATE: (B)(6) 2019 THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [MICRA], PRODUCT ID [MC2AVR1-DELSYS], (SERIAL: (B)(6)). D9: NO MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21524645&pc=BZD | 2518422-2025-102622 | 2024-10-07 04:00 | Death | RESPIRONICS, INC. | 2025-03-05 05:00 | BZD | DREAMSTATION AUTO BIPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS NOTIFIED THAT A PATIENT WHO HAD BEEN USING A REP DREAMSTATION AUTO BIPAP DEVICE HAS PASSED AWAY. THE PATIENT'S WIFE CONTACTED THE MANUFACTURER, STATING THAT HER HUSBAND HAD PASSED AWAY AND THAT SHE NO LONGER REQUIRED THE DEVICE. SINCE THE DEVICE HAD BEEN USED, THE RESPIRATORY THERAPIST INFORMED THE CALLER THAT IT WAS NOT NECESSARY TO RETURN THE DEVICE. THE CALLER INQUIRED ABOUT THE POSSIBILITY OF FINANCIAL COMPENSATION, AS SHE HAD HEARD ABOUT A SETTLEMENT. UPON REVIEWING THE X PRIORITY NOTIFICATION, THE RESPIRATORY THERAPIST CONFIRMED THAT THE AFFECTED DEVICE HAD NOT BEEN RETURNED. THE CALLER INDICATED THAT SHE DID NOT HAVE ANY TRACKING INFORMATION FOR THE DEVICE, EXPLAINING THAT THE DURABLE MEDICAL EQUIPMENT (DME) PROVIDER HAD TAKEN POSSESSION OF IT. THE CALLER WAS SUBSEQUENTLY INFORMED THAT PHILIPS' ONLY REMEDIATION FOR THE AFFECTED DEVICE WAS A REPLACEMENT UNIT, WHICH THE CALLER ACKNOWLEDGED AND UNDERSTOOD. NO ADDITIONAL DETAILS WERE PROVIDED, AND THERE IS NO INDICATION OR ALLEGATION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE PATIENT'S DEATH. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21524837&pc=BZD | 2518422-2025-102623 | 2025-01-09 05:00 | Death | RESPIRONICS, INC. | 2025-03-05 05:00 | BZD | DREAMSTATION AUTO BIPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS NOTIFIED THAT A PATIENT WHO HAD BEEN USING A DREAMSTATION AUTO BIPAP DEVICE HAS PASSED AWAY. THE PATIENT'S WIFE CONTACTED THE MANUFACTURER, STATING THAT HER HUSBAND HAD PASSED AND REQUESTED A REFUND FOR THE DEVICE. IN RESPONSE, THE MANUFACTURER CLARIFIED THAT DEVICES ARE BEING PROVIDED TO PATIENTS THROUGH THE REPAIR AND REPLACEMENT PROGRAM, AND THAT FINANCIAL COMPENSATION, REIMBURSEMENT, OR PAYMENT IS NOT BEING OFFERED. ADDITIONALLY, THE WIFE WAS PROVIDED WITH THE CONTACT INFORMATION (PHONE NUMBER AND EMAIL ADDRESS) OF A THIRD-PARTY SETTLEMENT GROUP, WHICH HAS BEEN RETAINED TO ADMINISTER AN ECONOMIC LOSS SETTLEMENT, SHOULD SHE HAVE ANY FURTHER QUESTIONS. NO ADDITIONAL DETAILS WERE PROVIDED, AND THERE IS NO INDICATION OR ALLEGATION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE PATIENT'S DEATH. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21524917&pc=OZD | 1220648-2025-26480 | 2025-02-10 05:00 | Death | ABIOMED, INC. | 2025-03-05 05:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Hematoma; Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT DURING IMPELLA 5.5 AXILLARY CUTDOWN AND ANASTOMOSIS, THERE WAS INITIALLY SOME BLEEDING FROM THE ANASTOMOSIS THAT WAS THEN REPAIRED. ADDITIONALLY, THE PATIENT HAD A HEMATOMA AND BLEEDING FROM THE ACCESS SITE. THE PATIENT WAS TAKEN TO THE OPERATING ROOM FOR A WASHOUT. TWO UNITS OF BLOOD WERE ADMINISTERED, AND HEPARIN WAS HELD. THE PATIENT REMAINS ON SUPPORT.Event Description: IT WAS FURTHER REPORTED THAT CARE WAS WITHDRAWN, AND THE PATIENT EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. THE ROOT CAUSE OF ACCESS SITE BLEEDING WAS MOST LIKELY PATIENT CONDITION SINCE THE PATIENT HAD BLEEDING FROM MULTIPLE SITES. Manufacturer Narrative: B2 OUTCOME REVISED BASED ON THE ADDITIONAL INFORMATION RECEIVED FROM THE CLINICAL TO SITE TO INCLUDE DEATH AND DATE OF DEATH B5 REVISED BASED ON THE ADDITIONAL INFORMATION RECEIVED FROM THE CLINICAL TO SITE. D6B EXPLANT DATE ADDED H1 TYPE OF REPORT REVISED BASED ON THE ADDITIONAL INFORMATION RECEIVED FROM THE CLINICAL TO SITE. H6 REVISED BASED ON THE ADDITIONAL INFORMATION RECEIVED FROM THE CLINICAL TO SITE TO INCLUDE DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21525539&pc=OZD | 1220648-2025-26487 | 2025-02-04 05:00 | Death | ABIOMED, INC. | 2025-03-05 05:00 | OZD | IMPELLA CP | Increase in Pressure | Ischemia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. THE COMPLAINANT REPORTED THAT A PURGE PRESSURE HIGH ALARM PRESENTED. LYSIS PROTOCOL WAS SENT TO THE SITE TO MANAGE THE CONDITION. ADDITIONALLY, THE PHYSICIAN REPORTED THAT THE PATIENT'S RIGHT LEG HAD TO BE AMPUTATED DUE TO THE IMPELLA. THE PUMP WAS EXPLANTED A FEW DAYS LATER AND THE PATIENT EXPIRED. IT WAS NOTED THAT THE PATIENT'S TREATMENT WAS PALLIATIVE. THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA DEVICE AS AN ASSOCIATED FACTOR IN THE PATIENT'S DEMISE. Manufacturer Narrative: THE INVESTIGATION OF HIGH PURGE PRESSURE AND LIMB ISCHEMIA HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. THE ROOT CAUSE OF THE HIGH PURGE PRESSURE WAS MOST LIKELY DUE TO BIOMATERIAL AS LYSIS PROTOCOLS WERE ABLE TO RESOLVE ISSUE ACCORDING TO CLINICAL DESCRIPTION. AS ALL THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE LIMB ISCHEMIA WAS NOT DETERMINED. B1 SELECTION OF PRODUCT PROBLEM WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26487. B2 REQUIRED INTERVENTION SELECTION WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26487. E4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26487 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. G1 REPORTING CONTACT EMAIL REVISED ON MANUFACTURER DEVICE REPORT 1220648-2025-26487 IN ACCORDANCE WITH UPDATED PROCEDURES. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA IFU: IMPELLA CP WITH SMARTASSIST SYSTEM: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): "ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21525594&pc=OZD | 1220648-2025-26491 | 2025-02-06 05:00 | Death | ABIOMED, INC. | 2025-03-05 05:00 | OZD | IMPELLA CP | Malposition of Device | Hemolysis | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT IN CARDIOGENIC SHOCK FROM AN ACUTE MYOCARDIAL INFARCTION IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS) DEVELOPED HEMOLYSIS AND WOULD LATER EXPIRE. THE PATIENT WAS BROUGHT IN BY AMBULANCE AFTER BEING DOWN WITH VENTRICULAR FIBRILLATION ARREST FOR APPROXIMATELY TEN MINUTES. CARDIOPULMONARY RESUSCITATION WAS INITIATED AND PATIENT SHOCKED THREE TIMES WITH RETURN OF SPONTANEOUS CIRCULATION ACHIEVED. THE PATIENT WAS INTUBATED IN FIELD AND THREE PUSHES OF EPINEPHRINE WAS GIVEN; THE PATIENT WAS BROUGHT TO CARDIOVASCULAR LAB ON 1MCG AMIODARONE DRIP. ANGIOGRAPHY REVEALED MULTI-VESSEL DISEASE INCLUDING CHRONIC TOTAL OCCLUSION OF THE LEFT ANTERIOR DESCENDING ARTERY. ULTRASOUND GUIDED RFA ACCESS OBTAINED AND 14FR MEDTRONIC SENTRANT SHEATH PLACED. IMPELLA PREPPED AND PLACED THROUGH MEDTRONIC SHEATH AND STARTED WITHOUT COMPLICATIONS, REACHING MEAN FLOWS OF 3.6L/MIN. NO PERCUTANEOUS CORONARY INTERVENTION WAS PERFORMED; THE PATIENT WAS REFERRED TO SURGICAL CONSULT. REPOSITIONING SHEATH WAS ADVANCED INTO MEDTRONIC SHEATH AND SUTURED TO LEG. THE PATIENT WAS TRANSPORTED TO THE UNIT ON IMPELLA MCS. ONCE ON THE UNIT IT WAS REPORTED LACTATE WAS TRENDING DOWN AND URINE OUTPUT WAS DARKENING. THE PHYSICIAN ACKNOWLEDGED THE PUMP WAS TOO DEEP AND DID NOT WANT TO REPOSITION AT SUCH TIME AS THE PUMP WAS ADEQUATELY FUNCTIONING. THIS DAY THE PATIENT WAS MADE A DO NOT RESUSCITATE (DNR) AND CARE PLANS WITH FAMILY WERE DISCUSSED. THE FOLLOWING DAY THE PATIENT WAS REPORTED TO BE STABLE BUT LABS X4 HEMOLYZED. PERFORMANCE LEVEL WAS DECREASED TO P-7 FROM P-9 AND POSITION CHECK UNDER ECHOCARDIOGRAM WAS RECOMMENDED. THE STATUS OF THE PATIENT FOLLOWING THE EVENT WAS NOTED TO BE UNKNOWN. HOWEVER, LATER THIS DAY, THE PATIENT WOULD EXPIRE. THE FAMILY DECIDED TO WITHDRAW CARE; ALL THERAPIES WERE DISCONTINUED AND THE PATIENT PASSED AWAY AFTER 16 TOTAL HOURS OF IMPELLA MCS. THE CAUSE OF DEATH IS UNKNOWN. THE IMPELLA RELATED EVENT MAY HAVE POSSIBLY ADDED TO AN ALREADY COMPLEX SITUATION. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: IMPELLA CP WITH SMARTASSIST SYSTEM. SECTION: POTENTIAL ADVERSE EVENTS: ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY¿ (INCLUDING VENTRICULAR PERFORATION). SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER: ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE: ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS: ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS.¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION: ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21525833&pc=OZD | 1220648-2025-26504 | 2025-02-08 05:00 | Death | ABIOMED, INC. | 2025-03-05 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Atrial Fibrillation; Tachycardia | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS INITIALLY IMPLANTED WITH AN IMPELLA CP VIA THE RIGHT FEMORAL ARTERY FOR MECHANICAL CIRCULATORY SUPPORT (MCS). FOUR DAYS INTO THE IMPELLA CP DEVICE SUPPORT, THE PATIENT DEVELOPED LIMB ISCHEMIA, AND THE PATIENT WAS PLANNED FOR ESCALATION TO AN IMPELLA 5.5 DEVICE IN RESPONSE TO THE CONDITION. THE ATTEMPT WITH THE IMPELLA 5.5 WAS ULTIMATELY UNSUCCESSFUL AND A NEW IMPELLA CP WAS IMPLANTED VIA THE AXILLARY ARTERY. TWO DAYS AFTER THE START OF SUPPORT ON THE AXILLARY IMPELLA CP DEVICE, THE PATIENT WAS NOW IN ATRIAL FIBRILLATION WITH RAPID VENTRICULAR RESPONSE AT RATE IN 170¿S. IT WAS NOTED THE PATIENT HAD BEEN CARDIOVERTED FOR UNSTABLE VENTRICULAR TACHYCARDIA. ECHOCARDIOGRAM WAS COMPLETED AND PHYSICIAN HAD PULLED BACK THE IMPELLA CP DEVICE 2CM BECAUSE IT WAS MEASURING GREATER THAN 6 CM AND WAS UP AGAINST THE WALL. ECHOCARDIOGRAM POST REPOSITIONING NOTED THE IMPELLA WAS MEASURING ABOUT 3.8 AND WAS FREE FLOATING IN THE LEFT VENTRICLE, AND NOT UP AGAINST ANY STRUCTURE. THE PATIENT CONTINUED TO GO IN AND OUT OF RAPID ATRIAL FIBRILLATION, HOWEVER. A COUPLE OF DAYS FOLLOWING, THE PATIENT STILL REQUIRED VASOPRESSORS AND INOTROPES TO MAINTAIN HEMODYNAMIC STABILITY. LACTIC NORMALIZING, HOWEVER, THE QUESTION REMAINED HOW TO ADDRESS LONG TERM SUPPORT FOR THE PATIENT AS A BRIDGE TO TRANSPLANT OR OTHER DECISION DUE TO THE INABILITY TO PLACE IMPELLA 5.5. DISCUSSION MADE WITH THE FAMILY AND CARE WAS WITHDRAWN AFTER A TOTAL OF SIX DAYS ON SUPPORT. THE CAUSE OF THE DEATH WAS NOT PROVIDED. MEDICAL SAFETY REVIEW NOTED THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE AXILLARY IMPELLA CP AS A POSSIBLE CONTRIBUTING FACTOR TO THE OUTCOME. ALTHOUGH IT SHOULD BE NOTED LONG-TERM SUSTAINING TREATMENT OPTIONS TO BRIDGE WERE LIMITED, AND THE FAMILY WITHDREW CARE FROM THE PATIENT. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THIS IS ONE OF THREE DEVICE MANUFACTURING REPORTS RELATED TO THE PATIENT'S IMPLANT EXPERIENCE. REFER TO THE FOLLOWING: MEDICAL DEVICE REPORT FOR IMPELLA CP SERIAL NUMBER (B)(6) WITH DATE OF EVENT 02-FEB-2025 AND PATIENT IDENTIFIER ENDING IN (B)(6). MEDICAL DEVICE REPORT FOR IMPELLA 5.5 SERIAL NUMBER (B)(6) WITH DATE OF EVENT 06-FEB-2025 AND PATIENT IDENTIFIER ENDING IN (B)(6). MEDICAL DEVICE REPORT FOR IMPELLA CP SERIAL NUMBER (B)(6) WITH DATE OF EVENT 08-FEB-2025 AND PATIENT IDENTIFIER ENDING IN (B)(6) (THIS REPORT).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21525898&pc=OZD | 1220648-2025-26509 | 2025-02-10 05:00 | Death | ABIOMED, INC. | 2025-03-05 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemolysis | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED 65-YEAR-OLD FEMALE PATIENT IMPLANTED WITH A CP PUMP FOR SUPPORT DUE TO CARDIOGENIC SHOCK/CARDIOMYOPATHY. THE PATIENT HAD SIGNS OF HEMOLYSIS BUT THE TEAM MADE A CHOICE TO EVALUATE PATIENT BEFORE REPOSITIONING OR OTHER TROUBLESHOOTING MEASURES. THE PUMP REMAINED ON WITH THE HEMOLYSIS UNTIL THE FAMILY MADE A DECISION TO WITHDRAW CARE AND PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. IMPELLA CP WITH SMARTASSIST SYSTEM, SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER: ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE: ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS: ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION: ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿ Manufacturer Narrative: THE INVESTIGATION OF HEMOLYSIS HAS BEEN COMPLETED. THE DEVICE AND THE DATA WERE NOT RETURNED FOR REVIEW. BASED ON CLINICAL DESCRIPTION, THE ROOT CAUSE OF HEMOLYSIS WAS MOST LIKELY DUE TO PUMP WAS NOT IN THE PROPER POSITION. E4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26509 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. G1 REPORTING CONTACT FACILITY NAME AND FAX NUMBER WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26509.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21525918&pc=OZD | 1220648-2025-26511 | 2025-02-07 05:00 | Death | ABIOMED, INC. | 2025-03-05 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT A PATIENT HAD AN IMPELLA CP AND EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) SUPPORT ONGOING FOR CARDIOGENIC SHOCK. THE PUMP WAS SUPPORTING FOR SEVEN HOURS WHEN THE PATIENT EXPIRED FROM BLEEDING THAT ORIGINATED AT THE ECMO SITE. THE CONTRIBUTION OF THE HEPARIN IN THE IMPELLA PURGE AND IMPELLA CP USE AND ANY CONTRIBUTION TO THE DEATH IS NOT KNOWN. THE BLEEDING MADE STABILIZATION OF THE PATIENT DIFFICULT AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21525990&pc=NGV | 2124215-2025-13304 | 2025-02-26 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-03-05 05:00 | NGV | WATCHMAN FLX PRO | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Thrombosis/Thrombus | P130013 |  | 1 | Event Description: IT WAS REPORTED THROMBOSIS AND CEREBRAL VASCULAR ACCIDENT OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS PERFORMED AND A 31MM WATCHMAN FLX PRO CLOSURE DEVICE WAS IMPLANTED AFTER MEETING ALL RELEASE CRITERIA, WITH COMPLETE SEAL NOTED, AND WITHOUT ANY COMPLICATIONS. THE PROCEDURE WAS CONCLUDED. THE PATIENT WAS SENT TO THE RECOVERY AREA AND THEN EXPERIENCED A STROKE. IMAGING PERFORMED FOUND A THROMBUS IN ONE OF THE MAJOR VESSELS IN THE NECK. POST STROKE HE WAS RESPONDING TO NEUROLOGICAL ASSESSMENT AND WAS ADMITTED INTO INPATIENT INTENSIVE CARE UNIT (ICU) FOR MONITORING. THE PATIENT IS EXPECTED TO MAKE A FULL RECOVERY WITH REHABILITATION.Event Description: IT WAS REPORTED THROMBOSIS AND CEREBRAL VASCULAR ACCIDENT OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS PERFORMED AND A 31MM WATCHMAN FLX PRO CLOSURE DEVICE WAS IMPLANTED AFTER MEETING ALL RELEASE CRITERIA, WITH COMPLETE SEAL NOTED, AND WITHOUT ANY COMPLICATIONS. THE PROCEDURE WAS CONCLUDED. THE PATIENT WAS SENT TO THE RECOVERY AREA AND THEN EXPERIENCED A STROKE. IMAGING PERFORMED FOUND A THROMBUS IN ONE OF THE MAJOR VESSELS IN THE NECK. POST STROKE HE WAS RESPONDING TO NEUROLOGICAL ASSESSMENT AND WAS ADMITTED INTO INPATIENT INTENSIVE CARE UNIT (ICU) FOR MONITORING. THE PATIENT IS EXPECTED TO MAKE A FULL RECOVERY WITH REHABILITATION. IT WAS FURTHER REPORTED THE PATIENT SYMPTOMS WHERE THEY WERE NOT EASILY AROUSABLE A FEW HOURS POST INDEX PROCEDURE WHILE IN THE RECOVERY AREA. THE PATIENT REQUIRED A THROMBECTOMY PROCEDURE IN RESPONSE TO THE STROKE. THE PATIENT DIED 24 HOURS AFTER STROKE INTERVENTION. Manufacturer Narrative: B2: DATE OF DEATH IS ESTIMATED. B5: INFORMATION ADDED. H1: UPDATED FROM SERIOUS INJURY TO DEATH. H6 PATIENT CODE ADDED: DECREASED LEVEL OF CONSCIOUSNESS. H6 IMPACT CODES ADDED: UNEXPECTED MEDICAL INTERVENTION AND DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21526023&pc=OZD | 1220648-2025-26514 | 2025-02-08 05:00 | Death | ABIOMED, INC. | 2025-03-05 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. DURING SUPPORT, THE PATIENT DEVELOPED LIMB ISCHEMIA. THE DECISION WAS MADE TO CONSULT VASCULAR SURGERY TO ASSIST WITH A RIGHT SUPERFICIAL FEMORAL ARTERY DISTAL PERFUSION SHEATH. LEFT TO RIGHT EXTERNAL FEMORO-FEMORAL BYPASS WAS PERFORMED SUCCESSFULLY WITH IMMEDIATE IMPROVEMENT IN PALLOR OF THE RIGHT LOWER EXTREMITY. SUPPORT WAS CONTINUED. ADDITIONAL INFORMATION NOTED THAT THE DECISION WAS MADE TO WITHDRAW CARE AND THE PATIENT EXPIRED. THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA CP DEVICE AS AN ASSOCIATED FACTOR IN THE PATIENT'S DEMISE. Manufacturer Narrative: THE INVESTIGATION OF LIMB ISCHEMIA HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM, SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21526402&pc=OAE | 3008452825-2025-00097 | 2025-01-06 05:00 | Death | ST. JUDE MEDICAL | 2025-03-05 05:00 | OAE | TACTIFLEX¿ ABLATION CATHETER, SENSOR ENABLED¿ | Adverse Event Without Identified Device or Use Problem | Perforation of Esophagus | P220013 |  | 1 | Event Description: FOLLOWING A SUCCESSFUL EPIC RESEARCH ATRIAL FIBRILLATION PROCEDURE PERFORMED ON (B)(6) 2025, THE PATIENT IS ALLEGED TO HAVE DEVELOPED AN ESOPHAGEAL FISTULA RESULTING IN DEATH ON (B)(6) 2025. DURING THE PROCEDURE THERE WERE NO ALLEGED PERFORMANCE ISSUES AND PATIENT HAD BEEN DISCHARGED. Manufacturer Narrative: THE PRODUCT'S LOT NUMBER COULD NOT BE OBTAINED; THEREFORE THE PRIMARY DI NUMBER IS PROVIDED (D4), BUT FULL UDI INFORMATION IS NOT AVAILABLE. Manufacturer Narrative: THE RESULTS OF THE INVESTIGATION ARE INCONCLUSIVE SINCE THE DEVICE WAS NOT RETURNED FOR ANALYSIS. HOWEVER, THE CASE STUDY AND TACTISYS LOG FILES WERE RETURNED. LOG FILE AND CASE STUDY ANALYSIS CONCLUDED THAT THE CATHETER PERFORMED AS INTENDED. THE LOG FILES SHOW THAT THE OPTICAL FIBERS MET SPECIFICATIONS, THE RECORDED TEMPERATURES INDICATED COOLING DURING RF ABLATION, AND CONTACT FORCE MEASUREMENTS WERE DISPLAYED THROUGHOUT THE DURATION OF THE FILES. NO POWER, TEMPERATURE, CONTACT FORCE, OR IMPEDANCE ANOMALIES WERE NOTED DURING THE CASE STUDY REVIEW. FOR ABLATIONS ON THE POSTERIOR WALL OF THE HEART, THE CONTACT FORCE ARROW DID NOT APPEAR TO BE POINTED INTO THE EPICARDIAL SPACE. THE CASE STUDY INDICATED THAT EPICARDIAL LESIONS WERE PERFORMED ON BOTH THE ANTERIOR AND POSTERIOR SURFACES OF THE HEART. THE ABLATIONS PERFORMED ON THE POSTERIOR WALL OF THE HEART WERE WITHIN THE RECOMMENDED POWER AND DURATION SETTINGS. HOWEVER, ABLATIONS ON THE ANTERIOR SURFACE OF THE HEART WERE NOTED TO BE UP TO 20 SECONDS AT 50W. THE TACTIFLEX ABLATION CATHETER, SENSOR ENABLED INSTRUCTIONS FOR USE (IFU) STATES ¿THE CONSECUTIVE DURATION OF AN ABLATION IN A GIVEN LOCATION (FOCAL AND DRAG LESIONS) SHOULD NOT EXCEED THE RECOMMENDED TIME (E.G. AT 50W, THE CONSECUTIVE SECONDS OF ABLATION ENERGY DELIVERY AT A SPECIFIC SITE SHOULD BE 10 SECONDS OR LESS).¿ ADDITIONALLY, ABLATIONS TO THE LEFT PULMONARY VEIN (LPV) AND THE ROOF AND FLOOR OF THE POSTERIOR WALL WERE PERFORMED UP TO 45W. THE IFU WARNS ¿APPLICATION OF RF ENERGY ON THE LEFT ATRIAL POSTERIOR WALL EXCEEDING 40W IN POWER, OR USE OF CONTACT FORCE =15 G, INCREASES THE RISK OF ESOPHAGEAL PERFORATING COMPLICATIONS INCLUDING ATRIO-ESOPHAGEAL FISTULA AND DEATH.¿ REVIEW OF THE DEVICE HISTORY RECORD WAS NOT POSSIBLE AS THE LOT NUMBER IS UNKNOWN. DUE TO UNKNOWN PROCEDURAL CONDITIONS, WE ARE UNABLE TO CONCLUSIVELY DETERMINE THE CAUSE OF THE ALLEGED ESOPHAGEAL FISTULA AND SUBSEQUENT DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21526438&pc=NPT | 2025587-2025-01648 | 2024-06-24 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-05 05:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Insufficient Information | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 9 MONTHS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT PASSED AWAY. THE CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21526626&pc=LWS | 9614453-2025-00911 | 2025-02-17 05:00 | Death | MEDTRONIC EUROPE SARL | 2025-03-05 05:00 | LWS | COBALT¿ XT DR MRI SURESCAN¿ | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Tachycardia; Ventricular Fibrillation; Asystole | P980016 |  | 1 | Event Description: IT WAS REPORTED THAT THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) DETECTED AN ATRIAL TACHYCARDIA AND APPROPRIATELY DELIVERED ATRIAL ANTI-TACHYCARDIA PACING (A-ATP). SUBSEQUENTLY, VENTRICULAR TACHYCARDIA (VT) OCCURRED DURING AN A-ATP EPISODE, AND THE RHYTHM SWITCHED TO VENTRICULAR FIBRILLATION (VF) AND WAS DETECTED; HOWEVER, THE INTERVAL OF THE VENTRICULAR EVENT WAS EXTENDED DURING CHARGING, AND THE THERAPY WAS SUSPENDED. FIFTY MINUTES AFTER THE TIME OF THE FIRST EPISODE, THE PATIENT WAS TRANSPORTED TO THE HOSPITAL WITH CARDIAC ARREST WHERE THE PATIENT EXPERIENCED ASYSTOLE AND PASSED AWAY. IT WAS NOTED THAT THE ICD FUNCTION WAS CONSIDERED "NORMAL OPERATION". Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21526756&pc=NVZ | 2182208-2025-01131 | 2024-08-12 04:00 | Death | MEDTRONIC, INC. | 2025-03-05 05:00 | NVZ | MEDTRONIC IPG | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Sepsis |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING CARDIOVASCULAR IMPLANTABLE ELECTRONIC DEVICE (CIED) EXTRACTION AND CONCOMITANT LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) IMPLANTATION. THE CAUSES FOR CIED EXTRACTIONS WERE DUE TO POCKET INFECTIONS, LEAD INFECTIONS/ENDOCARDITIS, POCKET PERFORATIONS, SEVERE TRICUSPID REGURGITATION DUE TO PACEMAKER LEAD, CHRONIC PAIN DUE TO THE PACEMAKER, AND UNKNOWN LEAD FAILURES. MICROBIOLOGICAL CULTURES OF THE EXTRACTED DEVICES AND BLOOD CULTURES WERE POSITIVE IN SOME PATIENTS. THE MOST FREQUENT PATHOGENS WERE STAPHYLOCOCCUS SPECIES (STAPHYLOCOCCUS AUREUS, METHICILLIN-RESISTANT, EPIDERMIDIS, HOMINIS, HEMOLYTICUS), BETA-HEMOLYTIC STREPTOCOCCUS, PSEUDOMONAS AERUGINOSA, ARTHROBACTER SPECIES, AND CO-INFECTION WITH CANDIDA ALBICANS. THE AUTHORS DESCRIBED PATIENT DEATHS; ONE PATIENT EXPERIENCED CARDIAC ARREST DUE TO SEVERE SEPSIS WHICH OCCURRED IMMEDIATELY AFTER CIED EXTRACTION AND DIED BEFORE LEADLESS IMPLANTATION. THERE WAS ANOTHER PATIENT THAT DIED DUE TO SEPSIS, AND ANOTHER DUE TO A COMPLICATED PULMONARY ABSCESS; BOTH OCCURRED WITHIN THIRTY DAYS OF CIED EXTRACTION WITH LEADLESS IMPLANTATION. OTHER PATIENTS' CAUSES OF DEATH INCLUDED MYOCARDIAL INFARCTION, SEPSIS, PULMONARY EDEMA AND SEVERE TRICUSPID REGURGITATION. DURING EXTRACTION PROCEDURES THERE WERE PATIENTS WHO EXPERIENCED PERICARDIAL EFFUSION WITH PERICARDIOCENTESIS, AND HEMATOMA WHICH REQUIRED REOPERATION. THE STATUS OF THE DEVICES AND LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/79 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: FEASIBILITY, TIMING AND OUTCOME OF LEADLESS CARDIAC PACEMAKER IMPLANTATION IN PATIENTS UNDERGOING CARDIAC IMPLANTABLE ELECTRONIC DEVICE EXTRACTION. CLINICAL RESEARCH IN CARDIOLOGY. 2024. 1-11. DOI: 10.1007/S00392-024-02516-0. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21526767&pc=DXY | 2182208-2025-01132 | 2024-08-12 04:00 | Death | MEDTRONIC, INC. | 2025-03-05 05:00 | DXY | MEDTRONIC LEAD | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Sepsis |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING CARDIOVASCULAR IMPLANTABLE ELECTRONIC DEVICE (CIED) EXTRACTION AND CONCOMITANT LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) IMPLANTATION. THE CAUSES FOR CIED EXTRACTIONS WERE DUE TO POCKET INFECTIONS, LEAD INFECTIONS/ENDOCARDITIS, POCKET PERFORATIONS, SEVERE TRICUSPID REGURGITATION DUE TO PACEMAKER LEAD, CHRONIC PAIN DUE TO THE PACEMAKER, AND UNKNOWN LEAD FAILURES. MICROBIOLOGICAL CULTURES OF THE EXTRACTED DEVICES AND BLOOD CULTURES WERE POSITIVE IN SOME PATIENTS. THE MOST FREQUENT PATHOGENS WERE STAPHYLOCOCCUS SPECIES (STAPHYLOCOCCUS AUREUS, METHICILLIN-RESISTANT, EPIDERMIDIS, HOMINIS, HEMOLYTICUS), BETA-HEMOLYTIC STREPTOCOCCUS, PSEUDOMONAS AERUGINOSA, ARTHROBACTER SPECIES, AND CO-INFECTION WITH CANDIDA ALBICANS. THE AUTHORS DESCRIBED PATIENT DEATHS; ONE PATIENT EXPERIENCED CARDIAC ARREST DUE TO SEVERE SEPSIS WHICH OCCURRED IMMEDIATELY AFTER CIED EXTRACTION AND DIED BEFORE LEADLESS IMPLANTATION. THERE WAS ANOTHER PATIENT THAT DIED DUE TO SEPSIS, AND ANOTHER DUE TO A COMPLICATED PULMONARY ABSCESS; BOTH OCCURRED WITHIN THIRTY DAYS OF CIED EXTRACTION WITH LEADLESS IMPLANTATION. OTHER PATIENTS' CAUSES OF DEATH INCLUDED MYOCARDIAL INFARCTION, SEPSIS, PULMONARY EDEMA AND SEVERE TRICUSPID REGURGITATION. DURING EXTRACTION PROCEDURES THERE WERE PATIENTS WHO EXPERIENCED PERICARDIAL EFFUSION WITH PERICARDIOCENTESIS, AND HEMATOMA WHICH REQUIRED REOPERATION. THE STATUS OF THE DEVICES AND LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/79 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: FEASIBILITY, TIMING AND OUTCOME OF LEADLESS CARDIAC PACEMAKER IMPLANTATION IN PATIENTS UNDERGOING CARDIAC IMPLANTABLE ELECTRONIC DEVICE EXTRACTION. CLINICAL RESEARCH IN CARDIOLOGY. 2024. 1-11. DOI: 10.1007/S00392-024-02516-0. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21526775&pc=PNJ | 9612164-2025-01169 | 2024-08-12 04:00 | Death | MEDTRONIC IRELAND | 2025-03-05 05:00 | PNJ | MICRA | Adverse Event Without Identified Device or Use Problem | Abscess; Myocardial Infarction; Pulmonary Edema; Sepsis; Tricuspid Valve Insufficiency/ Regurgitation |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING CARDIOVASCULAR IMPLANTABLE ELECTRONIC DEVICE (CIED) EXTRACTION AND CONCOMITANT LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) IMPLANTATION. THE CAUSES FOR CIED EXTRACTIONS WERE DUE TO POCKET INFECTIONS, LEAD INFECTIONS/ENDOCARDITIS, POCKET PERFORATIONS, SEVERE TRICUSPID REGURGITATION DUE TO PACEMAKER LEAD, CHRONIC PAIN DUE TO THE PACEMAKER, AND UNKNOWN LEAD FAILURES. MICROBIOLOGICAL CULTURES OF THE EXTRACTED DEVICES AND BLOOD CULTURES WERE POSITIVE IN SOME PATIENTS. THE MOST FREQUENT PATHOGENS WERE STAPHYLOCOCCUS SPECIES (STAPHYLOCOCCUS AUREUS, METHICILLIN-RESISTANT, EPIDERMIDIS, HOMINIS, HEMOLYTICUS), BETA-HEMOLYTIC STREPTOCOCCUS, PSEUDOMONAS AERUGINOSA, ARTHROBACTER SPECIES, AND CO-INFECTION WITH CANDIDA ALBICANS. THE AUTHORS DESCRIBED PATIENT DEATHS; ONE PATIENT EXPERIENCED CARDIAC ARREST DUE TO SEVERE SEPSIS WHICH OCCURRED IMMEDIATELY AFTER CIED EXTRACTION AND DIED BEFORE LEADLESS IMPLANTATION. THERE WAS ANOTHER PATIENT THAT DIED DUE TO SEPSIS, AND ANOTHER DUE TO A COMPLICATED PULMONARY ABSCESS; BOTH OCCURRED WITHIN THIRTY DAYS OF CIED EXTRACTION WITH LEADLESS IMPLANTATION. OTHER PATIENTS' CAUSES OF DEATH INCLUDED MYOCARDIAL INFARCTION, SEPSIS, PULMONARY EDEMA AND SEVERE TRICUSPID REGURGITATION. DURING EXTRACTION PROCEDURES THERE WERE PATIENTS WHO EXPERIENCED PERICARDIAL EFFUSION WITH PERICARDIOCENTESIS, AND HEMATOMA WHICH REQUIRED REOPERATION. THE STATUS OF THE DEVICES AND LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/79 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: FEASIBILITY, TIMING AND OUTCOME OF LEADLESS CARDIAC PACEMAKER IMPLANTATION IN PATIENTS UNDERGOING CARDIAC IMPLANTABLE ELECTRONIC DEVICE EXTRACTION. CLINICAL RESEARCH IN CARDIOLOGY. 2024. 1-11. DOI: 10.1007/S00392-024-02516-0. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21526851&pc=BTR | 3000219639-2025-00027 | 2025-01-28 05:00 | Death | PARKER MEDICAL | 2025-03-05 05:00 | BTR | PARKER | Gas/Air Leak | Low Oxygen Saturation | N/A |  | 1 | Event Description: IT WAS REPORTED THAT THE CUFF WAS MISALIGNED TOWARD THE TIP, EXPOSING THE AIRFLOW HOLE IN THE INFLATING LINE, WHICH IS NORMALLY LOCATED INSIDE THE CUFF, AND THAT AN AIR LEAK HAD OCCURRED AT THIS POINT. A PATIENT (80-YEAR-OLD MALE) WHO WAS INTUBATED WITH THIS PRODUCT ON JANUARY 14 DEVELOPED VOICE LEAKAGE (AIR LEAK) IMMEDIATELY AFTER POSITIONAL CHANGE ON JANUARY 19, AND SINCE THE LEAKAGE DID NOT IMPROVE EVEN AFTER AIR WAS SUPPLIED TO THE CUFF, THE TUBE WAS REPLACED. WHEN AIR WAS SUPPLIED TO THE PRODUCT AFTER EXTUBATION, IT WAS CONFIRMED THAT AIR COULD NO LONGER ENTER THE CUFF. THE PATIENT SUBSEQUENTLY DIED. HOWEVER, ACCORDING TO THE HOSPITAL, THERE WAS NO DIRECT CAUSAL RELATIONSHIP BETWEEN THE LEAK AND THE DEATH.Event Description: IT WAS REPORTED THAT THE CUFF WAS MISALIGNED TOWARD THE TIP, EXPOSING THE AIRFLOW HOLE IN THE INFLATING LINE, WHICH IS NORMALLY LOCATED INSIDE THE CUFF, AND THAT AN AIR LEAK HAD OCCURRED AT THIS POINT. A PATIENT (80-YEAR-OLD MALE) WHO WAS INTUBATED WITH THIS PRODUCT ON JANUARY 14 DEVELOPED VOICE LEAKAGE (AIR LEAK) IMMEDIATELY AFTER POSITIONAL CHANGE ON JANUARY 19, AND SINCE THE LEAKAGE DID NOT IMPROVE EVEN AFTER AIR WAS SUPPLIED TO THE CUFF, THE TUBE WAS REPLACED. WHEN AIR WAS SUPPLIED TO THE PRODUCT AFTER EXTUBATION, IT WAS CONFIRMED THAT AIR COULD NO LONGER ENTER THE CUFF. THE PATIENT SUBSEQUENTLY DIED. HOWEVER, ACCORDING TO THE HOSPITAL, THERE WAS NO DIRECT CAUSAL RELATIONSHIP BETWEEN THE LEAK AND THE DEATH. Manufacturer Narrative: THE PRODUCT INVOLVED IN THE REPORT HAS NOT BEEN RETURNED. A REVIEW OF THE DEVICE HISTORY RECORD IS IN-PROGRESS. ALL INFORMATION REASONABLY KNOWN AS OF 05 MARCH 2025 HAS BEEN INCLUDED IN THE HEALTH AUTHORITY REPORT. SHOULD ADDITIONAL INFORMATION BE OBTAINED A FOLLOW-UP HEALTH AUTHORITY REPORT WILL BE PROVIDED. THE INFORMATION PROVIDED BY AIRLIFE. REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO AIRLIFE. AIRLIFE HAS NO INDEPENDENT KNOWLEDGE OF THE EVENT REPORTED BUT IS RELAYING THE INFORMATION THAT WAS PROVIDED BY THE USER FACILITY WHERE THE INCIDENT OCCURRED. THIS PRODUCT INCIDENT IS DOCUMENTED IN THE AIRLIFE COMPLAINT DATABASE AND IS IDENTIFIED AS COMPLAINT (B)(4). THIS INFORMATION IS SUBMITTED PURSUANT TO 21CFR803, IN COMPLIANCE WITH THE MEDICAL DEVICE REPORTING REQUIREMENT AND SHOULD NOT BE CONSIDERED TO BE AN ADMISSION THAT A AIRLIFE PRODUCT IS DEFECTIVE OR CAUSES SERIOUS INJURY. Manufacturer Narrative: THE PRODUCT INVOLVED IN THE REPORT HAS NOT BEEN RETURNED. A REVIEW OF THE DEVICE HISTORY RECORD IS IN-PROGRESS. ALL INFORMATION REASONABLY KNOWN AS OF 05 MARCH 2025 HAS BEEN INCLUDED IN THE HEALTH AUTHORITY REPORT. SHOULD ADDITIONAL INFORMATION BE OBTAINED A FOLLOW-UP HEALTH AUTHORITY REPORT WILL BE PROVIDED. THE INFORMATION PROVIDED BY AIRLIFE. REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO AIRLIFE. AIRLIFE HAS NO INDEPENDENT KNOWLEDGE OF THE EVENT REPORTED BUT IS RELAYING THE INFORMATION THAT WAS PROVIDED BY THE USER FACILITY WHERE THE INCIDENT OCCURRED. THIS PRODUCT INCIDENT IS DOCUMENTED IN THE AIRLIFE COMPLAINT DATABASE AND IS IDENTIFIED AS COMPLAINT (B)(4). THIS INFORMATION IS SUBMITTED PURSUANT TO 21CFR803, IN COMPLIANCE WITH THE MEDICAL DEVICE REPORTING REQUIREMENT AND SHOULD NOT BE CONSIDERED TO BE AN ADMISSION THAT A AIRLIFE PRODUCT IS DEFECTIVE OR CAUSES SERIOUS INJURY. COMPLAINT WAS CONSIDERED CONFIRMED BASED ON CUSTOMER NARRATIVE OF EVENTS. A 24 MONTH REVIEW WAS CONDUCTED AND 0 ADDITIONAL COMPLAINTS WERE IDENTIFIED RELATED TO THIS ISSUE. NO TREND WAS OBSERVED. THE LOT NUMBER WAS NOT PROVIDED, AND THE VENDOR WAS NOTIFIED OF THE INCIDENT. THIS INCIDENT WAS NOT IDENTIFIED AS AN INCREASING TREND. NO FURTHER ACTION TO BE TAKEN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21526870&pc=DSQ | 3007042319-2025-00609 | 2025-02-06 05:00 | Death | HEARTWARE, INC. | 2025-03-05 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Sepsis | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WITH A HISTORY OF HEART FAILURE, WHO HAD A VENTRICULAR ASSIST DEVICE (VAD), DEVELOPED A DRIVELINE (DL) INFECTION. THE PATIENT WAS HOSPITALIZED FOR THE INFECTION AND SUBSEQUENTLY DIED OF SEPSIS. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR INVESTIGATION SUMMARY. PRODUCT EVENT SUMMARY: THE VENTRICULAR ASSIST DEVICE (VAD) (B)(6) WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE MANUFACTURING DOCUMENTATION WAS NOT PERFORMED AS THE REPORTED EVENT AND ANALYSIS ARE NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT WAS HOSPITALIZED FOR A DRIVELINE INFECTION AND SUBSEQUENTLY DIED OF SEPSIS. BASED ON THE LIMITED INFORMATION AVAILABLE, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, DRIVELINE INFECTION, SEPSIS, AND DEATH ARE KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. BASED ON A REVIEW OF PAST ADVERSE EVENTS FOR THIS PATIENT, IT WAS NOTED THE PATIENT HAD A HISTORY OF DRIVELINE INFECTION. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE, INCLUDING CARE OF THE DRIVELINE EXIT SITE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21527039&pc=DSQ | 3007042319-2025-00612 | 2025-02-28 05:00 | Death | HEARTWARE, INC. | 2025-03-05 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Power Problem; Pumping Problem | Thrombosis/Thrombus | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WITH A HISTORY OF CANCER, WHO HAD A VENTRICULAR ASSIST DEVICE (VAD) IMPLANT, WAS ADMITTED FOR AN EXTENDED PERIOD AND WAS IN THE END-OF-LIFE STAGE. THE PATIENT EXPERIENCED A SUSPECTED THROMBUS, ELEVATED POWER LEVELS, AND INCREASED LACTIC ACID DEHYDROGENASE (LDH). REVIEW OF LOG FILES DATA CONFIRMED ABOVE NORMAL POWER CONSUMPTION. LOG FILES DATA ALSO SHOWED A MOTOR START EVENT WITH PARAMETERS OUTSIDE OF THE TYPICAL RANGE THAT OCCURRED FOLLOWING A LOSS OF POWER TO THE CONTROLLER. ULTIMATELY, SUPPORT WAS WITHDRAWN AND THE PATIENT PASSED AWAY. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR ANALYSIS AND INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: THE VENTRICULAR ASSIST DEVICE (VAD) (B)(6) AND THE CONTROLLER (B)(6) WERE NOT RETURNED FOR EVALUATION AS THE VAD REMAINS IN PA TIENT AND THE CONTROLLER DEVICE LOCATION IS UNKNOWN. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT AND ANALYSIS ARE NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. LOG FILE ANALYSIS REVEALED A MOTOR START EVENT RECORDED ON (B)(6) 2025 AT 09:26:53, IN WHICH THE MOTOR START PARAMETERS WERE ATYPICAL. IN ADDITION, LOG FILES REVEALED AN INCREASE IN POWER CONSUMPTION AND ESTIMATED FLOW STARTING ON (B)(6) 2025, LEADING TO PARAMETERS ABOVE THE NORMAL OPERATING RANGE. LOG FILE ANALYSIS REVEALED ONE (1) CONTROLLER POWER-UP WITH ASSOCIATED MOTOR START EVENT LOGGED ON (B)(6) ON (B)(6) 2025 AT 09:26:44, INDICATING A LOSS OF POWER TO THE CONTROLLER. THE DATA POINT PRIOR TO THE LOSS OF POWER REVEALED THAT (B)(6) WAS CONNECTED TO POWER PORT ONE (1) WITH 79% RELATIVE STATE OF CHARGE (RSOC) AND (B)(6) WAS CONNECTED TO POWER PORT TWO (2) WITH 24% RSOC. THE DATA POINT RECORDED AFTER THE LOSS OF POWER REVEALED THAT A POWER ADAPTER WAS CONNECTED TO POWER PORT ONE (1) AND (B)(6) WAS CONNECTED TO POWER PORT TWO (2). THE CONTROLLER WAS WITHOUT POWER FOR FOURTEEN (14) SECONDS. NO ANOMALIES WERE OBSERVED LEADING UP TO THE EVENT. AS A RESULT, THE REPORTED HIGH-POWER EVENT, ATYPICAL MOTOR START PARAMETERS EVENT, AND CONTROLLER LOSS OF POWER EVENT WERE CONFIRMED. BASED ON THE AVAILABLE INFORMATION AND HISTORICAL REVIEW OF SIMILAR EVENTS, THE MOST LIKELY ROOT CAUSE OF THE HIGH-POWER EVENT CAN BE ATTRIBUTED TO EXTERNAL FACTORS SUCH AS THROMBUS FORMATION/INGESTION. THE MOST LIKELY ROOT CAUSE OF THE LOSS OF POWER TO THE CONTROLLER CAN BE ATTRIBUTED TO A DISCONNECTION OF BOTH POWER SOURCES AND/OR TO AN INTERMITTENT DISCONNECTION ON ONE OR BOTH POWER SOURCES. CAPA PR00551638 INVESTIGATED CONTROLLER LOSSES OF POWER. EVEN THOUGH THIS CAPA IS NOW CLOSED, (B)(6) FALLS IN SCOPE OF THIS CAPA. BASED ON THE RISK DOCUMENTATION, THE MOST LIKELY ROOT CAUSE OF ATYPICAL MOTOR START PARAMETERS MAY BE ATTRIBUTED, BUT NOT LIMITED, TO OUTER SHROUD CONTACT THAT CREATES MORE FRICTION AT THE HOUSING TO IMPELLER INTERFACE DURING PUMP STARTUP. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED HIGH-POWER EVENT. PER THE INSTRUCTIONS FOR USE, DEVICE THROMBUS AND DEATH ARE KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. THERE WAS NO EVIDENCE THAT THE PATIENT HAD A HISTORY OF THROMBUS EVENTS. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. ADDITIONAL PRODUCTS: D1: HEARTWARE VENTRICULAR ASSIST SYSTEM - 1420-CONTROLLER 2.0 D4: SERIAL#: (B)(6) H3: YES H6: THE CODES PRESENT IN SECTION H6 CORRESPOND TO COMPONENTS/PRODUCTS THAT COMPRISE THE REPORTED EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. ADDITIONAL PRODUCTS: D1: HEARTWARE VENTRICULAR ASSIST SYSTEM - CONTROLLER 2.0 D4: MODEL#: 1420 / CATALOG#: 1420 / EXPIRATION DATE: 30-NOV-2024 / SERIAL#: (B)(6), D9: NO, H3: NO, H4: MFG DATE: 17-NOV-2023, H5: NO. H6: THE CODES PRESENT IN SECTION H6 CORRESPOND TO COMPONENTS/PRODUCTS THAT COMPRISE THE REPORTED EVENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21527109&pc=JJE | 2122870-2025-00072 | 2025-02-06 05:00 | Death | BECKMAN COULTER, INC | 2025-03-05 05:00 | JJE | DXI 9000 | Missing Test Results | Insufficient Information |  |  | 1 | Event Description: ON 06FEB2025 THE CUSTOMER REPORTED THAT PATIENT SAMPLES WERE QUEUING ON THE AUTOMATION LINE TRACK BEHIND DXA 5000 INSTRUMENT SERIAL NUMBER (S/N) (B)(6) CONNECTED TO DXI 9000 INSTRUMENT S/N (B)(6). FIELD SERVICE ENGINEER (FSE) NOTICED THAT NO HARDWARE ERROR MESSAGES WERE GENERATED ON THE DXI 9000 INSTRUMENT INDICATING THAT SAMPLES WERE NOT BEING ASPIRATED. THIS ISSUE CAUSED A DELAY IN REPORTING PATIENT RESULTS. THE CUSTOMER REPORTED A PATIENT DEATH WHILE WAITING FOR THE RESULTS BEING PROCESSED. UPON THE INITIAL REPORT OF THE EVENT, THE CUSTOMER DID NOT ALLEGE WHETHER THE DELAY IN OBTAINING THE RESULTS CONTRIBUTED TO THE PATIENT'S DEATH. THE HSTNI RESULT FOR THIS PATIENT SAMPLE WAS ELEVATED (236.26 NG/L) COMPARED WITH THE EXPECTED VALUES FOR OVERALL POPULATION (99TH PERCENTILE URL (95% CL): 17.5 PG/ML (12.6 - 20.7 PG/ML)). PER DISCUSSION BETWEEN FIELD APPLICATION SPECIALIST (FAS) AND THE CUSTOMER ON 11FEB2025, THE CUSTOMER STATED THE PHYSICIANS WERE NOT WAITING ONLY FOR THE HSTNI RESULT TO DETERMINE A DIAGNOSIS. THE CUSTOMER DID NOT CONFIRM WHETHER THE DELAY IN OBTAINING THE RESULT LED TO A DELAY IN DIAGNOSIS OR TREATMENT; THE CUSTOMER ALSO DID NOT PROVIDE THE PATIENT¿S DIAGNOSIS AND DID NOT PROVIDE ANY ADDITIONAL INFORMATION REGARDING THE PATIENT¿S PRESENTATION OR HISTORY. Manufacturer Narrative: AN INTERNAL INVESTIGATION WAS PERFORMED AND SOFTWARE DESIGN ERRORS WERE IDENTIFIED. A FIELD ACTION (FA-25027) WILL ADDRESS THE ISSUE. H6: DUE TO SMARTSOLVE LIMITATIONS, THE EVALUATION CODE FOR INVESTIGATION CONCLUSION IS "4316: APPROPRIATE TERM/CODE NOT AVAILABLE". THE MOST ACCURATE CODING WOULD HAVE BEEN "58" FOR "CAUSE TRACED TO SOFTWARE CODING". NOTES: THE FOLLOWING SECTIONS WERE UPDATED FROM THE INITIAL MEDWATCH REPORT: H1/H2: DEVICE EVALUATION. H6: COMPONENT CODE, INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS. Manufacturer Narrative: A1: THE FULL PATIENT IDENTIFIER IS (B)(4). A2, A3, A4 AND A5: THE CUSTOMER DID NOT SUPPLY PATIENT DEMOGRAPHICS SUCH AS AGE, DATE OF BIRTH, SEX, WEIGHT, ETHNICITY OR RACE. H3 AND H6: THE BECKMAN COULTER R&D TEAM HAS IMPLEMENTED A TEMPORARY COUNTERMEASURE WHICH HAS BEEN VERIFIED BY THE ONSITE FIELD SERVICE ENGINEER TO ADDRESS THE ISSUE. THE STEPS FOR THIS SHORT-TERM SOLUTION ARE AS FOLLOWS: NAVIGATE TO SYSTEM CONFIGURATION > LIS CONFIGURATION, DISABLE THE LIS PROTOCOL AND ACTIVATE THE NEW CONFIGURATION, REBOOT THE PC AND WAIT FOR INITIALIZATION, RE-ENABLE THE LIS CONFIGURATION AND ACTIVATE THE CONFIGURATION. FOR A LONG-TERM SOLUTION, TWO SOFTWARE ANOMALIES HAVE BEEN IDENTIFIED AND WILL BE CORRECTED IN THE UPCOMING DXI 9000 V1.22 SOFTWARE RELEASE. THE IMPACTED USER IN IRELAND HAS NOT SUBMITTED ANY ADDITIONAL COMPLAINTS RELATED TO THIS ISSUE. BECKMAN COULTER IS CURRENTLY FOLLOWING ESTABLISHED PROCEDURES TO DETERMINE THE ROOT CAUSE. AN INVESTIGATION REPORT IS IN PROGRESS WITH A TARGET COMPLETION DATE OF 11 MARCH 2025. UPON COMPLETION OF THE INVESTIGATION REPORT, ADDITIONAL INVESTIGATION ACTIVITIES WILL BE CONDUCTED TO SUPPORT A CAPA TO IDENTIFY CORRECTIVE ACTIONS TO ADDRESS THE ROOT CAUSE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21527119&pc=NPT | 2025587-2025-01652 | 2024-06-01 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-05 05:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Stroke/CVA; Endocarditis | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 6 MONTHS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT DEVELOPED A FEVER AND WAS DIAGNOSED WITH INFECTIOUS ENDOCARDITIS DUE TO A VERRUCOUS ULCER IN THE TRANSCATHETER VALVE AND THE TRICUSPID VALVE. SUBSEQUENTLY, ANTIBIOTICS WERE ADMINISTERED. DESPITE TREATMENT, THE PATIENT'S CONDITION DID NOT IMPROVE. A CEREBRAL INFARCTION WAS NOTED APPROXIMATELY 2 MONTHS LATER. THE PATIENT PASSED AWAY AND THE CAUSE OF DEATH WAS REPORTED TO BE INFECTION. PER THE PHYSICIAN, THE CAUSE OF DEATH IS NOT RELATED TO THE PROCEDURE. INVOLVEMENT WITH THE VALVE CANNOT BE DENIED AS A CAUSE OF UNSUCCESSFUL ANTIBIOTIC TREATMENT, HOWEVER.Event Description: IT WAS REPORTED THAT APPROXIMATELY 6 MONTHS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT DEVELOPED A FEVER AND WAS DIAGNOSED WITH INFECTIOUS ENDOCARDITIS DUE TO A VERRUCOUS ULCER IN THE TRANSCATHETER VALVE AND THE TRICUSPID VALVE. SUBSEQUENTLY, ANTIBIOTICS WERE ADMINISTERED. DESPITE TREATMENT, THE PATIENT'S CONDITION DID NOT IMPROVE. A CEREBRAL INFARCTION WAS NOTED APPROXIMATELY 2 MONTHS LATER. THE PATIENT PASSED AWAY AND THE CAUSE OF DEATH WAS REPORTED TO BE INFECTION. PER THE PHYSICIAN, THE CAUSE OF DEATH IS NOT RELATED TO THE PROCEDURE. INVOLVEMENT WITH THE VALVE CANNOT BE DENIED AS A CAUSE OF UNSUCCESSFUL ANTIBIOTIC TREATMENT, HOWEVER. ADDITIONAL INFORMATION WAS RECEIVED INDICATING THAT THE PATIENT'S N-R FUSED TYPE 1 BICUSPID VALVE AND HORIZONTAL AORTA OF 68 DEGREES MAY HAVE BEEN CONTRIBUTING FACTORS TO THE CEREBRAL INFARCTION. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 (SECOND PARAGRAPH), D4, AND H4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21527571&pc=NPT | 2015691-2025-01508 | 2025-02-20 05:00 | Death | EDWARDS LIFESCIENCES | 2025-03-05 05:00 | NPT | EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Rupture | P140031 |  | 1 | Event Description: AS REPORTED, IT WAS A CASE OF AN IMPLANT OF A 26MM SAPIEN 3 ULTRA VALVE, IN AORTIC POSITION BY TRANSFEMORAL APPROACH. AFTER VALVE DEPLOYMENT, THE PATIENT BECAME HYPOTENSIVE AND THE ANGIOGRAPHIC IMAGES SHOWED A PERICARDIAL EFFUSION AND TAMPONADE. PERICARDIOCENTESIS WAS PERFORMED WITH CONCOMITANT CPR. THE PATIENT EXPIRED. AS PER MEDICAL OPINION, THE PERCEIVED ROOT CAUSE OF THE ANNULAR RUPTURE WAS DUE TO CALCIUM FROM THE VBR TOWARDS THE LVOT. ALTHOUGH THE CALCIFICATION WAS NOT VERY THICK, IT WAS FOUND TO BE A POINT OF FRAGILITY OF THE HEART. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. ACCORDING TO THE DEVICE TRAINING MANUALS, RISK FACTORS FOR AORTIC DISSECTION, CARDIAC/AORTIC HEMATOMA, OR ANNULAR RUPTURE DURING THE TVR PROCEDURE INCLUDE SIGNIFICANT VALVE OVER SIZING, SEVERELY OBLITERATED SINUSES OF VALSALVA, PORCELAIN AORTA AND/OR PRESENCE OF BULKY CALCIFICATION, AND NARROW CALCIFIED SINOTUBULAR JUNCTION. IN ADDITION, ADVANCED AGE, FEMALE GENDER, SMALL BODY WEIGHT, AND STEROID DEPENDENCY CAN ALSO BE CONTRIBUTING FACTORS. THE SAPIEN TRANSCATHETER HEART VALVE (ALL MODELS) RELIES CALCIUM TO SECURELY ANCHOR IN THE LANDING ZONE. DESPITE THIS BENEFICIAL ASPECT OF CALCIUM, BULKY CALCIUM CAN INCREASE THE RISK OF CALCIFIC NODULE DISPLACEMENT INTO THE VASCULATURE, WHICH CAN LEAD TO VASCULAR INJURY. AT TIMES, THE EXTENT AND DISTRIBUTION OF CALCIUM CAN IMPAIR EASE OF DELIVERY OF THE VALVE, CORRECT POSITIONING OF THE VALVE, DEPLOYMENT OF THE VALVE, AND PROCEDURAL SUCCESS. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST PER MEDICAL OPINION, THE PERCEIVED ROOT CAUSE OF THE ANNULAR RUPTURE WAS DUE TO CALCIUM FROM THE VBR TOWARDS THE LVOT. ALTHOUGH THE CALCIFICATION WAS NOT VERY THICK, IT WAS FOUND TO BE A POINT OF FRAGILITY OF THE HEART. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21528435&pc=OZD | 1220648-2025-26524 | 2025-02-11 05:00 | Death | ABIOMED, INC. | 2025-03-05 05:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT ON IMPELLA 5.5 SUPPORT EXPERIENCED A STROKE. THE PATIENT WAS UNRESPONSIVE, AND PUPILS WERE UNEQUAL. THE PATIENT WAS ON AND OFF ANTICOAGULATION. IT IS UNKNOWN WHETHER STROKE WAS IMPELLA 5.5 RELATED. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. THE RELATIONSHIP BETWEEN THE IMPELLA 5.5 AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE INVESTIGATION OF STROKE/CVA HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21528518&pc=DSQ | 2916596-2025-01488 | 2025-02-21 05:00 | Death | THORATEC CORPORATION | 2025-03-05 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORT THAT PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT'S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. DUE TO HOSPITAL POLICY, NO ADDITIONAL INFORMATION WAS PROVIDED. THE HEARTMATE 3 LVAS IFU, REV. G, IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21528522&pc=DSQ | 2916596-2025-01285 | 2021-03-01 05:00 | Death | THORATEC CORPORATION | 2025-03-05 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE "ORAL AMIODARONE AND PROPRANOLOL IN MAINTENANCE THERAPY OF POST IMPLANTATION TACHYCARDIA: AN OBSERVATIONAL STUDY" THAT THE HM3 MAY BE RELATED TO RIGHT HEART FAILURE AND DEATH. CANNULA-RELATED VTS WERE OBSERVED IN >50% OF PATIENTS IN THE FIRST MONTH AFTER LEFT VENTRICULAR ASSIST DEVICE (LVAD) IMPLANTATION. IN ADDITION TO CHRONIC VENTRICULAR SCARS, DE NOVO SCARS AT LVAD ENTRY CANNULATION SITES PROVIDED POTENTIAL SUBSTRATES FOR REENTRY ARRHYTHMIA. THUS, VT ORIGINATING FROM THE LATERAL WALL OF THE LEFT VENTRICLE (LV) IN THE EVENT OF NON-MYOCARDITIS, FROM THE LV SEPTUM DUE TO SARCOIDOSIS, AND FROM THE APICAL ANEURYSM-INDUCED SCAR TISSUE IN THE EVENT OF ANTERIOR MYOCARDIAL INFARCTION CAN OCCUR, AND NOT ALL VTS WERE CAUSED BY DE NOVO SCAR TISSUE. PROPRANOLOL IS WELL-TOLERATED BY PATIENTS WITH LVAD IMPLANTATION AND CAN BE ADMINISTERED IN HIGHER DOSES. THUS, PROPRANOLOL IS THE PREFERRED BETA-BLOCKER FOR THIS PATIENT GROUP. THIS MEDICATION CAN BE USED IN HIGH DOSES IN PATIENTS WITH LVAD IMPLANTATION, ACT ON THE CENTRAL NERVOUS SYSTEM RECEPTORS THROUGH ITS LIPOSOLUBLE CHARACTERISTICS, AND PROVIDE SIGNIFICANTLY REDUCED CATECHOLAMINE DISCHARGE BY NON SELECTIVE BETA-BLOCKADE. IT WAS REPORTED THAT 1 PATIENT DIED ONE DAY AFTER LVAD IMPLANTATION BECAUSE OF LOW RIGHT HEART OUTPUT. TABLE 2 REPORTED PATIENT 11, AS A 56-YEAR-OLD IMPLANTED WITH THE HEARTMATE 3 THAT PASSED AWAY ONE DAY LATER. THE PATIENT HAD A POSITIVE HISTORY OF ARRHYTHMIA, NONISCHEMIC HEART FAILURE AND MODERATE-SEVERELY DEPRESSED RIGHT VENTRICULAR FUNCTION. THE PATIENT WAS FOLLOWED UP AFTER NOT RECEIVING AMIODARONE TREATMENT. Manufacturer Narrative: (B)(6). MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED EVENTS COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE HEARTMATE 3 LVAS WAS NOT RETURNED FOR EVALUATION. THE DEVICE SERIAL NUMBER WAS NOT PROVIDED AND WAS UNABLE TO BE DETERMINED THROUGH THIS EVALUATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ OF THIS DOCUMENT LISTS RIGHT HEART FAILURE, CARDIAC ARRHYTHMIA, AND DEATH AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. ADDITIONALLY, SECTION 6 ¿PATIENT CARE AND MANAGEMENT¿ LISTS RIGHT HEART FAILURE AS A POTENTIAL RISK/ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS. THIS SECTION (UNDER ¿RIGHT HEART FAILURE¿) ALSO OUTLINES INDICATIONS OF RIGHT HEART FAILURE AS WELL AS POSSIBLE TREATMENTS. SECTION 6 ALSO LISTS ARRHYTHMIA AS A POTENTIAL LATE POSTIMPLANT COMPLICATION THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SECTION A, D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. ARSLAN, A., ÖZKAYA, Ö. G., AYTEMIZ, F., ALTAY, H., ÖZCAN, E. E., KÜÇÜKAKSU, D. S., BAKUY, V., KOZAN, Ö., & PEHLIVANOGLU, S. (2024). ORAL AMIODARONE AND PROPRANOLOL IN MAINTENANCE THERAPY OF POSTIMPLANTATION TACHYCARDIA: AN OBSERVATIONAL STUDY. MEDICINE, 103(28), E38839. HTTPS://DOI.ORG/10.1097/MD.0000000000038839. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21528862&pc=DSQ | 2916596-2025-01258 | 2025-02-10 05:00 | Death | THORATEC CORPORATION | 2025-03-05 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Loss of consciousness; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD SOME ISSUES OVER THE PAST SEVERAL MONTHS WITH INCREASING DIURETIC REQUIREMENTS DUE PRESUMED RIGHT HEART FAILURE. IT WAS NOTED THAT THE PATIENT HAD COVID PNEUMONIA ON (B)(6) 2024. ON (B)(6) 2025 RAMP STUDY WAS DONE. THE PATIENT WAS GIVEN IV FUROSEMIDE 100MG IN CLINIC ON (B)(6) 2025 AS A RESULT OF THE CATHETERIZATION WITH PLANS TO SEE AGAIN IN THE UPCOMING WEEK. LATE ON (B)(6) 2025, THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) STARTED ALARMING WHILE SITTING ON A CHAIR. THE PATIENT CHANGED THEIR BATTERIES, BUT THE ALARM CONTINUED. THE PATIENT WAS THEN TO CHANGE THE SYSTEM CONTROLLER. THEIR SPOUSE THOUGHT THE PATIENT SHOULD CALL FOR HELP, BUT THE PATIENT SAID THEY HAD THINGS UNDER CONTROL. IN FRUSTRATION, THE SPOUSE WENT TO THE KITCHEN TO MAKE DINNER AND WHEN COMING BACK, THE PATIENT WAS UNRESPONSIVE. EMERGENCY MEDICAL SERVICES WERE CALLED WHO CAME AND STARTED CARDIOPULMONARY NECESSITATION. IN THE MEANTIME, THE SPOUSE DESPERATELY TRIED TO CHANGE THE CONTROLLER AND WAS EVENTUALLY ABLE TO GET THE PATIENT ON THE NEW CONTROLLER. THE PATIENT WAS AWAKENED AT SOME POINT ALONG THE WAT. IT WAS UNKNOWN HOW LONG THE PATIENT WAS DISCONNECTED BUT IT APPEARED TO BE 3 ATTEMPTS. IN THE EMERGENCY ROOM (ER) THE PATIENT WAS RESPONSIVE, AND THE DEVICE WAS WORKING. THE PATIENT WAS INTUBATED TO PROTECT THEIR AIRWAY AS THEY COUGHED UP BLOOD. A USUAL BLOODWORK WAS DONE WHICH WAS NOT CONCERNING EITHER THAN THE INTERNATIONAL NORMALIZED RATIO (INR). LACTATE 3.3, CREATININE WAS 172 (CLOSE TO THE PATIENT'S USUAL), AND HEMOGLOBIN WAS UNCHANGED FROM PREVIOUS. COMPUTED TOMOGRAPHY WAS DONE WHICH BASICALLY SHOWED SOME INFILTRATES IN RIGHT LUNG AND SOME BLOOD. THE PATIENT DID RECEIVE 2 LITERS OF FLUID FROM THE ER AND WAS PLACED ON BROAD SPECTRUM ANTIBIOTICS. ADDITIONAL INFORMATION RECEIVED ON 24FEB2025 CONFIRMED THAT THE PATIENT WAS STARTED ON BROAD SPECTRUM ANTIBIOTICS DUE TO PNEUMONIA. NO ARRYTHMIAS. IN THE INTENSIVE CARE UNIT (ICU), INITIAL ASSESSMENT SHOWED A LOW PULSATILITY INDEX (PI), MEAN ARTERIAL PRESSURE (MAP) WAS 65. THERE WAS NO EVIDENCE FOR BLOOD LOSS OR DISTRIBUTE SHOCK. THE PATIENT WAS NOT DEHYDRATED, AND THEIR CENTRAL VENOUS PRESSURE (CVP) WAS APPROXIMATELY 16 AND THEIR CARDIOMEMS WAS ASSESSED: PA 47/ 24 AT 4800 RPM. THESE VALUES WERE NOT MUCH DIFFERENT THAN THE ASSESSMENT FROM A WEEK AGO. DESPITE THE HEMODYNAMICS, RIGHT VENTRICULAR FAILURE WAS SUSPECTED. IN THE EVENING, THE RPM WAS DOWN FROM 4800 TO 4600. IT WAS NOTED THAT THERE WAS REALLY NO DATA FROM THE DEVICE AS IT WAS BRAND NEW. OVERNIGHT THE DEVICE STARTED TO ALARM AGAIN AT AROUND 0300. THE PATIENT WAS GIVEN SOME MORE FLUID AS THE PI REMAINED LOW. IN ADDITION, NOREPINEPHRINE WAS STARTED TO ATTAIN MAP OF 70 AND SUBSEQUENTLY 75. THIS WOULD HELP FOR A WHILE AND THEN THE ALARMS WOULD GO OFF. THE PATIENT WAS THEN STARTED WITH DOBUTAMINE, AND DIURESIS WAS DOWN TO CVP 12. NOREPINEPHRINE WAS WEANED. THE ALARMS ALSO STOPPED. THE HEARTMATE TOUCH WAS REVIEWED, AND IT LOOKED LIKE MULTIPLE SPIN DOWNS. THE DOBUTAMINE WAS BUMPED TO 5UG. GRADUALLY THE PI CAME UP TO 4-6 AND FLOWS INCREASED. DESPITE THE CARDIOMEMS, THE FINDINGS WERE THOUGHT TO SUPPORT RV FAILURE. THE PATIENT WAS IN STABLE CONDITION. THEY WERE COOLED AND WAS BEING WARMED UP ON (B)(6) 2025 REQUIRING MORE DIURESIS AND WAS TO BE SWITCHED TO MILRINONE. MOST OF THE EVENTS WERE COMMUNICATED BY THE SPOUSE AND WHEN THE ORIGINAL CONTROLLER THE SPOUSE PRESENTED WAS INTERROGATED, NOTHING MUCH WAS SEEN EXCEPT FOR THE POWER DISCONNECTIONS. THE EVENT LOG FILE FROM THE ORIGINAL PRIMARY CONTROLLER THAT WAS EXCHANGED CAPTURED SOME LOW VOLTAGE ADVISORY EVENTS ON (B)(6) 2025 AT 18:58-18:59 WHILE USING BATTERY POWER. THE PATIENT APPEARED TO HAVE ATTEMPTED TO EXCHANGE THE BATTERIES BUT ATTACHED A BATTERY THAT WAS ALREADY IN LOW VOLTAGE STATE RESULTING IN CONTINUED LOW VOLTAGE ADVISORY EVENTS. THE CONTROLLER WAS EXCHANGED SHORTLY AFTER AT 19:04 AND THEN THE ORIGINAL RECONNECTED AGAIN AT 19:09 THEN WAS DISCONNECTED ONCE AGAIN AT 19:18.Event Description: ADDITIONAL INFORMATION WAS RECEIVED ON 28MAR2025 THAT THE PATIENT PASSED AWAY ON (B)(6) 2025 DUE TO WITHDRAWAL OF SUPPORT IN PALLIATIVE CARE. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED AND DEVICE PARAMETERS WERE WITHIN NORMAL LIMITS. THE PUMP WAS NOT EXPLANTED, AND AN AUTOPSY WAS NOT PERFORMED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE (HM) 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. REVIEW OF THE SUBMITTED LOG FILES CONFIRMED PUMP STOP EVENTS ASSOCIATED WITH THE DISCONNECTION OF THE DRIVELINE, WHICH APPEARED CONSISTENT WITH THE REPORT OF ATTEMPTED SYSTEM CONTROLLER EXCHANGES. A SPECIFIC CAUSE FOR THE PATIENT¿S INFECTION COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. ADDITIONALLY, ANALYSIS OF THE SUBMITTED LOG FILES ALSO CONFIRMED LOW FLOW ALARMS; HOWEVER, A SPECIFIC CAUSE FOR THESE EVENTS COULD NOT CONCLUSIVELY BE DETERMINED THROUGH THIS EVALUATION. THE CONTROLLER EVENT LOG FILES COLLECTIVELY CONTAINED EVENTS FROM (B)(6) 2025 THROUGH (B)(6) 2025. THE DRIVELINE WAS DISCONNECTED TWICE ON (B)(6) 2025, CAUSING THE PUMP TO STOP. FOLLOWING EACH RECONNECTION OF THE DRIVELINE THE PUMP RESUMED OPERATION AT THE SET SPEED WITHOUT ISSUE. THE DRIVELINE WAS DISCONNECTED AGAIN ON (B)(6) 2025, CAUSING THE PUMP TO STOP AND THE DRIVELINE REMAINED DISCONNECTED FOR THE REMAINDER OF THE FILE. THESE EVENTS APPEARED CONSISTENT WITH THE REPORT OF ATTEMPTED SYSTEM CONTROLLER EXCHANGES PERFORMED BY THE PATIENT AND WIFE. LOW FLOW HAZARD ALARMS WERE CAPTURED ON (B)(6) 2025. NO OTHER NOTABLE EVENTS OR ALARMS ASSOCIATED WITH THE PUMP WERE CAPTURED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, DOCUMENT AND THE HEARTMATE 3 PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING RIGHT HEART FAILURE, BLEEDING, INFECTION AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 1 ALSO ADDRESSES ALL PUMP PARAMETERS. SECTION 4 OF THE IFU, ¿SYSTEM MONITOR¿, DESCRIBES THE PUMP FLOW DISPLAY AND THE HAZARD ALARMS AND EXPLAINS THAT CHANGES IN PATIENT CONDITION CAN RESULT IN LOW FLOW. PER DESIGN, WHEN THE ESTIMATED FLOW VALUE IS CALCULATED AT LESS THAN 2.5 LITERS PER MINUTE (LPM), A LOW FLOW STATUS IS POSTED TO THE LOG FILE. IF THE FLOW REMAINS BELOW 2.5 LPM FOR 10 SECONDS, A LOW FLOW HAZARD ALARM IS TRIGGERED. SECTION 2 OF THE IFU, "SYSTEM OPERATIONS", AND SECTION 2 OF THE PATIENT HANDBOOK, "HOW YOUR HEART PUMP WORKS", WARN THAT IF THE DRIVELINE DISCONNECTS FROM THE SYSTEM CONTROLLER, THE PUMP STOPS. THESE SECTIONS FURTHER INSTRUCT THAT IF THE DRIVELINE DISCONNECTS FROM THE SYSTEM CONTROLLER, IT SHOULD BE PROMPTLY RECONNECTED TO RESUME PUMP OPERATION. THE PATIENT HANDBOOK FURTHER EXPLAINS THAT THE PUMP CANNOT RUN WITHOUT POWER. SECTION 2 OF THE IFU ALSO CONTAINS A SECTION TITLED "CONNECTING THE DRIVELINE TO THE SYSTEM CONTROLLER", WHICH PROVIDES INSTRUCTIONS ON CONNECTING/DISCONNECTING THE DRIVELINE TO/FROM THE SYSTEM CONTROLLER AND MAKING SURE THAT IT IS FULLY AND PROPERLY INSERTED INTO THE SYSTEM CONTROLLER SOCKET. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿ANTICOAGULATION¿), PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) VALUES, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. SECTION 6 OF THE IFU (UNDER "RIGHT HEART FAILURE") DISCUSSES THE POTENTIAL DEVELOPMENT OF RIGHT HEART FAILURE FOLLOWING IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS, INCLUDING INOTROPES AND RIGHT VENTRICULAR ASSIST DEVICE (RVAD) PLACEMENT. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, ALSO LISTS INFECTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. SECTION 5 OF THE PATIENT HANDBOOK, ¿ALARMS AND TROUBLESHOOTING¿, AND SECTION 7 OF THE IFU, ¿ALARMS AND TROUBLESHOOTING¿, PROVIDES INFORMATION ON ALL SYSTEM ALARM CONDITIONS AS WELL AS THE APPROPRIATE ACTIONS ASSOCIATED WITH EACH CONDITION. FURTHERMORE, SECTION 8 OF THE PATIENT HANDBOOK, ¿HANDLING EMERGENCIES¿, ALSO PROVIDES EXAMPLES OF EMERGENCIES AND THE PROPER ACTIONS TO TAKE IN THE EVENT AN EMERGENCY OCCURS. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21528886&pc=OAE | MW5167328 |  | Death | BIOSENSE WEBSTER | 2025-03-05 05:00 | OAE | CATHETER, PERCUTANEOUS, CARDIAC ABLATION, FOR TREATMENT OF ATRIAL FIBRILLATION |  | Cardiac Arrest; Pleural Effusion; Ventricular Fibrillation; Cardiac Tamponade |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED CARDIAC TAMPONADE, VENTRICULAR FIBRILLATION, AND CARDIAC ARREST. THE PATIENT EXPIRED AFTER THE PROCEDURE. A LEFT ACCESSORY PATHWAY MAP WAS PERFORMED FOLLOWED BY ABLATION OF THE LEFT LATERAL LEFT ATRIUM (LA) WITH A NON-BOSTON SCIENTIFIC RF CATHETER. AFTER PERFORMING THIS ABLATION, A PLEURAL EFFUSION WAS NOTICED THAT BECAME VERY LARGE VERY FAST. A PERICARDIAL PERFUSION WAS PERFORMED TO CLEAR THE EFFUSION, BUT THEN THE PATIENT WENT INTO CARDIAC ARREST WITH VF. CARDIOPULMONARY RESUSCITATION WAS PERFORMED ALONG WITH MULTIPLE SHOCKS FROM CARDIOVERSION. THE PATIENT WAS PLACED ON EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO), THEN TRANSFERRED TO ANOTHER HOSPITAL AND ADMITTED. THE PATIENT WAS NOT CONSIDERED STABLE AT THE TIME OF ADMISSION AND REMAINED ON LIFE SUPPORT. THE PATIENT LATER EXPIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21528909&pc=DRF | 3008452825-2025-00071 | 2025-02-17 05:00 | Death | ST. JUDE MEDICAL | 2025-03-05 05:00 | DRF | INQUIRY¿ STEERABLE DIAGNOSTIC CATHETER | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Cardiac Arrest | K961924 |  | 1 | Event Description: DURING AN ATRIAL FLUTTER PROCEDURE, THE PATIENT BECAME BRADYCARDIC, WENT INTO CARDIAC ARREST, AND EXPIRED. THE PROCEDURE BEGAN BY POSITIONING THE INQUIRY CATHETER IN THE CORONARY SINUS AND USING A NON-ABBOTT INTRACARDIAC ECHO PROBE TO MAP THE RIGHT ATRIUM. THE MAP AND GEOMETRY WERE COLLECTED WITH A NON-ABBOTT CIRCULAR CATHETER. THE TYPICAL ATRIAL FLUTTER ABLATION WAS PERFORMED WITH A NON-ABBOTT 8MM RF ABLATION CATHETER. AFTER MAPPING AND ABLATION IN THE RIGHT ATRIUM WERE COMPLETED, THE NON-ABBOTT CIRCULAR CATHETER WAS MOVED TO THE LEFT ATRIUM VIA TRANSSEPTAL PUNCTURE USING SWARTZ AND AGILIS SHEATHS AND A BRK NEEDLE. GEOMETRY OF THE LEFT ATRIUM AND A VOLTAGE MAP WERE COLLECTED AND SHOWED AN ELECTRICAL SIGNAL IN THE PULMONARY VEINS. ISOLATION OF THE PULMONARY VEINS BEGAN WITH A NON-ABBOTT (FARAPULSE, BOSTON SCIENTIFIC) CATHETER AND NON-ABBOTT (BOSTON SCIENTIFIC) EQUIPMENT AND TECHNICAL SUPPORT. IMMEDIATELY AFTER ABLATION OF THE PULMONARY VEINS, A NEW MAP OF THE LEFT ATRIUM WAS COLLECTED USING THE NON-ABBOTT CIRCULAR CATHETER TO CHECK WHETHER THE VEINS WERE ISOLATED. AT THAT MOMENT, AFTER THE GEOMETRY COLLECTION BEGAN ON THE POSTERIOR WALL OF THE LEFT ATRIUM, THE PATIENT BECAME BRADYCARDIC AND WENT INTO CARDIAC ARREST. THE TEAM PERFORMED RESUSCITATION MANEUVERS, BUT AFTER A PERIOD OF ATTEMPTS, THE PATIENT DIED. Manufacturer Narrative: ADDITIONAL INFORMATION: G3, H2, H3, H6, H11. THE RESULTS OF THE INVESTIGATION ARE INCONCLUSIVE SINCE THE DEVICE WAS NOT RETURNED FOR ANALYSIS. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED CARDIAC ARREST AND DEATH COULD NOT BE CONCLUSIVELY DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21528918&pc=DSQ | 2916596-2025-01338 | 2025-02-18 05:00 | Death | THORATEC CORPORATION | 2025-03-05 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO RIGHT HEART FAILURE AND CARDIAC ARREST. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. Manufacturer Narrative: SECTION A4: PATIENT WEIGHT WAS NOT PROVIDED. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE DEVICE WAS NOT RETURNED FOR EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING RIGHT HEART FAILURE, CARDIAC ARRYTHMIA, AND DEATH THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 OF THE IFU, UNDER "RIGHT HEART FAILURE", DISCUSSES THE POTENTIAL DEVELOPMENT OF RIGHT HEART FAILURE FOLLOWING IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS. ADDITIONALLY, SECTION 6, ¿PATIENT CARE AND MANAGEMENT¿, LISTS ARRHYTHMIA AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21529118&pc=MNS | 2518422-2025-027444 | 2024-07-10 04:00 | Death | RESPIRONICS, INC. | 2025-03-05 05:00 | MNS | DREAMSTATION BIPAP AVAPS | Degraded | No Clinical Signs, Symptoms or Conditions | K102465 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21529152&pc=DSQ | 2916596-2025-01269 | 2025-02-17 05:00 | Death | THORATEC CORPORATION | 2025-03-05 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Cardiomyopathy; Hemorrhage/Bleeding; High Blood Pressure/ Hypertension; Hypothermia; Cardiogenic Shock; Respiratory Acidosis; Gastrointestinal Hemorrhage | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS FOUND UNRESPONSIVE IN THEIR HOME. THE PATIENT'S DOWNTIME WAS UNKNOWN. THEY WERE INTUBATED IN THE EMERGENCY ROOM OF THE OUTSIDE HOSPITAL. THEY WERE NOTED TO HAVE DARK BLOODY SECRETIONS IN THE ENDOTRACHEAL TUBE, AND THERE WAS CONCERN FOR GASTROINTESTINAL BLEEDING GIVEN THEIR HISTORY. THEY RECEIVED 2 GRAMS CALCIUM GLUCONATE, 1 LITER NORMAL SALINE, 1 LITER LACTATED RINGER'S SOLUTION, 1 UNIT PACKED RED BLOOD CELLS, EPINEPHRINE 200 MICROGRAM PUSH, AND STARTED ON EPINEPHRINE INFUSION. THE PATIENT ARRIVED HYPOTHERMIC, WITH NO NEUROLOGICAL RESPONSE WITHOUT ANY SEDATION, AND VERY FAINT PULSES. THE SITE WAS UNABLE TO OBTAIN TEMPERATURE GIVEN SEVERE HYPOTHERMIA NOR BLOOD PRESSURE GIVEN SEVERELY LOW MEAN ARTERIAL PRESSURES (MAPS). RADIAL ARTERIAL LINE WAS ATTEMPTED WITHOUT SUCCESS. BEDSIDE POINT OF CARE ULTRASOUND SHOWED VERY LITTLE, IF ANY, CARDIAC ACTIVITY. FEMORAL VENOUS AND ARTERIAL LINES WERE PLACED AND SHOWED VERY FAINT ARTERIAL ACTIVITY, ESTIMATED MAP IN 10-20 RANGE. THEIR FEMORAL ARTERIAL LINE CLOTTED SOON AFTER PLACEMENT. GIVEN THEIR SEVERE LABS, CONCERN FOR DISSEMINATED INTRAVASCULAR COAGULATION AND SUSPECTED PROLONGED HYPOPERFUSION WITH POOR NEUROLOGICAL EXAM, PATIENT WAS DECLARED DECEASED ON 24DEC2024 AT 18:47. THE CAUSE OF DEATH IN THIS CASE WAS CARDIOGENIC SHOCK DUE TO NON-ISCHEMIC DILATED CARDIOMYOPATHY. OTHER SIGNIFICANT CONTRIBUTORY FACTORS INCLUDED HYPERTENSION AND TYPE II DIABETES MELLITUS. THE SITE IS WAITING ON DEVICE EVALUATION FROM ABBOTT AND COULD NOT PROVIDE WHETHER THE DEATH WAS DEVICE RELATED OR WHETHER THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. ADDITIONALLY, REVIEW OF THE RETRIEVED LOG FILES REVEALED DATA CONSISTENT WITH A TOTAL LOSS OF POWER TO THE PUMP, WHICH WOULD HAVE RESULT IN A PUMP STOP EVENT. (B)(6) WAS RETURNED ASSEMBLED WITH THE PUMP CABLE SEVERED APPROXIMATELY 18¿ FROM THE PUMP HEADER. THE REMAINDER OF THE PUMP CABLE, AS WELL AS THE MODULAR CABLE, WERE NOT RETURNED. THE SEALED OUTFLOW GRAFT WAS RETURNED SECURED TO THE PUMP COVER OUTLET PORT. THE OUTFLOW GRAFT BEND RELIEF WAS ATTACHED AT THE GRAFT HARDWARE. THE APICAL CUFF WAS RETURNED WITH THE CUFF LOCK FULLY ENGAGED AND INCLUDED TISSUE FROM THE VENTRICLE. UPON DISASSEMBLY OF THE RETURNED PUMP, EXAMINATION OF THE BLOOD-CONTACTING SURFACES DID NOT REVEAL ANY ADHERED DEPOSITIONS THAT WOULD HAVE CONTRIBUTED TO A FLOW OR FUNCTIONAL ISSUE. VISUAL INSPECTION OF THE PUMP ROTOR AND ROTOR WELL DID NOT REVEAL ANY OBVIOUS SURFACE SCRATCHES OR DEFECTS. THE LVAD EVENT LOG FILE RETRIEVED FROM THE RETURNED DEVICE CONTAINED RELEVANT DATA SPANNING APPROXIMATELY 2 HOURS (21:20:23 THROUGH 23:07:15 WITH A DEFAULT DATE TIMESTAMP OF (B)(6) 2000). REVIEW OF THE LVAD CLOCK REVEALED THAT IT HAD PREVIOUSLY BEEN RESET TO DEFAULT TIMESTAMPS, INDICATING A COMPLETE LOSS OF POWER TO THE SYSTEM THAT WOULD HAVE CAUSED THE PUMP TO STOP. MEAN FLOW WAS CAPTURED RANGING FROM 1.4-4.5 LITERS PER MINUTE (LPM) THROUGHOUT THE RELEVANT DATA, WHICH APPEARED CONSISTENT WITH THE REPORTED EVENT AND PATIENT OUTCOME. NO OTHER NOTABLE EVENTS OR ALARMS WERE CAPTURED, AND THE DEVICE APPEARED TO OPERATE AS INTENDED AT THE SET SPEED. (B)(6) WAS CLEANED, REBUILT, AND FUNCTIONALLY TESTED ON A MOCK CIRCULATORY LOOP. THE DEVICE OPERATED AS INTENDED AND IN ACCORDANCE WITH MANUFACTURING SPECIFICATIONS. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV D, AND THE HEARTMATE 3 PATIENT HANDBOOK, REV A, ARE CURRENTLY AVAILABLE. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING BLEEDING, RESPIRATORY FAILURE, HYPERTENSION, AND DEATH. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿ANTICOAGULATION¿), PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) RANGE, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. THE IFU AND PATIENT HANDBOOK CONTAIN INFORMATION ON SYSTEM CONTROLLER ALARMS, AS WELL AS THE PROPER ACTIONS ASSOCIATED WITH THEM. THESE DOCUMENTS ALSO WARN OF EVENTS THAT MAY CAUSE THE PUMP TO STOP AND HOW TO PREVENT PUMP STOPS FROM OCCURRING. THE ¿HANDLING EMERGENCIES¿ SECTION OF THE PATIENT HANDBOOK LISTS EXAMPLES OF EMERGENCIES AND WHAT ACTIONS TO TAKE. FURTHERMORE, THE PATIENT HANDBOOK CAUTIONS THE USER TO CALL THEIR HOSPITAL CONTACTS IF THEY THINK, FOR ANY REASON, ANY PORTION OF THEIR EQUIPMENT IS NOT FUNCTIONING AS USUAL, IS BROKEN, OR THEY ARE UNCOMFORTABLE WITH THE OPERATION OF THE EQUIPMENT. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21529339&pc=DSQ | 2916596-2025-01560 | 2022-01-01 05:00 | Death | THORATEC CORPORATION | 2025-03-05 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE ¿PREDICTIVE ACCURACY OF HEARTMATE 3 RISK SCORE AFTER THE HEART TRANSPLANT ALLOCATION CHANGE¿ THAT HEARTMATE 3 (HM3) MAY BE ASSOCIATED WITH TRANSPLANT, PUMP EXCHANGE/EXPLANT, AND DEATH. THIS WAS A MULTICENTER RETROSPECTIVE ANALYSIS OF CONTEMPORARY HM3 RECIPIENTS. THE UNITED NETWORK FOR ORGAN SHARING (UNOS) REVISED ITS DONOR HEART ALLOCATION SYSTEM ON 18OCT2018 TO REDUCE WAITLIST MORTALITY AND IMPROVE ORGAN DISTRIBUTION EQUITY. SINCE THEN, THERE HAD BEEN A DRASTIC INCREASE IN TEMPORARY MECHANICAL CIRCULATORY SUPPORT DEVICES AND A CONCURRENT DECREASE IN DURABLE LVAD USE AS A BRIDGE-TO-TRANSPLANT STRATEGY. ADULT PATIENTS WHO RECEIVED HM3 AS INITIAL IMPLANTATION BETWEEN 2015 AND 2022 AT TWO ACADEMIC MEDICAL INSTITUTIONS (N = 315) WERE ANALYZED. PARTICIPANTS IN THE MOMENTUM 3 CLINICAL TRIAL WERE INCLUDED. PATIENTS WHO UNDERWENT PREEMPTIVE RIGHT VENTRICULAR (RV) ASSIST DEVICE IMPLANT IN A CONCURRENT STUDY (N = 17) AND THOSE WITH INSUFFICIENT DATA (N = 56) WERE EXCLUDED. PATIENTS WERE CHARACTERIZED BY THEIR HM3RS AS LOW (<2.41), MEDIUM (2.41¿2.97), OR HIGH (>2.97) RISK. PATIENTS IN THIS STUDY UNDERWENT TRANSPLANT, PUMP EXCHANGE/EXPLANT, OR DEATH. THE PRIMARY OUTCOME WAS DEATH AT 1 AND 2 YEARS, STRATIFIED BY THOSE WHO RECEIVED HM3 BEFORE AND AFTER THE HEART TRANSPLANT ALLOCATION CHANGE. THE OVERALL INCIDENCE OF DEATH IN THIS STUDY¿S COHORT AT 1 AND 2 YEARS WAS 17.4% AND 22.3%, RESPECTIVELY. IN COMPARISON, THE 1- AND 2-YEAR MORTALITY IN MOMENTUM 3 WERE 13% AND 18.2%. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 DEVICES AND THE REPORTED OUTCOMES COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RESEARCH ARTICLE TITLED ¿PREDICTIVE ACCURACY OF HEARTMATE 3 RISK SCORE AFTER THE HEART TRANSPLANT ALLOCATION CHANGE¿ WAS RECEIVED. THE ARTICLE CONCLUDED THAT, FOLLOWING THE ALLOCATION SYSTEM CHANGE, THE HM3RS DID NOT PREDICT MORTALITY IN PATIENTS. THE HEARTMATE 3 DEVICE SERIAL NUMBERS, AS WELL AS OTHER SPECIFIC CASE/PATIENT INFORMATION, ARE NOT AVAILABLE. THE SERIAL NUMBER OR OTHER IDENTIFYING INFORMATION OF THE PRODUCT WAS NOT REPORTED AND WAS NOT ABLE TO BE DETERMINED DURING THE INVESTIGATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. VINCENT, J., YANG, B. Q., DEYCH, E., & VADER, J. (2024). PREDICTIVE ACCURACY OF HEARTMATE 3 RISK SCORE AFTER HEART TRANSPLANT ALLOCATION CHANGE. JOURNAL OF CARDIAC FAILURE, 30(1), 172¿173. HTTPS://DOI.ORG/10.1016/J.CARDFAIL.2023.10.136. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21529693&pc=NPT | 2015691-2025-01540 | 2025-02-13 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-05 05:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA VALVE | Adverse Event Without Identified Device or Use Problem | Vascular Dissection | P140031 |  | 1 | Event Description: AS REPORTED BY OUR EDWARDS LIFESCIENCES JAPANESE AFFILIATE, DURING TAVR PROCEDURE OF A 26MM SAPIEN 3 ULTRA RESILIA VALVE, AORTIC DISSECTION OCCURRED AND EVENTUALLY THE PATIENT PASSED AWAY. THE VALVE WAS DEPLOYED IN 80:20 AORTIC/VENTRICULAR POSITION WITH 1ML LESS THAN NOMINAL VOLUME. ALTHOUGH THERE WERE NO PARTICULAR ABNORMALITIES IMMEDIATELY AFTER VALVE DEPLOYMENT, AN ECHO REVEALED THAT ACCEDING AORTIC DISSECTION WHILE PREPARING TO RETURN TO THE OPERATING ROOM. SINCE THE PATIENT'S VITAL SIGNS WERE STABLE, IT WAS DECIDED THAT TREATMENT PLAN WOULD BE DECIDED AFTER TAKING A CT SCAN. ONE DAY AFTER TAVR PROCEDURE, THE SALES REP RECEIVED THE INFORMATION THAT THE PATIENT WAS EVENTUALLY PASSED AWAY AFTER OPEN HEART SURGERY FOR TREATMENT OF AORTIC DISSECTION. PER THE MEDICAL OPINION, DISSECTION OF THE STJ (3 MM ABOVE THE RCA) WAS OBSERVED DURING OPEN HEART SURGERY. THE PHYSICIAN WONDERS IF IT WOULD HAVE BEEN BETTER TO EXPAND THE 26MM VALVE WITH -2ML LESS THAN NOMINAL VOLUME. ADDITIONAL INFORMATION WAS OBTAINED FROM THE SALES REP THAT THE AORTIC DISSECTION WAS TRIGGERED BY STJ INJURY DURING VALVE DILATION. IT WAS NARROW STJ CASE AND CALCIFICATION OF THE AORTIC VALVE AND SINUSES OF VALSALVA WAS MINIMAL. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. ACCORDING TO THE DEVICE TRAINING MANUALS, RISK FACTORS FOR AORTIC DISSECTION, CARDIAC/AORTIC HEMATOMA, OR ANNULAR RUPTURE DURING THE TVR PROCEDURE INCLUDE SIGNIFICANT VALVE OVER SIZING, SEVERELY OBLITERATED SINUSES OF VALSALVA, PORCELAIN AORTA AND/OR PRESENCE OF BULKY CALCIFICATION, AND NARROW CALCIFIED SINOTUBULAR JUNCTION. IN ADDITION, ADVANCED AGE, FEMALE GENDER, SMALL BODY WEIGHT, AND STEROID DEPENDENCY CAN ALSO BE CONTRIBUTING FACTORS. THE SAPIEN TRANSCATHETER HEART VALVE (ALL MODELS) RELIES CALCIUM TO SECURELY ANCHOR IN THE LANDING ZONE. DESPITE THIS BENEFICIAL ASPECT OF CALCIUM, BULKY CALCIUM CAN INCREASE THE RISK OF CALCIFIC NODULE DISPLACEMENT INTO THE VASCULATURE, WHICH CAN LEAD TO VASCULAR INJURY. AT TIMES, THE EXTENT AND DISTRIBUTION OF CALCIUM CAN IMPAIR EASE OF DELIVERY OF THE VALVE, CORRECT POSITIONING OF THE VALVE, DEPLOYMENT OF THE VALVE, AND PROCEDURAL SUCCESS. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST/INDICATE THAT THE AORTIC DISSECTION WAS TRIGGERED BY STJ INJURY DURING VALVE DILATION, WHICH MAY HAVE CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21529810&pc=FKX | 0008030665-2025-00526 | 2025-02-24 05:00 | Death | ERIKA DE REYNOSA, S.A. DE C.V. | 2025-03-05 05:00 | FKX | LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | K222318 |  | 1 | Event Description: IT WAS REPORTED THAT THIS PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON CONTINUOUS CYCLIC PERITONEAL DIALYSIS [CC(PD)] UTILIZING A LIBERTY SELECT CYCLER FOR RENAL REPLACEMENT THERAPY (RRT) EXPIRED ON (B)(6) 2025 WHILE UNDERGOING CCPD THERAPY. NO ADDITIONAL INFORMATION WAS PROVIDED DURING THE INTAKE PROCESS. FOLLOW-UP WITH THE PATIENT¿S PD REGISTERED NURSE (PDRN) REPORTED THE PATIENT EXPIRED PEACEFULLY IN HIS SLEEP DUE TO A CARDIAC ARREST ON (B)(6) 2025. THE PDRN STATED THE PATIENT'S TREATMENT WAS COMPLETED WITHOUT ANY KNOWN ALARMS/ISSUES (VERIFIED THROUGH TREATMENT DATA). THE PDRN CONFIRMED THE PATIENT'S PRIMARY CAUSE OF DEATH WAS CARDIAC ARREST, SECONDARY TO THE PATIENT'S ADVANCED AGE AND CARDIAC COMORBID CONDITIONS. PER THE PDRN, THE SERIOUS ADVERSE EVENTS WERE UNRELATED TO THE PATIENT'S UTILIZATION OF ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S).Event Description: IT WAS REPORTED THAT THIS PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON CONTINUOUS CYCLIC PERITONEAL DIALYSIS [CC(PD)] UTILIZING A LIBERTY SELECT CYCLER FOR RENAL REPLACEMENT THERAPY (RRT) EXPIRED ON (B)(6) 2025 WHILE UNDERGOING CCPD THERAPY. NO ADDITIONAL INFORMATION WAS PROVIDED DURING THE INTAKE PROCESS. FOLLOW-UP WITH THE PATIENT¿S PD REGISTERED NURSE (PDRN) REPORTED THE PATIENT EXPIRED PEACEFULLY IN HIS SLEEP DUE TO A CARDIAC ARREST ON (B)(6) 2025. THE PDRN STATED THE PATIENT¿S TREATMENT WAS COMPLETED WITHOUT ANY KNOWN ALARMS/ISSUES (VERIFIED THROUGH TREATMENT DATA). THE PDRN CONFIRMED THE PATIENT¿S PRIMARY CAUSE OF DEATH WAS CARDIAC ARREST, SECONDARY TO THE PATIENT¿S ADVANCED AGE AND CARDIAC COMORBID CONDITIONS. PER THE PDRN, THE SERIOUS ADVERSE EVENTS WERE UNRELATED TO THE PATIENT¿S UTILIZATION OF ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S). Manufacturer Narrative: THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL REVIEW: A TEMPORAL RELATIONSHIP EXISTS BETWEEN CCPD THERAPY UTILIZING A LIBERTY SELECT CYCLER, AND THE SERIOUS ADVERSE EVENTS OF CARDIAC ARREST AND DEATH, AS THE PATIENT WAS UNDERGOING CCPD THERAPY WHEN HE EXPIRED. THE DEFINITIVE CAUSE OF THE PATIENT¿S CARDIAC ARREST IS UNKNOWN; THEREFORE, CAUSALITY CANNOT FIRMLY BE ESTABLISHED. HOWEVER, THE PATIENT¿S PDRN REPORTED THE PATIENT WAS ELDERLY AND POSSESSED A SIGNIFICANT CARDIAC HISTORY. THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. CARDIOVASCULAR DISEASE (OR SUDDEN CARDIAC DEATH) ACCOUNTS FOR THE MOST DEATHS AMONG THE ESRD POPULATION. ADDITIONALLY, ADULTS WITH ESRD HAVE MORTALITY RATES UP TO 30-FOLD HIGHER THAN THE GENERAL POPULATION. PER THE PDRN, THE SERIOUS ADVERSE EVENTS WERE UNRELATED TO THE PATIENT¿S UTILIZATION OF ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S). BASED ON THE INFORMATION AVAILABLE, THE LIBERTY SELECT CYCLER CAN BE DISASSOCIATED FROM THE SERIOUS ADVERSE EVENTS. THERE IS NO ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) CAUSED OR CONTRIBUTED TO THE SERIOUS ADVERSE EVENTS. FURTHERMORE, THERE WAS NO REPORT A FRESENIUS PRODUCT(S) AND/OR DEVICE(S) FAILED TO MEET THE USERS¿ EXPECTATIONS AND/OR THE MANUFACTURERS¿ SPECIFICATIONS. Manufacturer Narrative: ADDITIONAL INFORMATION PROVIDED IN D9 AND H3. PLANT INVESTIGATION: A VISUAL INSPECTION OF THE RETURNED CYCLER EXTERIOR SHOWED SIGNS OF PHYSICAL DAMAGE: DAMAGED/CRACKED BOTTOM COVER. THERE WERE NO VISUAL INDICATIONS OF DRIED FLUID WITHIN THE CASSETTE COMPARTMENT ON THE PUMP. THERE WERE VISUAL INDICATIONS OF PARTICULATES WITHIN THE CASSETTE AREA. THERE WERE NO BURRS OR SHARP EDGES IN CASSETTE AREA THAT MAY HAVE PUNCTURED A CASSETTE MEMBRANE. AN (AS-RECEIVED WITH REDUCED DWELL) SIMULATED TREATMENT WAS PERFORMED AND COMPLETED. VALVE ACTUATION TEST PASSED. SYSTEM AIR LEAK TEST PASSED. A REVIEW OF THE DEVICE MANUFACTURING RECORDS WAS CONDUCTED BY THE MANUFACTURER. THERE WERE NO DEVIATIONS OR NON-CONFORMANCES DURING THE MANUFACTURING PROCESS. IN ADDITION, A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND VERIFIED THAT THE RESULTS OF THE IN-PROGRESS AND FINAL QUALITY CONTROL (QC) TESTING MET ALL REQUIREMENTS. UPON COMPLETION OF THE EVALUATION, THERE WERE NO MALFUNCTIONS THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE CYCLER PERFORMED AS DESIGNED AND AN ASSOCIATED CAUSE COULD NOT BE DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21529883&pc=DSQ | 2916596-2025-01398 | 2023-05-10 04:00 | Death | THORATEC CORPORATION | 2025-03-05 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Pleural Effusion; Renal Failure; Sepsis; Multiple Organ Failure; Heart Failure/Congestive Heart Failure; Liver Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE "OUTCOMES OF HEARTMATE 3 IN PEDIATRIC PATIENTS WITH END-STAGE HEART FAILURE: A SINGLE-CENTER PRELIMINARY EXPERIENCE FROM TURKEY" THAT HEARTMATE 3 (HM3) MAY BE ASSOCIATED WITH PLEURAL EFFUSION, CARDIAC TAMPONADE, POLYURIA, POSITIVE BLOOD CULTURES, THROMBUS, INFECTION, RIGHT VENTRICULAR FAILURE AND DEATH. PATIENT #3 EXPERIENCED AN ICU STAY OF 28 DAYS, PLEURAL EFFUSION, RIGHT HEART FAILURE, LIVER FAILURE, KIDNEY FAILURE, MYOCYTE HYPERTROPHY, AND DEATH DUE TO SEPSIS.Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE "OUTCOMES OF HEARTMATE 3 IN PEDIATRIC PATIENTS WITH END-STAGE HEART FAILURE: A SINGLE-CENTER PRELIMINARY EXPERIENCE FROM TURKEY" THAT SEVERAL HEARTMATE 3 PEDIATRIC PATIENTS WERE EVALUATED FOR PALLIATION WITH END-STAGE HEART FAILURE. PATIENT 3 EXPERIENCED AN INTENSIVE CARE UNIT (ICU) STAY OF 28 DAYS, PLEURAL EFFUSION, RIGHT HEART FAILURE, AND DEATH DUE TO SEPSIS. THE PATIENT HAD NOTED PRE-EXISTING HEPATIC AND LIVER FAILURE PRIOR TO IMPLANT.Event Description: ADDITIONAL INFORMATION STATED THAT THE PATIENT UNDERWENT LEFT VENTRICULAR ASSIST DEVICE (LVAD) IMPLANTATION UNDER EMERGENCY CONDITIONS WHEN THEIR PEDIMACS SCORE WAS 1 AND THEY HAD ADVANCED HEART FAILURE. ON THE 7TH POSTOPERATIVE DAY, BORELLIA BULDOGFERIA GROWTH WAS DETECTED IN TRANSTRACHEAL ASPIRATION. ON THE 8TH POSTOPERATIVE DAY, THE SAME AGENT GREW IN THE CENTRAL VENOUS CATHETER. ON THE 9TH DAY AFTER SURGERY, BLOOD CULTURE REVEALED GROWTH OF STAPHYLOCOCCUS EPIDERMIDIS. THEY PASSED AWAY DUE TO MULTIORGAN FAILURE FOLLOWING SEPSIS. Manufacturer Narrative: SECTION B3: DATE OF EVENT AND DATE OF DEATH CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE HEARTMATE 3 LVAS IFU, REV. G AND THE HEARTMATE 3 PATIENT HANDBOOK, REV. G ARE CURRENTLY AVAILABLE. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING INFECTION (LOCAL, DRIVELINE, PUMP POCKET), SEPSIS, RIGHT HEART FAILURE, AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿ANTICOAGULATION¿), PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) RANGE, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. THIS SECTION ALSO LISTS INFECTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. SECTION 6 OF THE IFU (UNDER "RIGHT HEART FAILURE") DISCUSSES THE POTENTIAL DEVELOPMENT OF RIGHT HEART FAILURE FOLLOWING IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS, INCLUDING INOTROPES AND RIGHT VENTRICULAR ASSIST DEVICE (RVAD) PLACEMENT. SEVERAL SECTIONS OF THE IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SECTION A, D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. SECTION D4: DEVICE SERIAL NUMBER AND LOT NUMBER WERE NOT PROVIDED, AND EXPIRATION DATE AND MANUFACTURE DATE (H4) CANNOT BE DETERMINED. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION. DOGAN, E., ULGER TUTAR, Z., TUNCER, O. N., LEVENT, R. E., ENGIN, Ç., YAGDI, T., ATAY, Y., & ÖZBARAN, M. (2024). OUTCOMES OF HEARTMATE 3 IN PEDIATRIC PATIENTS WITH END-STAGE HEART FAILURE: A SINGLE-CENTER PRELIMINARY EXPERIENCE FROM TURKEY. FRONTIERS IN CARDIOVASCULAR MEDICINE, 11. HTTPS://DOI.ORG/10.3389/FCVM.2024.1472663 NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: SECTION B2: OUTCOMES ATTRIBUTED TO ADVERSE CORRECTED. SECTION D1: BRAND NAME CORRECTED. SECTION D4: MODEL NUMBER AND CATALOG NUMBER CORRECTED. SECTION E: INITIAL REPORTER INFORMATION CORRECTED. SECTION G2: REPORT SOURCE CORRECTED. SECTION H1: TYPE OF REPORTABLE EVENT CORRECTED. SECTION H6: HEALTH EFFECT - CLINICAL CODE CORRECTED. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21530026&pc=NKM | 2135147-2025-01119 | 2009-01-01 05:00 | Death | ABBOTT MEDICAL | 2025-03-05 05:00 | NKM | MITRACLIP® SYSTEM CLIP DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P100009 |  | 1 | Event Description: THE ARTICLE "MACHINE-LEARNING PHENOTYPING OF PATIENTS WITH FUNCTIONAL MITRAL REGURGITATION UNDERGOING TRANSCATHETER EDGE-TO-EDGE REPAIR - THE MITRA-AI STUDY" WAS REVIEWED. THE ARTICLE PRESENTED A PROSPECTIVE MULTI CENTER STUDY, TO USE MACHINE-LEARNING (ML) APPROACHES TO UNCOVER CONCEALED CONNECTIONS BETWEEN CLINICAL, ECHOCARDIOGRAPHIC, AND HEMODYNAMIC DATA ASSOCIATED WITH PATIENTS' OUTCOMES. DEVICES MENTIONED INCLUDE MITRACLIP. THE ARTICLE CONCLUDED THAT ML ANALYSIS IDENTIFIED MEANINGFUL CLINICAL PHENOTYPIC PRESENTATIONS IN FUNCTIONAL MITRAL REGURGITATION UNDERGOING TEER, WITH SIGNIFICANT DIFFERENCES IN TERMS OF CARDIOVASCULAR DEATH AND HEART FAILURE HOSPITALIZATIONS, CONFIRMED IN AN EXTERNAL VALIDATION COHORT. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS DR. FABRIZIO D¿ASCENZO AT CITTA DELLA SALUTE E DELLA SCIENZA HOSPITAL, TURIN, ITALY]. THE TIME FRAME OF THE STUDY WAS 2009 TO 2020. A TOTAL OF 824 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH ALL RECEIVED AN ABBOTT DEVICE. COMORBIDITIES INCLUDED HEART FAILURE, DIABETES, MYOCARDIAL INFARCTION, PAROXYSMAL ATRIAL FIBRILLATION, PERSISTENT ATRIAL FIBRILLATION, HYPERTENSION, HYPERLIPIDEMIA, SMOKING, COPD, PERIPHERAL ARTERY DISEASE, CORONARY ARTERY DISEASE, CHRONIC KIDNEY DISEASE, CARDIAC RESYNCHRONIZATION THERAPY, HEART FAILURE MEDICATION, MITRAL REGURGITATION. PERI AND POST-PROCEDURAL COMPLICATIONS INCLUDED DEATH, HEART FAILURE HOSPITALIZATION. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF MITRACLIP WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING HEART FAILURE, DIABETES, MYOCARDIAL INFARCTION, PAROXYSMAL ATRIAL FIBRILLATION, PERSISTENT ATRIAL FIBRILLATION, HYPERTENSION, HYPERLIPIDEMIA, SMOKING, COPD, PERIPHERAL ARTERY DISEASE, CORONARY ARTERY DISEASE, CHRONIC KIDNEY DISEASE, CARDIAC RESYNCHRONIZATION THERAPY, HEART FAILURE MEDICATION, MITRAL REGURGITATION. COMPLICATIONS REPORTED INCLUDED DEATH, HEART FAILURE HOSPITALIZATION; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. B2 DEATH DATE IS ESTIMATED. B3 EVENT DATE IS ESTIMATED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21530789&pc=OZD | 1220648-2025-26532 | 2025-02-11 05:00 | Death | ABIOMED, INC. | 2025-03-05 05:00 | OZD | IMPELLA CP | Positioning Failure | Cardiac Tamponade | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT IMPELLA CP WAS PLACED FOR PRIMARY PERCUTANEOUS CORONARY INTERVENTION. AFTER THE PROCEDURE WAS COMPLETED, AN TRANSTHORACIC ECHOCARDIOGRAM (TTE) WAS OBTAINED SHOWING NO AORTIC (VALVE) INSUFFICIENCY (AI) WITH IMPELLA, SO DECISION MADE TO LEAVE CP IN PLACE. DURING SHEATH EXCHANGE IMPELLA WENT DEEP INTO VENTRICLE AND QUICKLY WITHDRAWN. DURING THIS TIME PATIENT WAS LOSING PULSATILITY AND WAS VERY HYPOTENSIVE, PATIENT STARTED ON VASOACTIVE, LABS DRAWN TO RULE OUT BLEEDING AND TTE SHOWED CARDIAC TAMPONADE. THE TEAM QUICKLY PERFORMED PERICARDIOCENTESIS. A MASSIVE BLOOD TRANSFUSION PROTOCOL WAS FOLLOWED. TTE SHOWED LEFT VENTRICLE (LV) HEMATOMAS AND CONTRAST ADMINISTRATION WITH WASHOUT IN THE PERICARDIUM. UNSURE OF LV PERFORATION ETIOLOGY. THE PATIENT WAS PUT ON PALLIATIVE CARE AND AFTER THE DEVICE WAS WITHDRAWN THAT PATIENT EXPIRED SHORTLY AFTER. THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM AND IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿ SECTION: WARNINGS & CAUTIONS: WARNINGS SECTION: PRE-SUPPORT EVALUATION SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER SECTION: DIRECT AORTIC INSERTION SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION. [ADDITION FOR 5.0, 5.5 ONLY] IN INSTANCES WHERE THE IMPELLA PUMP HAS BEEN PLACED PRIOR TO PERFORMING CARDIAC SURGERY WITH AORTIC CROSS CLAMPING AND CARDIOPLEGIC ARREST, CARE SHOULD BE TAKEN WHEN MANIPULATING THE HEART WHEN THE PUMP POSITION IS FIXED WITH APPLICATION OF THE AORTIC CROSS CLAMP ACROSS THE PUMP CATHETER.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE. SECTION: WARNINGS & CAUTIONS: WARNINGS SECTION: PRE-SUPPORT EVALUATION SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER SECTION: DIRECT AORTIC INSERTION SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES ¿TO REDUCE THE RISK OF CARDIAC INJURY (INCLUDING VENTRICULAR PERFORATION), PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY IN THE SETTINGS OF ACUTE INFARCTION WITH TISSUE NECROSIS.¿ ¿TO REDUCE THE RISK OF VASCULAR INJURY, PHYSICIANS SHOULD EXERCISE CAUTION WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX PERIPHERAL VASCULAR ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: UNREPAIRED ABDOMINAL AORTIC ANEURYSM, SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM, DISSECTION OF THE ASCENDING/ TRANSVERSE/DESCENDING AORTA, CHRONIC ANATOMICAL CHANGES IN THE RELATIONSHIP OF THE AORTA/AORTIC VALVE/VENTRICULAR ALIGNMENT, SIGNIFICANT MOBILE ATHEROMATOUS DISEASE IN THE THORACIC OR ABDOMINAL AORTA OR PERIPHERAL VESSELS.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING A RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY IN THE LEFT VENTRICLE AFTER CPR WITH ECHOCARDIOGRAPHY GUIDANCE. SECTION: WARNINGS & CAUTIONS: CAUTIONS ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ Manufacturer Narrative: THE INVESTIGATION OF VENTRICLE PERFORATION AND POSITIONING ISSUES HAS BEEN COMPLETED. THE PUMP WAS NOT RETURNED FOR INVESTIGATION, AND A DATA LOG REVIEW WAS NOT REQUIRED FOR THIS CASE. ACCORDING TO CLINICAL FINDINGS, THE LEFT VENTRICLE (LV) IS FRIABLE IN NATURE, AND DURING THE PERFORATION, THE IMPELLA DEVICE WAS INSERTED DEEPLY. THE POSITIONING ISSUE FAILURE MODE WAS REMOVED SINCE IT WAS ADDRESSED UNDER THE ROOT CAUSE OF VENTRICULAR PERFORATION. THE CAUSE OF THE VENTRICLE PERFORATION ISSUE WAS MOST LIKELY IMPROPER POSITIONING OF THE DEVICE. E4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26532 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21531237&pc=DSQ | 2916596-2025-01260 | 2024-01-01 05:00 | Death | THORATEC CORPORATION | 2025-03-05 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Pseudoaneurysm | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE "EMERGENT IN SITU FENESTRATION IN THE ASCENDING AORTA FOR THE ENDOVASCULAR REPAIR OF A LARGE PSEUDOANEURYSM: A TECHNICAL NOTE" THAT THE IN SITU FENESTRATION (ISF) TECHNIQUE MAY BE A VALID OPTION TO ADOPT IN EMERGENCY CASES IN WHICH NO OTHER TECHNICAL SOLUTIONS ARE SUITABLE. THE PURPOSE OF THIS ARTICLE WAS TO DESCRIBE AN EMERGENT ISF TECHNIQUE IN THE ASCENDING AORTA FOR THE ENDOVASCULAR REPAIR OF A LARGE PSEUDOANEURYSM USING A TRANS-SEPTAL NEEDLE DEVICE THROUGH DIRECT RIGHT COMMON CAROTID ARTERY ACCESS, IN A PATIENT WITH LEFT VENTRICULAR ASSIST DEVICE (LVAD). THE STUDY WAS PERFORMED, IN A MULTIDISCIPLINARY TEAM-WORK APPROACH, AN EMERGENT ISF TO CORRECT THE DISPLACEMENT OF A PHYSICIAN-MODIFIED THORACIC ENDOGRAFT RELEASED IN THE ASCENDING AORTA TO CORRECT A LARGE ANASTOMOTIC PSEUDOANEURYSM IN A 34-YEAR-OLD PATIENT WHO UNDERWENT ASCENDING AORTA REPLACEMENT AND SUBSEQUENT LVAD IMPLANTATION. AFTER MULTIDISCIPLINARY EVALUATION WITH THE CARDIAC SURGEONS, THE PATIENT FIRST UNDERWENT SURGICAL AORTIC VALVE AND ASCENDING AORTA REPLACEMENT. THEN, DUE TO PERSISTENCE OF SEVERE SYSTOLIC DYSFUNCTION OF THE LEFT VENTRICLE WITH POSTOPERATIVE CARDIOGENIC SHOCK, THE PATIENT UNDERWENT A MECHANICAL CIRCULATORY SUPPORT WITH EXTRA CORPOREAL MEMBRANE OXYGENATION (ECMO) AND ESCALATED THEN TO HEARTMATE 3 LEFT VENTRICULAR ASSIST DEVICE (LVAD). A THORACIC ANGIO-CT PERFORMED AT 3 MONTHS FOR THE FIRST INTERVENTION SHOWED THE PRESENCE OF A LARGE PSEUDOANEURYSM OF THE DISTAL GRAFT-AORTIC ANASTOMOSIS (14 CM MAXIMUM DIAMETER). A TRANS-SEPTAL NEEDLE DEVICE INSERTED THROUGH A DIRECT ACCESS TO THE RIGHT CAROTID ARTERY AND PERFORMED AN ISF TO RESTORE THE PATENCY OF THE OUTFLOW OSTIUM OF THE LVAD. WINDOW WAS THEN COMPLETED AND STABILIZED WITH A NITINOL BALLOON EXPANDABLE COVERED STENT GRAFT OBTAINING AN EFFECTIVE EXCLUSION OF THE ANASTOMOTIC AORTIC ANEURISM AND THE REGULAR PATENCY OF THE LVAD OUTFLOW GRAFT WITH NO SIGNS OF LEAKS. THE TOTAL DURATION OF THE INTERVENTION HAS BEEN 380 MINUTES. THE PERIOPERATIVE PERIOD WAS REGULAR WITH NORMAL LVAD FUNCTIONING, NO NEUROLOGICAL DEFECTS, AND SATISFACTORY RESULT AT ANGIO-CT SCAN (FIGURE 4). HOWEVER, THE PATIENT DID NOT MAKE IT TO HEART TRANSPLANTATION DUE TO A LATE PROGRESSIVE WORSENING OF GENERAL CONDITIONS. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: SECTION E1: REPORTER ESTABLISHMENT NAME: CORRECTION. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 DEVICE AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RESEARCH ARTICLE TITLED ¿EMERGENT IN SITU FENESTRATION IN THE ASCENDING AORTA FOR THE ENDOVASCULAR REPAIR OF A LARGE PSEUDOANEURYSM: A TECHNICAL NOTE¿ WAS RECEIVED. THE HEARTMATE 3 DEVICE SERIAL NUMBERS AS WELL AS OTHER SPECIFIC CASE/PATIENT INFORMATION, IS NOT AVAILABLE. THE SERIAL NUMBER OR OTHER IDENTIFYING INFORMATION OF THE PRODUCT WAS NOT REPORTED AND WAS NOT ABLE TO BE DETERMINED DURING THE INVESTIGATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. THE IFU LISTS POTENTIAL ADVERSE EVENTS, INCLUDING BLEEDING AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. THE IFU OUTLINES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING THE INTERNATIONAL NORMALIZED RATIO RANGE FOR PATIENTS USING THE HEARTMATE 3 LVAS, AS WELL AS THE SUGGESTED ANTICOAGULATION MODIFICATIONS IN THE EVENT THAT THERE IS A RISK OF BLEEDING. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SECTION D: DEVICE INFORMATION WAS REQUESTED BUT NOT YET PROVIDED. DATE OF EVENT HAS BEEN ESTIMATED AS THE DATE OF DATA COLLECTION WAS NOT SPECIFIED. AUTHOR INFORMATION: ANTONELLO, M., SPERTINO, A., GIULIO RODINÒ, & TARANTINI, G. (2022). EMERGENT IN SITU FENESTRATION IN THE ASCENDING AORTA FOR THE ENDOVASCULAR REPAIR OF A LARGE PSEUDOANEURYSM: A TECHNICAL NOTE. JOURNAL OF ENDOVASCULAR THERAPY, 31(3), 366¿370. HTTPS://DOI.ORG/10.1177/15266028221125587. DEPARTMENT OF CARDIAC, THORACIC, VASCULAR SCIENCES AND PUBLIC HEALTH, UNIVERSITY OF PADUA, PADUA, ITALY. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21531703&pc=OMP | 3009897021-2025-00009 | 2025-02-06 05:00 | Death | KINETIC CONCEPTS, INC | 2025-03-05 05:00 | OMP | ACTIV.A.C.¿ THERAPY SYSTEM WITH ION PROGRESS¿ REMOTE THERAPY MONITORING | Insufficient Information | Hemorrhage/Bleeding | K201571 |  | 1 | Event Description: ON 14-FEB-2025, THE FOLLOWING INFORMATION WAS PROVIDED TO KCI BY THE FDA VIA MEDWATCH REPORT MW5166398: "SUBJECT HAD A SARCOMA REMOVED FROM THE LEFT THIGH AND HAD AN [ACTIV.A.C.¿ THERAPY SYSTEM WITH ION PROGRESS¿ REMOTE THERAPY MONITORING] PLACED OVER THE WOUND. SCENE PICTURES SHOW ABUNDANT BLOOD LOSS IN THE BATHROOM AND BEDROOM SOAKING THE CARPET. SUBJECT WAS FOUND DEAD BY WIFE UPON RETURN FROM WORK IN THE AFTERNOON. AT AUTOPSY, BLOOD-SOAKED TOWEL WRAPPED AROUND THE LEG, ABUNDANT BLOOD CLOTS UNDER THE DRESSING, VERY LITTLE BLOOD 10 CC RETRIEVABLE FROM THE HEART, AND [ACTIV.A.C.¿ THERAPY SYSTEM WITH ION PROGRESS¿ REMOTE THERAPY MONITORING] FULL OF BLOOD. HE BLEEDS FROM THE WOUND APPROXIMATELY 200 CC AND WAS MANAGED AT THE ER.¿ ON (B)(6) 2025, THE FOLLOWING INFORMATION WAS PROVIDED BY THE PHYSICIAN: V.A.C.® THERAPY ALLEGEDLY CONTRIBUTED TO THE PATIENT¿S DEATH BY EXSANGUINATION AS THE GREAT SAPHENOUS VEIN WAS PUNCTURED. ON (B)(6)2025, PATIENT WAS HOSPITALIZED FOR AN EPISODE OF BLEEDING FROM THE SAME AREA WITH ESTIMATED BLOOD LOSS OF 200 CC AND HEMOGLOBIN DROPPED TO 10.6. PER REVIEW OF CLINICAL RECORDS RECEIVED: PATIENT WAS SEEN ON (B)(6) 2025 FOR FOLLOW-UP AND TO HAVE HIS VAC PLACED. ¿WOUND BED DEBRIDED. VAC PLACED - WHITE FOAM TO TUNNELING ASPECT IN 12 O'CLOCK POSITION AND GRANUFOAM TO REMAINDER OF THE WOUND BED. PERIWOUND HAS SOFT TISSUE RADIATION INJURY. WILL USE BENZOIN TO PROTECT THE SKIN FROM THE DRAPE. WILL START THE VAC ON 100MMHG CONTINUOUS, IF HE DOES WELL WITH THIS SETTING, WILL CHANGE TO 125 AT HIS NEXT VAC CHANGE ON FRIDAY. HE WAS SENT WITH ADDITIONAL CANISTERS AND AN ALTERNATE DRESSING TO BE USED IF HE HAS ANY ISSUES WITH THE VAC BETWEEN VISITS.¿ PATIENT WAS SEEN ON (B)(6) 2025 FOR FOLLOW-UP. ¿WOUND IS MEASURING SMALLER, NO AREA OF BLEEDING IDENTIFIED. THERE IS A SMALL AREA THAT LOOKS LIKE IT WAS POSSIBLY THE PROBLEMS AREA TO THE POSTERIOR/MEDIAL WOUND MARGIN. NO ACTIVE BLEEDING IDENTIFIED IN THE CLINIC TODAY. WILL HOLD VAC UNTIL NEXT WEEK¿WILL REFER TO VASCULAR FOR WOUNDS LEFT LOWER LEG IN THE SETTING OF EDEMA (LYMPHEDEMA VS VENOUS ETIOLOGY) AND TO EVALUATE FOR SOURCE OF BLEEDING.¿ PATIENT WAS SEEN ON (B)(6) 2025 FOR FOLLOW-UP. ¿NO FURTHER EPISODES OF BLEEDING SINCE HOLDING THE WOUND VAC. HE IS SCHEDULED TO SEE VASCULAR PHYSICIAN ON WEDNESDAY. WILL HOLD THE VAC UNTIL AFTER THAT APPOINTMENT. HE IS SCHEDULED FOR A NURSE VISIT ON WEDNESDAY AFTERNOON. REFERRAL TO VASCULAR FOR EDEMA AND WOUNDS TO LEFT LOWER LEG, AND FOR CONSULTATION FOR BLEEDING FROM WOUND. I DO NOT SEE A VESSEL VISIBLE IN THE WOUND BED. BLEEDING COULD BE FROM FRIABLE WOUND BED, STATUS POST RADIATION. NO WOUND ODOR TODAY, HE IS STILL ON HIS ANTIBIOTIC.¿ ADDITIONALLY, THE PHYSICIAN PERFORMED AN EXCISIONAL SKIN/SUBCUTANEOUS TISSUE DEBRIDEMENT WITH A TOTAL AREA OF 75.83 SQ CM BY CURETTE TO REMOVE VIABLE AND NON-VIABLE TISSUE/MATERIAL. VAC WAS RESUMED ON (B)(6) 2025. A DEVICE EVALUATION OF THE ACTIV.A.C.¿ THERAPY SYSTEM WITH ION PROGRESS¿ REMOTE THERAPY MONITORING IS PENDING RETURN OF THE DEVICE.Event Description: ON 15-APR-2025, A DEVICE EVALUATION WAS COMPLETED BY SOLVENTUM QUALITY ENGINEERING. ON 30-DEC-2024, THE DEVICE WAS TESTED PER QUALITY CONTROL PROCEDURE BY THE SOLVENTUM SERVICE CENTER, AND THE DEVICE PASSED AND MET SPECIFICATION PRIOR TO PATIENT PLACEMENT. ON (B)(6) 2025, THE DEVICE WAS PLACED WITH THE PATIENT. ON 15-APR-2025, THE DEVICE WAS TESTED PER QUALITY CONTROL PROCEDURE BY SOLVENTUM QUALITY ENGINEERING AND NO FAILURES WERE FOUND WITH THE DEVICE RELATED TO THIS EVENT. Manufacturer Narrative: BASED ON THE INFORMATION PROVIDED, IT CANNOT BE DETERMINED THAT THE ALLEGED PATIENT EXPIRATION IS RELATED TO THE ACTIV.A.C.¿ THERAPY SYSTEM WITH ION PROGRESS¿ REMOTE THERAPY MONITORING. THE PATIENT'S PREDISPOSITION TO BLEEDING IS ATTRIBUTED TO THE EXTENSIVE DEEP TISSUE INFECTION, WHICH RESULTED IN FRIABLE AND WEAKENED VASCULATURE. A DEVICE EVALUATION IS PENDING RETURN OF THE DEVICE. DEVICE LABELING, AVAILABLE IN PRINT AND ONLINE, STATES: CONTRAINDICATIONS DO NOT PLACE FOAM DRESSINGS OF THE V.A.C.® THERAPY SYSTEM DIRECTLY IN CONTACT WITH EXPOSED BLOOD VESSELS, ANASTOMOTIC SITES, ORGANS, OR NERVES. WARNINGS BLEEDING: WITH OR WITHOUT USING V.A.C.® THERAPY, CERTAIN PATIENTS ARE AT HIGH RISK OF BLEEDING COMPLICATIONS. THE FOLLOWING TYPES OF PATIENTS ARE AT INCREASED RISK OF BLEEDING, WHICH, IF UNCONTROLLED, COULD BE POTENTIALLY FATAL: PATIENTS WHO HAVE WEAKENED OR FRIABLE BLOOD VESSELS OR ORGANS IN OR AROUND THE WOUND AS A RESULT OF, BUT NOT LIMITED TO: SUTURING OF THE BLOOD VESSEL (NATIVE ANASTOMOSIS OR GRAFTS) / ORGAN INFECTION, TRAUMA, RADIATION, PATIENTS WITHOUT ADEQUATE WOUND HEMOSTASIS, PATIENTS WHO HAVE BEEN ADMINISTERED ANTICOAGULANTS OR PLATELET AGGREGATION INHIBITORS, PATIENTS WHO DO NOT HAVE ADEQUATE TISSUE COVERAGE OVER VASCULAR STRUCTURES. IF V.A.C.® THERAPY IS PRESCRIBED FOR PATIENTS WHO HAVE AN INCREASED RISK OF BLEEDING COMPLICATIONS, THEY SHOULD BE TREATED AND MONITORED IN A CARE SETTING DEEMED APPROPRIATE BY THE TREATING PHYSICIAN. IF ACTIVE BLEEDING DEVELOPS SUDDENLY OR IN LARGE AMOUNTS DURING V.A.C.® THERAPY, OR IF FRANK (BRIGHT RED) BLOOD IS SEEN IN THE TUBING OR IN THE CANISTER, IMMEDIATELY STOP V.A.C.® THERAPY, LEAVE DRESSING IN PLACE, TAKE MEASURES TO STOP THE BLEEDING AND SEEK IMMEDIATE MEDICAL ASSISTANCE. THE V.A.C.® THERAPY UNITS AND DRESSINGS SHOULD NOT BE USED TO PREVENT, MINIMIZE OR STOP VASCULAR BLEEDING. PROTECT VESSELS AND ORGANS: ALL EXPOSED OR SUPERFICIAL VESSELS AND ORGANS IN OR AROUND THE WOUND MUST BE COMPLETELY COVERED AND PROTECTED PRIOR TO THE ADMINISTRATION OF V.A.C.® THERAPY. ALWAYS ENSURE THAT V.A.C.® FOAM DRESSINGS DO NOT COME IN CONTACT WITH VESSELS OR ORGANS. USE A THICK LAYER OF NATURAL TISSUE SHOULD PROVIDE THE MOST EFFECTIVE PROTECTION. IF A THICK LAYER OF NATURAL TISSUE IS NOT AVAILABLE OR IS NOT SURGICALLY POSSIBLE, MULTIPLE LAYERS OF NON-ADHERENT DRESSING MATERIAL MAY BE CONSIDERED AS AN ALTERNATIVE, IF DEEMED BY THE TREATING PHYSICIAN TO PROVIDE A COMPLETE PROTECTIVE BARRIER. IF USING NON-ADHERENT MATERIALS, ENSURE THEY ARE SECURED IN A MANNER THAT WILL MAINTAIN THEIR PROTECTIVE POSITION THROUGHOUT THERAPY. DISCLAIMER: THIS INFORMATION IS SUBMITTED PURSUANT TO 21 CFR 803, IN COMPLIANCE WITH THE MEDICAL DEVICE REPORTING REQUIREMENT AND SHOULD NOT BE CONSIDERED TO BE AN ADMISSION THAT A KINETIC CONCEPTS, INC. PRODUCT MALFUNCTIONED, IS DEFECTIVE OR HAS CAUSED SERIOUS INJURY. Manufacturer Narrative: BASED ON THE ADDITIONAL INFORMATION PROVIDED REGARDING THE DEVICE, THE ASSESSMENT REMAINS THE SAME; IT CANNOT BE DETERMINED THAT THE ALLEGED PATIENT EXPIRATION IS RELATED TO THE ACTIV.A.C.¿ THERAPY SYSTEM WITH ION PROGRESS¿ REMOTE THERAPY MONITORING. THE DEVICE MET SPECIFICATIONS BEFORE PATIENT PLACEMENT AND NO FAILURES WERE FOUND WITH THE DEVICE RELATED TO THIS EVENT POST-PLACEMENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21531851&pc=LPB | 2029046-2025-00690 | 2025-02-10 05:00 | Death | BIOSENSE WEBSTER INC | 2025-03-05 05:00 | LPB | THERMOCOOL SMARTTOUCH SF | Adverse Event Without Identified Device or Use Problem | Cardiac Perforation | P030031 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN ATRIAL FIBRILLATION ABLATION WITH A THERMOCOOL SMARTTOUCH SF AND THE PATIENT EXPERIENCED A PERICARDIAL EFFUSION THAT REQUIRED PERICARDIOCENTESIS. THE PATIENT DIED. DURING THE REPEATED ISOLATION OF THE PULMONARY VEINS, PERICARDIAL EFFUSION WAS DETECTED. PATIENT UNDERWENT PERICARDIAL PUNCTURE, THE PATIENT WAS INTUBATED, HEMODYNAMICS WERE STABILIZED AND PATIENT WAS TRANSFERRED TO ANOTHER HOSPITAL FOR FURTHER TREATMENT. THE PATIENT EVENTUALLY DIED. ABLATION WAS PERFORMED WITH A THERMOCOOL SMARTTOUCH SF CATHETER, MAPPING WITH A PENTARAY CATHETER, AND A WEBSTER CORONARY SINUS CATHETER WAS ALSO USED. THE INTRODUCERS AND TRANSSEPTAL NEEDLE WERE FROM OTHER MANUFACTURERS. THERE WERE NO ERRORS IN THE CARTO 3 SYSTEM AND THE NGEN GENERATOR DURING THE PROCEDURE. THERE WERE NO ERRORS IN THE CATHETERS EITHER. SUBSEQUENT ANALYSIS OF THE ABLATION DATA ALSO DID NOT REVEAL ANY DEVIATIONS DURING THE ABLATION. ACCORDING TO THE DOCTOR, HE DID NOT NOTICE ANYTHING UNUSUAL DURING THE ABLATION, AND MOSTLY ASSOCIATED THE COMPLICATION WITH THE ATRIAL SEPTAL PUNCTURE. HOWEVER, DURING THE ABLATION, THE DOCTOR USED A POWER OF 60 WATTS (THE DOCTOR WAS WARNED BEFORE THE START OF ABLATION ABOUT THE RISKS OF USING SUCH A HIGH POWER AND THAT THIS IS NOT RECOMMENDED BY THE MANUFACTURER). ADDITIONAL INFORMATION WAS RECEIVED. PATIENT REQUIRED CARDIAC SURGERY OF A STERNOTOMY WITH REVISION OF THE PERICARDIAL CAVITY. THE PROCEDURAL ACT WAS KEPT ABOVE 300. THERE WERE NO SIGNS OF STEAM POP. EVIDENCE WAS RECEIVED THAT THERE MAY HAVE BEEN PERFORATION ALONG THE POSTERIOR WALL, BUT WITHOUT SIGNS OF BLEEDING. BUT THE MAIN BLEEDING WAS FROM THE RIGHT VENTRICLE, WHICH WAS DAMAGED AS A RESULT OF PERICARDIAL PUNCTURE. ONE TRANSSEPTAL PUNCTURE SITE WAS PERFORMED DURING THE PROCEDURE. Manufacturer Narrative: D4: UDI: AS THE LOT NUMBER FOR THE DEVICE INVOLVED IN THE EVENT WAS NOT PROVIDED, THE FULL UDI IS CURRENTLY NOT AVAILABLE. AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. AN EVALUATION OF THE MANUFACTURING RECORD COULD NOT BE PERFORMED AS THE REQUIRED PRODUCT IDENTIFICATION NUMBER WAS NOT PROVIDED TO COMPLETE THE EVALUATION. HOWEVER, IF THE PRODUCT OR PRODUCT ID NUMBERS ARE RECEIVED AT A LATER DATE, THE INVESTIGATION WILL BE UPDATED AS APPLICABLE. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. MANUFACTURER'S REFERENCE NUMBER: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21532534&pc=MIH | 3013164176-2025-02417 | 2023-01-01 05:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-03-05 05:00 | MIH | GORE® EXCLUDER® ILIAC BRANCH ENDOPROSTHESIS | Adverse Event Without Identified Device or Use Problem | Unspecified Infection | P020004 |  | 1 | Event Description: THE FOLLOWING PUBLICATION WAS REVIEWED BY GORE: TITLE: MID-TERM TREATMENT OUTCOMES USING EXCLUDER IBE FOR AAA COMPLICATED BY ILIAC ARTERY ANEURYSM SOURCE: JOURNAL OF THE JAPANESE SOCIETY FOR CARDIOVASCULAR SURGERY, ABSTRACTS OF THE 53RD ANNUAL MEETING, PAGE 381 OBJECTIVE: TO INVESTIGATE THE LONG-TERM OUTCOMES OF EVAR TREATMENT USING THE GORE® EXCLUDER® ILIAC BRANCH ENDOPROSTHESIS (IBE) FOR ABDOMINAL AORTIC ANEURYSM COMPLICATED BY ILIAC ARTERY ANEURYSM. METHODS: FROM SEPTEMBER 2017 TO APRIL 2022, A RETROSPECTIVE STUDY WAS CONDUCTED ON 29 CASES (27 MALES, 2 FEMALES, AVERAGE AGE 71±7 YEARS (RANGE 59-90)) OF EVAR TREATMENT USING THE IBE FOR ABDOMINAL AORTIC ANEURYSM COMPLICATED BY ILIAC ARTERY ANEURYSM. THE IBE WAS USED UNILATERALLY IN ALL CASES. THE AVERAGE FOLLOW-UP PERIOD WAS 28±18 MONTHS (RANGE 0-54 MONTHS), WITH A FOLLOW-UP RATE OF 100%. CONTRALATERAL INTERNAL ILIAC ARTERY EMBOLIZATION WAS PERFORMED IN 11 CASES (37.9%). IN ONE CASE, AN ENDURANT LEG WAS USED AS A BRIDGE DEVICE. SURVIVAL RATE, CHANGES IN ANEURYSM DIAMETER OF THE ABDOMINAL AORTA AND COMMON ILIAC ARTERY BY MORE THAN 5 MM, AND PATENCY RATE OF THE INTERNAL ILIAC ARTERY ON THE TREATED SIDE WERE INVESTIGATED. RESULTS: TECHNICAL SUCCESS WAS ACHIEVED IN 96.6% OF CASES, WITH ONE CASE OF INTERNAL ILIAC ARTERY OCCLUSION DUE TO INTRAOPERATIVE DISSECTION. INTRAOPERATIVE COMPLICATIONS INCLUDED ONE CASE OF INTERNAL ILIAC ARTERY OCCLUSION. ENDOLEAKS WERE OBSERVED IN 15 CASES: 14 OF TYPE II AND 1 OF TYPE IV (USING AN ENDURANT LEG). THERE WERE NO IN-HOSPITAL DEATHS. POSTOPERATIVE COMPLICATIONS INCLUDED GLUTEAL CLAUDICATION IN THREE CASES, ALL OBSERVED ON THE SIDE WHERE INTERNAL ILIAC ARTERY EMBOLIZATION WAS PERFORMED. ONE CASE HAD MICROEMBOLISM OF THE INTERNAL ILIAC ARTERY ON THE IBE SIDE, WHICH IMPROVED WITH CONSERVATIVE TREATMENT. TWO CASES OF LONG-TERM MORTALITY WERE OBSERVED: ONE DUE TO STENT-GRAFT INFECTION AND ONE DUE TO PNEUMONIA. REINTERVENTION WAS PERFORMED IN ONE CASE FOR ABDOMINAL AORTIC ANEURYSM RE-ENLARGEMENT, WHERE ANEURYSM RESECTION WAS CONDUCTED. AT THE FINAL FOLLOW-UP, THE ABDOMINAL AORTIC ANEURYSM DIAMETER HAD DECREASED IN SEVEN CASES, REMAINED UNCHANGED IN 19 CASES, AND INCREASED IN THREE CASES (ONE OF WHICH UNDERWENT ANEURYSM RESECTION). ILIAC ARTERY ANEURYSMS DECREASED IN SIZE IN 10 CASES AND REMAINED UNCHANGED IN 19 CASES. EXCEPT FOR FOUR CASES WHERE EVALUATION WAS NOT POSSIBLE, ALL CASES ON THE IBE-TREATED SIDE SHOWED INTERNAL ILIAC ARTERY PATENCY. THE OVERALL SURVIVAL RATES WERE 93.8% AT 24 MONTHS AND 87.5% AT 48 MONTHS. THE REINTERVENTION-FREE RATES WERE 100% AT 24 MONTHS AND 92.9% AT 48 MONTHS. CONCLUSION: THE MID-TERM TREATMENT RESULTS FOR ABDOMINAL AORTIC ANEURYSMS WITH ILIAC ARTERY ANEURYSMS USING IBE WERE FAVORABLE. Manufacturer Narrative: A1: NO PATIENT SPECIFIC DETAILS HAVE BEEN PROVIDED. THEREFORE, THE PATIENT INITIALS REFLECT THE W. L. GORE INTERNAL CASE NUMBER. B2: THE DATE OF PATIENT DEATH IS UNKNOWN. H6: A REVIEW OF THE MANUFACTURING RECORDS FOR THE DEVICE COULD NOT BE PERFORMED AS A VALID LOT NUMBER WAS NOT PROVIDED. H6: CODE D12: ACCORDING TO THE GORE® EXCLUDER® ILIAC BRANCH ENDOPROSTHESIS INSTRUCTIONS FOR USE, ADVERSE EVENTS THAT MAY OCCUR AND / OR REQUIRE INTERVENTION OR ADDITIONAL INTRAOPERATIVE PROCEDURE TIME INCLUDE, BUT ARE NOT LIMITED TO: INFECTION (E.G., ANEURYSM, DEVICE OR ACCESS SITES), DEATH. LITERATURE TITLE: MID-TERM TREATMENT OUTCOMES USING EXCLUDER IBE FOR AAA COMPLICATED BY ILIAC ARTERY ANEURYSM SOURCE: JOURNAL OF THE JAPANESE SOCIETY FOR CARDIOVASCULAR SURGERY, ABSTRACTS OF THE 53RD ANNUAL MEETING, PAGE 381 W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21511731&pc=OZP | 2032227-2025-130872 | 2025-02-04 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-03-04 05:00 | OZP | PUMP MMT-1780KPK 670G PATHWAY BLACK MG | Appropriate Term/Code Not Available | Insufficient Information | P150001 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2025. INFORMATION WAS OBTAINED FROM THE REPORTING PERSON WHILE RESPONDING TO THE SAFETY NOTIFICATION FOR THE PUMP DELIVERY VOLUME ACCURACY (DVA) DURING CHANGES IN AIR PRESSURE. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: MMT-1780KPK, MMT-397A, MMT-332A. TROUBLESHOOTING WAS NOT PERFORMED. IT WAS UNKNOWN WHETHER THE CUSTOMER HAD USED THE INSULIN PUMP SYSTEM WITHIN 48 HOURS OF THE REPORTED EVENT AND WHETHER THE CUSTOMER USED THE SMARTGUARD/AUTO MODE OF THE INSULIN PUMP. NO PRODUCT RETURN IS REQUIRED FOR MMT-1780KPK. NO PRODUCT RETURN IS REQUIRED FOR MMT-397A. NO PRODUCT RETURN IS REQUIRED FOR MMT-332A. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21512199&pc=PRC | 3006513822-2025-00011 | 2025-01-01 05:00 | Death | LUTONIX, INC | 2025-03-04 05:00 | PRC | LUTONIX 035AV DRUG COATED PTA DILATION CATHETER | Adverse Event Without Identified Device or Use Problem; Insufficient Information | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P170003 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESULTS OF A CLINICAL TRIAL THAT APPROXIMATELY FIVE MONTHS AND ELEVEN DAYS POST INDEX PROCEDURE COMPLETION USING A DRUG-COATED BALLOON CATHETER FOR TARGET LESION STENOSIS IN THE LEFT UPPER ARM BASILIC VEIN, THE SUBJECT HAD AN ADVERSE EVENT OF DEATH OF UNKNOWN CAUSE. THE PRIMARY CAUSE OF DEATH WAS UNKNOWN. REPORTEDLY, THE RELATIONSHIP OF THE DEATH WITH THE STUDY DEVICE, PROCEDURE, AND ARTERIOVENOUS ACCESS CIRCUIT WERE NOT RELATED.Event Description: IT WAS REPORTED THROUGH THE RESULTS OF A CLINICAL TRIAL THAT FIVE MONTHS AND ELEVEN DAYS POST AN INDEX PROCEDURE COMPLETION USING A DRUG-COATED BALLOON CATHETER FOR TARGET LESION STENOSIS IN THE LEFT UPPER ARM BASILIC VEIN, THE SUBJECT HAD AN ADVERSE EVENT OF DEATH OF UNKNOWN CAUSE. THE PRIMARY CAUSE OF DEATH WAS UNKNOWN. REPORTEDLY, THE DEATH WAS NOT RELATED TO STUDY DEVICE, PROCEDURE, AND ARTERIOVENOUS ACCESS CIRCUIT. Manufacturer Narrative: H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WAS PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. Manufacturer Narrative: H10: ADDITIONAL INFORMATION WAS RECEIVED, AND THE FILE WAS REASSESSED FOR REPORTABILITY AND DETERMINED TO BE NO LONGER REPORTABLE. SINCE AN INITIAL MDR WAS SUBMITTED, THEREFORE, THE FILE WILL REMAIN ASSESSED AS A DEATH. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21512356&pc=OAE | 2029046-2025-00659 | 2025-02-03 05:00 | Death | BIOSENSE WEBSTER INC | 2025-03-04 05:00 | OAE | QDOT MICRO, BI, TC, D-F | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Respiratory Failure | P210027 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN ATRIAL FIBRILLATION ABLATION, AND THE PATIENT EXPERIENCED RESPIRATORY EVENT / CARDIAC ARREST THAT RESULTED IN DEATH. THE PATIENT EXPIRED WITHIN AN HOUR AFTER THE PROCEDURE WAS COMPLETED. CASE WAS COMPLETED WITH NO SIGNS OF EFFUSION OR OTHER ADVERSE EVENTS. ONCE THEY STARTED TO WAKE THE PATIENT UP, PRESSURES STARTED TO DROP. DURING EXTUBATION, THE PATIENT WAS RESPONSIVE. AFTER EXTUBATION THE PATIENT STARTED TO CODE SO THEY TRIED TO RE-INTUBATE AND GET AN A-LINE WHICH SHOWED PRESSURES OF 120/30. LABS SHOWED A DROP IN HEMOGLOBIN (FROM AROUND ~11 TO 4). THIS NUMBER IS NOT COMPLETELY ACCURATE AS IT WAS PULLED FROM A LINE THAT WAS DILUTED. WHEN THE ANESTHESIA TEAM EXTUBATED THE PATIENT, THE PATIENT CODED. THE PHYSICIAN ADDED THAT THE PATIENT HAD A RESPIRATORY EVENT, AND THE CAUSE OF THE DEATH IS UNKNOWN SINCE THE FAMILY DECIDED NOT TO COMPLETE AN AUTOPSY. THE PHYSICIAN DID NOT THINK THE DEATH WAS CAUSED BY ANY BIOSENSE WEBSTER INC. PRODUCTS. ALL OF THE INTERVENTIONS WERE REPORTED AS UNKNOWN. HOWEVER, IT WAS NOTED THAT THEY DID PERFORM CARDIOPULMONARY RESUSCITATION (CPR), RE-INTUBATE THE PATIENT, AND CHECKED THE ECHO FOR EFFUSION. THE PHYSICIAN'S OPINION ON THE CAUSE OF THE ADVERSE EVENT IS EITHER THE PROCEDURE OR PATIENT CONDITION. THEY ARE NOT 100% SURE WHAT THE CAUSE WAS BUT THEY BELIEVED IT COULD HAVE BEEN A RESPIRATORY ARREST AFTER THE PROCEDURE OR A RETROPERITONEAL BLEED. THE PROCEDURAL ACTIVATED CLOTTING TIME (ACT) WAS ABOVE 300. Manufacturer Narrative: AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. AN EVALUATION OF THE MANUFACTURING RECORD COULD NOT BE PERFORMED AS THE REQUIRED PRODUCT IDENTIFICATION NUMBER WAS NOT PROVIDED TO COMPLETE THE EVALUATION. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. D4: UDI: AS THE LOT NUMBER FOR THE DEVICE INVOLVED IN THE EVENT WAS NOT PROVIDED, THE FULL UDI IS CURRENTLY NOT AVAILABLE. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. MANUFACTURER'S REFERENCE NUMBER: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21512369&pc=PNJ | 9612164-2025-01125 | 2025-02-28 05:00 | Death | MEDTRONIC IRELAND | 2025-03-04 05:00 | PNJ | MICRA AV | Difficult to Remove; Use of Device Problem; Device Dislodged or Dislocated; Adverse Event Without Identified Device or Use Problem | Pulmonary Embolism; Cardiac Arrest; Dyspnea; Thrombosis/Thrombus; Asystole | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE IMPLANT OF A LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) THE PATIENT DIED. INTERVENTIONS USED TO HELP THE PATIENT INCLUDED RESUSCITATIVE MEASURES, CARDIOPULMONARY RESUSCITATION, MEDICATIONS, INTUBATION AND A TEMPORARY PACER, HOWEVER THESE WERE NOT SUCCESSFUL. THE IPG WAS EXPLANTED. THE CAUSE OF DEATH IS CURRENTLY UNKNOWN HOWEVER THE PHYSICIAN SUSPECTED A BLOOD LUNG EMBOLISM.Event Description: IT WAS FURTHER REPORTED THE IMPLANTATION OF THE IPG WAS PROBLEM FREE UNTIL THE RELEASE, THERE WAS A PROBLEM WITH THE LOCK MECHANISM. THE SUTURE COULD INITIALLY ONLY BE PULLED OUT VERY SLOWLY AND WITH RESISTANCE (DESPITE PRIOR IRRIGATION AND REPEATED IRRIGATION DURING THE PROCESS), BUT COULD THEN BE PULLED MORE AND MORE CONTINUOUSLY WITH THE HEARTBEAT. IT WAS NOTED "ALL OF A SUDDEN THE DELIVERY CATHETER WAS PULLED CLOSE TO THE DEVICE AGAIN AND COULD NO LONGER BE BROUGHT AT A DISTANCE". AFTER SLOWLY PULLING OUT COMPLETELY, THE IPG DISLOCATED INTO THE LEFT PULMONARY ARTERY AND WAS RECOVERED BY SNARING.Event Description: IT WAS FURTHER REPORTED THAT THE AFTER THE LEADLESS IPG WAS RECOVERED BY THE FIRST SNARE IT WAS RELEASED AGAIN IN THE RIGHT ATRIUM, VERY MOBILE AND SNARE WAS RENEWED, BUT ONLY ON THE TINES. THEREFORE A SECOND SNARE WAS USED AND THE LEADLESS IPG WAS SUCCESSFULLY RECOVERED. A LARGE THROMBI WAS OBSERVED ON THE IPG AND THE SNARE. AIRLOCK OF THE DELIVERY SYSTEM WAS RINSED IN ORDER TO REMOVE THE THROMBI. MEDICATION WAS THEN ADMINISTERED THROUGH THE AIRLOCK. "A FEW MINUTES LATER, THE PATIENT PRESENTS WITH ACUTE DYSPNEA AND THEN CARDIAC ARREST WITH ASYSTOLE". Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE DEVICE WAS RETURNED AND ANALYZED. RETURNED PRODUCT ANALYSIS WAS PERFORMED AND NO ANOMALIES WERE FOUND. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME MICRA PRODUCT ID MC1AVR1-DELSYS (SERIAL: (B)(6)). D9: NO. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21512382&pc=DYB | 9612164-2025-01126 | 2025-02-28 05:00 | Death | MEDTRONIC IRELAND | 2025-03-04 05:00 | DYB | MICRA | Use of Device Problem; Adverse Event Without Identified Device or Use Problem | Pulmonary Embolism; Cardiac Arrest; Dyspnea; Thrombosis/Thrombus; Asystole | K132030 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE IMPLANT OF A LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) THE PATIENT DIED. INTERVENTIONS USED TO HELP THE PATIENT INCLUDED; RESUSCITATIVE MEASURES, CARDIOPULMONARY RESUSCITATION , MEDICATIONS, INTUBATION AND A TEMPORARY PACER, HOWEVER THESE WERE NOT SUCCESSFUL. THE IPG WAS EXPLANTED. THE CAUSE OF DEATH IS CURRENTLY UNKNOWN HOWEVER THE PHYSICIAN SUSPECTED A BLOOD LUNG EMBOLISM.Event Description: IT WAS FURTHER REPORTED THE IMPLANTATION OF THE IPG WAS PROBLEM FREE UNTIL THE RELEASE, THERE WAS A PROBLEM WITH THE LOCK MECHANISM. THE SUTURE COULD INITIALLY ONLY BE PULLED OUT VERY SLOWLY AND WITH RESISTANCE (DESPITE PRIOR IRRIGATION AND REPEATED IRRIGATION DURING THE PROCESS), BUT COULD THEN BE PULLED MORE AND MORE CONTINUOUSLY WITH THE HEARTBEAT. IT WAS NOTED "ALL OF A SUDDEN THE DELIVERY CATHETER WAS PULLED CLOSE TO THE DEVICE AGAIN AND COULD NO LONGER BE BROUGHT AT A DISTANCE". AFTER SLOWLY PULLING OUT COMPLETELY, THE IPG DISLOCATED INTO THE LEFT PULMONARY ARTERY AND WAS RECOVERED BY SNARING.Event Description: IT WAS FURTHER REPORTED THAT THE AFTER THE LEADLESS IPG WAS RECOVERED BY THE FIRST SNARE IT WAS RELEASED AGAIN IN THE RIGHT ATRIUM, VERY MOBILE AND SNARE WAS RENEWED, BUT ONLY ON THE TINES. THEREFORE, A SECOND SNARE WAS USED AND THE LEADLESS IPG WAS SUCCESSFULLY RECOVERED. A LARGE THROMBI WAS OBSERVED ON THE IPG AND THE SNARE. AIRLOCK OF THE DELIVERY SYSTEM WAS RINSED IN ORDER TO REMOVE THE THROMBI. MEDICATION WAS THEN ADMINISTERED THROUGH THE AIRLOCK. "A FEW MINUTES LATER, THE PATIENT PRESENTS WITH ACUTE DYSPNEA AND THEN CARDIAC ARREST WITH ASYSTOLE". Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21512465&pc=QSK | 9612515-2025-00019 | 2025-02-03 05:00 | Death | VASCUTEK LTD | 2025-03-04 05:00 | QSK | THORAFLEX HYBRID | Adverse Event Without Identified Device or Use Problem; Insufficient Information | Hemorrhage/Bleeding | P210006 |  | 1 | Event Description: DOCTOR WAS UNABLE TO CONTROL BLEEDING AT THE PROXIMAL END OF THE AORTA DUE TO CONTINUOUS FILLING OF THE FALSE LUMEN. AS AN ATTEMPT TO REGAIN CONTROL, AT RELAY PRO WAS DEPLOYED TO ELIMINATE THE DISTAL PREFUSION OF THE FALSE LUMEN. THE RELAY PRO WAS SUCCESSFULLY DEPLOYED AND SUCCESSFULLY COVERED THE FENESTRATIONS WITHIN ITS LENGTH. ADDITIONAL FENESTRATIONS WERE NOTED DISTAL TO THE RELAY GRAFT THAT WERE STILL FILLING THE FALSE LUMEN. IT IS BELIEVED THAT THIS LED TO THE BLEEDING THAT COULD NO LONGER BE CONTROLLED PROXIMALLY. PATIENT OUTCOME: THE PATIENT DID NOT SURVIVE THE SURGERY. AFTER SPEAKING WITH DOCTOR, THEY WERE NOT LEFT WITH THE IMPRESSION THAT THE DEVICE DID NOT HAVE A ROLE IN THE OUTCOME. IT WAS MENTIONED THAT HE SIMPLY WAS UNABLE TO CONTROL/STOP THE BLEEDING.Event Description: DOCTOR WAS UNABLE TO CONTROL BLEEDING AT THE PROXIMAL END OF THE AORTA DUE TO CONTINUOUS FILLING OF THE FALSE LUMEN. AS AN ATTEMPT TO REGAIN CONTROL, AT RELAY PRO WAS DEPLOYED TO ELIMINATE THE DISTAL PREFUSION OF THE FALSE LUMEN. THE RELAY PRO WAS SUCCESSFULLY DEPLOYED AND SUCCESSFULLY COVERED THE FENESTRATIONS WITHIN ITS LENGTH. ADDITIONAL FENESTRATIONS WERE NOTED DISTAL TO THE RELAY GRAFT THAT WERE STILL FILLING THE FALSE LUMEN. IT IS BELIEVED THAT THIS LED TO THE BLEEDING THAT COULD NO LONGER BE CONTROLLED PROXIMALLY. PATIENT OUTCOME: THE PATIENT DID NOT SURVIVE THE SURGERY. AFTER SPEAKING WITH DOCTOR, THEY WERE NOT LEFT WITH THE IMPRESSION THAT THE DEVICE DID NOT HAVE A ROLE IN THE OUTCOME. IT WAS MENTIONED THAT HE SIMPLY WAS UNABLE TO CONTROL/STOP THE BLEEDING. THIS REPORT IS BEING SUBMITTED AS FOLLOW UP #1 FOR MANUFACTURING REPORT NUMBER 9612515-2025-00019 TO PROVIDE EVENT CLOSURE INFORMATION FOR (B)(4). Manufacturer Narrative: MANUFACTURER NARRATIVE: CLINICAL CODE: E. 1888 - HEMORRHAGE/BLOOD LOSS/BLEEDING - DOCTOR WAS UNABLE TO CONTROL BLEEDING AT THE PROXIMAL END OF THE AORTA DUE TO CONTINUOUS FILLING OF THE FALSE LUMEN. IMPACT CODE: F. 4624 - SURGICAL INTERVENTION- AS AN ATTEMPT TO REGAIN CONTROL, AT RELAY PRO WAS DEPLOYED TO ELIMINATE THE DISTAL PREFUSION OF THE FALSE LUMEN. THE RELAY PRO WAS SUCCESSFULLY DEPLOYED AND SUCCESSFULLY COVERED THE FENESTRATIONS WITHIN ITS LENGTH. 1802- DEATH- UNFORTUNATELY, THE PATIENT DID NOT SURVIVE THE SURGERY. DEVICE PROBLEM CODE :A. 2993: ADVERSE EVENT WITHOUT IDENTIFIED DEVICE OR USE PROBLEM- ADDITIONAL INFORMATION RECEIVED FROM THE SITE ON 11 FEB 25 CONFIRMING THERE WERE NO VISUAL ISSUES RELATED TO THE GRAFT. THE GRAFT WAS SOAKED IN SALINE PRIOR TO IMPLANTATION, THOUGH THE EXACT TIME IS UNKNOWN, AND IT WAS NOT ALLOWED TO DRY OUT. THE GELATIN COATING APPEARED UNDAMAGED. THE PROCEDURE WAS EXTENDED DUE TO UNCONTROLLED BLEEDING, BUT AT NO POINT WAS THE GRAFT IDENTIFIED AS THE ISSUE. THERE WERE NO SIZING ISSUES INDICATED, AND NO BLEEDING WAS REPORTED FROM THE GELATIN PORTION OF THE THORAFLEX. COMPONENT CODE: G. 4755 - PART/COMPONENT/SUB-ASSEMBLY TERM NOT APPLICABLE. TYPE OF INVESTIGATION: B. 4110 - TREND ANALYSIS: A 4 -YEAR REVIEW OF SIMILAR COMPLAINTS THORAFLEX HYBRID SALES JAN 21 - DEC 2024 VS BLOOD OOZING BODY TYPE JAN 2021 - FEB 2025 GAVE AN OCCURRENCE RATE OF 0.014% NO TREND REQUIRING ACTION HAS BEEN IDENTIFIED. 4111 - COMMUNICATION INTERVIEW: ADDITIONAL INFORMATION RECEIVED FROM THE SITE ON 11 FEB 2025 CONFIRMING THERE WERE NO VISUAL ISSUES RELATED TO THE GRAFT. THE GRAFT WAS SOAKED IN SALINE PRIOR TO IMPLANTATION, THOUGH THE EXACT TIME IS UNKNOWN, AND IT WAS NOT ALLOWED TO DRY OUT. THE GELATIN COATING APPEARED UNDAMAGED. THE PROCEDURE WAS EXTENDED DUE TO UNCONTROLLED BLEEDING, BUT AT NO POINT WAS THE GRAFT IDENTIFIED AS THE ISSUE. THERE WERE NO SIZING ISSUES INDICATED, AND NO BLEEDING WAS REPORTED FROM THE GELATIN PORTION OF THE THORAFLEX. FURTHER INFORMATION RECEIVED FROM THE SITE ON 25 FEB 25 CONFIRMED THE PATIENT HAD A PREVIOUS TYPE A DISSECTION EXTENDING TO AND BEYOND THE CELIAC ARTERY, WHICH WAS NOT CAUSED BY THE DEVICE. THE AREAS OF BLEEDING WERE NOT FROM THE GRAFT OR DEVICE. THERE WAS A TEAR IN THE FALSE LUMEN JUST DISTAL TO THE THORAFLEX CUFF, WHICH WAS STILL BEING FED FROM A FENESTRATION IN THE DISTAL AORTA, INDICATING THAT THE PATIENT'S ANATOMY WAS WEAK. THIS FALSE LUMEN WAS TEARING AND FALLING APART JUST DISTAL TO THE CUFF. THERE WAS NO REPORT OF BLEEDING FROM THE GELATIN PORTION OF THE THORAFLEX. THORAFLEX HYBRID US IFU REF 301-213-USEN REV 2.0: BLEEDING, DEATH IS AN EXPECTED OR FORESEEABLE SIDE EFFECTS AS PER CURRENT LABELLING AND CLINICALLY WELL KNOWN. A FULL BATCH REVIEW WAS PERFORMED FOR TAG0149, AND NO ISSUES WERE IDENTIFIED DURING MANUFACTURING PROCESS FROM RAW MATERIALS TO FINISHED PRODUCTS. SCAN REVIEW WAS COMPLETED ON (B)(6) 2025, HOWEVER ONLY PRE-OP SCANS WERE RECEIVED. THE PATIENT HAS A TYPE A AORTIC DISSECTION EXTENDING TO THE AORTIC BIFURCATION. THERE IS GREATER CONTRAST OPACIFICATION OF THE FALSE LUMEN. 3331 - ANALYSIS OF PRODUCTION RECORDS: FULL BATCH REVIEW PERFORMED WHICH SHOWED NO ISSUES WITH THE DEVICE DURING MANUFACTURE 4114 - DEVICE NOT RETURNED: THE DEVICE IS NOT AVAILABLE FOR RETURN INVESTIGATION FINDINGS: C 4256- MALFUNCTION OBSERVED WITHOUT CONCLUSIVE FINDING- NO DEVICE DEFICIENCY WAS IDENTIFIED, AS A TEAR IN THE FALSE LUMEN JUST DISTAL TO THE THORAFLEX CUFF WAS STILL BEING FED FROM A FENESTRATION IN THE DISTAL AORTA, INDICATING THAT THE PATIENT'S ANATOMY WAS WEAK. INVESTIGATION CONCLUSION: D 4315 - CAUSE NOT ESTABLISHED- THE COMPLAINT INVESTIGATION OF BLOOD OOZING FROM BODY EVENT COULD NOT ESTABLISH A CAUSAL LINK TO A DEVICE DEFICIENCY. Manufacturer Narrative: MANUFACTURER NARRATIVE: CLINICAL CODE: E. 1888 - HEMORRHAGE/BLOOD LOSS/BLEEDING - DOCTOR WAS UNABLE TO CONTROL BLEEDING AT THE PROXIMAL END OF THE AORTA DUE TO CONTINUOUS FILLING OF THE FALSE LUMEN. IMPACT CODE: F. 4624 - SURGICAL INTERVENTION- AS AN ATTEMPT TO REGAIN CONTROL, AT RELAY PRO WAS DEPLOYED TO ELIMINATE THE DISTAL PREFUSION OF THE FALSE LUMEN. THE RELAY PRO WAS SUCCESSFULLY DEPLOYED AND SUCCESSFULLY COVERED THE FENESTRATIONS WITHIN ITS LENGTH. 1802- DEATH- UNFORTUNATELY, THE PATIENT DID NOT SURVIVE THE SURGERY. DEVICE PROBLEM CODE: A. 3190- INSUFFICIENT INFORMATION- AWAITING ADDITIONAL INFORMATION. COMPONENT CODE: G. 4755 - PART/COMPONENT/SUB-ASSEMBLY TERM NOT APPLICABLE. TYPE OF INVESTIGATION: B. 4110 - TREND ANALYSIS: A 4 -YEAR REVIEW OF SIMILAR COMPLAINTS THORAFLEX HYBRID SALES (B)(6)2021 - (B)(6)2024 VS BLOOD OOZING BODY TYPE (B)(6)2021 - (B)(6)2025 GAVE AN OCCURRENCE RATE OF(B)(4). NO TREND REQUIRING ACTION HAS BEEN IDENTIFIED. 4111 - COMMUNICATION INTERVIEW: THE GRAFT WAS SOAKED IN SALINE PRIOR TO IMPLANT. THE GRAFT WAS NOT ALLOWED TO DRY OUT. THERE WAS NO SIZING ISSUES INDIACTED. THE PROCEDURE WAS EXTENDED BECAUSE OF THE UNCONTROLLED BLEEDING BUT AT NO POINT WAS THE GRAFT IDENTIFIED AT THE ISSUE. AWAITING ADDITIONAL INFORMATION. 3331 - ANALYSIS OF PRODUCTION RECORDS: FULL BATCH REVIEW PERFORMED WHICH SHOWED NO ISSUES WITH THE DEVICE DURING MANUFACTURE. 4114 - DEVICE NOT RETURNED: THE DEVICE IS NOT AVAILABLE FOR RETURN. INVESTIGATION FINDINGS: C. 3233 - RESULTS PENDING COMPLETION OF INVESTIGATION. INVESTIGATION CONCLUSION: D. 11 - CONCLUSION NOT YET AVAILABLE AS INVESTIGATION IS ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21513253&pc=NVZ | 2182208-2025-01102 | 2025-02-08 05:00 | Death | MEDTRONIC, INC. | 2025-03-04 05:00 | NVZ | ADVISA MRI SURESCAN | Data Problem | Spinal Cord Injury | P980035 |  | 1 | Event Description: IT WAS REPORTED THAT THE IMPLANTABLE PULSE GENERATOR (IPG) EXHIBITED PRODUCT PERFORMANCE ISSUE POSSIBLE RELATED TO BATTERY FAILURE. PATIENT WAS HOSPITALIZED. EVENTUALLY IT WAS NOTED THAT THE PATIENT PASSED AWAY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21513582&pc=OZD | 1220648-2025-26457 | 2025-01-23 05:00 | Death | ABIOMED, INC. | 2025-03-04 05:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Disseminated Intravascular Coagulation (DIC); Sepsis | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WAS INITIALLY IMPLANTED WITH AN IMPELLA CP DEVICE AS PREOPERATIVE PLACEMENT AND THEN ESCALATED TO AN IMPELLA 5.5 DEVICE AFTER CORONARY ARTERY BYPASS GRAFT SURGERY. THE PATIENT SUBSEQUENTLY EXPIRED FROM SEPSIS AND DISSEMINATED INTRAVASCULAR COAGULATION (DIC) DUE TO INFECTION APPROXIMATELY 27 DAYS AFTER START OF SUPPORT. POSTOPERATIVE CARDIAC FUNCTION DID NOT RECOVER WELL, AND THEREFORE THE PATIENT WAS ESCALATED TO AN IMPELLA 5.5. REPORT STATED BEFORE THE EMERGENCY BYPASS, THE PATIENT HAD DEVELOPED "80% OF INFECTIONS," WHICH WAS NOTED TO BE LIKELY DUE TO THE PATIENT'S CONDITION RATHER THAN THE IMPELLA INSERTION SITE. HOWEVER, NO ALTERNATIVE CAUSE FOR THE EVENT WAS IDENTIFIED AS A LIKELY CAUSE OF SUCH EVENT. ADDITIONAL INFORMATION WAS SUBSEQUENTLY RECEIVED AND NOTED THE PATIENT'S DEMISE WAS RELATED TO THE IMPELLA 5.5 DEVICE AS THE PATIENT EXPIRED FROM SEPSIS AND DIC (NOTED AS THE CAUSE OF DEATH). Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. SECTION: TABLE 3.2 IMPELLA CATHETER COMPONENTS ¿THE INFUSION FILTER PREVENTS BACTERIAL CONTAMINATION AND AIR FROM ENTERING THE PURGE LUMEN.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21513627&pc=OZD | 1220648-2025-26459 | 2025-02-09 05:00 | Death | ABIOMED, INC. | 2025-03-04 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemolysis; Ischemia | P140003 |  | 1 | Event Description: THE COMPLAINANT HAD PLACED AN IMPELLA CP PUMP TO SUPPORT A PATIENT. THE PATIENT WAS SEEN TO HAVE SIGNS OF HEMOLYSIS WITH ELEVATED PLASMA FREE HEMOGLOBIN LABS TWICE. AN ECHOCARDIOGRAM WAS PERFORMED AND SHOWED THE IMPELLA CP WAS POSITIONED IN THE PAPILLARY MUSCLE. NO ATTEMPTS WERE MADE TO REPOSITION. THE TEAM ALSO SAW LIMB ISCHEMIA AND CONSULTED WITH THE VASCULAR TEAM IN WHICH THEY SAID THAT THE RIGHT LEG WAS ¿DEAD¿ AFTER ASSESSING DOPPLER STUDIES AND PHYSICAL ASSESSMENT. THE PATIENT DID GET BLOOD PRODUCTS PRIOR TO THE EXPLANT. THE PATIENT WAS PLACED ON EXTRACORPOREAL MEMBRANE OXYGENATION SUPPORT AND THE IMPELLA CP WAS EXPLANTED. THE DAY AFTER EXPLANT, THE PATIENT EXPIRED. THE RELATIONSHIP BETWEEN THE IMPELLA CP AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE INVESTIGATION OF HEMOLYSIS AND LIMB ISCHEMIA HAS BEEN COMPLETED. THE PRODUCT WAS RETURNED AND EVALUATED. THERE WAS BIOMATERIAL ON THE IMPELLER, BUT NO OTHER ABNORMALITIES WERE FOUND WITH THE RETURNED PRODUCT. THE DATA LOGS SHOWED NO MOTOR CURRENT SPIKES AND THEREFORE NO SIGNS OF BIOMATERIAL INGESTION. CLINICAL DETAILS STATED THAT THE PUMP WAS POSITIONED IN THE PAPILLARY MUSCLE UPON ARRIVAL TO THE SECOND HOSPITAL, BUT THE PUMP WAS NOT REPOSITIONED. THE ROOT CAUSE OF THE HEMOLYSIS WAS MOST LIKELY IMPROPER POSITIONING SINCE THE CLINICAL DETAILS CONFIRM THE POSITIONING ISSUE AND NO SIGNS OF BIOMATERIAL INGESTION WERE FOUND IN THE DATA LOGS. AS ALL THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE LIMB ISCHEMIA WAS NOT DETERMINED. Manufacturer Narrative: THE IMPELLA PUMP AND DATA STICK WERE RETURNED, AND AN EVALUATION IS UNDERWAY. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21513757&pc=OZD | 1220648-2025-26461 | 2025-02-10 05:00 | Death | ABIOMED, INC. | 2025-03-04 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hematoma | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT DURING IMPELLA CP PLACEMENT, THE PHYSICIAN ATTEMPTED THE RIGHT FEMORAL ARTERY, BUT LOST WIRE ACCESS. MICRO PUNCTURE ACCESS WAS GAINED TO LEFT FEMORAL ARTERY (LFA) AND ACCESS LOST OCCURRED AGAIN. THE PHYSICIAN REATTEMPTED MICRO PUNCTURE ACCESS TO LFA AND WAS SUCCESSFUL. THE IMPELLA CP WAS INSERTED AND STARTED. THE PERCUTANEOUS CORONARY INTERVENTION WAS PERFORMED AND ONCE COMPLETED A SIGNIFICANT HEMATOMA WAS NOTED TO THE RIGHT GROIN. THERE WAS ALSO A HEMATOMA ON THE PREVIOUS STICK SITE TO LFA. BLOOD PRODUCTS WERE ADMINISTERED. THE PATIENT EXPIRED AFTER FOUR HOURS OF SUPPORT. THE RELATIONSHIP BETWEEN THE IMPELLA CP AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE INVESTIGATION FOR THE REPORTED ACCESS SITE BLEED HAS BEEN COMPLETED. THE DEVICE WAS NOT RETURNED FOR EVALUATION. HOWEVER, CLINICAL INFORMATION WAS SUFFICIENT TO DETERMINE THE ROOT CAUSE. THE ROOT CAUSE OF THE ACCESS SITE BLEEDING ISSUE WAS MOST LIKELY USE RELATED, THE DOCTOR LOST WIRE ACCESS MULTIPLE TIMES THIS LED TO GROIN SITE HEMATOMA¿S REQUIRING BLOOD ADMINISTRATION. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM, SECTION: GENERAL PATIENT CARE CONSIDERATIONS: ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT, USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ , RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21513857&pc=BZD | 2518422-2025-023360 | 2023-12-07 05:00 | Death | RESPIRONICS, INC. | 2025-03-04 05:00 | BZD | PHILIPS CPAP DEVICE | Degraded | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21513918&pc=GDW | 3005075853-2025-01708 | 2024-02-15 05:00 | Death | ETHICON ENDO-SURGERY, LLC. | 2025-03-04 05:00 | GDW | ECHELON FLEX 60 POWERED PLUS ENDOPATH STAPLER ARTICULATING ENDOSCOPIC LINEAR CUT | Patient Device Interaction Problem | Insufficient Information | K110385 |  | 1 | Event Description: IT WAS REPORTED VIA JOURNAL ARTICLE: TITLE: PROPOSAL OF THE NOVEL PREDICTIVE MODEL FOR POSTOPERATIVE PANCREATIC FISTULA IN DISTAL PANCREATECTOMY FOR PANCREATIC TUMOR BASED ON PREOPERATIVE IMAGING PARAMETERS AND STAPLER HANDLING. AUTHORS: KOSHIRO MORINO I KENZO NAKANO I TOSHIHIKO GOTO I TAKAFUMI MACHIMOTO. CITATION: WORLD J SURG. 2024;1-11. DOI: 10.1002/WJS.12113. THE PURPOSE OF THIS SINGLE-CENTER RETROSPECTIVE STUDY WAS TO INVESTIGATE PREOPERATIVE RISK FACTORS AND ESTABLISH THE PREDICTIVE MODEL FOR POPF. FURTHER, WE ESTIMATED THE ASSOCIATION BETWEEN SURGICAL VARIABLES AND POPF, THEN EXPLORED SAFE AND FAVORABLE STAPLING TECHNIQUES FOR PANCREATIC TRANSECTION. BETWEEN 2014 AND 2022, A TOTAL OF 81 PATIENTS (37 FEMALES AND 44 MALES) INVOLVED IN THIS STUDY. TRANSECTION OF THE PANCREAS WAS PERFORMED USING BLUE/GOLD CARTRIDGE OF ENDOPATH¿ STAPLER OR GST¿ SYSTEM OF THE ECHELON FLEX¿ (ETHICON ENDOSURGERY) AND PURPLE/BLACK CARTRIDGE OF TRI-STAPLE TM TECHNOLOGY OF ENDO GIA¿ (COVIDIEN MEDTRONIC). 24 REINFORCEMENT USING POLYGLYCOLIC ACID MESH (NEOVEIL® GUNZE LTD.) WAS USED, ACCORDING TO SURGEONS' PREFERENCE, WITHOUT ANY SET INDICATIONS. AT LEAST ONE DRAIN WAS PLACED NEAR THE STUMP OF THE REMNANT PANCREAS. THE D-AMY WAS MEASURED ON POSTOPERATIVE DAY (POD) 1, 2, 3 ROUTINELY AND LATER DEPENDING ON THE CLINICAL SITUATION. THE FOLLOW-UP COMPUTED TOMOGRAPHY (CT) WAS PERFORMED POD 3-5 ROUTINELY. THE DRAIN WAS REMOVED WHEN THE CT SHOWED NO RESIDUAL FLUID AND DRAINAGE FLUID AMYLASE (D-AMY) WAS LESS THAN ONE THOUSAND ON POD 3-5. SOMATOSTATIN ANALOGS WERE NOT ADMINISTERED POSTOPERATIVELY TO PREVENT POPF. THE FOLLOW-UP DATA RELATED TO THESE PATIENTS WAS UPDATED IN OCTOBER 2023. REPORTED COMPLICATIONS INCLUDED. ECHELON FLEX¿ (ETHICON ENDOSURGERY). ECHELON FLEX BLUE/GOLD CARTRIDGE (ETHICON ENDOSURGERY). POPF (POSTOPERATIVE PANCREATIC FISTULA) (N=?). TREATMENT: DRAIN PLACEMENT, MONITORING D-AMY LEVELS, CT SCAN FOR FOLLOW-UP, AND REMOVAL OF THE DRAIN WHEN CRITERIA WERE MET MORBIDITY AS DEFINED BY CLAVIEN-DINDO CLASSIFICATION VERSION 2.0 (N=?). TREATMENT: NECESSITY FOR REPEAT SURGERY OR RE-ADMISSION. MORTALITY (N=?). TREATMENT: NOT REPORTED. IN CONCLUSION, THE NOVEL PREOPERATIVE MODEL CONSISTING OF IMAGING PARAMETERS EXHIBITED EXCELLENT PREDICTABILITY, AND FOUND THICK CARTRIDGE USE AND LONG PRE-FIRING COMPRESSION TO BE PROTECTIVE FACTORS AGAINST POPF. THIS MODEL MAY FACILITATE THE DEVELOPMENT OF PREVENTIVE SURGICAL STRATEGIES TO REDUCE POPF. Manufacturer Narrative: (B)(4). DATE SENT: 3/4/2025. D4: BATCH # UNK. D4: UDI: AS THE LOT NUMBER FOR THE DEVICE INVOLVED IN THE EVENT WAS NOT PROVIDED, THE FULL UDI IS CURRENTLY NOT AVAILABLE. D4: UDI: THE EXPIRATION DATE IS CURRENTLY NOT AVAILABLE. THEREFORE, THE FULL UDI IS CURRENTLY NOT AVAILABLE. THIS REPORT IS RELATED TO A JOURNAL ARTICLE, THEREFORE NO PRODUCT WILL BE RETURNED FOR ANALYSIS AND THE BATCH HISTORY RECORDS CANNOT BE REVIEWED AS THE LOT/BATCH NUMBER HAS NOT BEEN PROVIDED. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT: DOES THE AUTHOR/SURGEON BELIEVE THAT THE ETHICON DEVICE CAUSED OR CONTRIBUTED TO THE PATIENT COMPLICATIONS MENTIONED IN THE ARTICLE? IF YES, PLEASE EXPLAIN. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. Manufacturer Narrative: (B)(4). DATE SENT: 4/1/2025. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: NO FURTHER INFORMATION WILL BE PROVIDED BECAUSE WE CAN¿T GET ANY ADDITIONAL INFORMATION FROM THE AUTHOR.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21514317&pc=MGB | 3004182619-2025-00005 | 2025-01-24 05:00 | Death | CARDIVA MEDICAL, INC. | 2025-03-04 05:00 | MGB | CARDIVA MEDICAL VASCADE 6/7F VCS | Insufficient Information | Hemorrhage/Bleeding | P120016 |  | 1 | Event Description: THE NURSE REPORTED THAT A PATIENT UNDERWENT AN INTERVENTIONAL CORONARY PROCEDURE USING THE VASCADE 6/7F DEVICE, WHICH WAS INSERTED INTO A 6F SHEATH. DURING THE PROCEDURE, THE PATIENT RECEIVED THE ANTICOAGULANT ANGIOMAX (BIVALIRUDIN), AND THERE WERE NO KNOWN ISSUES WITH THE DEVICE, AS THE PROCEDURE WAS COMPLETED SUCCESSFULLY. AFTER THE PROCEDURE, THE PATIENT WAS TRANSFERRED TO RECOVERY, WHERE A HEMATOMA WAS OBSERVED. WHILE MANUAL COMPRESSION WAS APPLIED, REBLEEDING OCCURRED AT THE ACCESS SITE. ATTEMPTS AT CONTINUED MANUAL COMPRESSION WERE UNSUCCESSFUL, LEADING TO THE USE OF A FEMOSTOP DEVICE. THE PATIENT REQUIRED A TRANSFUSION, RECEIVING FOUR LITERS OF BLOOD PRODUCTS, AND SUBSEQUENTLY EXPIRED. Manufacturer Narrative: THE VASCADE 6/7F DEVICE WILL NOT BE RETURNED TO HAEMONETICS FOR EVALUATION; THEREFORE, NO PHYSICAL INVESTIGATION WILL BE CONDUCTED. AS PART OF OUR INVESTIGATION A REVIEW OF THE COMPLAINT WAS CONDUCTED AND FOUND THAT THE DEVICE WAS DEPLOYED SUCCESSFULLY AND THERE WAS NO REPORT OF ANY ANOMALIES WITH THE COLLAGEN DISK. TEMPORARY HEMOSTASIS WAS ACHIEVED WITH THE DEVICE. HAEMONETICS FOLLOWED UP WITH THE ACCOUNT TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT AND LEARNED THAT THE PATIENT WAS STABLE THROUGHOUT THE PROCEDURE AND HAD A HISTORY OF ANTICOAGULANT THERAPY. NO AUTOPSY WILL BE PERFORMED. A MEDICAL ASSESSMENT WAS CONDUCTED BY THE HAEMONETICS MEDICAL STAFF, AND IT WAS DETERMINED THAT THE VASCULAR CLOSURE PERFORMED ACCORDING TO THE INSTRUCTIONS FOR USE (IFU). THE COLLAGEN WAS SUCCESSFULLY DEPLOYED, ACHIEVING TEMPORARY HEMOSTASIS; HOWEVER, FINAL HEMOSTASIS WAS NOT ACHIEVED. IT WAS DETERMINED THAT WHILE THE MAJOR BLEEDING THAT NECESSITATED A BLOOD TRANSFUSION COULD POSSIBLY BE ASSOCIATED WITH THE SUBJECT VASCADE 6/7 DEVICE OR THE PERCUTANEOUS CORONARY INTERVENTION (PCI) PROCEDURE, PARTICULARLY REGARDING THE ANTICOAGULATION DOSAGE FOR THIS PATIENT. THE EXACT ROOT CAUSE OF THIS EVENT IS LIKELY MULTIFACTORIAL. IN ADDITION, THE VASCULAR CLOSURE SYSTEM (VASCADE VCS) INSTRUCTIONS FOR USE 5F AND 6/7 STATES THAT BOTH A HEMATOMA AND BLEEDING AT THE SITE ARE COMPLICATIONS THAT MAY OCCUR AND MAY BE RELATED TO THE ENDOVASCULAR PROCEDURE OR THE VASCULAR CLOSURE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21514416&pc=DSQ | 2916596-2025-01416 | 2020-10-01 04:00 | Death | THORATEC CORPORATION | 2025-03-04 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE ¿PLATELET REACTIVITY IS ASSOCIATED WITH PUMP THROMBOSIS IN PATIENTS WITH LEFT VENTRICULAR ASSIST DEVICES¿ THAT HEARTMATE 3 (HM3) MAY BE ASSOCIATED WITH BLEEDING, AND DEATH. THE PREVALENCE OF HIGH ON-TREATMENT RESIDUAL PLATELET REACTIVITY (HRPR) AND ITS IMPACT ON THROMBOTIC AND BLEEDING OUTCOMES IN LEFT VENTRICULAR ASSIST DEVICE (LVAD) PATIENTS WAS UNKNOWN. THE PREVALENCE AND POSSIBLE CLINICAL IMPACT OF HRPR IN A COHORT OF STABLE LVAD PATIENTS TREATED WITH A VITAMIN K ANTAGONIST (VKA) AND LOW-DOSE ASPIRIN WAS INVESTIGATED IN THIS STUDY. 62 LVAD PATIENTS WERE RECRUITED FOR THIS STUDY BETWEEN JANUARY 2018 AND OCTOBER 2020. THESE PATIENTS WERE RECRUITED FROM THE OUTPATIENT CLINIC OF THE DEPARTMENT OF CARDIAC SURGERY, MEDICAL UNIVERSITY OF VIENNA. THESE PATIENTS HAD A MEDIAN AGE OF 62 YEARS, WHICH RANGE FROM 55 TO 70 YEARS. THE PATIENTS IN THIS STUDY HAVE THE FOLLOWING PREEXISTING CONDITIONS; 34 PATIENTS HAD ARTERIAL HYPERTENSION, 30 HAD HYPERLIPOPROTEINEMIA, 4 HAD PERIPHERAL ARTERY DISEASE, 15 HAD TYPE 2 DIABETES, 2 WERE SMOKERS, 8 HAD PRIOR HISTORY OF STROKE AND 44 HAD PRIOR MYOCARDIAL INFARCTIONS. 56 OF THESE PATIENTS WERE MALE, AND 19 PATIENTS WERE IMPLANTED WITH HEARTWARE (HVAD) AND 43 WERE IMPLANTED WITH HEARTMATE 3 (HM3). PATIENTS WERE EXCLUDED BASED ON ASPIRIN INTOLERANCE, KNOWN BLEEDING DISORDERS, KNOWN DEFECTS IN THROMBOCYTE FUNCTION, A PLATELET COUNT OF 450.000/L, A HEMATOCRIT LEVEL OF <30%, ACUTE OR CHRONIC INFECTION, MALIGNANT PARAPROTEINEMIA, MYELOPROLIFERATIVE DISORDERS, SEVERE HEPATIC FAILURE, AND A MAJOR SURGICAL PROCEDURE WITHIN 7 DAYS BEFORE ENROLLMENT. THE CLINICAL ENDPOINTS FOR THE PATIENTS IN THIS STUDY WERE GASTROINTESTINAL BLEEDING, HEMORRHAGIC STROKE, MAJOR HEMOLYSIS, AND HEART FAILURE NOT EXPLAINED BY STRUCTURAL HEART DISEASE OR ABNORMAL PUMP PARAMETERS TOGETHER WITH AN ACCOMPANYING INTERVENTION. PUMP REPLACEMENT/EXCHANGE/DEACTIVATION, URGENT TRANSPLANT LISTING, STROKE, TRANSIENT ISCHEMIC ATTACK, OR DEATH WERE ALSO LISTED AS ENDPOINT FOR PATIENTS IN THIS STUDY. PLATELET SAMPLING WAS CONDUCTED BY HAVING THE SAME PHYSICIAN DRAW BLOOD FROM EACH PATIENT FROM AN ANTECUBITAL VEIN. THESE SAMPLES WERE IMMEDIATELY BROUGHT TO THE LAB AND CENTRIFUGED AFTER COLLECTION. PLASMA SAMPLES WERE STORED AT -80 C, PLATELET FUNCTION TESTING WAS CONDUCTED 30 MINUTES AFTER BLOOD SAMPLING. THE MULTIPLE ELECTRODE AGGREGOMETRY (MEA) WAS CONDUCTED USING A WHOLE BLOOD IMPEDANCE AGGREGOMETRY WITH A MULTIPLE ANALYZER. AGGREGATION UNITS (AU) WERE PLOTTED AGAINST TIME. THE AU AT 6 MINUTES WERE USED FOR ALL CALCULATIONS. ONE AU CORRESPONDS TO 10 AU X MINUTES. P-SELECTIN (SP-SELECTIN) WAS MEASURED ACCORDING TO THE MANUFACTURER¿S INSTRUCTIONS. HIGH RESIDUAL PLATELET REACTIVITY (HRPR) WAS DETECTED IN 29 PATIENTS, DUE TO THE NUMBER OF HM3 PATIENTS IN THIS STUDY, THERE WERE MORE HM3 PATIENTS WITH HRPR THAN PATIENTS WITH HVAD. PATIENTS THAT WERE FOUND TO HAVE HRPR, HAD AN AVERAGE MEA MEASUREMENT WERE: ARACHIDONIC ACID (AA), (AU) 26, ADENOSINE DIPHOSPHATE (ADP), (AU) 68, AND THROMBIN RECEPTOR¿ACTIVATING PEPTIDE (TRAP), (AU) 107. PATIENTS WITH NO HRPR, HAD AN AVERAGE MEA MEASUREMENT WERE: AA (AU) 15, ADP (AU) 56, AND TRAP (AU) 78. SOLUBLE P-SELECTIN WAS MEASURED AT 40.6 FOR PATIENTS WITH NO HRPR, AND 47.89 FOR PATIENTS WITH HRPR. OVER 24 MONTHS, PATIENTS WITHOUT HRPR HAD ADVERSE EVENTS AT THE FOLLOWING RATES: 4 PATIENTS AT 6 MONTHS, 6 AT 12 MONTHS, 8 AT 24 MONTHS, 6 HAD A MAJOR BLEEDING EVENT AT 24 MONTHS, NO PATIENTS HAD PUMP THROMBOSIS, AND 5 PATIENTS HAD PASSED AWAY. PATIENTS THAT HAD HRPR HAD ADVERSE EVENTS AT THE FOLLOWING RATES: 3 PATIENTS AT 6 MONTHS, 3 AT 12 MONTHS, 5 AT 24 MONTHS, 4 HAD A MAJOR BLEEDING EVENT AT 24 MONTHS, 3 PATIENTS HAD PUMP THROMBOSIS, AND 4 PATIENTS HAD PASSED AWAY. THE 10 PATIENTS THAT HAD A BLEEDING EVENT HAD MET THE CLINICAL ENDPOINT OF THE STUDY. 1 PATIENT WITH PUMP THROMBOSIS DUE TO SYSTEMIC EMBOLISM, AND 2 PATIENTS NEEDING PUMP REPLACEMENTS, WHICH ALSO MET THE CLINICAL ENDPOINT. PATIENTS WITH THROMBUS HAD HIGHER MEASUREMENTS WITH THE MEA AA. THERE WERE NO SIGNIFICANT DIFFERENCES IN BLEEDING EVENTS BETWEEN HM3 AND HVAD DURING THE FOLLOW UP PERIOD, BUT HVAD HAD HIGHER RATES OF PUMP THROMBUS AND MEA AA. ALL THROMBUS EVENTS OCCURRED WITH PATIENTS WITH HRPR. PATIENTS WITH HRPR ALSO SHOWED INCREASED PLATELET REACTIVITY TO ADP AND TRAP AS WELL AS HIGHER LEVELS OF SP-SELECTIN. THERE WAS NO SIGNIFICANT DIFFERENCE IN THE NUMBER OF DEATHS BETWEEN HM3 AND HVAD PATIENTS. WITH HM3 PATIENTS, 7 HAD BLEEDING EVENTS, 6 HAD PASSED AWAY, AND 16 HAD HRPR. PLATELET REACTIVITY WAS ASSOCIATED WITH PUMP THROMBOSIS IN LVAD PATIENTS. HRPR MAY REPRESENT A RISK MARKER FOR PUMP THROMBOSIS, PARTICULARLY IN HVAD PATIENTS. LARGE PROSPECTIVE CLINICAL TRIALS ARE NEEDED TO FURTHER STUDY THE CLINICAL IMPACT OF HRPR IN LVAD PATIENTS AND EVALUATE DIFFERENCES BETWEEN AVAILABLE PUMP TYPES AS WELL AS POTENTIAL THERAPEUTIC STRATEGIES. Manufacturer Narrative: SECTION A, D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. SECTION D4: DEVICE SERIAL NUMBER AND LOT NUMBER WERE NOT PROVIDED, AND EXPIRATION DATE AND MANUFACTURE DATE (H4) CANNOT BE DETERMINED. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION. MUTSCHLECHNER, D., TSCHARRE, M., WITTMANN, F., KITZMANTL, D., SCHLÖGLHOFER, T., WADOWSKI, P. P., LAUFER, G., EICHELBERGER, B., LEE, S., WIEDEMANN, D., PANZER, S., ZIMPFER, D., & GREMMEL, T. (2024). PLATELET REACTIVITY IS ASSOCIATED WITH PUMP THROMBOSIS IN PATIENTS WITH LEFT VENTRICULAR ASSIST DEVICES. RESEARCH AND PRACTICE IN THROMBOSIS AND HAEMOSTASIS, 8(6), 102564. HTTPS://DOI.ORG/10.1016/J.RPTH.2024.102564 LANDESKLINIKUM MISTELBACH-GÄNSERNDORF, MISTELBACH, AUSTRIA; KARL LANDSTEINER SOCIETY, ST. PÖLTEN, AUSTRIA; KARL LANDSTEINER UNIVERSITY OF HEALTH SCIENCES, KREMS, AUSTRIA. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED PATIENT DEATHS COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. G, IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. THE SERIAL NUMBERS OR OTHER IDENTIFYING INFORMATION OF THE PRODUCTS WERE NOT REPORTED AND WERE NOT ABLE TO BE DETERMINED DURING THE INVESTIGATION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21514484&pc=QBJ | MW5167182 |  | Death | ABBOTT DIABETES CARE INC. | 2025-03-04 05:00 | QBJ | FREESTYLE LIBRE 2 SENSOR |  | Insufficient Information |  |  | 1 | Event Description: PATIENT IS DECEASED. HCP DID CONSENT TO FOLLOW UP (B)(6), NO FURTHER INFORMATION PROVIDED OR AVAILABLE REGARDING THIS REPORT. INC-50690.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21514497&pc=DSQ | 2916596-2025-01399 | 2024-01-31 05:00 | Death | THORATEC CORPORATION | 2025-03-04 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE ¿DEFINING OPTIMAL LEFT VENTRICULAR ASSIST DEVICE SHORT-TERM OUTCOMES MAY PROVIDE INSIGHT INTO PROGRAMMATIC QUALITY ASSESSMENT¿ THAT HEARTMATE 3 (HM3) MAY BE ASSOCIATED WITH SEVERE TRICUSPID REGURGITATION, RIGHT HEART FAILURE (RHF), REOPERATION FOR INFECTION/BLEEDING, STROKE, PROLONGED RESPIRATORY FAILURE, PROLONGED/PERMANENT POSTOPERATIVE DIALYSIS, URGENT TRANSPLANTATION FOR DEVICE MALFUNCTION, NON-RECOVERY CESSATION OF SUPPORT, DEVICE EXCHANGE FOR MALFUNCTION, AND DEATH. THE ARTICLE AIMED TO GAIN A BETTER UNDERSTANDING OF THE DETERMINANTS OF OPTIMAL SHORT-TERM (180 DAYS) OUTCOMES FOR PATIENTS ON CONTINUOUS-FLOW DURABLE LEFT VENTRICULAR ASSIST DEVICE (DLVAD) SUPPORT. THIS WAS A RETROSPECTIVE COHORT ANALYSIS OF ADULT PATIENTS (AGE =19 YEARS) UNDERGOING PRIMARY CONTINUOUS-FLOW MAGNETICALLY LEVITATED DLVAD IMPLANT (HEARTMATE 3, ABBOTT, ABBOTT PARKWAY, IL) FROM 01JAN2017 THROUGH 31JAN2024 IN STS INTERMACS. FOLLOW-UP ENDED ON 31JAN2024. 1,294 PATIENTS HAD SEVERE TRICUSPID REGURGITATION. THE HIGHEST HAZARD FOR MORTALITY AFTER LVAD WAS IN THE FIRST 6 MONTHS FOLLOWING IMPLANT. AT 180 DAYS POST-IMPLANT, 13.6% OF PATIENTS EXHIBITED RHF, 7.3% REQUIRED REOPERATION FOR INFECTION/BLEEDING, 5.4% HAD STROKE, 6.4% HAD PROLONGED RESPIRATORY FAILURE, 9.9% PASSED AWAY, 1.0% REQUIRED DIALYSIS, 0.4% REQUIRED DEVICE EXPLANATION, AND 0.3% HAD CESSATION OF SUPPORT. OF THE 706 PATIENTS WITH A STROKE, MORTALITY AT 180 DAYS WAS 41.1%. PATIENTS WITH PROLONGED RESPIRATORY FAILURE AND/OR NEED FOR PROLONGED/PERMANENT POSTOPERATIVE DIALYSIS HAD 40.0% AND 41.5% MORTALITIES, RESPECTIVELY, AT 180 DAYS, WHICH WERE APPROXIMATELY 4 TO 5 TIMES THAT OF PATIENTS WITHOUT COMPLICATIONS. THE NEED FOR REOPERATION FOR BLEEDING/INFECTION WAS ASSOCIATED WITH A MORTALITY OF 18.4%. 57 PATIENTS UNDERWENT DEVICE EXCHANGE. AMONG THE ENTIRE COHORT, 69.5% (N = 9,138) OF PATIENTS HAD AN OPTIMAL OUTCOME (ALIVE WITH NO AES AT 180 DAYS), 20.3% WERE ALIVE WITH ONE OR MORE AES BY 180 DAYS (A SCORE OF 1-12.5), 7.7% OF PATIENTS HAD A MORTALITY END-POINT WITH NO TALLIED AES (SCORE OF 13.0), AND 2.5% HAD A MORTALITY END-POINT WITH ONE OR MORE AES CONTRIBUTING TO MORTALITY. AT 180 DAYS, 11,808 (89.8%) PATIENTS WERE ALIVE AND ON DEVICE SUPPORT, AND 1,340 (10.2%) PATIENTS MET A MORTALITY END-POINT. INCLUDED IN THE MORTALITY-DEFINED END-POINT WERE 1,296 DEATHS, 2 URGENT TRANSPLANTS DUE TO DEVICE MALFUNCTION, AND 42 PATIENTS WHO UNDERWENT NONRECOVERY CESSATION OF SUPPORT (E.G., HOSPICE/WITHDRAWAL OF CARE). Manufacturer Narrative: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. AUTHOR INFORMATION: COWGER, J. A., MOLINA, E., DENG, L., KANWAR, M., SHAH, P., COGSWELL, R., GOSEV, I., CANTOR, R. S., DARDAS, T. F., KIRKLIN, J. K., ROGERS, J. G., CLEVELAND, J. C., SANDAU, K. E., MCILVENNAN, C. K., KACZOROWSKI, D., ESTEP, J. D., & PAGANI, F. D. (2024). DEFINING OPTIMAL LEFT VENTRICULAR ASSIST DEVICE SHORT-TERM OUTCOMES MAY PROVIDE INSIGHT INTO PROGRAMMATIC QUALITY ASSESSMENT. THE JOURNAL OF HEART AND LUNG TRANSPLANTATION, 43(11), 1777¿1787. HTTPS://DOI.ORG/10.1016/J.HEALUN.2024.08.006. HENRY FORD HEART AND VASCULAR INSTITUTE, DETROIT, MICHIGAN. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 DEVICES AND THE REPORTED PATIENT OUTCOMES COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE HEARTMATE 3 DEVICE SERIAL NUMBERS, AS WELL AS OTHER SPECIFIC CASE/PATIENT INFORMATION, ARE NOT AVAILABLE. THE SERIAL NUMBER OR OTHER IDENTIFYING INFORMATION OF THE PRODUCT WAS NOT REPORTED AND WAS NOT ABLE TO BE DETERMINED DURING THE INVESTIGATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C IS CURRENTLY AVAILABLE. THE IFU LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21514973&pc=NPT | 2025587-2025-01612 | 2024-01-27 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-04 05:00 | NPT | EVOLUT PRO PLUS VALVE | Insufficient Information | Intracranial Hemorrhage | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 8 MONTHS AND 2 WEEKS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT EXPERIENCED A CEREBRAL HEMORRHAGE AND SUBSEQUENTLY PASSED AWAY. Manufacturer Narrative: UPDATED D.4 AND H.4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21515123&pc=OZD | 1220648-2025-26472 | 2025-02-10 05:00 | Death | ABIOMED, INC. | 2025-03-04 05:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT HAD CARE BEING ESCALATED FROM A IMPELLA CP TO THE 5.5 PUMP. THE 5.5 PUMP PLACEMENT TOOK TIME AND BLEEDING OCCURRED AT THE SITE. THE BLEEDING PROMPTED THE TEAM TO REDUCE/REVERSE THE ANTICOAGULATION. ADDITIONALLY, DURING THE DELAYED AND LONG PUMP PLACEMENT THE TEAM OBSERVED ARRHYTHMIAS. THE PATIENT REMAINED STABLE AND ON SUPPORT.Event Description: ADDITIONAL INFORMATION RECEIVED THAT THE PATIENT REMAINED ON THE 5.5 PUMP FOR 69 HOURS UNTIL THE DECISION WAS MADE TO WITHDRAW CARE AND THE PATIENT EXPIRED. IT IS UNKNOWN IF THE USE OF THE IMPELLA WAS RELATED TO PATIENT DEATH, THUS THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING AND ARRYTHMIA HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. THE ROOT CAUSE OF THE ACCESS SITE BLEEDING ISSUE WAS MOST LIKELY USE RELATED AS LENGTHY CUT DOWN/ANASTAMOSIS DUE TO BLEEDING WAS OBSERVED. AS THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE ARRYTHMIA WAS NOT DETERMINED. Manufacturer Narrative: B2 OUTCOME REVISED TO REFLECT THE ADDITIONAL INFORMATION RECEIVED FROM THE CLINICAL SITE TO INCLUDE DEATH AND DATE OF DEATH. B5 REVISED TO REFLECT THE ADDITIONAL INFORMATION RECEIVED FROM THE CLINICAL SITE. D6B EXPLANT DATE ADDED. H1 TYPE OF REPORT REVISED TO REFLECT THE ADDITIONAL INFORMATION RECEIVED FROM THE CLINICAL SITE. H6 REVISED TO REFLECT THE ADDITIONAL INFORMATION RECEIVED FROM THE CLINICAL SITE TO INCLUDE DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21515158&pc=NPT | 2025587-2025-01613 | 2024-09-17 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-04 05:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Intracranial Hemorrhage; Sepsis | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 1 YEAR, 2 MONTHS, AND 19 DAYS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS REPORTED TO BE SEPSIS AND CEREBRAL HEMORRHAGE. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21515172&pc=MKJ | 3030677-2025-000709 | 2025-02-26 05:00 | Death | PHILIPS NORTH AMERICA LLC | 2025-03-04 05:00 | MKJ | FRX (W B) DEFIB, US ENGLISH, EXCHG | Failure to Deliver Shock/Stimulation | Cardiac Arrest | P180028 |  | 1 | Event Description: IT HAS BEEN REPORTED THAT THE DEVICE DID NOT SHOCK DURING A CARDIAC EVENT. THE DEVICE SIGNALED THE ERROR "NO SHOCK ADVISED". THE DEVICE WAS IN USE ON A PATIENT HOWEVER, THE PATIENT DID NOT SURVIVE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21515198&pc=NPT | 2025587-2025-01615 | 2025-01-19 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-04 05:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Sepsis | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 10 MONTHS AND 2 WEEKS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT WAS FOUND ON THE FLOOR OF A NURSING HOME WITH A POOR COMPLEXION. NO PULSE COULD BE FOUND AND CHEST COMPRESSIONS WERE INITIATED. AN AUTOMATED EXTERNAL DEFIBRILLATOR WAS USED AND EMERGENCY SERVICES WERE CALLED. AT THE HOSPITAL ADDITIONAL LIFE-SAVING MEASURE WERE TAKEN HOWEVER, IT WAS DETERMINED THAT FURTHER RESUSCITATION WOULD BE DIFFICULT. THE PATIENT'S DEATH WAS CONFIRMED THE SAME DAY. A BLOOD TEST REVEALED AN INCREASE IN INFLAMMATORY RESPONSE INDICATING THE CAUSE OF DEATH WAS SEPSIS. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21515413&pc=NPT | 2015691-2025-01470 | 2025-02-14 05:00 | Death | EDWARDS LIFESCIENCES | 2025-03-04 05:00 | NPT | SAPIEN 3 TRANSCATHETER HEART VALVE | Degraded; Device Stenosis | Heart Failure/Congestive Heart Failure; Insufficient Information | P140031 |  | 1 | Event Description: AS REPORTED BY A FIELD CLINICAL SPECIALIST (FCS), APPROXIMATELY 5 YEARS AND 8 MONTHS POST-IMPLANT OF A 26 MM SAPIEN 3 VALVE IN THE AORTIC POSITION, THE VALVE POSSIBLY FAILED. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED FOR CORRECTION AND ADDITIONAL INFORMATION. THE FOLLOWING SECTIONS OF THIS REPORT HAVE BEEN UPDATED: CORRECTED B.2 OUTCOME ATTRIBUTED TO ADVERSE EVENT, H.1 TYPE OF REPORT H.6 IMPACT CODE, CLINICAL CODE, DEVICE CODE AND INVESTIGATION FINDINGS. ADDED NEW INFORMATION TO H.11 NARRATIVE. ADDITIONAL INFORMATION RECEIVED. THE VALVE POSSIBLY FAILED DUE STENOSIS BASED ON THE AVA OF 0.34 CM2. THE PATIENT PRESENTED WITH SOB AND FATIGUE. THE PATIENT EXPIRED DURING VALVE IN VALVE PROCEDURE. PER MEDICAL OPINION, THE CAUSE OF DEATH WAS CARDIOVASCULAR, BIVENTRICULAR FAILURE. PER THE INSTRUCTION FOR USE (IFU), VALVE STENOSIS IS A POTENTIAL RISK ASSOCIATED WITH BIOPROSTHETIC HEART VALVES. PER THE VALVE ACADEMIC RESEARCH CONSORTIUM (VARC), VALVE STENOSIS CAN RESULT FROM A NUMBER OF FACTORS, INCLUDING PANNUS, CALCIFICATION, SUPPORT STRUCTURE DEFORMATION (OUT-OF-ROUND CONFIGURATION), TRAUMA (CARDIA-PULMONARY RESUSCITATION, BLUNT CHEST TRAUMA), ENDOCARDITIS, PROSTHETIC VALVE THROMBOSIS, AND NATIVE LEAFLET PROLAPSE IMPEDING PROSTHETIC LEAFLET MOTION. VALVE STENOSIS MAY RESULT IN SYMPTOMS SUCH AS SHORTNESS OF BREATH AND DECREASED EXERCISE TOLERANCE, WHICH MAY BE ACCOMPANIED BY AN INCREASED GRADIENT ACROSS THE VALVE. PER THE INSTRUCTIONS FOR USE (IFU), STRUCTURAL VALVE DETERIORATION (WEAR, FRACTURE, CALCIFICATION, LEAFLET TEAR/TEARING FROM THE STENT POSTS, LEAFLET RETRACTION, SUTURE LINE DISRUPTION OF COMPONENTS OF A PROSTHETIC VALVE, THICKENING, STENOSIS), AND DEVICE DEGENERATION ARE KNOWN POTENTIAL RISKS ASSOCIATED WITH BIOPROSTHETIC HEART VALVES. THE IFU CAUTIONS THAT ACCELERATED DETERIORATION OF THE VALVE MAY OCCUR IN PATIENTS WITH AN ALTERED CALCIUM METABOLISM. LONG-TERM DURABILITY HAS NOT BEEN ESTABLISHED FOR THE VALVE. REGULAR MEDICAL FOLLOW-UP IS ADVISED TO EVALUATE VALVE PERFORMANCE. STRUCTURAL VALVE DETERIORATION (SVD) MAY BE MANIFESTED AS STENOSIS WITH THICKENED LEAFLETS. SVD MAY BE MILD AND NOT REQUIRE ANY INTERVENTION OR IT MAY BE MODERATE TO SEVERE. IT CAN CAUSE THE HEART TO WORK HARDER TO EJECT BLOOD FROM THE VENTRICLE. DEPENDING ON THE SEVERITY IT COULD BE AN INDICATION FOR VALVE REPLACEMENT OR MEDICAL INTERVENTION. TISSUE CALCIFICATION IS A COMMON FAILURE MODE. THE MECHANISMS FOR BIOPROSTHETIC HEART VALVE TISSUE CALCIFICATION ARE NOT YET FULLY UNDERSTOOD. MANY FACTORS CAN CONTRIBUTE TO THE ONSET AND PROPAGATION OF CALCIFICATION INCLUDING PATIENT-RELATED (E.G., PATIENT AGE, DISEASE STATE, IMMUNE STATUS, AND OTHER CO-MORBIDITIES), PHARMACOLOGICAL, AND INTRINSIC PROPERTIES OF THE VALVE ITSELF. IT IS WIDELY UNDERSTOOD THAT PATIENTS WITH CHRONIC RENAL DISEASE AND PRIOR HISTORY OF CALCIFIC STENOSIS OF THE NATIVE VALVE MAY BE PREDISPOSED TO BIOPROSTHETIC CALCIFICATION. DURING THE MANUFACTURING PROCESS, ALL SAPIEN THV VALVES ARE 100% VISUALLY INSPECTED FOR DEFECTS AND 100% FUNCTIONALLY TESTED FOR PROPER COAPTATION UNDER PHYSIOLOGICAL BACKPRESSURE CONDITIONS PRIOR TO RELEASE FOR DISTRIBUTION. THEREFORE, IT IS HIGHLY UNLIKELY THAT A MANUFACTURING DEFECT OR DEVICE MALFUNCTION CONTRIBUTED TO THE EVENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A MANUFACTURING NON-CONFORMANCE CONTRIBUTED TO THIS PRODUCT PROBLEM. INVESTIGATION RESULTS SUGGEST THE CAUSE OF THE STENOSIS IS UNKNOWN DUE TO THE LIMITED INFORMATION AVAILABLE. THE PATIENT EXPIRED DURING VALVE IN VALVE PROCEDURE. PER MEDICAL OPINION, THE CAUSE OF DEATH WAS CARDIOVASCULAR, BIVENTRICULAR FAILURE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED FOR CORRECTION AND ADDITIONAL INFORMATION. THE FOLLOWING SECTIONS OF THIS REPORT HAVE BEEN UPDATED: CORRECTED H.6 DEVICE CODE, INVESTIGATION CONCLUSIONS AND INVESTIGATION FINDINGS. PER MEDICAL RECORD REVIEW, THE PATIENT HAD SEVERE AORTIC STENOSIS, MILD AORTIC INSUFFICIENCY, AND REDUCED LEAFLET EXCURSION. PER THE INSTRUCTIONS FOR USE (IFU), STRUCTURAL VALVE DETERIORATION (WEAR, FRACTURE, CALCIFICATION, LEAFLET TEAR/TEARING FROM THE STENT POSTS, LEAFLET RETRACTION, SUTURE LINE DISRUPTION OF COMPONENTS OF A PROSTHETIC VALVE, THICKENING, STENOSIS), AND DEVICE DEGENERATION ARE KNOWN POTENTIAL RISKS ASSOCIATED WITH THE USE OF BIOPROSTHETIC HEART VALVES. THE IFU CAUTIONS THAT ACCELERATED DETERIORATION OF THE VALVE MAY OCCUR IN PATIENTS WITH AN ALTERED CALCIUM METABOLISM. LONG-TERM DURABILITY HAS NOT BEEN ESTABLISHED FOR THE VALVE. REGULAR MEDICAL FOLLOW-UP IS ADVISED TO EVALUATE VALVE PERFORMANCE. STRUCTURAL VALVE DETERIORATION (SVD) MAY BE MANIFESTED AS STENOSIS WITH THICKENED LEAFLETS. SVD MAY BE MILD AND NOT REQUIRE ANY INTERVENTION OR IT MAY BE MODERATE TO SEVERE. IT CAN CAUSE THE HEART TO WORK HARDER TO EJECT BLOOD FROM THE VENTRICLE. DEPENDING ON THE SEVERITY IT COULD BE AN INDICATION FOR VALVE REPLACEMENT OR MEDICAL INTERVENTION. TISSUE CALCIFICATION IS A COMMON FAILURE MODE. THE MECHANISMS FOR BIOPROSTHETIC HEART VALVE TISSUE CALCIFICATION ARE NOT YET FULLY UNDERSTOOD. MANY FACTORS CAN CONTRIBUTE TO THE ONSET AND PROPAGATION OF CALCIFICATION INCLUDING PATIENT-RELATED (E.G., PATIENT AGE, DISEASE STATE, IMMUNE STATUS, AND OTHER CO-MORBIDITIES), PHARMACOLOGICAL, AND INTRINSIC PROPERTIES OF THE VALVE ITSELF. IT IS WIDELY UNDERSTOOD THAT PATIENTS WITH CHRONIC RENAL DISEASE AND PRIOR HISTORY OF CALCIFIC STENOSIS OF THE NATIVE VALVE MAY BE PREDISPOSED TO BIOPROSTHETIC CALCIFICATION. DURING THE MANUFACTURING PROCESS, ALL SAPIEN THV VALVES ARE 100% VISUALLY INSPECTED FOR DEFECTS AND 100% FUNCTIONALLY TESTED FOR PROPER COAPTATION UNDER PHYSIOLOGICAL BACKPRESSURE CONDITIONS PRIOR TO RELEASE FOR DISTRIBUTION. THEREFORE, IT IS HIGHLY UNLIKELY THAT A MANUFACTURING DEFECT OR DEVICE MALFUNCTION CONTRIBUTED TO THE EVENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A MANUFACTURING NON-CONFORMANCE CONTRIBUTED TO THIS PRODUCT PROBLEM. INVESTIGATION RESULTS SUGGEST THE CAUSE OF THE STENOSIS IS UNKNOWN DUE TO THE LIMITED INFORMATION AVAILABLE. THE PATIENT EXPIRED DURING VALVE IN VALVE PROCEDURE. PER MEDICAL OPINION, THE CAUSE OF DEATH WAS CARDIOVASCULAR, BIVENTRICULAR FAILURE. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), STRUCTURAL VALVE DETERIORATION (WEAR, FRACTURE, CALCIFICATION, LEAFLET TEAR/TEARING FROM THE STENT POSTS, LEAFLET RETRACTION, SUTURE LINE DISRUPTION OF COMPONENTS OF A PROSTHETIC VALVE, THICKENING, STENOSIS), AND DEVICE DEGENERATION ARE KNOWN POTENTIAL RISKS ASSOCIATED WITH THE USE OF BIOPROSTHETIC HEART VALVES. THE IFU CAUTIONS THAT ACCELERATED DETERIORATION OF THE VALVE MAY OCCUR IN PATIENTS WITH AN ALTERED CALCIUM METABOLISM. LONG-TERM DURABILITY HAS NOT BEEN ESTABLISHED FOR THE VALVE. REGULAR MEDICAL FOLLOW-UP IS ADVISED TO EVALUATE VALVE PERFORMANCE. STRUCTURAL VALVE DETERIORATION (SVD) MAY BE MANIFESTED AS STENOSIS WITH THICKENED LEAFLETS. SVD MAY BE MILD AND NOT REQUIRE ANY INTERVENTION OR IT MAY BE MODERATE TO SEVERE. IT CAN CAUSE THE HEART TO WORK HARDER TO EJECT BLOOD FROM THE VENTRICLE. DEPENDING ON THE SEVERITY IT COULD BE AN INDICATION FOR VALVE REPLACEMENT OR MEDICAL INTERVENTION. TISSUE CALCIFICATION IS A COMMON FAILURE MODE. THE MECHANISMS FOR BIOPROSTHETIC HEART VALVE TISSUE CALCIFICATION ARE NOT YET FULLY UNDERSTOOD. MANY FACTORS CAN CONTRIBUTE TO THE ONSET AND PROPAGATION OF CALCIFICATION INCLUDING PATIENT-RELATED (E.G., PATIENT AGE, DISEASE STATE, IMMUNE STATUS, AND OTHER CO-MORBIDITIES), PHARMACOLOGICAL, AND INTRINSIC PROPERTIES OF THE VALVE ITSELF. IT IS WIDELY UNDERSTOOD THAT PATIENTS WITH CHRONIC RENAL DISEASE AND PRIOR HISTORY OF CALCIFIC STENOSIS OF THE NATIVE VALVE MAY BE PREDISPOSED TO BIOPROSTHETIC CALCIFICATION. DURING THE MANUFACTURING PROCESS, ALL SAPIEN THV VALVES ARE 100% VISUALLY INSPECTED FOR DEFECTS AND 100% FUNCTIONALLY TESTED FOR PROPER COAPTATION UNDER PHYSIOLOGICAL BACKPRESSURE CONDITIONS PRIOR TO RELEASE FOR DISTRIBUTION. THEREFORE, IT IS HIGHLY UNLIKELY THAT A MANUFACTURING DEFECT OR DEVICE MALFUNCTION CONTRIBUTED TO THE EVENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A MANUFACTURING NON-CONFORMANCE CONTRIBUTED TO THIS PRODUCT PROBLEM. INVESTIGATION RESULTS SUGGEST THE CAUSE OF THE EVENT IS UNKNOWN DUE TO THE LIMITED INFORMATION AVAILABLE. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21515479&pc=DSQ | 3007042319-2025-00595 | 2025-01-01 05:00 | Death | HEARTWARE, INC. | 2025-03-04 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P100047 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING THE OCCURRENCE OF VENTRICULAR ARRHYTHMIAS IN PATIENTS WITH PREEXISTING IMPLANTABLE CARDIOVERTER DE FIBRILLATORS (ICDS) UNDERGOING VENTRICULAR ASSIST DEVICE (VAD) IMPLANTATION. THE AUTHORS DESCRIBED PATIENT DEATHS; HOWEVER, THE CAUSES OF DEATH WERE UNKNOWN. THERE WERE PATIENTS WHO EXPERIENCED SUSTAINED VENTRICULAR ARRHYTHMIAS AND SUCTION EVENTS. THE STATUS OF THE DEVICES IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/59 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: ECHOCARDIOGRAPHIC PREDICTORS OF VENTRICULAR ARRHYTHMIAS IN PATIENTS WITH LEFT VENTRICULAR ASSIST DEVICES AND IMP LANTABLE CARDIOVERTER-DEFIBRILLATOR. JOURNAL OF CARDIOVASCULAR ELECTROPHYSIOLOGY. 2025; 36:387¿395. DOI: 10.1111/JCE.16539 D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE PRODUCT WAS DISTRIBUTED OUTSIDE OF THE UNITED STATES AND D4 UDI IS THEREFORE BLANK. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21515726&pc=DSQ | 2916596-2025-01405 | 2014-11-01 04:00 | Death | THORATEC CORPORATION | 2025-03-04 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE ¿UNPLANNED HOSPITAL READMISSIONS FOLLOWING HEARTMATE 3 IMPLANTATION: READMISSION RATES, CAUSES, AND IMPACT ON SURVIVAL¿ THAT HEARTMATE 3 (HM3) MAY BE ASSOCIATED WITH HOSPITAL READMISSION DUE TO INFECTION, STROKE, BLEEDING, ARRHYTHMIA, VOLUME OVERLOAD (DEFINED AS CLINICAL EVIDENCE OF VOLUME OVERLOAD AND NEED FOR INTRAVENOUS DIURETIC), RIGHT HEART FAILURE, DEVICE MALFUNCTION, PUMP THROMBOSIS, AND DEATH. THIS RETROSPECTIVE STUDY EVALUATED ALL 182 CONSECUTIVE PATIENTS WHO WERE IMPLANTED WITH THE HM3 LEFT VENTRICULAR ASSIST DEVICE (LVAD) AT COLUMBIA UNIVERSITY IRVING MEDICAL CENTER BETWEEN NOVEMBER 2014 TO DECEMBER 2019. ALL READMISSIONS TO THE IMPLANTING CENTER AFTER THE INDEX HOSPITAL DISCHARGE WERE REVIEWED AND THE PRIMARY REASON FOR READMISSION WAS RECORDED. 10 PATIENTS DIED DURING THE IMPLANT HOSPITALIZATION AND 5 UNDERWENT A HEART TRANSPLANTATION. THE MEDIAN AGE OF THE COHORT WAS 61.2 YEARS, 62.3% OF THE COHORT WERE WHITE, 83% WERE MALE, 62% HAD AN INTERMACS PROFILE LESS THAN OR EQUAL TO 2, AND 82% RECEIVED HM3 AS DESTINATION THERAPY. DURING A MEDIAN FOLLOW UP OF 727 DAYS, THERE WERE 407 UNPLANNED READMISSIONS. 125 PATIENTS HAD AT LEAST ONE READMISSION. 7 PATIENTS HAD AT LEAST ONE POSTOPERATIVE GASTROINTESTINAL BLEED (GIB), 6 HAD POSTOPERATIVE STROKE (4 HEMORRHAGIC, 2 ISCHEMIC), AND 21 HAD POSTOPERATIVE RIGHT VENTRICULAR ASSIST DEVICE (RVAD). 29.3% OF READMISSIONS WERE DUE TO MAJOR INFECTION, INCLUDING 10.6% FOR LVAD-RELATED INFECTION. 13.2% WERE DUE TO BLEEDING, INCLUDING 11.5% GIB AND 1.7% OTHER CAUSES. 12.5% WERE DEVICE RELATED, INCLUDING 9.1% DEVICE ALARMS, 3.4% DEVICE MALFUNCTION, AND 0.2% PUMP THROMBOSIS. 9.8% WERE FOR ARRHYTHMIA, 7.1% WERE FOR VOLUME OVERLOAD, AND 28% WERE FOR OTHER CAUSES. ONE YEAR AND 900-DAY SURVIVAL WERE UNAFFECTED BY THE NUMBER OF READMISSIONS AND REMAINED >90%. UNPLANNED HOSPITAL READMISSIONS AFTER HM3 ARE COMMON, WITH INFECTIONS AND BLEEDING ACCOUNTING FOR THE MAJORITY OF READMISSIONS. Manufacturer Narrative: SECTIONS A AND D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. SHIH, H., MONDELLINI, G. M., KURLANSKY, P. A., SUN, J., NING, Y., FELDMAN, V. R., TIBURCIO, M., MAGUIRE, C. W., LADANYI, A., CLERKIN, K., NAKA, Y., SAYER, G. T., URIEL, N., COLOMBO, P. C., TAKEDA, K., & YUZEFPOLSKAYA, M. (2024). UNPLANNED HOSPITAL READMISSIONS FOLLOWING HEARTMATE 3 IMPLANTATION: READMISSION RATES, CAUSES, AND IMPACT ON SURVIVAL. ARTIFICIAL ORGANS, 48(9), 1049¿1059. HTTPS://DOI.ORG/10.1111/AOR.14763. DEPARTMENT OF SURGERY, DIVISION OF CARDIOTHORACIC AND VASCULAR SURGERY, COLUMBIA UNIVERSITY MEDICAL CENTER, NEW YORK, NEW YORK, USA; DIVISION OF CARDIOVASCULAR MEDICINE, DEPARTMENT OF CARDIOLOGY, NEW YORK PRESBYTERIAN HOSPITAL, COLUMBIA UNIVERSITY, NEW YORK, NEW YORK, USA; CENTER FOR INNOVATION AND OUTCOMES RESEARCH, DEPARTMENT OF SURGERY, COLUMBIA UNIVERSITY, NEW YORK, NEW YORK, USA MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED PATIENT DEATHS COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C, IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, INCLUDING DEATH. THE SERIAL NUMBERS OR OTHER IDENTIFYING INFORMATION OF THE PRODUCTS WERE NOT REPORTED AND WERE NOT ABLE TO BE DETERMINED DURING THE INVESTIGATION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21516044&pc=NPT | 2135147-2025-01094 | 2025-02-07 05:00 | Death | ABBOTT MEDICAL | 2025-03-04 05:00 | NPT | NAVITOR TRANSCATHETER AORTIC VALVE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Cardiac Tamponade; Pericardial Effusion | P190023 |  | 1 | Event Description: IT WAS REPORTED THAT A 25MM NAVITOR TRANSCATHETER AORTIC VALVE WAS SELECTED FOR IMPLANT ON (B)(6) 2025. THE PATIENT'S NATIVE VALVE WAS MEASURED UNDER THE FOLLOWING: 22.7MM DIAMETER, 71.4MM PERIMETER, AND AN AREA OF 402.7MM^2. THE PATIENT HAD SEVERE CALCIFICATION ON ALL THREE CUSPS. A NON-ABBOTT GUIDEWIRE WAS POSITIONED IN THE LEFT VENTRICLE (LV). A PRE-DILATION WAS PERFORMED WITH AN 18MM NON-ABBOTT BALLOON. AFTER PRE-DILATION AND PRIOR TO THE DELIVERY SYSTEM ENTERING THE PATIENT, THE PATIENT PRESENTED WITH PULSELESS ELECTRICAL ACTIVITY (PEA). ADVANCED CARDIAC LIFE SUPPORT (ACLS) WAS INITIATED. CARDIOPULMONARY RESUSCITATION WAS PERFORMED AS PART OF ACLS. THE VALVE WAS POSITIONED AT 80% AND LEFT IN SITU. A TRANSTHORACIC ECHOCARDIOGRAM (TTE) WAS PERFORMED AND SHOWED A PERICARDIAL EFFUSION CLOSE TO THE LEFT VENTRICLE. A PERICARDIAL DRAINAGE WAS PLACED WITH PARTIAL REDUCTION OF A CARDIAC TAMPONADE, HOWEVER THE PATIENT REMAINED IN PEA. THE VALVE WAS RELEASED FROM THE DELIVERY CABLE. THE FINAL IMPLANT DEPTH WAS 3-4MM FROM THE NON-CORONARY CUSP (NCC). EXTRACORPOREAL CIRCULATION (EC) WAS STARTED AND A LV APICAL LESION WAS SURGICALLY TREATED. AFTER THE SURGICAL REPAIR, A RETURN OF SPONTANEOUS CIRCULATION (ROSC) WAS ACHIEVED AND MAINTAINED. AFTER ROSC EC WAS STOPPED AND REMOVED. A TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) SHOWED THE VALVE WAS IN A STABLE POSITION WITH A TRIVIAL PERIVALVULAR LEAK. PER THE PHYSICIAN, THE PERICARDIAL EFFUSION WAS DUE TO THE NON-ABBOTT GUIDEWIRE IN THE LV. THE PATIENT WAS HOSPITALIZATION IN THE INTENSIVE CARE UNIT (ICU). ON (B)(6) 2025 THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS THE PEA. PER THE PHYSICIAN, THE PEA AND SUBSEQUENT DEATH WERE RELATED TO THE PROCEDURE AND PATIENT'S COMORBIDITIES. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. AN EVENT OF CARDIAC ARREST AND DEATH WAS REPORTED. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED, AND THE PRODUCT MET ALL SPECIFICATIONS. BASED ON ALL AVAILABLE INFORMATION, CAUSES FOR THE REPORTED CARDIAC ARREST COULD NOT BE CONCLUSIVELY DETERMINED. IT IS POSSIBLE THAT THEY ARE RELATED TO THE IMPLANT PROCEDURE. HOWEVER, THIS CANNOT BE CONFIRMED. THE REPORTED DEATH APPEARS TO BE RELATED TO THE CARDIAC ARREST. THE REPORTED UNEXPECTED MEDICAL INTERVENTION WAS THE RESULT OF CASE-SPECIFIC CIRCUMSTANCES. THE EVENT AND IMAGING WAS REVIEWED BY AN ABBOTT SENIOR DIRECTOR OF MEDICAL AFFAIRS. THE REVIEWER STATED, THIS APPEARS TO BE A PROCEDURAL RELATED DEATH AND NOT RELATED TO DEVICE MALFUNCTION." CODES REMOVED: HEALTH EFFECT-CLINICAL CODE: 2226 CARDIAC TAMPONADE;3271 PERICARDIAL EFFUSION. HEALTH EFFECT- IMPACT CODE: 4624 SURGICAL INTERVENTION. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21516287&pc=NPT | 2025587-2025-01626 | 2023-12-01 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-04 05:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Insufficient Information | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 6 MONTHS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT PASSED AWAY. THE CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21516289&pc=DSQ | 2916596-2025-01438 | 2024-05-01 04:00 | Death | THORATEC CORPORATION | 2025-03-04 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE LITERATURE ARTICLE ¿PATIENT-SPECIFIC HEMODYNAMIC MODELING TO OPTIMIZE LVAD SPEED AND RIGHT HEART HEALTH¿ THAT THE HEARTMATE 3 (HM3) MAY BE ASSOCIATED WITH ADVERSE OUTCOMES, SUCH AS HOSPITAL READMISSION DUE TO LEFT VENTRICULAR ASSIST DEVICE (LVAD) RELATED ISSUES OR DEATH. THIS SINGLE-CENTER RETROSPECTIVE STUDY EVALUATED 16 PATIENTS IMPLANTED WITH A HM3 LVAD OVER THE AGE OF 18 WHO UNDERWENT RIGHT HEART CATHETERIZATION (RHC) SPEED TITRATION AS AN OUTPATIENT WERE REVIEWED RETROSPECTIVELY. MEDIAN AGE WAS 63 YEARS, 69% OF PATIENTS WERE MALE AND 75% WERE WHITE. EXACTLY 69% WERE IMPLANTED AS DESTINATION THERAPY. ALL PATIENTS WHO UNDERWENT RHC SPEED TITRATION HAD A CLINICIAN¿GUIDED SPEED INCREASE, WITH A MEDIAN INCREASE OF 300 REVOLUTIONS PER MINUTE (RPM). A VIRTUAL HEMODYNAMIC MODEL (VHM) PREDICTION WAS BUILT AND CUSTOMIZED FOR EACH PATIENT BASED ON RHC MEASUREMENTS. 7 PATIENTS REQUIRED A SPEED INCREASE WHILE THE REMAINING 9 PATIENTS NEEDED A SPEED DECREASE COMPARED TO CLINICIAN-GUIDED OPTIMIZATION, DEMONSTRATING THE PATIENT- SPECIFIC NATURE OF HEMODYNAMIC OPTIMIZATION. FOR THEIR COHORT OF 16 PATIENTS, THERE WAS AN OVERALL CONGRUENCE BETWEEN CLINICIAN¿GUIDED AND PATIENT¿SPECIFIC VHM-PREDICTED OPTIMAL LVAD SPEEDS. EXACTLY 75% OF PATIENTS WHO HAD NONCONGRUENCE (I.E., A DIFFERENCE OF OVER 200 RPM BETWEEN CLINICIAN-GUIDED AND MODEL-PREDICTED SPEEDS) HAD A HEART FAILURE OR LVAD-RELATED ADMISSION WITHIN 6 MONTHS OF THE SPEED TITRATION STUDY COMPARED TO 12.5% OF THE PATIENTS WHO HAD CONGRUENCE. EXACTLY 37.5% OF PATIENTS WHO HAD NONCONGRUENCE DIED WITHIN 1 YEAR OF THE SPEED TITRATION STUDY COMPARED TO 0 PATIENTS WHO HAD CONGRUENCY. SPECIFICALLY, 7 PATIENTS REQUIRED A SPEED INCREASE WHILE THE REMAINING 9 PATIENTS NEEDED A SPEED DECREASE COMPARED TO CLINICIAN-GUIDED OPTIMIZATION, DEMONSTRATING THE PATIENT- SPECIFIC NATURE OF HEMODYNAMIC OPTIMIZATION. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. NO PRODUCT WAS RETURNED FOR EVALUATION. THE SERIAL NUMBER OR OTHER IDENTIFYING INFORMATION OF THE PRODUCT WAS NOT REPORTED AND WAS NOT ABLE TO BE DETERMINED DURING THE INVESTIGATION. THE HEARTMATE 3 LVAS IFU, REV. C AND THE HEARTMATE 3 PATIENT HANDBOOK, REV. D ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. AUTHOR INFORMATION: AHMED, M. M., GRANT, H., MARTINEZ, J., THOMAS, J., AL-ANI, M., PARKER, A., VILARO, J., ARANDA, J., & CHIVUKULA, V. K. (2024). PATIENT-SPECIFIC HEMODYNAMIC MODELING TO OPTIMIZE LVAD SPEED AND RIGHT HEART HEALTH. JHLT OPEN, 100190. HTTPS://DOI.ORG/10.1016/J.JHLTO.2024.100190. DIVISION OF CARDIOVASCULAR MEDICINE HEART FAILURE, UNIVERSITY OF FLORIDA, GAINESVILLE, FLORIDA; BIOMEDICAL ENGINEERING AND SCIENCE, FLORIDA INSTITUTE OF TECHNOLOGY, MELBOURNE, FLORIDA NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21516512&pc=DSQ | 3007042319-2025-00601 | 2024-08-29 04:00 | Death | HEARTWARE, INC. | 2025-03-04 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P100047 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING THE ASSOCIATION BETWEEN CASELOAD VOLUME AND OUTCOMES IN LEFT VENTRICULAR ASSIST DEVICE (VAD) IMPLANTATIONS. THE AUTHORS DESCRIBED PATIENT DEATHS; HOWEVER, THE CAUSES OF DEATH WERE UNKNOWN. THERE WERE PATIENTS WHO EXPERIENCED PUMP THROMBOSIS, STROKES, RIGHT HEART FAILURE, GASTROINTESTINAL BLEEDS (GIB), DRIVELINE INFECTIONS, OTHER DEVICE RELATED INFECTIONS, NEUROLOGICAL DYSFUNCTIONS, AND OUTFLOW GRAFT OBSTRUCTION. PATIENTS WERE RE-HOSPITALIZED, IMPLANTED WITH A RIGHT VAD, AND SOME UNDERWENT PUMP EXCHANGES. THERE WERE ALSO UNSPECIFIED MAJOR DEVICE MALFUNCTIONS. THE STATUS OF THE DEVICES IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/57 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: ASSOCIATION BETWEEN CASELOAD VOLUME AND OUTCOMES IN LEFT VENTRICULAR ASSIST DEVICE IMPLANTATIONS ¿ A EUROMACS ANALYSIS. EUROPEAN JOURNAL OF HEART FAILURE. 2024. 26, 2400¿2409. DOI:10.1002/EJHF.3418. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE PRODUCT WAS DISTRIBUTED OUTSIDE OF THE UNITED STATES AND D4 UDI IS THEREFORE BLANK. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21516800&pc=NPT | 2025587-2025-01627 | 2025-01-23 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-04 05:00 | NPT | EVOLUT FX VALVE | Patient Device Interaction Problem | Endocarditis; Fever; Unspecified Infection | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT THREE DAYS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT DEVELOPED A FEVER, AND ENDOCARDITIS WAS DIAGNOSED. ADDITIONALLY, AN UNSPECIFIED MITRAL VALVE DEFECT WAS OBSERVED. TREATMENT WAS NOT REPORTED.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT 20 DAYS FOLLOWING THE VALVE IMPLANT, THE PATIENT DEVELOPED ENDOCARDITIS DUE TO METHICILLIN-RESISTANT STAPHYLOCOCCUS AUREUS (MRSA) AFTER THE VALVE IMPLANT AND RECEIVED ANTIBIOTIC TREATMENT. SUBSEQUENTLY THE PATIENT DIED. PER THE PHYSICIAN, THE CAUSE OF DEATH INFECTION AND WAS RELATED TO THE VALVE.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT CLARIFIED PREVIOUSLY REPORTED INFORMATION INDICATING THAT THE ENDOCARDITIS WAS FIRST CONFIRMED 1 DAY FOLLOWING THE IMPLANT OF THE VALVE. ADDITIONALLY, THE AORTIC VALVE DID NOT INTERACT WITH THE MITRAL VALVE. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. H6. CORRECTED DATA: B3. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA B2 B5 H1 H2 H6 ADDITIONAL CODES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21516974&pc=LDD | 3003832357-2025-000157 | 2025-02-27 05:00 | Death | SCHILLER AG | 2025-03-04 05:00 | LDD | TEMPUS LS | Unable to Obtain Readings | Cardiac Arrest | K200849 |  | 1 | Event Description: THIS REPORT IS BASED ON INFORMATION PROVIDED BY REMOTE SERVICE ENGINEER RSE AND IS BEING INVESTIGATED BY THE PHILIPS COMPLAINT HANDLING TEAM. PHILIPS RECEIVED A COMPLAINT ON THE TEMPUS LS MANUAL INDICATING SHOWED THE DEVICE DISPLAYED AN ORANGE BANNER STATING, "CHECK ECG LEADS", EVEN THOUGH THEY WERE CONNECTED TO THE DEVICE. MULTIPLE KNOWN WORKING ECG LEADS WERE ATTEMPTED, BUT NO CHANGE. THE CUSTOMER STATED THEY WERE TRANSITIONING FROM MOTORING THE PATIENT, 18-YEAR-OLD FEMALE, TO ATTEMPTING TO PACE AND WERE "UNABLE TO DO SO BECAUSE DEVICE WOULD NOT READ ECG LEAD 2 OR 3". THE PATIENT WAS IN BRADYCARDIC RHYTHM BEFORE SWITCHING TO PACING. THE PATIENT WENT INTO CARDIAC ARREST AND SUBSEQUENTLY DIED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21517227&pc=LWS | 2649622-2025-06235 | 2025-02-20 05:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-03-04 05:00 | LWS | SPRINT QUATTRO SECURE S MRI SURESCAN | Over-Sensing | Insufficient Information; No Clinical Signs, Symptoms or Conditions | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE RIGHT ATRIAL (RA) LEAD EXHIBITED UNDERSENSING TRIGGERING DEVICE CLASSIFIED FALSE TERMINATION OF ATRIAL TACH YARRHYTHMIAS/ATRIAL FIBRILLATION (AT/AF) EVENT(S) AND INAPPROPRIATELY PACING AT TIMES. IT WAS FURTHER REPORTED THAT THE RIGHT VENTRICULAR (RV) LEAD EXHIBITED OCCASIONAL OVERSENSING. THE LEADS REMAIN IN USE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED THE PATIENT IS DECEASED AND PASSED AWAY APPROXIMATELY THREE MONTHS LATER. THE CAUSE OF DEATH WAS REQUESTED AND NOT RECEIVED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21517235&pc=LWR | 2025587-2025-01628 | 2024-07-24 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-04 05:00 | LWR | AVALUS AORTIC TISSUE VALVE | Insufficient Information | Insufficient Information | P170006 |  | 1 | Event Description: MEDTRONIC RECEIVED INFORMATION THAT 9 YEARS AND 4 MONTHS POST IMPLANT OF THIS AVALUS AORTIC VALVE, THE PATIENT DIED. THE CAUSE OF DEATH WAS NOT RECEIVED. THEREFORE, RELATEDNESS OF THE DEATH TO THE AVALUS AORTIC VALVE COULD NOT BE EXCLUDED. IT IS UNKNOWN WHETHER AN AUTOPSY WAS PERFORMED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21517236&pc=DTB | 2649622-2025-06237 | 2025-02-20 05:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-03-04 05:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Under-Sensing | Insufficient Information; No Clinical Signs, Symptoms or Conditions | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE RIGHT ATRIAL (RA) LEAD EXHIBITED UNDERSENSING TRIGGERING DEVICE CLASSIFIED FALSE TERMINATION OF ATRIAL TACH YARRHYTHMIAS/ATRIAL FIBRILLATION (AT/AF) EVENT(S) AND INAPPROPRIATELY PACING AT TIMES. IT WAS FURTHER REPORTED THAT THE RIGHT VENTRICULAR (RV) LEAD EXHIBITED OCCASIONAL OVERSENSING. THE LEADS REMAIN IN USE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED THE PATIENT IS DECEASED AND PASSED AWAY APPROXIMATELY THREE MONTHS LATER. THE CAUSE OF DEATH WAS REQUESTED AND NOT RECEIVED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21517319&pc=NPT | 2025587-2025-01631 | 2025-02-19 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-04 05:00 | NPT | EVOLUT FX DCS | Failure to Advance; Adverse Event Without Identified Device or Use Problem; Patient Device Interaction Problem | Hematoma; Hemorrhage/Bleeding; Hypovolemia; Obstruction/Occlusion; No Clinical Signs, Symptoms or Conditions | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE OF THIS TRANSCATHETER BIOPROSTHETIC VALVE INTO A NON-MEDTRONIC VALVE, NO PRE-BALLOON DILATION WAS PERFORMED. THE ACCESS ON THE LEFT FEMORAL WAS FOR DIAGNOSTIC PIGTAIL WHICH BLEEDING OCCURRED, CREATING A HEMATOMA. THE VALVE WAS ADVANCED VIA THE RIGHT FEMORAL AND WAS SUCCESSFULLY ADVANCED ACROSS THE AORTIC ARCH. THE VALVE WAS NOT ABLE TO PASS BEYOND THE DISTAL CROWN OF THE NON-MEDTRONIC VALVE. AFTER SEVERAL UNSUCCESSFUL ATTEMPTS OF ADVANCING THE VALVE, THE VALVE WAS RETRIEVED FROM THE PATIENT. THE PROCEDURE WAS CONVERTED TO AN APICAL APPROACH USING A NON-MEDTRONIC VALVE. AFTER THE VALVE WAS IMPLANTED, THERE WAS GOOD CORONARY FLOW CONFIRMED VIA ANGIOGRAM. THE PATIENT BECAME HEMODYNAMICALLY UNSTABLE AND RESUSCITATION EFFORTS WERE INITIATED. AN ANGIOGRAM CONFIRMED THE LEFT CORONARY WAS OCCLUDED. PER THE PHYSICIAN, THE PATIENT DID NOT HAVE ANY FURTHER VENTRICULAR ACTION CONFIRMED VIA ECHOCARDIOGRAM. THE PATIENT DIED. IT WAS REPORTED THAT THE BLEEDING FROM THE LEFT FEMORAL ACCESS CONTRIBUTED TO THE HEMODYNAMIC INSTABILITY FOLLOWING THE DEPLOYMENT OF THE NON-MEDTRONIC VALVE.Event Description: IT WAS REPORTED THAT DURING A PROCEDURE OF THIS TRANSCATHETER BIOPROSTHETIC VALVE INTO A NON-MEDTRONIC VALVE, NO PRE-BALLOON DILATION WAS PERFORMED. THE ACCESS ON THE LEFT FEMORAL WAS FOR DIAGNOSTIC PIGTAIL WHICH BLEEDING OCCURRED, CREATING A HEMATOMA. THE VALVE WAS ADVANCED VIA THE RIGHT FEMORAL AND WAS SUCCESSFULLY ADVANCED ACROSS THE AORTIC ARCH. THE VALVE WAS NOT ABLE TO PASS BEYOND THE DISTAL CROWN OF THE NON-MEDTRONIC VALVE. AFTER SEVERAL UNSUCCESSFUL ATTEMPTS OF ADVANCING THE VALVE, THE VALVE WAS RETRIEVED FROM THE PATIENT. THE PROCEDURE WAS CONVERTED TO AN APICAL APPROACH USING A NON-MEDTRONIC VALVE. AFTER THE VALVE WAS IMPLANTED, THERE WAS GOOD CORONARY FLOW CONFIRMED VIA ANGIOGRAM. THE PATIENT BECAME HEMODYNAMICALLY UNSTABLE AND RESUSCITATION EFFORTS WERE INITIATED. AN ANGIOGRAM CONFIRMED THE LEFT CORONARY WAS OCCLUDED. PER THE PHYSICIAN, THE PATIENT DID NOT HAVE ANY FURTHER VENTRICULAR ACTION CONFIRMED VIA ECHOCARDIOGRAM. THE PATIENT DIED. IT WAS REPORTED THAT THE BLEEDING FROM THE LEFT FEMORAL ACCESS CONTRIBUTED TO THE HEMODYNAMIC INSTABILITY FOLLOWING THE DEPLOYMENT OF THE NON-MEDTRONIC VALVE. ADDITIONAL INFORMATION WAS RECEIVED TO CLARIFY THE PIGTAIL CATHETER USED WAS NOT A MEDTRONIC PRODUCT. THE HEMATOMA CAUSED BY THE PIGTAIL CATHETER CONTRIBUTED TO THE UNSTABLE HEMODYNAMIC SITUATION. PER PHYSICIAN THE COMBINATION OF VOLUME LOSS CAUSED BY THE HEMATOMA AND THE CARDIOPULMONARY RESUSCITATION AFTER THE NON-MEDTRONIC VALVE IMPLANT LED TO A CORONARY OCCLUSION. THE CAUSE OF DEATH WAS A CORONARY OCCLUSION. Manufacturer Narrative: UPDATED DATA: B5. A SECOND PARAGRAPH WAS ADDED. H6. PRODUCT ANALYSIS: THE SUSPECT DEVICE WAS DISCARDED AND; THEREFORE, NOT ANALYZED. HOWEVER, PROCEDURAL IMAGES INCLUDING THREE STATIC IMAGES AND ELEVEN MEDIA FILES WERE SUBMITTED TO MEDTRONIC FOR REVIEW. IN ADDITION, IMAGES FROM THE PATIENT¿S EXECUTIVE SUMMARY WERE PROVIDED FOR ANATOMICAL REVIEW. UPON REVIEW OF THE PROCEDURAL IMAGES, EVIDENCE SHOWED THE PATIENT HAD A PREVIOUSLY IMPLANTED NON-MEDTRONIC SURGICAL VALVE WITH A PERIMETER MEASURING 70.8MM AND A PERIMETER DERIVED DIAMETER OF 22.5MM SUGGESTING A 26MM EVOLUT. DOCUMENTATION SUPPORTS THE IMPLANT OF A 29MM EVOLUT WAS ATTEMPTED. MEASUREMENTS OF THE ILIOFEMORAL ARTERIES WERE NOT PROVIDED; THEREFORE, IT WAS NOT POSSIBLE TO VALIDATE ADEQUATE VESSEL SIZE TO ACCOMMODATE THE SUSPECT DEVICE. THE RIGHT FEMORAL ACCESS WAS REPORTEDLY USED AS THE PRIMARY ACCESS AND THE LEFT FEMORAL WAS SECONDARY. FLUOROSCOPIC INSPECTION OF THE VALVE LOAD WAS PERFORMED AND CONFIRMED A GOOD LOAD. THE MEDIA FILES SHOWED THE NOSE CONE OF THE SUSPECT DEVICE INTERACTED WITH THE SUPERIOR PORTION OF THE SURGICAL VALVE AND FAILED TO CROSS. EVIDENCE SUPPORTS FURTHER ATTEMPTS WERE ABORTED AND A NON-MEDTRONIC TRANSCATHETER AORTIC WAS IMPLANTED VIA AN ALTERNATIVE APPROACH. AN ANGIOGRAM PERFORMED POST-IMPLANT OF THE VALVE SHOWED CORONARY PERFUSION; HOWEVER, THE SUBSEQUENT ANGIOGRAM CONFIRMED ABSENCE OF BLOOD FLOW TO THE LEFT CORONARY ARTERY. IT WAS REPORTED BLEEDING FROM THE LEFT FEMORAL APPROACH OCCURRED LEADING TO A HEMATOMA AND SUBSEQUENTLY THE PATIENT DIED. BASED ON THE INFORMATION PROVIDED, IT APPEARED POSSIBLE THAT SECONDARY ACCESS SITE BLEEDING AND CORONARY OCCLUSION MAY HAVE CONTRIBUTED TO THE DEATH, BUT EVIDENCE SUGGESTS THAT THE COMPLICATIONS WERE UNRELATED TO THE MEDTRONIC EVOLUT SYSTEM. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED H.6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID EVFXPLUS-29, PRODUCT SERIAL NUMBER (B)(6), IMPLANT DATE NA, EXPLANT DATE NA, PRODUCT ID L-EVOLUTFX-2329, PRODUCT LOT NUMBER 0012570961, IMPLANT DATE NA, EXPLANT DATE NA. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21517346&pc=KNS | 1220592-2025-00001 | 2024-12-17 05:00 | Death | HOBBS MEDICAL, INC. | 2025-03-04 05:00 | KNS | TOUCHSOFT COAGULATOR | Adverse Event Without Identified Device or Use Problem | Bowel Perforation | K093079 |  | 1 | Event Description: IT WAS REPORTED THAT HOSPITAL "HAD A PATIENT THAT HAD 2 DUODENAL AVMS, FOR WHICH THE TOUCHSOFT PROBE WAS USED. THE SETTINGS WERE APPROPRIATE, WITH IT BEING ON SOFT COAG. AND AT 30-40 WATTS. THE COAGULATION APPEARED TO BE APPROPRIATE. THE PATIENT WAS DISCHARGED, AND A DAY LATER WAS HAVING ABDOMINAL PAIN. THE PATIENT HAD A DUODENAL PERFORATION REQUIRING SURGERY." UPON FOLLOW-UP, THE HOSPITAL EXPLAINED, TWO DAYS AFTER THE SURGERY, THE PATIENT CALLED THE DOCTOR COMPLAINING ABOUT PAIN. THE PATIENT RETURNED TO THE HOSPITAL, AND THE DOCTOR FOUND AN INFECTION IN THE DUODENUM. THE PATIENT WAS ADMITTED TO THE HOSPITAL WHERE A CT SCAN WAS PERFORMED, AND A PERFORATION WAS FOUND. THE PATIENT WAS TAKEN TO SURGERY AND THEN WAS ADMITTED TO THE ICU, WHERE THE PATIENT REMAINED FOR SIX WEEKS UNTIL PASSING AWAY. IN ADDITION, THE HOSPITAL REPRESENTATIVE STATED THAT THE PATIENT HAD SEVERAL OTHER CO-MORBIDITIES AND DID NOT KNOW EXACTLY WHAT CAUSED THIS EVENT TO HAPPEN. THE HOSPITAL REPRESENTATIVE STATED THAT THE DOCTOR DID NOT THINK THAT THE PROBE CONTRIBUTED TO THE PATIENT'S PASSING BUT CANNOT RULE IT OUT. AVMS WERE 2 SMALL, NONBLEEDING ANGIODYSPLASIAS LOCATED IN THE SECOND PORTION OF THE DUODENUM. TOUCHSOFT PROBE SETTINGS USED WERE 30, INCREASED TO 40 DURING TREATMENT. NO ABNORMALITIES WERE NOTED DURING TREATMENT. PATIENT ALSO RECEIVED BIOPSY OF 10MM NODULE IN DUODENAL BULB. PERFORATION WAS LOCATED IN THIRD PORTION OF THE DUDENUM PER OPERATIVE REPORT. Manufacturer Narrative: MANUFACTURING RECORDS COULD NOT BE REVIEWED AS THE FACILITY WAS NOT ABLE TO PROVIDE LOT INFORMATION FOR THE DEVICE INVOLVED IN THE EVENT. PER THE DESCRIPTION OF THE PROCEDURE PROVIDED BY THE FACILITY, GENERATOR SETTINGS USED WERE WITHIN THE PARAMETERS RECOMMENDED IN THE IFU.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21517622&pc=DSQ | 2916596-2025-01208 | 2022-04-06 04:00 | Death | THORATEC CORPORATION | 2025-03-04 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PASSED AWAY DUE TO MULTI ORGAN FAILURE. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED. THE DEVICE OPERATED AS EXPECTED. AN AUTOPSY WAS NOT PERFORMED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. AN AUTOPSY WAS NOT PERFORMED, AND THE PUMP WAS NOT EXPLANTED FOR EVALUATION. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REV. C IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS, INCLUDING VARIOUS TYPES OF ORGAN FAILURES/DYSFUNCTIONS AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21517919&pc=MNS | 2518422-2025-026504 | 2024-04-11 04:00 | Death | RESPIRONICS, INC. | 2025-03-04 05:00 | MNS | RENTAL, DREAMSTATION ST30 | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K102465 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION REGARDING A RENTAL, DREAMSTATION ST30. THE PATIENT WAS PASSED AWAY AND THAT WAS THE REASON FOR RETURNING THE DEVICE. THERE WAS NO ALLEGATION THAT THE PATIENT'S DEATH WAS RELATED TO THE DEVICE. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE DEVICE WAS RETURNED TO MANUFACTURER INVESTIGATION LABORATORY. THE DEVICE PASSED ALL THE TEST. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21517987&pc=NVN | 1028232-2025-01136 |  | Death | BIOTRONIK SE & CO. KG | 2025-03-04 05:00 | NVN | SOLIA S 45 | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Cardiac Perforation; Vascular Dissection | P950037 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IMPLANTED WITH A PACEMAKER AND FOUR LEADS (TWO CAPPED/ABANDONED AND TWO ACTIVE,) UNDERWENT A LEAD EXTRACTION PROCEDURE DUE TO VEGETATION ON THE LEADS. UPON REMOVAL OF THE ABANDONED RIGHT ATRIAL (RA) COMPETITOR LEAD, MODEL 1018T, THE PATIENTS BLOOD PRESSURE DROPPED, AND AN EMERGENCY THORACOTOMY WAS PERFORMED. THE PATIENT HAD A TEAR IN THE RIGHT ATRIUM AND SUPERIOR VENA CAVA (SVC). THE PATIENT DEVELOPED DISSEMINATED INTRAVASCULAR COAGULATION (DIC). THE DAMAGED AREAS OF THE HEART COULD NOT BE REPAIRED, AND THE PATIENT DIED ON THE OPERATING TABLE. THIS REPORT REFLECTS INFORMATION RECEIVED BY FDA IN THE FORM OF A NOTIFICATION PER 803.22 (B)(2). NO DEVICE DETAILS ARE AVAILABLE. RECEIVED VOLUNTARY ANONYMOUS MEDWATCH REPORT, MW5166913. Manufacturer Narrative: COMBINATION PRODUCT: YES. AN INFECTION WAS OBSERVED FOLLOWING THE IMPLANTATION OF THIS BIOTRONIK DEVICE. THE STERILIZATION PROCESS WAS INVESTIGATED. THE VALIDATED PROCESS ASSURES THAT ALL STERILIZATION PARAMETERS, SUCH AS GAS CONCENTRATION, TEMPERATURE, HUMIDITY, ETC., ARE WITHIN ITS SPECIFIED RANGES FOR EACH DISTRIBUTED DEVICE. ADDITIONALLY, AN ANALYSIS OF VALIDATED MICROBIOLOGICAL INDICATORS IS PERFORMED AFTER EVERY STERILIZATION PROCEDURE AS EVIDENCE OF SUCCESSFUL COMPLETION OF THE STERILIZATION PROCESS. REVIEW OF THE BIOTRONIK COMPLAINT DATABASE DID NOT REVEAL ANY CHANGES REGARDING THE TREND FOR THIS TYPE OF INCIDENT. IN SUMMARY, THE INFECTION WAS NOT DEVICE RELATED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21517992&pc=NVN | 1028232-2025-01137 |  | Death | BIOTRONIK SE & CO. KG | 2025-03-04 05:00 | NVN | SOLIA S 53 | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Cardiac Perforation; Vascular Dissection | P950037 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IMPLANTED WITH A PACEMAKER AND FOUR LEADS (TWO CAPPED/ABANDONED AND TWO ACTIVE,) UNDERWENT A LEAD EXTRACTION PROCEDURE DUE TO VEGETATION ON THE LEADS. UPON REMOVAL OF THE ABANDONED RIGHT ATRIAL (RA) COMPETITOR LEAD, MODEL 1018T, THE PATIENTS BLOOD PRESSURE DROPPED, AND AN EMERGENCY THORACOTOMY WAS PERFORMED. THE PATIENT HAD A TEAR IN THE RIGHT ATRIUM AND SUPERIOR VENA CAVA (SVC). THE PATIENT DEVELOPED DISSEMINATED INTRAVASCULAR COAGULATION (DIC). THE DAMAGED AREAS OF THE HEART COULD NOT BE REPAIRED, AND THE PATIENT DIED ON THE OPERATING TABLE. THIS REPORT REFLECTS INFORMATION RECEIVED BY FDA IN THE FORM OF A NOTIFICATION PER 803.22 (B)(2). NO DEVICE DETAILS ARE AVAILABLE. RECEIVED VOLUNTARY ANONYMOUS MEDWATCH REPORT, MW5166912. Manufacturer Narrative: COMBINATION PRODUCT: YES. AN INFECTION WAS OBSERVED FOLLOWING THE IMPLANTATION OF THIS BIOTRONIK DEVICE. THE STERILIZATION PROCESS WAS INVESTIGATED. THE VALIDATED PROCESS ASSURES THAT ALL STERILIZATION PARAMETERS, SUCH AS GAS CONCENTRATION, TEMPERATURE, HUMIDITY, ETC., ARE WITHIN ITS SPECIFIED RANGES FOR EACH DISTRIBUTED DEVICE. ADDITIONALLY, AN ANALYSIS OF VALIDATED MICROBIOLOGICAL INDICATORS IS PERFORMED AFTER EVERY STERILIZATION PROCEDURE AS EVIDENCE OF SUCCESSFUL COMPLETION OF THE STERILIZATION PROCESS. REVIEW OF THE BIOTRONIK COMPLAINT DATABASE DID NOT REVEAL ANY CHANGES REGARDING THE TREND FOR THIS TYPE OF INCIDENT. IN SUMMARY, THE INFECTION WAS NOT DEVICE RELATED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21518017&pc=DSQ | 2916596-2025-01206 | 2025-02-07 05:00 | Death | THORATEC CORPORATION | 2025-03-04 05:00 | DSQ | HEARTMATE II LEFT VENTRICULAR ASSIST DEVICE | Adverse Event Without Identified Device or Use Problem | Sepsis; Tachycardia | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO SEPSIS AND VENTRICULAR TACHYCARDIA. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: CORRECTED DATA: SECTION D4: MODEL NUMBER AND PRIMARY UNIQUE DEVICE IDENTIFICATION (UDI) NUMBER CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT'S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. FURTHERMORE, A SPECIFIC CAUSE FOR THE REPORTED SEPSIS COULD NOT BE DETERMINED. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE II LVAS IFU, REV. C AND THE HEARTMATE II LVAS PATIENT HANDBOOK, REV. C ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU LISTS, SEPSIS, CARDIAC ARRYTHMIA, AND DEATH AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM. ALTHOUGH ONLY SEPSIS WAS REPORTED BY THE ACCOUNT, SEVERAL SECTIONS OF THE HEARTMATE II LVAS IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION, AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, LISTS ARRHYTHMIA AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21518564&pc=QMG | 3015053838-2025-00045 | 2025-02-05 05:00 | Death | SHOCKWAVE MEDICAL, INC. | 2025-03-04 05:00 | QMG | SHOCKWAVE C2+ CORONARY INTRAVASCULAR LITHOTRIPSY (IVL) CATHETER | Material Split, Cut or Torn | Chest Pain | P200039/S008 |  | 1 | Event Description: A SHOCKWAVE C2+ CORONARY INTRAVASCULAR LITHOTRIPSY (IVL) CATHETER WAS USED TO TREAT A DENOVO LESION WITH NODULAR CALCIUM IN THE LEFT ANTERIOR DESCENDING ARTERY (LAD). AFTER 40 PULSES WERE DELIVERED, THE BALLOON LOST PRESSURE. POST IVL, A CUTTING/SCORING BALLOON WAS USED FOLLOWED BY THE PLACEMENT OF A DRUG ELUTING STENT (DES). THE PROCEDURE WAS FINISHED SUCCESSFULLY, AND THE DRUG ELUTING STENTS WERE IMPLANTED AS INTENDED. THE PATIENT, WHO HAD TRIPLE VESSEL DISEASE AND DIABETES, PASSED AWAY POST-PROCEDURE. IT WAS REPORTED THAT THE PATIENT MAY HAVE LOST SEPTAL FLOW, WHICH WAS SUPPORTING THE RIGHT CORONARY ARTERY (RCA). THE PHYSICIAN PERFORMED A PERCUTANEOUS CORONARY INTERVENTION (PCI) BEFORE THE PATIENT EXPIRED. THE PATIENT DID NOT PASS AWAY IMMEDIATELY AFTER THE IVL TREATMENT AND SHOWED ONLY SLIGHT CHEST PAIN WHEN THE SEPTAL FLOW DECREASED. THE PATIENT EXPIRED DURING RECOVERY. THE ACCESS SITE WAS FEMORAL AND THE INTERVAL BETWEEN THE IVL TREATMENT AND THE PATIENT'S DEATH IS UNKNOWN. THE CAUSE OF THE DEATH IS UNKNOWN. IT WAS REPORTED THAT THE PATIENT HAD MULTIPLE CO-MORBIDITIES AND THE DEATH WAS NOT RELATED TO THE PCI PROCEDURE. Manufacturer Narrative: THE DEVICE REFERENCED IN THIS REPORT WAS NOT RETURNED TO SHOCKWAVE MEDICAL, INC. THEREFORE, A PHYSICAL EXAMINATION CANNOT BE PERFORMED. THE CAUSE OF THE REDUCED BLOOD FLOW, CHEST PAIN, AND SUBSEQUENT DEATH COULD NOT BE DEFINITIVELY DETERMINED BASED ON THE INFORMATION AVAILABLE. THE INTERVAL BETWEEN THE IVL TREATMENT AND THE PATIENT'S DEATH IS UNKNOWN. THE PROCEDURE WAS FINISHED SUCCESSFULLY, AND THE DRUG ELUTING STENTS WERE IMPLANTED AS INTENDED. THE CAUSE OF THE DEATH IS UNKNOWN. IT WAS REPORTED THAT THE PATIENT HAD MULTIPLE CO-MORBIDITIES AND THE DEATH WAS NOT RELATED TO THE PCI PROCEDURE. SHOCKWAVE MEDICAL HAS CONTROLS IN PLACE TO ENSURE DEVICES ARE BUILT TO APPROVED PROCEDURES AND MEET LOT RELEASE ACCEPTANCE CRITERIA PRIOR TO BEING DISTRIBUTED. A REVIEW OF THE MANUFACTURING AND TEST DOCUMENTATION FOR SUBJECT LOT DOES NOT REVEAL ANY ISSUES WITH THE MANUFACTURING OF THE DEVICE. THE DEVICE PASSED ALL OF SHOCKWAVE MEDICAL, INC. ACCEPTANCE CRITERIA PRIOR TO SHIPPING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21518635&pc=OAE | 2029046-2025-00663 | 2025-02-06 05:00 | Death | BIOSENSE WEBSTER INC | 2025-03-04 05:00 | OAE | QDOT MICRO¿ CATHETER | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Cardiac Tamponade | P210027 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN IDIOPATHIC VENTRICULAR TACHYCARDIA ABLATION WITH A QDOT MICRO¿ CATHETER AND THE PATIENT EXPERIENCED A PERICARDIAL EFFUSION THAT REQUIRED PERICARDIOCENTESIS AND DRAIN PLACEMENT. A PERICARDIAL EFFUSION WAS NOTICED AT THE END OF THE CASE. WHILE CHECKING THE INTRACARDIAC ECHOCARDIOGRAM POST-CASE WHILE WITHDRAWING SHEATHS AND CATHETERS FROM THE HEART, A PERICARDIAL EFFUSION SURROUNDING THE VENTRICLE WAS NOTICED AND CONFIRMED. THE PATIENT'S BLOOD PRESSURE THEN DROPPED. A PERICARDIOCENTESIS AND DRAIN PLACEMENT WAS DONE. THE LAST KNOWN STATUS WAS STABLE. THE PHYSICIAN'S OPINION ON THE CAUSE OF THE ADVERSE EVENT WAS THE PROCEDURE. PATIENT REQUIRED EXTENDED HOSPITALIZATION FOR PERICARDIAL EFFUSION MANAGEMENT. HOWEVER, THE OUTCOME WAS DEATH. PROCEDURAL ACT (ACTIVATED CLOTTING TIME) WAS 350. ABLATION WAS PERFORMED PRIOR TO NOTING THE PERICARDIAL EFFUSION. NO STEAM POP. Manufacturer Narrative: ON 7-MAR-2025, THE BWI PRODUCT ANALYSIS LAB RECEIVED THE DEVICE FOR EVALUATION. THE PRODUCT INVESTIGATION WAS SUBSEQUENTLY COMPLETED. ON 14-MAR-2025, BWI RECEIVED ADDITIONAL INFORMATION REGARDING THE EVENT. THE PATIENT'S DATE OF DEATH WAS (B)(6) 2025. THE PERICARDIAL EFFUSION APPEARED TO HAVE REACCUMULATED ABRUPTLY DURING AN ECHOCARDIOGRAM. HE BECAME ACUTELY SYMPTOMATIC AND EVENTUALLY REQUIRED CARDIOPULMONARY RESUSCITATION. B2 DATE OF DEATH WAS UPDATED ACCORDINGLY. IT WAS REPORTED THAT A PATIENT UNDERWENT AN IDIOPATHIC VENTRICULAR TACHYCARDIA ABLATION WITH A QDOT MICRO¿ CATHETER AND THE PATIENT EXPERIENCED A PERICARDIAL EFFUSION THAT REQUIRED PERICARDIOCENTESIS AND DRAIN PLACEMENT. A PERICARDIAL EFFUSION WAS NOTICED AT THE END OF THE CASE. WHILE CHECKING THE INTRACARDIAC ECHOCARDIOGRAM POST-CASE WHILE WITHDRAWING SHEATHS AND CATHETERS FROM THE HEART, A PERICARDIAL EFFUSION SURROUNDING THE VENTRICLE WAS NOTICED AND CONFIRMED. THE PATIENT'S BLOOD PRESSURE THEN DROPPED. A PERICARDIOCENTESIS AND DRAIN PLACEMENT WAS DONE. THE LAST KNOWN STATUS WAS STABLE. DEVICE EVALUATION DETAILS: THE PRODUCT WAS RETURNED TO JOHNSON & JOHNSON MEDTECH (J&J MEDTECH) FOR EVALUATION. A VISUAL INSPECTION AND REVISION OF ALL FEATURES WERE PERFORMED FOLLOWING J&J MEDTECH PROCEDURES. VISUAL ANALYSIS REVEALED NO DAMAGE OR ANOMALIES ON THE DEVICE. THE DEVICE FEATURES WERE REVIEWED AND THE DEVICE WAS VISUALIZED AND RECOGNIZED CORRECTLY HOWEVER, NO TEMPERATURE WAS DISPLAYED ON THE GENERATOR, CAUSING THAT THE ABLATION COULD NOT BE PERFORMED, DUE TO AN OPEN CIRCUIT ON THE CONNECTOR AREA. NO OTHER ERRORS OR ISSUES WERE OBSERVED. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE NUMBER LOT 31518061L AND NO INTERNAL ACTIONS RELATED TO THE COMPLAINT WAS FOUND DURING THE REVIEW. THE TEMPERATURE ISSUE OBSERVED WAS NOT RELATED TO THE ADVERSE EVENT REPORTED. THE POTENTIAL CAUSE OF THE OPEN CIRCUIT COULD NOT BE CONCLUSIVELY DETERMINED. THIS ISSUES WERE NOR RELATED TO THE ADVERSE EVENT REPORTED BY THE CUSTOMER. THE ROOT CAUSE OF THE ADVERSE EVENT REMAINS UNKNOWN. THE INSTRUCTIONS FOR USE OF THE DEVICE CONTAIN THE FOLLOWING RECOMMENDATIONS: CAREFUL CATHETER MANIPULATION MUST BE PERFORMED TO AVOID CARDIAC DAMAGE, PERFORATION, OR TAMPONADE; IF THE RF GENERATOR DOES NOT DISPLAY THE TEMPERATURE, VERIFY THAT THE APPROPRIATE CABLE IS PLUGGED INTO THE RF (RADIOFREQUENCY) GENERATOR. IF THE TEMPERATURE IS STILL NOT DISPLAYED, THERE MAY BE A MALFUNCTION IN THE TEMPERATURE SENSING SYSTEM THAT MUST BE CORRECTED PRIOR TO APPLYING RF ENERGY. AS PART OF JOHNSON & JOHNSON MEDTECH¿S QUALITY PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND RELEASED TO APPROVED SPECIFICATIONS. IF ADDITIONAL INFORMATION IS RECEIVED REGARDING THIS EVENT, A SUPPLEMENTAL 3500A REPORT WILL BE SUBMITTED TO THE FDA. MANUFACTURER'S REFERENCE NUMBER: (B)(4). Manufacturer Narrative: THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, BIOSENSE WEBSTER, INC. OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. MANUFACTURER'S REFERENCE NUMBER: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21518755&pc=BZD | 2518422-2025-026658 | 2023-08-01 04:00 | Death | RESPIRONICS, INC. | 2025-03-04 05:00 | BZD | DS2ADV AUTO CPAP | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K200480 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION REGARDING A DS2ADV AUTO CPAP. THE PATIENT WAS PASSED AWAY. THE MANUFACTURE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURE'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER RECEIVED INFORMATION REGARDING A DS2ADV AUTO CPAP. THE PATIENT WAS PASSED AWAY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. IN BOX H: EVALUATION METHOD CODE, EVALUATION RESULT CODE, COMPONENT CODE AND CONCLUSION CODE HAS BEEN UPDATED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21518773&pc=NPT | 2025587-2025-01635 | 2024-01-11 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-04 05:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Renal Failure | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 3 MONTHS AND 8 DAYS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT WAS FOUND IN A STATE OF CARDIOPULMONARY ARREST WHILE TAKING A BATH. EDEMA IN THE LOWER LIMBS WAS OBSERVED. THE FAMILY DECLINED LIFE-PROLONGING TREATMENT, AND THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS NOTED TO BE CHRONIC RENAL FAILURE. IT WAS NOTED THAT PANCYTOPENIA WAS PRESENT AND POSSIBLY RELATED. PER THE PHYSICIAN, IT IS UNKNOWN WHETHER THE DEVICE OR THE PROCEDURE CONTRIBUTED TO THE CAUSE OF DEATH. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME EVOLUTFX DELIVERY CATHETER SYSTEM (DCS); PRODUCT ID D-EVOLUTFX-2329; LOT UNKNOWN; USE BY DATE UNKNOWN; UDI UNKNOWN. UPDATED: H6, H11. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21520290&pc=DSQ | 2916596-2025-01191 | 2021-12-01 05:00 | Death | THORATEC CORPORATION | 2025-03-04 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE "CLINICAL CHARACTERISTIC OF LVAD RECIPIENTS REQUIRED OVER AND LOWER THAN MEDIAN DOSAGE OF LOOP DIURETIC DURING 12 MONTH FOLLOWING THE IMPLANTATION" THAT 30 PATIENTS DIED DURING THE STUDY. THE STUDY AIMED TO DETERMINE WHICH CLINICAL OR LABORATORY FACTORS AT BASELINE MAY PREDICT HIGHER DOSAGE OF LOOP DIURETICS (LD) DURING 12 MONTHS FOLLOWING THE IMPLANTATION. BETWEEN OCTOBER 2015 AND DECEMBER 2021 YEAR 113 PATIENTS HAVE BEEN IMPLANTED WITH LVAD IN A SINGLE CENTER IN POLAND. MEDIAN AGE OF THE STUDY POPULATIONS WAS 57.2 ± 10.1 YEARS AND THERE WERE 94.6% MEN. AFTER 12 MONTHS OF OBSERVATIONS THERE WERE 83 PATIENTS ALIVE (73.5%) AND 30 PATIENTS DIED (26.5 %). 41% OF AMBULATORY CARE PATIENTS REQUIRED OVER MEDIAN DOSAGE OF BOTH LD DURING THE FIRST 12 MONTHS FOLLOWING THE IMPLANTATION. THERE WERE NO DIFFERENCES IN TERMS OF AGE [57 (50-64) VS. 63 (52-66), P = 0.44]; GENDER AND TYPE OF LVAD (61.8 VS. 53.2% HEART MATE3, P = 44) BETWEEN GROUPS. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEMS (LVAS) AND THE REPORTED EVENTS COULD NOT CONCLUSIVELY BE ESTABLISHED THROUGH THIS EVALUATION. IT WAS REPORTED THROUGH THE RESEARCH ARTICLE "CLINICAL CHARACTERISTIC OF LVAD RECIPIENTS REQUIRED OVER AND LOWER THAN MEDIAN DOSAGE OF LOOP DIURETIC DURING 12 MONTH FOLLOWING THE IMPLANTATION" THAT 30 PATIENTS DIED DURING THE STUDY. THE STUDY AIMED TO DETERMINE WHICH CLINICAL OR LABORATORY FACTORS AT BASELINE MAY PREDICT HIGHER DOSAGE OF LOOP DIURETICS (LD) DURING 12 MONTHS FOLLOWING THE IMPLANTATION. BETWEEN (B)(6) 2015 AND (B)(6) 2021 YEAR 113 PATIENTS HAVE BEEN IMPLANTED WITH LVAD IN A SINGLE CENTER IN POLAND. MEDIAN AGE OF THE STUDY POPULATIONS WAS 57.2 ± 10.1 YEARS AND THERE WERE 94.6% MEN. AFTER 12 MONTHS OF OBSERVATIONS THERE WERE 83 PATIENTS ALIVE (73.5%) AND 30 PATIENTS DIED (26.5 %). 41% OF AMBULATORY CARE PATIENTS REQUIRED OVER MEDIAN DOSAGE OF BOTH LD DURING THE FIRST 12 MONTHS FOLLOWING THE IMPLANTATION. THERE WERE NO DIFFERENCES IN TERMS OF AGE [57 (50-64) VS. 63 (52-66), P = 0.44]; GENDER AND TYPE OF LVAD (61.8 VS. 53.2% HEART MATE3, P = 44) BETWEEN GROUPS. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REVISION D, IS CURRENTLY AVAILABLE. SECTION 1 OF THIS DOCUMENT LISTS THE ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. THE SERIAL NUMBERS OF THE HEARTMATE 3 LEFT VENTRICULAR SYSTEMS IN USE WERE NOT PROVIDED. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SECTION A, D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. SECTION B3: DATE OF EVENT HAS BEEN ENTERED AS 01DEC2021 AS PATIENTS WERE IMPLANTED BETWEEN OCTOBER 2015 AND DECEMBER 2021. SECTION D4: DEVICE SERIAL NUMBER AND LOT NUMBER WERE NOT PROVIDED, AND EXPIRATION DATE AND MANUFACTURE DATE (H4) CANNOT BE DETERMINED. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION. WISNIOWSKA-SMIALEK, S., KALETA, M., RUBIS, P., WIERZBICKI, K., GÓRKIEWICZ-KOT, I., & WASILEWSKI, G. (2024). CLINICAL CHARACTERISTIC OF LVAD RECIPIENTS REQUIRED OVER AND LOWER THAN MEDIAN DOSAGE OF LOOP DIURETIC DURING 12 MONTHS FOLLOWING THE IMPLANTATION. KARDIOLOGIA POLSKA, 82, 166-167. DOI:HTTP://DX.DOI.ORG/10.33963/V.PHJ.102611 NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21520372&pc=QLG | 2954323-2025-07903 | 2025-02-10 05:00 | Death | ABBOTT DIABETES CARE INC | 2025-03-04 05:00 | QLG | FREESTYLE LIBRE 2 | Insufficient Information | Insufficient Information | K193371 |  | 1 | Event Description: ABBOTT DIABETES CARE (ADC) RECEIVED A MEDWATCH REPORT WHICH REPORTED THE FOLLOWING INFORMATION: A HEALTHCARE PROFESSIONAL (HCP) REPORTED THAT THE CUSTOMER, WHO WAS USING A FREESTYLE LIBRE 2 SENSOR HAD "PASSED AWAY". AN OFFICIAL CAUSE OF DEATH HAS NOT BEEN PROVIDED AT THIS TIME AND NO ALLEGATION OF DEVICE DEFICIENCY HAS BEEN MADE. THE REPORT INDICATED THAT THE HCP DID NOT CONSENT TO FOLLOW UP. ADC IS THEREFORE UNABLE TO ATTEMPT ANY FOLLOW UP ACTIVITIES RELATED TO THIS EVENT. THEREFORE, BASED ON THE LIMITED INFORMATION PROVIDED, IT CANNOT BE REASONABLY CONCLUDED THAT AN ADC PRODUCT HAS OR MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT. Manufacturer Narrative: THIS ISSUE DOES NOT MEET REPORTABILITY CRITERIA; HOWEVER, IT IS BEING REPORTED TO THE FDA AS THE COMPLAINT WAS RECEIVED VIA USER REPORT. A FOLLOW-UP REPORT WILL BE SUBMITTED IF ADDITIONAL INFORMATION IS OBTAINED. THE DATE OF THE INCIDENT IS UNKNOWN. THE DATE ENTERED IN SECTION B3 IS THE DATE ABBOTT DIABETES CARE BECAME AWARE OF THE EVENT. THE DEVICE MFG DATE IS UNKNOWN. THE DATE ENTERED IN SECTION H4 IS THE DATE ABBOTT DIABETES CARE BECAME AWARE OF THE EVENT. ALL PERTINENT INFORMATION AVAILABLE TO ABBOTT DIABETES CARE HAS BEEN SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21521260&pc=NPV | 2025587-2025-01640 | 2025-01-17 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-04 05:00 | NPV | MELODY TRANSCATHETER PULMONARY VALVE | Insufficient Information; Patient Device Interaction Problem | Endocarditis | P140017 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING INCIDENCE AND OUTCOMES OF PROSTHETIC VALVE ENDOCARDITIS IN ADULTS WITH CONGENITAL HEART DISEASE. THE STUDY POPULATION INCLUDED 3,150 PATIENTS WHO WERE PREDOMINANTLY MALE WITH A MEAN AGE OF 38 YEARS OLD. OF THOSE PATIENTS, 149 WERE IMPLANTED WITH A PROSTHETIC VALVE. MULTIPLE MANUFACTURER¿S DEVICES WERE IMPLANTED IN THE STUDY POPULATION; 73 PATIENTS RECEIVED A MEDTRONIC MELODY BIOPROSTHETIC VALVE. AMONG ALL PATIENTS, 21 WITH PROSTHETIC VALVE ENDOCARDITIS DIED DURING FOLLOW-UP. THE AUTHORS STATED THAT PROSTHETIC VALVE ENDOCARDITIS WAS ASSOCIATED WITH A MORE THAN A TWO-FOLD INCREASE IN ALL-CAUSE MORTALITY AFTER ADJUSTMENT FOR DEMOGRAPHIC/ANATOMIC INDICES AND COMORBIDITIES. AMONG ALL PATIENTS, NON-DEATH CLINICAL OBSERVATIONS INCLUDED: PROSTHETIC VALVE ENDOCARDITIS (AVERAGE TIME FROM IMPLANT TO DIAGNOSIS WAS 7.58 YEARS), PERIVALVULAR ABSCESS, PERIVALVULAR FISTULA, MYCOTIC ANEURYSMS, SEPTIC EMBOLI, ATRIAL ARRHYTHMIA, MODERATE TO SEVERE AORTIC/PULMONARY/TRICUSPID/MITRAL REGURGITATION, AND SURGICAL INTERVENTION TO EXPLANT THE PROSTHETIC VALVE. NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. Manufacturer Narrative: CITATION: ANDREA R. HSU ET AL. INCIDENCE AND OUTCOMES OF PROSTHETIC VALVE ENDOCARDITIS IN ADULTS WITH CONGENITAL HEART DISEASE. AM HEART J JAN 16:S0002-8703(25)00006-7. 2025. 10.1016/J.AHJ.2025.01.005. EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT AND DATE OF DEATH. NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21501483&pc=EOQ | 2955842-2025-04126 | 2025-02-04 05:00 | Death | INTUITIVE SURGICAL, INC | 2025-03-03 05:00 | EOQ | ION | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Hemorrhage/Bleeding | K182188 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN ION ENDOLUMINAL LUNG BIOPSY PROCEDURE. THE PHYSICIAN OBTAINED BIOPSY SAMPLES FROM THE RIGHT LOWER LOBE TARGET LESION. THERE WERE NO ISSUES WITH THE ION SYSTEM. AFTER COMPLETING THE ION PORTION OF THE PROCEDURE, THE ION SCOPE WAS REMOVED AND MANUAL BRONCHOSCOPY WAS PERFORMED. BLEEDING WAS NOTED AT THE BIOPSY SITE WHILE OBTAINING A BRONCHOALVEOLAR LAVAGE SAMPLE. THE ESTIMATED BLOOD LOSS WAS 60ML. THE BLOOD ENTERED THE ENDOTRACHEAL TUBE (ETT) AND THE PATIENT SUFFERED A CARDIAC ARREST, PROMPTING ADVANCED CARDIAC LIFE SUPPORT INTERVENTIONS. THE PATIENT WAS REVIVED AND HEMOSTASIS WAS ACHIEVED. THE ETT WAS REPLACED WITH A DOUBLE LUMEN ETT AND INTRAVENOUS TRANEXAMIC ACID AND AN EPINEPHRINE DRIP WERE ADMINISTERED. THE PATIENT EXPERIENCED A SECOND CARDIAC ARREST 2 DAYS POST-PROCEDURE AND EXPIRED. THE PHYSICIAN REPORTED THAT THE BLEEDING AND PATIENT OUTCOME WERE DUE TO SIGNIFICANT RISKS ASSOCIATED WITH THE BIOPSY PROCEDURE AND COMPOUNDED BY THE PATIENT'S SEVERE PREEXISTING COMORBID CONDITIONS, INCLUDING HEART FAILURE WITH REDUCED EJECTION FRACTION, PULMONARY HYPERTENSION, AND SEVERE COPD. Manufacturer Narrative: REVIEW OF THE ION SYSTEM LOGS FOUND NO RELATED ERRORS OCCURRED DURING THE PROCEDURE THAT WERE RELEVANT TO THE REPORTED EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT UNDERWENT AN ION LUNG BIOPSY; THE ION PORTION OF THE CASE WAS COMPLETED, THE ION SCOPE WAS REMOVED AND MANUAL BRONCHOSCOPY WAS PERFORMED. BLEEDING WAS NOTED WHILE OBTAINING A BRONCHOALVEOLAR LAVAGE SAMPLE. THIS WAS COMPLICATED BY A CARDIAC ARREST. RETURN OF SPONTANEOUS CIRCULATION (ROSC) WAS ACHIEVED WITH RESUSCITATIVE EFFORTS AND THE PATIENT WAS HOSPITALIZED IN THE CCU. THE PATIENT SUFFERED ANOTHER ARREST 2 DAYS AFTER THE PROCEDURE AND DIED. THERE IS NO ALLEGATION THAT A MALFUNCTION OF THE ION SYSTEM, INSTRUMENT, OR ACCESSORY OCCURRED. GIVEN THE AVAILABLE DATA THE REPORTED EVENT WAS LIKELY PROCEDURE-RELATED AND NOT DEVICE-RELATED. BRONCHOSCOPY IS A MINIMALLY INVASIVE PROCEDURE WITH A LOW-RISK PROFILE FOR SIGNIFICANT BLEEDING. A RETROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES REPORTED 21 EPISODES OF HEMOPTYSIS OF > 50 ML (0.38%) AND 19 EPISODES < 50 ML (0.34%). A PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 NAVIGATIONAL BRONCHOSCOPY CASES REPORTED A BLEEDING RATE OF 2.5% OVERALL AND A CTCAE GRADE 2 OR GREATER BLEEDING RATE OF 1.5%. A SINGLE-CENTER RETROSPECTIVE REVIEW OF 19,017 BRONCHOSCOPIC BIOPSIES REPORTED A SEVERE BLEEDING RATE OF 0.79% WITH A HIGHER RATE IN MORE CENTRAL LESIONS. ANOTHER LARGE CASE SERIES OF 26,895 BRONCHOSCOPIES WITH BIOPSIES OF ANY TYPE REPORTED ANY BLEEDING IN 41.4% OF CASES BUT A SEVERE BLEEDING RATE OF 1.47%. A RECENT META-ANALYSIS OF NAVIGATIONAL BRONCHOSCOPY IN 10,381 PATIENTS REPORTED AN OVERALL ADVERSE EVENT RATE OF 5.6% WITH 1 DEATH. BLEEDING OF ANY SEVERITY WAS REPORTED IN 2.1% OF ALL CASES. THESE DATA REFLECT VARIOUS DEFINITIONS OF ANY BLEEDING IN THE LITERATURE RANGING FROM 0.72-41.4% WITH POSSIBLY MORE CONSISTENT DEFINITIONS OF RELATIVELY SEVERE CLINICALLY SIGNIFICANT BLEEDING RANGING FROM 0.38-1.47%. THE RATE OF DEATHS ASSOCIATED WITH BLEEDING WOULD BE SOME FRACTION OF THE REPORTED BLEEDING RATES AND WILL THUS BE MORE RARE THAN SIGNIFICANT BLEEDING. A RETROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES REPORTED 4 TOTAL DEATHS (0.02%). ANOTHER PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 REPORTED 1 DEATH (0.08%). A RECENT META-ANALYSIS OF NAVIGATIONAL BRONCHOSCOPY IN 10,381 PATIENTS REPORTED AN OVERALL ADVERSE EVENT RATE OF 5.6% WITH 1 DEATH (0.01%). A CASE SERIES OF ION ROBOTIC ASSISTED BRONCHOSCOPIES PUBLISHED AFTER THE META-ANALYSIS INCLUDING 415 CASES REPORTED NO DEATHS. --FOLCH EE, PRITCHETT MA, NEAD MA, ET AL. ELECTROMAGNETIC NAVIGATION BRONCHOSCOPY FOR PERIPHERAL PULMONARY LESIONS: ONE-YEAR RESULTS OF THE PROSPECTIVE, MULTICENTER NAVIGATE STUDY. JOURNAL OF THORACIC ONCOLOGY. 2019. --FACCIOLONGO N, PATELLI M, GASPARINI S, ET AL. INCIDENCE OF COMPLICATIONS IN BRONCHOSCOPY. MULTICENTRE PROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES. MONALDI ARCHIVES FOR CHEST DISEASE. 2009. --KOPS SEP, HEUS P, KOREVAAR DA, ET AL. DIAGNOSTIC YIELD AND SAFETY OF NAVIGATION BRONCHOSCOPY: A SYSTEMATIC REVIEW AND META-ANALYSIS. LUNG CANCER. 2023. --BO L, SHI L, JIN F, LI C. THE HEMORRHAGE RISK OF PATIENTS UNDERGOING BRONCHOSCOPIC EXAMINATIONS OR TREATMENTS. AM J TRANSL RES. 2021. --BROWNLEE AR, WATSON JJJ, AKHMEROV A, ET AL. ROBOTIC NAVIGATIONAL BRONCHOSCOPY IN A THORACIC SURGICAL PRACTICE: LEVERAGING TECHNOLOGY IN THE MANAGEMENT OF PULMONARY NODULES. JTCVS. 2023.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21501960&pc=CBK | 3001421318-2025-00360 | 2025-02-25 05:00 | Death | HAMILTON MEDICAL AG | 2025-03-03 05:00 | CBK | HAMILTON G5 | Insufficient Information | Insufficient Information | K193228 |  | 1 | Event Description: HAMILTON MEDICAL AG RECEIVED THE FOLLOWING EVENT DESCRIPTION: QUOTATION FROM THE REPORTER: " HAMILTON G5 SN (B)(6) WAS ON A PATIENT WHO EXPIRED, UNKNOWN OF ANY PROBLEMS OR ISSUES THAT LED TO THIS. " "TECH DID INSPECT MACHINE AFTER INCIDENT AND MENTIONED THAT EXP. VALVE MAY BE STICKY. " ACCORDING TO THE INFORMATION RECEIVED IT IS UNKNOWN IF THE PATIENT WAS HARMED BY MACHINE. A MEDICAL INTERVENTION WAS NECESSARY, BUT THE SPECIFIC TYPE OF INTERVENTION REQUIRED WAS NOT SPECIFIED. MOREOVER, THE MANUFACTURER WAS INFORMED THAT THE USER REPORTED THIS EVENT TO LOCAL AUTHORITIES, ALTHOUGH NO REFERENCE NUMBER WAS PROVIDED. FURTHER INQUIRIES HAVE BEEN SENT TO THE CUSTOMER TO OBTAIN ADDITIONAL DETAILS ABOUT THE EVENT. THE CAUSALITY OF THE ADVERSE EVENT ASSOCIATED WITH THE HAMILTON-G5 HAS NOT YET BEEN DETERMINED AND IS CURRENTLY UNDER INVESTIGATION Manufacturer Narrative: HAMILTON MEDICAL AG COMPLAINT NUMBER: (B)(4). INVESTIGATION ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21502494&pc=NVN | 2649622-2025-05939 | 2025-02-02 05:00 | Death | MPRI | 2025-03-03 05:00 | NVN | SELECTSECURE MRI SURESCAN | Failure to Capture; Impedance Problem; High Capture Threshold | Cerebral Hyperperfusion Syndrome | P030036 |  | 1 | Event Description: IIT WAS REPORTED THAT THE PATIENT DECEASED. THE RIGHT VENTRICULAR (RV) LEAD EXHIBITED HIGH THRESHOLDS AND LOSS OF CAPTURE WAS NOTED. RISE IN IMPEDANCE WAS NOTED ON THE RV LEAD. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21502713&pc=NVZ | 3008973940-2025-01554 | 2025-02-11 05:00 | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-03-03 05:00 | NVZ | AZURE XT DR MRI SURESCAN | Appropriate Term/Code Not Available | Insufficient Information | P980035 |  | 1 | Event Description: IT WAS REPORTED BY THE PATIENT REPRESENTATIVE THAT THE PATIENT DECEASED AND THEIR IMPLANTABLE PULSE GENERATOR (IPG) WAS NOT WORKING. IT WAS REPORTED THAT THE MOBILE MONITORING APPLICATION DID NOT NOTIFY THAT THE DEVICE WAS NOT WORKING. IT WAS ALSO REPORTED THAT THE MOBILE MONITORING APPLICATION DID NOT SEND ANY INFORMATION. CONFIRMED THAT THE APPLICATION REMAINED CONNECTED UNTIL THE PATIENT PASSED AND THE TRANSMISSION ALERT WAS SENT (B)(6). Manufacturer Narrative: CONTINUATION OF H10: EVOLUTFX-26TRANSCATHETER VALVE IMPLANTED ON (B)(6) 2023. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21502922&pc=KRH | 9612164-2025-01113 | 2024-11-21 05:00 | Death | MEDTRONIC MEXICO | 2025-03-03 05:00 | KRH | RING ANNULOPLSTY TRICUSP CONTOUR 3D 28MM | Insufficient Information | Insufficient Information | K101212 |  | 1 | Event Description: MEDTRONIC RECEIVED INFORMATION FROM A PATIENT REPRESENTATIVE THAT 1 DAY POST IMPLANT OF THIS 28MM ANNULOPLASTY RING, THE PATIENT DIED. THE CAUSE OF DEATH WAS NOT RECEIVED. THEREFORE, RELATEDNESS OF THE DEATH TO THE RING COULD NOT BE EXCLUDED. IT IS UNKNOWN WHETHER AN AUTOPSY WAS PERFORMED. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21503280&pc=NVZ | 3008973940-2025-01565 | 2025-02-11 05:00 | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-03-03 05:00 | NVZ | ATTESTA L DR MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Asystole | P980035 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT'S IMPLANTABLE PULSE GENERATOR (IPG) WAS REPROGRAMMED TO VOO FOR FISTULA SURGERY. TWO HOURS POST MODE CHANGE THE PATIENT SUFFERED ASYSTOLE AND WAS DECEASED. CARDIOPULMONARY RESUSCITATION WAS PERFORMED AND PATIENT COULD NOT BE REVIVED.Event Description: IT WAS FURTHER REPORTED THAT THE INCREASED RATE OF OPERATION OF THE IPG PRIOR TO THE FISTULA SURGERY CAUSED A TEMPORARY ELEVATION IN CAPTURE THRESHOLDS. Manufacturer Narrative: CORRECTED H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21503302&pc=DTB | 3008973940-2025-01566 | 2025-02-11 05:00 | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-03-03 05:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Asystole | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT'S IMPLANTABLE PULSE GENERATOR (IPG) WAS REPROGRAMMED TO VOO FOR FISTULA SURGERY. TWO HOURS POST MODE CHANGE THE PATIENT SUFFERED ASYSTOLE AND WAS DECEASED. CARDIOPULMONARY RESUSCITATION WAS PERFORMED AND PATIENT COULD NOT BE REVIVED.Event Description: IT WAS FURTHER REPORTED THAT THE INCREASED RATE OF OPERATION OF THE IPG PRIOR TO THE FISTULA SURGERY CAUSED A TEMPORARY ELEVATION IN CAPTURE THRESHOLDS. Manufacturer Narrative: CORRECTION H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21503323&pc=DTB | 3008973940-2025-01568 | 2025-02-11 05:00 | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-03-03 05:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Adverse Event Without Identified Device or Use Problem; High Capture Threshold | Asystole | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT'S IMPLANTABLE PULSE GENERATOR (IPG) WAS REPROGRAMMED TO VOO FOR FISTULA SURGERY. TWO HOURS POST MODE CHANGE THE PATIENT SUFFERED ASYSTOLE AND WAS DECEASED. CARDIOPULMONARY RESUSCITATION WAS PERFORMED AND PATIENT COULD NOT BE REVIVED.Event Description: IT WAS FURTHER REPORTED THAT THE INCREASED RATE OF OPERATION OF THE IPG PRIOR TO THE FISTULA SURGERY CAUSED A TEMPORARY ELEVATION IN CAPTURE THRESHOLDS. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE LEAD WAS NOT RETURNED FOR ANALYSIS, HOWEVER, PERFORMANCE DATA COLLECTED FROM THE DEVICE WAS RECEIVED AND ANALYZED. ANALYSIS OF THE DEVICE MEMORY WAS PERFORMED AND NO ANOMALIES WERE FOUND. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21503627&pc=DSQ | 2916596-2025-01138 | 2025-02-11 05:00 | Death | THORATEC CORPORATION | 2025-03-03 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2025. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE CAUSE OF DEATH WAS REPORTED AS UNKNOWN, AND WHETHER THE PATIENT'S OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. THE DEVICE WAS NOT RETURNED FOR EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21504148&pc=FWM | MW5167119 | 2010-10-01 04:00 | Death | MENTOR CORP. / MENTOR WORLDWIDE, LLC. | 2025-03-03 05:00 | FWM | MENTOR SMOOTH RND MOD PLUS PROFILE SALINE |  | Congenital Defect/Deformity; Fatigue; Headache; Failure of Implant; Rash; Diaphoresis; Sleep Dysfunction; Cognitive Changes; Spontaneous Abortion |  |  | 1 | Event Description: IN (B)(6) 2010, I RECEIVED BREAST IMPLANTS. VERY SHORTLY AFTER IMPLANTATION, I BEGAN TO DEVELOP BREAST IMPLANT ILLNESS SYMPTOMS THAT HAVE GREATLY INCREASED WITH TIME. SYMPTOMS AND DATES OF ONSET: CHILD WITH BIRTH DEFECT - 2011. LATE TERM MISCARRIAGE - 2013. FACIAL RASH - 2014. LATE TERM MISCARRIAGE - 2014. LATE TERM MISCARRIAGE - 2014. FATIGUE - 2015. CHILD WITH BIRTH DEFECT - 2016. MIGRAINES - 2020. COVID VACCINE IN 2021 - SYMPTOMS WORSENED. CHRONIC FATIGUE - 2021. BODY ACHES - 2021. NIGHT SWEATS - 2021. INSOMNIA - 2021. WORSENING RASH ON FACE AND CHEST - 2022. HAIR LOSS - 2022. BRAIN FOG/WORD RETRIEVAL DIFFICULTIES/COGNITIVE DYSFUNCTION - 2022. ON (B)(6) 2025, I HAD MY IMPLANTS EXPLANTED BY A PLASTIC SURGEON WHO SPECIALIZES IN EXPLANATION (DR. (B)(6)). I AM ALREADY NOTICING IMPROVEMENTS WITH SOME OF MY SYMPTOMS, BUT I HAVE LOST YEARS AND YEARS OF MY LIFE TO WHAT I KNOW HAS BEEN CAUSING ME TO SUFFER WITHOUT ANSWERS. REF REPORT: MW5167120.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21504178&pc=FWM | MW5167120 | 2010-10-01 04:00 | Death | MENTOR CORP. / MENTOR WORLDWIDE, LLC. | 2025-03-03 05:00 | FWM | MENTOR SMOOTH RND MOD PLUS PROFILE SALINE |  | Abortion; Congenital Defect/Deformity; Fatigue; Hair Loss; Headache; Pain; Rash; Diaphoresis; Sleep Dysfunction; Cognitive Changes |  |  | 1 | Event Description: IN (B)(6) 2010, I RECEIVED BREAST IMPLANTS. VERY SHORTLY AFTER IMPLANTATION, I BEGAN TO DEVELOP BREAST IMPLANT ILLNESS SYMPTOMS THAT HAVE GREATLY INCREASED WITH TIME. SYMPTOMS AND DATES OF ONSET: CHILD WITH BIRTH DEFECT - 2011. LATE TERM MISCARRIAGE - 2013. FACIAL RASH - 2014. LATE TERM MISCARRIAGE - 2014. LATE TERM MISCARRIAGE - 2014. FATIGUE - 2015. CHILD WITH BIRTH DEFECT - 2016. MIGRAINES - 2020. COVID VACCINE IN 2021 - SYMPTOMS WORSENED. CHRONIC FATIGUE - 2021. BODY ACHES - 2021. NIGHT SWEATS - 2021. INSOMNIA - 2021. WORSENING RASH ON FACE AND CHEST - 2022. HAIR LOSS - 2022. BRAIN FOG/WORD RETRIEVAL DIFFICULTIES/COGNITIVE DYSFUNCTION - 2022. ON (B)(6) 2025, I HAD MY IMPLANTS EXPLANTED BY A PLASTIC SURGEON WHO SPECIALIZES IN EXPLANATION (DR. (B)(6)). I AM ALREADY NOTICING IMPROVEMENTS WITH SOME OF MY SYMPTOMS, BUT I HAVE LOST YEARS AND YEARS OF MY LIFE TO WHAT I KNOW HAS BEEN CAUSING ME TO SUFFER WITHOUT ANSWERS. REF REPORT: MW5167119.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21504246&pc=DSQ | 2916596-2025-01261 | 2025-02-12 05:00 | Death | THORATEC CORPORATION | 2025-03-03 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. THE ACCOUNT HAS DECLINED TO PROVIDE PATIENT OUTCOME INFORMATION FOR THIS EVENT, DUE TO PATIENT PRIVACY LAWS. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND THE REPORTED PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21504289&pc=OZD | 1220648-2025-26447 | 2025-02-09 05:00 | Death | ABIOMED, INC. | 2025-03-03 05:00 | OZD | IMPELLA CP | Failure to Advance; Patient-Device Incompatibility | Insufficient Information | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT COMPLICATIONS WERE ENCOUNTERED DURING ATTEMPTED DELIVERY OF AN IMPELLA CP PUMP FOR MECHANICAL CIRCULATORY SUPPORT. THE IMPELLA IMPLANT WAS INITIATED AFTER THE PATIENT BEGAN TO RAPIDLY DECOMPENSATE FOLLOWING A PCI ATTEMPT. FOLLOWING SUCCESSFUL INSERTION, THE PUMP WAS UNABLE TO CROSS THE AORTIC VALVE. THE IMPLANT WAS ABORTED AND THE PATIENT PASSED AWAY. IT IS UNKNOWN IF THE DELIVERY ISSUE CONTRIBUTED TO THE PATIENT'S PASSING. Manufacturer Narrative: THE INVESTIGATION OF DELIVERY ISSUES HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. BASED ON THE CLINICAL DETAILS THE ROOT CAUSE OF THE DELIVERY ISSUE WAS MOST LIKELY PATIENT CONDITION RELATED AS IMPELLA CP INSERTED BUT UNABLE TO CROSS AORTIC VALVE. H10 INVESTIGATION RESULTS WERE INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26447. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH RISK PCI SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿ SECTION: WARNINGS & CAUTIONS: WARNINGS SECTION: PRE-SUPPORT EVALUATION SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER SECTION: DIRECT AORTIC INSERTION SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES ¿TO REDUCE THE RISK OF CARDIAC INJURY (INCLUDING VENTRICULAR PERFORATION), PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY IN THE SETTINGS OF ACUTE INFARCTION WITH TISSUE NECROSIS.¿ ¿TO REDUCE THE RISK OF VASCULAR INJURY, PHYSICIANS SHOULD EXERCISE CAUTION WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX PERIPHERAL VASCULAR ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: UNREPAIRED ABDOMINAL AORTIC ANEURYSM, SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM, DISSECTION OF THE ASCENDING/ TRANSVERSE/DESCENDING AORTA, CHRONIC ANATOMICAL CHANGES IN THE RELATIONSHIP OF THE AORTA/AORTIC VALVE/VENTRICULAR ALIGNMENT, SIGNIFICANT MOBILE ATHEROMATOUS DISEASE IN THE THORACIC OR ABDOMINAL AORTA OR PERIPHERAL VESSELS.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING A RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY IN THE LEFT VENTRICLE AFTER CPR WITH ECHOCARDIOGRAPHY GUIDANCE. SECTION: WARNINGS & CAUTIONS: CAUTIONS ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21504509&pc=OZD | 1220648-2025-26449 | 2025-02-09 05:00 | Death | ABIOMED, INC. | 2025-03-03 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia | P140003 |  | 1 | Event Description: THE COMPLAINANT HAD PLACED A IMPELLA CP TO SUPPORT A PATIENT ADMITTED IN WITH ST-SEGMENT ELEVATION MYOCARDIAL INFARCTION AND IN NEED OF CORONARY ARTERY BYPASS GRAFT. THE CP WAS SUPPORTING WHEN LIMB ISCHEMIA PROMPTED THE TEAM TO PLACE A FEM-FEM BYPASS. THE PUMP REMAINS ON FOR SUPPORT. Manufacturer Narrative: THE INVESTIGATION OF LIMB ISCHEMIA HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. B2 LIFE THREATENING WAS ERRONEOUSLY SELECTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26449. E4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26449 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM, SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES), ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21504650&pc=NVY | MW5167123 |  | Death | MEDTRONIC, INC. | 2025-03-03 05:00 | NVY | RIGHT ATRIAL (RA) LEAD |  | Low Blood Pressure/ Hypotension; Vascular Dissection |  |  | 1 | Event Description: IT WAS REPORTED THAT THIS PATIENT IMPLANTED WITH THIS CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) DIED WHILE UNDERGOING A LEAD EXTRACTION PROCEDURE FOR THEIR NON-BOSTON SCIENTIFIC RIGHT ATRIAL (RA) LEAD. THE LEAD WAS STARTING TO LOSE INSULATION AND NOISE ARTIFACTS WERE PRESENT ON THE RIGHT ATRIAL (RA) LEAD ELECTROGRAM. DURING THE PROCEDURE, THE PATIENT LOST BLOOD PRESSURE AND OUTPUT. THE CARDIOTHORACIC AND VASCULAR TEAM OPENED THE PATIENT'S CHEST AND REPAIRED A TEAR FROM THE LASER SHEATH BUT THE PATIENT DID NOT SURVIVE THE REST OF THE PROCEDURE. THE DEVICE WAS NOT PLANNED TO BE RETURNED. THIS REPORT REFLECTS INFORMATION RECEIVED BY FDA IN THE FORM OF A NOTIFICATION PER 803.22 (B)(2).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21504660&pc=DTB | MW5167124 |  | Death | ABBOTT MEDICAL | 2025-03-03 05:00 | DTB | RIGHT ATRIAL (RA) LEAD |  | Low Blood Pressure/ Hypotension; Cardiac Tamponade |  |  | 1 | Event Description: IT WAS REPORTED THAT THIS PATIENT IMPLANTED WITH THIS CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) DIED WHILE UNDERGOING A LEAD EXTRACTION PROCEDURE FOR THEIR NON-(B)(6) RIGHT ATRIAL (RA) LEAD. THE LEAD WAS STARTING TO LOSE INSULATION AND NOISE ARTIFACTS WERE PRESENT ON THE RIGHT ATRIAL (RA) LEAD ELECTROGRAM. DURING THE PROCEDURE, THE PATIENT LOST BLOOD PRESSURE AND OUTPUT. THE CARDIOTHORACIC AND VASCULAR TEAM OPENED THE PATIENT'S CHEST AND REPAIRED A TEAR FROM THE LASER SHEATH BUT THE PATIENT DID NOT SURVIVE THE REST OF THE PROCEDURE. THE DEVICE WAS NOT PLANNED TO BE RETURNED. THIS REPORT REFLECTS INFORMATION RECEIVED BY FDA IN THE FORM OF A NOTIFICATION PER 803.22 (B)(2).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21504682&pc=DSQ | 2916596-2025-00983 | 2012-04-01 04:00 | Death | THORATEC CORPORATION | 2025-03-03 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE ¿EVALUATION OF MORBIDITY AND MORTALITY IN PATIENTS OVER 65 YEARS OF AGE WITH LEFT VENTRICULAR ASSIST DEVICE IMPLANTATION¿ THAT THE HEARTMATE 3 (HM3) MAY BE ASSOCIATED WITH STROKE, DEVICE THROMBOSIS, DRIVELINE INFECTION, AND DEATH. 64 PATIENTS WERE INCLUDED IN THE STUDY AND WERE GROUPED ACCORDING TO AGE, AND THEY WERE IMPLANTED WITH HEARTMATE 3, HEARTMATE II, OR HEARTWARE LEFT VENTRICULAR ASSIST DEVICES (LVADS). IN THE CONTROL GROUP, SVO (CEREBROVASCULAR EVENT) OCCURRED IN 23 OUT OF 42 PATIENTS (55%), WHILE IN THE STUDY GROUP, SVO OCCURRED IN 3 OUT OF 12 PATIENTS (25%). DEVICE THROMBOSIS WAS OBSERVED IN 12 PATIENTS (29%) IN THE CONTROL GROUP AND IN 1 PATIENT (8.3%) IN THE STUDY GROUP. DRIVELINE INFECTION DEVELOPED IN 20 PATIENTS (48%) IN THE CONTROL GROUP AND IN 4 PATIENTS (25%) IN THE STUDY GROUP. THE ONE-YEAR MORTALITY RATE WAS 14% (6 PATIENTS) IN THE CONTROL GROUP AND 8.3% (1 PATIENT) IN THE STUDY GROUP. OUTCOMES WERE NOT BROKEN DOWN BY LVAD TYPE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: SECTIONS A AND D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. DEPARTMENT OF CARDIOVASCULAR SURGERY, BASKENT UNIVERSITY, ANKARA, TURKEY; DEPARTMENT OF ANAESTHIOLOGY AND REANIMATION, BASKENT UNIVERSITY, ANKARA, TURKEY. SARP BEYAZPINAR, D., BALLA, E., OKAY KARSLIOGLU, A., AKPINAR, D., TUGBA YAMAC, E., GULTEKIN, B., HAS HASIRCI, S., KESIMCI, E., TANKUT AKAY, H., & SEZGIN, A. (2024). EVALUATION OF MORBIDITY AND MORTALITY IN PATIENTS OVER 65 YEARS OF AGE WITH LEFT VENTRICULAR ASSIST DEVICE IMPLANTATION. TRANSPLANTATION, 108, 326.6. A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEMS (LVAS) AND THE REPORTED EVENTS COULD NOT CONCLUSIVELY BE ESTABLISHED THROUGH THIS EVALUATION. THE SERIAL NUMBERS OF THE HEARTMATE 3 LEFT VENTRICULAR SYSTEMS IN USE WERE NOT PROVIDED. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21504686&pc=DRA | MW5167125 | 2024-11-06 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-03-03 05:00 | DRA | FARADRIVE SHEATH |  | Insufficient Information |  |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT FROM A FARAWAVE CASE BACK ON (B)(6) 2024, EXPERIENCED GROIN ACCESS SITE COMPLICATIONS AFTER USING THE FARADRIVE SHEATH. THE COMPLICATION CAUSED THE PATIENT TO EXPIRE THE DAY OR TWO AFTER THE PROCEDURE. THE REPORTER DID NOT HAVE A CONFIRMED DATE OF DEATH OR ANY OTHER DETAILS. THE PHYSICIAN'S OPINION ON THE CAUSE OF THE ADVERSE EVENT IS GROIN COMPLICATION FROM FARADRIVE SHEATH. THIS REPORT REFLECTS INFORMATION RECEIVED BY FDA IN THE FORM OF A NOTIFICATION PER 803.22 (B)(2).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21504832&pc=DSQ | 2916596-2025-01037 | 2025-01-27 05:00 | Death | THORATEC CORPORATION | 2025-03-03 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Pumping Stopped; Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY AT HOME DUE TO AN UNKNOWN CAUSE. LOG FILES WERE SUBMITTED FOR REVIEW. THE EVENT LOG CAPTURED A LOW POWER ADVISORY STARTING 27JAN2025 AT 19:35 DUE TO BATTERY DEPLETION. THE LOW POWER ADVISORY TURNED INTO A LOW POWER HAZARD FROM 22:26 TO 23:03. POWER CABLE DISCONNECT AND NO EXTERNAL POWER ALARMS WERE CAPTURED FROM 23:05 TO 01:17. THE BACKUP BATTERY AUTOMATICALLY ACTIVATED AT THAT TIME. AT 01:17 TO 01:18 THE LOG CAPTURED AN INTENTIONAL PUMP STOP, LOW SPEED ADVISORY AND LOW SPEED HAZARD ALARM DUE TO LOSS OF ALL POWER. THE LOSS OF ALL POWER RESULTED IN A SYSTEM CONTROLLER CLOCK RESET. IN ADDITION, THE EVENT LOG INDICATED THAT THE PATIENT DID NOT UTILIZE THE POWER MODULE/MOBILE POWER UNIT, WHICH SUGGESTED THAT THE PATIENT WAS SLEEPING ON BATTERY POWER. OTHERWISE, THERE WERE NO NOTABLE ALARM CONDITIONS OR PARAMETER CHANGES PRIOR ON (B)(6) 2025. BASED ON THE VISIBLE DATA IN THE LOG FILE THE MECHANICAL CIRCULATORY SYSTEM EQUIPMENT WAS OPERATING AS INTENDED ELECTRONICALLY AND MECHANICALLY. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. A SPECIFIC CAUSE FOR THE LOG FILE FINDING OF THE PATIENT¿S BATTERIES BEING ALLOWED TO FULLY DEPLETED RESULTING IN A PUMP STOP COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. REVIEW OF THE CONTROLLER EVENT LOG FILE EVENTS FROM 17JAN2025 THROUGH 28JAN2025. THE LOG FILE CAPTURED THE 14-VOLT BATTERIES BEING DEPLETED ON (B)(6) 2025. THE CONTROLLER BACKUP BATTERY POWERED THE SYSTEM UNTIL IT ALSO DEPLETED, AND THE PUMP STOPPED ON (B)(6) 2025. A SPECIFIC CAUSE FOR THE BATTERIES BEING ALLOWED TO BE DEPLETED COULD NOT BE CONCLUSIVELY DETERMINED. NO OTHER NOTABLE EVENTS OR ALARMS WERE CAPTURED. THE PUMP APPEARED TO FUNCTION AS INTENDED AT THE SET SPEED. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU,) REV. C AND PATIENT HANDBOOK, REV. D, ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿ OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. SECTION 2 OF THE IFU, ¿SYSTEM OPERATIONS¿ STATES THAT THE 11 VOLT LITHIUM-ION BACKUP BATTERY SHOULD BE USED ONLY FOR TEMPORARY SUPPORT DURING A POWER-LOSS EMERGENCY. THE 11 VOLT LITHIUM-ION BACKUP BATTERY INSIDE THE HEARTMATE 3 SYSTEM CONTROLLER PROVIDES ENOUGH POWER TO RUN THE IMPLANTED HEARTMATE 3 PUMP FOR AT LEAST 15 MINUTES IF THE MAIN POWER SOURCE (EITHER THE POWER MODULE, MOBILE POWER UNIT, OR TWO HEARTMATE 14 VOLT LITHIUM-ION BATTERIES) IS DISCONNECTED OR FAILS. INAPPROPRIATE USE OF THE 11 VOLT LITHIUM-ION BACKUP BATTERY MAY RESULT IN DIMINISHED RUN TIME DURING A POWER-LOSS EMERGENCY. ADDITIONALLY, EVERY HEARTMATE 14 VOLT LITHIUM-ION BATTERY ALSO HAS ITS OWN ON-BATTERY GAUGE. IT SHOWS THE POWER LEVEL FOR THAT BATTERY. 4 GREEN BARS = 75¿100% OF BATTERY POWER REMAINS. 3 GREEN BARS = 50¿75% OF BATTERY POWER REMAINS. 2 GREEN BARS = 25¿50% OF BATTERY POWER REMAINS. 1 GREEN BAR = LESS THAN 25% OF BATTERY POWER REMAINS. SECTION 3 ¿POWERING THE SYSTEM¿ AND SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿ WARN THAT THE PATIENT MUST ALWAYS CONNECT TO THE POWER MODULE OR MPU FOR SLEEPING OR WHEN THERE IS A CHANCE OF SLEEP. A SLEEPING PATIENT MAY NOT HEAR SYSTEM CONTROLLER ALARMS. THESE SECTIONS ALSO WARN NOT TO USE BATTERIES TO POWER THE SYSTEM WHEN THE PATIENT IS SLEEPING. SECTION 6 OF THE IFU, ALSO INCLUDES A SECTION ON ¿PREPARING FOR SLEEP¿, WHICH DETAILS THE STEPS PATIENTS SHOULD TAKE WHEN GOING TO SLEEP. THESE STEPS INCLUDE SWITCHING TO THE PM OR MPU AND STATES THAT IF A PATIENT FALLS ASLEEP DURING BATTERY-POWERED OPERATING, THE LOW BATTERY ALARMS MAY NOT AWAKEN THE PATIENT BEFORE BATTERY DEPLETION. ALSO, SECTION 7 OF THE IFU ¿ALARMS AND TROUBLESHOOTING¿ DETAILS SYSTEM CONTROLLER ALARMS AND HOW TO RESOLVE THEM. SECTION 3 ¿POWERING THE SYSTEM¿ OF THE PATIENT HANDBOOK DETAILS HOW TO CHECK A BATTERY'S CHARGE LEVEL, REPLACING LOW BATTERIES WITH FULLY CHARGED BATTERIES, AND SWITCHING POWER SOURCES. TWO, NEW, FULLY CHARGED LI-ION BATTERIES PROVIDE 17 HOURS OF SUPPORT. BATTERIES LAST FOR LESS TIME IF THE USER IS ACTIVE OR EMOTIONALLY STRESSED. SECTION 4 OF THE PATIENT HANDBOOK ¿LIVING WITH THE HEARTMATE 3¿ (UNDER ¿SLEEPING¿) PROVIDES INSTRUCTIONS FOR PREPARING TO SLEEP AND INSTRUCTS THE PATIENT TO ALWAYS CONNECT TO THE MPU WHEN SLEEPING (OR WHEN SLEEP IS LIKELY). ¿IF YOU FALL ASLEEP ON BATTERY POWER, YOU MIGHT NOT HEAR LOW POWER ALARMS. THE BATTERIES COULD RUN OUT OF POWER, AND THE PUMP COULD STOP BEFORE YOU HEAR THE ALARMS.¿ SECTION 5 OF THE PATIENT HANDBOOK, ¿ALARMS AND TROUBLESHOOTING¿ PROVIDES INFORMATION REGARDING SYSTEM CONTROLLER ALARMS AND THE PROPER ACTIONS ASSOCIATED WITH THEM. THIS SECTION ALSO INCLUDES DETAILED INSTRUCTIONS ON CONNECTING AND DISCONNECTING TO POWER UNDER ¿GUIDELINE FOR POWER CABLE CONNECTORS.¿ THE PATIENT HANDBOOK STATES, "CALL YOUR HOSPITAL CONTACT IF YOU THINK THAT, FOR ANY REASON, ANY PORTION OF YOUR EQUIPMENT IS NOT FUNCTIONING AS USUAL, IS BROKEN, OR YOU ARE UNCOMFORTABLE WITH THE OPERATION OF THE EQUIPMENT." THE PATIENT HANDBOOK ALSO CONTAINS A SECTION ON HANDLING EMERGENCIES. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21504848&pc=DSQ | 2916596-2025-01007 | 2021-10-01 04:00 | Death | Unknown Manufacturer | 2025-03-03 05:00 | DSQ | HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE ¿SINGLE CENTER OUTCOMES AFTER TEMPORARY MECHANICAL CIRCULATORY ASSIST DEVICE PRIOR TO HEARTMATE 3 IMPLANTATION ¿ A RETROSPECTIVE COHORT STUDY¿ THAT THE HEARTMATE 3 (HM3) MAY BE ASSOCIATED WITH DEATH, UNKNOWN TRANSPLANT, RESPIRATORY FAILURE, BLEEDING, RENAL DYSFUNCTION, RIGHT VENTRICLE (RV) FAILURE, INFECTION, PUMP THROMBOSIS, NEUROLOGICAL COMPLICATIONS, AND STROKE. THIS RETROSPECTIVE COHORT STUDY EVALUATED 63 HM3 PATIENTS IN THE PERIOD BETWEEN NOVEMBER 2015 AND OCTOBER 2021. THE AIM OF THE STUDY WAS TO COMPARE OUTCOMES AFTER HM3 IMPLANTATION IN PATIENTS WITH AND WITHOUT TEMPORARY MECHANICAL CIRCULATORY SUPPORT (TMCS) PRIOR TO HM3 IMPLANT. THERE WERE 41 PATIENTS IN THE NON-TMCS GROUP AND 22 IN THE TMCS GROUP. 11 PATIENTS DIED DURING THE PERIOD OF THE STUDY, AND 3 OF THE NON-TMCS PATIENTS AND NONE OF THE TMCS PATIENTS DIED WITHIN 30 DAYS POST-LEFT VENTRICULAR ASSIST DEVICE (LVAD) IMPLANTATION (P VALUE WAS 0.2). SURVIVAL FOUR YEARS AFTER HM3 IMPLANTATION WAS 80 AND 82% IN THE TMCS AND NON-TMCS GROUP, RESPECTIVELY. 11 PATIENTS (27%) IN THE NON-TMCS GROUP AND 8 (36%) IN THE TMCS GROUP WERE TRANSPLANTED. ONLY 1 (2%) PATIENT IN THE NON-TMCS GROUP ACHIEVED RECOVERY AND THE LVAD WAS EXPLANTED. EARLY POSTOPERATIVE OUTCOMES AND COMPLICATIONS LESS THAN 30 DAYS POST-LVAD IMPLANTATION WERE RECORDED. THE MEDIAN NUMBER OF DAYS IN THE INTENSIVE CARE UNIT (ICU) AFTER HM3 IMPLANTATION WAS 4.5 FOR THE NON-TMCS GROUP AND 12 FOR THE TMCS GROUP (P VALUE WAS .005). THE MEDIAN HOURS IN VENTILATOR WAS 23.5 FOR THE NON-TMCS GROUP AND 127 FOR THE TMCS GROUP (P VALUE WAS 0.005). THE NUMBER OF READMISSIONS WAS 6 FOR THE NON-TMCS GROUP AND 1 FOR THE TMCS GROUP (P VALUE WAS 0.22). THE MEDIAN IN HOSPITAL LENGTH OF STAY (LOS) WHEN DISCHARGED TO HOME WAS 29 FOR THE NON-TMCS GROUP AND 68.5 FOR THE TMCS GROUP (P VALUE WAS 0.00). THE MEDIAN DAYS WITH INOTROPES WAS 4 FOR THE NON-TMCS GROUP AND 5 FOR THE TMCS GROUP (P VALUE WAS 0.43). THE NUMBER OF PATIENTS WITH DELAYED CLOSURE OF STERNUM WAS 10 FOR THE NON-TMCS GROUP AND 12 FOR THE TMCS GROUP (P VALUE WAS 0.017). THE NUMBER OF PATIENTS WITH REOPERATION FOR BLEEDING WAS 12 FOR THE NON-TMCS GROUP AND 3 FOR THE TMCS GROUP (P VALUE WAS 0.17). THE MEDIAN POSTOPERATIVE BLEEDING IN ML WAS 1097.5 FOR THE NON-TMCS GROUP AND 1270 FOR THE TMCS GROUP (P VALUE WAS 0.46). THE MEDIAN RED BLOOD CELLS IN SODIUM CHLORIDE, ADENINE, GLUCOSE, AND MANNITOL MEDIUM (SAGM) IN ML WAS 735 FOR THE NON-TMCS GROUP AND 1470 FOR THE TMCS GROUP (P VALUE WAS 0.02). THE MEDIAN FRESH FROZEN PLASMA (FFP) IN ML WAS 0 FOR THE NON-TMCS GROUP AND 135 FOR THE TMCS GROUP (P VALUE WAS 0.34). THE MEDIAN PLATELETS IN ML WAS 0 FOR THE NON-TMCS GROUP AND 180 FOR THE TMCS GROUP (P VALUE WAS 0.4). THE NUMBER OF PATIENTS NEEDING DIALYSIS WAS 7 FOR THE NON-TMCS GROUP AND 13 FOR THE TMCS GROUP (P VALUE WAS 0.001). THE NUMBER OF PATIENTS WHO GOT RV FAILURE WAS 3 FOR THE NON-TMCS GROUP AND 1 FOR THE TMCS GROUP (P VALUE WAS 0.66). THE NUMBER OF PATIENTS WHO HAD TEMPORARY RIGHT VENTRICULAR ASSIST DEVICE (RVAD) IMPLANTATION WAS 1 FOR THE NON-TMCS GROUP AND 1 FOR THE TMCS GROUP (P VALUE WAS 0.65). THE NUMBER OF PATIENTS WHO GOT WOUND INFECTIONS WAS 0 FOR THE NON-TMCS GROUP AND 3 FOR THE TMCS GROUP (P VALUE WAS 0.015). OUT OF THE 3 TMCS PATIENTS WHO HAD THE INFECTIONS, 1 HAD STERNUM INFECTION, 1 HAD FEMORAL CANNULATION SITE INFECTION, AND 1 HAD SUBCLAVIAN CANNULATION SITE INFECTION. THE NUMBER OF PATIENTS WHO GOT DRIVELINE INFECTION WAS 1 FOR THE NON-TMCS GROUP AND 1 FOR THE TMCS GROUP (P VALUE WAS 0.65). THE NUMBER OF PATIENTS WHO HAD PNEUMONIA WAS 8 FOR THE NON-TMCS GROUP AND 5 FOR THE TMCS GROUP (P VALUE WAS 0.76). THE NUMBER OF PATIENTS WITH PUMP-RELATED THROMBOSIS WAS 1 FOR THE NON-TMCS GROUP AND 0 FOR THE TMCS GROUP (P VALUE WAS 0.46). THE NUMBER OF PATIENTS WITH NEUROLOGICAL COMPLICATIONS (STROKE) WAS 2 FOR THE NON-TMCS GROUP AND 3 FOR THE TMCS GROUP (P VALUE WAS 0.22). THE NUMBER OF PATIENTS WITH GASTRO-INTESTINAL BLEEDING (GIB) WAS 3 FOR THE NON-TMCS GROUP AND 2 FOR THE TMCS GROUP (P VALUE WAS 0.8). ADVERSE EVENTS AFTER LONG-TERM FOLLOW-UP WERE RECORDED IN EVENT PER PATIENTS-YEARS (EPPY). THE EPPY OF WOUND INFECTIONS WAS 0.009 IN THE NON-TMCS GROUP (41 PATIENTS; 105 PATIENT-YEARS), 0.086 IN THE TMCS GROUP (22 PATIENTS; 46 PATIENT-YEARS), AND 0.033 IN THE WHOLE POPULATION (63 PATIENTS; 150 PATIENT-YEARS). THE EPPY OF STERNUM INFECTIONS WAS 0 IN THE NON-TMCS GROUP, 0.021 IN THE TMCS GROUP, AND 0.021 IN THE WHOLE POPULATION. THE EPPY OF FEMORAL CANNULATION SITE INFECTIONS WAS 0.009 IN THE NON-TMCS GROUP, 0.043 IN THE TMCS GROUP, AND 0.02 IN THE WHOLE POPULATION. THE EPPY OF SUBCLAVIAN CANNULATION SITE INFECTIONS WAS 0 IN THE NON-TMCS GROUP, 0.021 IN THE TMCS GROUP, AND 0.021 IN THE WHOLE POPULATION. THE EPPY OF DRIVELINE INFECTIONS WAS 0.171 IN THE NON-TMCS GROUP, 0.195 IN THE TMCS GROUP, AND 0.18 IN THE WHOLE POPULATION. THE EPPY OF PNEUMONIA WAS 0.123 IN THE NON-TMCS GROUP, 0.217 IN THE TMCS GROUP, AND 0.153 IN THE WHOLE POPULATION. THE EPPY OF PUMP THROMBOSIS WAS 0.019 IN THE NON-TMCS GROUP, 0 IN THE TMCS GROUP, AND 0.013 IN THE WHOLE POPULATION. THE EPPY OF NEUROLOGICAL COMPLICATIONS WAS 0.057 IN THE NON-TMCS GROUP, 0.173 IN THE TMCS GROUP, AND 0.093 IN THE WHOLE POPULATION. THE EPPY OF STROKE WAS 0.038 IN THE NON-TMCS GROUP, 0.152 IN THE TMCS GROUP, AND 0.073 IN THE WHOLE POPULATION. THE EPPY OF HEMORRHAGE WAS 0.019 IN THE NON-TMCS GROUP, 0.021 IN THE TMCS GROUP, AND 0.02 IN THE WHOLE POPULATION. THE EPPY OF GIB WAS 0.057 IN THE NON-TMCS GROUP, 0.130 IN THE TMCS GROUP, AND 0.08 IN THE WHOLE POPULATION. THE EPPY FOR OPERATION OF ANY PUMP COMPLICATION WAS 0.028 IN THE NON-TMCS GROUP, 0.021 IN THE TMCS GROUP, AND 0.026 IN THE WHOLE POPULATION. PATIENTS ON TMCS HAD AN ACCEPTABLE SHORT AND LONG-TERM SURVIVAL AND COMPARABLE TO PATIENTS RECEIVING HM3 WITHOUT PRIOR TMCS. HOWEVER, THEY HAD A MORE COMPLICATED POSTOPERATIVE COURSE. Manufacturer Narrative: SECTION A, D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. IVERSEN, I. J., GUSTAFSSON, F., ROSSING, K., MØLLER-SØRENSEN, P. H., OLSEN, P. S., & MØLLER, C. H. (2024). SINGLE CENTER OUTCOMES AFTER TEMPORARY MECHANICAL CIRCULATORY ASSIST DEVICE PRIOR TO HEARTMATE 3 IMPLANTATION - A RETROSPECTIVE COHORT STUDY. SCANDINAVIAN CARDIOVASCULAR JOURNAL: SCJ, 58(1), 2353066. HTTPS://DOI.ORG/10.1080/14017431.2024.2353066. DEPARTMENT OF CARDIOTHORACIC SURGERY, COPENHAGEN UNIVERSITY HOSPITAL, RIGSHOSPITALET, COPENHAGEN, DENMARK; DEPARTMENT OF CARDIOLOGY, COPENHAGEN UNIVERSITY HOSPITAL, RIGSHOSPITALET, COPENHAGEN, DENMARK, DEPARTMENT OF CLINICAL MEDICINE, UNIVERSITY OF COPENHAGEN, COPENHAGEN, DENMARK; DEPARTMENT OF CARDIOTHORACIC ANAESTHESIOLOGY, COPENHAGEN UNIVERSITY HOSPITAL, RIGSHOSPITALET, COPENHAGEN, DENMARK; DEPARTMENT OF CARDIOTHORACIC SURGERY, COPENHAGEN UNIVERSITY HOSPITAL, RIGSHOSPITALET, COPENHAGEN, DENMARK, DEPARTMENT OF CLINICAL MEDICINE, UNIVERSITY OF COPENHAGEN, COPENHAGEN, DENMARK. MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEMS (LVAS) AND THE REPORTED EVENTS COULD NOT CONCLUSIVELY BE ESTABLISHED THROUGH THIS EVALUATION. THE SERIAL NUMBERS OF THE HEARTMATE 3 LEFT VENTRICULAR SYSTEMS IN USE WERE NOT PROVIDED. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REVISION D, IS CURRENTLY AVAILABLE. SECTION 1 OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21505431&pc=DQY | 2124215-2025-09219 | 2025-02-13 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-03-03 05:00 | DQY | WATCHMAN TRUSTEER ACCESS SYSTEM | Adverse Event Without Identified Device or Use Problem | Non specific EKG/ECG Changes; Low Blood Pressure/ Hypotension; Hypoxia; Brain Injury; Loss of consciousness | K240018 |  | 1 | Event Description: IT WAS REPORTED THAT ST ELEVATION, BRAIN INJURY, AND DEATH OCCURRED. A CONCOMITANT PROCEDURE WAS BEING PERFORMED WITH INTRACARDIAC ECHOCARDIOGRAM (ICE) IMAGING UNDER DEEP SEDATION WITH NO INTUBATION PERFORMED. A PULSE FIELD ABLATION (PFA) PROCEDURE WAS PERFORMED SUCCESSFULLY FIRST USING FARAPULSE PFA SYSTEM. SOMETIME BETWEEN THE CAVOTRICUSPID ISTHMUS LINE (CTI) LINE ABLATION PORTION AND THE INSERTION OF THE WATCHMAN TRUSTEER ACCESS SYSTEM (WAS) FOR THE INITIATION OF THE LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE, THE PATIENT HAD SUDDEN ST ELEVATION AND HYPOTENSION. AIR WAS SUSPECTED, SO NITRO MEDICATION WAS ADMINISTERED VIA INTRACARDIAC DELIVERY. A 31 MM WATCHMAN FLX PRO DELIVERY SYSTEM AND CLOSURE DEVICE (WDS) WAS QUICKLY INSERTED THROUGH THE WAS AND DEPLOYED SUCCESSFULLY, JUST IN CASE AIR WAS PRESENT IN THE DISTAL LAA. THE INTERVENTIONAL CARDIOLOGIST PERFORMED A DIAGNOSTIC INTERVENTION TO LOOK FOR PRESENT AIR IN THE CORONARY ARTERIES, BUT NO AIR WAS EVER VISUALIZED IN THE PATIENT. THE ST ELEVATIONS SELF-RESOLVED AND THE PATIENT REGAINED NORMAL HEMODYNAMICS. THE PROCEDURE WAS CONCLUDED. THE PATIENT WAS HOSPITALIZED YET NEVER REGAINED CONSCIOUSNESS POST INDEX PROCEDURE. MAGNETIC RESONANCE IMAGING (MRI) SCAN AND MULTIPLE COMPUTED TOMOGRAPHY (CT) SCANS WERE PERFORMED AND REVEALED NO AREA OF CONCERN, AND NO AIR WAS FOUND. HOSPICE CARE WAS INITIATED AND THE PATIENT DIED SEVERAL DAYS POST INDEX PROCEDURE ON AN UNKNOWN DATE. THE OFFICIAL CAUSE OF DEATH IS ANOXIC BRAIN INJURY DUE TO GLOBAL HYPOXIA. THE PHYSICIAN STATED THE DEATH HAD NO OFFICIAL RELATIONSHIP TO THE DEVICES USED OR THE PROCEDURE ITSELF AND IT WAS UNKNOWN WHAT CAUSED THE ANOXIC BRAIN INJURY OTHER THAN IT WAS POSSIBLE THE NON-INTUBATED PATIENT HAD TAKEN A DEEP BREATH WHILE THE WAS WAS IN THE GROIN. Manufacturer Narrative: B3: DATE OF DEATH IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21505459&pc=DSQ | 2916596-2025-01325 | 2025-01-20 05:00 | Death | THORATEC CORPORATION | 2025-03-03 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT'S OUTCOME COULD NOT CONCLUSIVELY BE ESTABLISHED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS FOR ADDITIONAL INFORMATION, INCLUDING PRODUCT STATUS, WERE SENT TO THE CUSTOMER; HOWEVER, NO FURTHER INFORMATION HAS BEEN PROVIDED AT THIS TIME. HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), WAS NOT RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21505504&pc=DSQ | 2916596-2025-01403 | 2025-02-01 05:00 | Death | THORATEC CORPORATION | 2025-03-03 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. DUE TO PATIENT PRIVACY LAWS THE MANAGING HOSPITAL WOULD NOT COMMUNICATE PATIENT OUTCOME INFORMATION. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. THE OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. NO AUTOPSY WAS PERFORMED. THE DEVICE WAS NOT EXPLANTED. THE DEVICE WOULD NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. DUE TO PATIENT PRIVACY LAWS THE CUSTOMER COMMUNICATED THAT NO ADDITIONAL INFORMATION WILL BE PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, REV. G IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21505970&pc=DSQ | 3007042319-2025-00579 | 2025-02-21 05:00 | Death | HEARTWARE, INC. | 2025-03-03 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Failure to Pump; Battery Problem; Adverse Event Without Identified Device or Use Problem; Power Problem; Pumping Problem | Loss of consciousness | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HADN'T FELT GREAT DURING THE DAY AND WENT TO BED. THE PATIENT'S SPOUSE WOKE UP TO A VENTRICULAR ASSIST DEVICE (VAD) ALARM AND THE PATIENT LAYING PRONE AND UNRESPONSIVE. THE SPOUSE CALLED 911. THE BATTERIES WERE THOUGHT TO BE FINE, BUT THERE WAS A RED ALARM ON THE CONTROLLER AND THE PATIENT WAS PRONOUNCED DEAD AT HOME. CAUSE OF DEATH IS UNKNOWN AT THIS TIME.Event Description: IT WAS FURTHER REPORTED THAT A REVIEW OF LOG FILES DATA REVEALED MOTOR START EVENTS WITH PARAMETERS OUTSIDE OF THE TYPICAL RANGE AND FAILED MOTOR START ATTEMPTS. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR ADDITIONAL INFORMATION WAS RECEIVED. ADDITIONAL PRODUCTS: D1: CONTROLLER D4: MODEL#: 1420 / CATALOG#: 1420 / EXPIRATION DATE: 31-JAN-2023 / SERIAL#: (B)(6), H4: MFG DATE: 24-JAN-2022 H6: THE CODES PRESENT IN SECTION H6 CORRESPOND TO COMPONENTS/PRODUCTS THAT COMPRISE THE REPORTED EVENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR DETAILS ABOUT ADDITIONAL PRODUCTS. ADDITIONAL PRODUCTS: D1: HEARTWARE VENTRICULAR ASSIST SYSTEM ¿ BATTERY D4: MODEL #: 1650DE / CATALOG #: 1650DE / EXPIRATION DATE: 30-DEP-2024 / SERIAL #: (B)(6) D9: NO H3: NO H4: MFG DATE: 12-SEP-2023 H5: NO D1: HEARTWARE VENTRICULAR ASSIST SYSTEM ¿ BATTERY D4: MODEL #: 1650DE / CATALOG #: 1650DE / EXPIRATION DATE: 30-JUN-2024 / SERIAL #: (B)(6) D9: NO H3: NO H4: MFG DATE: 28-JUN-2023 H5: NO D1: HEARTWARE VENTRICULAR ASSIST SYSTEM ¿ CONTROLLER D4: MODEL #: 1420 / CATALOG #: 1420 / EXPIRATION DATE: 31-JAN-2023 / SERIAL #: (B)(6) D9: NO H3: NO H4: MFG DATE: 24-JAN-2022 H5: NO MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR INVESTIGATION SUMMARY. ADDITIONAL PRODUCTS: SERIAL# (B)(6), H3: YES, SERIAL# (B)(6), H3: YES, SERIAL# (B)(6), H3: YES. H6: THE CODES PRESENT IN SECTION H6 CORRESPOND TO COMPONENTS/PRODUCTS THAT COMPRISE THE REPORTED EVENT. PRODUCT EVENT SUMMARY: THE VENTRICULAR ASSIST DEVICE (VAD) (B)(6), ONE (1) CONTROLLER (B)(6) AND TWO (2) BATTERIES (B)(6) WERE NOT RETURNED FOR EVALUATION. A REVIEW OF THE MANUFACTURING DOCUMENTATION WAS NOT PERFORMED AS THE REPORTED EVENT AND ANALYSIS ARE NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. LOG FILE ANALYSIS REVEALED THAT (B)(6) WAS THE PATIENT'S PRIMARY CONTROLLER IN USE DURING THE REPORTED EVENT. REVIEW OF THE CONTROLLER LOG FILES ASSOCIATED WITH (B)(6) REVEALED SEVERAL VAD STOPPED ALARMS, CONTROLLER FAULT ALARMS, A CRITICAL BATTERY ALARM, LOW FLOW ALARM, HIGH WATT ALARMS, AND CONTROLLER POWER UP EVENTS ON (B)(6) 2025. REVIEW OF THE EVENT LOG FILE AND MOTOR START REPORT ASSOCIATED WITH (B)(6) REVEALED TWO MOTOR START EVENTS RECORDED ON (B)(6) 2025 AT 01:22:22 AND 02:07:38 IN WHICH THE MOTOR START PARAMETERS WERE ATYPICAL. REVIEW OF THE EVENT LOG FILE REVEALED THAT DURING AN ATTEMPTED PUMP START EVENT, SEVERAL CONTROLLER RESET EVENTS WERE LOGGED ON (B)(6) 2025. DURING THE CONTROLLER RESET EVENTS, LOGS REVEALED INSTANCES THAT ONLY ONE POWER SOURCE WAS CONNECTED TO THE CONTROLLER DURING THE ATTEMPTED PUMP START EVENT. CAPA PR00609659 IS INVESTIGATING CONTROLLER RESETS DURING PUMP START. THE CONTROLLER CAN ONLY STORE A MAXIMUM OF 250 ENTRIES OF EVENT DATA. AFTER REACHING THE LIMIT, THE CONTROLLER INITIATES A FIRST-IN FIRST-OUT METHOD WHEREBY THE OLDEST EVENT POINT IS DELETED TO ALLOW THE NEWEST ENTRIES TO BE RECORDED. THEREFORE, EVENT LOGS PRIOR TO THE FIRST AVAILABLE CONTROLLER POWER-UP SEQUENCE EVENT, LOGGED AT 00:58:29, WERE NOT AVAILABLE. LOG FILE ANALYSIS REVEALED A CONTROLLER FAULT ALARM LOGGED ON (B)(6) 2025 AT 01:12:31, INDICATING A POWER OUT OF RANGE CONDITION. THIS WAS FOLLOWED BY A VAD STOPPED ALARM AT 01:12:49, INDICATING THE PUMP FAILED TO RESTART AFTER MULTIPLE ATTEMPTS. THIS WAS FOLLOWED BY ADDITIONAL VAD STOPPED AND CONTROLLER FAULT ALARMS. ONE (1) CRITICAL BATTERY ALARM, INVOLVING (B)(6), WAS LOGGED AT 01:22:22. THE CRITICAL BATTERY ALARM WAS DUE TO THE BATTERY DEPLETING BELOW 10% RELATIVE STATE OF CHARGE (RSOC). THE CONTROLLER FAULT ALARMS WERE DUE TO A POWER OUT OF RANGE CONDITION DUE TO (B)(6) APPROACHING 0% RSOC. A CONTROLLER FAULT ALARM WILL OCCUR IF THE BATTERY IS APPROACHING 0% RSOC (OR A CAPACITY BELOW 12 VOLTS). REVIEW OF THE EVENT LOG FILE REVEALED A CONTROLLER POWER UP WITH AN ASSOCIATED SUCCESSFUL MOTOR START EVENT LOGGED ON (B)(6) 2025 AT 01:22:22. FURTHERMORE, ANALYSIS OF THE ALARM LOG FILES REVEALED ONE (1) LOW FLOW AND FIVE (5) HIGH WATT ALARMS LOGGED BETWEEN 01:22:27 AND 01:55:05. REVIEW OF THE EVENT LOG FILE REVEALED A CONTROLLER POWER UP WITH ASSOCIATED MOTOR START EVENT LOGGED ON (B)(6) 2025 AT 02:06:52, INDICATING A LOSS OF POWER TO THE CONTROLLER. THEN ONE (1) VAD STOPPED ALARM WAS LOGGED AT 02:07:22, INDICATING THE PUMP FAILED TO RESTART AFTER MULTIPLE ATTEMPTS. SHORTLY AFTER THIS THIRD FAILURE TO RESTART EVENT, A SUCCESSFUL MOTOR START EVENT WAS LOGGED AT 02:07:38. THE DATA POINT RECORDED PRIOR TO THE LOSS OF POWER REVEALED THAT (B)(6) WAS CONNECTED TO POWER PORT 1 (PS1) WITH 91% RELATIVE STATE OF CHARGE (RSOC) AND (B)(6) WAS CONNECTED TO POWER PORT 2 (PS2) WITH 92% RSOC. DURING THIS PERIOD, SEVERAL SAFETY ALERT WORD (SAW) VALUES WERE RECORDED ON (B)(6), INDICATING AN OVERCURRENT ALERT. THE LOG FILES RECORDED A HIGH-POWER CONSUMPTION DURING THE ATTEMPTED MOTOR STARTS, WHICH REQUIRED MORE CURRENT FROM THE BATTERY. PER THE INSTRUCTIONS FOR USE (IFU), THE CAPACITY OF A BATTERY IS DEPENDENT ON THE CONTROLLER AND PUMP OPERATING POWER CONSUMPTION. CONSIDERING THAT THE PUMP'S POWER CONSUMPTION WAS SIGNIFICANTLY ABOVE THE NORMAL OPERATING RANGE DURING THE REPORTED EVENTS, IT IS EXPECTED THAT THE BATTERIES WOULD BE ABLE TO PROVIDE POWER TO THE CONTROLLER FOR A SHORTER PERIOD OF TIME. OF NOTE, CRITICAL BATTERY AND VAD STOPPED ALARMS ARE HIGH PRIORITY ALARMS, WHICH WILL RESULT IN THE CONTROLLER'S ALARM INDICATOR FLASHING RED. THE MOST LIKELY ROOT CAUSE OF THE VAD STOPPED ALARMS AND REPORTED ¿RED ALARMS¿ CAN BE ATTRIBUTED TO FAILURES OF THE PUMP TO RESTART AFTER SEVERAL ATTEMPTS. AS A RESULT, THE REPORTED RED ALARMS, ATYPICAL MOTOR STARTS, FAILURE TO RESTART EVENTS WERE CONFIRMED. (B)(6) WAS IN SCOPE OF FCA CVG-21-Q3-21 (NON-SUBGROUP). CAPA PR00532915 IS INVESTIGATING PUMP FAILURES TO RESTART. BASED ON AN INVESTIGATION CONDUCTED UNDER CAPA PR00532915, THE MOST LIKELY ROOT CAUSE OF THE FAILURE TO RESTART EVENTS AND ATYPICAL MOTOR START PARAMETERS MAY BE ATTRIBUTED TO OUTER SHROUD CONTACT THAT CREATED MORE FRICTION AT THE HOUSING TO IMPELLER INTERFACE. POSSIBLE ROOT CAUSES OF THE LOSS OF POWER CAN BE ATTRIBUTED TO A DISCONNECTION OF BOTH POWER SOURCES, AN INTERMITTENT DISCONNECTION ON ONE OR BOTH POWER SOURCES, AND/OR TROUBLESHOOTING THE OBSERVED CONTROLLER RESETS. CAPA PR00551638 IS INVESTIGATING CONTROLLER LOSSES OF POWER. BASED ON THE AVAILABLE INFORMATION, THE MOST LIKELY ROOT CAUSE OF THE CONTROLLER FAULT ALARMS CAN BE ATTRIBUTED TO A BATTERY CONNECTED TO THE CONTROLLER WITH A LOW RSOC VALUE DURING A PUMP START ATTEMPT WITH HIGH POWER CONSUMPTION, DRAWING MORE CURRENT FROM THE BATTERY AND RESULTING IN A POWER OUT OF RANGE CONDITION. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED HIGH-POWER AND LOW FLOW EVENT. BASED ON RISK DOCUMENTATION, MULTIPLE FACTORS MAY HAVE CONTRIBUTED TO THE REPORTED HIGH POWER EVENT INCLUDING BUT NOT LIMITED TO EXTERNAL FACTORS SUCH AS THROMBUS FORMATION/INGESTION, PATIENT RELATED FACTORS, AND/OR INAPPROPRIATE PUMP ROTATIONAL SPEED. BASED ON THE RISK DOCUMENTATION, POSSIBLE CAUSES OF THE REPORTED LOW FLOW EVENT MAY BE ATTRIBUTED TO MULTIPLE FACTORS INCLUDING BUT NOT LIMITED TO THROMBUS AT THE INFLOW CANNULA/OUTFLOW GRAFT, CONSTRICTION AT THE OUTFLOW GRAFT, POOR VAD FILLING, AND/OR INAPPROPRIATE PUMP ROTATIONAL SPEED. PER THE INSTRUCTIONS FOR USE, DEATH IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. BASED ON A REVIEW OF PAST ADVERSE EVENTS FOR THIS PATIENT, IT WAS NOTED THAT THE PATIENT HAD A HISTORY OF CVA AND BLEEDING. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. ADDITIONAL PRODUCTS: D1: HEARTWARE VENTRICULAR ASSIST SYSTEM ¿ BATTERY. D9: NO. H3: NO. H4: MFG DATE: H5: NO. THE CODES PRESENT IN SECTION H6 CORRESPOND TO COMPONENTS/PRODUCTS THAT COMPRISE THE REPORTED EVENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21505991&pc=DSQ | 2916596-2025-01250 | 2022-06-01 04:00 | Death | THORATEC CORPORATION | 2025-03-03 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE LITERATURE REVIEW ARTICLE, "LEFT VENTRICULAR ASSIST DEVICE EXCHANGE IS A GOOD ALTERNATIVE TO HEART TRANSPLANT IN THE CURRENT ORGAN ALLOCATION SYSTEM" THAT LEFT VENTRICULAR ASSIST DEVICE (LVAD) COMPLICATIONS INCLUDING THROMBOSIS, INFECTION, AND DRIVELINE FRACTURES ARE NOT INFREQUENT AND HAVE A POOR PROGNOSIS UNLESS EXPEDITIOUSLY MANAGED. DATA WAS PRESENTED FOR PATIENTS THAT RECEIVED DEVICE EXCHANGES DUE TO DEVICE COMPLICATION INSTEAD OF HEART TRANSPLANTATION. BETWEEN JANUARY 2010 AND JUNE 2022, 49 PATIENTS UNDERWENT LVAD EXCHANGE FOR DEVICE COMPLICATIONS, AMOUNTING TO 52 PATIENT EVENTS. MEAN AGE WAS 53±13 YEARS. REASONS FOR LVAD EXCHANGE WERE BIVENTRICULAR FAILURE REQUIRING UPGRADE TO TOTAL ARTIFICIAL HEART (TAH) IN 2 PATIENTS, DRIVELINE FRACTURE IN 2 PATIENTS, BLOOD-CULTURE POSITIVE INFECTIONS IN 2 PATIENTS, INTERNAL SHORT TO SHIELD IN 3 PATIENTS, TAH EXCHANGE FOR COMPRESSION IN 1 PATIENT, AND DEVICE THROMBOSIS IN THE REST, AS DIAGNOSED BY ECHOCARDIOGRAPHIC EVIDENCE OF CLOT, RISING LACTATE DEHYDROGENASE AND PLASMA-FREE HEMOGLOBIN. DEVICE EXCHANGES INCLUDED HEARTMATE 3 (HM3) TO HM3 IN 1 PATIENT, HM2 TO HM3 IN 8 PATIENTS, HM2 TO HM2 IN 28 PATIENTS, HM2 TO HEARTWARE IN 7 PATIENTS, HEARTWARE TO HEARTWARE IN 5 PATIENTS, AND UPGRADE/CHANGE TO TAH IN 3 PATIENTS. COMPLICATIONS NOTED IN THE IMMEDIATE POSTOPERATIVE PERIOD INCLUDED CENTRAL NERVOUS SYSTEM (CNS) ISCHEMIC STROKE IN 5 PATIENTS, GASTROINTESTINAL (GI) BLEED IN 4 PATIENTS, GI ISCHEMIA IN 3 PATIENTS AND LONG-TERM HEMODIALYSIS IN 5 PATIENTS. EARLY MORTALITY WAS 12% (N=6). ACTUARIAL SURVIVAL AT 5-YEARS WAS 36.3±7.5%. Manufacturer Narrative: B3: THE EVENT DATE HAS BEEN ENTERED AS (B)(6) 2022 AS PATIENTS WERE IMPLANTED BETWEEN (B)(6) 2010 AND (B)(6) 2022. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 DEVICES AND THE REPORTED OUTCOMES COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE ARTICLE REPORTED THAT EARLY MORTALITY WAS 12% (N=6). ACTUAL SURVIVAL AT 5-YEARS WAS 36.3% WITH A STANDARD DEVIATION OF 7.5%. THE ARTICLE CONCLUDED THAT LVAD EXCHANGES FOR DEVICE-RELATED COMPLICATIONS CAN BE PERFORMED SAFELY AND EFFECTIVELY AND REMAINS A GOOD ALTERNATIVE TO A HEART TRANSPLANT. THE HEARTMATE 3 DEVICE SERIAL NUMBERS, AS WELL AS OTHER SPECIFIC CASE/PATIENT INFORMATION, ARE NOT AVAILABLE. THE SERIAL NUMBER OR OTHER IDENTIFYING INFORMATION OF THE PRODUCT WAS NOT REPORTED AND WAS NOT ABLE TO BE DETERMINED DURING THE INVESTIGATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) REV. C, IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SECTION A AND D - SPECIFIC DEVICE AND PATIENT INFORMATION WAS UNKNOWN. MAHESH, B., OBLENDER, R., SCANLON, T., SHOUEY, J., & SOLEIMANI, B. (2024). LEFT VENTRICULAR ASSIST DEVICE EXCHANGE IS A GOOD ALTERNATIVE TO HEART TRANSPLANT IN THE CURRENT ORGAN ALLOCATION SYSTEM. ASAIO JOURNAL, 70, 116. DOI:HTTP://DX.DOI.ORG/10.1097/MAT.0000000000002076. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21506080&pc=BZD | 2518422-2025-102550 | 2024-12-16 05:00 | Death | RESPIRONICS, INC. | 2025-03-03 05:00 | BZD | NO EQUIPMENT NUMBER | Adverse Event Without Identified Device or Use Problem | Insufficient Information | UNKNOWN |  | 1 | Event Description: IT WAS REPORTED THAT THE USER OF THE UNSPECIFIED DEVICE PASSED AWAY. THERE IS NO ALLEGATION OF MALFUNCTION, OR THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21506099&pc=BZD | 2518422-2025-102544 | 2024-09-09 04:00 | Death | RESPIRONICS, INC. | 2025-03-03 05:00 | BZD | NO EQUIPMENT NUMBER | Adverse Event Without Identified Device or Use Problem | Insufficient Information | UNKNOWN |  | 1 | Event Description: IT WAS REPORTED THAT THE UNSPECIFIED DEVICE WAS BEING RETURNED DUE TO THE USER PASSING AWAY. THERE IS NO ALLEGATION OF MALFUNCTION, OR THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21506114&pc=BZD | 2518422-2025-102552 | 2025-01-23 05:00 | Death | RESPIRONICS, INC. | 2025-03-03 05:00 | BZD | REP DREAMSTATION AUTO BIPAP, DOM | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: IT WAS REPORTED THAT THE REP DREAMSTATION AUTO BIPAP DEVICE WAS BEING RETURNED DUE TO THE USER PASSING AWAY. THERE IS NO ALLEGATION OF MALFUNCTION, OR THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.Event Description: THIS REPORT IS BEING SUBMITTED TO PROVIDE THE DEVICE'S EVALUATION RESULTS, UPDATED CODING, AND TO INCLUDE THE PRODUCT UDI. IT WAS REPORTED THAT THE REP DREAMSTATION AUTO BIPAP DEVICE WAS BEING RETURNED DUE TO THE USER PASSING AWAY. THERE IS NO ALLEGATION OF MALFUNCTION, OR THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH. Manufacturer Narrative: THIS REPORT IS BEING SUBMITTED TO PROVIDE THE DEVICE'S EVALUATION RESULTS, UPDATED CODING, AND TO INCLUDE THE PRODUCT UDI. THE REP DREAMSTATION AUTO BIPAP DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER. THE DEVICE WAS VISUALLY INSPECTED DURING THE EVALUATION, AND THE COMPLAINT COULD NOT BE REPRODUCED. THE EVALUATION FOUND THE DEVICE WAS MISSING FILTERS. NO OTHER FAILURES WERE OBSERVED. THE DEVICE WAS SCRAPPED DUE TO THE CUSTOMER'S REQUEST. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21506175&pc=OZD | 1220648-2025-26454 | 2025-02-05 05:00 | Death | ABIOMED, INC. | 2025-03-03 05:00 | OZD | IMPELLA CP | Restricted Flow rate | No Clinical Signs, Symptoms or Conditions | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. SHORTLY AFTER IMPLANT THE FLOWS ON THE PUMP DROPPED AND THE TEAM MADE THE DECISION TO EXPLANT THE REPLACE THE PUMP. THE SECOND PUMP WAS PLACED BUT, THE PATIENT EXPIRED LATER ON THAT SECOND PUMP WHILE IN THE INTENSIVE CARE UNIT. THE TEAM NOTED THE DEATH WAS NOT DUE TO THE PUMP PER THE FIELD REP. THE SECOND PUMP SUPPORTED THE PATIENT FOR UNDER 3 HOURS BEFORE THE PATIENT EXPIRED.Event Description: MEDICAL SAFETY REVIEW OF THE EVENT NOTED THAT THE PATIENT DID NOT RECEIVE SUPPORT IN CRITICAL NEED. THERE WAS A DELAY IN TREATMENT. A NEW IMPELLA CP WAS IMPLANTED AND SUPPORTED THE PATIENT ABOUT THREE HOURS BEFORE THE PATIENTS DEMISE. THE NOTED DEVICE ISSUE MAY HAVE ADDED TO AN ALREADY COMPLEX SITUATION. THERE IS NOT ENOUGH INFORMATION TO DISSOCIATE THE DELAY IN THERAPY. Manufacturer Narrative: MEDICAL SAFETY REVIEW OF THE EVENT NOTED THAT THE PATIENT DID NOT RECEIVE SUPPORT IN CRITICAL NEED. THERE WAS A DELAY IN TREATMENT. A NEW IMPELLA CP WAS IMPLANTED AND SUPPORTED THE PATIENT ABOUT THREE HOURS BEFORE THE PATIENTS DEMISE. THE NOTED DEVICE ISSUE MAY HAVE ADDED TO AN ALREADY COMPLEX SITUATION. THERE IS NOT ENOUGH INFORMATION TO DISSOCIATE THE DELAY IN THERAPY. B.1 REVISED AS PRODUCT PROBLEM WAS INADVERTENTLY OMITTED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26454. B.2 REVISED TO ADD DEATH AND DATE OF DEATH AS THEY WERE INADVERTENTLY OMITTED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26454. B.5 ADDED MEDICAL SAFETY OFFICER REVIEW AS IT WAS INADVERTENTLY OMITTED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26454. E.1 REVISED CITY AND ADDRESS LINE 1 AS THEY WERE PREVIOUSLY ENTERED INCORRECTLY ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26454. ADDED FACILITY NAME AS IT WAS INADVERTENTLY OMITTED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26454. E.4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26454 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. F.6 AND F.8 DATES WERE REPORTED ON MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26454 AND SHOULD NOT HAVE BEEN. G.1 REVISED REPORTING CONTACT FIRST AND LAST NAME AND MANUFACTURING SITE NAME AND ADDRESS SECTION AS THEY WERE PREVIOUSLY ENTERED INCORRECTLY ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26454. G.2 ADDED HEALTH PROFESSIONAL AS IT WAS INADVERTENTLY OMITTED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26454. G.3 30-DAY SHOULD HAVE BEEN SELECTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26454 AND WAS NOT. H.1 REVISED AS IT WAS PREVIOUSLY SELECTED INCORRECTLY ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26454 H.6 ADDED DEATH CODE AS IT WAS INADVERTENTLY OMITTED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26454. H.10 ADDED MEDICAL SAFETY OFFICER REVIEW AS IT WAS INADVERTENTLY OMITTED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26454. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21506439&pc=DSQ | 2916596-2025-01313 | 2018-08-01 04:00 | Death | THORATEC CORPORATION | 2025-03-03 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Multiple Organ Failure; Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE TITLED ¿PREOPERATIVE ANATOMICAL LANDMARKS AND LONGITUDINAL HEARTMATE 3 PUMP POSITION IN X- RAYS: RELEVANCE FOR ADVERSE EVENTS¿ THAT HEARTMATE 3 MALPOSITION MAY BE ASSOCIATED WITH HEMOCOMPATIBILITY-RELATED ADVERSE EVENTS (HRAES), RIGHT HEART FAILURE (RHF), TRANSPLANTATION, AND DEATH. ANY NEUROLOGICAL EVENT (NE), SUCH AS ISCHEMIC STROKE (IS) AND HEMORRHAGIC STROKE (HS), NONSURGICAL BLEEDING (NSB), INCLUDING GASTROINTESTINAL BLEEDING (GIB), AND SUSPECTED OR CONFIRMED PUMP THROMBOSIS (PT), CATEGORIZED IN THROMBOEMBOLIC EVENT, WERE CONSOLIDATED INTO A COMPOSITE ENDPOINT REFERRED TO AS ANY HRAE. THIS RETROSPECTIVE SINGLE- CENTER STUDY INCLUDED 33 HM3 PATIENTS (15.2% FEMALE, AGE 66 (9.5) YEARS), IMPLANTED BETWEEN JUNE 2014 AND AUGUST 2018. IT WAS NOTED THAT OVER THE 24- MONTH FOLLOW- UP PERIOD, THE STUDY COHORT FURTHER REDUCED IN SIZE AS PATIENTS FAILED TO REACH THE DEFINED LONGITUDINAL TIMEPOINTS, MAINLY DUE TO TRANSPLANTATION OR DEATH. PRE AND POSTOPERATIVE CHEST X-RAY WERE TAKEN (=14 DAYS PRE- IMPLANTATION, FIRST POSTOPERATIVE STANDING, 6-, 12-, 18-, AND 24-MONTHS POST- IMPLANTATION) WERE ANALYZED FOR THEIR ASSOCIATION WITH HRAES OVER 24 MONTHS. OF THE 33 PATIENTS, 15 EXPERIENCED HRAES WITHIN A TWO-YEAR FOLLOW-UP. 10 OF THE 15 PATIENTS EXPERIENCED NE, 6 EXPERIENCED NSB, AND 1 EXPERIENCED THROMBOEMBOLIC EVENT. 8 OF THE 33 PATIENTS REQUIRED A TEMPORARY RVAD. FURTHERMORE, OF THE 33 PATIENTS 9 WERE REPORTED AS DECEASED WITHIN THE TWO YEAR FOLLOW UP. 4 DIED DUE TO NEUROLOGICAL EVENTS, 3 DIED DIE TO MULTIORGAN FAILURE, AND 2 DUE TO OTHER CAUSES. 8 OUT OF THE 9 PATIENTS THAT WERE DECEASED HAD EXPERIENCED HRAE¿S. HM3 PATIENTS WITH ANY HRAE SHOWED SIGNIFICANTLY LOWER PREOPERATIVE DISTANCES BETWEEN LEFT VENTRICLE AND THORACIC OUTLINE (DLVT) (25.3 ± 10.2 MM VS. 40.3 ± 15.5 MM, P = 0.004). A ROC- DERIVED CUTOFF DLVT = 29.2 MM PROVIDED 85.7% SENSITIVITY AND 72.2% SPECIFICITY PREDICTING ANY HRAE DURING HM3 SUPPORT (76.2% (>29.2 MM) VS. 16.7% (=29.2 MM) FREEDOM FROM HRAE, P < 0.001) AND SIGNIFICANT DIFFERENCES IN CARDIOTHORACIC RATIO (0.58 ± 0.04 VS. 0.62 ± 0.04, P = 0.045). POSTOPERATIVE X- RAYS INDICATED LOWER PUMP DEPTHS IN PATIENTS WITH ISCHEMIC STROKES (9.1 ± 16.2 MM VS. 38.0 ± 18.5 MM, P = 0.007), REDUCED FREEDOM FROM ANY NEUROLOGICAL EVENT (PUMP DEPTH = 28.7 MM: 45.5% VS. 94.1%, P = 0.004), AND A SIGNIFICANT CORRELATION BETWEEN PUMP DEPTH AND INFLOW CANNULA (IC) ANGLE (R = 0.66, P < 0.001). IN CONCLUSION, THE IDENTIFICATION OF PREOPERATIVE X- RAY MARKERS, WHICH HELP QUANTIFY CARDIAC ENLARGEMENT IN RELATION TO THORACIC DIMENSIONS, SERVES AS A RISK FACTOR FOR HRAES DURING HM3 SUPPORT AND SIGNIFIES AN ADVANCEMENT IN UNDERSTANDING PATIENT OUTCOMES IN THIS CONTEXT. ADDITIONALLY, THE POSTOPERATIVE PUMP DEPTH NOT ONLY SERVES AS A PREDICTOR FOR IS AND CORRELATES WITH THE IC ANGLE BUT ALSO REMAINS CONSISTENT OVER TIME. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE (HM) 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS INVESTIGATION. THE HEARTMATE 3 DEVICE SERIAL NUMBERS AND OTHER SPECIFIC CASE/PATIENT INFORMATION ARE NOT AVAILABLE AND WERE NOT REQUESTED. NO PRODUCT WAS EVALUATED UNDER THIS COMPLAINT. THE SERIAL NUMBERS OR OTHER IDENTIFYING INFORMATION OF THE PRODUCTS WERE NOT REPORTED AND WERE NOT ABLE TO BE DETERMINED DURING THE INVESTIGATION. THE HEARTMATE 3 LVAS IFU AND THE HEARTMATE 3 PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING STROKE, MULTI ORGAN FAILURE (RESPIRATORY FAILURE, RIGHT HEART FAILURE, RENAL DYSFUNCTION, AND HEPATIC DYSFUNCTION), AND DEATH. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿ PROVIDES INFORMATION REGARDING ANTICOAGULATION, INCLUDING RECOMMENDED INTERNATIONAL NORMALIZED RATIO (INR) VALUES. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: B5:SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE DEVICE WAS IMPLANTED AT TIME OF EVENT. DATE OF EVENT HAS BEEN ENTERED AS: 01AUG2018 AS PATIENTS WERE IMPLANTED BETWEEN JUN2014 AND AUG2018. DEPARTMENT OF CARDIAC SURGERY, MEDICAL UNIVERSITY OF VIENNA, VIENNA, AUSTRIA 2CENTER FOR MEDICAL PHYSICS AND BIOMEDICAL ENGINEERING, MEDICAL UNIVERSITY OF VIENNA, VIENNA, AUSTRIA 3LUDWIG BOLTZMANN INSTITUTE FOR CARDIOVASCULAR RESEARCH, VIENNA, AUSTRIA 4AUSTRIAN CLUSTER FOR TISSUE REGENERATION, VIENNA, AUSTRIA WIDHALM, G., AIGNER, P., GRUBER, B., MOSCATO, F., MOAYEDIFAR, R., SCHAEFER, A., DIMITROV, K., ZIMPFER, D., RIEBANDT, J., & SCHLÖGLHOFER, T. (2024). PREOPERATIVE ANATOMICAL LANDMARKS AND LONGITUDINAL HEARTMATE 3 PUMP POSITION IN X-RAYS: RELEVANCE FOR ADVERSE EVENTS. ARTIFICIAL ORGANS, 48(12), 1502¿1512. HTTPS://DOI.ORG/10.1111/AOR.14837 NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21506443&pc=BZD | 2518422-2025-102547 | 2024-12-17 05:00 | Death | RESPIRONICS, INC. | 2025-03-03 05:00 | BZD | DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT OF THE DREAMSTATION AUTO CPAP DEVICE HAD PASSED AWAY. THERE IS NO ALLEGATION OF MALFUNCTION, OR THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION AND THE DEVICE IS CURRENTLY REGISTERED TO A NEW USER. AT THIS TIME, THE DEVICE WILL NOT BE RETURNING, AND NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21506487&pc=DSQ | 2916596-2025-01170 | 2025-02-09 05:00 | Death | THORATEC CORPORATION | 2025-03-03 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Respiratory Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY (B)(6) 2025 DUE TO ACUTE RESPIRATORY FAILURE RELATED TO A NEUROLOGICAL EVENT. THE OUTCOME WAS STATED NOT TO BE DEVICE OR THERAPY RELATED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS RESPIRATORY FAILURE, NEUROLOGICAL EVENTS, AND DEATH AS POTENTIAL ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, LISTS NEUROLOGIC DYSFUNCTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21507291&pc=DSQ | 2916596-2025-01069 | 2025-02-09 05:00 | Death | THORATEC CORPORATION | 2025-03-03 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS MULTI ORGANS FAILURE AND WAS NOT DEVICE RELATED. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. THE DEVICE FUNCTIONED AS EXPECTED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED AND WOULD NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. IT WAS COMMUNICATED THAT NO AUTOPSY WAS PERFORMED AND THAT THE PUMP WAS NOT EXPLANTED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) REV. G IS CURRENTLY AVAILABLE. THE IFU LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21507298&pc=DSQ | 3007042319-2025-00584 | 2024-11-08 05:00 | Death | HEARTWARE, INC. | 2025-03-03 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Dyspnea; Pleural Effusion; Septic Shock; Ascites; Heart Failure/Congestive Heart Failure; Tricuspid Valve Insufficiency/ Regurgitation | P100047 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING TRICUSPID VALVE REPAIR (TVR) IN PATIENTS WITH RIGHT HEART DYSFUNCTION AFTER LEFT VENTRICULAR ASSIST DEVICE (LVAD) IMPLANTATION. THE AUTHORS DESCRIBED A PATIENT WHO EXPERIENCED SEVERE RIGHT VENTRICULAR (RV) DYSFUNCTION POST VAD IMPLANT. TEMPORARY RIGHT VAD WAS REQUIRED, AND THE RVAD WAS REMOVED FIFTEEN DAYS LATER. AFTER BEING STABLE FOR MORE THAN TWO YEARS, THE PATIENT SHOWED INCREASED TRICUSPID REGURGITATION (TR) AND RECURRENTLY PRESENTED TO THE CLINIC. MORE THAN THREE YEARS AFTER VAD IMPLANTATION, THE PATIENT PRESENTED WITH MASSIVE ASCITES AND PLEURAL EFFUSION. ECHOCARDIOGRAPHY SHOWED AN IMPAIRED RV FUNCTION WITH SEVERE TR ALONG WITH INCREASED DYSPNEA, WHICH LED TO TVR. POSTOPERATIVELY, THE TRICUSPID VALVE SHOWED ONLY MINIMAL INSUFFICIENCY. HOWEVER, DURING FURTHER TREATMENT IN THE INTENSIVE CARE UNIT, THE PATIENT DEVELOPED A PARALYTIC ILEUS WHICH REQUIRED SURGICAL INTERVENTION (ILEOSTOMY). SUBSEQUENT COMPLICATIONS LED TO SEPTIC SHOCK, AND THE PATIENT DIED. THE STATUS OF THE DEVICE IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. THE EVENT ONLY OCCURRED WITH ONE PATIENT BUT SPECIFIC DETAILS ON THE PATIENT WERE NOT PROVIDED. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: TRICUSPID VALVE RECONSTRUCTION IN PATIENTS WITH RIGHT HEART DECOMPENSATION DUE TO SEVERE TRICUSPID REGURGITATION ON LVAD SUPPORT. JOURNAL OF CLINICAL MEDICINE. 2024, 13, 6705. DOI: 10.3390/JCM13226705 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21507926&pc=MVK | 3008642652-2025-02231 | 2025-02-05 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-03-03 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Incorrect Interpretation of Signal | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED AN INAPPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 17:19:39 ON (B)(6) 2025. AT 10:48:29 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS SINUS RHYTHM @ 80 BPM WITH TALL T WAVES. AT 10:49:20, THE PATIENT RECEIVED THE INAPPROPRIATE TREATMENT. DOUBLE COUNTING CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS SINUS RHYTHM @ 80 BPM WITH TALL T WAVES. POST SHOCK RHYTHM WAS SINUS RHYTHM 80 BPM WITH PVC¿S. THE ELECTRODE BELT WAS DISCONNECTED AT 11:17:39 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21507961&pc=MVK | 3008642652-2025-02232 | 2025-02-18 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-03-03 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED AN APPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 07:07:29 ON (B)(6) 2025. AT 08:05:31, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF. AT 08:06:08, THE PATIENT RECEIVED THE APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS ASYSTOLE. POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY. THE ELECTRODE BELT WAS DISCONNECTED AT 08:54:36 ON (B)(6) 2025. Manufacturer Narrative: THE ELECTRODE BELT AND MONITOR HAVE NOT BEEN RECOVERED. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21508876&pc=DSQ | 2916596-2025-01211 | 2024-04-01 04:00 | Death | THORATEC CORPORATION | 2025-03-03 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE RESEARCH ARTICLE ¿BEST PRACTICES TO MANAGE RIGHT VENTRICULAR DYSFUNCTION IN PATIENTS WITH LEFT VENTRICULAR ASSIST DEVICE THERAPY¿ CONDUCTED A RETROSPECTIVE STUDY OF 139 PATIENTS COLLECTED FROM THE CHRIST HOSPITAL, CINCINNATI. INVESTIGATORS SOUGHT TO IDENTIFY STRATEGIES IN WHICH TO AVOID RIGHT VENTRICULAR ASSIST DEVICE (RVAD) USE IN LVAD PATIENTS. DATA INCLUDE DEMOGRAPHICS, HEMODYNAMICS, ECHOCARDIOGRAM, LABORATORIES, INTRAOPERATIVE AND INTENSIVE CARE UNIT INFORMATION, POST OPERATIVE HEMODYNAMICS AND ECHOCARDIOGRAM AND SURVIVAL. OF THE 139 HEARTMATE 3 LVAD PATIENTS ENROLLED, 5 (3.6%), REQUIRED RV MECHANICAL SUPPORT. AVERAGE DURATION OF RVAD SUPPORT WAS 4.8 ± 7.3 DAYS (RANGE 0.18-17 DAYS). THOSE WHO AVOIDED AN RVAD HAD CHARACTERISTICS THAT REFLECTED THE HOSPITALS GENERAL PRACTICE PATTERN, INCLUDING CONCOMITANT TRICUSPID VALVE (TV) REPAIR IN 47%. THE 5 PATIENTS RECEIVING RVAD SUPPORT TENDED TO BE OLDER, WITH LOWER PAPI AT BASELINE. THEY RECEIVED MORE BLOOD PRODUCTS, AND ONLY 1 OF 5 RECEIVED TV REPAIR. INITIAL PUMP SPEED WAS LOWER, AND THEY WERE TREATED FOR LONGER PERIOD WITH INOTROPE(S). CVP ALSO TENDED TO BE HIGHER AS WELL. THIRTY-DAY SURVIVAL WAS 98.5% AND 80% FOR THE NO RVAD GROUP AND RVAD GROUPS RESPECTIVELY. CONCLUSION: IN PATIENTS RECEIVING LVAD THERAPY, CLINICAL STRATEGY OF REDUCING TRANSFUSIONS, TRICUSPID VALVE REPAIR, MAINTAIN INTRA AND POST OPERATIVE MEAN ARTERIAL PRESSURE > 75MMHG, AND BALANCING THE VENTRICULAR SEPTUM APPEARS TO MINIMIZE THE NEED FOR RVAD SUPPORT. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. NO PRODUCT WAS EVALUATED UNDER THIS COMPLAINT. THE HEARTMATE 3 DEVICE SERIAL NUMBERS, AS WELL AS OTHER SPECIFIC CASE/PATIENT INFORMATION, ARE NOT AVAILABLE. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. IT WAS REPORTED THAT THE HEARTMATE 3 LVAS WAS USED AS A RIGHT VENTRICULAR ASSIST DEVICE (RVAD) WHICH IS NOT AN APPROVED INDICATION. THE HEARTMATE 3 IS INDICATED FOR LVAD SUPPORT, AND ALL LABELING, PHYSICIAN TRAINING AND CUSTOMER MARKETING MATERIALS ARE ALIGNED WITH THIS INDICATION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. DATE OF EVENT HAS BEEN ENTERED AS THE DATE OF PUBLICATION, 01APR2024, SINCE THE DATE OF DATA COLLECTION WAS NOT PROVIDED. STOKES, J., BRANDT, A., GUPTA, S., EMANI, S., ANSWINI, G., MEHAL, A., BARTONE, C., CHUNG, E. S., & EGNACZYK, G. (2024). BEST PRACTICES TO MANAGE RIGHT VENTRICULAR DYSFUNCTION IN PATIENTS WITH LEFT VENTRICULAR ASSIST DEVICE THERAPY. THE JOURNAL OF HEART AND LUNG TRANSPLANTATION, 43(4). HTTPS://DOI.ORG/10.1016/J.HEALUN.2024.02.1261 CHRIST HOSPITAL, CINCINNATI, OH; UNIVERSITY OF CINCINNATI, CINCINNATI, OH; THE CHRIST HOSPITAL, CINCINNATI, OH. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21508997&pc=OUT | 2029214-2025-00544 | 2025-02-25 05:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-03-03 05:00 | OUT | PIPELINE VANTAGE WITH SHIELD TECHNOLOGY | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage; Obstruction/Occlusion; Vascular Dissection; Cerebral Edema; Hemorrhagic Stroke |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING THE STUDY TITLED "EFFICACY AND SAFETY IN THE USE OF PIPELINE VANTAGE SHIELD STENT VERSUS PIPELINE FLEX SHIELD STENT IN THE TREATMENT OF PATIENTS WITH UNRUPTURED INTRACRANIAL ANEURYSMS: A MULTICENTER STUDY." THE OBJECTIVE IS TO COMPARE THE EFFICACY AND SAFETY OF THE PIPELINE VANTAGE EMBOLIZATION DEVICE WITH SHIELD TECHNOLOGY (PEDV) AND THE PIPELINE FLEX EMBOLIZATION DEVICE WITH SHIELD TECHNOLOGY (PED-SHIELD) IN MANAGING UNRUPTURED INTRACRANIAL ANEURYSMS. THE TIME FRAME OF THIS STUDY WAS FROM JANUARY 2017 UNTIL JUNE 2023. FIVE HUNDRED AND FORTY-SIX PATIENTS WITH 574 ANEURYSMS WERE TREATED, AND 571 PROCEDURES WITH 586 STENTS (445 FOR PED-SHIELD AND 141 FOR PEDV). THE MEAN AGE WAS 58 YEARS (18-85) WITH 85% FEMALE PATIENTS. THE FOLLOWING MEDTRONIC DEVICES WERE USED: PIPELINE VANTAGE EMBOLIZATION DEVICE WITH SHIELD TECHNOLOGY (PEDV) AND THE PIPELINE FLEX EMBOLIZATION DEVICE WITH SHIELD TECHNOLOGY (PED-SHIELD). DEATHS OCCURRED IN THE STUDY POPULATION. THE OVERALL MORTALITY RATE WAS 0.9%, WITH FIVE PATIENTS DECEASED: -TWO IN THE PEDV GROUP (DUE TO STENT OCCLUSION WITH POSTERIOR INFECTION AND MALIGNANT EDEMA AND A HEMORRHAGIC EVENT AFTER ICA DISSECTION). -THREE IN THE PED-SHIELD GROUP (TWO DUE TO HEMORRHAGIC EVENTS AFTER ANEURYSM RUPTURE AND INTRAPARENCHYMAL HEMATOMA, AND ONE NON-NEUROLOGICAL DEATH (GUNSHOT)). AMONG PATIENT ADVERSE EVENTS INCLUDED: -28 COMPLICATIONS WERE PRESENTED (24 NEUROLOGICAL AND 4 NON-NEUROLOGICAL), ALSO REGISTERED THROMBOEMBOLIC (18) AND HEMORRHAGIC (10) COMPLICATIONS. -THROMBOEMBOLIC EVENTS AT 3.1% OVERALL, WITH SPECIFIC RATES OF 2.1% FOR PED-SHIELD AND 6.7% FOR PEDV. DISABLING SEQUELAE WERE ESTABLISHED IN SIX PATIENTS. TWO PATIENTS HAD IN-STENT THROMBOSIS TREATED SUCCESSFULLY WITH TIROFIBAN DURING EMBOLIZATION, AND ONE OF THEM HAD ANOTHER DISABLING EVENT AT SIX MONTHS WITH MILD DYSARTHRIA AND RIGHT UPPER PARESIS (MRS=2). THE OTHER FOUR CASES INCLUDED TWO PUTAMINAL LEFT AND RIGHT MILD STROKES (MRS=1), ONE PERFORANT STROKE (MRS=1), AND ONE RIGHT MCA OCCLUSION FOUR DAYS POST-PROCEDURE, DESPITE DOUBLE ANTI-AGGREGATION. THROMBECTOMY FAILED, LEAVING THE STENT OCCLUDED (MRS=4). -OVERALL PERIPROCEDURAL EVENTS WERE 1.4%, FOR THE PEDV STENT IN SIX PATIENTS (4.3%) AND TWO (0.4%) FOR THE PED-SHIELD STENT (P=0<(> <<)>.001). DISABLING SEQUELAE WERE REGISTERED IN THREE PATIENTS (0.3%), TWO IN THE PEDV STENT GROUP WITH MILD DYSARTHRIA DUE TO C LOPIDOGREL RESISTANCE (MRS=1) AND LEFT UPPER ARM PARESIS (MRS=1) AFTER THROMBOEMBOLIC EVENT IN TWISTED STENTS WHILE DEPLOYMENT NECESSITATING A THIRD STENT. ANOTHER PATIENT IN THE PED-SHIELD GROUP WITH MILD RIGHT UPPER ARM PARESIS (MRS=1) AT 12 MONTHS AFTER CLOPIDOGREL RESISTANCE SWITCHED TO TICAGRELOR AND ASPIRIN. -AT 6 MONTHS, OVERALL NEW ISCHEMIC COMPLICATIONS WERE REGISTERED IN NINE PATIENTS (1.6%), FOR THE PEDV STENT TWO PATIENTS (1.4%), AND SEVEN PATIENTS (1.6%) FOR THE PED-SHIELD STENT (P=0.63). -AT 12 MONTHS, ONLY THE PEDV STENT REGISTERED ONE THROMBOEMBOLIC COMPLICATION (0.1%), WHICH OCCURRED DUE TO AN OCCLUSION OF TWO TELE SCOPED PEDS 8 MONTHS AFTER THE OCCLUSION OF A DISSECTING ICA CAVERNOUS ANEURYSM. THIS COMPLICATION WAS TREATED WITH INTRA-ARTERIAL THROMBOLYSIS AND ASPIRATION THROMBECTOMY, RESULTING IN NO DISABILITIES. -ELEVEN PATIENTS HAD MILD STROKES WITH NO DISABLING OUTCOMES. -OVERALL MORBIDITY AT 0.5%, THREE PATIENTS HAVING MRS=3-4: ONE IN THE PED-SHIELD GROUP AND TWO IN THE PEDV GROUP (0.3% FOR PEDV AND 0.2% FOR PED-SHIELD). -BLEEDING COMPLICATIONS AT 1.7% OVERALL, WITH SPECIFIC RATES OF 0.9% FOR PED-SHIELD AND 1.4% FOR PEDV. -TEN HEMORRHAGIC EVENTS WERE RECORDED, FOUR NON-NEUROLOGICAL COMPLICATIONS (BRACHIAL DISSECTION, UPPER DIGESTIVE BLEEDING, AND TWO F EMORAL PSEUDOANEURYSMS TREATED WITH STENTING AND COILING) AND SIX HEMORRHAGIC NEUROLOGICAL EVENTS. DISABLING SEQUELAE WERE REGISTERED IN FOUR PATIENTS. ONE PATIENT DEVELOPED MILD RIGHT UPPER ARM PARESIS (MRS=1) AT 12 MONTHS AFTER CLOPIDOGREL RESISTANCE. ONE ASYMPT OMATIC HYPERTENSIVE HEMATOMA RESOLVED SPONTANEOUSLY, AND ONE SYMPTOMATIC CASE RESOLVED AFTER FOUR WEEKS WITHOUT SEQUELAE (MRS=0). -OVERALL, 83% OF ANEURYSMS ACHIEVED OKM:C-D RESULTS (ADEQUATE ANGIOGRAPHIC OCCLUSION): 97% IN THE PEDV GROUP AND 80% IN THE PED-SHIELD GROUP (P=<(><<)>0.001). 17% HAD OKM B-A RESULTS (PERSISTENT REMNANT OR ABSENCE OF OCCLUSION). -TWO TELESCOPED STENTS IN 13 PATIENTS (4 PEDV, 9 PED-SHIELD) AND THREE STENTS IN ONE PATIENT (PED-SHIELD) WERE USED AS AN INITIAL TREATMENT FOR WIDE-NECK OR LARGE ANEURYSMS, IN CASES OF INADEQUATE OPENING OR MALAPPOSITION. INADEQUATE STENT WALL APPOSITION WERE FOUND IN HIGHER RATES WITH PEDV STENT 12% AND PEDSHIELD STENTS 3%, PROBABLY BECAUSE OF INADEQUATE ANCHORING DUE TO VESSEL TORTUOSITY AND INADEQUATE DIAMETER OF THE STENT SELECTED, MANAGED WITH ADDITIONAL STENTING, BALLOON ANGIOPLASTY AND IN ONE CASE WITH STENTECTOMY DUE TO INADEQUATE DEPLOYMENT. -INTIMAL HYPERPLASIA/IN STENT STENOSIS: PEDV 2.4% AND 1.9% FOR PED-SHIELD.Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED THAT BASED ON THE ADVERSE EVENTS MENTIONED IN THE ARTICLE, SOME CASES WERE ASSOCIATED WITH MEDTRONIC DEVICES, SPECIFICALLY THE PIPELINE EMBOLIZATION DEVICE (PED) VANTAGE AND PED SHIELD FLOW DIVERTERS. HOWEVER, IT IS IMPORTANT TO NOTE THAT WHILE THESE DEVICES WERE USED IN THE TREATMENT OF COMPLEX ANEURYSMS, THE ADVERSE EVENTS WERE PRIMARILY RELATED TO UNDERLYING PATHOLOGY, PROCEDURAL FACTORS, AND PATIENT-SPECIFIC CONDITIONS RATHER THAN DIRECT DEVICE FAILURE. ONE OF THE CENTERS HAS SUBMITTED TWO REPORTS TO MEDTRONIC. HOWEVER, THERE WERE NO FURTHER DETAILS ON THE SPECIFIC EVENTS. THE DEVICES INVOLVED WERE PEDV. THE MEDTRONIC DEVICES INVOLVED IN TWO SPECIFIC CASES FROM ONE CENTER INCLUDE THE FOLLOWING: PIPELINE¿ VANTAGE EMBOLIZATION DEVICE WITH SHIELD TECHNOLOGY (PEDV) PED3-027-550-30 (B)(6), PED3-027-600-20 (B)(6), PIPELINE¿ FLEX EMBOLIZATION DEVICE WITH SHIELD TECHNOLOGY (PED-SHIELD) PED2-450-12 (B)(6), PED2-450-20 (B)(6). THE DATES OF DEATH FOR ALL PATIENTS HAVE BEEN OBTAINED: PATIENT 1 ((B)(6) 2020), PATIENT 2 ((B)(6) 2020), PATIENT 3 ((B)(6) 2023), PATIENT 4 ((B)(6) 2021), AND PATIENT 5 ((B)(6) 2020). NO SPECIFIC INFORMAT ION ON WHICH SECTION OF THE PIPELINE FAILED TO OPEN WAS KNOWN. HOWEVER, POOR STENT OPENING OCCURRED IN 28 PATIENTS, THE MAJORITY OF WHICH WERE PROXIMAL: 15 WITH PEDV AND 13 WITH PED SHIELD. IN THE MAJORITY OF CASES, THERE WAS A SMALL GAP BETWEEN THE STENT AND THE ARTERIAL WALL, WHICH REQUIRED INTRASTENT ANGIOPLASTY. ADDITIONALLY, THERE WAS A LACK OF CONTACT BETWEEN THE LOWER WALL OF THE ARTERY, REQUIRING INTRASTENT ANGIOPLASTY. PEDV: TWISTING OF THE TWO STENTS OCCURRED, REQUIRING THE IMPLEMENTATION OF A THIRD STENT (CORONARY). PED-SHIELD: NONE OF THE INADEQUATE OPENINGS WERE RELATED TO THE ADVERSE EVENTS MENTIONED. CASE SUMMARIES 1. FUSIFORM ANEURYSM OF THE LEFT INTERNAL CAROTID ARTERY (ICA): A FLOW-DIVERTING PED-SHIELD STENT WAS DEPLOYED FROM THE TOP OF THE LEFT ICA, COVERING THE ANEURYSM NECK. DIMINISHED FLOW WAS OBSERVED ON CONTROL ANGIOGRAPHY, PROMPTING MULTIPLE RETRIEVAL ATTEMPTS USING LOOPS AND ALLIGATOR RETRIEVAL DEVICES, ALL OF WHICH WERE UNSUCCESSFUL. DUE TO THE RISK OF FURTHER COMPLICATIONS, THE PROCEDURE WAS DISCONTINUED, BUT THE PATIENT ACHIEVED AN EXCELLENT FUNCTIONAL OUTCOME (MRS 0). 2. INTRASTENT THROMBUS FORMATION IN FLOW DIVERTER (PEDV) MESH: SUCCESSFULLY MANAGED WITH INTRAVENOUS TIROFIBAN ADMINISTRATION DURING THE PROCEDURE. NO FURTHER COMPLICATIONS OCCURRED, AND THE PATIENT ACHIEVED AN EXCELLENT FUNCTIONAL OUTCOME (MRS 0). 3. SYMPTOMATIC LARGE RIGHT CAVERNOUS SINUS ANEURYSM: TREATED WITHTWO PED VANTAGE DEVICES (5.5X30 MM AND 6X40 MM), WHICH EXPERIENCED STENT TWISTING. A 4.5 MM CORONARY STENT WAS DEPLOYED TO CORRECT THE TWISTING, ALONG WITH COILS FOR ANEURYSM OCCLUSION. THE PATIENT ACHIEVED A FAVORABLE FUNCTIONAL OUTCOME (MRS 1). 4. SPHENOIDECTOMY + SEPTOPLASTY WITH SUSPECTED INTRACRANIAL VASCULAR INJURY: DURING SURGERY, ARTERIAL-LIKE BLEEDING RAISED SUSPICION OF AN INTRACRANIAL VASCULAR INJURY. CT IMAGING SHOWED FILLING DEFECTS IN THE SUPRACLINOID RIGHT ICA. ENDOVASCULAR REPAIR WAS PERFORMED USING A FLOW-DIVERTING PEDV STENT, BUT A HEMODYNAMIC DEFECT WAS OBSERVED IN THE OPHTHALMIC ARTERY. FIVE HOURS POST-PROCEDURE, THE PATIENT DEVELOPED RIGHT EYE VISUAL IMPAIRMENT (LEFT INFERIOR QUADRANTANOPIA), PROMPTING STROKE CODE ACTIVATION. DOPPLER ULTRASOUND: ABSENT INTRACONAL FLOW AND NO FLOW IN THE RIGHT CENTRAL RETINAL ARTERY. THE PATIENT WAS TREATED WITH CLOPIDOGREL, TIROFIBAN INFUSION, AND ASA AND REMAINED UNDER ICU MONITORING. DISCHARGED FOUR DAYS LATER WITH IMPROVED SYMPTOMS (MRS 1). 5. INTRASTENT THROMBOSIS IN THE LEFT ICA POST-PROCEDURE WITH PEDV STENT: MANAGED WITH ANGIOPLASTY AND TIROFIBAN INFUSION. CLOPIDOGREL RESISTANCE WAS SUSPECTED, LEADING TO A SWITCH TO TICAGRELOR OR PRASUGREL. THE PATIENT ACHIEVED AN EXCELLENT FUNCTIONAL OUTCOME (MRS 0). 6. RUPTURED GIANT WIDE-NECK ANEURYSM TREATED WITH PEDV: HIGH-GRADE EXTENSIVE SUBARACHNOID HEMORRHAGE (SAH, FISHER IV). INTRASTENT THROMBOSIS WAS OBSERVED BUT IMPROVED WITH TIROFIBAN ADMINISTRATION. THE PATIENT HAD A FATAL OUTCOME DUE TO CEREBRAL EDEMA AND EXTERNAL VENTRICULAR DRAIN (EVD) OCCLUSION (MRS 6). THROMBOEMBOLIC EVENTS AND PERIOPERATIVE COMPLICATIONS WERE OBSERVED AT VARYING RATES ACROSS TREATMENT GROUPS, WITH HIGHER INCIDENCE IN THE PEDV COHORT (6.7%) COMPARED TO PED-SHIELD (2.1%). MOST CASES WERE SUCCESSFULLY MANAGED WITH TIROFIBAN INFUSION, ANGIOPLASTY, AND ANTIPLATELET ADJUSTMENTS, LEADING TO GOOD FUNCTIONAL OUTCOMES (MRS 0¿1). HOWEVER, A FATAL CASE OCCURRED IN A PATIENT WITH A RUPTURED GIANT WIDE-NECK ANEURYSM, WHERE INTRASTENT THROMBOSIS AND CEREBRAL EDEMA CONTRIBUTED TO MORTALITY. SIX-MONTH THROMBOEMBOLIC COMPLICATIONS RELATED TO DEVICE IMPLEMENTATION THROMBOEMBOLIC EVENT RATE: OVERALL INCIDENCE: 1.6%, PED SHIELD: 1.6%, PED VANTAGE (PEDV): 1.4% CASE SUMMARIES 1. STROKE IN THE RIGHT PUTAMEN WITH TERMINAL INTERNAL CAROTID ARTERY (TICA) INVOLVEMENT: CAUSED BY OCCLUSION OF A PERFORATING ARTERY IN A PATIENT TREATED WITH A PED SHIELD (2.5X14 MM). THE PATIENT ACHIEVED A FAVORABLE FUNCTIONAL OUTCOME (MRS 1). 2. ACUTE ISCHEMIC TRANSIENT (AIT) EVENT DUE TO INTIMAL HYPERPLASIA: MANAGED BY INTENSIFYING STATIN THERAPY FOR LIPID CONTROL AND PLAQUE STABILIZATION. TREATED WITH PEDV (4.5X20 MM). THE PATIENT HAD AN EXCELLENT FUNCTIONAL OUTCOME AT FOLLOW-UP (MRS 0). 3. STROKE IN THE LEFT PUTAMEN WITH TERMINAL ICA INVOLVEMENT: CAUSED BY OCCLUSION OF PERFORATING ARTERIES. TREATED WITH A PED SHIELD (2.5X20 MM). THE PATIENT ACHIEVED A FAVORABLE FUNCTIONAL OUTCOME (MRS 1). 4. LEFT-SIDED HEMIPARESIS, DYSARTHRIA, AND VISUAL ACUITY LOSS AFTER ANEURYSM TREATMENT WITH PEDV: NIHSS SCORE: 10 ON PRESENTATION. TREATED WITH INTRAVENOUS THROMBOLYSIS, RESULTING IN A NIHSS SCORE REDUCTION TO 3 POST-TREATMENTS. MODERATE FUNCTIONAL OUTCOME AT FOLLOW-UP (MRS 3). 5. PERFORATOR STROKE IN THE LEFT M1 SEGMENT: TREATED WITH PED SHIELD (2.75X20 MM) FOR A 7 MM LEFT MCA ANEURYSM INVOLVING THE M2 BRANCH. DESPITE THE PERFORATOR STROKE, THE PATIENT HAD NO RESIDUAL DEFICITS AND ACHIEVED A FAVORABLE FUNCTIONAL OUTCOME (MRS 1). 6. PERFORATOR STROKES IN THE LEFT MCA TERRITORY DURING DUAL ANTIPLATELET THERAPY (DAPT): THE PATIENT WAS TREATED WITH A PED SHIELD WHILE ON DAPT, WHICH WAS TEMPORARILY SUSPENDED DUE TO AN UPCOMING ORTHOPEDIC HAND SURGERY. FOLLOWING THE EVENT, THE PATIENT WAS TR ANSITIONEDTO TICAGRELOR MONOTHERAPY. EXCELLENT FUNCTIONAL OUTCOME (MRS 0). 7. FUSIFORM ANEURYSM OF THE RIGHT MCA (5X3 MM) INVOLVING THE M2 BRANCH: TREATED WITH TWO PED SHIELD DEVICES (2.75X16 MM AND 3X20 MM). THE FIRST STENT SHORTENED AND DID NOT FULLY COVER THE ANEURYSM, REQUIRING A SECOND STENT DEPLOYMENT. THE PROCEDURE WAS SUCCESSFUL, AND THE PATIENT ACHIEVED AN EXCELLENT FUNCTIONAL OUTCOME (MRS 1). 8. RIGHT MCA OCCLUSION POST-PROCEDURE WHILE ON ASPIRIN AND CLOPIDOGREL TREATED WITH PED SHIELD: MECHANICAL THROMBECTOMY (MTB) WITH THE ADAPT TECHNIQUE WAS PERFORMED. INTRASTENT OCCLUSION OCCURRED, LEADING TO A POOR FUNCTIONAL OUTCOME (MRS 4). 9. CHOROIDAL ARTERY STROKE SIX WEEKS POST-PROCEDURE TREATED WITH PED SHIELD: DELAYED PRESENTATION (24 HOURS AFTER SYMPTOM ONSET) AT A LOCAL CLINIC, RESULTING IN A MISSED TREATMENT WINDOW. DESPITE THIS, THE PATIENT ACHIEVED A FAVORABLE FUNCTIONAL OUTCOME (MRS 1). 12-MONTH THROMBOEMBOLIC COMPLICATIONS RELATED TO DEVICE IMPLEMENTATION: OVERALL INCIDENCE: 0.1% CASE SUMMARY: EIGHT MONTHS POST-ANEURYSM TREATMENT, OCCLUSION OF THE PEDV DEVICE OCCURRED. SUCCESSFULLY MANAGED WITH INTRA-ARTERIAL THROMBOLYSIS AND ASPIRATION THROMBECTOMY. THE PATIENT ACHIEVED RECOVERY, WITH A MODIFIED RANKIN SCALE (MRS) SCORE OF 2. HEMORRHAGIC EVENT RATE: BLEEDING COMPLICATIONS AT 1.7% OVERALL, WITH SPECIFIC RATES OF 0.9% FOR PED-SHIELD AND 1.4% FOR PEDV. NONE OF THE EVENTS WERE DIRECTLY RELATED TO THE DEVICES. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: G2: CITATION: AUTHORS: CATALÁ, A. J., ORTIZ-GIRALDO, A. F., ESTÉVEZ-OCHOA, M. F., ALVARADO-BEDOYA, C. S., MUÑOZ ORDOÑEZ, J. R., MEJÍ A, J. A., PERILLA, L. C., PACHECO, C. D., RUIZ, A. H., MUÑOZ, J. G., DE GRACIA, S. G., SERR. EFFICACY AND SAFETY IN THE USE OF PIPELINE VANTAGE SHIELD STENT VERSUS PIPELINE FLEX SHIELD STENT IN THE TREATMENT OF PATIENTS WITH UNRUPTURED INTRACRANIAL ANEURYSMS: A MU LTICENTER STUDY. AJNR. AMERICAN JOURNAL OF NEURORADIOLOGY. 2025. HTTPS://DOI.ORG/10.3174/AJNR.A8719 A.2. THIS VALUE IS THE AVERAGE AGE OF THE PATIENTS REPORTED IN THE ARTICLE AS SPECIFIC PATIENTS COULD NOT BE IDENTIFIED. A.3. THIS VALUE REFLECTS THE SEX OF THE MAJORITY OF THE PATIENTS REPORTED IN THE ARTICLE AS SPECIFIC PATIENTS COULD NOT BE IDENTIFIED. EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21509051&pc=OUT | 2029214-2025-00543 | 2025-02-25 05:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-03-03 05:00 | OUT | PIPELINE FLEX W/SHIELD TECHNOLOGY | Adverse Event Without Identified Device or Use Problem | Hemorrhagic Stroke; Ruptured Aneurysm |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING THE STUDY TITLED "EFFICACY AND SAFETY IN THE USE OF PIPELINE VANTAGE SHIELD STENT VERSUS PIPELINE FLEX SHIELD STENT IN THE TREATMENT OF PATIENTS WITH UNRUPTURED INTRACRANIAL ANEURYSMS: A MULTICENTER STUDY." THE OBJECTIVE IS TO COMPARE THE EFFICACY AND SAFETY OF THE PIPELINE VANTAGE EMBOLIZATION DEVICE WITH SHIELD TECHNOLOGY (PEDV) AND THE PIPELINE FLEX EMBOLIZATION DEVICE WITH SHIELD TECHNOLOGY (PED-SHIELD) IN MANAGING UNRUPTURED INTRACRANIAL ANEURYSMS. THE TIME FRAME OF THIS STUDY WAS FROM JANUARY 2017 UNTIL JUNE 2023. FIVE HUNDRED AND FORTY-SIX PATIENTS WITH 574 ANEURYSMS WERE TREATED, AND 571 PROCEDURES WITH 586 STENTS (445 FOR PED-SHIELD AND 141 FOR PEDV). THE MEAN AGE WAS 58 YEARS (18-85) WITH 85% FEMALE PATIENTS. THE FOLLOWING MEDTRONIC DEVICES WERE USED: PIPELINE VANTAGE EMBOLIZATION DEVICE WITH SHIELD TECHNOLOGY (PEDV) AND THE PIPELINE FLEX EMBOLIZATION DEVICE WITH SHIELD TECHNOLOGY (PED-SHIELD). DEATHS OCCURRED IN THE STUDY POPULATION. THE OVERALL MORTALITY RATE WAS 0.9%, WITH FIVE PATIENTS DECEASED: -TWO IN THE PEDV GROUP (DUE TO STENT OCCLUSION WITH POSTERIOR INFECTION AND MALIGNANT EDEMA AND A HEMORRHAGIC EVENT AFTER ICA DISSECTION). -THREE IN THE PED-SHIELD GROUP (TWO DUE TO HEMORRHAGIC EVENTS AFTER ANEURYSM RUPTURE AND INTRAPARENCHYMAL HEMATOMA, AND ONE NON-NEUROLOGICAL DEATH (GUNSHOT)). AMONG PATIENT ADVERSE EVENTS INCLUDED: -28 COMPLICATIONS WERE PRESENTED (24 NEUROLOGICAL AND 4 NON-NEUROLOGICAL), ALSO REGISTERED THROMBOEMBOLIC (18) AND HEMORRHAGIC (10) COMPLICATIONS. -THROMBOEMBOLIC EVENTS AT 3.1% OVERALL, WITH SPECIFIC RATES OF 2.1% FOR PED-SHIELD AND 6.7% FOR PEDV. DISABLING SEQUELAE WERE ESTABLISHED IN SIX PATIENTS. TWO PATIENTS HAD IN-STENT THROMBOSIS TREATED SUCCESSFULLY WITH TIROFIBAN DURING EMBOLIZATION, AND ONE OF THEM HAD ANOTHER DISABLING EVENT AT SIX MONTHS WITH MILD DYSARTHRIA AND RIGHT UPPER PARESIS (MRS=2). THE OTHER FOUR CASES INCLUDED TWO PUTAMINAL LEFT AND RIGHT MILD STROKES (MRS=1), ONE PERFORANT STROKE (MRS=1), AND ONE RIGHT MCA OCCLUSION FOUR DAYS POST-PROCEDURE, DESPITE DOUBLE ANTI-AGGREGATION. THROMBECTOMY FAILED, LEAVING THE STENT OCCLUDED (MRS=4). -OVERALL PERIPROCEDURAL EVENTS WERE 1.4%, FOR THE PEDV STENT IN SIX PATIENTS (4.3%) AND TWO (0.4%) FOR THE PED-SHIELD STENT (P=0<(> <<)>.001). DISABLING SEQUELAE WERE REGISTERED IN THREE PATIENTS (0.3%), TWO IN THE PEDV STENT GROUP WITH MILD DYSARTHRIA DUE TO C LOPIDOGREL RESISTANCE (MRS=1) AND LEFT UPPER ARM PARESIS (MRS=1) AFTER THROMBOEMBOLIC EVENT IN TWISTED STENTS WHILE DEPLOYMENT NECESSITATING A THIRD STENT. ANOTHER PATIENT IN THE PED-SHIELD GROUP WITH MILD RIGHT UPPER ARM PARESIS (MRS=1) AT 12 MONTHS AFTER CLOPIDOGREL RESISTANCE SWITCHED TO TICAGRELOR AND ASPIRIN. -AT 6 MONTHS, OVERALL NEW ISCHEMIC COMPLICATIONS WERE REGISTERED IN NINE PATIENTS (1.6%), FOR THE PEDV STENT TWO PATIENTS (1.4%), AND SEVEN PATIENTS (1.6%) FOR THE PED-SHIELD STENT (P=0.63). -AT 12 MONTHS, ONLY THE PEDV STENT REGISTERED ONE THROMBOEMBOLIC COMPLICATION (0.1%), WHICH OCCURRED DUE TO AN OCCLUSION OF TWO TELE SCOPED PEDS 8 MONTHS AFTER THE OCCLUSION OF A DISSECTING ICA CAVERNOUS ANEURYSM. THIS COMPLICATION WAS TREATED WITH INTRA-ARTERIAL THROMBOLYSIS AND ASPIRATION THROMBECTOMY, RESULTING IN NO DISABILITIES. -ELEVEN PATIENTS HAD MILD STROKES WITH NO DISABLING OUTCOMES. -OVERALL MORBIDITY AT 0.5%, THREE PATIENTS HAVING MRS=3-4: ONE IN THE PED-SHIELD GROUP AND TWO IN THE PEDV GROUP (0.3% FOR PEDV AND 0.2% FOR PED-SHIELD). -BLEEDING COMPLICATIONS AT 1.7% OVERALL, WITH SPECIFIC RATES OF 0.9% FOR PED-SHIELD AND 1.4% FOR PEDV. -TEN HEMORRHAGIC EVENTS WERE RECORDED, FOUR NON-NEUROLOGICAL COMPLICATIONS (BRACHIAL DISSECTION, UPPER DIGESTIVE BLEEDING, AND TWO F EMORAL PSEUDOANEURYSMS TREATED WITH STENTING AND COILING) AND SIX HEMORRHAGIC NEUROLOGICAL EVENTS. DISABLING SEQUELAE WERE REGISTERED IN FOUR PATIENTS. ONE PATIENT DEVELOPED MILD RIGHT UPPER ARM PARESIS (MRS=1) AT 12 MONTHS AFTER CLOPIDOGREL RESISTANCE. ONE ASYMPT OMATIC HYPERTENSIVE HEMATOMA RESOLVED SPONTANEOUSLY, AND ONE SYMPTOMATIC CASE RESOLVED AFTER FOUR WEEKS WITHOUT SEQUELAE (MRS=0). -OVERALL, 83% OF ANEURYSMS ACHIEVED OKM:C-D RESULTS (ADEQUATE ANGIOGRAPHIC OCCLUSION): 97% IN THE PEDV GROUP AND 80% IN THE PED-SHIELD GROUP (P=<(><<)>0.001). 17% HAD OKM B-A RESULTS (PERSISTENT REMNANT OR ABSENCE OF OCCLUSION). -TWO TELESCOPED STENTS IN 13 PATIENTS (4 PEDV, 9 PED-SHIELD) AND THREE STENTS IN ONE PATIENT (PED-SHIELD) WERE USED AS AN INITIAL TREATMENT FOR WIDE-NECK OR LARGE ANEURYSMS, IN CASES OF INADEQUATE OPENING OR MALAPPOSITION. INADEQUATE STENT WALL APPOSITION WERE FOUND IN HIGHER RATES WITH PEDV STENT 12% AND PEDSHIELD STENTS 3%, PROBABLY BECAUSE OF INADEQUATE ANCHORING DUE TO VESSEL TORTUOSITY AND INADEQUATE DIAMETER OF THE STENT SELECTED, MANAGED WITH ADDITIONAL STENTING, BALLOON ANGIOPLASTY AND IN ONE CASE WITH STENTECTOMY DUE TO INADEQUATE DEPLOYMENT. -INTIMAL HYPERPLASIA/IN STENT STENOSIS: PEDV 2.4% AND 1.9% FOR PED-SHIELDEvent Description: ADDITIONAL INFORMATION RECEIVED REPORTED THAT BASED ON THE ADVERSE EVENTS MENTIONED IN THE ARTICLE, SOME CASES WERE ASSOCIATED WITH MEDTRONIC DEVICES, SPECIFICALLY THE PIPELINE EMBOLIZATION DEVICE (PED) VANTAGE AND PED SHIELD FLOW DIVERTERS. HOWEVER, IT IS IMPORTANT TO NOTE THAT WHILE THESE DEVICES WERE USED IN THE TREATMENT OF COMPLEX ANEURYSMS, THE ADVERSE EVENTS WERE PRIMARILY RELATED TO UNDERLYING PATHOLOGY, PROCEDURAL FACTORS, AND PATIENT-SPECIFIC CONDITIONS RATHER THAN DIRECT DEVICE FAILURE. ONE OF THE CENTERS HAS SUBMITTED TWO REPORTS TO MEDTRONIC. HOWEVER, THERE WERE NO FURTHER DETAILS ON THE SPECIFIC EVENTS. THE DEVICES INVOLVED WERE PEDV. THE MEDTRONIC DEVICES INVOLVED IN TWO SPECIFIC CASES FROM ONE CENTER INCLUDE THE FOLLOWING: PIPELINE¿ VANTAGE EMBOLIZATION DEVICE WITH SHIELD TECHNOLOGY (PEDV)PED3-027-550-30 (B)(6), PED3-027-600-20 (B)(6), PIPELINE¿ FLEX EMBOLIZATION DEVICE WITH SHIELD TECHNOLOGY (PED-SHIELD) PED2-450-12 (B)(6), PED2-450-20 (B)(6). THE DATES OF DEATH FOR ALL PATIENTS HAVE BEEN OBTAINED: PATIENT 1 ((B)(6) 2020), PATIENT 2 ((B)(6) 2020), PATIENT 3 ((B)(6) 2023), PATIENT 4 ((B)(6) 2021), AND PATIENT 5 ((B)(6) 2020). NO SPECIFIC INFORMATION ON WHICH SECTION OF THE PIPELINE FAILED TO OPEN WAS KNOWN. HOWEVER, POOR STENT OPENING OCCURRED IN 28 PATIENTS, THE MAJORITY OF WHICH WERE PROXIMAL: 15 WITH PEDV AND 13 WITH PED SHIELD. IN THE MAJORITY OF CASES, THERE WAS A SMALL GAP BETWEEN THE STENT AND THE ARTERIAL WALL, WHICH REQUIRED INTRASTENT ANGIOPLASTY. ADDITIONALLY, THERE WAS A LACK OF CONTACT BETWEEN THE LOWER WALL OF THE ARTERY, REQUIRING INTRASTENT ANGIOPLASTY. PEDV: TWISTING OF THE TWO STENTS OCCURRED, REQUIRING THE IMPLEMENTATION OF A THIRD STENT (CORONARY). PED-SHIELD: NONE OF THE INADEQUATE OPENINGS WERE RELATED TO THE ADVERSE EVENTS MENTIONED. CASE SUMMARIES 1. FUSIFORM ANEURYSM OF THE LEFT INTERNAL CAROTID ARTERY (ICA): A FLOW-DIVERTING PED-SHIELD STENT WAS DEPLOYED FROM THE TOP OF THE LEFT ICA, COVERING THE ANEURYSM NECK. DIMINISHED FLOW WAS OBSERVED ON CONTROL ANGIOGRAPHY, PROMPTING MULTIPLE RETRIEVAL ATTEMPTS USING LOOPS AND ALLIGATOR RETRIEVAL DEVICES, ALL OF WHICH WERE UNSUCCESSFUL. DUE TO THE RISK OF FURTHER COMPLICATIONS, THE PROCEDURE WAS DISCONTINUED, BUT THE PATIENT ACHIEVED AN EXCELLENT FUNCTIONAL OUTCOME (MRS 0). 2. INTRASTENT THROMBUS FORMATION IN FLOW DIVERTER (PEDV) MESH: SUCCESSFULLY MANAGED WITH INTRAVENOUS TIROFIBAN ADMINISTRATION DURING THE PROCEDURE. NO FURTHER COMPLICATIONS OCCURRED, AND THE PATIENT ACHIEVED AN EXCELLENT FUNCTIONAL OUTCOME (MRS 0). 3. SYMPTOMATIC LARGE RIGHT CAVERNOUS SINUS ANEURYSM: TREATED WITH TWO PED VANTAGE DEVICES (5.5X30 MM AND 6X40 MM), WHICH EXPERIENCED STENT TWISTING. A 4.5 MM CORONARY STENT WAS DEPLOYED TO CORRECT THE TWISTING, ALONG WITH COILS FOR ANEURYSM OCCLUSION. THE PATIENT ACHIEVED A FAVORABLE FUNCTIONAL OUTCOME (MRS 1). 4. SPHENOIDECTOMY + SEPTOPLASTY WITH SUSPECTED INTRACRANIAL VASCULAR INJURY: DURING SURGERY, ARTERIAL-LIKE BLEEDING RAISED SUSPICION OF AN INTRACRANIAL VASCULAR INJURY. CT IMAGING SHOWED FILLING DEFECTS IN THE SUPRACLINOID RIGHT ICA. ENDOVASCULAR REPAIR WAS PERFORMED USING A FLOW-DIVERTING PEDV STENT, BUT A HEMODYNAMIC DEFECT WAS OBSERVED IN THE OPHTHALMIC ARTERY. FIVE HOURS POST-PROCEDURE, THE PATIENT DEVELOPED RIGHT EYE VISUAL IMPAIRMENT (LEFT INFERIOR QUADRANTANOPIA), PROMPTING STROKE CODE ACTIVATION. DOPPLER ULTRASOUND: ABSENT INTRACONAL FLOW AND NO FLOW IN THE RIGHT CENTRAL RETINAL ARTERY. THE PATIENT WAS TREATED WITH CLOPIDOGREL, TIROFIBAN INFUSION, AND ASA AND REMAINED UNDER ICU MONITORING. DISCHARGED FOUR DAYS LATER WITH IMPROVED SYMPTOMS (MRS 1). 5. INTRASTENT THROMBOSIS IN THE LEFT ICA POST-PROCEDURE WITH PEDV STENT: MANAGED WITH ANGIOPLASTY AND TIROFIBAN INFUSION. CLOPIDOGREL RESISTANCE WAS SUSPECTED, LEADING TO A SWITCH TO TICAGRELOR OR PRASUGREL. THE PATIENT ACHIEVED AN EXCELLENT FUNCTIONAL OUTCOME (MRS 0). 6. RUPTURED GIANT WIDE-NECK ANEURYSM TREATED WITH PEDV: HIGH-GRADE EXTENSIVE SUBARACHNOID HEMORRHAGE (SAH, FISHER IV). INTRASTENT THROMBOSIS WAS OBSERVED BUT IMPROVED WITH TIROFIBAN ADMINISTRATION. THE PATIENT HAD A FATAL OUTCOME DUE TO CEREBRAL EDEMA AND EXTERNAL VENTRICULAR DRAIN (EVD) OCCLUSION (MRS 6). THROMBOEMBOLIC EVENTS AND PERIOPERATIVE COMPLICATIONS WERE OBSERVED AT VARYING RATES ACROSS TREATMENT GROUPS, WITH HIGHER INCIDENCE IN THE PEDV COHORT (6.7%) COMPARED TO PED-SHIELD (2.1%). MOST CASES WERE SUCCESSFULLY MANAGED WITH TIROFIBAN INFUSION, ANGIOPLASTY, AND ANTIPLATELET ADJUSTMENTS, LEADING TO GOOD FUNCTIONAL OUTCOMES (MRS 0¿1). HOWEVER, A FATAL CASE OCCURRED IN A PATIENT WITH A RUPTURED GIANT WIDE-NECK ANEURYSM, WHERE INTRASTENT THROMBOSIS AND CEREBRAL EDEMA CONTRIBUTED TO MORTALITY. SIX-MONTH THROMBOEMBOLIC COMPLICATIONS RELATED TO DEVICE IMPLEMENTATION THROMBOEMBOLIC EVENT RATE: OVERALL INCIDENCE: 1.6%, PED SHIELD: 1.6%, PED VANTAGE (PEDV): 1.4% CASE SUMMARIES 1. STROKE IN THE RIGHT PUTAMEN WITH TERMINAL INTERNAL CAROTID ARTERY (TICA) INVOLVEMENT: CAUSED BY OCCLUSION OF A PERFORATING ARTERY IN A PATIENT TREATED WITH A PED SHIELD (2.5X14 MM). THE PATIENT ACHIEVED A FAVORABLE FUNCTIONAL OUTCOME (MRS 1). 2. ACUTE ISCHEMIC TRANSIENT (AIT) EVENT DUE TO INTIMAL HYPERPLASIA: MANAGED BY INTENSIFYING STATIN THERAPY FOR LIPID CONTROL AND PLAQUE STABILIZATION. TREATED WITH PEDV (4.5X20 MM). THE PATIENT HAD AN EXCELLENT FUNCTIONAL OUTCOME AT FOLLOW-UP (MRS 0). 3. STROKE IN THE LEFT PUTAMEN WITH TERMINAL ICA INVOLVEMENT: CAUSED BY OCCLUSION OF PERFORATING ARTERIES. TREATED WITH A PED SHIELD (2.5X20 MM). THE PATIENT ACHIEVED A FAVORABLE FUNCTIONAL OUTCOME (MRS 1). 4. LEFT-SIDED HEMIPARESIS, DYSARTHRIA, AND VISUAL ACUITY LOSS AFTER ANEURYSM TREATMENT WITH PEDV: NIHSS SCORE: 10 ON PRESENTATION. TREATED WITH INTRAVENOUS THROMBOLYSIS, RESULTING IN A NIHSS SCORE REDUCTION TO 3 POST-TREATMENTS. MODERATE FUNCTIONAL OUTCOME AT FOLLOW-UP (MRS 3). 5. PERFORATOR STROKE IN THE LEFT M1 SEGMENT: TREATED WITH PED SHIELD (2.75X20 MM) FOR A 7 MM LEFT MCA ANEURYSM INVOLVING THE M2 BRANCH. DESPITE THE PERFORATOR STROKE, THE PATIENT HAD NO RESIDUAL DEFICITS AND ACHIEVED A FAVORABLE FUNCTIONAL OUTCOME (MRS 1). 6. PERFORATOR STROKES IN THE LEFT MCA TERRITORY DURING DUAL ANTIPLATELET THERAPY (DAPT): THE PATIENT WAS TREATED WITH A PED SHIELD WHILE ON DAPT, WHICH WAS TEMPORARILY SUSPENDED DUE TO AN UPCOMING ORTHOPEDIC HAND SURGERY. FOLLOWING THE EVENT, THE PATIENT WAS TRANSITIONED TO TICAGRELOR MONOTHERAPY. EXCELLENT FUNCTIONAL OUTCOME (MRS 0). 7. FUSIFORM ANEURYSM OF THE RIGHT MCA (5X3 MM) INVOLVING THE M2 BRANCH: TREATED WITH TWO PED SHIELD DEVICES (2.75X16 MM AND 3X20 MM). THE FIRST STENT SHORTENED AND DID NOT FULLY COVER THE ANEURYSM, REQUIRING A SECOND STENT DEPLOYMENT. THE PROCEDURE WAS SUCCESSFUL, AND THE PATIENT ACHIEVED AN EXCELLENT FUNCTIONAL OUTCOME (MRS 1). 8. RIGHT MCA OCCLUSION POST-PROCEDURE WHILE ON ASPIRIN AND CLOPIDOGREL TREATED WITH PED SHIELD: MECHANICAL THROMBECTOMY (MTB) WITH THE ADAPT TECHNIQUE WAS PERFORMED. INTRASTENT OCCLUSION OCCURRED, LEADING TO A POOR FUNCTIONAL OUTCOME (MRS 4). 9. CHOROIDAL ARTERY STROKE SIX WEEKS POST-PROCEDURE TREATED WITH PED SHIELD: DELAYED PRESENTATION (24 HOURS AFTER SYMPTOM ONSET) AT A LOCAL CLINIC, RESULTING IN A MISSED TREATMENT WINDOW. DESPITE THIS, THE PATIENT ACHIEVED A FAVORABLE FUNCTIONAL OUTCOME (MRS 1). 12-MONTH THROMBOEMBOLIC COMPLICATIONS RELATED TO DEVICE IMPLEMENTATION: OVERALL INCIDENCE: 0.1% CASE SUMMARY: EIGHT MONTHS POST-ANEURYSM TREATMENT, OCCLUSION OF THE PEDV DEVICE OCCURRED. SUCCESSFULLY MANAGED WITH INTRA-ARTERIAL THROMBOLYSIS AND ASPIRATION THROMBECTOMY. THE PATIENT ACHIEVED RECOVERY, WITH A MODIFIED RANKIN SCALE (MRS) SCORE OF 2. HEMORRHAGIC EVENT RATE: BLEEDING COMPLICATIONS AT 1.7% OVERALL, WITH SPECIFIC RATES OF 0.9% FOR PED-SHIELD AND 1.4% FOR PEDV. NONE OF THE EVENTS WERE DIRECTLY RELATED TO THE DEVICES. Manufacturer Narrative: G2: CITATION: AUTHORS: CATALÁ, A. J., ORTIZ-GIRALDO, A. F., ESTÉVEZ-OCHOA, M. F., ALVARADO-BEDOYA, C. S., MUÑOZ ORDOÑEZ, J. R., MEJÍ A, J. A., PERILLA, L. C., PACHECO, C. D., RUIZ, A. H., MUÑOZ, J. G., DE GRACIA, S. G., SERR. EFFICACY AND SAFETY IN THE USE OF PIPELINE VANTAGE SHIELD STENT VERSUS PIPELINE FLEX SHIELD STENT IN THE TREATMENT OF PATIENTS WITH UNRUPTURED INTRACRANIAL ANEURYSMS: A MU LTICENTER STUDY. AJNR. AMERICAN JOURNAL OF NEURORADIOLOGY. 2025. HTTPS://DOI.ORG/10.3174/AJNR.A8719 A.2. THIS VALUE IS THE AVERAGE AGE OF THE PATIENTS REPORTED IN THE ARTICLE AS SPECIFIC PATIENTS COULD NOT BE IDENTIFIED. A.3. THIS VALUE REFLECTS THE SEX OF THE MAJORITY OF THE PATIENTS REPORTED IN THE ARTICLE AS SPECIFIC PATIENTS COULD NOT BE IDENTIFIED. EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21509504&pc=DSQ | 2916596-2025-01242 | 2025-02-13 05:00 | Death | THORATEC CORPORATION | 2025-03-03 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Disconnection; Mechanical Problem; Pumping Stopped | Brain Injury; Loss of consciousness; No Clinical Signs, Symptoms or Conditions | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT ARRIVED AT THE EMERGENCY DEPARTMENT VIA AMBULANCE. IT WAS NOTED THAT THE LOGFILE CAPTURED DRIVELINE DISCONNECT, LOW VOLTAGES, LOW FLOWS, AND PUMP STOPS. LOG FILES WERE SUBMITTED FOR REVIEW AND CAPTURED A STRING OF LOW POWER ADVISORIES ON (B)(6) 2025 AT 8:09 AM WHILE TETHERED ON BATTERIES WHERE THE WHITE POWER LEAD WAS DISCONNECTED FROM BATTERY POWER. THE BLACK POWER LEAD WAS CONNECTED TO A NEAR FULL CHARGED BATTERY FROM 8:10 AM TO 9:38 AM, WITH INTERMITTENT EPISODES OF LOW POWER HAZARD ALARMS. THERE WAS NO INTERRUPTION IN PUMP SUPPORT DURING THESE EVENTS. THE EVENT LOG RECORDED LOW_FLOW / DRIVELINE_DISCONNECT / LVAD_OFF ALARMS ON (B)(6) 2025 AT 9:40 AM FOR AROUND 4 SECONDS DUE TO POSSIBLE ELECTROSTATIC DISCHARGE (ESD) OR INTENTIONAL DRIVELINE DISCONNECT FROM THE SYSTEM CONTROLLER. IN THIS CASE IT APPEARED TO BE RELATED TO A DRIVELINE DISCONNECT EVENT. THE BLACK POWER LEAD WAS CONNECTED TO BATTERY POWER DURING THESE EVENTS. THE EVENT LOG RECORDED ADDITIONAL DRIVELINE_DISCONNECT / LVAD_OFF ALARMS ON 13FEB2025 AT 9:43 AM, INCLUDING NO EXTERNAL POWER ALARMS WHERE THE DRIVELINE WAS DISCONNECTED FROM THE SYSTEM CONTROLLER. IN ADDITION, THE CONTROLLER POWER LEADS WERE DISCONNECTED FROM EXTERNAL POWER WHERE THE SYSTEM CONTROLLER OPERATED ON EMERGENCY BACKUP BATTERY (EBB) POWER. THE SYSTEM CONTROLLER RECORDED SHUTDOWN_SEQUENCE_STARTED EVENTS AT AROUND 10:05 AM, WHERE IT APPEARS THE CONTROLLER WAS ATTEMPTED TO BE PLACED IN SLEEP MODE. THE CONTROLLER RECORDED SHUTDOWN_SEQUENCE_STARTED & SHUTDOWN_SEQUENCE_ABORTED EVENTS 13FEB2025 FROM 10:05 AM THROUGH 10:27 AM. IT WAS NOTED THE PATIENT WAS SLEEPING ON BATTERY POWER. IT WAS NOTED THAT THE BACKUP SYSTEM CONTROLLER WAS REPORTEDLY CONNECTED TO THE PATIENT AT SOME POINT. LOG FILES FOR THE BACKUP SYSTEM CONTROLLER WERE SUBMITTED FOR REVIEW AND DID NOT INDICATE THAT THE DRIVELINE WAS CONNECTED TO THE BACKUP SYSTEM CONTROLLER. IT APPEARED THAT THE BACKUP CONTROLLER WAS CONNECTED TO EXTERNAL POWER ON (B)(6)2025 FROM 10:06 AM THROUGH 12:33 PM.Event Description: IT WAS ADDITIONALLY REPORTED THAT A RESCUE SQUAD FOUND THE PATIENT AT HOME NOT CONNECTED TO A POWER SOURCE AND UNCONSCIOUS/UNRESPONSIVE WITH AGONAL BREATHING RHYTHM. THEY PERFORMED CARDIOPULMONARY RESUSCITATION (CPR) AND THEN PROCEEDED TO INITIATE ADVANCED CARDIAC LIFE SUPPORT (ACLS). THERE WAS NO LEFT VENTRICULAR ASSIST DEVICE (LVAD) HUM NOTED, AND THE PUMP RUNNING SYMBOL WAS NOT ILLUMINATED ON THE SYSTEM CONTROLLER. THE SYSTEM CONTROLLER DISPLAYED "CHARGING" MODE. THE DRIVELINE WAS CONNECTED TO THE PATIENT AND THE SYSTEM CONTROLLER HAD BOTH FULLY CHARGED 14V BATTERIES CONNECTED. AS ACLS CONTINUED, A SYSTEM CONTROLLER EXCHANGE WAS PERFORMED SUCCESSFULLY AND THE PUMP RESTARTED, AND THERE WAS AUSCULTATION OF AN LVAD HUM. THE PATIENT WAS NOTED TO BE IN A SINUS RHYTHM WITH UNKNOWN BLOOD PRESSURE AND WAS THEN TRANSFERRED TO A FACILITY AND ULTIMATELY PASSED AWAY DUE TO AN ANOXIC BRAIN INJURY. THE OUTCOME WAS CONSIDERED TO HAVE BEEN DEVICE RELATED OR POSSIBLY USER ERROR, SINCE THE DEVICE SEEMED TO HAVE OPERATED AS EXPECTED. AT THE TIME OF EXPLANT, THE PUMP WAS NOTED BY THE SITE TO BE BROKEN. Manufacturer Narrative: B2 - OUTCOMES ATTRIBUTED TO ADVERSE - CORRECTION. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE (HM) 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE CONTROLLER EVENT LOG FILE CONTAINED EVENTS FROM (B)(6) 2025 THROUGH (B)(6) 2025. ON (B)(6) 2024, THE DRIVELINE WAS DISCONNECTED TWICE, CAUSING THE PUMP TO STOP. FOLLOWING THE RECONNECTIONS OF THE DRIVELINE, THE PUMP RESUMED OPERATION AT THE SET SPEED WITHOUT ISSUE. THE PUMP APPEARED TO FUNCTION AS INTENDED AT THE FIXED SPEED WHILE THE DRIVELINE WAS CONNECTED. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU AND THE PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS NEUROLOGICAL EVENTS (NOT STROKE RELATED), AND DEATH AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 2 OF THE IFU, "SYSTEM OPERATIONS" (UNDER "SYSTEM CONTROLLER WARNINGS AND CAUTIONS"), AND SECTION 2 OF THE PATIENT HANDBOOK, ¿HOW YOUR HEART PUMP WORKS¿, INSTRUCT THE USER TO CHECK THE SYSTEM CONTROLLER DRIVELINE CONNECTOR TO CONFIRM THAT THE DRIVELINE IS SECURELY INSERTED IN THE SOCKET. IF THE DRIVELINE DISCONNECTS FROM THE SYSTEM CONTROLLER, THE PUMP STOPS. IF THE DRIVELINE DISCONNECTS FROM THE SYSTEM CONTROLLER, PROMPTLY RECONNECT IT TO RESUME PUMP OPERATION. THE PATIENT HANDBOOK FURTHER EXPLAINS THAT THE PUMP CANNOT RUN WITHOUT POWER. SECTION 2 OF THE IFU ALSO CONTAINS A SECTION TITLED "CONNECTING THE DRIVELINE TO THE SYSTEM CONTROLLER", WHICH PROVIDES INSTRUCTIONS ON CONNECTING/DISCONNECTING THE DRIVELINE TO/FROM THE SYSTEM CONTROLLER AND MAKING SURE THAT IT IS FULLY AND PROPERLY INSERTED INTO THE SYSTEM CONTROLLER SOCKET. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER "EDUCATING AND TRAINING PATIENTS, FAMILIES, AND CAREGIVERS"), EXPLAINS THAT DURING THE PATIENT SELECTION, PREIMPLANT, AND POSTOPERATIVE PERIOD, THE PATIENT MUST RECEIVE INSTRUCTIONS REGARDING THE OPERATION AND CARE OF EVERY SYSTEM COMPONENT (INCLUDING THE DRIVELINE AND WHAT TO DO IN AN EMERGENCY). SECTION 6 OF THE IFU ALSO LISTS NEUROLOGIC DYSFUNCTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. SECTION 7 OF THE IFU, ¿ALARMS AND TROUBLESHOOTING¿, AND SECTION 5 OF THE PATIENT HANDBOOK, ¿ALARMS AND TROUBLESHOOTING¿, DESCRIBE ALL ALARM CONDITIONS AS WELL AS THE APPROPRIATE ACTIONS ASSOCIATED WITH THEM. THE PATIENT HANDBOOK ALSO CONTAINS A SECTION ON HANDLING EMERGENCIES AND FURTHER INSTRUCTS THE USER TO CALL THEIR HOSPITAL CONTACT IF THE USER THINKS, FOR ANY REASON, ANY PORTION OF THE EQUIPMENT IS NOT FUNCTIONING AS USUAL, IS BROKEN, OR THE USER IS UNCOMFORTABLE WITH THE OPERATION OF THE EQUIPMENT. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: B5 NARRATIVE INFORMATION. H1 - TYPE OF REPORTABLE EVENT UPDATED. H6 - ADVERSE EVENT PROBLEM CODES UPDATED. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21509635&pc=DSQ | 2916596-2025-01177 | 2024-09-16 04:00 | Death | THORATEC CORPORATION | 2025-03-03 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY FROM UNKNOWN CAUSES ON (B)(6) 2024. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT'S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATION. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) IFU IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION,¿ LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21509882&pc=DSQ | 3007042319-2025-00585 | 2024-01-01 05:00 | Death | HEARTWARE, INC. | 2025-03-03 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P100047 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING CONCOMITANT AORTIC VALVE REPLACEMENT IN PATIENTS WITH PREOPERATIVE MILD-TO-MODERATE AORTIC REGURGITATION UNDERGOING LEFT VENTRICULAR ASSIST DEVICE (VAD) IMPLANTATION. THE AUTHORS DESCRIBED PATIENT DEATHS; HOWEVER, THE CAUSES OF DEATH WERE UNKNOWN. THERE WERE PATIENTS WHO EXPERIENCED MAJOR DEVICE-RELATED INFECTIONS, PUMP THROMBOSES, NEUROLOGICAL DYSFUNCTIONS, HEMORRHAGIC STROKES, ISCHEMIC STROKES, RIGHT HEART FAILURE, AND BLEEDING EVENTS WHICH REQUIRED SURGICAL INTERVENTION. SOME PATIENTS WERE REINTUBATED OR REQUIRED DIALYSIS. THERE WERE ALSO UNKNOWN/UNSPECIFIED DEVICE MALFUNCTIONS. THE STATUS OF THE DEVICES IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: ### THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/59 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: IMPACT OF CONCOMITANT AORTIC VALVE REPLACEMENT IN PATIENTS WITH MILD-TO- MODERATE AORTIC VALVE REGURGITATION UNDERGOING LEFT VENTRICULAR ASSIST DEVICE IMPLANTATION: EUROMACS ANALYSIS. ARTIFICIAL ORGANS. 2024; 00:1¿14. DOI: 10.1111/AOR.14926 D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE PRODUCT WAS DISTRIBUTED OUTSIDE OF THE UNITED STATES AND D4 UDI IS THEREFORE BLANK. ### MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21509901&pc=DSQ | 2916596-2025-01077 | 2022-01-01 05:00 | Death | THORATEC CORPORATION | 2025-03-03 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Sepsis; Multiple Organ Failure; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE, "DRIVELINE INFECTION IN PATIENTS WITH LEFT VENTRICULAR ASSIST DEVICES IMPLANTED AS DESTINATION THERAPY" THAT ONE OF THE MOST COMMON COMPLICATIONS FOLLOWING IMPLANTATION WITH THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) WAS DRIVELINE INFECTIONS (DLIS). 120 PATIENTS IMPLANTED BETWEEN 2015 AND 2022 WITH THE HEARTWARE (47 PATIENTS) AND HEARTMATE 3 (73 PATIENTS) LVADS WERE EXAMINED FOR THIS STUDY. MORE THAN HALF OF THE PATIENTS (N=52; 58%) EXPERIENCED A DRIVELINE INFECTION. THERE WERE 25 PATIENTS THAT DIED DURING THE FOLLOW-UP OBSERVATION PERIOD (39%): 5 DUE TO SEPSIS AFTER DIAGNOSIS OF A BLOODSTREAM INFECTION, AND 10 DUE TO MULTISYSTEM ORGAN FAILURE CONNECTED WITH AN INFECTION. IN ALL CASES, THE PATIENTS WITH DRIVELINE INFECTIONS WERE ADMITTED TO THE HOSPITAL AND TREATED WITH TARGETED ANTIBIOTICS AND ADVANCED WOUND CLEANING. THERE WERE A FEW CASES THAT REQUIRED EXTENSIVE WOUND DEBRIDEMENT. THE MEDIAN STAY IN THE INTENSIVE CARE UNIT WAS 6 DAYS AND THE MEDIAN TIME OF EXTUBATION AFTER THE PROCEDURE WAS 24 HOURS. THE CAUSES OF DEATH IN THE COHORT WERE DUE TO CEREBROVASCULAR EVENTS (N=14; 40%) AND RIGHT HEART FAILURE (N=5; 14%). Manufacturer Narrative: SECTIONS A AND D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. DEPARTMENT OF CARDIOVASCULAR SURGERY AND TRANSPLANTOLOGY, JAGIELLONIAN UNIVERSITY, MEDICAL COLLEGE, JOHN PAUL II HOSPITAL, CRACOW, POLAND. WASILEWSKI, G., WISNIOWSKA-SMIALEK, S., GÓRKIEWICZ-KOT, I., MILANIAK, I., KALETA, M., HYMCZAK, H., WASILEWSKA, A., & WIERZBICKI, K. (2024). DRIVELINE INFECTION IN PATIENTS WITH LEFT VENTRICULAR ASSIST DEVICES IMPLANTED AS DESTINATION THERAPY. TRANSPLANTATION PROCEEDINGS, 56(4), 860¿863. HTTPS://DOI.ORG/10.1016/J.TRANSPROCEED.2024.03.029. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: SECTION B3: DATE OF EVENT HAS BEEN ENTERED AS 01JAN2022 AS PATIENTS WERE IMPLANTED BETWEEN 2015 AND 2022. SECTION D4: MODEL NUMBER AND CATALOG NUMBER CORRECTED. SECTION G2: REPORT SOURCE CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEMS (LVAS) AND THE REPORTED EVENTS COULD NOT CONCLUSIVELY BE ESTABLISHED THROUGH THIS EVALUATION. THE SERIAL NUMBERS OF THE HEARTMATE 3 LEFT VENTRICULAR SYSTEMS IN USE WERE NOT PROVIDED. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C, AND THE HEARTMATE 3 PATIENT HANDBOOK, REV. D, ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (INCLUDING RIGHT HEART FAILURE), BLEEDING, STROKE, SEPSIS, AND DEATH AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6, ¿PATIENT CARE AND MANAGEMENT¿, UNDER ¿ANTICOAGULATION¿, PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INR VALUES, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. SECTION 6 "PATIENT CARE AND MANAGEMENT" (UNDER "RIGHT HEART FAILURE") DISCUSSES THE POTENTIAL DEVELOPMENT OF RIGHT HEART FAILURE FOLLOWING IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS, INCLUDING INOTROPES AND RIGHT VENTRICULAR ASSIST DEVICE (RVAD) PLACEMENT. SECTION 6 (UNDER CAUTION!) STATES: "RIGHT HEART FAILURE CAN OCCUR FOLLOWING IMPLANTATION OF THE PUMP. RIGHT VENTRICULAR DYSFUNCTION, ESPECIALLY WHEN COMBINED WITH ELEVATED PULMONARY VASCULAR RESISTANCE, MAY LIMIT THE EFFECTIVENESS OF THE LVAS DUE TO REDUCED FILLING OF THE PUMP." BOTH THE IFU AND PATIENT HANDBOOK CONTAIN VARIOUS SECTIONS REGARDING INFECTION AND HOW TO PREVENT IT. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21511434&pc=NPT | 2025587-2025-01604 | 2025-02-26 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-03 05:00 | NPT | EVOLUT PRO PLUS DCS | Failure to Advance; Patient Device Interaction Problem | Cardiac Arrest; Chest Pain; Low Blood Pressure/ Hypotension; Vascular Dissection | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER HEART VALVE PROCEDURE, THE ANATOMY OF THE THORACIC AORTA WAS NOTED TO BE VERY TORTUOUS. PRE-IMPLANT BALLOON DILATION OF THE VALVE WAS PERFORMED, BUT WHEN THE DELIVERY CATHETER SYSTEM (DCS) WAS INSERTED, IT WAS NOT ABLE TO CROSS THE AORTIC ARCH DUE TO THE DIFFICULTY AT THE END OF THE TRANSVERSE AORTA AND THE TORTUOSITY, WHICH DID NOT ALLOW THE TRANSMISSION OF TENSION. TWO HIGH-SUPPORT GUIDEWIRES WERE PLACED, ONE INSIDE THE PIGTAIL CATHETER AND THE OTHER IN THE DCS, BUT THE DCS WAS STILL NOT ABLE TO ADVANCE TO THE ANNULUS. THE PATIENT EXPERIENCED CHEST PAIN AND A DROP IN BLOOD PRESSURE. SUBSEQUENTLY, THE PATIENT DIED. IT WAS NOTED THAT THERE WAS A POSSIBLE DISSECTION OF THE DESCENDING THORACIC AORTA, WHICH WAS IDENTIFIED AS THE CAUSE OF DEATH.Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER HEART VALVE PROCEDURE, THE ANATOMY OF THE THORACIC AORTA WAS NOTED TO BE VERY TORTUOUS. PRE-IMPLANT BALLOON DILATION OF THE VALVE WAS PERFORMED, BUT WHEN THE DELIVERY CATHETER SYSTEM (DCS) WAS INSERTED, IT WAS NOT ABLE TO CROSS THE AORTIC ARCH DUE TO THE DIFFICULTY AT THE END OF THE TRANSVERSE AORTA AND THE TORTUOSITY, WHICH DID NOT ALLOW THE TRANSMISSION OF TENSION. TWO HIGH-SUPPORT GUIDEWIRES WERE PLACED, ONE INSIDE THE PIGTAIL CATHETER AND THE OTHER IN THE DCS, BUT THE DCS WAS STILL NOT ABLE TO ADVANCE TO THE ANNULUS. THE PATIENT EXPERIENCED CHEST PAIN AND A DROP IN BLOOD PRESSURE. SUBSEQUENTLY, THE PATIENT DIED. IT WAS NOTED THAT THERE WAS A POSSIBLE DISSECTION OF THE DESCENDING THORACIC AORTA, WHICH WAS IDENTIFIED AS THE CAUSE OF DEATH. ADDITIONAL INFORMATION WAS RECEIVED TO REPORT THE ATTEMPTED VALVE IMPLANT PROCEDURE ENDED AT 6PM WITH NO VALVE IMPLANTED. THE PATIENT WAS TRANSFERRED TO THE INTENSIVE CARE UNIT, "AS USUAL" AT THIS FACILITY, WITHOUT ANY FURTHER INCIDENCE NOTED. HOWEVER, THE NEXT MORNING, THE PHYSICIAN REPORTED THE PATIENT DIED IN THE EARLY MORNING AND INDICATED "THE AORTA MAY HAVE BEEN DISSECTED". NO FURTHER DETAILS WERE PROVIDED.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT DURING THE VALVE IMPLANT, THE DIFFICULTY CROSSING THE AORTIC ARCH "TRAUMATIZED THE AORTA." CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED. PER THE PHYSICIAN, THE CAUSE OF DEATH WAS CARDIORESPIRATORY ARREST.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT ACCORDING TO THE PHYSICIAN, THE TENSION GENERATED BY THE INABILITY TO ADVANCE THE DCS CONTRIBUTED TO THE PATIENT'S DESCENT, AND THE CAUSE OF THE DISSECTION WAS THE STRAIN OF THE LUNGS WHILE ASCENDING THROUGH THE TORTUOSITY. Manufacturer Narrative: UPDATED DATA: H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SECTION D REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: PRODUCT ID: EVPROPLUS-34; PRODUCT SERIAL NUMBER: (B)(6); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: NON-IMPLANTABLE DEVICE PRODUCT ID: L-EVPROP34; PRODUCT LOT NUMBER: 0012408654; PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: NON-IMPLANTABLE DEVICE SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5 H6 ADDITIONAL CODES MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B2. OUTCOME ATTRIBUTED TO ADVERSE EVENT "HOSPITALIZATION" WAS ADDED. B5. A SECOND PARAGRAPH WAS ADDED. H3. DEVICE NOT EVALUATED AS THE SUSPECT DEVICE WILL NOT BE RETURNED TO MEDTRONIC FOR ANALYSIS. H6. EVAL CODE METHOD, EVAL CODE RESULT, AND EVAL CODE CONCLUSION WERE UPDATED. ADDITIONAL CODES - ANNEX F "INTENSIVE CARE" WAS ADDED. CORRECTED DATA: B2. OUTCOME ATTRIBUTED TO ADVERSE EVENT "LIFE THREATENING" WAS ADDED AS IT WAS INADVERTENTLY OMITTED FROM THE INITIAL MEDWATCH REPORT. ADDITIONAL CODES - ANNEX F LIFE THREATENING ILLNESS OR INJURY AND ABSENCE OF TREATMENT WERE ADDED AS THEY WERE INADVERTENTLY OMITTED FROM THE INITIAL MEDWATCH REPORT. PRODUCT ANALYSIS: THE SUSPECT DELIVERY CATHETER SYSTEM (DCS) WAS NOT RETURNED TO MEDTRONIC FOR ANALYSIS AND THE PATIENT¿S EXECUTIVE SUMMARY WAS NOT PROVIDED FOR ANATOMICAL REVIEW. HOWEVER, INTRA-PROCEDURAL FLUOROSCOPIC IMAGES WERE SUBMITTED FOR REVIEW WHICH SHOWED EVIDENCE OF TORTUOSITY IN THE AORTA. FLUOROSCOPIC INSPECTION OF THE VALVE LOAD WAS PERFORMED AND CONFIRMED A GOOD LOAD. A PRE-IMPLANT BALLOON AORTIC VALVULOPLASTY WAS PERFORMED PRIOR TO ADVANCEMENT OF THE DCS. THERE WAS EVIDENCE OF DIFFICULTY ADVANCING THE DCS ACROSS AORTIC ARCH. PER MEDTRONIC BEST PRACTICES, PHYSICIANS SHOULD CONSIDER THAT COMPLEX ANATOMICAL CONFIGURATIONS INCREASE THE RISK OF VASCULAR TRAUMA. THESE INCLUDE BUT ARE NOT LIMITED TO MULTIPLANAR CURVATURE OF THE AORTA, ACUTE ANGULATION OF THE AORTA, AND SEVERE CALCIFICATION. PER THE DEVICE INSTRUCTIONS FOR USE, IF TWO OR MORE OF THESE FACTORS ARE PRESENT, CONSIDER AN ALTERNATIVE ACCESS ROUTE TO MITIGATE POTENTIAL FOR VASCULAR COMPLICATIONS. CONCLUSION: THE INVESTIGATION REMAINS IN PROGRESS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21511474&pc=DTB | 2124215-2025-12944 | 2025-02-06 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-03-03 05:00 | DTB | FINELINE II EZ STEROX | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Cardiac Perforation; Pericardial Effusion | P960004 |  | 1 | Event Description: IT WAS REPORTED THAT THIS CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) SYSTEM WAS BEING EXPLANTED DUE TO THE PATIENT'S INFECTION. DURING THE EXTRACTION PROCEDURE, THE NON-BOSTON SCIENTIFIC RIGHT VENTRICULAR (RV) LEAD WAS REMOVED FIRST. THIS RIGHT ATRIAL (RA) LEAD AND THE NON-BOSTON SCIENTIFIC LEFT VENTRICULAR (LV) LEAD CAME OUT TOGETHER. AFTER THE LV LEAD WAS REMOVED, AN EFFUSION WAS FOUND. A PERICARDIOCENTESIS WAS PERFORMED, THEN THE CHEST WAS OPENED. A CARDIAC TEAR HAD OCCURRED NEAR THE CORONARY SINUS AND THE PATIENT DIED ON THE OPERATING TABLE. NO EXPLANTED PRODUCTS WERE EXPECTED FOR RETURN. Manufacturer Narrative: IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21511686&pc=OZP | 2032227-2025-130859 | 2025-01-29 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-03-03 05:00 | OZP | MINIMED 780G US SYSTEM BLE CONNECT 3.0 MG/DL | Adverse Event Without Identified Device or Use Problem | Hypoglycemia; Heart Failure/Congestive Heart Failure | P160017 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY AT HOME ON (B)(6) 2025 DUE TO HEART FAILURE AND COMPLICATIONS ON DIABETES. CUSTOMER HAD HYPOGLYCEMIA WITH A BLOOD GLUCOSE VALUE OF LESS THAN 50 MG/DL AT THE TIME OF THE EVENT. THE LAST KNOWN PRODUCT(S) FOR THE CUSTOMER IS MMT-1884, UNOMEDICAL, UNK_RESERVOIR. INFORMATION WAS OBTAINED FROM THE REPORTING PERSON AS PER THE DECEASED REPORTING TROUBLESHOOTING GUIDELINES. THE CUSTOMER HAS WORN THE PUMP AT THE TIME OF PASSING AND IT WAS UNKNOWN WHETHER THE AUTO MODE FEATURE WAS ACTIVE AT THE TIME OF THE EVENT. NO PRODUCT RETURN IS REQUIRED FOR MMT-1884. NO PRODUCT RETURN IS REQUIRED FOR UNOMEDICAL. NO PRODUCT RETURN IS REQUIRED FOR UNK_RESERVOIR. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21500992&pc=DSQ | 2916596-2025-01165 | 2021-01-01 05:00 | Death | THORATEC CORPORATION | 2025-03-02 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE ¿ESCALATING SURGICAL TREATMENT FOR LEFT VENTRICULAR ASSIST DEVICE INFECTION AND EXPECTED MORTALITY: CLINICAL RISK PREDICTION SCORE¿ THAT THE HEARTMATE 3 (HM3) MAY BE ASSOCIATED WITH RESPIRATORY FAILURE, BLEEDING, DRIVELINE INFECTION, PUMP POCKET INFECTION, BACTERIAL INFECTION, HEMODYNAMIC INSTABILITY, PUMP EXCHANGE, AND DEATH. THIS WAS A SINGLE-CENTER, RETROSPECTIVE, OBSERVATIONAL COHORT STUDY THAT INCLUDED CONSECUTIVE ADULT PATIENTS WHO UNDERWENT LEFT VENTRICULAR ASSIST DEVICE (LVAD) PLACEMENT AT A LARGE, TERTIARY ACADEMIC CENTER IN THE MIDWESTERN US BETWEEN 2011 AND 2021. PATIENTS WHO WERE UNDERGOING PUMP EXCHANGE WERE EXCLUDED. A TOTAL OF 760 PATIENTS UNDERGOING PRIMARY LVAD IMPLANT WERE IDENTIFIED, OF WHOM 255 (34%) DEVELOPED A LVAD-SPECIFIC INFECTION (VSI). OF THE 255 VSI PATIENTS, 130 HAD A HEARTMATE II, 71 HAD A HEARTWARE, AND 54 HAD A HM3. 16 PATIENTS HAD EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) AFTER LVAD. 28 PATIENTS HAD BLEEDING THAT REQUIRED REOPERATION AFTER LVAD. THE MEDIAN TIME TO FIRST VSI WAS 1.00 YEARS (0.26¿3.52), AND THERE WERE 24 VSIS PER 100 PATIENT-YEARS OF LVAD SUPPORT. THE MEDIAN AGE AT VSI WAS 56 YEARS (44 TO 64), 189 (74%) PATIENTS WERE MALE, 176 (69%) WERE WHITE, AND 132 (52%) WERE INITIATED AS DESTINATION THERAPY. VSIS WERE LOCALIZED TO THE EXTRA-ABDOMINAL DRIVELINE IN 188 (73%) PATIENTS, WHEREAS 67 (27%) INVOLVED THE LVAD PUMP POCKET AND OUTFLOW CANNULA. GRAM-POSITIVE SPECIES ACCOUNTED FOR 198 (78%) OF ISOLATED ORGANISMS AND 61 (24%) WERE POLYMICROBIAL. 36 (14%) PATIENTS HAD CONCURRENT POSITIVE BLOOD CULTURES AND 198 (78%) WERE LOCALIZED ON IMAGING. THE MOST HIGHLY REPRESENTED ORGANISMS INCLUDED 79 (31%) METHICILLIN-SENSITIVE S. AUREUS (MSSA), 47 (18%) CORYNEBACTERIUM SPECIES, 38 (15%) METHICILLIN-RESISTANT STAPHYLOCOCCUS AUREUS (MRSA), AND 37 (15%) STAPHYLOCOCCUS EPIDERMIDIS. AFTER INITIAL VSI, THE MEDIAN SURVIVAL WAS 2.40 YEARS (1.16¿6.79), AND THE MEDIAN FOLLOW-UP WAS 6.11 YEARS (3.30¿8.08). MANAGEMENT STRATEGIES FOR VSI INCLUDED 91 (36%) PATIENTS WITH MEDICAL THERAPY ALONE, 134 (53%) WITH SURGICAL DEBRIDEMENT, AND 30 (12%) WITH SURGICAL DEBRIDEMENT WITH SUBSEQUENT FLAP RECONSTRUCTION. ONE YEAR AFTER VSI, 12 (13%) MEDICALLY MANAGED, 20 (15%) SURGICALLY DEBRIDED, AND 5 (16%) FLAP RECONSTRUCTION PATIENTS DIED. WITHIN 1 YEAR OF INITIAL VSI, 90 (46%) PATIENTS WERE READMITTED WITH RECURRENT INFECTION; 53 (32%) DEVELOPED A SECONDARY INFECTION WITH A NEW ORGANISM, AND 31 (16%) DEVELOPED INCREASED ANTIMICROBIAL RESISTANCE OF THE PRIMARY PATHOGEN. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. NO PRODUCT WAS RETURNED FOR EVALUATION. THE SERIAL NUMBER OR OTHER IDENTIFYING INFORMATION OF THE PRODUCT WAS NOT REPORTED AND WAS NOT ABLE TO BE DETERMINED DURING THE INVESTIGATION. THE HEARTMATE 3 LVAS IFU AND THE HEARTMATE 3 PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SECTIONS A AND D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DEVICE WAS IMPLANTED AT TIME OF EVENT. SECTION B3: DATE OF EVENT HAS BEEN ENTERED AS 01JAN2021 AS DATA WAS COLLECTED BETWEEN 2011 AND 2021. AUTHOR INFORMATION: FINNAN, M. J., CHI, D., CHIANG, S. N., VADER, J. M., CABRERA, N. L., KELLS, A. F., MASOOD, M. F., KOTKAR, K. D., & FOX, I. K. (2024). ESCALATING SURGICAL TREATMENT FOR LEFT VENTRICULAR ASSIST DEVICE INFECTION AND EXPECTED MORTALITY: CLINICAL RISK PREDICTION SCORE. JOURNAL OF THE AMERICAN COLLEGE OF SURGEONS, 239(3), 263¿275. HTTPS://DOI.ORG/10.1097/XCS.0000000000001096. DIVISIONS OF PLASTIC AND RECONSTRUCTIVE SURGERY, DEPARTMENT OF MEDICINE, WASHINGTON UNIVERSITY, SCHOOL OF MEDICINE, ST LOUIS, MO, UNITED STATES; DEPARTMENT OF SURGERY, DIVISIONS OF CARDIOLOGY, DEPARTMENT OF MEDICINE, WASHINGTON UNIVERSITY, SCHOOL OF MEDICINE, ST LOUIS, MO, UNITED STATES; INFECTIOUS DISEASES, DEPARTMENT OF MEDICINE, WASHINGTON UNIVERSITY, SCHOOL OF MEDICINE, ST LOUIS, MO, UNITED STATES; CARDIOTHORACIC SURGERY, DEPARTMENT OF MEDICINE, WASHINGTON UNIVERSITY, SCHOOL OF MEDICINE, ST LOUIS, MO, UNITED STATES THE REPORTABLE AWARE DATE IS THE DATE THE SJM NOTIFIER COMPLETED READING THE ARTICLE AND ENTERED IN THE COMPLAINT DATABASE (PER SECTION 6.3.4 OF 90979616).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21501323&pc=MIH | 9680654-2025-00007 | 2025-02-17 05:00 | Death | WILLIAM A. COOK AUSTRALIA, PTY LTD | 2025-03-02 05:00 | MIH | ZENITH FENESTRATED AAA ENDOVASCULAR GRAFT PROXIMAL BODY | Difficult to Insert; Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P020018 |  | 1 | Event Description: UNABLE TO CANNULATE THE RENAL ARTERIES THROUGH THE FENESTRATIONS. (THERE WAS SIGNIFICANT TORTUOSITY AND THROMBUS WITHIN THE VISCERAL SECTION). THE DEVICE WAS APPROPRIATELY DELIVERED AND DEPLOYED BUT THE ABILITY TO ROTATE THE DEVICE WAS MINIMAL. AFTER NUMEROUS ATTEMPTS TO CANNULATE, REPOSITION AND UTILIZING MULTIPLE TROUBLESHOOTING TECHNIQUES, THE SURGEON HAD TO PERFORM A LEFT RENAL ARTERY/EXTERNAL ILIAC BYPASS. THE RIGHT RENAL ARTERY HAD A KNOWN SUBTOTAL OSTIAL STENOSIS AND WAS NOT BYPASSED. THE ZFEN-D AND BOTH ZSLES WERE DEPLOYED WITHOUT INCIDENT." ZFEN-P-2-32-124-R, ZFEN-D-12-28-76, ZSLE-16-56-ZT, ZSLE-20-74-ZT.Event Description: UNABLE TO CANNULATE THE RENAL ARTERIES THROUGH THE FENESTRATIONS. (THERE WAS SIGNIFICANT TORTUOSITY AND THROMBUS WITHIN THE VISCERAL SECTION). THE DEVICE WAS APPROPRIATELY DELIVERED AND DEPLOYED BUT THE ABILITY TO ROTATE THE DEVICE WAS MINIMAL. AFTER NUMEROUS ATTEMPTS TO CANNULATE, REPOSITION AND UTILIZING MULTIPLE TROUBLESHOOTING TECHNIQUES, THE SURGEON HAD TO PERFORM A LEFT RENAL ARTERY/EXTERNAL ILIAC BYPASS. THE RIGHT RENAL ARTERY HAD A KNOWN SUBTOTAL OSTIAL STENOSIS AND WAS NOT BYPASSED. THE ZFEN-D AND BOTH ZSLES WERE DEPLOYED WITHOUT INCIDENT." ZFEN-P-2-32-124-R LOT NUMBER AC1176074, ZFEN-D-12-28-76 LOT NUMBER AC1176077, ZSLE-16-56-ZT LOT NUMBER 15980020, ZSLE-20-74- ZT LOT NUMBER 15727540. UPDATED DESCRIPTION OF EVENT: THE DEVICE FUNCTIONED AS EXPECTED WHILE FOLLOWING THE INSTRUCTIONS FOR USE (IFU). NO DIFFICULTY WAS EXPERIENCED DURING DEPLOYMENT. THE PROXIMAL AND DISTAL FENESTRATED DEVICES WERE DEPLOYED, AS WELL AS BILATERAL ZSLE GRAFTS WITHOUT INCIDENT. DIFFICULTY WAS EXPERIENCED IN ROTATING THE DEVICE TO ALIGN THE FENESTRATIONS. THE PROXIMAL FENESTRATED DEVICE HAD BEEN ADVANCED TO THE APPROPRIATE POSITION IN THE AORTA. THE ANTERIOR/POSTERIOR MARKERS (CROSS) HAD BEEN IDENTIFIED AND POSITIONED APPROPRIATELY. THE PROXIMAL FENESTRATED DEVICE HAD BEEN UNSHEATHED, AND THE PHYSICIAN WAS ATTEMPTING TO CANNULATE THE RENAL FENESTRATIONS. HE HAD DIFFICULTY IN ROTATING THE DEVICE AWAY FROM THE AORTIC WALL AND COULD NOT GAIN ACCESS INTO THE RENAL ARTERIES. THE PHYSICIAN ATTEMPTED TO ROTATE THE DEVICE, HE ATTEMPTED TO CANNULATED WITH MULTIPLE DIFFERENT SELECTIVE CATHETERS AND WIRES, INCLUDING .035 AND .018 PLATFORMS. HE TRIED NUMEROUS SHEATH CONFIGURATIONS. THE SURGEON HAD TO PERFORM A LEFT RENAL ARTERY/EXTERNAL ILIAC BYPASS. THE RIGHT RENAL ARTERY HAD A KNOWN SUBTOTAL OSTIAL STENOSIS AND WAS NOT BYPASSED. THE INABILITY TO MINIMALLY ROTATE THE DEVICE WAS INFLUENCED BY THE PATIENT'S ANATOMY, MURAL THROMBUS AND ANGULATION. THE INFRARENAL NECK LENGTH WAS LESS THAN 4MM. THE PATIENT WAS ON HEPARIN ANTICOAGULANT DURING THE PROCEDURE. THE PATIENT EXPIRED A FEW DAYS AFTER THE PROCEDURE. NO INFORMATION ON CAUSE OF DEATH HAS BEEN SUPPLIED. Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED: THE DEVICE FUNCTIONED AS EXPECTED WHILE FOLLOWING THE INSTRUCTIONS FOR USE (IFU). NO DIFFICULTY WAS EXPERIENCED DURING DEPLOYMENT. THE PROXIMAL AND DISTAL FENESTRATED DEVICES WERE DEPLOYED, AS WELL AS BILATERAL ZSLE GRAFTS WITHOUT INCIDENT. DIFFICULTY WAS EXPERIENCED IN ROTATING THE DEVICE TO ALIGN THE FENESTRATIONS. AT THE TIME OF THE DIFFICULTY WAS EXPERIENCED, THE PROXIMAL FENESTRATED DEVICE HAD BEEN ADVANCED TO THE APPROPRIATE POSITION IN THE AORTA. THE ANTERIOR/POSTERIOR MARKERS (CROSS) HAD BEEN IDENTIFIED AND POSITIONED APPROPRIATELY. THE DEVICE HAD BEEN UNSHEATHED. THE DISTAL END OF THE DEVICE HAD BEEN CANNULATED AND CONFIRMED TO BE WITHIN THE DEVICE. THE PHYSICIAN WAS ATTEMPTING TO CANNULATE THE RENAL FENESTRATIONS. HE HAD DIFFICULTY IN ROTATING THE DEVICE AWAY FROM THE AORTIC WALL AND COULD NOT GAIN ACCESS INTO THE RENAL ARTERIES. THE PHYSICIAN ATTEMPTED TO ROTATE THE DEVICE, HE ATTEMPTED TO CANNULATED WITH MULTIPLE DIFFERENT SELECTIVE CATHETERS AND WIRES, INCLUDING .035 AND .018 PLATFORMS. HE TRIED NUMEROUS SHEATH CONFIGURATIONS. THE SURGEON HAD TO PERFORM A LEFT RENAL ARTERY/EXTERNAL ILIAC BYPASS. THE INABILITY TO MINIMALLY ROTATE THE DEVICE WAS INFLUENCED BY THE PATIENT'S ANATOMY; MURAL THROMBUS AND ANGULATION. DILATORS WERE NOT USED IN THE ACCESS VESSELS. THE INFRARENAL NECK LENGTH WAS LESS THAN 4MM. THE PATIENT WAS ON HEPARIN ANTICOAGULANT DURING THE PROCEDURE. THE PATIENT EXPIRED A FEW DAYS AFTER THE PROCEDURE. THE PHYSICIAN THINKS THE INABILITY TO ROTATE THE DEVICE WAS DUE TO MURAL THROMBUS AND ANGULATION. THE GRAFT WAS NOT RETURNED FOR EVALUATION. HOWEVER, IMAGES WERE PROVIDED AND REVIEWED BY THE MEDICAL DIRECTOR WHO STATED "I¿VE REVIEWED THIS CASE, ALONG WITH THE IMAGING WE HAVE, WHICH IS ONLY THE PRE-OP IMAGING. THE SITE WAS UNABLE TO CANNULATE EITHER OF THE RENAL ARTERIES AND ENDED UP HAVING TO DO AN OPEN SURGICAL BYPASS FROM THE L. EXTERNAL ILIAC ARTERY TO THE L. RENAL ARTERY. THE R. RENAL ARTERY WAS SACRIFICED DUE TO PRE-EXISTING TIGHT STENOSIS OF ITS OSTIUM (ORIFICE) THAT MEANT IT WAS ALMOST OCCLUDED PRIOR TO THE PROCEDURE. THE PRE-OP CT SCAN CONFIRMS THE COMMENTS MADE BY THE SITE, THAT THE MAIN PROBLEM WAS THE RENO-VISCERAL SEGMENT OF THE ANEURYSM NECK WAS VERY TORTUOUS AND HAD THROMBUS WITHIN IT. THIS PRE-EXISTING DISEASE, ALONG WITH THE ALMOST OCCLUDED R. RENAL ARTERY OSTIUM, MEANT THAT THEY WERE UNABLE TO CANNULATE EITHER OF THE RENAL ARTERIES. SUMMARY: NOT A FAULT OR FAILURE OF ANY OF THE ENDOGRAFT DEVICES. PROBLEMS DUE TO PRE-EXISTING DISEASE (TORTUOSITY, THROMBUS, AND PRE-OCCLUSIVE R. RENAL ARTERY OSTIUM)." REVIEW OF THE DEVICE HISTORY RECORD FOR LOT NUMBER AC1176074 FOUND THE WORK ORDER APPEARED COMPLETE AND THE QUALITY CONTROL INSPECTION RECORD WAS VERIFIED TO ENSURE THAT THE DEVICE PASSED INSPECTION. THERE WERE NO RELATED TEMPORARY DEVIATIONS IN PLACE AT THE TIME OF MANUFACTURING. UPON REVIEWING THE PHOTOS TAKEN OF THE COMPLAINT DEVICE DURING MANUFACTURE, IT APPEARS THAT THE DEVICE WAS MANUFACTURED TO SPECIFICATION. REVIEW OF THE INSTRUCTIONS FOR USE (IFU) SUPPLIED WITH THE DEVICE FOR GENERAL INFORMATION FOUND IT TO CONTAIN APPROPRIATE WARNINGS, PRECAUTIONS, AND INSTRUCTIONS TO THE USER, INCLUDING: 4. WARNINGS AND PRECAUTIONS 4.1 GENERAL USE INFORMATION AFTER ENDOVASCULAR GRAFT PLACEMENT, PATIENTS SHOULD BE REGULARLY MONITORED FOR PERIGRAFT FLOW, ANEURYSM GROWTH, PATENCY OF VESSELS ACCOMMODATED BY A FENESTRATION/SCALLOP, OR CHANGES IN THE STRUCTURE OR POSITION OF THE ENDOVASCULAR GRAFT. AT A MINIMUM, ANNUAL IMAGING IS RECOMMENDED, INCLUDING: 1) ABDOMINAL RADIOGRAPHS TO EXAMINE DEVICE INTEGRITY (SEPARATION BETWEEN COMPONENTS, STENT FRACTURE OR BARB SEPARATION) AND 2) CONTRAST AND NON-CONTRAST CT TO EXAMINE ANEURYSM CHANGES, PERIGRAFT FLOW, PATENCY, TORTUOSITY AND PROGRESSIVE DISEASE. IF RENAL COMPLICATIONS OR OTHER FACTORS PRECLUDE THE USE OF IMAGE CONTRAST MEDIA, ABDOMINAL RADIOGRAPHS AND DUPLEX ULTRASOUND MAY PROVIDE SIMILAR INFORMATION. 4.2 PATIENT SELECTION, TREATMENT AND FOLLOW-UP KEY ANATOMIC ELEMENTS THAT MAY AFFECT SUCCESSFUL EXCLUSION OF THE ANEURYSM INCLUDE SEVERE PROXIMAL NECK ANGULATION (> 45 DEGREES FOR INFRARENAL NECK TO AXIS OF AAA OR > 45 DEGREES FOR SUPRARENAL NECK RELATIVE TO THE IMMEDIATE INFRARENAL NECK); SHORT PROXIMAL AORTIC NECK (<4 MM); GREATER THAN 10% INCREASE IN DIAMETER OVER 15 MM OF PROXIMAL AORTIC NECK LENGTH; AND CIRCUMFERENTIAL THROMBUS AND/OR CALCIFICATION AT THE ARTERIAL IMPLANTATION SITES, SPECIFICALLY THE PROXIMAL AORTIC NECK AND DISTAL ILIAC ARTERY INTERFACE. IRREGULAR CALCIFICATION AND/OR PLAQUE MAY COMPROMISE THE FIXATION AND SEALING OF THE IMPLANTATION SITES. NECKS EXHIBITING THESE KEY ANATOMIC ELEMENTS MAY BE MORE CONDUCIVE TO GRAFT MIGRATION. 5. ADVERSE EVENTS POTENTIAL ADVERSE EVENTS THAT MAY OCCUR AND/OR REQUIRE INTERVENTION INCLUDE BUT ARE NOT LIMITED TO: DEATH 7. PATIENT SELECTION AND TREATMENT 7.1 INDIVIDUALIZATION OF TREATMENT EACH PATIENT MUST BE EVALUATED ON AN INDIVIDUAL BASIS TO DETERMINE THE SPECIFIC LOCATION OF THE GRAFT FENESTRATIONS (REFER TO PLANNING AND SIZING SHEET), WITH CAREFUL CONSIDERATION ALSO GIVEN TO BOTH THE POTENTIAL BENEFITS AND SPECIFIC RISKS ASSOCIATED WITH THE PROCEDURE. CONSIDERATIONS REGARDING THE USE OF THE ZENITH FENESTRATED AAA ENDOVASCULAR GRAFT (SEE WARNINGS) INCLUDE: ¿ ILIOFEMORAL ACCESS VESSEL SIZE AND MORPHOLOGY (MINIMAL THROMBUS, CALCIFICATION AND/OR TORTUOSITY) SHOULD BE COMPATIBLE WITH VASCULAR ACCESS TECHNIQUES AND ACCESSORIES OF THE DELIVERY PROFILE OF A 14 FRENCH TO 22 FRENCH VASCULAR INTRODUCER SHEATH. THERE IS NO EVIDENCE TO SUGGEST THAT THE USER DID NOT FOLLOW THE INSTRUCTIONS FOR USE. A REVIEW OF THE MANUFACTURING RECORDS AND SPECIFICATIONS DID NOT REVEAL ANY DISCREPANCIES THAT COULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INVESTIGATION CONCLUDED THAT THE COMPLAINT DEVICE WAS MANUFACTURED TO SPECIFICATION. BASED ON THE MEDICAL DIRECTOR'S ASSESSMENT, THE ROOT CAUSE FOR THE INABILITY TO CANNULATE WAS DETERMINED TO BE DUE TO THE PATIENT'S PRE-EXISTING DISEASE. AFTER CONSIDERING THIS EVENT THE RISK ASSOCIATED WITH THE USE OF THIS DEVICE IS STILL DEEMED ADEQUATE. THE APPROPRIATE PERSONNEL WILL CONTINUE TO MONITOR FOR SIMILAR COMPLAINTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21501380&pc=NVZ | 2017865-2025-15505 | 2025-02-10 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-03-02 05:00 | NVZ | ELLIPSE VR | Over-Sensing; Under-Sensing | Arrhythmia | P910023 |  | 1 | Event Description: IT WAS REPORTED THAT AT 11.28AM ON (B)(6)2025, THE PATIENT EXPERIENCED A CARDIAC ARREST EPISODE. THE SIGNAL AMPLITUDE OF THE ARRHYTHMIA WAS VERY LOW, THE DISCRIMINATION CHANNEL WAS UNDERSENSING AND AS A RESULT NON SUSTAINED RIGHT VENTRICULAR OVERSENSING (NSRVO) EPISODES WERE STORED. BY 11:41AM THE SENSE AMP CHANNEL WAS UNDERSENSING. THE LAST RECORDED EPISODE OF NSRVO WAS AT 11:55AM. AN ALERT WAS TRANSMITTED REMOTELY VIA MERLIN.NET AT 2:00AM ON (B)(6)2025. THE CLINIC REACHED OUT TO THE PATIENT'S RELATIVE WHO DID A WELLNESS CHECK AND FOUND THE PATIENT TO BE DECEASED. THERE WERE NO COMPLAINTS RELATED TO THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) PRIOR TO THE EVENT AND A RECENT REMOTE TRANSMISSION SHOWED APPROPRIATE FUNCTION. TECHNICAL SERVICES REVIEW OF AVAILABLE EPISODES SURROUNDING THE EVENT (B)(6) 2025 - (B)(6)2025, REVEALED THAT ON THE VENTRICULAR TACHYCARDIA (VT) AND NSRVO EPISODES THERE APPEARED TO BE A SMALL NOISY ARTIFACT ON THE VSENSEAMP CHANNEL NEAR THE R WAVES THAT IN SOME CASES WOULD BE OVERSENSED. IT WAS DIFFICULT TO ASSESS IF THIS WAS DUE TO LEAD NOISE OR ADDITIONAL ECTOPIC OR MYOPOTENTIAL ACTIVITY FROM THE HEART. THE DEVICE APPEARED TO CLASSIFY SOME OF THAT ACTIVITY AS NOISE AND EVEN FLAGGED IT ON A FEW VT EPISODES, HOWEVER, AS PROGRAMMED, VT ZONE WAS PROGRAMMED AS A MONITOR ZONE ONLY THUS THIS WOULD HAVE NOT RESULTED IN THERAPY. OTHERWISE, THE DEVICE APPEARED TO BEHAVE AS PROGRAMMED WITHIN THE VT ZONE EPISODES. ADDITIONALLY, IT WAS OBSERVED THAT THERE WERE SOME NSRVO EPISODES WHERE IT APPEARED THE R WAVE AMPLITUDE WAS TOO SMALL TO BE SEEN, RESULTING IN UNDERSENSING. SECURESENSE ALGORITHM WAS ALSO CHANGED TO PASSIVE DUE TO UNDERSENSING. IN SUMMARY, THE DEVICE APPEARED TO HAVE REACTED AS PROGRAMMED, BUT THERE WAS SOME OVERSENSING AND UNDERSENSING OBSERVED. Manufacturer Narrative: THE PROVIDED SESSION RECORDS WERE REVIEWED BY TECHNICAL SERVICES. IT WAS OBSERVED THAT THERE WAS A SMALL NOISY ARTIFACT THAT WAS OCCASIONALLY OVERSENSED. THE DEVICE CLASSIFIED SOME OF THE ACTIVITY AS NOISE BUT THE VENTRICULAR TACHYCARDIA (VT) ZONE WAS SET AS A MONITOR ZONE ONLY AND WOULD NOT HAVE RESULTED IN THERAPY. ADDITIONALLY, NON-SUSTAINED RIGHT VENTRICULAR OVERSENSING (NSRVO) EPISODES WERE OBSERVED WHERE THE R WAVE AMPLITUDE WAS TOO SMALL TO BE SEEN ¿ RESULTING IN UNDERSENSING. THE DEVICE APPEARED TO HAVE BEHAVED PER PROGRAMMED SETTINGS. A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED, AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21501386&pc=DSQ | 2916596-2025-01222 | 2023-06-01 04:00 | Death | THORATEC CORPORATION | 2025-03-02 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE ARTICLE TITLED ¿LEFT VENTRICULAR ASSIST DEVICE IN THE PRESENCE OF SUBCUTANEOUS IMPLANTABLE CARDIOVERTER DEFIBRILLATOR: DATA FROM A MULTICENTER EXPERIENCE¿ THAT THE HEARTMATE 3 MAY BE ASSOCIATED WITH CARDIAC ARRHYTHMIAS, ELECTROMAGNETIC INTERFERENCE (EMI), KIDNEY DISEASE, HEART TRANSPLANT (HT), AND DEATH. THIS WAS A MULTICENTER RETROSPECTIVE COHORT STUDY AMONG PATIENTS WITH A PREEXISTING S-ICD WHO UNDERWENT LVAD IMPLANTATION FOR ADVANCED HF DESPITE OPTIMAL MEDICAL THERAPY FROM MARCH 2015 AND JUNE 2023. THE STUDY POPULATION INCLUDED 30 PATIENTS (N = 25 MALES, 83.3%; MEDIAN AGE: 45 [38¿52] YEARS). THE MEDIAN TIME FROM S-ICD IMPLANTATION TO LVAD DEVICE IMPLANTATION WAS 12 [6¿25] MONTHS. BASELINE CLINICAL CHARACTERISTICS OF THE STUDY POPULATION ARE REPORTED IN TABLE 1. SEVEN (23%) PATIENTS RECEIVED AN S-ICD FOR SECONDARY PREVENTION. MOST PATIENTS HAD NON-ISCHEMIC CARDIOMYOPATHY (N = 23; 76.7%) WITH A MEDIAN LEFT VENTRICULAR (LV) EJECTION FRACTION OF 15 [15¿22]. INDICATIONS FOR LVAD IMPLANTATION WERE BRIDGE TO HT (N = 15; 50%), BRIDGE TO HT CANDIDACY (N = 13; 43%) AND DESTINATION THERAPY (N = 2; 7%). IMPLANTED LVAD MODELS WERE HEARTMATE III (N =19; 63%), HEARTWARE (N =5; 17%), JARVIK 2000 (N =5; 17%) AND LEVITRONIX, PARACORPOREAL LVAD (N =1; 3%). FROM THE BASELINE PATIENT CHARACTERISTICS PROVIDED. 7 OF THE 30 PATIENTS HAD KIDNEY DISEASE, AND 2 HAD ATRIAL FIBRILLATION (AF). THE ONSET WAS NOT SPECIFIED AS BEING PRE-LVAD IMPLANTATION. EMI OVERSENSING IN AT LEAST 1 SENSING VECTOR OCCURRED IN 21 (70%) PATIENTS: 14 OF 19 (73%) WITH THE HEARTMATE III, 4 OF 5 (80%) WITH THE JARVIK 2000, 3 OF 5 (60%) WITH THE HEARTWARE (SEE FIG. 1). SIX PATIENTS (20%) EXPERIENCED A TOTAL OF 95 INAPPROPRIATE SHOCKS (IS) DUE TO EMI. SIX PATIENTS (20%) EXPERIENCED A TOTAL OF 14 APPROPRIATE SHOCKS FOR VENTRICULAR TACHYCARDIA (VT)/ VENTRICULAR FIBRILLATION (VF). NO COMPLICATIONS OCCURRED. NO S-ICD EXTRACTION BECAUSE OF NEED FOR ATP, INEFFECTIVE THERAPY, OR INFECTION WAS REPORTED. DURING FOLLOW-UP, 5 PATIENTS (16%) UNDERWENT HT AND 7 PATIENTS (23%) DIED BECAUSE OF CARDIAC DEATH DUE TO WORSENING HF. HEARTMATE 3 AND HEARTWARE HVAD OPERATE AT A LOWER ROTATIONAL SPEED, GENERATING LOW FREQUENCY EMI (30¿150 HZ) THAT IS LESS LIKELY TO BE REMOVED BY DEVICE FILTERS RESULTING IN EMI OVERSENSING AND INAPPROPRIATE THERAPY [11]. IN CONCLUSION, CONCOMITANT USE OF LVADS AND S-ICDS WAS FEASIBLE IN MOST PATIENTS. HOWEVER, THE POTENTIAL RISK OF EMI OVERSENSING, IS, AND UNDER SENSING IN THE POST-OPERATIVE PERIOD FOLLOWING LVAD IMPLANTATION SHOULD BE CONSIDERED. Manufacturer Narrative: SECTION B3: DATE OF EVENT CORRECTED. DATE OF DEATH WAS NOT PROVIDED. DATE OF EVENT HAS BEEN ENTERED AS 01JUN2023 AS PATIENTS WERE IMPLANTED FROM (B)(6) 2015 TO (B)(6) 2023. SECTION E: REPORTER INFORMATION CORRECTED. SECTION G2: REPORT SOURCE CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED PATIENT OUTCOMES COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. NO DEVICES WERE RETURNED FOR THIS EVALUATION. THE DEVICE SERIAL NUMBERS WERE NOT REPORTED AND WERE UNABLE TO BE DETERMINED DURING THIS EVALUATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: B5: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. DATE OF EVENT HAS BEEN ENTERED AS: 01JUN2023 AS PATIENTS WERE IMPLANTED FROM MAR2015 TO JUN2023. DEPARTMENT OF CARDIAC, THORACIC, VASCULAR SCIENCES AND PUBLIC HEALTH, UNIVERSITY OF PADOVA, PADUA, ITALY, DEPARTMENT OF CLINICAL ELECTROPHYSIOLOGY & CARDIAC PACING, CENTRO CARDIOLOGICO MONZINO, IRCCS, MILAN, ITALY, DEPARTMENT OF SYSTEMS MEDICINE, UNIVERSITY OF ROME TOR VERGATA, ROME, ITALY, CARDIOLOGY UNIT, LUIGI SACCO UNIVERSITY HOSPITAL, MILAN, ITALY, CARDIOLOGY UNIT, IRCCS, DEPARTMENT OF EXPERIMENTAL, DIAGNOSTIC AND SPECIALTY MEDICINE, SANT'ORSOLA HOSPITAL, UNIVERSITY OF BOLOGNA, BOLOGNA, ITALY, ARRHYTHMOLOGY AND ELECTROPHYSIOLOGY UNIT, SAN RAFFAELE HOSPITAL, IRCCS, MILAN, ITALY,CARDIOLOGY UNIT, UNIVERSITY MEDICAL CENTRE MANNHEIM, MANNHEIM, GERMANY, DEPARTMENT OF CARDIOLOGY AND ANGIOLOGY, FACULTY OF MEDICINE, HEART, CENTER FREIBURG UNIVERSITY, UNIVERSITY OF FREIBURG, GERMANY,CARDIOLOGY UNIT, FONDAZIONE IRCCS SAN GERARDO DEI TINTORI, MONZA, ITALY, DEPARTMENT OF RHYTHMOLOGY, UNIVERSITY HEART CENTER LUBECK, LUBECK, GERMANY, DEPARTMENT OF BIOMEDICAL SURGICAL AND DENTAL SCIENCES, UNIVERSITY OF MILAN, MILAN, ITALY, MONTEFIORE-EINSTEIN CENTER FOR HEART AND VASCULAR CARE, MONTEFIORE MEDICAL CENTER, ALBERT EINSTEIN COLLEGE OF MEDICINE AT MONTEFIORE HEALTH SYSTEM, BRONX, NY, USA, DIVISION OF CARDIOLOGY, DEPARTMENT OF MEDICINE, THE JOHNS HOPKINS UNIVERSITY SCHOOL OF MEDICINE, BALTIMORE, USA MIGLIORE, F., SCHIAVONE, M., PITTORRU, R., FORLEO, G. B., DE LAZZARI, M., MITACCHIONE, G., BIFFI, M., GULLETTA, S.,KUSCHYK, J., DALL¿AGLIO, P. B., ROVARIS, G., TILZ, R., MASTRO, F. R., ILICETO, S., TONDO, C., DI BIASE, L., GASPERETTI, A., TARZIA, V., & GEROSA, G. (2024). LEFT VENTRICULAR ASSIST DEVICE IN THE PRESENCE OF SUBCUTANEOUS IMPLANTABLE CARDIOVERTER DEFIBRILLATOR: DATA FROM A MULTICENTER EXPERIENCE. INTERNATIONAL JOURNAL OF CARDIOLOGY, 400, 131807. HTTPS://DOI.ORG/10.1016/J.IJCARD.2024.131807 THE REPORTABLE AWARE DATE IS THE DATE THE SJM NOTIFIER COMPLETED READING THE ARTICLE AND ENTERED IN THE COMPLAINT DATABASE (PER SECTION 6.3.4 OF (B)(4)). NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21501398&pc=DSQ | 2916596-2025-01390 | 2021-01-07 05:00 | Death | THORATEC CORPORATION | 2025-03-02 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE ¿COMPARISON OF INTRAOPERATIVE BLOOD PRODUCT USE DURING HEART TRANSPLANTATION IN PATIENTS BRIDGED WITH IMPELLA 5.5 VERSUS DURABLE LEFT VENTRICULAR ASSIST DEVICES¿ THAT HEARTMATE 3 MAY BE ASSOCIATED WITH STROKE, INFECTION, RV FAILURE, BLEEDING, PUMP THROMBOSIS, TRANSPLANT AND DEATH. INDIVIDUALS WITH DURABLE LVADS ALSO EXPERIENCED HEMATOMA, HEMOLYSIS, AND CONTROLLER FAILURE. THIS WAS A SINGLE-CENTER, RETROSPECTIVE, OBSERVATIONAL STUDY OF PATIENTS WHO UNDERWENT ORTHOTOPIC HEART TRANSPLANT OR COMBINED HEART AND KIDNEY TRANSPLANT FROM NOVEMBER 2018 TO JULY 2021. THE DATA WERE COLLECTED VIA REVIEW OF ELECTRONIC MEDICAL RECORDS. THIS STUDY AIMED TO DETERMINE IF THE INTRAOPERATIVE TRANSFUSION REQUIREMENTS DIFFER BASED ON THE MECHANICAL CIRCULATORY DEVICE USED AS A BRIDGE TO HEART TRANSPLANTATION. 24 PATIENTS WERE BRIDGED WITH IMPELLA 5.5 AND 19 PATIENTS WERE BRIDGED WITH DURABLE LVAD (11 PATIENTS HAD HEARTWARE LVADS, 4 HAD HEARTMATE 2 LVADS AND 4 HAD HEART MATE 3 LVADS.) OF THE 19 DURABLE LVAD PATIENTS IN THE STUDY, 12 WERE TRANSPLANTED AFTER A COMPLICATION SUCH AS THROMBOSIS OR DRIVELINE INFECTION. INCLUSION CRITERIA INCLUDED ALL PATIENTS WHO UNDERWENT ORTHO TOPIC HEART OR HEART/KIDNEY TRANSPLANTATION BETWEEN NOVEMBER 2018 AND JULY2021 BRIDGED TO TRANSPLANTATION WITH EITHER AN IMPELLA 5.5 OR DURABLE LVAD. PATIENTS WHO DID NOT SURVIVE TO HEART TRANSPLANTATION WERE EXCLUDED FROM THE STUDY. OUT OF THE 19 INDIVIDUALS WITH DURABLE LVAD, 1 EXPERIENCED STROKE, 7 HAD INFECTIONS, 3 EXPERIENCED RV FAILURE, 3 EXPERIENCED BLEEDING, AND 2 HAD PUMP THROMBOSIS. MINOR COMPLICATIONS AMONG THE LVAD POPULATION INCLUDED 9 CASES OF HEMATOMA, 4 CASES OF HEMOLYSIS AND 6 CASES OF CONTROLLER FAILURE. THIS STUDY HIGHLIGHTED THE SIGNIFICANT DIFFERENCES IN INTRAOPERATIVE TRANSFUSIONS BETWEEN THE IMPELLA 5.5 AND THE DURABLE LVAD, INCLUDING A REDUCTION IN CRYOPRECIPITATE, AUTOLOGOUS BLOOD TRANSFUSION, AND PLATELETS IN THOSE BRIDGED WITH THE IMPELLA 5.5. IT ALSO FOUND A TREND TOWARD A REDUCTION OF PRBC AND FFP IN THE INTRAOPERATIVE TRANSFUSION REQUIREMENTS OF PATIENTS BRIDGED TO HEART TRANSPLANT WITH AN IMPELLA 5.5 DEVICE VERSUS THOSE BRIDGED USING A DURABLE LVAD. ADDITION ALLY, PATIENTS BRIDGED WITH THE IMPELLA 5.5 HAD SHORTER CPB TIMES AND A LOWER VIS POST HEART TRANSPLANTATION. THE IMPELLA 5.5 AS A TEMPORARY MCS DEVICE MAY OFFER AN ALTERNATIVE PATHWAY COMPARED TO THE TRADITIONAL DURABLE LVAD PLACEMENT AND ASSOCIATED COMPLICATIONS. Manufacturer Narrative: D4: UDI NUMBER CORRECTED. E1: CUSTOMER SITE CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 DEVICES AND REPORTED OUTCOMES COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE HEARTMATE 3 DEVICE SERIAL NUMBERS, AS WELL AS OTHER SPECIFIC CASE/PATIENT INFORMATION, ARE NOT AVAILABLE. THE SERIAL NUMBER OR OTHER IDENTIFYING INFORMATION OF THE PRODUCT WAS NOT REPORTED AND WAS NOT ABLE TO BE DETERMINED DURING THE INVESTIGATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. THE IFU LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: B5: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DEVICE WAS IMPLANTED AT TIME OF EVENT. DATE OF EVENT HAS BEEN ENTERED AS: 01JUL2021 AS PATIENTS UNDERWENT HEART OR HEART/KIDNEY TRANSPLANTATION BETWEEN NOV2018 AND JUL2021. DEPARTMENT OF ANESTHESIOLOGY AND PERIOPERATIVE MEDICINE, MAYO CLINIC, JACKSONVILLE, FL, DEPARTMENT OF CRITICAL CARE MEDICINE, MAYO CLINIC, JACKSONVILLE, FL, DIVISION OF CARDIOVASCULAR AND THORACIC ANESTHESIOLOGY, DEPARTMENT OF ANESTHESIOLOGY AND PERIOPERATIVE MEDICINE, MAYO CLINIC, JACKSONVILLE, FL DIVISION OF TRANSPLANTATION AND ADVANCED HEART FAILURE, MAYO CLINIC, JACKSONVILLE, FL, DEPARTMENT OF CARDIOLOGY, OREGON HEALTH AND SCIENCE UNIVERSITY, PORTLAND, OR. SHAPIRO, A. B., FRITZ, A. V., KILEY, S., SHARMA, S., PATEL, P., HECKMAN, A., MARTIN, A. K., & GOSWAMI, R. (2024). COMPARISON OF INTRAOPERATIVE BLOOD PRODUCT USE DURING HEART TRANSPLANTATION IN PATIENTS BRIDGED WITH IMPELLA 5.5 VERSUS DURABLE LEFT VENTRICULAR ASSIST DEVICES. JOURNAL OF CARDIOTHORACIC AND VASCULAR ANESTHESIA, 38(11), 2567¿2575. HTTPS://DOI.ORG/10.1053/J.JVCA.2024.04.047. THE REPORTABLE AWARE DATE IS THE DATE THE SJM NOTIFIER COMPLETED READING THE ARTICLE AND ENTERED IN THE COMPLAINT DATABASE (PER SECTION 6.3.4 OF 90979616). D4: DEVICE SERIAL NUMBER AND LOT NUMBER WERE NOT PROVIDED, AND EXPIRATION DATE AND MANUFACTURE DATE (H4) CANNOT BE DETERMINED. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION. H10: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21499586&pc=DSQ | 2916596-2025-01075 | 2024-10-20 04:00 | Death | THORATEC CORPORATION | 2025-03-01 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Mechanical Problem; Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Hemorrhage/Bleeding; Thromboembolism; Thrombosis/Thrombus; Renal Impairment | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE PATIENT SAW THEIR CARDIOLOGIST IN (B)(6) AND EVER SINCE THEN THEY HAD BEEN FEELING WEAKER. THE PATIENT HAD A HEAD COMPUTED TOMOGRAPHY (CT) AND RESULTS INCLUDED NEW PUNCTATE HYPERDENSITY ALONG THE LEFT POSTCENTRAL GYRUS, FAVOR CORTICAL SUCH AS SECONDARY TO A TINY INTRINSICALLY BRIGHT LESION SUCH AS A CAVERNOMA OR OTHER TINY PARENCHYMAL HEMORRHAGE AS OPPOSED TO REFLECTING TRACE FOCAL SUBARACHNOID HEMORRHAGE. CONTRAST-ENHANCED MRI EVALUATION FOR CHARACTERIZATION WAS CONSIDERED. CT OF THE HEART WITH 4D WAS PERFORMED ON (B)(6) 2025. THERE WAS 1 CM OF FOCAL THROMBUS IN THE PROXIMAL OUTFLOW CANNULA OF THE LVAD, OF UNCERTAIN HEMODYNAMIC SIGNIFICANCE. MINIMAL PULMONARY EDEMA WAS SEEN. THE PATIENT HAD A HISTORY OF ISCHEMIC CARDIOMYOPATHY (ICM), CORONARY ARTERY DISEASE (CAD) STATUS POST CORONARY ARTERY BYPASS GRAFTING (CABG), HYPERLIPIDEMIA (HLD), HEART FAILURE WITH REDUCED EJECTION (EJECTION FRACTION (EF) LESS THAN 20%) STATUS POST LVAD ON WARFARIN, CHRONIC DRIVELINE INFECTION ON CEPHALEXIN PRESENTED TO THE EMERGENCY DEPARTMENT DUE TO RIGHT EYE PAIN, REDNESS. THE PATIENT HAD CHRONIC METHICILLIN-SUSCEPTIBLE STAPHYLOCOCCUS AUREUS (MSSA) DRIVELINE INFECTION AND WAS FEELING WEAK FOR TWO WEEKS. PATIENT ALSO HAD RIGHT OCULAR VISION LOSS OVER LAST DAY AND WAS BROUGHT IN FOR CONCERN FOR CEREBROVASCULAR ACCIDENT (CVA). PATIENT DID NOT HAVE A FEVER, BUT WHITE BLOOD CELL COUNT (WBC) WAS HIGH (FROM 9 TO 14). LACTATE DEHYDROGENASE (LDH) WAS FROM 241 TO 1397. N-TERMINAL PRO-B-TYPE NATRIURETIC PEPTIDE WAS FROM 10059 TO 24027 (CLINICALLY LOSING WEIGHT SINCE (B)(6) 2024), LIKELY DUE TO LOW OUTPUT STATE DUE TO FAILING LVAD. PATIENT HAD BROWN URINE AND CONCERN FOR LVAD THROMBUS. 4D CT HEART DONE TODAY WHICH NOTED POSSIBLE LVOD OUTFLOW TRACT 1 CM THROMBUS AND INFECTED THROMBUS. ECHOCARDIOGRAM RAMP WAS PERFORMED WITH FAILURE TO DECREASE LEFT VENTRICLE SIZE. PATIENT WOKE UP THIS MORNING AND COMPLAINED OF EYE REDNESS AND BLURRY VISION. LAST KNOWN NORMAL WAS LAST NIGHT. NO UNILATERAL WEAKNESS, NUMBNESS, TINGLING, HEADACHE, EYE PAIN, COLD, COUGH, SORE THROAT, FEVERS, CHILLS. PATIENT DENIED NAUSEA VOMITING OR DIARRHEA AND HAD BEEN TAKING MEDICATIONS AS PRESCRIBED. THERE WERE NO CHANGES TO PATIENT CONDITION OR ANTICOAGULATION STATUS THAT MAY HAVE CONTRIBUTED. INTERNATIONAL NORMALIZED RATIO (INR) ON ADMISSION WAS 2.4 AND REPEAT WAS 2.9; THE PATIENT WAS NOW ON INTRAVENOUS HEPARIN AND PLANNED FOR PERIPHERALLY INSERTED CENTRAL CATHETER (PICC) LINE. MIXED VENOUS OXYGEN SATURATION WAS CHECKED TO ADD INOTROPES TO SUPPORT CO AND ALSO GET BLOOD CULTURES TO RULE OUT BACTEREMIA. THE PATIENT WAS PLANNED FOR SURGERY AFTER ONE WEEK IN LIGHT OF RECENT CVA. LEFT VENTRICULAR ASSIST DEVICE (LVAD) THROMBUS WAS THOUGHT TO BE INFECTED AND ETIOLOGY OF PIGMENT NEPHROPATHY THAT LED TO REQUIRING CONTINUOUS RENAL REPLACEMENT THERAPY (CRRT). PATIENT EXPERIENCED RAPID DECOMPENSATION WITH WORSENING LACTATE AND CONCERN FOR ISCHEMIC BOWL FROM EMBOLI. AUTOPSY WAS NOT PERFORMED, AND THE PUMP WOULD NOT BE SENT BACK. Manufacturer Narrative: DUE TO SYSTEM LIMITATIONS THE FOLLOWING CODES WERE ADDED. EYE E08: EYE PAIN E0820, NERVOUS SYSTEM E16: MUSCLE WEAKNESS E1621. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: H6 CORRECTION: MEDICAL DEVICE PROBLEM 1384 - MECHANICAL PROBLEM ADDED. MANUFACTURER¿S INVESTIGATION CONCLUSION: THE REPORTED OUTFLOW GRAFT THROMBUS WAS UNABLE TO BE CONFIRMED THROUGH REVIEW OF THE SUBMITTED IMAGES, AND THE PRODUCT WAS NOT RETURNED FOR EVALUATION. FURTHERMORE, A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE ACCOUNT SUBMITTED THREE CT IMAGES OF THE HEART FOR REVIEW. THE IMAGES CAPTURED AN AREA OF HYPODENSITY AND NARROWING ALONG THE PROXIMAL PORTION OF THE OUTFLOW GRAFT, HOWEVER THE GRAFT'S TRAJECTORY AND ANGULATION COULD NOT BE CONCLUSIVELY DETERMINED. THE REPORTED OUTFLOW GRAFT THROMBUS WAS UNABLE TO BE CONFIRMED BASED ON THE SUBMITTED IMAGES. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C, AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK, REV. D, ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEVICE THROMBOSIS, RENAL FAILURE, THROMBOEMBOLISM (INCLUDING VENOUS AND ARTERIAL NON-CENTRAL NERVOUS SYSTEM), STROKE, BLEEDING, INFECTION (LOCAL, DRIVELINE, OR PUMP POCKET INFECTION), HEMOLYSIS, AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, LISTS THROMBOEMBOLISM AND INFECTION AS POTENTIAL LATE POSTIMPLANT COMPLICATIONS. ADDITIONALLY, THIS SECTION, UNDER ¿ANTICOAGULATION¿ PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) VALUES, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. SEVERAL SECTIONS OF THE IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING INFECTION PREVENTION, AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21499606&pc=QMG | 3015053858-2025-00043 | 2025-01-31 05:00 | Death | SHOCKWAVE MEDICAL, INC. | 2025-03-01 05:00 | QMG | SHOCKWAVE C2+ CORONARY INTRAVASCULAR LITHOTRIPSY (IVL) CATHETER | Adverse Event Without Identified Device or Use Problem | Vascular Dissection; Pericardial Effusion | P200039/S008 |  | 1 | Event Description: A SHOCKWAVE C2+ CORONARY INTRAVASCULAR LITHOTRIPSY (IVL) CATHETER WAS USED TO TREAT A LESION IN THE OSTIAL LEFT MAIN ARTERY WITH IMPELLA SUPPORT. PRIOR TO THE PROCEDURE, THE PATIENT WAS TRANSFERRED TO INTENSIVE CARE WITH A BALLOON PUMP, WHICH WAS SWITCHED TO IMPELLA BEFORE USING SHOCKWAVE. PRE-DILATATION WAS DONE WITH A 2.0 NON-COMPLIANT BALLOON FOLLOWED BY ROTABLATOR (ROTA) DUE TO HEAVY CALCIFICATION IN THE LEFT MAIN ARTERY. AFTER A FEW PASSES WITH ROTA, A 3.5MM SHOCKWAVE IVL BALLOON WAS USED, DELIVERING THREE CYCLES AT 2 AND 4 ATMOSPHERES (ATM). DURING THE PROCEDURE, INTRAOPERATIVE IMAGING SHOWED A DISSECTION, WHICH WAS TREATED WITH THE PLACEMENT OF A COVERED STENT. THE PATIENT DID NOT SHOW ANY CLINICAL SYMPTOMS, HOWEVER, FURTHER IMAGING SHOWED THAT THE DISSECTION WAS GETTING LARGER. THE PATIENT DEVELOPED PERICARDIAL EFFUSION, AND DESPITE PERICARDIOCENTESIS, THE PATIENT EXPIRED ON THE TABLE. NO FURTHER INFORMATION WAS PROVIDED. THE PHYSICIAN WAS UNSURE IF THE DISSECTION WAS CAUSED BY THE ROTA OR SHOCKWAVE. THE SHOCKWAVE DEVICE WAS USED ACCORDING TO THE INSTRUCTIONS FOR USE (IFU), WITH NO MALFUNCTION REPORTED WITH THE SHOCKWAVE CATHETER OR GENERATOR. Manufacturer Narrative: THE DEVICE REFERENCED IN THIS REPORT WAS NOT RETURNED TO SHOCKWAVE MEDICAL, INC. THEREFORE, A PHYSICAL EXAMINATION COULD NOT BE PERFORMED. BASED ON THE REPORTED INFORMATION, THERE WAS A DISSECTION AND DEVELOPMENT OF PERICARDIAL EFFUSION, WHICH MAY HAVE CONTRIBUTED TO THE SUBSEQUENT DEATH; HOWEVER, THIS COULD NOT BE DEFINITIVELY DETERMINED BASED ON THE INFORMATION AVAILABLE. THERE WAS NO IVL MALFUNCTION REPORTED. SHOCKWAVE MEDICAL HAS CONTROLS IN PLACE TO ENSURE DEVICES ARE BUILT TO APPROVED PROCEDURES AND MEET LOT RELEASE ACCEPTANCE CRITERIA PRIOR TO DISTRIBUTION. A REVIEW OF THE MANUFACTURING AND TEST DOCUMENTATION FOR THE SUBJECT LOT DOES NOT REVEAL ANY ISSUES WITH THE MANUFACTURING OF THE DEVICE. THE DEVICE PASSED ALL OF SHOCKWAVE MEDICAL, INC.'S ACCEPTANCE CRITERIA PRIOR TO SHIPPING. Manufacturer Narrative: INVESTIGATION EFFORTS ARE ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21499636&pc=MKJ | 3030677-2025-000666 | 2025-02-11 05:00 | Death | PHILIPS GOLDWAY (SHENZHEN) INDUSTRIAL INC. | 2025-03-01 05:00 | MKJ | EFFICIA DFM100 | Incorrect Measurement | Unspecified Heart Problem; Insufficient Information | K110825 |  | 1 | Event Description: PHILIPS RECEIVED A COMPLAINT ON THE EFFICIA DFM100 DEFIBRILLATOR MONITOR INDICATING THE CUSTOMER WAS USING ECG ON A DECEASED PATIENT AND DOTTED LINES AND JUMPING WAVE PATTERNS APPEARED ON THE SCREEN.Event Description: PHILIPS RECEIVED A COMPLAINT ON THE DFM100 DEVICE INDICATING THAT THE USER WAS USING THE DEVICE TO MONITOR PATIENT EKG, AND DOTTED LINES AND JUMPING WAVE PATTERNS APPEARED ON THE SCREEN. THE PATIENT SUFFERED AN OUT-OF-HOSPITAL CARDIAC ARREST AND ARRIVED AT THE HOSPITAL WITHOUT PULSE. THE HOSPITAL STAFF ATTEMPTED TO VERIFY THE PATIENT'S CONDITION AND ATTACHED THE DEVICE TO THE PATIENT. THE HOSPITAL STAFF NOTICED DOTTED LINES AND WAVES ON THE SCREEN AND SWITCHED TO ANOTHER DEVICE. NO SHOCK WAS DELIVERED TO THE PATIENT. THE EFFICIA DFM100 DEFIBRILLATOR, MODEL# 866199, IS SUBSTANTIALLY SIMILAR TO THE HEARTSTART XL+ DEFIBRILLATOR (MODEL# 861290) AND WILL BE REPORTED IN THE UNITED STATES UNDER DEVICE MODEL# 861290. Manufacturer Narrative: THIS REPORT IS BASED ON INFORMATION PROVIDED BY PHILIPS FIELD SERVICE ENGINEER (FSE) AND HAS BEEN INVESTIGATED BY THE PHILIPS COMPLAINT HANDLING TEAM. THE REPORTS FROM THE CUSTOMER WERE CONFLICTING AS TO THE DATE THE EVENT OCCURRED. ALTHOUGH THE CUSTOMER INITIALLY REPORTED THAT THE EVENT OCCURRED ON (B)(6) 2025, THE PATIENT EVENT FILE (PEF) PROVIDED BY THE CUSTOMER WAS FROM (B)(6) 2025. PHILIPS ATTEMPTED TO CONFIRM THE DATE THE CUSTOMER¿S REPORTED ISSUE OCCURRED, BUT NO CONFIRMATION WAS RECEIVED. A PHILIPS CLINICIAN REVIEWED THE PEF PROVIDED FROM (B)(6) 2025 AND DETERMINED THAT THE DEVICE WAS IN MONITORING MODE THE ENTIRE EVENT; MANUAL MODE WAS NEVER SELECTED, AND NO THERAPY DELIVERY WAS ATTEMPTED AS PART OF A RESUSCITATION. ADDITIONALLY, NO DEFIBRILLATION PADS OR PADDLES WERE APPLIED TO THE PATIENT, WHICH WOULD ALSO PREVENT THE DELIVERY OF THERAPY. THE PEF INDICATED THAT ECG CAPTURE WAS INTERMITTENT, WHICH ALIGNS WITH CUSTOMER REPORTS OF ¿DOTTED LINES AND JUMPING WAVE PATTERNS.¿ IT IS POSSIBLE THAT THE CUSTOMER IS REFERRING TO INTERMITTENT ECG SIGNAL SHOWN IN THE PEF FROM (B)(6) 2025, HOWEVER, THIS FILE IS NOT FROM THE EVENT DATE REPORTED BY THE CUSTOMER. THEREFORE, A CONCLUSION ABOUT THE REPORTED FAILURE ¿ECG DOTTED LINES AND JUMPING WAVE PATTERNS ON THE SCREEN¿ CANNOT BE DETERMINED. A PHILIPS FIELD SERVICE ENGINEER (FSE) EVALUATED THE DEVICE ONSITE AT THE CUSTOMER LOCATION, AND DID NOT FIND ANY FAULT WITH THE DEVICE. THE FSE FURTHER TESTED THE DEVICE WITH A SIMULATOR AND ALL WAVEFORMS WERE NORMAL. FURTHERMORE, A PHILIPS PRODUCT SUPPORT ENGINEER (PSE) EVALUATED LOG FILES EXPORTED FROM THE DEVICE FROM BOTH THE EVENT DATE GIVEN BY THE CUSTOMER ((B)(6) 2025) AND THE DATE ON THE PEF ((B)(6) 2025) AND FOUND NO DEVICE ERRORS OR MALFUNCTIONS. THE DEVICE WAS DETERMINED TO BE FUNCTIONING AS PER SPECIFICATION AND CONTINUES TO BE IN USE AT THE CUSTOMER SITE. BASED ON THE INFORMATION AVAILABLE AND THE TESTING CONDUCTED, THE CAUSE OF THE REPORTED PROBLEM WAS NOT DETERMINED. THE REPORTED PROBLEM WAS NOT CONFIRMED. NO DEVICE PROBLEM WAS FOUND. IT HAS BEEN CONCLUDED THAT NO FURTHER ACTION IS REQUIRED AT THIS TIME. Manufacturer Narrative: REPORTING INSTITUTION PHONE # AND REPORTER PHONE # UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21500848&pc=BZD | 2518422-2025-102522 | 2025-02-03 05:00 | Death | RESPIRONICS, INC. | 2025-03-01 05:00 | BZD | BIPAP SYSTEM ONE 60 SERIES | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED ABOUT A BIPAP AUTO, DOM DEVICE. THE PATIENT WAS REACHED OUT TO ABOUT RETURNING HIS RECALLED DEVICE AFTER HE RECEIVED HIS REPLACEMENT. HE MENTIONED THAT HE HAS BEEN USING HIS MOTHER'S CPAP DEVICE SINCE SHE PASSED AWAY AND HAS ONLY RECENTLY OPENED HIS REPLACEMENT DEVICE. HE RECEIVED A NOTIFICATION THAT HE NEEDED TO RETURN HIS RECALLED DEVICE AND HAS RECEIVED THE SHIPPING LABEL. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE MOTHER'S DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21706121&pc=OZD | 1220648-2025-26989 | 2025-03-02 05:00 | Death | ABIOMED, INC. | 2025-03-27 04:00 | OZD | AUTOMATED IMPELLA CONTROLLER | Pumping Stopped; Defective Device | Insufficient Information | P140003 |  | 1 | Event Description: THE COMPLAINANT HAD AN AUTOMATED IMPELLA CONTROLLER (AIC)/IMPELLA CP SUPPORT ONGOING AND THE PATIENT WAS UNDERGOING PERCUTANEOUS CORONARY INTERVENTION FOR THE THROMBOTIC OCCLUSION OF THE CORONARY. THE CONSOLE FAILED DURING THE SUPPORT AND WAS SMELLING BURNT. THE CONSOLE WAS REPLACED AND SUPPORT CONTINUED ON THE SECOND AIC. THE PATIENTS CONDITION CONTINUED TO DETERIORATE OVER SEVERAL HOURS ON THE UNIT POST IMPLANT AND THE PATIENT WAS MADE DO NOT RESUSCITATE AND EXPIRED SHORTLY AFTER. THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE INVESTIGATION OF AUTOMATED IMPELLA CONTROLLER (AIC) - DAMAGE DURING THERAPY HAS BEEN COMPLETED. DATA ANALYSIS: CONSOLE LOGS FROM THE DAY OF THE INCIDENT ALIGN WITH THE COMPLAINT, SHOWING MULTIPLE MOTOR ERROR MESSAGES AFTER APPROXIMATELY 1 HOUR AND 16 MINUTES OF PUMP OPERATION. THE PUMP STOPPED AT (B)(6) 2025 AT 18:55:06. THE CONSOLE WAS TURNED OFF AT AROUND 18:58:12, AND THE AIC WAS EXCHANGED. DEVICE ANALYSIS: THE CONSOLE WAS SENT BACK FOR FURTHER ANALYSIS. DURING THE INVESTIGATION, A BURNT CAPACITOR WAS FOUND ON THE IMPELLATRONIC BOARD, WHICH LED TO THE PUMP STOP. UPON EXAMINING THE AIC, IT WAS NOTED THAT THE LCD FOR BATTERY 1 WAS NOT DISPLAYING ANY CHARGE INDICATION, AND THE POWER BATTERY MANAGER (PBM) WAS FOUND TO BE DEFECTIVE, AS IT WAS UNABLE TO CHARGE THE OTHER BATTERY (BATT 2) AS WELL. THIS WAS NOT RELATED TO THE REPORTED COMPLAINT. AFTER REPLACING THE IMPELLATRONIC PCA, PBM PCA AND BOTH BATTERIES WITH A KNOWN-GOOD, THE AIC OPERATED CORRECTLY WITH BOTH THE PUMP AND PURGE CASSETTE CONNECTED, FUNCTIONING WITHOUT ANY ISSUES DURING 2 HOURS OF TESTING. ROOT CAUSE: THE REPORTED COMPLAINT OF THE WAVEFORM BEGAN TO DISAPPEAR AND FLOW DROPS (I.E. PUMP STOP) WAS DUE TO A COMPONENT MALFUNCTION (FAILED TANTALUM CAPACITOR) ON THE IMPELLATRONIC BOARD. NO ISSUE RELATED TO THE REPORTED FAILURE MODE OF "AIC - DAMAGE DURING THERAPY" WAS SEEN WITHIN THE AIC. Manufacturer Narrative: THE IMPELLA CONSOLE WAS RETURNED, AND AN EVALUATION IS UNDERWAY. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21706547&pc=NPT | 2015691-2025-02292 | 2025-03-12 04:00 | Death | EDWARDS LIFESCIENCES | 2025-03-27 04:00 | NPT | EDWARDS COMMANDER DELIVERY SYSTEM | Difficult or Delayed Positioning; Detachment of Device or Device Component; Material Split, Cut or Torn | Device Embedded In Tissue or Plaque | P140031 |  | 1 | Event Description: AS REPORTED BY OUR UNITED KINGDOM AFFILIATES, DURING IMPLANT OF A 26 MM SAPIEN 3 ULTRA RESILIA VALVE IN THE AORTIC POSITION BY TRANSFEMORAL APPROACH IN A PATIENT WITH HORIZONTAL AORTA, THERE WERE DIFFICULTIES TRYING TO CROSS THE NATIVE ANNULUS. THE SYSTEM WAS RETRACTED SLIGHTLY AND RE-ADVANCED, JUMPING ACROSS THE ANNULUS WITH THE PATIENT STABLE. AT THAT POINT, THE VALVE WAS SUCCESSFULLY IMPLANTED USING -2 ML BUT THE PATIENT'S PRESSURE DID NOT RECOVER. CPR COMMENCED AND PATIENT OUTPUT RETURNED. POST IMPLANT ANGIOGRAM SHOWED A SMALL, LOCALIZED DISSECTION OF THE NON-CORONARY CUSP (NCC). THE PATIENT WENT INTO ARREST AGAIN AND CPR WAS STARTED AGAIN. AN ECHO SHOWED A LARGE PERICARDIAL EFFUSION AND A PERICARDIOCENTESIS WAS PERFORMED DRAINING 2 LITERS OF BLOOD. THE PATIENT WAS INTUBATED AND REGAINED OUTPUT AND WAS STABILIZED. FURTHER IMAGING DEMONSTRATED AN EXTENSION OF THE DISSECTION TO THE ASCENDING AORTA WITH DOUBLE LUMEN. A DECISION WAS TO NOT RESUSCITATE AGAIN, AND THE PATIENT PASSED AWAY A FEW HOURS LATER. THE PERCEIVED ROOT CAUSE OF THE EVENT WAS FELT TO BE A TRAUMA WITH THE DELIVERY SYSTEM DEVICE DURING THE RETRACTION AND RE-ADVANCEMENT OF THE SYSTEM. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED FOR COMPLETION TO THE ENGINEERING EVALUATION. THE FOLLOWING SECTIONS OF THIS REPORT HAS BEEN UPDATED: UPDATE TO B4, G3, G6, H2, H6, H11. THE EVENT REPORTED IS AN ANTICIPATED IN THE RISK MANAGEMENT DOCUMENTATION FOR TRANSCATHETER HEART VALVE PROCEDURES. A PREVIOUS INVESTIGATION INTO THIS TYPE OF EVENT IS CAPTURED IN AN EDWARDS LIFESCIENCES TECHNICAL SUMMARY AND APPLIES TO THIS COMPLAINT. ADDITIONAL ASSESSMENT OF THE FAILURE MODE IS NOT REQUIRED AT THIS TIME. DUE TO THE UNAVAILABILITY OF THE COMPLAINT DEVICE, ENGINEERING WAS UNABLE TO PERFORM ANY VISUAL INSPECTION, FUNCTIONAL TESTING, OR DIMENSIONAL ANALYSIS. PER THE TECHNICAL SUMMARY, THE IFU, CURRENT RISK MITIGATIONS INCLUDE DESIGN AND MANUFACTURING CONTROLS, AND TRAINING MANUALS HAVE BEEN REVIEWED, NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. IMAGERY WAS PROVIDED AND THE FOLLOWING WAS OBSERVED: THV IMPLANTED IN THE NATIVE ANNULUS. THROUGH EXTENSIVE COMPLAINT INVESTIGATIONS, DIFFICULTY OR INABILITY IN TRACKING THE DELIVERY SYSTEM TO THE LANDING ZONE, INCLUDING TRAVERSING THE AORTIC ARCH AND CROSSING THE NATIVE ANNULUS/BIOPROSTHESIS HAS NOT BEEN ASSOCIATED WITH DEVICE MALFUNCTIONS OR MANUFACTURING NONCONFORMANCES. RATHER, THE ROOT CAUSE FOR THESE EVENTS HAVE HISTORICALLY BEEN DUE TO PATIENT FACTORS, POOR GUIDEWIRE MANAGEMENT, DAMAGED GUIDEWIRE, AND/OR EXCESSIVE MANIPULATION OF THE FLEX WHEEL. IN THIS CASE, THE COMPLAINT FOR DIFFICULTY OR INABILITY TO CROSS NATIVE ANNULUS AND/OR BIOPROSTHESIS WAS UNABLE TO BE CONFIRMED DUE TO UNAVAILABILITY OF APPLICABLE IMAGERY OR MEDICAL RECORDS. AVAILABLE INFORMATION SUGGESTS PATIENT FACTORS (CALCIFICATION, HORIZONTAL AORTA) LIKELY CONTRIBUTED TO THE EVENT AS REPORTED INFORMATION INDICATED A 'MODERATE CALCIUM IN THE ANNULUS/LEAFLETS' AND 'HORIZONTAL AORTA.'. PATIENT FACTORS SUCH AS CALCIFICATION AND HORIZONTAL AORTA MAY CAUSE CONSTRAINED SECTIONS OF THE ANATOMY THAT INTERFERE WITH PASSAGE OF THE DELIVERY SYSTEM AND CAUSES DIFFICULTY DURING TRACKING/CROSSING. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. SINCE NO EDWARDS PRODUCT DEFECTS OR LABELING DEFICIENCIES WERE IDENTIFIED, NO CORRECTIVE OR PREVENTATIVE ACTION IS REQUIRED. Manufacturer Narrative: THE INVESTIGATION IS ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21706814&pc=OZD | 1220648-2025-26992 | 2025-02-13 05:00 | Death | ABIOMED, INC. | 2025-03-27 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA CP DUE TO ACUTE MYOCARDIAL INFARCTION. THE PATIENT WAS UNABLE TO MAINTAIN CIRCULATION DUE TO INTRAPERITONEAL BLEEDING. THE BLEEDING WAS LIKELY CAUSED BY INJURY DURING INSERTION OF THE EXTRACORPOREAL MEMBRANE OXYGENATION OR IMPELLA. THERE WERE NO PROBLEMS WITH THE FUNCTIONALITY OF THE PUMP. A BLOOD TRANSFUSION WAS PERFORMED AS A TREATMENT FOR THE BLEEDING BUT THE AMOUNT OF BLOOD TRANSFUSED IS UNKNOWN. THE PATIENT EXPIRED ON SUPPORT. THE CAUSE OF DEATH WAS MULTIPLE ORGAN FAILURE DUE TO BLEEDING.THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21706848&pc=LJT | MW5168306 | 2025-03-14 04:00 | Death | UNK | 2025-03-27 04:00 | LJT | CATHETER PORT |  | Exsanguination; Ventricular Fibrillation; Diminished Pulse Pressure; Asystole |  |  | 1 | Event Description: PER REPORT, VENOUS PORT DISCONNECTED DURING DIALYSIS WHICH LED TO SIGNIFICANT BLOOD LOSS. PATIENT BECAME PULSELESS AND CPR STARTED. NOTED RHYTHM WAS VFIB (VENTRICULAR FIBRILLATION) AND THEN ASYSTOLE. PATIENT UNDERWENT MASSIVE BLOOD TRANSFUSION PROTOCOL. ROSC (RETURN OF SPONTANEOUS CIRCULATION) UNABLE TO BE ACHIEVED. PATIENT PASSED AWAY (B)(6) 2025 1641. CATHETER WAS PLACED (B)(6) 2024 - REPAIR WAS COMPLETED (B)(6) 2024.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21707606&pc=MNS | 2518422-2025-103461 | 2025-03-19 04:00 | Death | RESPIRONICS, INC. | 2025-03-27 04:00 | MNS | PHILIPS RESPIRONICS E30 WITH HUMIDIFIER | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K102465 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A "PATIENT DEATH WAS REPORTED SEVERAL DAYS AFTER USING THE E30 VENTILATOR. THE HOSPITAL STATES THAT PHILIPS DID NOT PERFORM SERVICE TO THE SYSTEM NOR PROCEED WITH THE SCHEDULED RECALL AND COLLECTION, AS PER THE OFFICIAL DOCUMENTATION. THE FAMILY HAS SUBMITTED THE CORRESPONDING LEGAL DEATH CERTIFICATE (ACTA DE DEFUNCIÓN / DEATH CERTIFICATE) THAT MIGHT DIRECTLY LINK THE INCIDENT TO THE USE OF THE DEVICE. THE FAMILY IS NOW REPORTING THE LACK OF FOLLOW UP ON THE RECALL, THE SERVICE OR EXCHANGE NOT BEING PERFORMED, AND OF COURSE THE CASE OF THEIR DECEASED PATIENT PLEASE BE ADVISED THAT PHILIPS MEXICO, UNDER THE SIGNATURE OF (B)(6), ISSUED AN OFFICIAL COMMUNICATION INDICATING THAT ALL E30 SYSTEMS WERE TO BE RECOLLECTED BEFORE (B)(6) 2024. AS OF TODAY, THIS RECOLLECTION HAS NOT BEEN EXECUTED, AND ADDITIONAL PATIENT RISK REPORTS CONTINUE TO BE FILED ON DIFFERENT E-30 SYSTEMS". THE COMPLAINT HAS BEEN REVIEWED BY A CLINICAL EXPERT, AND IT HAS BEEN DETERMINED THAT THE REPORTED EVENT MAKES NO ALLEGATION OF ANY SPECIFIC DEVICE MALFUNCTION, USE ERROR, FAILURE, LABELING, OR OTHER DEFICIENCY THAT IS RELEVANT TO THE REPORTED EVENT. ADDITIONALLY, THERE IS NO OTHER CLINICAL INFORMATION PROVIDED TO INDICATE ANY CAUSAL RELATIONSHIP BETWEEN THE DEVICE AND THE REPORTED DEATH. THEREFORE, BASED ON AVAILABLE INFORMATION THE REPORTED DEATH IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. THEREFORE, BASED ON INFORMATION AVAILABLE AT THIS TIME, THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO A SERIOUS INJURY OR DEATH. NO FURTHER ACTION IS REQUIRED. THE COMPLAINT REVIEW HAS DETERMINED THAT DUE TO INSUFFICIENT INFORMATION THIS WILL BE REPORTED. ALTHOUGH THERE IS NOT AN ALLEGATION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE MENTIONED "DEATH". THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. NO SERIAL NUMBER HAS BEEN PROVIDED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21708330&pc=MFE | 2029214-2025-00787 | 2024-06-27 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-03-27 04:00 | MFE | ONYX | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Intracranial Hemorrhage |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING: ENDOVASCULAR TREATMENT OF COMPLEX VASCULAR PATHOLOGIES IN THE PEDIATRIC POPULATION, OFTEN PERFORMED BY NONPEDIATRIC SUBSPECIALISTS WITH ADAPTATION OF EQUIPMENT AND TECHNIQUES DEVELOPED FOR ADULT PATIENTS. THE STUDY AIMED TO REPORT ON THE CENTER¿S EXPERIENCE WITH SAFETY AND OUTCOMES OF ENDOVASCULAR TREATMENTS FOR PEDIATRIC VASCULAR PATHOLOGIES. THE TIME FRAME OF THIS STUDY WAS: JANUARY 1, 2004, TO DECEMBER 1, 2022. MULTIPLE MANUFACTURER¿S DEVICES WERE USED IN THE STUDY POPULATION THE FOLLOWING MEDTRONIC DEVICES WERE USED: PIPELINE EMBOLIZATION DEVICE (MEDTRONIC, DUBLIN, IRELAND), ONYX (MEDTRONIC, DUBLIN, IRELAND), REACT (MEDTRONIC, DUBLIN, IRELAND), SOLITAIRE (MEDTRONIC, DUBLIN, IRELAND), AXIUM COILS (MEDTRONIC, DUBLIN, IRELAND) DEATHS OCCURRED IN THE STUDY POPULATION THE CAUSES OF DEATH WERE: RUPTURED ANEURYSM WITH FISHER GRADE IV SUBARACHNOID HEMORRHAGE (ONE DEATHS), HEMODYNAMIC FAILURE SECONDARY TO THE DISEASE (ONE DEATH), INTRAVENTRICULAR HEMORRHAGE AFTER EMBOLIZATION (ONE DEATH) AMONG PATIENT ADVERSE EVENTS INCLUDED: ACCESS-SITE COMPLICATIONS: 2 PATIENTS (1.7%), INTRAPROCEDURAL COMPLICATIONS: 3 PATIENTS (2.5%) , TRANSIENT NEUROLOGICAL DEFICITS: 2 PATIENTS (1.7%), SEVERE IN-STENT STENOSIS OBSERVED IN 1 PATIENT (12.5%) AFTER FLOW-DIVERTER TREATMENT, TREATED WITH ANGIOPLASTY, LEFT VERTEBRAL ARTERY DISSECTION TREATED WITH A PIPELINE EMBOLIZATION DEVICE IN 1 PATIENT (4.8%), VESSEL PERFORATION IN ANOTHER PATIENT (4.8%), POSTOPERATIVE TRANSIENT UNILATERAL NEUROLOGICAL DEFICITS IN 2 PATIENTS (9.5%), MILD L EFT-SIDED WEAKNESS IN 1 PATIENT, RESIDUAL NIDUS IN 2 PATIENTS, ALLERGIC REACTION TO CONTRAST MATERIAL IN 1 PATIENT NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21708335&pc=HCG | 2029214-2025-00788 | 2024-06-27 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-03-27 04:00 | HCG | AXIUM | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Intracranial Hemorrhage |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING: ENDOVASCULAR TREATMENT OF COMPLEX VASCULAR PATHOLOGIES IN THE PEDIATRIC POPULATION, OFTEN PERFORMED BY NONPEDIATRIC SUBSPECIALISTS WITH ADAPTATION OF EQUIPMENT AND TECHNIQUES DEVELOPED FOR ADULT PATIENTS. THE STUDY AIMED TO REPORT ON THE CENTER¿S EXPERIENCE WITH SAFETY AND OUTCOMES OF ENDOVASCULAR TREATMENTS FOR PEDIATRIC VASCULAR PATHOLOGIES. THE TIME FRAME OF THIS STUDY WAS: JANUARY 1, 2004, TO DECEMBER 1, 2022. MULTIPLE MANUFACTURER¿S DEVICES WERE USED IN THE STUDY POPULATION THE FOLLOWING MEDTRONIC DEVICES WERE USED: PIPELINE EMBOLIZATION DEVICE (MEDTRONIC, DUBLIN, IRELAND), ONYX (MEDTRONIC, DUBLIN, IRELAND), REACT (MEDTRONIC, DUBLIN, IRELAND), SOLITAIRE (MEDTRONIC, DUBLIN, IRELAND), AXIUM COILS (MEDTRONIC, DUBLIN, IRELAND) DEATHS OCCURRED IN THE STUDY POPULATION THE CAUSES OF DEATH WERE: RUPTURED ANEURYSM WITH FISHER GRADE IV SUBARACHNOID HEMORRHAGE (ONE DEATHS), HEMODYNAMIC FAILURE SECONDARY TO THE DISEASE (ONE DEATH), INTRAVENTRICULAR HEMORRHAGE AFTER EMBOLIZATION (ONE DEATH) AMONG PATIENT ADVERSE EVENTS INCLUDED: ACCESS-SITE COMPLICATIONS: 2 PATIENTS (1.7%), INTRAPROCEDURAL COMPLICATIONS: 3 PATIENTS (2.5%) , TRANSIENT NEUROLOGICAL DEFICITS: 2 PATIENTS (1.7%), SEVERE IN-STENT STENOSIS OBSERVED IN 1 PATIENT (12.5%) AFTER FLOW-DIVERTER TR EATMENT, TREATED WITH ANGIOPLASTY, LEFT VERTEBRAL ARTERY DISSECTION TREATED WITH A PIPELINE EMBOLIZATION DEVICE IN 1 PATIENT (4.8%), VESSEL PERFORATION IN ANOTHER PATIENT (4.8%), POSTOPERATIVE TRANSIENT UNILATERAL NEUROLOGICAL DEFICITS IN 2 PATIENTS (9.5%), MILD L EFT-SIDED WEAKNESS IN 1 PATIENT, RESIDUAL NIDUS IN 2 PATIENTS, ALLERGIC REACTION TO CONTRAST MATERIAL IN 1 PATIENT NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21708444&pc=PYX | 1220648-2025-26997 | 2025-02-02 05:00 | Death | ABIOMED, INC. | 2025-03-27 04:00 | PYX | IMPELLA RP FLEX | Adverse Event Without Identified Device or Use Problem | Hemolysis | P170011 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK IMPLANTED WITH AN IMPELLA RP FLEX DEVICE EXPERIENCED HEMOLYSIS AND WOULD LATER EXPIRE. THE PATIENT WAS ADMITTED IN WITH RIGHT VENTRICULAR DYSFUNCTION. THE IMPELLA RP FLEX WAS IMPLANTED, THE PATIENT WAS PLACED ON CONTINUOUS RENAL REPLACEMENT THERAPY AND WAS SUBSEQUENTLY TRANSFERRED TO THE TERTIARY CENTER FOR CARE. AFTER APPROXIMATELY FOUR DAYS ON SUPPORT, THE PATIENT STARTED HAVING PINK-TINGED URINE AND LABS CAME BACK HEMOLYZED; A CHEST X-RAY THE FOLLOWING DAY CONFIRMED GOOD IMPELLA RP FLEX PLACEMENT. HOWEVER, THE HEMOLYSIS EVENT WAS NOT RESOLVED. THE FOLLOWING DAY IMPELLA ALARMING PURGE SYSTEM BLOCKED; ECHOCARDIOGRAM SHOWED SOME IMPROVEMENT TO RIGHT VENTRICLE. HOWEVER, THE DECISION WAS MADE TO REMOVE THE DEVICE AND THE PATIENT EXPIRED. FURTHER DETAILS SURROUNDING THE EVENT INCLUDING TREATMENT FOR THE HEMOLYSIS ALONG WITH THE CAUSE OF DEATH ARE UNKNOWN. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21708456&pc=FDT | 9610595-2025-04765 | 2024-11-07 05:00 | Death | AIZU OLYMPUS CO., LTD. | 2025-03-27 04:00 | FDT | EVIS EXERA DUODENOVIDEOSCOPE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | K954451 |  | 1 | Event Description: OLYMPUS REVIEWED THE FOLLOWING LITERATURE TITLED "ENDOSCOPIC RETROGRADE CHOLANGIOPANCREATOGRAPHY IN PATIENTS WITH DIFFERENT TYPES OF TOTAL AND PARTIAL GASTRECTOMY." BACKGROUND AND STUDY AIMS: ENDOSCOPIC RETROGRADE CHOLANGIOPANCREATOGRAPHY (ERCP) IN SURGICALLY ALTERED ANATOMY PATIENTS IS CHALLENGING. WE ANALYZED ERCP PROCEDURES AFTER GASTRIC SURGERY WITH MAINTAINED VATER¿S PAPILLA: TOTAL/PARTIAL GASTRECTOMY ROUX-EN-Y, BILLROTH II GASTRECTOMY, SLEEVE GASTRECTOMY AND GASTROJEJUNOSTOMY. METHODS: MONOCENTRIC RETROSPECTIVE ANALYSIS OF PROSPECTIVE ERCP COHORT IN SURGICALLY ALTERED GASTRIC ANATOMY (SAGA) PATIENTS WITH MAINTAINED VATER¿S PAPILLA BETWEEN 2006 AND 2024. RESULTS: 186 ERCP PROCEDURES IN 110 PATIENTS WITH SAGA PROPORTIONS: 37% TOTAL GASTRECTOMY ROUX-EN-Y, 32% BILLROTH II PARTIAL GASTRECTOMY, 15% PARTIAL GASTRECTOMY ROUX-EN-Y, 11% GASTROJEJUNOSTOMY, 4% SLEEVE GASTRECTOMY AND 1% LESS COMMON RECONSTRUCTIONS. ERCP INDICATIONS WERE MAINLY BILIARY (87%). TOTAL TECHNICAL SUCCESS WAS 82% WITH LOWER TECHNICAL SUCCESS IN ROUX-EN-Y GASTRIC SURGERY (73%) VS GASTRIC SURGERY WITHOUT ROUXEN-Y (92%). 13 ADVERSE EVENTS (GRADE II ¿ IV) WERE ENCOUNTERED (7%) WITH 1 MORTALITY DUE TO PREPROCEDURAL CARDIAC ARREST. SLEEVE GASTRECTOMY ALLOWED THE USE OF CONVENTIONAL DUODENOSCOPES WITH 100% TECHNICAL SUCCESS. IN BILLROTH II PATIENTS BOTH DUODENOSCOPES, GASTROSCOPES AND COLONOSCOPES AS WELL AS BALLOON-ASSISTED ENTEROSCOPES WERE USED WITH HIGH TECHNICAL SUCCESS (94-100%), AS WELL AS IN GASTROJEJUNOSTOMY PATIENTS (57-100%). FOR ROUX-EN-Y RECONSTRUCTIONS, ONLY SINGLE-BALLOON ENTEROSCOPES WERE USED WITH ACCEPTABLE SUCCESS (75% IN PARTIAL AND 74% IN TOTAL GASTRECTOMY ROUX-EN-Y). CONCLUSIONS: ERCP IN SAGA PATIENTS IS CHALLENGING REQUIRING DIFFERENT TYPES OF ENDOSCOPES, ESPECIALLY AFTER ROUX-EN-Y TOTAL OR PARTIAL GASTRECTOMY WITH LOWER TECHNICAL SUCCESS AS COMPARED TO BILLROTH II GASTRECTOMY, SLEEVE GASTRECTOMY AND GASTROJEJUNOSTOMY. ADVERSE EVENT RATES ARE COMPARABLE TO ERCP PROCEDURES IN PATIENTS WITH NORMAL ANATOMY. (ACTA GASTROENTERAL BELG., 2025, 88, 23-30). TYPE OF ADVERSE EVENTS/NUMBER OF PATIENTS. CARDIAC ARREST (DEATH) (1 CASE) - A PATIENT DIED WITHIN 30 DAYS OF THE ERCP PROCEDURE THAT WAS PREMATURELY INTERRUPTED BECAUSE OF A CARDIAC ARREST AND A SUCCESSFUL CARDIOPULMONARY RESUSCITATION.Event Description: ADDITIONAL INFORMATION WAS RECEIVED FROM THE AUTHOR THAT THERE WAS NO MALFUNCTION OF THE OLYMPUS DEVICES. IT WAS REPORTED THERE WAS NO LINK BETWEEN THE PATIENT'S DEATH AND THE DUODENOSCOPE. THE PATIENT'S DEATH WAS RELATED TO A COMPLICATION OF THE ANESTHESIA. Manufacturer Narrative: THIS SUPPLEMENTAL REPORT IS BEING SUBMITTED TO PROVIDE ADDITIONAL INFORMATION RECEIVED FROM THE AUTHOR. UPDATED FIELDS: B5. Manufacturer Narrative: THE REPORTS OF THE FOLLOWING PATIENT IDENTIFIERS ARE LINKED: (B)(6). E1: FACILITY NAME: (B)(6). THIS REPORT IS BEING SUBMITTED AS A REPRESENTATIVE DEVICE FOR THE EVENT OF PATIENT DEATH. THE LITERATURE STUDY DID NOT SPECIFY WHICH OF THE FOLLOWING DEVICES WAS INVOLVED WITH THE EVENT. TJF140R, TJF-145, TJF-160R, TJF-160VR, TJF-Q180V, TJF-Q190V, GIF-2TH180, PCF-H190I/L, PCF-HQ190I/L, PCFH190TI, XSIF-180JY, SIF-Y0011, SIF-Q180, SIF-H190, SIF-H290S. THE DEVICE WAS NOT RETURNED FOR EVALUATION; THEREFORE, A DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED. THE MOST PROBABLE CAUSE WAS NOT ESTABLISHED; THE INVESTIGATION FINDINGS DO NOT LEAD TO A CLEAR CONCLUSION ABOUT THE CAUSE OF THE REPORTED ADVERSE EVENT. OLYMPUS WILL CONTINUE TO MONITOR FIELD PERFORMANCE FOR THIS DEVICE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21708580&pc=DRA | 2124215-2025-18981 | 2025-03-10 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-03-27 04:00 | DRA | FARADRIVE STEERABLE SHEATH CLEAR | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Low Blood Pressure/ Hypotension; Tachycardia; Cardiac Tamponade |  |  | 1 | Event Description: IT WAS REPORTED THAT DURING A PULSE FIELD ABLATION (PFA) PROCEDURE USING A FARADRIVE SHEATH THE PATIENT PASSED AWAY. THE CTI LINE ABLATION WAS PERFORMED IN THE RIGHT ATRIUM USING A NON-BOSTON SCIENTIFIC MAPPING CATHETER. AFTER THE CTI LINE WAS PERFORMED, THE PHYSICIAN PERFORMED TRANSEPTAL PUNCTURE. ONCE ACROSS THE SEPTUM, THE VERSACROSS SHEATH WAS EXCHANGED FOR THE FARADRIVE SHEATH AND NON-BOSTON SCIENTIFIC MAPPING CATHETER WAS INSERTED. IT WAS AT THIS POINT THAT THE PATIENT'S HEART RATE BECAME TACHYCARDIC AND HYPOTENSIVE WITH ONLY AN AVAILABLE READING OF MAP OF 24. THE BLOOD PRESSURE WAS THEN REATTEMPTED WITH A SYSTOLIC IN THE 60'S. THE PATIENT'S HEART RATE THEN DECLINED TO THE 50'S-40'S WHERE THEY WERE THEN PULSELESS ELECTRICAL ACTIVITY (PEA) AND REQUIRED RESUSCITATIVE EFFORTS. AN INTRACARDIAC ECHOCARDIOGRAPHY (ICE) SCAN DETECTED A PERICARDIAL EFFUSION WHICH LED TO CARDIAC TAMPONADE. A PERICARDIOCENTESIS WAS PERFORMED AND BLOOD WAS ADMINISTERED. THE PATIENT PASSED AWAY DESPITE THE EFFORTS OF THE LAB STAFF. THE DEVICE IS EXPECTED TO BE RETURNED FOR ANALYSIS. Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21708811&pc=NPT | 2025587-2025-02166 | 2024-12-11 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-27 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem; Central Regurgitation | Endocarditis; Heart Failure/Congestive Heart Failure; Mitral Valve Insufficiency/ Regurgitation | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT ON THE DAY OF IMPLANT OF A TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT'S MITRAL REGURGITATION SEVERITY INCREASED. THE NEXT DAY, THE PATIENT WAS HOSPITALIZED DUE TO CONGESTIVE HEART FAILURE. APPROXIMATELY A MONTH AND A HALF AFTER THE VALVE IMPLANT, THE PATIENT DEVELOPED INFECTIVE ENDOCARDITIS. SUBSEQUENTLY, THE PATIENT DIED APPROXIMATELY ONE MONTH LATER.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE SEVERITY OF THE PATIENT'S MITRAL REGURGITATION PRIOR TO THE IMPLANT OF THE TRANSCATHETER VALVE WAS MODERATE, AND THE MITRAL REGURGITATION REMAINED THE SAME FOLLOWING IMPLANT OF THE TRANSCATHETER AORTIC VALVE. HOWEVER, PER THE PHYSICIAN, THE TRANSCATHETER VALVE CONTRIBUTED TO THE WORSENING MITRAL REGURGITATION.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT CLARIFIED PREVIOUSLY REPORTED INFORMATION INDICATING THAT FOLLOWING THE IMPLANT OF THE VALVE, THE SEVERITY OF THE MITRAL REGURGITATION WAS MODERATE. SUBSEQUENTLY, THE MITRAL REGURGITATION WORSENED FOLLOWING IMPLANT. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTED DATA: H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. CORRECTED DATA: D4. THE VALVE SERIAL NUMBER WAS OMITTED FROM THIS REPORT AS IT WAS REPORTED THAT THE CORRECT SERIAL NUMBER WAS UNKNOWN. H4. DEVICE MANUFACTURE DATE WAS OMITTED FROM THIS REPORT AS IT WAS REPORTED THAT THE CORRECT SERIAL NUMBER OF THE VALVE WAS UNKNOWN. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. D4. H4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21708948&pc=LWP | 2017865-2025-38720 | 2025-03-07 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-03-27 04:00 | LWP | ENDURITY MRI PACEMAKER, SR | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P140033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED WITHOUT CLEAR INFORMATION AS TO WHETHER THE ADVISORY PULSE GENERATOR CAUSED OR CONTRIBUTED TO THE DEATH. IT WAS ALSO NOTED THAT THE RIGHT VENTRICULAR (RV) LEAD IMPLANTED IN 2007 WAS INACTIVE AND REPLACED ON (B)(6) 2020. NO FURTHER INFORMATION WAS AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21708977&pc=NVN | 2017865-2025-38721 | 2025-03-07 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-03-27 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED WITHOUT CLEAR INFORMATION AS TO WHETHER THE ADVISORY PULSE GENERATOR CAUSED OR CONTRIBUTED TO THE DEATH. IT WAS ALSO NOTED THAT THE RIGHT VENTRICULAR (RV) LEAD IMPLANTED IN 2007 WAS INACTIVE AND REPLACED ON (B)(6) 2020. NO FURTHER INFORMATION WAS AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21708987&pc=NVN | 2017865-2025-38722 | 2025-03-07 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-03-27 04:00 | NVN | TENDRIL SDX LEAD | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED WITHOUT CLEAR INFORMATION AS TO WHETHER THE ADVISORY PULSE GENERATOR CAUSED OR CONTRIBUTED TO THE DEATH. IT WAS ALSO NOTED THAT THE RIGHT VENTRICULAR (RV) LEAD IMPLANTED IN 2007 WAS INACTIVE AND REPLACED ON (B)(6) 2020. NO FURTHER INFORMATION WAS AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21709183&pc=DQX | 2025587-2025-02169 | 2025-03-07 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-27 04:00 | DQX | CONFIDA BRECKER CURVE GUIDEWIRE | Patient Device Interaction Problem | Cardiac Arrest; Ventricular Fibrillation | K132623 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE IMPLANT OF THE TRANSCATHETER BIOPROSTHETIC VALVE INTO A FRAGILE PATIENT WITH TYPE IV LUNG CANCER, THE RIGHT CAROTID ACCESS WAS USED. ONCE THE GUIDEWIRE ENTERED THE VENTRICLE, THE PATIENT WENT INTO CARDIAC ARREST. AFTER 20 MINUTES OF RESUSCITATION, THE IMPLANT TEAM DECIDED TO IMPLANT THE VALVE TO SEE IF THE PATIENT WOULD RECOVER. THE PATIENT DIED. ADDITIONAL INFORMATION WAS RECEIVED THAT A MEDTRONIC (CONFIDA) GUIDEWIRE WAS USED DURING THE PROCEDURE. PER THE PHYSICIAN, THE CAUSE OF DEATH WAS VENTRICULAR FIBRILLATION AND CARDIORESPIRATORY ARREST, WHICH THE PATIENT'S UNDERLYING MEDICAL HISTORY CONTRIBUTED TO DUE TO BEING HIGH RISK FROM COMORBIDITY, SEVERE LUNG DISEASE, AND SEVERE MALNUTRITION. AFTER THE VALVE WAS IMPLANTED, NO OBSTRUCTION OF THE CORONARY ARTERIES OR RUPTURE WAS NOTED.Event Description: ADDITIONAL INFORMATION WAS RECEIVED WHICH INDICATED THAT PER THE PHYSICIAN, NEITHER THE GUIDEWIRE, VALVE, NOR THE DELIVERY CATHETER SYSTEM CAUSED OR CONTRIBUTED TO THE ADVERSE EVENTS.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE PATIENT WAS IN EXTREMELY FRAGILE CONDITION WITH STAGE IV PULMONARY ARTERY DISEASE PRIOR TO THE VALVE IMPLANT PROCEDURE. THEY WERE EVALUATED FOR A VALVE DUE TO SEVERE AORTIC STENOSIS. IT WAS NOT POSSIBLE TO USE VIABLE FEMORAL ACCESS, THEREFORE RIGHT CAROTID ACCESS WAS USED AS AN ALTERNATIVE STRATEGY. PRIOR TO THE 23 MILLIMETER (MM) VALVE BEING INSERTED INTO THE PATIENT, AS THE GUIDEWIRE CROSSED INTO THE LEFT VENTRICLE, THE PATIENT SUFFERED CARDIAC ARREST. ADVANCED CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED IMMEDIATELY, WHICH LASTED MORE THAN 20 MINUTES WITHOUT RESTORING SPONTANEOUS CIRCULATION. UP TO THIS POINT, THE VALVE HAD NOT BEEN INSERTED, IMPLANTED OR INTERACTED WITH THE PATIENT IN ANY WAY. IN AN ATTEMPT TO REVERSE THE PATIENT'S CRITICAL CONDITION, THE MEDICAL TEAM DECIDED TO PROCEED WITH THE IMPLANTATION OF THE VALVE AS A LAST RESORT HOPING TO RESTORE CARDIAC FUNCTION. HOWEVER, THE PATIENT NEVER REGAINED EFFECTIVE CIRCULATION AND THEY DIED DURING THE PROCEDURE. THE FATAL OUTCOME WAS NOTED TO BE DUE TO THE PATIENT'S UNDERLYING CRITICAL ILLNESS AND EXTREME INSTABILITY. THEREFORE, THE DEVICE DID NOT CONTRIBUTE TO OR WORSEN THE REPORTED EVENT. DURING THE PROCEDURE AND AFTERWARDS, THE LEAD IMPLANT PHYSICIAN STATED " THE VALVE DID NOT CAUSE ANY HARM TO THE PATIENT. THE DEVICE WAS NOT INVOLVED IN THE CARDIAC ARREST OR THE PATIENT'S HEMODYNAMIC DETERIORATION." Manufacturer Narrative: SECTION D REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: PRODUCT ID: EVOLUTFX-23 (SERIAL: (B)(6)); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE (B)(6) 2025; EXPLANT DATE N/A. PRODUCT ID: D-EVOLUTFX-2329 (LOT: 0012526061); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE N/A; EXPLANT DATE N/A. PRODUCT ID L-EVOLUTFX-2329 (LOT: 0012526084); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE N/A; EXPLANT DATE N/A. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21709569&pc=FJK | 8041145-2025-00005 | 2025-02-15 05:00 | Death | NIPRO (THAILAND) CORP. LTD. | 2025-03-27 04:00 | FJK | NIPRO BLOOD TUBING SET W/TP | Fluid/Blood Leak; Detachment of Device or Device Component | Hemorrhage/Bleeding | K072024 |  | 1 | Event Description: FEMALE PATIENT WITH CVC LINE AND FOR SOME UNKNOWN REASONS THERE WAS AN ACCIDENTAL DISCONNECTION FROM THE VENOUS BLOODLINE CONNECTED TO THE CVC LINE AND THEREFORE A MASSIVE BLEEDING, WITH THE DEATH OF THE PATIENT, CONFIRMED THE (B)(6) BY THE CUSTOMER TOWARDS THE SALES REP, WHO VISITED THE CENTER ON THE (B)(6). THE COSTUMER DOES NOT BLAME THE MACHINE, OR THE LINES, AS THEY THREW AWAY EVERYTHING AND USED THE MACHINE FOR 4 MORE TREATMENTS AFTERWARDS. THIS IS ALSO THE REASON THE COSTUMER DID NOT REPORT THIS ISSUE TO NIPRO. HOWEVER, BECAUSE THE MACHINE WAS USED AFTERWARDS FOR FOUR MORE TREATMENTS, WE COULD NOT SAVE ANY DATA, LOGS OR PICTURES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21709665&pc=LWS | 9614453-2025-01220 | 2025-02-13 05:00 | Death | MEDTRONIC EUROPE SARL | 2025-03-27 04:00 | LWS | EVERA MRI XT DR SURESCAN | Appropriate Term/Code Not Available | Insufficient Information | P980016 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE PATIENT REPRESENTATIVE REPORTED THAT THE DEVICE WAS NOT WORKING PROPERLY. CAUSE OF DEATH WAS REQUESTED BUT WAS NOT RECEIVED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE DEVICE WAS RETURNED AND ANALYZED. RETURNED PRODUCT ANALYSIS WAS PERFORMED AND NO ANOMALIES WERE FOUND. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21710380&pc=LGZ | 21710380 | 2025-03-24 04:00 | Death | BELMONT INSTRUMENT LLC | 2025-03-27 04:00 | LGZ | 3-SPIKE DISPOSABLE SET (1 PIECE) |  | Cardiac Arrest; Exsanguination; Respiratory Arrest |  |  | 1 | Event Description: DUE TO ACTIVE BLEEDING, A PATIENT WAS REQUIRING THE MASSIVE TRANSFUSION PROTOCOL. THE PATIENT WAS CRITICALLY ILL. THE STAFF SET-UP THE BELMONT RAPID INFUSER TO INFUSE THE BLOOD. AS THE BLOOD WAS INFUSING, THE STAFF NOTICED THE TUBING WAS LEAKING AND CAUSING THE BLOOD LEAK OUT. UPON EXAMINATION OF THE TUBING, THE STAFF NOTICED THE TUBING WAS CRACKED ABOVE THE BLUE ARROW. THE TUBING WAS REMOVED AND A 2ND SET OF TUBING WITH A DIFFERENT LOT NUMBER WAS UTILIZED. THE PATIENT DID NOT RECEIVE ALL THE BLOOD FROM THE UNIT TRANSFUSING. THE PATIENT DID HAVE A CARDIAC ARREST AND DID NOT SURVIVE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21710689&pc=DSQ | 2916596-2025-01815 | 2025-02-22 05:00 | Death | THORATEC CORPORATION | 2025-03-27 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Cardiomyopathy; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS HOSPITALIZED WITH END STAGE HEART FAILURE AND MADE THE DECISION TO HAVE THEIR LEFT VENTRICULAR DEVICE TURNED OFF. THE PATIENT PASSED AWAY DUE TO END STAGE CARDIOMYOPATHY. THE PATIENT'S DEATH WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED AND IT WAS CONFIRMED THAT THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: CORRECTED DATA: SECTION G2: CORRECTED. SECTION H6: HEALTH EFFECT - CLINICAL CODE AND HEALTH EFFECT - IMPACT CODE CORRECTED. MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. IT WAS REPORTED THAT THE PATIENT WAS HOSPITALIZED WITH END STAGE HEART FAILURE AND EVENTUALLY MADE THE DECISION TO HAVE THE LEFT VENTRICULAR ASSIST DEVICE TURNED OFF. THE PATIENT ULTIMATELY EXPIRED DUE TO END STAGE CARDIOMYOPATHY. THE PATIENT¿S OUTCOME WAS NOT CONSIDERED TO BE DEVICE RELATED AND THE DEVICE OPERATED AS EXPECTED. IT WAS COMMUNICATED THAT THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE ELECTRONIC IFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REV. D IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21710710&pc=OZD | 1220648-2025-27005 | 2025-02-18 05:00 | Death | ABIOMED, INC. | 2025-03-27 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Perforation | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT WAS ON IMPELLA CP SUPPORT AND EXPIRED FROM A LEFT VENTRICLE RUPTURE. THE LEFT VENTRICLE RUPTURE WAS FOUND AFTER IMPELLA INSERTION. IT IS UNCLEAR WHETHER IMPELLA INSERTION CAUSED THE PROBLEM OR NOT. THE RELATIONSHIP BETWEEN THE CAUSE OF DEATH AND THE IMPELLA DEVICE IS UNKNOWN. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE,¿AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿. SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION: SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER: SECTION: DIRECT AORTIC INSERTION: SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION. [ADDITION FOR 5.0, 5.5 ONLY] IN INSTANCES WHERE THE IMPELLA PUMP HAS BEEN PLACED PRIOR TO PERFORMING CARDIAC SURGERY WITH AORTIC CROSS CLAMPING AND CARDIOPLEGIC ARREST, CARE SHOULD BE TAKEN WHEN MANIPULATING THE HEART WHEN THE PUMP POSITION IS FIXED WITH APPLICATION OF THE AORTIC CROSS CLAMP ACROSS THE PUMP CATHETER.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE. ¿TO REDUCE THE RISK OF CARDIAC INJURY (INCLUDING VENTRICULAR PERFORATION), PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY IN THE SETTINGS OF ACUTE INFARCTION WITH TISSUE NECROSIS.¿ ¿TO REDUCE THE RISK OF VASCULAR INJURY, PHYSICIANS SHOULD EXERCISE CAUTION WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX PERIPHERAL VASCULAR ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: UNREPAIRED ABDOMINAL AORTIC ANEURYSM, SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM, DISSECTION OF THE ASCENDING/ TRANSVERSE/DESCENDING AORTA, CHRONIC ANATOMICAL CHANGES IN THE RELATIONSHIP OF THE AORTA/AORTIC VALVE/VENTRICULAR ALIGNMENT, SIGNIFICANT MOBILE ATHEROMATOUS DISEASE IN THE THORACIC OR ABDOMINAL AORTA OR PERIPHERAL VESSELS.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING A RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY IN THE LEFT VENTRICLE AFTER CPR WITH ECHOCARDIOGRAPHY GUIDANCE. SECTION: WARNINGS & CAUTIONS: CAUTIONS: ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21711359&pc=LWS | 2649622-2025-08446 | 2024-11-15 05:00 | Death | MEDTRONIC EUROPE SARL | 2025-03-27 04:00 | LWS | COBALT¿ XT VR MRI SURESCAN¿ | Over-Sensing; Impedance Problem; Adverse Event Without Identified Device or Use Problem | Insufficient Information | P980016 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS DECEASED. IT WAS REPORTED THAT THE RIGHT VENTRICULAR (RV) LEAD HAD EXHIBITED OVERSENSING AND A LEAD WARNING WAS TRIGGERED FOR OUT OF RANGE (OOR) PACING IMPEDANCE. ADDITIONAL INFORMATION RELATED TO THE CAUSE OF DEATH WAS REQUESTED AND NOT RECEIVED. THE DISPOSITION OF THE LEAD REMAINS UNKNOWN.Event Description: IT WAS REPORTED THAT THE PATIENT WAS IN THE INTENSIVE CARE UNIT WITH ATRIAL TACHYCARDIA/ATRIAL FIBRILLATION (AT/AF), THE PATIENT PASSED AWAY ONE WEEK LATER. ADDITIONAL INFORMATION RELATED TO THE CAUSE OF DEATH WAS REQUESTED AND NOT RECEIVED. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID DVPA2D4 ICD IMPLANTED: (B)(6) 2024 EXPLANTED: (B)(6) 2025. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTION: B5; D1; D3; D4; D10; H2; H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21712553&pc=NPT | 2025587-2025-02183 | 2024-03-11 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-27 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Respiratory Insufficiency | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 10 MONTHS AND 2 WEEKS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER VALVE, THE PATIENT WAS HOSPITALIZED DUE TO RESPIRATORY DISTRESS. SUBSEQUENTLY, THE PATIENT PASSED AWAY. THE CAUSE OF DEATH IS UNKNOWN. THE RELATIONSHIP BETWEEN THE DEATH AND THE VALVE OR PROCEDURE IS UNKNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTION : ADDED D.4 AND H.4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21712726&pc=FRN | 3000240707-2025-00218 | 2025-03-19 04:00 | Death | FRESENIUS VIAL S.A.S. | 2025-03-27 04:00 | FRN | INJECTOMAT AGILIA BR | Excess Flow or Over-Infusion | No Clinical Signs, Symptoms or Conditions | K121613 |  | 1 | Event Description: THE FOLLOWING HAS BEEN REPORTED BY CUSTOMER: THE CUSTOMER REPORTS THAT DURING THE ADMINISTRATION OF MEDICATION, A LARGER AMOUNT THAN PROGRAMMED WAS INFUSED, RESULTING IN THE DEATH OF A NEWBORN. REPORTING AS A CONSERVATIVE MEASURE. ADVERSE EFFECTS WERE REPORTED. MORE INFORMATION IS NEEDED TO COMPLETE THE INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21712731&pc=FRN | 3004548776-2025-00220 | 2025-03-19 04:00 | Death | FRESENIUS VIAL S.A.S. | 2025-03-27 04:00 | FRN | INJECTOMAT AGILIA BR | Excess Flow or Over-Infusion | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: THE FOLLOWING HAS BEEN REPORTED BY CUSTOMER: THE CUSTOMER REPORTS THAT DURING THE ADMINISTRATION OF MEDICATION, A LARGER AMOUNT THAN PROGRAMMED WAS INFUSED, RESULTING IN THE DEATH OF A NEWBORN. REPORTING AS A CONSERVATIVE MEASURE. ADVERSE EFFECTS WERE REPORTED. MORE INFORMATION IS NEEDED TO COMPLETE THE INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21712960&pc=OZD | 1220648-2025-27012 | 2025-01-06 05:00 | Death | ABIOMED, INC. | 2025-03-27 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Tachycardia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT WITH IMPELLA CP SUPPORT DEVELOPED BLEEDING AT THE RIGHT AXILLARY SITE. PRESSURE DRESSING WAS APPLIED, AND HEPARIN WAS HELD FOR 48 HOURS. BLEEDING WAS RESOLVED. THREE DAYS LATER, THE PATIENT HAD A 40-BEAT RUN OF VENTRICULAR TACHYCARDIA AND WAS CONVERTED. THE PATIENT WAS BACK IN PACED RHYTHM. THE PATIENT WAS MADE A DO NOT RESUSCITATE AND EXPIRED ON IMPELLA SUPPORT. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21713261&pc=QJP | 2029214-2025-00791 | 2023-12-14 05:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-03-27 04:00 | QJP | RIST | Adverse Event Without Identified Device or Use Problem | Thrombosis/Thrombus |  |  | 1 | Event Description: MOLINARO, S., RUSSO, R., MISTRETTA, F., RISI, G., GAVA, U. A., BERGUI, M.; NEUROINTERVENTION; 2024; 19:6-13; A ¿RADIAL READY¿ TRICOAXIAL SETUP FOR ANTERIOR CIRCULATION MECHANICAL THROMBECTOMY: TECHNICAL ASPECTS AND PRELIMINARY RESULTS.; DOI.ORG/10.5469/NEU ROINT.2023.00500 LITERATURE WAS REVIEWED REGARDING: A ¿RADIAL READY¿ TRICOAXIAL SETUP FOR ANTERIOR CIRCULATION MECHANICAL THROMBECTOMY: TECHNICAL ASPECTS AND PRELIMINARY RESULTS. THE TIME FRAME OF THIS STUDY WAS: NOVEMBER 2022 TO NOVEMBER 2023. MULTIPLE MANUFACTUR ER¿S DEVICES WERE USED IN THE STUDY POPULATION. THE FOLLOWING MEDTRONIC DEVICES WERE USED: RIST 7F GUIDE CATHETER. DEATHS OCCURRED IN THE STUDY POPULATION: THE OVERALL IN-HOSPITAL MORTALITY RATE WAS 2/20 (10.0%). THE CAUSES WERE NOT REPORTED. AMONG PATIENT ADVERSE EVENTS INCLUDED: FOUR PATIENTS (20%) SHOWED A HEMORRHAGE AT 24 HOURS POSTPROCEDURAL CT SCAN, OF WHICH 2 (10.0%) WERE SCORED AT HEMORRHAGIC INFARCTION TYPE 1 (HI1) AND 2 (10.0%) DEVELOPED SYMPTOMATIC PARENCHYMATOUS HEMATOMA TYPE 2 (PH2), ACCORDING TO HEIDELBERG BLEEDING CLASSIFICATION. RATE OF RADIAL ARTERY OCCLUSION AT 24 HOURS WAS 10.6%. NO ACCESS-SITE HEMORRHAGIC COMPLICATIONS WERE REPORTED. SUCCESSFUL REVASCULARIZATION TO MTICI 2B-3 WAS ACHIEVED IN 18/20 PATIENTS (90%). NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS.Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED THERE WERE NO MEDTRONIC PRODUCTS RELATED ISSUES. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21713314&pc=BZD | 2518422-2025-039605 | 2023-12-19 05:00 | Death | RESPIRONICS, INC. | 2025-03-27 04:00 | BZD | PHILIPS CPAP DEVICE | Degraded | Insufficient Information |  |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT PASSED AWAY. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972,Z-1973, AND Z-1974.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21713495&pc=QFG | 3019004087-2025-00327 | 2025-02-21 05:00 | Death | BETA BIONICS | 2025-03-27 04:00 | QFG | ILET BIONIC PANCREAS | Insufficient Information | Hypoglycemia | K231485 |  | 1 | Event Description: ON 2/25/25 A BETA BIONICS CLINICAL DIABETES SPECIALIST (CDS) REPORTED THAT AN ILET USER EXPERIENCED A LOW BLOOD GLUCOSE (BG) EVENT AND WAS HOSPITALIZED. IT WAS LATER NOTED THAT THE USER IS NOW DECEASED. HOWEVER, THE PHYSICIAN IS UNCERTAIN WHETHER THE CAUSE OF DEATH IS RELATED TO DIABETES. AT THE TIME OF THIS REPORT, MULTIPLE ATTEMPTS BY BETA BIONICS TO DETERMINE IF THE USER WAS ON THE ILET AT THE TIME OF THE EVENT AS WELL AS CAUSE OF DEATH HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: NO PRODUCT WAS RETURNED FOR EVALUATION. ATTEMPTS TO DOWNLOAD THE DEVICE ENGINEERING LOGS WERE UNSUCCESSFUL, INDICATING THAT THE LOGS HAVE NOT BEEN SYNCED SINCE DEVICE WAS SHIPPED FROM MANUFACTURING. AS SUCH, A LOG REVIEW WAS UNABLE TO BE CONDUCTED NO PRODUCT PERFORMANCE ISSUES CAN BE IDENTIFIED. IF THE PRODUCT IS RECEIVED AT A LATER DATE, OR THE ILET DATA IS UPLOADED, THE COMPLAINT WILL BE REOPENED AND INVESTIGATED ACCORDINGLY. THE DEVICE HISTORY RECORD (DHR) REVIEW WAS COMPLETED, AND THIS DEVICE PASSED ALL MANUFACTURING RELEASE CRITERIA FOR DISTRIBUTION. THERE WERE NO ISSUES IDENTIFIED THAT WOULD HAVE IMPACTED THIS EVENT. WITHOUT DEVICE DATA OR A DATE OF DEATH, IT CANNOT BE DETERMINED IF THE USER WAS WEARING THE ILET AT THEIR TIME OF DEATH OR IF THE ILET WAS RELATED TO THEIR PASSING. CHART NOTES ON FILE INDICATE A HISTORY OF UNCONTROLLED DIABETES WITH RECURRENT HYPOGLYCEMIA AND AN A1C OF 11.4% PRIOR TO USE OF THE ILET. IT IS POSSIBLE THAT THERE WERE COMORBIDITIES THAT MAY HAVE CONTRIBUTED TO THE DEMISE AND WITHOUT FURTHER INFORMATION, THIS CANNOT BE RULED OUT. THIS CASE WILL BE REVISITED AND INVESTIGATED FURTHER IF ANY NEW INFORMATION BECOMES AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21713547&pc=QEW | 3020347218-2025-00023 | 2024-01-01 05:00 | Death | INARI MEDICAL, INC. | 2025-03-27 04:00 | QEW | INARI MEDICAL THROMBECTOMY CATHETERS | Insufficient Information | Thrombosis/Thrombus |  |  | 1 | Event Description: ON FEBRUARY 27, 2025 INARI RECEIVED A BRAND EQUITY SURVEY RESPONSE WHERE AN ANONYMOUS INDIVIDUAL REPORTED THAT AFTER TREATMENT WITH INARI DEVICES SEVERAL OF THEIR PATIENTS RETHROMBOSED AND ONE HAD A FATAL PULMONARY EMBOLISM. Manufacturer Narrative: CORRECTION TO SECTION B4 IN INITAL REPORT 3020347218-2025-00023: "DATE OF THIS REPORT" SHOULD READ "27-MAR-2025. Manufacturer Narrative: THERE WERE NO SPECIFIC CLINICAL CASES, EVENT DATES, HOSPITAL, OR PHYSICIAN NAMES PROVIDED IN THIS SURVEY RESPONSE. ADDITIONAL INFORMATION FROM THE ANONYMOUS COMPLAINANT WAS REQUESTED, HOWEVER THEY DECLINED. DUE TO LACK OF INFORMATION INARI IS UNABLE TO INVESTIGATE THIS SPECIFIC CLAIM. MANUFACTURER REFERENCE: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21713590&pc=OZD | 1220648-2025-27013 | 2025-01-13 05:00 | Death | ABIOMED, INC. | 2025-03-27 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hematoma; Hemolysis; Hemorrhage/Bleeding; Low Blood Pressure/ Hypotension | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT FOLLOWING IMPELLA CP IMPLANT, A PATIENT EXPERIENCED BLEEDING AT THE TRIPLE LUMEN AND IMPELLA ACCESS SITE. A LARGE HEMATOMA WAS NOTED AT THE IMPELLA SITE. THE PATIENT HAD A DROP IN BLOOD PRESSURE AND CENTRAL VENOUS PRESSURE. MANUAL PRESSURE WAS HELD AND FIVE UNITS OF PACKED RED BLOOD CELLS WITH TWO UNITS OF FRESH FROZEN PLASMA WERE GIVEN TO COUNTERACT THE BLOOD LOSS. HEPARIN WAS DISCONTINUED IN RESPONSE TO THE BLEED AND FEMSTOP WAS APPLIED JUST ABOVE THE ACCESS SITE. WITH THE SITE REDRESSED AND SURGISEAL APPLIED, THE BLEEDING WAS CONTAINED. SUPPORT WAS CONTINUED FOLLOWING THE NOTED INTERVENTION. HOWEVER, THE PATIENT SUCCUMBED TO THE THEIR CONDITION AND PASSED AWAY THE FOLLOWING DAY. IT WAS NOTED THAT THE PATIENT HAD BEGUN TO SHOW SIGNS OF HEMOLYSIS PRIOR TO THEIR PASSING FOR WHICH SUPPORT LEVELS WERE TURNED DOWN. IT IS UNKNOWN IF ALTERING THE SUPPORT LEVEL RESOLVED THE HEMOLYSIS. IT IS UNKNOWN IF THE NOTED ISSUES CONTRIBUTED TO THE PATIENT'S PASSING. Manufacturer Narrative: THE INVESTIGATION FOR ACCESS SIT BLEEDING HAS BEEN COMPLETED. THE PUMP WAS RETURNED FOR INSPECTION. UPON INSPECTION, NO ABNORMALITIES WERE NOTED WITH THE PUMP NOR THE CATHETER. THE ROOT CAUSE OF THE ACCESS SITE BLEEDING WAS DETERMINED TO BE ANTICOAGULATION MANAGEMENT BECAUSE THE PARTIAL THROMBOPLASTIN TIME VALUES (200) WERE HIGHER THAN THERAPEUTIC VALUES AND THE HEPARIN WAS DISCONTINUED AND FEMSTOP WAS PLACED TO CONTROL THE BLEEDING. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM, SECTION: GENERAL PATIENT CARE CONSIDERATIONS, ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT, USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER: ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE: ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS: ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION: ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿ Manufacturer Narrative: E.4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27013 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. H.1 TYPE OF REPORTABLE EVENT WAS PREVIOUSLY ENTERED INCORRECTLY ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27013.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21713748&pc=OZD | 1220648-2025-27015 | 2025-03-01 05:00 | Death | ABIOMED, INC. | 2025-03-27 04:00 | OZD | AUTOMATED IMPELLA CONTROLLER | Failure to Read Input Signal | Insufficient Information | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THE PURGE DISC WAS NOT DETECTED WHEN CHANGING PURGE CASSETTE. AFTER SEVERAL ATTEMPTS TO SLIDE IT IN, THEY SWITCHED TO A NEW AUTOMATED IMPELLA CONTROLLER AND THERE WAS NO FURTHER ISSUES. AT DAY 55 OF PUMP SUPPORT CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE INVESTIGATION HAS BEEN COMPLETED. THE AUTOMATED IMPELLA CONTROLLER WAS RETURNED FOR INVESTIGATION. THE LOGS FROM THE COMPLAINT DATE REVEALED THAT THE CONSOLE ISSUED PURGE DISC NOT DETECTED ALARM AND OCCURRED MULTIPLE TIMES. THE PHYSICAL CONSOLE WAS EXAMINED AND FOUND TO HAVE A BROKEN PURGE DISC FLAG. THE PURGE DISC FLAG WAS OBSERVED TO BE BROKEN AND WAS THE ROOT CAUSE OF THE REPORTED PURGE DISC NOT DETECTED ALARM ISSUE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21713777&pc=OZD | 1220648-2025-27016 | 2025-01-31 05:00 | Death | ABIOMED, INC. | 2025-03-27 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT ON IMPELLA CP SUPPORT WAS OBSERVED TO HAVE A CEREBRAL HEMORRHAGE. THE PATIENT LATER EXPIRED DUE TO THE CEREBRAL HEMORRHAGE. IT WAS BELIEVED THAT A POTENTIAL RELATIONSHIP EXISTED BETWEEN THE CAUSE OF DEATH AND THE DEVICE BECAUSE ACTIVATED CLOTTING TIME (ACT) MANAGEMENT WAS PERFORMED, BUT POOR CONTROL OF OTHER COAGULATION SYSTEMS WAS OBSERVED. THE ACT VALUES WERE CONTROLLED AT 160-180. ACTIVATED PARTIAL THROMBOPLASTIN TIME WAS OVER 300. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21715427&pc=OZP | 2032227-2025-144800 | 2024-12-31 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-03-27 04:00 | OZP | MINIMED 780G US SYSTEM BLE CONNECT 3.0 MG/DL | Adverse Event Without Identified Device or Use Problem | Hypoglycemia; Heart Failure/Congestive Heart Failure | P160017 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY AT CARE CENTER ON (B)(6) 2024. THE CAUSE OF DEATH WAS HYPOGLYCEMIA AND CONGESTIVE HEART FAILURE. CUSTOMER BLOOD GLUCOSE WAS REPORTED AS 34 MG/DL AT THE TIME OF INCIDENT. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: MMT-7841ZN, MMT-386A, MMT-1884, MMT-1715K, AND MMT-332A. TROUBLESHOOTING WAS PERFORMED FOR DECEASED REPORTING, CUSTOMER WAS NOT WORN INSULIN PUMP AND SENSOR AT THE TIME OF PASSING. INSULIN PUMP AND SENSOR WAS REMOVED WHEN CUSTOMER WAS HOSPITALIZED. THERE IS NO MENTION OF THE AUTO MODE FEATURE AT THE TIME OF THE EVENT. IT WAS AGREED TO RETURN THE INSULIN PUMP. NO PRODUCT RETURN IS REQUIRED FOR MMT-7841ZN. NO PRODUCT RETURN IS REQUIRED FOR MMT-386A. MMT-1884 WAS REQUESTED, AND THE RESPONSE WAS THE DEVICE WILL BE RETURNED. MMT-1715K WAS REQUESTED, AND THE RESPONSE WAS THE DEVICE WILL BE RETURNED. NO PRODUCT RETURN IS REQUIRED FOR MMT-332A. Manufacturer Narrative: THE PUMP PASSED THE FUNCTIONAL TEST, INCLUDING THE SELF TEST, SLEEP CURRENT MEASUREMENT, ACTIVE CURRENT MEASUREMENT, REWIND TEST, PRIME/SEATING TEST, BASIC OCCLUSION TEST, OCCLUSION TEST, FORCE SENSOR TEST, DISPLACEMENT TEST, AND THE DAT AT 0.0872 INCHES. DELIVERY ANOMALY-POSSIBLE OVER DELIVERY NOT CONFIRMED. THE FOLLOWING WERE NOTED DURING VISUAL INSPECTION: MINOR SCRATCHED DISPLAY WINDOW, SCRATCHED CASE AND PILLOWING KEYPAD OVERLAY. TEST P-CAP AND RESERVOIR LOCKED PROPERLY INTO RESERVOIR COMPARTMENT DURING TESTING. HISTORY DOWNLOAD WAS SUCCESSFUL USING THUMP AND CARELINK UPLOAD WAS SUCCESSFUL. THE PUMP WAS RECEIVED WITHOUT A BATTERY. THE PUMP WAS RECEIVED WITHOUT THE BATTERY CAP. THE PUMP HISTORY FILE LISTS DATA FROM (B)(6) 2024. UNABLE TO VERIFY CUSTOMER'S DAILY TOTAL OF ALL INSULIN DELIVERED, DAILY TOTAL OF BASAL INSULIN DELIVERED AND DAILY TOTAL OF BOLUS INSULIN DELIVERED SURROUNDING THE EVENT DATE OF 31-DEC-2024 LISTED ON SMARTSOLVE DUE TO NO DATA AVAILABLE FOR THE PRIMARY SVN (B)(4) AND RELATED SVN (B)(4). UNABLE TO PERFORM THE HISTORY REVIEW 1 WEEK PRIOR TO THE EVENT DATE 31-DEC-2024 IN THE PUMP HISTORY FILE DUE TO NO DATA AVAILABLE FOR THE PRIMARY SVN (B)(4) AND RELATED SVN (B)(4). ON A RELATED SVN (B)(4), UNABLE TO VERIFY CUSTOMER'S DAILY TOTAL OF ALL INSULIN DELIVERED, DAILY TOTAL OF BASAL INSULIN DELIVERED AND DAILY TOTAL OF BOLUS INSULIN DELIVERED SURROUNDING THE EVENT DATE OF DEC 22, 2024, LISTED ON SMARTSOLVE DUE TO NO DATA AVAILABLE. ON A RELATED SVN (B)(4), PERFORM THE HISTORY REVIEW 1 WEEK PRIOR TO THE EVENT DATE DEC 22, 2024, IN THE PUMP HISTORY FILE AND FOUND A LOST SENSOR ALERT ON (B)(6) 2024, A FAILEDBATTTEST (58) ALERT ON 12/21/2024, AND 2 PUMP ERROR 23 ALARMS ON (B)(6) 2024 14:19:04.000 AND ON (B)(6) 2024 14:29:00.000. (THE PUMP HISTORY FILE LISTS DATA ON DECEMBER 16-21, 2024. HOWEVER, THERE WAS NO DATA AVAILABLE ON DEC 22, 2024.) THE PUMP PROPERLY PAIRED WITH THE GUARDIAN LINK 3 TRANSMITTER. AFTER THE PUMP COMPLETED WARM UP AND CALIBRATION, THE PUMP WAS ABLE TO DISPLAY THE TEST BG VALUE OF 240 MG/DL ON THE PUMP'S HOME SCREEN. NO COMMUNICATION ANOMALY, CALIBRATION ANOMALY OR LOST SENSOR ALERT NOTED. THE POWER MANAGEMENT TOOL CONFIRMED THE UNLOADED VOLTAGE (UL VLITH) AND LOADED VOLTAGE (LOADED VLITH) WAS WITHIN SPEC RANGE. FAILED BATTERY TEST ALARM WAS EXPECTED DUE TO THE CUSTOMER'S BATTERY INSERTED ON A BATTERY CHANGE DOES NOT HAVE SUFFICIENT VOLTAGE. PUMP ERROR 23 ALARMS WERE EXPECTED SINCE THE PUMP RESETS AFTER THE BATTERY WAS REMOVED FOR MORE THAN 10 MINUTES. PUMP WAS CUT OPEN TO PERFORM VISUAL INSPECTION AND FOUND NO EVIDENCE OF PHYSICAL OR MOISTURE DAMAGE ON THE PCBA1, PCBA2, FORCE SENSOR, MOTOR AND VIBRATOR ASSEMBLY NOTED. FORCE SENSOR ZERO OFFSET WITHIN SPECIFICATION (22.8 MV). THE PUMP PASSED THE FUNCTIONAL TESTING. UNABLE TO CONFIRM ALLEGED LOW BGS. DELIVERY ANOMALY-POSSIBLE OVER DELIVERY NOT CONFIRMED (RELATED SVN (B)(4)). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: THIS IS 1 OF 2 MEDWATCH REPORTS REGARDING THIS EVENT. THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. THE DEVICE WILL BE RETURNED FOR ANALYSIS AND FURTHER INFORMATION WILL FOLLOW ONCE THE ANALYSIS HAS BEEN COMPLETED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21715470&pc=QDS | 3002773840-2025-00018 | 2025-02-24 05:00 | Death | BIOFIRE DIAGNOSTICS, LLC | 2025-03-27 04:00 | QDS | FILMARRAY® PNEUMONIA PANEL PLUS | False Negative Result | Insufficient Information | K222601 |  | 1 | Event Description: SUMMARY: (B)(6) FRANCE, REPORTED A POTENTIAL FALSE NEGATIVE KLEBSIELLA OXYTOCA, KLEBSIELLA PNEUMONIAE GROUP, AND CTX-M RESULT ON THE FILMARRAY PNEUMONIA PANEL PLUS AFTER TESTING A PATIENT'S BRONCHOALVEOLAR LAVAGE (BAL) SAMPLE. IT IS UNKNOWN IF THERE WAS ANY IMPACT TO THE PATIENT DUE TO THE FILMARRAY PNEUMONIA PANEL PLUS RESULT. BIOFIRE HAS REQUESTED FURTHER INFORMATION FROM THE CUSTOMER AND IS CURRENTLY INVESTIGATING THIS EVENT.Event Description: SUMMARY: (B)(6) HÔPITAL (FRANCE) REPORTED A CONTROL FAILURE AND A POTENTIAL FALSE NEGATIVE KLEBSIELLA OXYTOCA, KLEBSIELLA PNEUMONIAE GROUP, AND CTX-M RESULT ON THE FILMARRAY PNEUMONIA PANEL PLUS AFTER TESTING A PATIENT'S BRONCHOALVEOLAR LAVAGE (BAL) SAMPLE. IT IS UNKNOWN IF THE PATIENT WAS IMPACTED DUE TO THE FILMARRAY PNEUMONIA PANEL PLUS RESULT. BASED ON THE INFORMATION PROVIDED, THE INVESTIGATION FOUND THAT THE MOST LIKELY CAUSE OF THE CONTROL FAILURE AND THE FALSE NEGATIVE K. OXYTOCA, K. PNEUMONIAE, AND CTX-M RESULTS WAS A REAGENT MANUFACTURING ERROR. Manufacturer Narrative: INVESTIGATION: THE PATIENT WAS A 65-YEAR-OLD MALE. THE CUSTOMER REPORTED THAT THE PATIENT WAS ADMITTED TO THE HOSPITAL ON (B)(6) 2025, WITH DYSPNEA AND HEMOPTYSIS. AT THE TIME OF TESTING, THE PATIENT WAS INTUBATED AND VENTILATED IN THE ICU AND ON TAZOBACTAM. ON (B)(6) 2025, A PATIENT'S BAL SAMPLE WAS TESTED ON FILMARRAY PNEUMONIA PANEL PLUS. THE FILMARRAY PNEUMONIA PANEL PLUS REPORTED ALL ANALYTES AS NOT DETECTED. ON (B)(6) 2025, THE PATIENT'S SAMPLE WAS RETESTED ON THE FILMARRAY PNEUMONIA PANEL PLUS. THE FILMARRAY PNEUMONIA PANEL PLUS REPORTED K. OXYTOCA, K. PNEUMONIAE GROUP, AND CTX-M AS DETECTED. THE CUSTOMER REPORTED CHEST COMPUTED TOMOGRAPHY (CT) SCAN WAS SUGGESTIVE OF CYTOMEGALOVIRUS INFECTION, CLINICAL WORSENING AND WORSENING OF LESIONS. ON AN UNKNOWN DATE, MEROPENEM FOR K.PNEUMONIAE ESBL WAS INTRODUCED. THE CUSTOMER STATED THAT THE PATIENT RAPIDLY IMPROVED ON (B)(6) 2025, WITH EXTUBATION. THE PATIENT DIED ON (B)(6) 2025. BIOFIRE'S INVESTIGATION INTO THIS EVENT IS ONGOING AND FURTHER INFORMATION HAS BEEN REQUESTED FROM THE CUSTOMER. THE FULL INVESTIGATION AND ASSOCIATED CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. CONCLUSION: INVESTIGATION ONGOING. Manufacturer Narrative: INVESTIGATION: THE PATIENT WAS A 65-YEAR-OLD MALE. THE CUSTOMER REPORTED THAT THE PATIENT WAS ADMITTED TO THE HOSPITAL ON (B)(6) 2025, WITH DYSPNEA AND HEMOPTYSIS. AT THE TIME OF TESTING, THE PATIENT WAS INTUBATED AND VENTILATED IN THE INTENSIVE CARE UNIT (ICU) AND ON TAZOBACTAM. ON (B)(6) 2025, A PATIENT'S BAL SAMPLE WAS TESTED ON FILMARRAY PNEUMONIA PANEL PLUS. NO RESULTS WERE OBTAINED DUE TO A CONTROL FAILURE. THE FILMARRAY PNEUMONIA PANEL PLUS WAS REPEATED 2 HOURS LATER. THE FILMARRAY PNEUMONIA PANEL PLUS REPORTED ALL ANALYTES AS NOT DETECTED. ON (B)(6) 2025, THE PATIENT'S SAMPLE WAS RETESTED ON THE FILMARRAY PNEUMONIA PANEL PLUS. THE FILMARRAY PNEUMONIA PANEL PLUS REPORTED K. OXYTOCA, K. PNEUMONIAE GROUP, AND CTX-M AS DETECTED. THE CUSTOMER REPORTED THAT THE NEGATIVE FILMARRAY PNEUMONIA PANEL PLUS RESULTS WERE COMMUNICATED TO THE PHYSICIAN IN THE EVENING OF (B)(6) 2025, AND THE MODIFIED RESULT WAS PROVIDED IN THE EARLY AFTERNOON OF (B)(6) 2025. THE SAME SAMPLE WAS USED FOR ADDITIONAL TESTING. AN UNKNOWN PCR TEST WAS POSITIVE FOR EPSTEIN-BARR VIRUS AND CYTOMEGALOVIRUS (CMV). THE CUSTOMER REPORTED THAT QUANTITATIVE CULTURE IDENTIFIED 3×10² K. PNEUMONIAE ESBL AND 10² STENOTROPHOMONAS MALTOPHILIA. NOTE, THE FILMARRAY PNEUMONIA PANEL PLUS DOES NOT DETECT S. MALTOPHILIA. THE PATIENT'S BLOOD SAMPLE WAS ALSO TESTED USING AN UNKNOWN PCR TEST AND WAS POSITIVE FOR CMV. WHEN ASKED IF THE PATIENT WAS AFFECTED BY THE FILMARRAY PNEUMONIA PANEL PLUS RESULT, THE CUSTOMER COMMUNICATED THE FOLLOWING: THE PATIENT WAS ADMITTED TO THE HOSPITAL ON (B)(6) 2025, WITH DYSPNEA AND HEMOPTYSIS, CHEST COMPUTED TOMOGRAPHY (CT) SCAN WAS SUGGESTIVE OF CYTOMEGALOVIRUS INFECTION, CLINICAL WORSENING AND WORSENING OF LESIONS ON CT SCAN, MEROPENEM FOR K.PNEUMONIAE ESBL WAS INTRODUCED, RAPID IMPROVEMENT WAS SEEN ON (B)(6) 2025, WITH EXTUBATION, AND THE PATIENT DIED ON (B)(6) 2025. IT IS UNKNOWN IF THE PATIENT WAS IMPACTED BY THE FILMARRAY PNEUMONIAE PANEL PLUS RESULT. BIOFIRE REQUESTED CLARIFICATION FROM THE CUSTOMER ON WHETHER THE PATIENT WAS AFFECTED BY THE FILMARRAY PNEUMONIAE PANEL PLUS RESULT, BUT DESPITE MULTIPLE ATTEMPTS, NO FURTHER INFORMATION WAS PROVIDED. CONCLUSION: BASED ON THE INFORMATION PROVIDED, THE INVESTIGATION FOUND THE MOST LIKELY CAUSE OF THE CONTROL FAILURE AND FALSE NEGATIVE K. OXYTOCA, K. PNEUMONIAE, AND CTX-M RESULTS WAS A REAGENT MANUFACTURING ERROR. PRELIMINARY INVESTIGATION INTO KIT LOT # 0979424/ POUCH LOT # 3BLF24 OF THE FILMARRAY PNEUMONIA PANEL PLUS IDENTIFIED INCREASED RISK OF CONTROL FAILURES AND FALSE NEGATIVE TEST RESULTS ON THIS SPECIFIC POUCH LOT. THIS WAS COMMUNICATED TO CUSTOMERS VIA FIELD SAFETY CORRECTIVE ACTION (FSCA) FA-TWD-000033, WHICH WAS ISSUED TO SUBSIDIARIES AND DISTRIBUTORS ON FEBRUARY 28, 2025. BIOFIRE IS CURRENTLY INVESTIGATING THE ISSUE AND TAKING APPROPRIATE CORRECTIVE AND PREVENTIVE ACTIONS. BIOFIRE RECOMMENDS THAT THE CUSTOMER FOLLOW THE INSTRUCTIONS IN FSCA FA-TWD-000033, INCLUDING STOPPING USE AND DISCARDING PRODUCT FROM KIT LOT # 0979424/ POUCH LOT # 3BLF24. CLINICAL PERFORMANCES CAN BE FOUND IN TABLES 22, 24, 56, AND 59 OF THE FILMARRAY PNEUMONIA PANEL PLUS IFU (WWW.ONLINE-IFU.COM/ITI0038).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21715487&pc=OZO | 2032227-2025-144801 | 2024-12-31 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-03-27 04:00 | OZO | 630G INSULIN PUMP MMT-1715K 630G BLACK MG | Adverse Event Without Identified Device or Use Problem | Hypoglycemia; Heart Failure/Congestive Heart Failure | P150001 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY AT CARE CENTER ON (B)(6) 2024. THE CAUSE OF DEATH WAS HYPOGLYCEMIA AND CONGESTED HEART FAILURE. CUSTOMER BLOOD GLUCOSE WAS REPORTED AS 34 MG/DL AT THE TIME OF INCIDENT. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: MMT-7841ZN, MMT-386A, MMT-1884, MMT-1715K, AND MMT-332A. TROUBLESHOOTING WAS PERFORMED FOR DECEASED REPORTING, CUSTOMER WAS NOT WORN INSULIN PUMP AND SENSOR AT THE TIME OF PASSING. INSULIN PUMP AND SENSOR WAS REMOVED WHEN CUSTOMER WAS HOSPITALIZED. THERE IS NO MENTION OF THE AUTO MODE FEATURE AT THE TIME OF THE EVENT. IT WAS AGREED TO RETURN THE INSULIN PUMP. NO PRODUCT RETURN IS REQUIRED FOR MMT-7841ZN. NO PRODUCT RETURN IS REQUIRED FOR MMT-386A. MMT-1884 WAS REQUESTED, AND THE RESPONSE WAS THE DEVICE WILL BE RETURNED. MMT-1715K WAS REQUESTED, AND THE RESPONSE WAS THE DEVICE WILL BE RETURNED. NO PRODUCT RETURN IS REQUIRED FOR MMT-332A. Manufacturer Narrative: THIS IS 2 OF 2 MEDWATCH REPORTS REGARDING THIS EVENT. THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. THE DEVICE WILL BE RETURNED FOR ANALYSIS AND FURTHER INFORMATION WILL FOLLOW ONCE THE ANALYSIS HAS BEEN COMPLETED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION HAS BEEN ADDED WHICH WAS NOT INCLUDED WITH THE INITIAL REPORT. THE INFORMATION HAS BEEN PROVIDED IN SECTION H7 AND H9 WITH THIS REPORT. THE PUMP PASSED THE FUNCTIONAL TEST, INCLUDING THE SELF TEST, SLEEP CURRENT MEASUREMENT, ACTIVE CURRENT MEASUREMENT, REWIND TEST, PRIME/SEATING TEST, BASIC OCCLUSION TEST, OCCLUSION TEST, FORCE SENSOR TEST, DISPLACEMENT TEST, AND THE DAT AT 0.0876 INCHES. THE FOLLOWING WERE NOTED DURING VISUAL INSPECTION: MINOR SCRATCHED DISPLAY WINDOW, MISSING DISPLAY WINDOW COVER, SCRATCHED CASE, CRACKED BATTERY TUBE THREADS, CRACKED CASE AT THE BATTERY TUBE SIDE, SCRATCHED SERIAL NUMBER LABEL, PILLOWING KEYPAD OVERLAY, CRACKED KEYPAD OVERLAY AT THE SELECT BUTTON AND CRACKED RETAINER RING. TEST P-CAP AND RESERVOIR LOCKED PROPERLY INTO RESERVOIR COMPARTMENT DURING TESTING. HISTORY DOWNLOAD WAS SUCCESSFUL USING THUS AND CARELINK UPLOAD WAS SUCCESSFUL. THE PUMP WAS RECEIVED WITHOUT A BATTERY. NO DAMAGE NOTED ON THE BATTERY CAP THE PUMP HISTORY FILE LISTS DATA FROM 09/24/2018 TO 03/03/2021. UNABLE TO VERIFY DAILY TOTAL OF ALL INSULIN DELIVERED, DAILY TOTAL OF BASAL INSULIN DELIVERED AND DAILY TOTAL OF BOLUS INSULIN DELIVERED FOR EVENT DATES DEC 22, 2024, AND 31-DEC-2024 DUE TO NO DATA AVAILABLE IN THE PUMP HISTORY FILE. UNABLE TO PERFORM THE HISTORY REVIEW 1 WEEK PRIOR TO THE EVENT DATES DEC 22, 2024, AND 31-DEC-2024 IN THE PUMP HISTORY FILE DUE TO NO DATA AVAILABLE. PUMP WAS CUT OPEN TO PERFORM VISUAL INSPECTION AND FOUND NO EVIDENCE OF PHYSICAL OR MOISTURE DAMAGE ON THE PCBA1, PCBA2, FORCE SENSOR, MOTOR AND VIBRATOR ASSEMBLY NOTED. FORCE SENSOR ZERO OFFSET WITHIN SPECIFICATION (22.4 MV). THE PUMP PASSED THE FUNCTIONAL TESTING. UNABLE TO CONFIRM ALLEGED LOW BGS. DAMAGE-RETAINER RING DAMAGE CONFIRMED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21716148&pc=OAE | 2029046-2025-00946 | 2025-02-17 05:00 | Death | BIOSENSE WEBSTER INC | 2025-03-27 04:00 | OAE | QDOT MICRO | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P210027 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A PAROXYSMAL ATRIAL FIBRILLATION ABLATION PROCEDURE WITH A QDOT MICRO AND THE PATIENT EXPERIENCED DEATH. PATIENT DIED 2 WEEKS AFTER ABLATION. IT IS UNCLEAR WHETHER THERE IS A CONNECTION BETWEEN THE EVENT AND THE PROCEDURE. THERE IS NO FURTHER INFORMATION. MULTIPLE ATTEMPTS HAVE BEEN MADE TO GAIN CLARIFICATION AND ADDITIONAL INFORMATION ABOUT THIS EVENT WITH NO RESPONSE. SHOULD ANY NEW INFORMATION BE OBTAINED IT WILL BE ASSESSED AND PROCESSED ACCORDINGLY. Manufacturer Narrative: DEVICE INVESTIGATION DETAILS: AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. AN EVALUATION OF THE MANUFACTURING RECORD COULD NOT BE PERFORMED AS THE REQUIRED PRODUCT IDENTIFICATION NUMBER WAS NOT PROVIDED TO COMPLETE THE EVALUATION. HOWEVER, IF THE PRODUCT OR PRODUCT ID NUMBERS ARE RECEIVED AT A LATER DATE, THE INVESTIGATION WILL BE UPDATED AS APPLICABLE. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. NOTE: EXACT DATE OF DEATH WAS NOT PROVIDED. AS SUCH, THE DATE HAS BEEN POPULATED TO MATCH THE EVENT DATE AS BEST GUESS DATE. NOTE: FOR FIELD D4: UDI: AS THE LOT NUMBER FOR THE DEVICE INVOLVED IN THE EVENT WAS NOT PROVIDED, THE FULL UDI IS CURRENTLY NOT AVAILABLE. IF ADDITIONAL INFORMATION IS RECEIVED REGARDING THIS EVENT, A SUPPLEMENTAL 3500A REPORT WILL BE SUBMITTED TO THE FDA. MANUFACTURER'S REF. # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21694889&pc=MEA | 2016493-2025-68779 | 2025-02-26 05:00 | Death | CAREFUSION SD | 2025-03-26 04:00 | MEA | ALARIS SYSTEM | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K211218 |  | 1 | Event Description: IT WAS INITIALLY REPORTED THAT AN EVENT OCCURRED WHILE THE PATIENT WAS ON HYDROMORPHONE INFUSION VIA PCA PUMP. THE EVENT OCCURRED ON (B)(6) 2025 AROUND 1925. PER REPORT, DURING SHIFT CHANGE, THE PATIENT WAS FOUND "CODING". THE PATIENT REPORTEDLY EXPIRED ON (B)(6) 2025 AROUND 1925 WITH UNDETERMINED CAUSE. THE CUSTOMER REACHED OUT TO THE MANUFACTURER SEEKING HELP WITH INTERPRETING THE DEVICE LOGS. BD RECEIVED ADDITIONAL INFORMATION THROUGH DUE DILIGENCE ATTEMPTS. IT WAS REPORTED BY THE FACILITY'S CLINICAL PHARMACY MANAGER (B)(6) THAT THE PATIENT WAS IN THE HOSPITAL DUE TO SICKLE CELL CRISIS. THE PCA MODULE WAS PROGRAMMED AS A PCA DOSE ONLY FOR HYDROMORPHONE. THE PATIENT "SUSTAINED CARDIAC ARREST (DUE TO) COMPLICATION FROM SICKLE CELL DISEASE." THE PRIMARY CAUSE OF DEATH WAS SICKLE CELL DISEASE, PER REPORT.Event Description: IT WAS INITIALLY REPORTED THAT AN EVENT OCCURRED WHILE THE PATIENT WAS ON HYDROMORPHONE INFUSION VIA PCA PUMP. THE EVENT OCCURRED ON 26 FEBRUARY 2025 AROUND 1925. PER REPORT, DURING SHIFT CHANGE, THE PATIENT WAS FOUND "CODING". THE PATIENT REPORTEDLY EXPIRED ON (B)(6) 2025 AROUND 1925 WITH UNDETERMINED CAUSE. THE CUSTOMER REACHED OUT TO THE MANUFACTURER SEEKING HELP WITH INTERPRETING THE DEVICE LOGS. BD RECEIVED ADDITIONAL INFORMATION THROUGH DUE DILIGENCE ATTEMPTS. IT WAS REPORTED BY THE FACILITY'S CLINICAL PHARMACY MANAGER (PHARMD) THAT THE PATIENT WAS IN THE HOSPITAL DUE TO SICKLE CELL CRISIS. THE PCA MODULE WAS PROGRAMMED AS A PCA DOSE ONLY FOR HYDROMORPHONE. THE PATIENT "SUSTAINED CARDIAC ARREST (DUE TO) COMPLICATION FROM SICKLE CELL DISEASE." THE PRIMARY CAUSE OF DEATH WAS SICKLE CELL DISEASE, PER REPORT.Event Description: IT WAS INITIALLY REPORTED THAT AN EVENT OCCURRED WHILE THE PATIENT WAS ON HYDROMORPHONE INFUSION VIA PCA PUMP. THE EVENT OCCURRED ON (B)(6) 2025 AROUND 1925. PER REPORT, DURING SHIFT CHANGE, THE PATIENT WAS FOUND "CODING". THE PATIENT REPORTEDLY EXPIRED ON (B)(6) 2025 AROUND 1925 WITH UNDETERMINED CAUSE. THE CUSTOMER REACHED OUT TO THE MANUFACTURER SEEKING HELP WITH INTERPRETING THE DEVICE LOGS. BD RECEIVED ADDITIONAL INFORMATION THROUGH DUE DILIGENCE ATTEMPTS. IT WAS REPORTED BY THE FACILITY'S CLINICAL PHARMACY MANAGER (PHARMD) THAT THE PATIENT WAS IN THE HOSPITAL DUE TO SICKLE CELL CRISIS. THE PCA MODULE WAS PROGRAMMED AS A PCA DOSE ONLY FOR HYDROMORPHONE. THE PATIENT "SUSTAINED CARDIAC ARREST (DUE TO) COMPLICATION FROM SICKLE CELL DISEASE." THE PRIMARY CAUSE OF DEATH WAS SICKLE CELL DISEASE, PER REPORT. DURING AN ONSITE VISIT, A BD CLINICAL PRACTICE CONSULTANT DISCUSSED THE EVENT WITH THE HOSPITAL CLINICIAN. THE CLINICIAN INDICATED: ENTIRE ALARIS SYSTEM SETUP WAS SEQUESTERED. EVENT OCCURRED ON A SICKLE CELL PATIENT AND A CODE BLUE WAS CALLED. THE AUTOPSY OF THE PATIENT SHOWED THE PATIENT HAD A CLOT AND HAD NOTHING TO DO WITH THE PCA SETTINGS THAT HAD BEEN CHANGED DURING THE DAY. THE CHANGE WAS A LOW CONTINUOUS DOSE TO A PCA DOSE. THE PATIENT HAD NEVER REQUESTED A DOSE AFTER THE CHANGE. Manufacturer Narrative: CORRECTION: DESCRIBE EVENT OR PROBLEM, MEDICAL DEVICE TYPE AND MANUFACTURER NARRATIVE. INVESTIGATION SUMMARY: THE REPORT OF A DEATH DURING A HYDROMORPHONE INFUSION WAS NOT ABLE TO BE CONFIRMED OR REPLICATED. INSPECTION, LOG REVIEW AND LABORATORY TESTING COULD NOT BE PERFORMED BECAUSE THE DEVICES AND THEIR ASSOCIATED LOGS WERE NOT RETURNED FOR INVESTIGATION DESPITE MULTIPLE ATTEMPTS TO OBTAIN THEM. THE FACILITY WAS UNABLE TO LOCATE THE DEVICES. THE FACILITY SENT AN INFUSION KNOWLEDGE PORTAL (IKP) REPORT OF THE INFUSIONS MADE ON THE DATE OF THE EVENT AT FROM 15:06 (3:06 PM) TO 19:07 (19:07 PM) USING GUARDRAILS DRUG ¿HYDROMORPHONE¿ UNDER THE PROFILE ¿.ADULT¿. THE LAST INFUSION PROGRAMMED RECORDED HAD A CONCENTRATION OF 1MG/ML, DOSE OF 1MG AND A DILUENT VOLUME OF 1ML. IT WAS OBSERVED THE INFUSION WAS COMPLETED 14,432 SECONDS (4 HOURS) LATER AT 19:07 (7:07PM). NO OTHER INFUSIONS WERE REPORTED ON THE DATE OF THE EVENT THE IKP REPORT ONLY PROVIDES MINIMAL INFUSION INFORMATION AND IS NOT VALIDATED FOR LOG ANALYSIS PURPOSES. THE DATA SET (NORTON ADULT 0125) WAS ALSO PROVIDED BY THE FACILITY; HOWEVER, WITHOUT THE SUSPECT AND CONCOMITANT LOGS IT¿S NOT POSSIBLE TO ASCERTAIN THE EVENT DETAILS ASSOCIATED TO THE REPORTED COMPLAINT. PROPER OPERATION OF THE BD ALARIS¿ SYSTEM REQUIRES THAT YOU ARE FAMILIAR WITH RELATED FEATURES, SETUP, PROGRAMMING, IV SETS, AND ACCESSORIES. READ ALL INSTRUCTIONS, INCLUDING THOSE FOR ALL ATTACHED MODULE(S) BEFORE USING THE BD ALARIS ¿ SYSTEM. THE DEVICE WAS IN USE FOR TREATMENT PURPOSES AS INTENDED PER 21 CFR 820.198(D)(2). PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER. Manufacturer Narrative: A FOLLOW UP REPORT WILL BE SUBMITTED WITH INVESTIGATION RESULTS SHOULD THE DEVICE BE REPAIRED OR THE DEVICE/LOGS BE RECEIVED FOR EVALUATION. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER. Manufacturer Narrative: OMIT : C20 - NO FINDINGS AVAILABLE, D15 - CAUSE NOT ESTABLISHED. ADDITIONAL INFORMATION : IMDRF ANNEX B, C, D CODES AND MANUFACTURER NARRATIVE. A DEVICE HISTORY RECORD, COMPLAINT HISTORY REVIEW, AND RISK REVIEW ON FAILURE MODES ARE PERFORMED ON EACH DEVICE WITH REPORTABLE MALFUNCTION(S) ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN SECTION H6 OF THIS MDR REPORT. INVESTIGATION SUMMARY: THE REPORT OF A DEATH DURING A HYDROMORPHONE INFUSION WAS NOT ABLE TO BE CONFIRMED OR REPLICATED. INSPECTION, LOG REVIEW AND LABORATORY TESTING COULD NOT BE PERFORMED BECAUSE THE DEVICES AND THEIR ASSOCIATED LOGS WERE NOT RETURNED FOR INVESTIGATION DESPITE MULTIPLE ATTEMPTS TO OBTAIN THEM. THE FACILITY WAS UNABLE TO LOCATE THE DEVICES. THE FACILITY SENT AN INFUSION KNOWLEDGE PORTAL (IKP) REPORT OF THE INFUSIONS MADE ON THE DATE OF THE EVENT AT FROM 15:06 (3:06 PM) TO 19:07 (19:07 PM) USING GUARDRAILS DRUG ¿HYDROMORPHONE¿ UNDER THE PROFILE ¿.ADULT¿. THE LAST INFUSION PROGRAMMED RECORDED HAD A CONCENTRATION OF 1MG/ML, DOSE OF 1MG AND A DILUENT VOLUME OF 1ML. IT WAS OBSERVED THE INFUSION WAS COMPLETED 14,432 SECONDS (4 HOURS) LATER AT 19:07 (7:07PM). NO OTHER INFUSIONS WERE REPORTED ON THE DATE OF THE EVENT THE IKP REPORT ONLY PROVIDES MINIMAL INFUSION INFORMATION AND IS NOT VALIDATED FOR LOG ANALYSIS PURPOSES. THE DATA SET (NORTON ADULT 0125) WAS ALSO PROVIDED BY THE FACILITY; HOWEVER, WITHOUT THE SUSPECT AND CONCOMITANT LOGS IT¿S NOT POSSIBLE TO ASCERTAIN THE EVENT DETAILS ASSOCIATED TO THE REPORTED COMPLAINT. ROOT CAUSE: THE ROOT CAUSE OF THE REPORT OF A DEATH DURING A HYDROMORPHONE INFUSION WAS NOT IDENTIFIED. INSPECTION AND LABORATORY TESTING OF THE DEVICES COULD NOT BE PERFORMED BECAUSE THEY WERE NOT RETURNED FOR INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21695303&pc=OZP | 2032227-2025-143146 | 2025-02-16 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-03-26 04:00 | OZP | MINIMED 780G US SYSTEM BLE CONNECT 3.0 MG/DL | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P160017 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY AT THE HOSPITAL ON (B)(6) 2025 DUE TO DIABETES COMPLICATION. THE CUSTOMER WAS ADMITTED TO A HOSPITAL ON (B)(6) 2025 DUE TO DIABETES COMPLICATION (WOUND). THE EVENT INVOLVED PRODUCT(S) MMT-1884, MMT-332A, UNOMEDICAL. THE CUSTOMER¿S BLOOD GLUCOSE WAS UNKNOWN. THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP AT THE TIME OF DEATH. THE INSULIN PUMP WAS LAST WORN ON (B)(6) 2025. NO PRODUCT RETURN IS REQUIRED FOR MMT-1884. NO PRODUCT RETURN IS REQUIRED FOR MMT-332A. NO PRODUCT RETURN IS REQUIRED FOR UNOMEDICAL. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21695405&pc=OZO | 2032227-2025-143193 | 2025-02-22 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-03-26 04:00 | OZO | 630G INSULIN PUMP MMT-1715KL 630G | Adverse Event Without Identified Device or Use Problem | Hyperglycemia; Heart Failure/Congestive Heart Failure | P150001 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2025 AT HOSPITAL. THE CAUSE OF DEATH WAS CUSTOMER FELL, POSSIBLE HEART ATTACK, HYPERGLYCEMIA AND PHENOMENA. THE CUSTOMER BLOOD GLUCOSE WAS MENTIONED AS 700 MG/DL AT THE TIME OF INCIDENT. THE INSULIN PUMP WAS NOT WORN AT THE TIME OF PASSING. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: MMT-397A, MMT-332A, MMT-1715KL, MMT-243A. TROUBLESHOOTING WAS PERFORMED FOR DECEASED REPORTING. CUSTOMER WAS NOT USING THE INSULIN PUMP SYSTEM WITHIN 48HRS OF REPORTED EVENT AND CUSTOMER WAS NOT USING SENSORS. NO PRODUCT RETURN IS REQUIRED FOR MMT-397A. NO PRODUCT RETURN IS REQUIRED FOR MMT-332A. PRODUCT RETURN REQUESTED FOR MMT-1715KL. CUSTOMER DECLINED TO RETURN OR IS UNABLE TO RETURN. NO PRODUCT RETURN IS REQUIRED FOR MMT-243A. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21696243&pc=OZD | 1220648-2025-26963 | 2025-02-26 05:00 | Death | ABIOMED, INC. | 2025-03-26 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemolysis; Hematuria | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT AND DEVELOPED HEMOLYSIS THE FOLLOWING DAY POST IMPLANT AND WOULD SUBSEQUENTLY EXPIRE. THE PATIENT PRESENTED WITH DIFFICULTY BREATHING AND SHORTNESS OF BREATH AND WAS FOUND TO BE IN ATRIAL FIBRILLATION. CATHETERIZATION WAS COMPLETED AND REVEALED 90 PERCENT OCCLUSION OF THE LEFT ANTERIOR DESCENDING, WHICH WAS TREATED WITH BALLOON ANGIOPLASTY AND STENTING. THE DECISION WAS MADE TO PLACE AN IMPELLA CP POSTOPERATIVELY AND CONTINUE WITH TRANSFER OUT TO A HIGHER LEVEL OF CARE FACILITY. AFTER THE PUMP WAS PLACED THE FOLLOWING DAY THERE WERE SIGNS OF HEMOLYSIS. THE PATIENT HAD A TRAUMATIC FOLEY AT THE FIRST HOSPITAL BEFORE ARRIVING TO THE TERTIARY CENTER. THE SIGNS OF HEMOLYSIS OBSERVED WERE HEMOLYZED LABS AND ELEVATED PLASMA HEMOGLOBIN. THE PERFORMANCE LEVEL FOR SUPPORT WAS DECREASED. TWO DAYS LATER, SEDATION WAS TURNED OFF. THE PATIENT WAS NOT RESPONDING, AND STATUS REMAINED UNCHANGED NOW FOUR DAYS LATER AFTER START OF MCS (STILL NO PATIENT RESPONSE). HEMOLYSIS WAS AGAIN NOTED. THE PERFORMANCE LEVEL HAD BEEN INCREASED, AND URINE AND SPUTUM BECAME BLOODY. THE PERFORMANCE LEVEL WAS REDUCED, TO MANAGE THE COMPLICATION. HOWEVER, CARE WAS WITHDRAWN, AND THE PATIENT WOULD EXPIRE THIS SAME DAY. ALTHOUGH CARE WAS WITHDRAWN THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA CP AS AN ASSOCIATED FACTOR THE EVENT. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM, SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER: ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE: ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS: ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION: ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿ Manufacturer Narrative: THE INVESTIGATION OF HEMOLYSIS HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21696264&pc=KRD | 2124215-2025-17223 | 2024-02-09 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-03-26 04:00 | KRD | OBSIDIO | Migration | Abdominal Pain; Ischemia; Necrosis; Sepsis; Multiple Organ Failure; Embolism/Embolus; Pancreatitis | K213385 |  | 1 | Event Description: IT WAS REPORTED VIA JOURNAL ARTICLE, THAT MIGRATION OCCURRED RESULTING IN PATIENT COMPLICATIONS AND DEATH. OBSIDIO WAS SELECTED FOR AN EMBOLIZATION PROCEDURE OF THE POSTERIOR SUPERIOR PANCREATICODUODENAL ARTERY (PSPDA). THE PATIENT WAS ADMITTED FOR HYPOXEMIC RESPIRATORY FAILURE AND SWEETS SYNDROME, AND THE HOSPITAL COURSE WAS COMPLICATED BY MELENA. THE PATIENTS HEMOGLOBIN DECLINED OVERNIGHT FROM 7.1 TO 5.8 G/DL WITH RECURRENT MELENA AND HEMATEMESIS REQUIRING A BLOOD TRANSFUSION. A COMPUTERIZED TOMOGRAPHY (CT) ANGIOGRAM DEMONSTRATED ACTIVE EXTRAVASATION INTO THE DUODENUM. ANGIOGRAPHY OF THE PSPDA WITH A 2.8 FRENCH NON-BOSTON SCIENTIFIC MICROCATHETER DEMONSTRATED A PSEUDOANEURYSM OF A SMALL BRANCH WHERE SUPER SELECTIVE CATHETERIZATION WAS IMPOSSIBLE. EMBOLIZATION OF THE PSPDA USING OBSIDIO WAS PERFORMED, WITH THE GOAL OF OCCLUDING BRANCHES BEYOND THE REACH OF THE MICROCATHETER. THE ALIQUOT TECHNIQUE WAS UTILIZED TO PRESERVE THE MICROCATHETER FOR LATER USE. WHILE APPLYING A SLOW INJECTION FORCE, 0.2 ML ALIQUOTS OF OBSIDIO WAS FLUSHED WITH A 3 ML SYRINGE AND IOHEXOL, WHICH WAS SURMISED TO BETTER RESEMBLE THE VISCOSITY OF OBSIDIO COMPARED TO NORMAL SALINE. THE EMBOLIC WAS ADMINISTERED UNTIL OPACIFICATION OF THE PSEUDOANEURYSM WAS ACHIEVED, DELIVERING A TOTAL OF 0.4 ML. FLUOROSCOPIC IMAGES DEMONSTRATED THE EMBOLIC DEPOSITED IN THE PSEUDOANEURYSM AND DISTAL BRANCHES MEASURING 0.2 TO 0.3 MM IN CALIBER, WHILE FORMING A CAST OF THE PARENT VESSEL. ANGIOGRAPHY OF THE GASTRODUODENAL ARTERY (GDA) DEMONSTRATED CONTINUED EXTRAVASATION FROM A SEPARATE BRANCH. OBSIDIO WAS THEREFORE ADMINISTERED WITH A MORE FORCEFUL INJECTION, WITH THE INTENTION OF ACHIEVING GDA OCCLUSION WITHOUT ADDITIONAL DISTAL EMBOLIZATION, USING 0.2 ML ALIQUOTS FOR A TOTAL OF 0.6 ML. ALTHOUGH DISTAL PENETRATION WITH SUBSEQUENT FLUSHING WAS NOT DIRECTLY OBSERVED, THE EMBOLIC PENETRATED INTO THE RIGHT GASTROEPIPLOIC ARTERY AND OTHER TERMINAL BRANCHES MEASURING 0.4 MM; NO REFLUX WAS APPRECIATED. THE PATIENTS POST PROCEDURE COURSE WAS COMPLICATED BY ABDOMINAL PAIN AND WHITE BLOOD CELL COUNT OF 52,000/MICROLITERS, LIPASE OF 240 U/L, AND LACTATE DEHYDROGENASE OF 601 U/L. A CT OF THE ABDOMEN/PELVIS OBTAINED ON POST PROCEDURE DAY 3, SHOWED REGIONAL NONENHANCEMENT OF THE DUODENUM AND PANCREATIC HEAD. THE PATIENT WAS EMERGENTLY TAKEN FOR EXPLORATORY LAPAROTOMY WHERE COMPLETE NECROSIS OF THE SECOND PORTION OF THE DUODENUM AND SEVERE PANCREATITIS OF THE PANCREATIC HEAD WERE OBSERVED. THE NECROTIC DUODENUM WAS RESECTED. THE PATIENT REMAINED IN CRITICAL CONDITION WITH THE POSTOPERATIVE COURSE COMPLICATED BY SEPSIS AND MULTI-ORGAN FAILURE. THE PATIENT WAS TRANSITIONED TO COMFORT CARE AND SUCCUMBED 18 DAYS AFTER EMBOLIZATION. Manufacturer Narrative: B3: DATE OF EVENT: THE PROCEDURE DATE WAS BETWEEN 9TH AND 12TH FEBRUARY 2024. JASON S. KIM, MD, R. PETER LOKKEN, MD, MPH, ALEXANDER LAM, MD, NESLIHAN KAYRAKLIOGLU, MD, PHD, K. PALLAV KOLLI, MD. "GASTRODUODENAL ARTERY EMBOLIZATION WITH SHEAR-THINNING CONFORMABLE EMBOLIC RESULTING IN PANCREATIC AND DUODENAL NECROSIS". JOURNAL OF VASCULAR AND INTERVENTIONAL RADIOLOGY, 1- 17
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21696763&pc=OUT | 2029214-2025-00772 | 2024-06-27 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-03-26 04:00 | OUT | PIPELINE | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Intracranial Hemorrhage |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING: ENDOVASCULAR TREATMENT OF COMPLEX VASCULAR PATHOLOGIES IN THE PEDIATRIC POPULATION, OFTEN PERFORMED BY NONPEDIATRIC SUBSPECIALISTS WITH ADAPTATION OF EQUIPMENT AND TECHNIQUES DEVELOPED FOR ADULT PATIENTS. THE STUDY AIMED TO REPORT ON THE CENTER¿S EXPERIENCE WITH SAFETY AND OUTCOMES OF ENDOVASCULAR TREATMENTS FOR PEDIATRIC VASCULAR PATHOLOGIES. THE TIME FRAME OF THIS STUDY WAS: JANUARY 1, 2004, TO DECEMBER 1, 2022. MULTIPLE MANUFACTURER¿S DEVICES WERE USED IN THE STUDY POPULATION THE FOLLOWING MEDTRONIC DEVICES WERE USED: PIPELINE EMBOLIZATION DEVICE (MEDTRONIC, DUBLIN, IRELAND), ONYX (MEDTRONIC, DUBLIN, IRELAND), REACT (MEDTRONIC, DUBLIN, IRELAND), SOLITAIRE (MEDTRONIC, DUBLIN, IRELAND), AXIUM COILS (MEDTRONIC, DUBLIN, IRELAND) DEATHS OCCURRED IN THE STUDY POPULATION THE CAUSES OF DEATH WERE: RUPTURED ANEURYSM WITH FISHER GRADE IV SUBARACHNOID HEMORRHAGE (ONE DEATHS), HEMODYNAMIC FAILURE SECONDARY TO THE DISEASE (ONE DEATH), INTRAVENTRICULAR HEMORRHAGE AFTER EMBOLIZATION (ONE DEATH) AMONG PATIENT ADVERSE EVENTS INCLUDED: ACCESS-SITE COMPLICATIONS: 2 PATIENTS (1.7%), INTRAPROCEDURAL COMPLICATIONS: 3 PATIENTS (2.5%) , TRANSIENT NEUROLOGICAL DEFICITS: 2 PATIENTS (1.7%), SEVERE IN-STENT STENOSIS OBSERVED IN 1 PATIENT (12.5%) AFTER FLOW-DIVERTER TR EATMENT, TREATED WITH ANGIOPLASTY, LEFT VERTEBRAL ARTERY DISSECTION TREATED WITH A PIPELINE EMBOLIZATION DEVICE IN 1 PATIENT (4.8%), VESSEL PERFORATION IN ANOTHER PATIENT (4.8%), POSTOPERATIVE TRANSIENT UNILATERAL NEUROLOGICAL DEFICITS IN 2 PATIENTS (9.5%), MILD L EFT-SIDED WEAKNESS IN 1 PATIENT, RESIDUAL NIDUS IN 2 PATIENTS, ALLERGIC REACTION TO CONTRAST MATERIAL IN 1 PATIENT NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID: NV UNK PIPELINE (LOT: UNKNOWN); IMPLANT DATE: N/A; EXPLANT DATE: N/A; PRODUCT ID: NV UNK PIPELINE (UNKNOWN); IMPLANT DATE: N/A; EXPLANT DATE: N/A. G2: CITATION: AUTHORS: BRIANNA M. DONNELLY, ANDRE MONTEIRO, MATTHEW J. RECKER, JAIMS LIM, PUI MAN ROSALIND LAI, WADY T. JACOBY, WASIQ I. KHAWAR, ALEXANDER B. BECKER, MUHAMMAD WAQAS, JUSTIN M. CAPPUZZO, JASON M. DAVIES, KENN. ENDOVASCULAR TREATMENT FOR COMPLEX VASCULAR PATHOLOGIES IN THE PEDIATRIC POPULATION: EXPERIENCE FROM A CENTER WITH DUAL-TRAINED NEUROSURGEONS. WORLD NEUROSURG 189: E696-E708 2024. DOI: 10.1016/J.WNEU.2024.06.151 EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21696788&pc=OZP | 2032227-2025-143545 | 2009-01-01 05:00 | Death | MEDTRONIC MINIMED | 2025-03-26 04:00 | OZP | UNKNOWN NGP | Excess Flow or Over-Infusion | Insufficient Information |  |  | 1 | Event Description: INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2009 DUE TO OVER-DELIVERY OF INSULIN FROM THE PUMP. THE CAUSE OF DEATH WAS UNKNOWN. THE CUSTOMER¿S BLOOD GLUCOSE WAS UNKNOWN. IT WAS NOT KNOWN IF THE PATIENT WAS WEARING THE PUMP AT THE TIME OF PASSING. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: UNK_NGP, MMT-397A, AND MMT-332A. TROUBLESHOOTING WAS NOT PERFORMED. UNK_NGP WAS REQUESTED AND THE RESPONSE WAS THE DEVICE WILL BE RETURNED. MMT-397A WAS REQUESTED AND THE RESPONSE WAS THE DEVICE WILL BE RETURNED. MMT-332A WAS REQUESTED AND THE RESPONSE WAS THE DEVICE WILL BE RETURNED. Manufacturer Narrative: CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. THE DEVICE WILL BE RETURNED FOR ANALYSIS AND FURTHER INFORMATION WILL FOLLOW ONCE THE ANALYSIS HAS BEEN COMPLETED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21696838&pc=QFG | 3013756811-2025-45659 | 2025-02-10 05:00 | Death | TANDEM DIABETES CARE | 2025-03-26 04:00 | QFG | TANDEM MOBI INSULIN PUMP WITH INTEROPERABLE TECHNOLOGY | Adverse Event Without Identified Device or Use Problem | Nausea; Vomiting; Diabetic Ketoacidosis | K233044 |  | 1 | Event Description: IT WAS REPORTED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2025. PER THE CONTACT, THE CAUSE OF DEATH WAS DIABETIC KETOACIDOSIS. PRIOR TO THE DEATH, THE CUSTOMER EXPERIENCED NAUSEA AND VOMITING. A REVIEW OF PUMP DATA WILL BE PERFORMED, AND THE RESULTS WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. Manufacturer Narrative: NO PRODUCT WAS RETURNED FOR EVALUATION. SHOULD NEW RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. H3 OTHER TEXT : DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21696916&pc=OZD | 1220648-2025-26967 | 2025-02-26 05:00 | Death | ABIOMED, INC. | 2025-03-26 04:00 | OZD | IMPELLA CP | Restricted Flow rate; Malposition of Device | Cardiac Arrest | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT (UNKNOWN RACE AND ETHNICITY) WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS) AND EXPERIENCED LOW PUMP FLOW AND DEVICE MALPOSITIONING. THE PATIENT WENT INTO CARDIAC ARREST AND EXPIRED. THE PATIENT PRESENTED TO HOSPITAL WITH FOUR-DAY HISTORY OF SIGNIFICANT CHEST PAIN AND FOUND HAVING AN ACUTE MYOCARDIAL INFARCTION. THE PATIENT WAS TAKEN TO THE CATHETERIZATION FOUND AN OCCLUDED LEFT ANTERIOR DESCENDING ARTERY, RIGHT CORONARY ARTERY AND 80% OCCLUDED LEFT CIRCUMFLEX ARTERY WITH EXTENSIVE LEFT VENTRICULAR WALL ISCHEMIA AND DECREASED RIGHT VENTRICULAR FUNCTION. DURING THE PERCUTANEOUS CORONARY INTERVENTION, THE PATIENT STARTED EXHIBITING SHOCK SIGNS AND THE DECISION MADE TO IMPLANT AN IMPELLA CP DEVICE. THE IMPLANT WAS UNCOMPLICATED, AND THE PATIENT WAS TAKEN TO THE UNIT ON MCS. OVER THE NEXT SIX HOURS THE PATIENTS CLINICAL STATE DETERIORATED, NEEDING FLUID RESUSCITATION, INOTROPIC SUPPORT. THE PATIENT WAS DETERMINED TO BE UNSUITABLE FOR ECMO, LVAD OR HEART TRANSPLANT GIVEN THE PRESENTATION AND CLINICAL STATE. THE PATIENT HAD INCREASING SUCTION ALARMS DUE TO VOLUME AND RIGHT VENTRICULAR FUNCTION AND AS THE PATIENT WORSENED. THE PATIENT HAD AN EPISODE OF HEMATEMESIS RESULTING IN EXCESSIVE MOVEMENT AND MALPOSITIONING OF THE IMPELLA CP DEVICE INTO THE VENTRICLE. ATTEMPTS WERE MADE TO REPOSITION IMPELLA CP PUMP UNDER ECHOCARDIOGRAM; HOWEVER, THE PATIENT SUBSEQUENTLY ARRESTED WITH A RHYTHM OF ASYSTOLE. AFTER 30 MINUTES OF RESUSCITATION ATTEMPTS INCLUDING RSI, THE PATIENT HAD NEVER REGAINED A HEART RHYTHM AND DECISION WAS MADE TO WITHDRAW CARE. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. "ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21697038&pc=NVZ | 2182208-2025-01512 | 2024-01-01 05:00 | Death | MEDTRONIC, INC. | 2025-03-26 04:00 | NVZ | MEDTRONIC IPG | Appropriate Term/Code Not Available | Heart Failure/Congestive Heart Failure |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING DELIVERY CATHETER-BASED AND STYLET-BASED RIGHT VENTRICULAR SEPTAL PACING. THE AUTHORS DESCRIBED PATIENT DEATHS; THE CAUSES OF DEATH WERE HEART FAILURE. THERE IS NO ALLEGATION OF DEVICE-DEATH RELATEDNESS INDICATED IN THE ARTICLE. THERE WERE PATIENTS WHO DEVELOPED PACING INDUCED CARDIOMYOPATHY, NEW ONSET ATRIAL FIBRILLATION (AF), AND PATIENTS WHO WERE HOSPITALIZED DUE TO HEART FAILURE. THE STATUS OF THE DEVICES AND LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: MID-TERM OUTCOMES OF DELIVERY CATHETER-BASED AND STYLET-BASED RIGHT VENTRICULAR SEPTAL PACING: FOLLOW-UP RESULTS FROM A MULTICENTER, PROSPECTIVE, RANDOMIZED STUDY. JOURNAL OF ARRHYTHMIA. 2024; 00:1¿9. DOI: 10.1002/JOA3.13034. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21697058&pc=NVN | 2182208-2025-01513 | 2024-01-01 05:00 | Death | MEDTRONIC, INC. | 2025-03-26 04:00 | NVN | SELECTSECURE MRI SURESCAN | Appropriate Term/Code Not Available | Heart Failure/Congestive Heart Failure | P030036 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING DELIVERY CATHETER-BASED AND STYLET-BASED RIGHT VENTRICULAR SEPTAL PACING. THE AUTHORS DESCRIBED PATIENT DEATHS; THE CAUSES OF DEATH WERE HEART FAILURE. THERE IS NO ALLEGATION OF DEVICE-DEATH RELATEDNESS INDICATED IN THE ARTICLE. THERE WERE PATIENTS WHO DEVELOPED PACING INDUCED CARDIOMYOPATHY, NEW ONSET ATRIAL FIBRILLATION (AF), AND PATIENTS WHO WERE HOSPITALIZED DUE TO HEART FAILURE. THE STATUS OF THE DEVICES AND LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: MID-TERM OUTCOMES OF DELIVERY CATHETER-BASED AND STYLET-BASED RIGHT VENTRICULAR SEPTAL PACING: FOLLOW-UP RESULTS FROM A MULTICENTER, PROSPECTIVE, RANDOMIZED STUDY. JOURNAL OF ARRHYTHMIA. 2024; 00:1¿9. DOI: 10.1002/JOA3.13034. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21697110&pc=DXY | 2182208-2025-01514 | 2024-01-01 05:00 | Death | MEDTRONIC, INC. | 2025-03-26 04:00 | DXY | MEDTRONIC LEAD | Appropriate Term/Code Not Available | Heart Failure/Congestive Heart Failure |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING DELIVERY CATHETER-BASED AND STYLET-BASED RIGHT VENTRICULAR SEPTAL PACING. THE AUTHORS DESCRIBED PATIENT DEATHS; THE CAUSES OF DEATH WERE HEART FAILURE. THERE IS NO ALLEGATION OF DEVICE-DEATH RELATEDNESS INDICATED IN THE ARTICLE. THERE WERE PATIENTS WHO DEVELOPED PACING INDUCED CARDIOMYOPATHY, NEW ONSET ATRIAL FIBRILLATION (AF), AND PATIENTS WHO WERE HOSPITALIZED DUE TO HEART FAILURE. THE STATUS OF THE DEVICES AND LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: MID-TERM OUTCOMES OF DELIVERY CATHETER-BASED AND STYLET-BASED RIGHT VENTRICULAR SEPTAL PACING: FOLLOW-UP RESULTS FROM A MULTICENTER, PROSPECTIVE, RANDOMIZED STUDY. JOURNAL OF ARRHYTHMIA. 2024; 00:1¿9. DOI: 10.1002/JOA3.13034. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21697256&pc=QFG | 3013756811-2025-44605 | 2024-02-29 05:00 | Death | TANDEM DIABETES CARE | 2025-03-26 04:00 | QFG | T:SLIM X2 INSULIN PUMP WITH INTEROPERABLE TECHNOLOGY (WITH T:CONNECT MOBILE APP) | Adverse Event Without Identified Device or Use Problem | Diabetic Ketoacidosis | K203234 |  | 1 | Event Description: IT WAS REPORTED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2024. PER THE CONTACT, THE CAUSE OF DEATH WAS DIABETIC KETOACIDOSIS. A REVIEW OF PUMP DATA WILL BE PERFORMED, AND THE RESULTS WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. Manufacturer Narrative: NO PRODUCT WAS RETURNED FOR EVALUATION. SHOULD NEW RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21697430&pc=OZD | 1220648-2025-26970 | 2025-02-28 05:00 | Death | ABIOMED, INC. | 2025-03-26 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A 64 YEAR OLD FEMALE PATIENT WITH AN IMPELLA CP BEING INSERTED FOR SUPPORT DURING CARDIOGENIC SHOCK, LIVER SHOCK, AND SEPTIC SHOCK, WAS FOUND TO HAVE NO DISTAL FLOW TO THE LEG WITH THE SHEATH. EXTERNAL BYPASS WITH 5FR ANTEGRADE SHEATH WAS PLACED AND THE LEG WAS WARM WITH GOOD COLOR. IT WAS FURTHER REPORTED THAT CARE WAS WITHDRAWN AND PATIENT EXPIRED WHILE ON CP SUPPORT. NO ADDITIONAL DETAILS WERE PROVIDED ABOUT THE CAUSE OF DEATH AS IT RELATES TO THE DEVICE. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMART ASSIST SYSTEM STATES THE FOLLOWING: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21697756&pc=OZD | 1220648-2025-26971 | 2025-03-03 05:00 | Death | ABIOMED, INC | 2025-03-26 04:00 | OZD | AUTOMATED IMPELLA CONTROLLER | Application Program Problem | Insufficient Information | P170011 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA RP FLEX AND IMPELLA CP BIPELLA DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. THE PUMP WAS PLACED AND HAD AN AUTOMATED IMPELLA CONTROLLER (AIC) ALARM FOR OPTICAL SIGNAL FOLLOWED BY A YELLOW CONTROLLER ALARM. AFTER 96 HOURS OF SUPPORT, THE DECISION WAS MADE BY THE FAMILY TO WITHDRAW CARE AND THE PATIENT EXPIRED. IT IS UNKNOWN IF THE USE OF THE IMPELLA WAS RELATED TO PATIENT DEATH, THUS THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: G.4 REVISED PMA/510(K) NUMBER AS IT WAS PREVIOUSLY ENTERED INCORRECTLY ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26971. Manufacturer Narrative: THE INVESTIGATION OF AUTOMATED IMPELLA CONTROLLER (AIC) - CONTROLLER ERROR (YELLOW ALARM) AND AIC - LOSS OF OPM DATA HAS BEEN COMPLETED. DATA ANALYSIS: ON 3/2 RP PUMP WITH CONSOLE IC5202, IMMEDIATELY UNPLUGGED, THEN PLUGGED BACK IN. AFTER THE PUMP WAS PLUGGED BACK IN, ON 3/3, THE LOGS SHOW TWO COMMANDS "PROCESSSAMPLINGDATAFROMOPTICALBENCH: SENTSTOPACQUISITIONCMD ACK" THEN "CLEARALLOPTICALDATASHAREDMEMORY: 0X4000". THIS RESULTS IN "PROCESSDUALSENSORSAMPLEFOROPMCOMMSTOPPEDERROR: CALCULATOR PLACEMENT SIGNAL 1058" WHICH IS INDICATIVE OF THE FAILURE MODE LOSS OF OPTICAL DATA DUE TO A SOFTWARE DEFECT. THE REAL TIME (RT) LOGS CONFIRMED NO OPTICAL DATA AT THIS TIME. THE PUMP STAYED ON THIS CONSOLE UNTIL THE END OF THE CASE ON 3/6 DESPITE THE LACK OF OPTICAL SIGNAL. DEVICE ANALYSIS: THE CONSOLE WAS TESTED ON AN ENGINEERING LAB BENCH. THE FAILURE MODE WAS NOT REPRODUCED WHILE RUNNING AN OPTICAL PUMP SIMULATOR. THERE WAS NO ISSUE WITH THE CONSOLE HARDWARE. ROOT CAUSE: THE ROOT CAUSE OF THE CONTROLLER ERROR WAS DUE TO AN AIC SOFTWARE ISSUE. THIS WAS ENTERED INTO JAMA WITH DEFECT/ BUG ID GID-DFT-2927090. Manufacturer Narrative: THE AUTOMATED IMPELLA CONTROLLER (AIC) WAS RETURNED, AND AN EVALUATION IS UNDERWAY. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21697832&pc=NGV | 2124215-2025-18468 | 2024-07-25 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-03-26 04:00 | NGV | WATCHMAN FLX LEFT ATRIAL APPENDAGE CLOSURE DEVICE WITH DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Hemorrhage/Bleeding; Transient Ischemic Attack; Cardiac Tamponade; Pericardial Effusion | P130013 |  | 1 | Event Description: REPORTED VIA JOURNAL ARTICLE. IT WAS REPORTED THAT SERIOUS INJURIES AND DEATH(S) OCCURRED. METHODS: THIS RETROSPECTIVE, MULTICENTER STUDY ANALYZED GENDER-SPECIFIC OUTCOMES IN 650 PATIENTS WHO UNDERWENT LEFT ATRIAL APPENDAGE CLOSURE (LAAC) WITH THE WATCHMAN FLX CLOSURE DEVICE BETWEEN MARCH 2019 AND MAY 2022, DRAWN FROM THE ITALIAN-FLX REGISTRY. RESULTS: PERIPROCEDURAL COMPLICATIONS INCLUDED 2 PATIENTS WHO EXPERIENCED A CEREBRAL VASCULAR ACCIDENT, 1 PATIENT WHO EXPERIENCED A TRANSIENT ISCHEMIC ATTACK, 14 PATIENTS EXPERIENCING MAJOR BLEEDING, 8 PATIENTS EXPERIENCED MINOR BLEEDING, AND 2 PATIENTS EXPERIENCED A PERICARDIAL EFFUSION WITH CARDIAC TAMPONADE. 2 PATIENTS WERE REPORTED TO HAVE DIED PERIPROCEDURAL. THE AVERAGE FOLLOW-UP TIME WAS 354 DAYS. A TOTAL OF 16 PATIENTS DIED OF CARDIOVASCULAR DEATH. 9 CASES OF ISCHEMIC STROKE WERE RECORDED, WHILE 11 PATIENTS SUFFERED FROM TIA. 15 PATIENTS EXPERIENCED MAJOR BLEEDING. A TOTAL OF 29 PATIENTS EXPERIENCED MINOR BLEEDING. Manufacturer Narrative: B2 DATE OF DEATH - ARTICLE PUBLISH DATE USED AS DEATH DATE IS UNKNOWN. B3 DATE OF EVENT - ARTICLE PUBLISH DATE USED AS EVENT DATE IS UNKNOWN. D3 MODEL NUMBER, D3 LOT NUMBER, D4 EXPIRATION DATE, D4 UNIQUE IDENTIFIER (UDI) # ARE UNKNOWN AS THE INFORMATION WAS NOT PROVIDED IN THE LITERATURE ARTICLE. D6A IMPLANT DATE WAS REPORTED AS MARCH 2019 THROUGH MAY 2022. LITERATURE CITATION: BONANNI, M., FRAZZETTO, M., NARDONE, A., MEUCCI, F., MUSTO, C., QUARANTA, G., ... & BERTI, S. (2024). GENDER DIFFERENCES IN OUTCOMES AFTER LEFT ATRIAL APPENDAGE CLOSURE WITH WATCHMAN FLX DEVICE: INSIGHTS FROM THE ITALIAN-FLX REGISTRY. FRONTIERS IN CARDIOVASCULAR MEDICINE, 11, 1419018. DOI: 10.3389/FCVM.2024.1419018.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21697991&pc=DSQ | 2916596-2025-01761 | 2025-03-08 05:00 | Death | THORATEC CORPORATION | 2025-03-26 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Shock; Syncope/Fainting | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE ON (B)(6) 2025 AT 13:35 THE PATIENTS FLOW DROPPED ON LEFT VENTRICULAR ASSIST DEVICE (LVAD) TO 0.4LPM AND THEY WENT UNRESPONSIVE. THREE ROUNDS OF EPINEPHRINE (EPI) WERE ADMINISTERED PRIOR TO INITIATING COMPRESSIONS. DUE TO THE PERSISTENT LOW FLOWS, THE CARDIAC SURGERY PHYSICIAN'S ASSISTANT (PA) DID A CONTROLLER EXCHANGE. ACCORDING TO THE NURSING STAFF, A FEW MORE MINUTES OF COMPRESSIONS WERE ADMINISTERED BEFORE FLOWS WERE REESTABLISHED ON THE LVAD. LOG FILES WERE SENT FOR REVIEW. THE LOG FILE CAPTURED THE START OF LOW FLOW EVENTS ON (B)(6) 2025 AT 13:28:38. THE CALCULATED FLOW APPEARS TO HAVE FLUCTUATED BELOW THE ALARM THRESHOLD OF 2.5 LPM DURING THE EVENTS. AT 13:28:38 THE FLOW WAS AT 4.3 LPM, THEN DROPPED TO 2.4 LPM. THE FLOW THEN CONTINUED TO DROP. THE PUMP WAS RUNNING AT THIS TIME. AT 13:34:50 THE DRIVELINE WAS DISCONNECTED. THE LOG FILES COULD NOT DETERMINE THE REASON FOR THE DROP IN FLOW VALUES.Event Description: IT WAS NOTED THAT THE FLOWS RETURNED FOLLOWING THE CONTROLLER EXCHANGE. THE CONTROLLER DID NOT REMAIN IN USE AS THE PATIENTS BACK UP. THE CAUSE OF THE LOW FLOWS/ CARDIAC ARREST WAS UNABLE TO BE DETERMINED. ADDITIONAL DIAGNOSTIC TESTING WAS NOT PERFORMED BECAUSE THE PATIENT WAS TOO UNSTABLE TO TRANSPORT. THE PATIENT HAD INCREASED PRESSOR REQUIREMENTS DUE TO PROFOUND VASODILATORY SHOCK STATE. THE PATIENT PASSED AWAY.Event Description: THE PATIENT PASSED AWAY ON (B)(6) 2025. THE FAMILY HAD DECLINED AN AUTOPSY. IT WAS UNKNOWN IF DEATH WAS DEVICE RELATED. THE DEVICE HAD OPERATED AS EXPECTED. THE PUMP WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT CONCLUSIVELY BE ESTABLISHED THROUGH THIS EVALUATION. ADDITIONALLY, REVIEW OF THE SUBMITTED LOG FILE CONFIRMED THE REPORTED DROP IN FLOW AND CORRESPONDING LOW FLOW ALARM; HOWEVER, A SPECIFIC CAUSE FOR THESE EVENTS COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. IT WAS REPORTED THAT THE PATIENT HAD A DROP IN FLOWS TO 0.4 LITERS PER MINUTE (LPM) AND WENT UNRESPONSIVE. THREE ROUNDS OF EPINEPHRINE WERE ADMINISTERED FOLLOWED BY COMPRESSIONS. DUE TO THE PERSISTENT LOW FLOWS, THE CARDIAC SURGERY PHYSICIAN¿S ASSISTANT PERFORMED A CONTROLLER EXCHANGE. COMPRESSIONS WERE ADMINISTERED AGAIN BEFORE FLOW WAS REESTABLISHED ON THE LEFT VENTRICULAR ASSIST DEVICE. IT WAS COMMUNICATED THAT THE PATIENT WAS TOO UNSTABLE TO TRANSPORT FOR ADDITIONAL DIAGNOSTIC TESTING. THE PATIENT WAS TREATED WITH INCREASED PRESSOR REQUIREMENTS DUE TO THEIR PROFOUND VASODILATORY SHOCK STATE. THE SUBMITTED CONTROLLER EVENT LOG FILE CONTAINED EVENTS FROM 06MAR2025 THROUGH 08MAR2025. THE FILE CAPTURED A PERSISTENT LOW FLOW HAZARD ALARM WITH FLOWS ULTIMATELY DECREASING TO 0.0 LPM. THE ALARM PERSISTED UNTIL THE DRIVELINE WAS DISCONNECTED ON (B)(6) 2025. THE CONTROLLER WAS THEN SHUTDOWN. THE DRIVELINE DISCONNECT AND CONTROLLER SHUTDOWN WERE CONSISTENT WITH THE REPORTED CONTROLLER EXCHANGE. NO OTHER NOTABLE EVENTS OR ALARMS WERE CAPTURED. IT WAS COMMUNICATED THAT THE PATIENT ULTIMATELY EXPIRED DUE TO UNKNOWN CAUSES. THE DEVICE OPERATED AS EXPECTED. IT WAS COMMUNICATED THAT AN AUTOPSY WAS NOT PERFORMED, AND THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. D, AND PATIENT HANDBOOK, REV. A, ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. THIS SECTION PROVIDES AN EXPLANATION OF PUMP PARAMETERS, INCLUDING PUMP FLOW. SECTION 4 OF THE IFU, ¿SYSTEM MONITOR¿, DESCRIBES THE PUMP FLOW DISPLAY AND THE HAZARD ALARMS. PER DESIGN, WHEN THE ESTIMATED FLOW VALUE IS CALCULATED AT LESS THAN 2.5 LITERS PER MINUTE (LPM), A LOW FLOW STATUS IS POSTED TO THE LOG FILE. IF THE FLOW REMAINS BELOW 2.5 LPM FOR 10 SECONDS, A LOW FLOW HAZARD ALARM IS TRIGGERED. SECTION 4 ALSO EXPLAINS THAT CHANGES IN PATIENT CONDITIONS CAN RESULT IN LOW FLOW. SECTION 7 OF THE IFU, ¿ALARMS AND TROUBLESHOOTING¿, AND SECTION 5 OF THE PATIENT HANDBOOK, ¿ALARMS AND TROUBLESHOOTING¿, OUTLINE SYSTEM CONTROLLER ALARMS, INCLUDING THE LOW FLOW HAZARD ALARM, AND PROVIDES INFORMATION REGARDING HOW TO RESPOND TO AND TROUBLESHOOT THE ALARMS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21698104&pc=OZD | 1220648-2025-26972 | 2025-02-17 05:00 | Death | ABIOMED, INC. | 2025-03-26 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Hemorrhagic Stroke | P140003 |  | 1 | Event Description: IT WAS REPORTED VIA ABIOMED'S LONG-TERM OUTCOME AND QUALITY INDICATOR (LOQI) DATABASE THAT A 57-YEAR-OLD FEMALE PATIENT IMPLANTED WITHA IMPELLA 5.5, HAD AN ADVERSE EVENT OF INTRACEREBRAL HEMORRHAGE WITH "PROBABLE" RELATEDNESS TO THE PUMP. IT WAS REPORTED THAT THE PATIENT HAD FACIAL WEAKNESS AFTER EXTUBATION, AND TONGUE DEVIATION TO LEFT. POTENTIALLY IN SETTING OF TRAUMATIC EXTUBATION, LEFT FACIAL DROOP NOTED, SLURRED SPEECH. PATIENT WAS ON ASPIRIN AND BRILINTA. MRI AND CTA CONFIRMED THE PATIENT HAD AN EVOLVING HEMORRHAGIC STROKE AND PATIENT HAD NO IMPROVEMENT. FAMILY INITIATED WITHDRAWING CARE DUE TO BEING FULLY DEPENDENT ON IMPELLA. EVENTUALLY CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK STATES THE FOLLOWING: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): "ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)".
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21699141&pc=DSQ | 2916596-2025-01996 | 2025-03-18 04:00 | Death | THORATEC CORPORATION | 2025-03-26 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Muscle Weakness; Hemorrhagic Stroke; Ischemia Stroke | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED TO THE HOSPITAL WITH RIGHT SIDED WEAKNESS AND INTELLIGIBLE SPEECH. A COMPUTED TOMOGRAPHY (CT) SCAN SHOWED SUSPICION FOR AN EVOLVING LEFT MIDDLE CEREBRAL ARTERY (MCA) TERRITORY STROKE WITH A HYPERDENSE LEFT MCA. IT WAS NOTED THAT THE PATIENT'S INTERNATIONAL NORMALIZED RATION (INR) HAD BEEN SUBTHERAPEUTIC SINCE ON (B)(6) 2025. THE PATIENT WAS GIVEN A COMPUTED TOMOGRAPHY ANGIOGRAPHY (CTA) SCAN, AND NO ACUTE HEMORRHAGE WAS FOUND. THE PATIENT WAS TAKEN IN FOR A THROMBECTOMY. THE EMBOLIC STROKE LATER CONVERTED TO A HEMORRHAGIC STROKE. THE PATIENT WAS INTUBATED IN THE INTENSIVE CARE UNIT (ICU) AND IS ABLE TO FOLLOW COMMANDS WITH THEIR LEFT SIDE BUT REMAINS UNABLE TO FOLLOW COMMANDS WITH THEIR RIGHT SIDE.Event Description: ADDITIONAL INFORMATION RECEIVED STATED THAT THE PATIENT PASSED AWAY ON (B)(6) 2025 WHILE AT HOME ON HOSPICE. THE CAUSE OF DEATH WAS A STROKE. THE DEVICE OPERATED AS EXPECTED. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. THE DEVICE WAS NOT EXPLANTED AND WOULD NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: THE PUMP REMAINS IN USE SUPPORTING THE PATIENT. A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF STROKE COULD NOT BE CONCLUSIVELY DETERMINED. STROKE IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. STROKE HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21699381&pc=HCF | 1320468-2025-00001 | 2025-02-05 05:00 | Death | LSI SOLUTIONS, INC. | 2025-03-26 04:00 | HCF | COR-KNOT | Mechanical Problem | Insufficient Information | K100593 |  | 1 | Event Description: THE HOSPITAL VERBALLY REPORTED THAT A PATIENT HAD UNDERGONE A SURGICAL MITRAL VALVE REPLACEMENT PROCEDURE WHERE A COR-KNOT® MIS DEVICE WAS USED TO SECURE SUTURE DURING THE OPERATION. THE VERBAL REPORT WAS THAT A SCRUB NURSE WAS HAVING DIFFICULTY LOADING THE COR-KNOT® DEVICE. IT WAS ALSO REPORTED THAT THE DEVICE APPEARED TO CUT THE SUTURE BUT DID NOT CRIMP A TITANIUM FASTENER ONTO THE SUTURE AND, THEREFORE, COULD NOT BE USED. IT IS UNKNOWN WHETHER THE SURGEON SUBSEQUENTLY USED A DIFFERENT COR-KNOT® DEVICE TO SECURE SUTURE OR IF THE SURGEON SWITCHED TO HAND-TIED KNOTS FOR SECURING THE SUTURE. THE PATIENT WAS REPORTED TO HAVE BEEN SENT TO THE INTENSIVE CARE UNIT (ICU) AFTER LEAVING THE OPERATING ROOM (OR). THE PATIENT WAS DESCRIBED AS "A VERY SICK PATIENT" AT THAT TIME. THE PATIENT LATER EXPIRED IN THE ICU. Manufacturer Narrative: THE EVENT AS REPORTED IN SECTION B5 OCCURRED ON FEBRUARY 5, 2025. ON FEBRUARY 24, 2025, OUR FIRST DATE OF AWARENESS FOR THIS REPORT, A REPRESENTATIVE OF THE HOSPITAL CALLED ONE OF OUR SALES REPRESENTATIVES SAYING THAT THERE HAD BEEN AN ISSUE WITH A COR-KNOT® MIS DEVICE USED DURING A SURGICAL PROCEDURE. THE HOSPITAL REPRESENTATIVE ASKED US TO HAVE A CALL WITH THE SURGEON IN QUESTION. OUR REGIONAL SALES MANAGER SPOKE TO THE SURGEON ON THE SAME DAY (ON (B)(6) 2025). IT WAS REPORTED DURING THIS CALL THAT A COR-KNOT® DEVICE HAD APPEARED TO CUT SUTURE BUT DID NOT CRIMP A TITANIUM FASTENER ONTO THE SUTURE. THE SURGEON WANTED TO KNOW IF IT WAS POSSIBLE FOR THE COR-KNOT® DEVICE TO CUT SUTURE WITHOUT CRIMPING A FASTENER ONTO THE SUTURE. (IT IS NOT POSSIBLE, AS WILL BE EXPLAINED IN MORE DETAIL BELOW.) WE WERE ALSO INFORMED THAT THE PATIENT WAS LATER SENT TO THE ICU AFTER LEAVING THE OR. THE PATIENT WAS DESCRIBED AS "A VERY SICK PATIENT," AND WE WERE FURTHER INFORMED THAT THE PATIENT LATER EXPIRED IN THE ICU. THERE IS NO KNOWN OR PROBABLE CONNECTION BETWEEN THE COR-KNOT® DEVICE AND THIS OUTCOME, NOR HAS SUCH AN ALLEGATION BEEN MADE OR INTIMATED. THE SURGEON FURTHER SHARED THAT HE CONTINUES TO USE COR-KNOT® DEVICES IN HIS SURGICAL PROCEDURES. IN ORDER TO BETTER UNDERSTAND THE ANSWER PROVIDED TO THE SURGEON, IT IS HELPFUL TO UNDERSTAND HOW THE COR-KNOT® MIS DEVICE IS USED. THE COR-KNOT® MIS DEVICE USES TITANIUM FASTENERS TO SECURE SUTURE IN THE APPROXIMATION OF SOFT TISSUE AND PROSTHETIC MATERIALS. IN PRACTICE, THE USER LOADS A COR-KNOT® QUICK LOAD® UNIT COMPRISING A HOLLOW COR-KNOT® TITANIUM FASTENER AND A WIRE SNARE INTO THE DISTAL TIP OF A COR-KNOT® MIS DEVICE. THE COR-KNOT® MIS DEVICE IS DESIGNED TO HOLD THE LOADED COR-KNOT® FASTENER UNTIL IT IS LATER CRIMPED BY THE DEVICE. THE USER THEN PULLS THE SNARE TO DRAW SUTURE ENDS THROUGH THE TITANIUM FASTENER LOADED IN THE COR-KNOT® MIS DEVICE. THE SNARE IS SET ASIDE, AND THE USER GENTLY SLIDES THE DISTAL TIP OF THE LOADED COR-KNOT® MIS DEVICE OVER THE SUTURE DOWN TO A TARGETED SITE. THE USER ORIENTS THE DEVICE APPROPRIATELY AND TENSIONS THE SUTURE AS DESIRED. THE USER THEN SQUEEZES A LEVER ON THE COR-KNOT® MIS DEVICE TO CRIMP THE TITANIUM FASTENER ONTO THE SUTURE AND GENTLY TUGS THE SUTURE TO TRIM THE SUTURE TAILS AGAINST AN INTERNAL CUTTING BLADE THAT MOVES INTO A CUTTING POSITION AFTER THE TITANIUM FASTENER IS CRIMPED. THE RESULTANT CRIMPED COR-KNOT® FASTENER SECURES THE SUTURE. THE USER RELEASES THE COR-KNOT® MIS DEVICE LEVER WHICH WAS PREVIOUSLY SQUEEZED TO CRIMP THE COR-KNOT® TITANIUM FASTENER ONTO THE SUTURE, ENSURING THAT THE LEVER IS FULLY RELEASED. AT THIS POINT, THE CRIMPED COR-KNOT® FASTENER IS NO LONGER HELD BY THE COR-KNOT® MIS DEVICE, AND THE COR-KNOT® MIS DEVICE CAN BE WITHDRAWN FROM THE CRIMPED FASTENER. THE CRIMPED FASTENER SECURES THE SUTURE ON WHICH IT WAS CRIMPED. CONSIDERING THE REPORT ASSOCIATED WITH THIS INCIDENT, THAT THE SUTURE APPEARED TO BE CUT BUT THE TITANIUM FASTENER DID NOT CRIMP, IT IS IMPORTANT TO NOTE THAT AS THE COR-KNOT® MIS DEVICE LEVER IS SQUEEZED, IT CAUSES A ROD TO SLIDE DISTALLY WITHIN THE COR-KNOT® MIS DEVICE SHAFT. AS THE ROD MOVES DISTALLY, IT ENGAGES A MECHANISM WHICH CRIMPS A TITANIUM FASTENER ONTO SUTURE LOADED THEREIN. AN INTERNAL CUTTING BLADE MOVES WITH THE ROD AS THE ROD ACTUATES THE CRIMPING MECHANISM, AND, AFTER THE FASTENER IS CRIMPED, THE BLADE REACHES A POSITION WHERE THE SUTURE ENDS EXTENDING FROM THE CRIMPED TITANIUM FASTENER CAN BE DRAWN AGAINST THE CUTTING BLADE BY THE USER. AT THIS POINT, A GENTLE TUG OF THE SUTURE ENDS CAUSES THE SUTURE TO BE TRIMMED AGAINST THE BLADE WHICH HAS BEEN MOVED INTO CUTTING POSITION. SINCE THE INTERNAL ROD THAT MOVES WITHIN THE DEVICE SHAFT CONTROLS THE SEQUENCE OF CRIMPING AND THEN CUTTING, IT IS NOT POSSIBLE FOR THE COR-KNOT® DEVICE TO CUT THE SUTURE BEFORE CRIMPING THE TITANIUM FASTENER ONTO THE SUTURE. IF THE ROD DOES NOT MOVE FAR ENOUGH TO CRIMP THE FASTENER, THE BLADE DOES NOT REACH A POSITION WHERE THE SUTURE CAN BE CUT BY THE COR-KNOT® DEVICE. WE HAVE INQUIRED WITH THE HOSPITAL REGARDING THE AVAILABILITY OF THE COMPLAINT DEVICE SO THAT WE MIGHT HAVE AN OPPORTUNITY TO EVALUATE THE DEVICE. UNFORTUNATELY, THE HOSPITAL INFORMED US THAT THE COR-KNOT® DEVICES (TWO COME IN ONE KIT) WERE DISCARDED AND ARE UNAVAILABLE FOR EVALUATION. ALTHOUGH WE ARE NOT ABLE TO EVALUATE THE COR-KNOT® MIS DEVICES, THE DESIGN OF THE COR-KNOT® MIS DEVICE, DESCRIBED ABOVE, IS SUCH THAT THE DEVICE CANNOT CUT THE SUTURE WITHOUT FIRST CRIMPING THE TITANIUM FASTENER ONTO THE SUTURE. THEREFORE, IT IS HIGHLY LIKELY THAT THE SUTURE BROKE DUE TO THE TENSION APPLIED TO THE SUTURE BEFORE THE COR-KNOT® MIS DEVICE LEVER WAS ACTUATED TO CRIMP THE TITANIUM FASTENER ONTO THE SUTURE. THIS COULD EITHER HAVE BEEN BECAUSE THE SUTURE WAS INADVERTENTLY OVER-TENSIONED, OR THE SUTURE WAS ALREADY WEAKENED IN SOME FASHION AND BROKE UNDER NORMAL TENSION. IN EITHER CASE, THE BREAKING OF THE SUTURE UNDER TENSION WAS LIKELY PERCEIVED AS IF THE COR-KNOT® DEVICE CUT THE SUTURE WITHOUT CRIMPING THE TITANIUM FASTENER. HOWEVER, AS NOTED ABOVE, THE COR-KNOT® DEVICE LEVER COULD NOT HAVE BEEN SQUEEZED BECAUSE, IF IT HAD BEEN SQUEEZED, THE TITANIUM FASTENER WOULD HAVE FIRST BEEN CRIMPED AND THEN THE SUTURE TAILS TRIMMED. IN THIS CASE, WE CAN BE SURE THE COR-KNOT® MIS DEVICE LEVER WAS NOT SQUEEZED BECAUSE IT WAS REPORTED THAT THE TITANIUM FASTENER WAS NOT CRIMPED. THEREFORE, SINCE THE COR-KNOT® MIS LEVER WAS NOT SQUEEZED, THE DEVICE COULD NOT HAVE CUT THE SUTURE. IN A SITUATION WHERE THE SUTURE BROKE BEFORE IT WAS SECURED, AS WAS REPORTED HERE, THE NORMAL COURSE WOULD BE TO REPLACE THE BROKEN SUTURE WITH A NEW SUTURE AND RESECURE THE REPLACEMENT SUTURE. IN THE CASE ON (B)(6) 2025, SURGICAL PROCEDURE, WE DO NOT KNOW IF THE SURGEON USED A DIFFERENT COR-KNOT® MIS DEVICE TO SECURE A REPLACEMENT SUTURE OR IF THE SURGEON USED A HAND-TIED KNOT TO SECURE THE REPLACEMENT SUTURE. HOWEVER, AS NOTED EARLIER, THE SURGEON DID SHARE THAT HE CONTINUES TO USE COR-KNOT® DEVICES IN HIS SURGICAL PROCEDURES. OUR SALES REPRESENTATIVES ALSO VISITED THE HOSPITAL ON (B)(6) 2025 FOR A FOLLOW-UP DISCUSSION. ALTHOUGH THE SURGEON WAS NOT AVAILABLE DURING THIS VISIT, OUR REPRESENTATIVES WERE ABLE TO SPEAK WITH ONE OF THE NURSES. AT THIS TIME, WE LEARNED OF A SECOND ASPECT OF THE REPORTED COMPLAINT: THAT THE SCRUB NURSE ON (B)(6) 2025 SURGICAL PROCEDURE HAD DIFFICULTY LOADING THE COR-KNOT® DEVICE. AS NOTED, ABOVE, THE COMPLAINT DEVICES ARE NOT AVAILABLE FOR EVALUATION, SO WE ARE NOT ABLE TO ASSESS THE LOADING PERFORMANCE OF THE COMPLAINT DEVICES. IF THE DEVICES COULD NOT BE LOADED AT ALL, THEY WOULD NOT HAVE BEEN ABLE TO BE USED, HOWEVER THIS DOES NOT APPEAR TO BE THE CASE SINCE THE LOADED DEVICES WERE PASSED TO THE SURGEON. EVEN THOUGH IT IS TYPICALLY A SCRUB NURSE WHO LOADS THE COR-KNOT® MIS DEVICE DURING A SURGICAL PROCEDURE, THE SURGEON IS INSTRUCTED IN THE DEVICE INSTRUCTIONS FOR USE TO ENSURE THAT THE TITANIUM FASTENER IS FULLY LOADED BEFORE USING THE DEVICE. ULTIMATELY WITH THIS COMPLAINT, THE DEVICE LEVER WAS NOT ACTUATED BECAUSE THE COR-KNOT® FASTENER WAS REPORTED TO REMAIN IN THE DEVICE TIP BUT WAS NOT CRIMPED. THEREFORE, THE REPORTED DIFFICULTY LOADING DOES NOT APPEAR TO BE RELEVANT TO THE OVERALL COMPLAINT BECAUSE THE DEVICE WAS ULTIMATELY LOADED. REGARDING THE PATIENT, WE HAVE NO INFORMATION ABOUT ANY PATIENT HISTORY OR THE COURSE IN THE OR THAT MAY HAVE CAUSED THE PATIENT TO BE "VERY SICK," OR POST-MORTEM ANALYSIS, OR AUTOPSY REPORTS AT THIS TIME. IT IS NOT POSSIBLE TO DETERMINE THE REASON OR REASONS FOR WHY THIS PATIENT EXPIRED. THERE IS NO INFORMATION KNOWN TO US TO SUGGEST THAT THE COR-KNOT® DEVICE OR FASTENER FAILED. IN FACT, THE INTACT, UNCRIMPED FASTENER SUGGESTS THAT THE COR-KNOT® DEVICE WAS NEVER ACTUATED. FURTHERMORE, THE SURGEON STATED THAT HE HAS CONTINUED AND WILL CONTINUE TO USE COR-KNOT® DEVICES AND FASTENERS. THE ROOT CAUSE OF THIS PATIENT'S TRAGIC DEMISE WILL LIKELY NEVER BE FULLY KNOWN. ALTHOUGH WE HAVE VERY LITTLE INFORMATION AT THIS TIME, THE PROPER USE OF A COR-KNOT® TITANIUM FASTENER TO MECHANICALLY SECURE 2-0 POLYESTER SUTURE FOR THE APPROXIMATION OF SOFT TISSUE AND PROSTHETIC MATERIAL IS HIGHLY LIKELY TO PROVIDE MORE THAN AMPLE SECURITY FOR PERMANENTLY HOLDING THIS SUTURE TOGETHER. A FAILURE OF A TITANIUM FASTENER UNDER THESE CONDITIONS OR ANY OTHER RELATED CLINICAL CONDITIONS HAS NEVER BEEN REPORTED DESPITE HAVING OVER 17.4 MILLION TITANIUM FASTENERS IN OVER 1,073,000 CARDIAC SURGERY PATIENTS OVER THE PAST FIFTEEN YEARS IN OVER 70 COUNTRIES. BASED ON THE EXCEPTIONALLY RELIABLE HISTORY OF THE USE OF TITANIUM FASTENERS TO FASTEN SUTURES IN SURGERY AND THE RECOGNIZED POTENTIAL RISKS OF CARDIAC SURGERY, WE BELIEVE THAT THE OUTCOME EXPERIENCED BY THIS PATIENT WAS VERY UNLIKELY TO BE CAUSED BY A FAILURE OF THE TITANIUM FASTENER OR THE COR-KNOT® MIS DEVICE TO PERFORM AS INTENDED. THE INFORMATION PROVIDED IN THIS FORM IS BASED ON WHAT IS REASONABLY KNOWN TO US AFTER CONTACTING THE USER FACILITY/INITIAL REPORTER AND ALSO INCLUDES INFORMATION IN OUR POSSESSION AND INFORMATION THAT WE WERE ABLE TO OBTAIN BY ANALYSIS OF THE TYPE OF DEVICE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21699611&pc=OUT | 2029214-2025-00780 |  | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-03-26 04:00 | OUT | PIPELINE VANTAGE WITH SHIELD TECHNOLOGY | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage; Ruptured Aneurysm; Thrombosis/Thrombus | P100018 |  | 1 | Event Description: IT WAS REPORTED THAT PRIOR TO TREATMENT, THE PATIENT EXPERIENCED A RUPTURED ANEURYSM AND HEMORRHAGIC STROKE. THE PHYSICIAN DECIDED TO TREAT THE WIDE-NECK ANEURYSM LOCATED FROM THE LEFT MIDDLE CEREBRAL ARTERY TO THE LEFT INTERNAL CAROTID ARTERY USING A PIPELINE VANTAGE WITH SHIELD TECHNOLOGY AND ADJUNCTIVE TARGET COILS. THE PROCEDURE WAS TECHNICALLY SUCCESSFUL WITH ANGIOGRAPHIC CONFIRMATION. HOWEVER, WITHIN A FEW WEEKS POST-PROCEDURE, THE ANEURYSM RE-RUPTURED, THE PIPELINE THROMBOSED, AND THE PATIENT SUFFERED ANOTHER HEMORRHAGIC STROKE, ULTIMATELY RESULTING IN DEATH. THE PIPELINE WAS USED OFF-LABEL WITH THE DISTAL LANDING ZONE IN THE MIDDLE CEREBRAL ARTERY. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. ANCILLARY DEVICES: SHEATH PENUMBRA BMX96, GUIDE CATHETER NAVIEN 6F .072ID 115CM, MICROCATHETER MEDTRONIC PHENOM21 150CM MICROCATHETER LOT#: 229655747.Event Description: ADDITIONAL INFORMATION WAS RECEIVED STATING THAT THE PATIENT HAD BEEN HOSPITALIZED PRIOR TO THEIR PASSING AND THEY WERE ADMITTED WITH SPONTANEOUS INTRACEREBRAL HEMORRHAGE (SICH). THE INFORMATION IS CONFIRMED BY THE PHYSICIAN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: B3: EVENT DATE IS NOT KNOWN. PLEASE SEE B5 FOR APPROXIMATE DATE RANGE, IF APPLICABLE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21699612&pc=NGV | 2124215-2025-18367 | 2025-03-18 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-03-26 04:00 | NGV | WATCHMAN FLX PRO | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Cardiac Tamponade; Cardiac Perforation; Pericardial Effusion | P130013 |  | 1 | Event Description: IT WAS REPORTED THAT PERICARDIAL EFFUSION (PE), CARDIAC PERFORATION, AND CARDIAC TAMPONADE OCCURRED. THE PATIENT DIED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED UNDER INTRACARDIAC ECHOCARDIOGRAM (ICE) IMAGING AND FLUOROSCOPY. PRE-PROCEDURE IMAGING CONFIRMED THERE WAS A SMALL RIGHT SIDED PE NOTED PRIOR TO THE PROCEDURE. VENOUS ACCESS WAS OBTAINED AND TRANSSEPTAL PUNCTURE (TSP) WAS PERFORMED USING VERSACROSS CONNECT (VCC) TRANSSEPTAL ACCESS SYSTEM. A WATCHMAN ACCESS SYSTEM (WAS) WAS ADVANCED WITH A 27 MM WATCHMAN FLX PRO CLOSURE DEVICE AND DELIVERY SYSTEM (WDS). THE 27 MM WDS WAS UNSHEATHED AND ADVANCED INTO THE LAA BUT THEN THE PHYSICIAN DECIDED TO REMOVE THE 27 MM WDS AND EXCHANGE IT FOR A LARGER DEVICE. A 31 MM WATCHMAN FLX PRO CLOSURE DEVICE WAS INSERTED, ADVANCED AND DEPLOYED. A PE LOCATED AT THE DISTAL BACK WALL OF THE LAA WAS NOTED, WHICH LED TO CARDIAC TAMPONADE. A PERICARDIOCENTESIS WAS PERFORMED AND BRIGHT RED BLOOD WAS REMOVED. THE CLOSURE DEVICE WAS IMPLANTED. THE PROCEDURE WAS CONCLUDED, AND THE PATIENT WAS SENT FOR CARDIAC SURGERY WHERE THE IMPLANTED CLOSURE DEVICE WAS EXPLANTED, THE CARDIAC PERFORATION REPAIRED, AND THE LAA WAS SURGICALLY LIGATED. THE PATIENT REQUIRED TWO (2) UNITS OF BLOOD TRANSFUSION ALONG WITH TWO (2) UNITS OF FRESH FROZEN PLASMA AND INTRAVENOUS FLUIDS. THE PATIENT CODED. THE PATIENT WAS ADMITTED IN THE INTENSIVE CARE UNIT (ICU). THE PATIENT WAS MADE A DO NOT RESUSCITATE (DNR) STATUS BY THE FAMILY. TWO (2) DAYS POST INDEX PROCEDURE, THE PATIENT DIED. IN THE PHYSICIAN'S OPINION, THE WAS WAS AGAINST THE BACK WALL OF THE LAA SO DURING DEPLOYMENT, THE CLOSURE DEVICE PERFORATED THE LAA WALL.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21699713&pc=KNT | MW5168243 | 2025-03-06 05:00 | Death | AVANOS MEDICAL / AVANOS MEDICAL, INC. (FORMERLY HALYARD HEALTH INC.) | 2025-03-26 04:00 | KNT | KIT INTRODUCER PEG 20FR PULL METHOD SAFETY MIC KEY LOW PROFILE TUBE |  | Pain; Sepsis; Peritonitis |  |  | 1 | Event Description: THE PATIENT'S PEG (PERCUTANEOUS ENDOSCOPIC GASTROSTOMY) WAS REMOVED AND REPLACED WITH A MICKEY BUTTON PER PROTOCOL WITHOUT COMPLICATION (MORE FORCE THAN USUAL WAS NEEDED, HOWEVER). PATIENT TOLERATED THIS WELL. THE PATIENT'S CAREGIVER CALLED LATER, STATING THE PATIENT WAS HAVING PAIN AND LOOKED PALE. PER REVIEW OF THE NOTES, THE PATIENT WAS EVALUATED AT HIS LOCAL ER AND THEN ADMITTED TO HIS LOCAL HOSPITAL FOR CONTINUED CARE AND TREATMENT. A CT SCAN REVEALED FINDINGS CONSISTENT WITH TUBE FEEDINGS WITHIN THE ABDOMINAL CAVITY. THE PATIENT DEVELOPED SEPSIS AND PERITONITIS AND PASSED AWAY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21699866&pc=DXY | 2182208-2025-01524 | 2024-12-01 05:00 | Death | MEDTRONIC, INC. | 2025-03-26 04:00 | DXY | MEDTRONIC ICD | Appropriate Term/Code Not Available | Insufficient Information |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING ARRHYTHMIA DETECTION EFFICACY OF IMPLANTABLE CARDIOVERTER DEFIBRILLATORS (ICDS). THE AUTHORS DESCRIBED PATIENT DEATHS; HOWEVER, THE CAUSES OF DEATH WERE UNKNOWN AND DESCRIBED AS ALL-CAUSE. THERE WERE PATIENTS WHO EXPERIENCED INAPPROPRIATE THERAPY DUE TO DETECTION OF ATRIAL ARRHYTHMIAS, SYSTEM, POCKET, OR SUPERFICIAL INFECTIONS, HEMATOMAS, ONE DEEP VENOUS THROMBOSIS, AND HEART FAILURE OR ARRHYTHMIA-RELATED HOSPITALIZATIONS. THERE WERE ALSO UNSPECIFIED LEAD-RELATED COMPLICATIONS IN WHICH INTERVENTION WAS PERFORMED. THE STATUS OF THE DEVICES AND LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/64 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: ARRHYTHMIA DETECTION IN ATRIOVENTRICULAR, SINGLE-LEAD, FLOATING ATRIAL DIPOLE ICD SYSTEMS COMPARED WITH CONVENTIONAL SINGLE AND DUAL-CHAMBER DEFIBRILLATORS. JOURNAL OF CARDIOVASCULAR DEVELOPMENT AND DISEASE. 2024, 11, 386. DOI: 10.3390/JCDD11120386. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21699868&pc=DXY | 2182208-2025-01525 | 2024-12-01 05:00 | Death | MEDTRONIC, INC. | 2025-03-26 04:00 | DXY | MEDTRONIC LEAD | Appropriate Term/Code Not Available | Insufficient Information |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING ARRHYTHMIA DETECTION EFFICACY OF IMPLANTABLE CARDIOVERTER DEFIBRILLATORS (ICDS). THE AUTHORS DESCRIBED PATIENT DEATHS; HOWEVER, THE CAUSES OF DEATH WERE UNKNOWN AND DESCRIBED AS ALL-CAUSE. THERE WERE PATIENTS WHO EXPERIENCED INAPPROPRIATE THERAPY DUE TO DETECTION OF ATRIAL ARRHYTHMIAS, SYSTEM, POCKET, OR SUPERFICIAL INFECTIONS, HEMATOMAS, ONE DEEP VENOUS THROMBOSIS, AND HEART FAILURE OR ARRHYTHMIA-RELATED HOSPITALIZATIONS. THERE WERE ALSO UNSPECIFIED LEAD-RELATED COMPLICATIONS IN WHICH INTERVENTION WAS PERFORMED. THE STATUS OF THE DEVICES AND LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/64 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: ARRHYTHMIA DETECTION IN ATRIOVENTRICULAR, SINGLE-LEAD, FLOATING ATRIAL DIPOLE ICD SYSTEMS COMPARED WITH CONVENTIONAL SINGLE AND DUAL-CHAMBER DEFIBRILLATORS. JOURNAL OF CARDIOVASCULAR DEVELOPMENT AND DISEASE. 2024, 11, 386. DOI: 10.3390/JCDD11120386. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21699909&pc=NPT | 2135147-2025-01565 | 2025-03-05 05:00 | Death | ABBOTT MEDICAL | 2025-03-26 04:00 | NPT | NAVITOR VISION | Migration or Expulsion of Device | Aortic Valve Insufficiency/ Regurgitation | P190023 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 25MM NAVITOR VALVE WAS SELECTED FOR IMPLANT. THE ANNULAR PERIMETER OF THE NATIVE VALVE WAS 71MM. PRIOR TO PROCEDURE, AND ANGIOGRAM SHOWED A SIGNIFICANT LESION IN THE LMS, SO IT WAS DECIDED TO PERFORM A PCI. A STENT WAS PLACED BUT SOON AFTER THE PATIENT LOST PRESSURE AND CPR WAS COMMENCED. AUTO PULSE CPR WAS CARRIED OUT WHILE ATTEMPTED TO TREAT CORONARIES. AFTER AROUND 30 MINUTES OF CPR, THE DECISION WAS MADE TO DEPLOY VALVE IN THE HOPE THAT TREATING THE AS AND IT WOULD STABILIZE THE PATIENT. THE VALVE DEPLOYED QUICKLY WITH LIMITED VIEWS AVAILABLE DUE TO THE AUTO PULSE CPR. THE IMPLANT DEPTH OF THE VALVE APPEARED TO BE 4-5MM. THE VALVE DIVED INTO THE LVOT LEADING TO TORRENTIAL AR. A 23MM NON-ABBOTT VALVE WAS THEN DEPLOYED INSIDE THE NAVITOR VALVE. THE AR RESOLVED BUT THERE WAS A DOWNTIME OF 50+ MINS, SO IT WAS DECIDED TO STOP CPR ON THE PATIENT. THE PHYSICIANS ALLEGE THAT THE CAUSE OF DEATH WAS DUE TO CRITICAL LMS STENOSIS, WHICH WHILE TREATING RESULTED IN A LOSS OF PRESSURE, AS WELL AS SEVERE AORTIC STENOSIS. THE PHYSICIAN DOESN'T ALLEGE THAT IT WAS THE VALVES' FAULT AND THAT IT WAS A LAST ATTEMPT TO SAVE THE PATIENT.Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 25MM NAVITOR VALVE WAS SELECTED FOR IMPLANT. THE ANNULAR PERIMETER OF THE NATIVE VALVE WAS 71MM. PRIOR TO PROCEDURE, AND ANGIOGRAM SHOWED A SIGNIFICANT LESION IN THE LMS, SO IT WAS DECIDED TO PERFORM A PCI. A STENT WAS PLACED BUT SOON AFTER THE PATIENT LOST PRESSURE AND CPR WAS COMMENCED. AUTO PULSE CPR WAS CARRIED OUT WHILE ATTEMPTED TO TREAT CORONARIES. AFTER AROUND 30 MINUTES OF CPR, THE DECISION WAS MADE TO DEPLOY VALVE IN THE HOPE THAT TREATING THE AS AND IT WOULD STABILIZE THE PATIENT. THE VALVE DEPLOYED QUICKLY WITH LIMITED VIEWS AVAILABLE DUE TO THE AUTO PULSE CPR. THE IMPLANT DEPTH OF THE VALVE APPEARED TO BE 4-5MM. THE VALVE DIVED INTO THE LVOT LEADING TO TORRENTIAL AR. A 23MM NON-ABBOTT VALVE WAS THEN DEPLOYED INSIDE THE NAVITOR VALVE. THE AR RESOLVED BUT THERE WAS A DOWNTIME OF 50+ MINS, SO IT WAS DECIDED TO STOP CPR ON THE PATIENT. THE PHYSICIANS ALLEGE THAT THE CAUSE OF DEATH WAS DUE TO CRITICAL LMS STENOSIS, WHICH WHILE TREATING RESULTED IN A LOSS OF PRESSURE, AS WELL AS SEVERE AORTIC STENOSIS. THE PHYSICIAN DOESN'T ALLEGE THAT IT WAS THE VALVES FAULT AND THAT IT WAS A LAST ATTEMPT TO SAVE THE PATIENT. Manufacturer Narrative: CORRECTION: B2, H1. AN EVENT OF MIGRATION OR EXPULSION OF DEVICE (VENTRICULAR DIRECTION) WAS REPORTED. A RETURNED DEVICE ASSESSMENT COULD NOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED FOR ANALYSIS. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED, AND THE PRODUCT MET ALL SPECIFICATIONS. BASED ON THE INFORMATION RECEIVED, THE CAUSE FOR THE REPORTED DEVICE MIGRATION IN VENTRICULAR DIRECTION COULD NOT BE DETERMINED. THE PATIENT EFFECTS OF DEATH WAS DUE TO CRITICAL LMS STENOSIS, WHICH WHILE TREATING RESULTED IN A LOSS OF PRESSURE, AS WELL AS SEVERE AORTIC STENOSIS. THE AORTIC VALVE REGURGITATION APPEARS TO BE RELATED TO THE DEVICE MIGRATION. THE REPORTED SURGICAL INTERVENTION WAS A RESULT OF CASE-SPECIFIC CIRCUMSTANCES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH REGARDS TO MANUFACTURE, DESIGN, OR LABELING. Manufacturer Narrative: INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21699984&pc=DXF | 2124215-2025-18769 | 2025-03-10 04:00 | Death | BAYLIS MEDICALE CIE INC | 2025-03-26 04:00 | DXF | VERSACROSS ACCESS SOLUTION | Adverse Event Without Identified Device or Use Problem | Bradycardia; Cardiac Arrest; Low Blood Pressure/ Hypotension; Tachycardia; Cardiac Tamponade; Pericardial Effusion | K150709 |  | 1 | Event Description: IT WAS REPORTED A PATIENT DEATH OCCURRED. THE PROCEDURE WAS CANCELLED. DURING AN ATRIAL FIBRILLATION ABLATION PROCEDURE, VERSACROSS ACCESS SOLUTION KIT WAS SELECTED FOR USE. THE CTI LINE WAS PERFORMED IN THE RIGHT ATRIUM USING A NON-BOSTON SCIENTIFIC MAPPING CATHETER. AFTER THE CTI LINE WAS PERFORMED, THE PHYSICIAN PERFORMED TRANSEPTAL PUNCTURE WITH NO ISSUES REPORTED. ONCE ACROSS THE SEPTUM, THE VERSACROSS SHEATH WAS EXCHANGED FOR THE FARADRIVE SHEATH AND THE MAPPING CATHETER WAS INSERTED. IT WAS AT THIS POINT THAT THE PATIENT'S HEART RATE BECAME TACHYCARDIC AND HYPOTENSIVE WITH ONLY AN AVAILABLE READING OF MAP OF 24. THE BLOOD PRESSURE WAS THEN REATTEMPTED WITH A SYSTOLIC IN THE 60'S. THE PATIENT'S HEART RATE THEN DECLINED TO THE 50'S-40'S WHERE THEY WERE THEN PULSELESS ELECTRICAL ACTIVITY (PEA) AND REQUIRED RESUSCITATIVE EFFORTS. AN INTRACARDIAC ECHOCARDIOGRAPHY (ICE) SCAN DETECTED A PERICARDIAL EFFUSION WHICH LED TO CARDIAC TAMPONADE. A PERICARDIOCENTESIS WAS PERFORMED AND BLOOD WAS ADMINISTERED. THE PATIENT PASSED AWAY DESPITE THE EFFORTS OF THE LAB STAFF. THE DEVICE IS EXPECTED TO BE RETURNED FOR ANALYSIS. THE PERICARDIAL EFFUSION WAS NOTED ON ICE IMAGE, BUT UNSURE OF EXACTLY WHERE IS STARTED. IT WAS REPORTED THAT THE PATIENT WAS ELDERLY AND WAS IN PERSISTENT ATRIAL FIBRILLATION (A FIB) FOR ALONG PERIOD OF TIME, ANATOMY COULD HAVE BEEN A FACTOR. THE PHYSICIAN DID NOT MAKE ANY COMMENTS THAT OUR DEVICE WAS THE CAUSE OF THE INCIDENT. Manufacturer Narrative: THE DEVICE HAS NOT BEEN RECEIVED FOR ANALYSIS. UPON RECEIPT AND COMPLETION OF THE FAILURE ANALYSIS OF THE COMPLAINT DEVICE, IF THERE IS ANY FURTHER RELEVANT INFORMATION FROM THAT REVIEW, A SUPPLEMENTAL MEDWATCH WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21700129&pc=NPT | 2015691-2025-02270 | 2025-02-13 05:00 | Death | EDWARDS LIFESCIENCES | 2025-03-26 04:00 | NPT | EDWARDS LIFESCIENCES SAPIEN 3 TRANSCATHETER HEART VALVE | Device Stenosis | Heart Failure/Congestive Heart Failure | P140031 |  | 1 | Event Description: IT WAS REPORTED BY THE VALVE CLINIC COORDINATOR (VCC) THAT APPROXIMATELY 7 YEARS, 9 MONTHS, AND 11 DAYS POST TRANSFEMORAL TAVR WITH A 20MM SAPIEN 3 VALVE, THE PATIENT HAS STENOSIS OF THEIR HEART VALVE. THE PATIENT WAS ADMITTED TO THE HOSPITAL FOR HEART FAILURE SYMPTOMS. WHILE BEING REVIEWED FOR INTERVENTION THE PATIENT PASSED AWAY. Manufacturer Narrative: PER THE INSTRUCTION FOR USE (IFU), VALVE STENOSIS IS A POTENTIAL RISK ASSOCIATED WITH BIOPROSTHETIC HEART VALVES. PER THE VALVE ACADEMIC RESEARCH CONSORTIUM (VARC), VALVE STENOSIS CAN RESULT FROM A NUMBER OF FACTORS, INCLUDING PANNUS, CALCIFICATION, SUPPORT STRUCTURE DEFORMATION (OUT-OF-ROUND CONFIGURATION), TRAUMA (CARDIA-PULMONARY RESUSCITATION, BLUNT CHEST TRAUMA), ENDOCARDITIS, PROSTHETIC VALVE THROMBOSIS, AND NATIVE LEAFLET PROLAPSE IMPEDING PROSTHETIC LEAFLET MOTION. VALVE STENOSIS MAY RESULT IN SYMPTOMS SUCH AS SHORTNESS OF BREATH AND DECREASED EXERCISE TOLERANCE, WHICH MAY BE ACCOMPANIED BY AN INCREASED GRADIENT ACROSS THE VALVE. PER THE INSTRUCTIONS FOR USE (IFU), STRUCTURAL VALVE DETERIORATION (WEAR, FRACTURE, CALCIFICATION, LEAFLET TEAR/TEARING FROM THE STENT POSTS, LEAFLET RETRACTION, SUTURE LINE DISRUPTION OF COMPONENTS OF A PROSTHETIC VALVE, THICKENING, STENOSIS), AND DEVICE DEGENERATION ARE KNOWN POTENTIAL RISKS ASSOCIATED WITH BIOPROSTHETIC HEART VALVES. THE IFU CAUTIONS THAT ACCELERATED DETERIORATION OF THE VALVE MAY OCCUR IN PATIENTS WITH AN ALTERED CALCIUM METABOLISM. LONG-TERM DURABILITY HAS NOT BEEN ESTABLISHED FOR THE VALVE. REGULAR MEDICAL FOLLOW-UP IS ADVISED TO EVALUATE VALVE PERFORMANCE. STRUCTURAL VALVE DETERIORATION (SVD) MAY BE MANIFESTED AS STENOSIS WITH THICKENED LEAFLETS. SVD MAY BE MILD AND NOT REQUIRE ANY INTERVENTION OR IT MAY BE MODERATE TO SEVERE. IT CAN CAUSE THE HEART TO WORK HARDER TO EJECT BLOOD FROM THE VENTRICLE. DEPENDING ON THE SEVERITY IT COULD BE AN INDICATION FOR VALVE REPLACEMENT OR MEDICAL INTERVENTION. TISSUE CALCIFICATION IS A COMMON FAILURE MODE. THE MECHANISMS FOR BIOPROSTHETIC HEART VALVE TISSUE CALCIFICATION ARE NOT YET FULLY UNDERSTOOD. MANY FACTORS CAN CONTRIBUTE TO THE ONSET AND PROPAGATION OF CALCIFICATION INCLUDING PATIENT-RELATED (E.G., PATIENT AGE, DISEASE STATE, IMMUNE STATUS, AND OTHER CO-MORBIDITIES), PHARMACOLOGICAL, AND INTRINSIC PROPERTIES OF THE VALVE ITSELF. IT IS WIDELY UNDERSTOOD THAT PATIENTS WITH CHRONIC RENAL DISEASE AND PRIOR HISTORY OF CALCIFIC STENOSIS OF THE NATIVE VALVE MAY BE PREDISPOSED TO BIOPROSTHETIC CALCIFICATION. DURING THE MANUFACTURING PROCESS, ALL SAPIEN THV VALVES ARE 100% VISUALLY INSPECTED FOR DEFECTS AND 100% FUNCTIONALLY TESTED FOR PROPER COAPTATION UNDER PHYSIOLOGICAL BACKPRESSURE CONDITIONS PRIOR TO RELEASE FOR DISTRIBUTION. THEREFORE, IT IS HIGHLY UNLIKELY THAT A MANUFACTURING DEFECT OR DEVICE MALFUNCTION CONTRIBUTED TO THE EVENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A MANUFACTURING NON-CONFORMANCE CONTRIBUTED TO THIS PRODUCT PROBLEM. ALTHOUGH THE EXACT CAUSE AND SOURCE OF THE STENOSIS CANNOT BE DETERMINED, IT IS POSSIBLE PATIENT FACTORS NOT PROVIDED MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21700320&pc=DSQ | 2916596-2025-01748 | 2025-03-05 05:00 | Death | THORATEC CORPORATION | 2025-03-26 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Liver Damage/Dysfunction; Renal Failure; Sepsis; Respiratory Failure; Multiple Organ Failure; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2025 DUE TO MULTISYSTEM ORGAN FAILURE (MSOF). WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED.Event Description: IT WAS LATER REPORTED THAT THE PATIENT HAD SEPSIS LEADING UP TO THEIR DEATH. THE PATIENT ALSO HAD CARDIAC FAILURE, RENAL FAILURE, AND RESPIRATORY FAILURE WITH HEPATIC DYSFUNCTION. IT WAS SAID THE PATIENT'S DEATH AND MULTISYSTEM ORGAN FAILURE WERE NOT DEVICE OR THERAPY RELATED. THE DEVICE WAS SAID TO HAVE BEEN OPERATING AS EXPECTED DURING THE EVENT. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN BE FOUND ON THE ELECTRONIC IFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REV. D AND THE HEARTMATE 3 PATIENT HANDBOOK, REV. D ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION, SEPSIS, AND DEATH, AS POTENTIAL ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SEVERAL SECTIONS OF THE HEARTMATE 3 IFU AND PATIENT HANDBOOK ALSO PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER: (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN BE FOUND ON THE ELECTRONIC IFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REV. D IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION, AND DEATH, AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21700653&pc=DSQ | 2916596-2025-01663 | 2020-01-20 05:00 | Death | THORATEC CORPORATION | 2025-03-26 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Ischemia; Pneumonia; Renal Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT 8 WAS IMPLANTED AT AGE 60.3 YEARS AS DESTINATION THERAPY AND PRESENTED WITH SYMPTOMATIC LOW FLOW ALARMS 1.1 YEARS FOLLOWING IMPLANT. COMPUTERIZED TOMOGRAPHY ANGIOGRAPHY, CTA, IDENTIFIED 100% OCCLUSION RESULTING IN CARDIOGENIC SHOCK AND NECESSITATING A PUMP EXCHANGE. PRE-INTERVENTION HM3 PARAMETERS WERE REPORTED AS; SPEED 5100 RPM, FLOW N/A, PI N/A AND POWER N/A. POST INTERVENTION PARAMETERS WERE REPORTED AS: SPEED 5200 RPM, FLOW N/A, PI N/A AND POWER N/A. DURATION OF FOLLOW-UP WAS 1.1 YEARS POST HM3. POST INTERVENTION 9 DAYS WAS COMPLICATED BY ACUTE LIMB ISCHEMIA, PNEUMONIA, VASODILATORY SHOCK, AND ACUTE RENAL FAILURE REQUIRING DIALYSIS. THE PATIENT ALSO HAD A RIGHT VENTRICULAR ASSIST DEVICE (RVAD) IMPLANTED POST-OPERATION. THE PATIENT PASSED AWAY ON (B)(6) 2020. Manufacturer Narrative: SECTIONS A AND D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DEVICE WAS IMPLANTED AT TIME OF EVENT. VINOGRADSKY, A., HYNDS, M., KAKU, Y., LEB, J., YUZEFPOLSKAYA, M., COLOMBO, P., SAYER, G., URIEL, N., NAKA, Y., & TAKEDA, K. (2024B). SURGICAL INTERVENTIONS FOR ACCUMULATION OF BIODEBRIS CAUSING OUTFLOW GRAFT OBSTRUCTION AND THROMBOSIS FOLLOWING HEARTMATE 3. THE JOURNAL OF HEART AND LUNG TRANSPLANTATION, 43(4), S287. HTTPS://DOI.ORG/10.1016/J.HEALUN.2024.02.1233. DIVISION OF CARDIAC, THORACIC & VASCULAR SURGERY, DEPARTMENT OF SURGERY, COLUMBIA UNIVERSITY IRVING MEDICAL CENTER, NEW YORK, NY; COLUMBIA UNIVERSITY IRVING MEDICAL CENTER, NEW YORK, NY; DIVISION OF CARDIOLOGY, DEPARTMENT OF MEDICINE, COLUMBIA UNIVERSITY IRVING MEDICAL CENTER, NEW YORK, NY. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER: (B)(6), AND THE REPORTED EVENTS COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. HEARTMATE 3 LVAS, SERIAL NUMBER: (B)(6), WAS NOT RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATION. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) AND THE HEARTMATE 3 PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING MULTIPLE ORGAN FAILURES/DYSFUNCTIONS (INCLUDING RENAL DYSFUNCTION) AND INFECTION. SEVERAL SECTIONS OF THE IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21700701&pc=DSQ | 2916596-2025-01805 | 2025-01-01 05:00 | Death | THORATEC CORPORATION | 2025-03-26 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE ARTICLE "VENTRICULAR ASSISTED DEVICE IN TRANSPOSITION OF GREAT ARTERIES: INDICATIONS, OUTCOMES, HEMODYNAMIC INSIGHTS" THAT AN ASSESSMENT LITERATURE WAS CONDUCTED INVOLVING VENTRICULAR ASSIST DEVICE (VAD) USE IN SYSTEMIC RIGHT VENTRICLE (SRV) PATIENTS. THE STUDY INVOLVED 89 PATIENTS WHO UNDERWENT VAD IN SRV. 66 MALES AND 14 FEMALES WERE INCLUDED AND 17 PATIENTS WERE IMPLANTED WITH A HEARTMATE 3. 47 RECEIVED THE DEVICE AS A BRIDGE TO TRANSPLANT AND 18 RECEIVED THE DEVICE AS DESTINATION THERAPY. 31 PATIENTS UNDERWENT HEART TRANSPLANT AND 13 PASSED AWAY, WITH THE AVERAGE DURATION OF VAD SUPPORT BEING 15 MONTHS. IMPLANTATION OF VAD RESULTED IN A SUSTAINED REDUCTION OF PULMONARY PRESSURES. THERE WERE NO DISCERNIBLE OUTCOME DIFFERENCES AMONG DIFFERENT DEVICE GENERATIONS. THE MEAN PULMONARY ARTERY PRESSURE DECREASED SIGNIFICANTLY FROM 43+/-16 MMHG TO 24+/-9 MMHG AFTER VAD IMPLANTATION (P=0.001). PULMONARY VASCULAR RESISTANCE (PVR) DECREASED SIGNIFICANTLY FROM 5WU TO 2.5WU AFTER VAD IMPLANT (P=0.001). THE VAD RESULTED IN A SUSTAINABLE REDUCTION IN PULMONARY AND FILLING PRESSURES. Manufacturer Narrative: SECTIONS A AND D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. SECTION B3: DATE OF EVENT HAS BEEN ENTERED AS THE DATE OF PUBLICATION SINCE THE DATE OF DATA COLLECTION WAS NOT PROVIDED. DATE OF DEATH WAS NOT PROVIDED. AUTHOR INFORMATION: INDIANA UNIVERSITY SCHOOL OF MEDICINE, INDIANAPOLIS, IN; INDIANA UNIVERSITY HEALTH, INDIANAPOLIS, IN. ABDELKADER, A. E., GUGLIN, M., & RAO, R. (2025). VENTRICULAR ASSISTED DEVICE IN TRANSPOSITION OF GREAT ARTERIES: INDICATIONS, OUTCOMES, HEMODYNAMIC INSIGHTS. JOURNAL OF CARDIAC FAILURE, 31(1), 316. HTTPS://DOI.ORG/10.1016/J.CARDFAIL.2024.10.342. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEMS (LVAS) AND THE REPORTED EVENTS COULD NOT CONCLUSIVELY BE ESTABLISHED THROUGH THIS EVALUATION. THE SERIAL NUMBERS OF THE HEARTMATE 3 LVAS IN USE WERE NOT PROVIDED. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THIS DOCUMENT LISTS THE ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21700710&pc=DSQ | MW5168268 | 2025-02-02 05:00 | Death | MEDTRONIC, INC. | 2025-03-26 04:00 | DSQ | MEDTRONIC HVAD |  | Cardiac Arrest |  |  | 1 | Event Description: PATIENT DRIVELINE NOTED TO BECOME DISCONNECTED FROM LVAD. EARLIER IN DAY, PATIENT WAS HAVING A CONTROLLER FAULT ALARM LEADING TO HIM ATTEMPTING TO CHANGE TO BACK UP CONTROLLER WHEN DRIVELINE BECAME DISCONNECTED. AS A RESULT, THE PATIENT ARRESTED AT A GAS STATION. EMS ARRIVED AND STARTED CHEST COMPRESSIONS. PATIENT WAS SUBSEQUENTLY TAKEN TO (B)(6) WHERE DRIVELINE WAS RECONNECTED. EMERGENCY ROOM WAS PUT IN CONTACT WITH (B)(6) FOR TRANSFER TO (B)(6). ONCE PATIENT ARRIVE TO (B)(6), VAD WAS INTERROGATED WHICH SHOWED THAT PATIENT HAD A DOUBLE DISCONNECT PRIOR TO LEAVING THE HOUSE ON (B)(6) 2025. PATIENT'S WIFE CONFIRMED THAT THE PATIENT WAS GETTING READY TO LEAVE AND ACCIDENTALLY DOUBLE DISCONNECTED BOTH POWERS. PUMP WAS OFF FOR 16 SECONDS. THIS WAS A SEPARATE INCIDENT FROM THE ATTEMPTED CONTROLLER CHANGE. THE CONTROLLER FAULT WAS DUE TO INTERNAL BATTERY END OF LIFE. WITH THIS TYPE OF ALARM, THE VAD IS STILL ON AND FUNCTIONING. IF POWERS ARE DOUBLE DISCONNECTED THERE IS A CHANCE THE ALARM WOULD NOT SOUND DUE TO INTERNAL BATTERY END OF LIFE BUT OTHERWISE THE CONTROLLER STILL FUNCTIONS. THIS WOULD EVENTUALLY REQUIRE A CONTROLLER CHANGE BUT THIS WOULD HAVE BEEN DONE IN A CLINIC SETTING. THE PATIENT NEVER SWITCHED TO HIS BACK UP CONTROLLER SO THE CONTROLLER HE WAS ON WAS STILL SHOWING CONTROLLER FAULT. THIS ALARM WAS PERMANENTLY SILENCED. THE VAD WAS OFF FOR 50 MINUTES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21700721&pc=NPT | 2025587-2025-02137 |  | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-26 04:00 | NPT | EVOLUT FX VALVE | Perivalvular Leak; Insufficient Information | Valvular Insufficiency/ Regurgitation | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING THE IMPLANT OF THE TRANSCATHETER AORTIC VALVE EVOLUTFX-34, PARAVALVULAR LEAK OCCURRED. SUBSEQUENTLY, THE DEVICE WAS EXPLANTED; HOWEVER, DURING THE EXPLANT PROCEDURE, THE PATIENT DIED. NO FURTHER DETAILS WERE PROVIDED. Manufacturer Narrative: PRODUCT ANALYSIS: UPON RECEIPT OF THE SUSPECT COMPLAINT DEVICE AT MEDTRONIC¿S QUALITY LABORATORY, IN A SPECIMEN JAR, FULLY SUBMERGED IN A CLOUDY SOLUTION ENCLOSED INSIDE TWO BAGS, THE VALVE WAS TIED WITH A GREEN MULTIFILAMENT SUTURE AT THE OUTFLOW CROWN. VISUAL ANALYSIS SHOWED ALL LEAFLETS WERE FLEXIBLE AND INTACT. ALL LEAFLETS WERE IN A CLOSED POSITION WITH A SLIGHT GAP BETWEEN THE FREE MARGINS OF LEAFLET 2 AND LEAFLET 3. ALL COMMISSURES WERE INTACT. NO DAMAGE WAS FOUND ON THE FRAME. ADDITIONAL OBSERVATIONS WERE A TEAR ON THE OUTER WRAP. RADIOGRAPHY SHOWED NO CALCIFICATION OR FRAME FRACTURES. CONCLUSION: THE DEVICE HISTORY RECORD WAS REVIEWED AND SHOWED THIS PRODUCT MET ALL MANUFACTURING SPECIFICATIONS FOR PRODUCT RELEASED FOR DISTRIBUTION. NO ISSUES WERE IDENTIFIED THAT WOULD HAVE IMPACTED THIS EVENT. PARAVALVULAR LEAK (PVL) IS A KNOWN POTENTIAL ADVERSE EFFECT PER THE DEVICE INSTRUCTIONS FOR USE (IFU) AND CAN BE CAUSED BY A VARIETY OF FACTORS, INCLUDING VALVE POSITIONING, PATIENT ANATOMY, OR PRESENCE OF PRE-EXISTING PATIENT CONDITIONS. A PROCEDURE- OR VALVE-RELATED DEATH IS AN INHERENT RISK WHEN THE PATIENT CONDITION IS SUCH THAT A TRANSCATHETER AORTIC VALVE IS NEEDED TO SUSTAIN CARDIAC FUNCTION, AND IT CAN OCCUR DESPITE AN IDEAL IMPLANT PROCEDURE OR DEVICE FUNCTIONALITY. THERE WAS NO INFORMATION TO INDICATE A QUALITY DEFICIENCY WAS RELATED TO THIS EVENT. UPDATED: D4, D9, H3, H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: B2 AND B3. EVENT DATE AND DATE OF DEATH ARE UNKNOWN. SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21701803&pc=NAY | 2955842-2025-11267 | 2025-02-28 05:00 | Death | INTUITIVE SURGICAL, INC | 2025-03-26 04:00 | NAY | DAVINCI XI | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | K131861 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A DA VINCI ASSISTED RIGHT HEMICOLECTOMY PROCEDURE, THE PATIENT WENT INTO CARDIAC ARREST AND THE ANESTHESIOLOGIST INFORMED THE OPERATING ROOM STAFF TO START COMPRESSIONS; THE PATIENT SUBSEQUENTLY DIED. THE SURGEON STATED THAT THE SURGERY WAS GOING WELL, THE COLON HAD BEEN RESECTED, THE COMMON CHANNEL WAS BEING CLOSED, AND THE SURGERY WAS ALMOST FINISHED. THE DA VINCI SYSTEM WAS UNDOCKED TO PERFORM CHEST COMPRESSIONS. THE SURGEON STATED THAT THE PATIENT LIKELY EXPERIENCED EITHER A MYOCARDIAL INFARCTION OR A PULMONARY EMBOLISM. THE EXACT CAUSE OF DEATH IS UNKNOWN. THE PATIENT HAD A CARDIAC HISTORY, FIVE PREVIOUSLY PLACED STENTS, AND BEEN CLEARED FROM A CARDIAC STANDPOINT AND PULMONOLOGY STANDPOINT PRIOR TO SURGERY. AN AUTOPSY WAS NOT PERFORMED. THE SURGEON STATED THAT THE DA VINCI PRODUCTS HAD NOTHING TO DO WITH THIS EVENT. Manufacturer Narrative: THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: 30 DEGREE ENDOSCOPE, MONOPOLAR CURVED SCISSORS, TIP-UP FENESTRATED GRASPER, SUREFORM STAPLER 60, FENESTRATED BIPOLAR FORCEPS, MEGA SUTURECUT NEEDLE DRIVER, VESSEL SEALER EXTEND. A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED FOR THE DEVICE(S) USED DURING THE PROCEDURE AND NO NON-CONFORMANCES WERE IDENTIFIED TO BE RELATED TO THIS EVENT. A REVIEW OF THE SYSTEM LOGS FOUND THAT NO RELEVANT ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT DIED DURING A RIGHT HEMICOLECTOMY AFTER THE COLON HAD ALREADY BEEN RESECTED. THE CAUSE OF DEATH HAS BEEN ATTRIBUTED TO A MYOCARDIAL EVENT AND NO OTHER INTRAOPERATIVE ISSUES WERE REPORTED. NO ALLEGATION HAS BEEN MADE ABOUT ANY INTUITIVE SURGICAL PRODUCTS. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, NO INTUITIVE SURGICAL PRODUCT OR INSTRUMENT CONTRIBUTED TO THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21701868&pc=BZD | 2518422-2025-038905 | 2025-03-03 05:00 | Death | RESPIRONICS, INC. | 2025-03-26 04:00 | BZD | DREAMSTATION AUTO CPAP | Degraded | Insufficient Information; No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE "ADVERSE EVENT/PRODUCT PROBLEM" FIELD, "TYPE OF REPORTED COMPLAINT", PATIENT OUTCOME CODE GRID AND EVALUATION METHOD CODE GRID HAS BEEN CORRECTED IN THIS REPORT. IN THIS REPORT, BOX B, H HAS BEEN UPDATED/CORRECTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21701883&pc=NVN | 2182208-2025-01534 | 2024-07-29 04:00 | Death | MEDTRONIC, INC. | 2025-03-26 04:00 | NVN | SELECTSECURE MRI SURESCAN | Appropriate Term/Code Not Available | Insufficient Information | P030036 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING LEFT BUNDLE BRANCH AREA PACING (LBBAP) COMPARED WITH RIGHT VENTRICULAR SEPTAL PACING. THE AUTHORS DESCRIBED PATIENT DEATHS; HOWEVER, THE CAUSES OF DEATH WERE UNKNOWN AND DEFINED AS ALL-CAUSE. THERE WERE PATIENTS IN BOTH GROUPS WITH COMPLICATIONS. TWO PATIENTS EXPERIENCED PNEUMOTHORAX, ONE WITH AN INFECTION, ONE SEPTAL PERFORATION IN WHICH THE LEAD WAS REPOSITIONED, WORSENING TRICUSPID REGURGITATION, AND HEART FAILURE HOSPITALIZATIONS. THERE WERE LEADS IN BOTH GROUPS WHICH WERE REVISED DUE TO DISLODGMENTS AND INCREASES IN THRESHOLDS. THE STATUS OF THE LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: MID-TERM CLINICAL OUTCOMES OF LEFT BUNDLE BRANCH AREA PACING COMPARED TO ACCURATE RIGHT VENTRICULAR SEPTAL PACING. JOURNAL OF INTERVENTIONAL CARDIAC ELECTROPHYSIOLOGY. 2025. 68:55¿63. DOI: 10.1007/S10840-024-01890-Z. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21702096&pc=FPA | 9617594-2025-00562 | 2025-02-12 05:00 | Death | ICU MEDICAL DE MEXICO, S. DE R.L. DE C.V. | 2025-03-26 04:00 | FPA | SEE SECTION H11 | Fluid/Blood Leak | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K100576 |  | 1 | Event Description: THE EVENT INVOLVED 4 INCIDENTS WHERE THE 7" (18 CM) APPX 0.31 ML, SMALLBORE PRESSURE INFUSION (400PSIG) EXT SET W/REMV MICROCLAVE® CLEAR, CLAMP, NANOCLAVE® T-CONNECTOR (PURPLE RING), ROTATING LUER DEVICES LEAKED DURING PATIENT ADMINISTRATION OF EMERGENCY MEDICATION/RAPID BLOOD ADMINISTRATION ON THE SAME PATIENT. THE PATIENT EXPIRED. THE CUSTOMER FURTHER REPORTED THAT IT WAS UNCLEAR IF THE CONNECTOR COULDN'T HANDLE THE PRESSURE FROM THE RAPID FLUID DEVICES.Event Description: ADDITIONAL INFORMATION RECEIVED STATING 2 T-CONNECTORS STARTED LEAKING WITH ADMINISTRATION OF RAPID BLOOD ADMINISTRATION, AND T-CONNECTOR STARTED LEAKING WITH ADMINISTRATION OF FLUSH. Manufacturer Narrative: RECEIVED FOUR (4) USED A1129 SMALLBORE PRESSURE INFUSION (400PSIG) EXT SETS AND ONE (1) USED LIST#: UNKNOWN CATHETER MATED TO ONE (1) OF THE SETS FOR INSPECTION. A CRACK WAS OBSERVED ON THE FEMALE LUER ON THREE (3) OF THE A1129 SETS. A CRACK WAS OBSERVED ON THE CATHETER MATED TO THE ONE (1) USED SET WITH NO CRACK. EACH SET WAS LEAK TESTED PER PRODUCT SPECIFICATIONS. THERE WAS LEAKAGE FROM THE CRACKS ON EACH SET. THE REPORTED COMPLAINT CAN BE CONFIRMED ON THREE (3) OF THE RETURNED USED A1129 SETS. THE PROBABLE CAUSE OF THE CRACK AND SUBSEQUENT LEAKAGE IS DUE TO A MATERIAL ISSUE ON A VENDOR SUPPLIED PART. THE REPORTED COMPLAINT COULD NOT BE CONFIRMED ON THE ONE (1) USED A1129 MATED TO THE LIST#: UNKNOWN CATHETER. THE CRACK AND SUBSEQUENT LEAKAGE WAS ON A NON-ICU MEDICAL PART. THERE WERE NO DEFECTS OR ANOMALIES ON THE A1129. THE REPORTED LOT HISTORY AND A LOT OF THE PACKAGING WAS REVIEWED; NO NONCONFORMITIES WERE IDENTIFIED THAT MAY HAVE CONTRIBUTED TO THE REPORTED COMPLAINT. CORRECTIVE ACTIONS ARE IN PROCESS. ADDITIONAL INFORMATION IN B5. D9 DEVICE RETURNED TO MANUFACTURER ON 3/6/2025. Manufacturer Narrative: SECTION D1 - BRAND NAME: 7" (18 CM) APPX 0.31 ML, SMALLBORE PRESSURE INFUSION (400PSIG) EXT SET W/REMV MICROCLAVE® CLEAR, CLAMP, NANOCLAVE® T-CONNECTOR (PURPLE RING), ROTATING LUER. THE DEVICE IS REPORTED TO BE AVAILABLE FOR EVALUATION, HOWEVER, IT HAS NOT YET BEEN RECEIVED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21702229&pc=BZD | 2518422-2025-103454 | 2024-08-23 04:00 | Death | RESPIRONICS, INC. | 2025-03-26 04:00 | BZD | RENTAL-DREAMSTATION AUTO CPAP, GB | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: IT WAS REPORTED THAT THE RENTAL DREAMSTATION AUTO CPAP DEVICE WAS BEING RETURNED AND SERVICED DUE TO THE USER PASSING AWAY. THERE WAS NO ALLEGATION OF MALFUNCTION OR THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH. THERE WERE NO REPORTS OF MEDICAL INTERVENTION. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT, AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.Event Description: CORRECTION: THIS REPORT IS BEING SUBMITTED TO INCLUDE THE PRODUCT UDI. Manufacturer Narrative: CORRECTION: THIS REPORT IS BEING SUBMITTED TO INCLUDE THE PRODUCT UDI.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21702303&pc=OZD | 1220648-2025-26982 | 2025-02-26 05:00 | Death | ABIOMED, INC. | 2025-03-26 04:00 | OZD | IMPELLA 5.5 | Infusion or Flow Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS INITIALLY IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT AND TRANSFERRED TO AN IMPELLA 5.5 DEVICE FOR ESCALATION IN CARE. FOLLOWING THE IMPELLA 5.5 IMPLANT, ACCESS SITE BLEEDING AND SATURATED FLOWS WERE EXPERIENCED REQUIRING INTERVENTION. BLEEDING AT ANASTOMOSIS DURING INSERTION WAS NOTED. CONSEQUENTLY, BLOOD PRODUCTS WERE ADMINISTERED, AND PROTAMINE WAS GIVEN. APPROXIMATELY FIVE DAYS LATER, THERE WERE SIGNS OF FLOW SATURATION. FLOWS WERE INCREASED BY 1 LITER. THE IMPELLA 5.5 PUMP POSITION WAS CHECKED UNDER TRANSTHORACIC ECHOCARDIOGRAM. REPLACEMENT OF THE IMPELLA 5.5 DEVICE WAS CONTEMPLATED BUT WAS WAITING FOR CONFIRMED NEUROLOGICAL STATUS. TEAM ACKNOWLEDGED AWARENESS OF THE POTENTIAL OF AN ACUTE PUMP STOP, AND THE PATIENT CONTINUED ON SUPPORT. APPROXIMATELY, 19 DAYS LATER, IT WAS NOTED HEPARIN WAS OFF FOR POTENTIAL BLEEDING CONCERNS. NO SOURCE WAS IDENTIFIED; NO BLOOD PRODUCTS WERE GIVEN. THE LATEST UPDATE NOTED THE PATIENT TO BE STABLE CONTINUING WITH IMPELLA 5.5 DEVICE SUPPORT.Event Description: THE DECISION WAS MADE TO WITHDRAW CARE AND THE PATIENT EXPIRED. IT IS UNKNOWN IF THE USE OF THE IMPELLA WAS RELATED TO PATIENT DEATH, THUS THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING AND SATURATED FLOW HAS BEEN COMPLETED. NO PUMP WAS RETURNED FOR EVALUATION. AS PER CLINICAL, BLEEDING FROM ANASTOMOSIS DURING INSERTION AND RESOLVED WITH ADMINISTERING BLOOD PRODUCTS/PROTAMINE. THE ROOT CAUSE OF THE ACCESS SITE BLEEDING ISSUE WAS NOT DETERMINED SINCE PRODUCT WAS NOT RETURNED AND DUE TO INSUFFICIENT CLINICAL DETAILS. DATA LOGS WERE RETURNED FOR REVIEW. THE LOG INDICATING A RISE IN MOTOR CURRENT (MC) FROM 450 TO 550 MA AT SAME P-LEVEL (P-8) BETWEEN (B)(6) 2025. SATURATED FLOWS OBSERVED WITH FLOW RISING UP TO 5. L/MIN. PURGE PARAMETERS WERE NORMAL. LATER P-LEVEL WAS DROPPED TO P-6 AND FLOWS WERE PARTIALLY LOW DURING THE REST OF CASE. ALSO, FLOWS WERE SLIGHTLY SATURATED AT P-6 DURING (B)(6) 2025 WITH FLOWS UP TO 4.4L/MIN. RT LOG INDICATING SATURATED FLOWS WITH DIASTOLIC FLOW WAVEFORM MAXED AT 5.6 L/MIN ON P-8. MC RISING FROM 418 TO 523 PER RT LOG MIN/MAX VALUES. NO OTHER ABNORMALITIES. THE ROOT CAUSE OF THE SATURATED ISSUE WAS MOST LIKELY BIOMATERIAL PER LOG REVIEW. B2 OUTCOME REVISED TO REFLECT DEATH BASED ON THE ADDITIONAL INFORMATION RECEIVED FROM THE CLINICAL SITE. B5 REVISED TO REFLECT DEATH BASED ON THE ADDITIONAL INFORMATION RECEIVED FROM THE CLINICAL SITE. D6B EXPLANT DATE ADDED. H1 TYPE OF REPORTABLE EVENT REVISED TO REFLECT DEATH BASED ON THE ADDITIONAL INFORMATION RECEIVED FROM THE CLINICAL SITE. H6 HEALTH EFFECT - IMPACT CODE REVISED TO REFLECT DEATH BASED ON THE ADDITIONAL INFORMATION RECEIVED FROM THE CLINICAL SITE. E4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26982 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED, SUPPORTING THE PATIENT AT THE TIME OF THE MDR WRITING, AND THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: SECTION: GENERAL PATIENT CARE CONSIDERATIONS: ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: POTENTIAL ADVERSE EVENTS: ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY¿ (INCLUDING VENTRICULAR PERFORATION).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21702500&pc=QZI | 2124215-2025-18842 | 2025-03-03 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-03-26 04:00 | QZI | FARAWAVE | Off-Label Use | Cardiac Arrest; Heart Block; Heart Failure/Congestive Heart Failure |  |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT DEATH OCCURRED. A PULSED FIELD ABLATION (PFA) PROCEDURE TO TREAT A PERMANENT ATRIAL FIBRILLATION WAS PERFORMED WITH A FARAWAVE CATHETER. THE PATIENT HAD HEART FAILURE AND A VERY IMPAIRED EJECTION FRACTION PRE-PROCEDURE. THE PROCEDURE STARTED, THE PATIENT WAS IN ATRIAL FIBRILLATION (AF) AND VITALS WERE STABLE. DURING THE PROCEDURE, THE AF SLOWED AND THE PATIENT EVENTUALLY WENT INTO 1:1 ATRIAL FLUTTER AND THE PATIENT WAS SHOCKED. THE SINUS RHYTHM WAS RESTORED BUT WAS SLOW DUE TO THE SINUS MODE THAT HAD DIFFICULTY IN RECOVERING. THE PHYSICIAN AND NURSES HAD TO MASSAGE THE PATIENT AND GIVE NORADRENALINE TO HELP THE RECOVER. THE PROCEDURE WAS CONTINUED, BUT AFTER A FEW MINUTES THE RHYTHM DECELERATED AND THE PATIENT WENT INTO AV BLOCK. CARDIAC MASSAGE WAS GIVEN FOR A FEW MINUTES BUT THE PATIENT DID NOT RECOVERED. A TOTAL OF 57 APPLICATIONS WERE GIVEN. Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION.. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21702698&pc=OWB | 1720753-2025-00002 |  | Death | GE OEC MEDICAL SYSTEMS, INC | 2025-03-26 04:00 | OWB | OEC ELITE | Poor Quality Image | Insufficient Information | K192819 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WAS A DEATH OF A PATIENT DURING A PROLONGED EVAR PROCEDURE. DURING A SITE VISIT FOLLOWING THIS PROCEDURE THE CLINICIAN EXPRESSED CONCERNS REGARDING THE IMAGE QUALITY DURING THE PROCEDURE WHILE USING THE DIGITAL SUBTRACTED ANGIOGRAPHY MODE. THE CAUSE (INCLUDING ANY CAUSAL RELATIONSHIP BETWEEN THE INCIDENT AND DEVICE) AND CONTRIBUTING FACTORS IN THE PATIENT DEATH HAVE NOT BEEN ESTABLISHED AND INVESTIGATION IS ONGOING RELATED TO THE REPORTED IMAGE QUALITY CONCERNS AND THE PROLONGED EVAR PROCEDURE. Manufacturer Narrative: GE HEALTHCARE HAS COMPLETED ITS INVESTIGATION AND DETERMINED THAT THE ROOT CAUSE WAS USE ERROR. THE FINDINGS INDICATED THAT THE IMAGING PULSE FRAME RATE (FPS) SETTINGS WERE USED SUB-OPTIMALLY, WITHOUT ANY EVIDENCE OF PRODUCT FAILURE OR MALFUNCTION. THIS USAGE WAS INCONSISTENT WITH THE WARNINGS PROVIDED IN THE USER MANUAL. THE GENERAL DIRECTOR OF THE HOSPITAL STATED AN INTERNAL REVIEW CONCLUDED THAT THE INCIDENT WAS MULTIFACTORIAL DUE TO POOR IMAGE QUALITY AND THE PATIENT'S MEDICAL HISTORY. THE HOSPITAL'S MMR (MORBIDITY AND MORTALITY REVIEW) ROOT CAUSE IS UNDETERMINED. Manufacturer Narrative: GE HEALTHCARE INVESTIGATION INTO THE REPORTED OCCURRENCE IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION HAS BEEN COMPLETED. LEGAL MANUFACTURER: SURGERY SLC - 384 N WRIGHT BROTHERS DR. USA SALT LAKE CITY, UT 84116.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21702775&pc=NPT | 2015691-2025-02282 | 2024-11-25 05:00 | Death | EDWARDS LIFESCIENCES | 2025-03-26 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Heart Block | P140031 |  | 1 | Event Description: IT WAS REPORTED BY A FAMILY MEMBER THAT APPROXIMATELY 6 DAYS POST INDEX TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) PROCEDURE USING A 23MM SAPIEN 3 ULTRA VALVE, THE PATIENT EXPIRED. IT WAS NOTED THAT THE PATIENT HAD EXPERIENCED A PREVIOUS CARDIAC ARREST TWO YEARS EARLIER AND WAS FREQUENTLY HOSPITALIZED FOR FLUID OVERLOAD COMPLICATED BY CONGESTIVE HEART FAILURE. THE PATIENT'S HEART TEAM DECIDED THAT HE WOULD UNDERGO TAVR DUE TO HIS AORTIC STENOSIS AND INCREASING SHORTNESS OF BREATH. THE PATIENT WAS ADMITTED TO (B)(6) HOSPITAL, WHERE THE TAVR PROCEDURE WAS PERFORMED, AND A PACEMAKER WAS ALSO PLACED. HOWEVER, THE CARE ADVOCATE REPORTED THAT THE PATIENT'S SHORTNESS OF BREATH DID NOT IMPROVE DUE TO FLUID ACCUMULATION, AND HE ALSO COMPLAINED OF PAIN AT THE PACEMAKER SURGICAL SITE. THREE DAYS AFTER POST TAVR PROCEDURE, THE PATIENT SUFFERED A SECOND CARDIAC ARREST AND WAS PLACED ON THE HIGHEST FORM OF LIFE SUPPORT. SIX DAYS POST TAVR PROCEDURE, THE MEDICAL TEAM DETERMINED THAT THE PATIENT HAD NO BRAIN ACTIVITY, AND THE DECISION WAS MADE TO DISCONTINUE LIFE SUPPORT, RESULTING IN THE PATIENT'S PASSING ON THE SAME DAY. THE CAUSE OF DEATH, AS STATED ON THE PATIENT'S CERTIFICATE, WAS CARDIAC ARREST AND HYPOXIA DUE TO AORTIC STENOSIS, CHRONIC KIDNEY DISEASE, AND OBSTRUCTIVE SLEEP APNEA. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), CARDIAC ARREST IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH BALLOON VALVULOPLASTY, THE USE OF LOCAL AND/OR GENERAL ANESTHESIA, BIOPROSTHETIC HEART VALVES, AND THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. CARDIAC ARREST OCCURS WHEN A MAJOR CARDIOVASCULAR, RESPIRATORY, OR METABOLIC EVENT RESULTS IN THE INABILITY OF CARDIAC MUSCLE TO GENERATE SUFFICIENT FORCE IN RESPONSE TO ELECTRICAL DEPOLARIZATION. IT MAY BE CAUSED BY A PROFOUND CARDIOVASCULAR INSULT. EXAMPLES INCLUDE SEVERE PROLONGED HYPOXIA OR ACIDOSIS, EXTREME HYPOVOLEMIA, FLOW-RESTRICTING PULMONARY EMBOLUS, CARDIAC TAMPONADE, THROMBOSIS (CORONARY OR PULMONARY). THESE EVENTS MAY BE RELATED TO THE EDWARDS DEVICE IF IT IS ASSOCIATED WITH CARDIAC TAMPONADE, ANNULAR RUPTURE, OR OTHER EDWARDS'S DEVICE-RELATED BLEED. PER THE INSTRUCTIONS FOR USE (IFU), ARRHYTHMIAS AND CONDUCTION SYSTEM DEFECTS WHICH MAY OR MAY NOT REQUIRE A PERMANENT PACEMAKER IMPLANTATION ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH BALLOON VALVULOPLASTY, THE USE OF LOCAL AND/OR GENERAL ANESTHESIA, BIOPROSTHETIC HEART VALVES, AND THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. ACCORDING TO THE VALVE ACADEMIC RESEARCH CONSORTIUM (VARC) GUIDELINES, THE CLOSE ANATOMICAL RELATIONSHIP BETWEEN THE AORTIC VALVE COMPLEX AND THE BRANCHING ATRIOVENTRICULAR BUNDLE MAY EXPLAIN THESE COMPLICATIONS OF THE TVR PROCEDURE. ACCORDING TO THE LITERATURE REVIEW, AND AS DOCUMENTED IN EW CLINICAL TECHNICAL SUMMARY FOR COMPLAINTS, CONDUCTION DISTURBANCES/ HEART BLOCK, ATRIOVENTRICULAR CONDUCTION DISTURBANCES AFTER TVR ARE ASSOCIATED WITH MANY PATIENT RELATED AND PROCEDURAL RELATED FACTORS, INCLUDING PRE-OPERATIVE COMORBID STATUS, THE DEGREE, AND BULKINESS OF ANNULAR CALCIFICATION, INTERVENTRICULAR SEPTAL THICKNESS, HISTORY OF ELECTROCARDIOGRAM ABNORMALITIES, THE DEPTH OF PROSTHESIS IMPLANTATION, AND THE PROFILE OF THE IMPLANTED PROSTHESIS. UNLIKE CONVENTIONAL AORTIC VALVE REPLACEMENT (AVR), WHERE THERE MAY BE LOCALIZED TRAUMA DUE TO DECALCIFICATION OF THE ANNULUS AND SUTURE PLACEMENT IN THE PROXIMITY OF THE ATRIOVENTRICULAR (AV) NODE OR THE BUNDLES, THE TRANSCATHETER HEART VALVE MAY CAUSE CONDUCTION ABNORMALITIES THROUGH MECHANICAL IMPINGEMENT OF THE CONDUCTION SYSTEM BY THE PROSTHESIS. THE MECHANISMS OF THE DEVELOPMENT OF HEART BLOCK AFTER TVR ARE WELL DOCUMENTED AND DESCRIBED IN THE LITERATURE. IT IS ALSO DOCUMENTED THAT PRE-EXISTING HEART BLOCK IS COMMON IN PATIENTS UNDERGOING THV OR SURGICAL AVR AND ANOTHER 4-6 % WILL DEVELOP POSTOPERATIVE HEART BLOCK, POTENTIALLY REQUIRING A PERMANENT PACEMAKER. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST/INDICATE THAT PATIENT FACTORS (HISTORY OF HEART FAILURE ALONG WITH MULTIPLE COMORBIDITIES) MAY HAVE CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED, AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21703394&pc=MIH | 2017233-2025-05990 | 2025-03-11 04:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-03-26 04:00 | MIH | GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM | Adverse Event Without Identified Device or Use Problem | Vascular Dissection | P040043 |  | 1 | Event Description: THE FIELD SALES ASSOCIATE REPORTED THE FOLLOWING: ON (B)(6) 2025, THE PATIENT WAS IMPLANTED WITH GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM WITH A GORE® TRI-LOBE BALLOON CATHETER TO TREAT A THORACIC DISSECTION WITH AN INTRAMURAL HEMATOMA PRESENT. IT WAS REPORTED THAT AFTER DEPLOYING THE GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM (TGMR313110/(B)(6)), IMAGES REVEALED A PROXIMAL TYPE I ENDOLEAK. THE PHYSICIAN CHOSE TO USE A GORE® TRI-LOBE BALLOON CATHETER TO BALLOON THE ENDOLEAK. REPORTEDLY THE PHYSICIAN WAS ¿AGGRESSIVELY¿ BALLOONING THE PROXIMAL ENDOLEAK AND PERFORATED THE AORTA. WHILE RETRACTING THE BALLOON THE PATIENT CODED. THE PHYSICIAN IMPLANTED AN ADDITIONAL GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM (TGMR313110/(B)(6)) TO TREAT THE PERFORATED AORTA. A BLOOD TRANSFUSION WAS REQUIRED. THE PATIENT NEVER RECOVERED FROM THE CODE AND EXPIRED IN THE OPERATING ROOM. REPORTEDLY THE PATIENT HAD A FRIABLE AORTA WITH HISTORY OF CARDIAC ISSUES AND PERIPHERAL VASCULAR DISEASE. DURING THE CODE THE PATIENT DEVELOPED A PLEURAL EFFUSION FROM THE BLEEDING. Manufacturer Narrative: PATIENT MEDICATIONS: ASPIRIN, MOXIFLOXACIN, DOXYCYCLINE, CIPRO, VANCOMYCIN, PERCOCET AND PENICILLIN. THE GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM INSTRUCTIONS FOR USE (IFU) STATES THAT ADVERSE EVENTS THAT MAY OCCUR AND/OR REQUIRE INTERVENTION INCLUDE, BUT ARE NOT LIMITED TO ENDOLEAK, VASCULAR TRAUMA, AND DEATH. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21703660&pc=DQY | 3007284313-2025-03857 | 2025-03-11 04:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-03-26 04:00 | DQY | GORE® TRI-LOBE BALLOON CATHETER | Adverse Event Without Identified Device or Use Problem | Vascular Dissection | K180919 |  | 1 | Event Description: THE FIELD SALES ASSOCIATE REPORTED THE FOLLOWING: ON (B)(6) 2025, THE PATIENT WAS IMPLANTED WITH GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM WITH A GORE® TRI-LOBE BALLOON CATHETER TO TREAT A THORACIC DISSECTION WITH AN INTRAMURAL HEMATOMA PRESENT. IT WAS REPORTED THAT AFTER DEPLOYING THE GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM (TGMR313110/(B)(6), IMAGES REVEALED A PROXIMAL TYPE I ENDOLEAK. THE PHYSICIAN CHOSE TO USE A GORE® TRI-LOBE BALLOON CATHETER TO BALLOON THE ENDOLEAK. REPORTEDLY THE PHYSICIAN WAS ¿AGGRESSIVELY¿ BALLOONING THE PROXIMAL ENDOLEAK AND PERFORATED THE AORTA. WHILE RETRACTING THE BALLOON THE PATIENT CODED. THE PHYSICIAN IMPLANTED AN ADDITIONAL GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM (TGMR313110/(B)(6) TO TREAT THE PERFORATED AORTA. A BLOOD TRANSFUSION WAS REQUIRED. THE PATIENT NEVER RECOVERED FROM THE CODE AND EXPIRED IN THE OPERATING ROOM. REPORTEDLY THE PATIENT HAD A FRIABLE AORTA WITH HISTORY OF CARDIAC ISSUES AND PERIPHERAL VASCULAR DISEASE. DURING THE CODE THE PATIENT DEVELOPED A PLEURAL EFFUSION FROM THE BLEEDING. Manufacturer Narrative: -PATIENT MEDICATIONS: ASPIRIN, MOXIFLOXACIN, DOXYCYCLINE, CIPRO, VANCOMYCIN, PERCOCET AND PENICILLIN. THE GORE® TRI-LOBE BALLOON CATHETER INSTRUCTIONS FOR USE (IFU) STATES THAT ADVERSE EVENTS INCLUDE, BUT ARE NOT LIMITED TO TRAUMA TO THE VESSEL WALL, INCLUDING PERFORATION, AND DEATH W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY HAVE NOT BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21703957&pc=GBO | 2124215-2025-18575 | 2021-03-14 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-03-26 04:00 | GBO | PERCUTANEOUS NEPHROSTOMY SET | Adverse Event Without Identified Device or Use Problem | Fatigue; Muscle Weakness; Pneumonia; Sepsis; Weight Changes; Respiratory Insufficiency; Unspecified Tissue Injury |  |  | 1 | Event Description: NOTE: THIS REPORT PERTAINS TO ONE OF TWO DEVICES USED FOR THE SAME PATIENT. REFER TO MANUFACTURER REPORT # 2124215-2025-18574 FOR THE ASSOCIATED DEVICE INFORMATION. IT WAS REPORTED TO BOSTON SCIENTIFIC CORPORATION THAT A PERCUTANEOUS NEPHROSTOMY SET WAS USED TO TREAT A LEFT NEPHROLITHIASIS DURING A LEFT PERCUTANEOUS NEPHROSTOLITHOTOMY (PCNL) OF STONE GREATER THAN 2 CM PROCEDURE PERFORMED SOMETIME IN (B)(6) 2021. TWO NAVIGUIDE DEVICES WERE UTILIZED AFTER DISCOVERING ANOTHER STONE DURING A RENAL ENDOSCOPY. THIRTEEN DAYS POST-PROCEDURE, THE PATIENT DEVELOPED SEPSIS AND WAS INTUBATED AND ADMITTED TO THE INTENSIVE CARE UNIT (ICU) DUE TO A RAPID DECLINE IN RESPIRATORY FUNCTION. NORMAL SALINE BOLUS, PHENYLEPHRINE, AND ADENOSINE WERE ADMINISTERED TO MAINTAIN PERFUSION. THE PATIENT, WHO WAS CACHECTIC, WAS TRANSFERRED TO HOSPICE CARE IN THE ICU. THE PATIENT REPORTEDLY EXPIRED, BUT THE CAUSE OF DEATH WAS NOT REPORTED. PER PHYSICIAN'S ASSESSMENT, THE EVENT WAS SERIOUS, WAS POSSIBLY RELATED TO THE DEVICE, AND POSSIBLY RELATED TO THE PROCEDURE.Event Description: NOTE: THIS REPORT PERTAINS TO ONE OF TWO DEVICES USED FOR THE SAME PATIENT. REFER TO MANUFACTURER REPORT # 2124215-2025-18574 FOR THE ASSOCIATED DEVICE INFORMATION. IT WAS REPORTED TO BOSTON SCIENTIFIC CORPORATION THAT A PERCUTANEOUS NEPHROSTOMY SET WAS USED TO TREAT A LEFT NEPHROLITHIASIS DURING A LEFT PERCUTANEOUS NEPHROSTOLITHOTOMY (PCNL) OF STONE GREATER THAN 2 CM PROCEDURE PERFORMED SOMETIME IN (B)(6) 2021. TWO NAVIGUIDE DEVICES WERE UTILIZED AFTER DISCOVERING ANOTHER STONE DURING A RENAL ENDOSCOPY. THIRTEEN DAYS POST-PROCEDURE, THE PATIENT DEVELOPED SEPSIS AND WAS INTUBATED AND ADMITTED TO THE INTENSIVE CARE UNIT (ICU) DUE TO A RAPID DECLINE IN RESPIRATORY FUNCTION. NORMAL SALINE BOLUS, PHENYLEPHRINE, AND ADENOSINE WERE ADMINISTERED TO MAINTAIN PERFUSION. THE PATIENT, WHO WAS CACHECTIC, WAS TRANSFERRED TO HOSPICE CARE IN THE ICU. THE PATIENT REPORTEDLY EXPIRED, BUT THE CAUSE OF DEATH WAS NOT REPORTED. PER PHYSICIAN'S ASSESSMENT, THE EVENT WAS SERIOUS, WAS POSSIBLY RELATED TO THE DEVICE, AND POSSIBLY RELATED TO THE PROCEDURE.Event Description: NOTE: THIS REPORT PERTAINS TO ONE OF TWO DEVICES USED FOR THE SAME PATIENT. REFER TO MANUFACTURER REPORT # 2124215-2025-18574 FOR THE ASSOCIATED DEVICE INFORMATION. IT WAS REPORTED TO BOSTON SCIENTIFIC CORPORATION THAT A PERCUTANEOUS NEPHROSTOMY SET WAS USED TO TREAT A LEFT NEPHROLITHIASIS DURING A LEFT PERCUTANEOUS NEPHROSTOLITHOTOMY (PCNL) OF STONE GREATER THAN 2 CM PROCEDURE PERFORMED SOMETIME IN (B)(6) 2021. TWO NAVIGUIDE DEVICES WERE UTILIZED AFTER DISCOVERING ANOTHER STONE DURING A RENAL ENDOSCOPY. THIRTEEN DAYS POST-PROCEDURE, THE PATIENT DEVELOPED SEPSIS AND WAS INTUBATED AND ADMITTED TO THE INTENSIVE CARE UNIT (ICU) DUE TO A RAPID DECLINE IN RESPIRATORY FUNCTION. NORMAL SALINE BOLUS, PHENYLEPHRINE, AND ADENOSINE WERE ADMINISTERED TO MAINTAIN PERFUSION. THE PATIENT, WHO WAS CACHECTIC, WAS TRANSFERRED TO HOSPICE CARE IN THE ICU. THE PATIENT REPORTEDLY EXPIRED, BUT THE CAUSE OF DEATH WAS NOT REPORTED. PER PHYSICIAN'S ASSESSMENT, THE EVENT WAS SERIOUS, WAS POSSIBLY RELATED TO THE DEVICE, AND POSSIBLY RELATED TO THE PROCEDURE. ***ADDITIONAL INFORMATION RECEIVED ON MAY 13, 2025*** ADDITIONAL CLINICAL DETAILS REVEALED THAT THE PATIENT WAS INTUBATED AND SEDATED IN THE ICU, RECEIVING VASOPRESSOR SUPPORT AND BROAD-SPECTRUM ANTIBIOTICS FOR SEPSIS, WITH A DIFFERENTIAL DIAGNOSIS OF ASPIRATION PNEUMONIA VERSUS A GENITOURINARY SOURCE FOLLOWING RECENT LEFT URETERAL STENT PLACEMENT. IN ALIGNMENT WITH THE FAMILY'S EXPRESSED WISHES AND THE PATIENT'S GOALS OF CARE, A PALLIATIVE EXTUBATION WAS PURSUED. THE HEALTHCARE PROVIDER REQUESTED COMFORT-FOCUSED CARE. THE PATIENT WAS SUBSEQUENTLY EXTUBATED IN ACCORDANCE WITH THE PALLIATIVE PLAN. Manufacturer Narrative: BLOCK H6 PATIENT CODES AND IMPACT CODES CORRECTED. BLOCK H11 ADDITIONAL MFR NARRATIVE CORRECTED. BLOCK B2: THE EXACT DATE OF THE PATIENT'S DEATH IS UNKNOWN. THE PROVIDED PATIENT'S DEATH DATE OF (B)(6) 2021, WAS CHOSEN AS BEST ESTIMATE BASED ON 30 DAYS FROM THE PROCEDURE DATE WHICH HAPPENED SOMETIME IN (B)(6) 2021. BLOCK B3: THE EXACT EVENT ONSET DATE IS UNKNOWN. THE PROVIDED EVENT DATE OF MARCH 14, 2021, WAS CHOSEN AS THE BEST ESTIMATE BASED ON THE PROCEDURE WHICH HAPPENED SOMETIME IN MARCH 2021 AND THE DAY OF ADVERSE EVENT REPORTED AT 13 DAYS (POD13) POST-PROCEDURE. BLOCKS D4, H4: DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER PRODUCT SPECIFIC INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. BLOCK H6: IMDRF PATIENT CODE E0306 CAPTURES THE REPORTABLE EVENT OF "THE PATIENT DEVELOPED SEPSIS." IMDRF PATIENT CODE E0743 CAPTURES THE REPORTABLE EVENT OF "RAPID DECLINE IN RESPIRATORY FUNCTION." IMDRF IMPACT CODE F23 CAPTURES THE REPORTABLE EVENT OF "PATIENT WAS INTUBATED." IMDRF IMPACT CODE F0801 CAPTURES THE REPORTABLE EVENT OF "PATIENT ADMITTED TO THE INTENSIVE CARE UNIT (ICU)." IMDRF IMPACT CODE F2303 CAPTURES THE REPORTABLE EVENT OF "NORMAL SALINE BOLUS, PHENYLEPHRINE, AND ADENOSINE WERE ADMINISTERED TO MAINTAIN PERFUSION." IMDRF IMPACT CODE F14 CAPTURES THE REPORTABLE EVENT OF "THE PATIENT WAS TRANSFERRED TO HOSPICE CARE." IMDRF IMPACT CODE F02 CAPTURES THE REPORTABLE EVENT OF "THE PATIENT REPORTEDLY EXPIRED, BUT THE CAUSE OF DEATH WAS NOT REPORTED." Manufacturer Narrative: BLOCKS B5, H6 (PATIENT AND IMPACT CODES) AND H11 HAVE BEEN UPDATED BASED ON THE ADDITIONAL INFORMATION RECEIVED ON MAY 13, 2025. BLOCK B2: THE EXACT DATE OF THE PATIENT'S DEATH IS UNKNOWN. THE PROVIDED PATIENT'S DEATH DATE OF MARCH 31, 2021, WAS CHOSEN AS BEST ESTIMATE BASED ON 30 DAYS FROM THE PROCEDURE DATE WHICH HAPPENED SOMETIME IN MARCH 2021. BLOCK B3: THE EXACT EVENT ONSET DATE IS UNKNOWN. THE PROVIDED EVENT DATE OF MARCH 14, 2021, WAS CHOSEN AS THE BEST ESTIMATE BASED ON THE PROCEDURE WHICH HAPPENED SOMETIME IN MARCH 2021, AND THE DAY OF ADVERSE EVENT REPORTED AT 13 DAYS (POD13) POST-PROCEDURE. BLOCKS D4, H4: DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BASED ON THE NATURE OF THE INFORMATION PROVIDED TO BSC, IT IS NOT POSSIBLE TO PERFORM A GOOD FAITH EFFORT TO OBTAIN ADDITIONAL INFORMATION. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER PRODUCT SPECIFIC INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. BLOCK H6: IMDRF PATIENT CODE E0306 CAPTURES THE REPORTABLE EVENT OF "THE PATIENT DEVELOPED SEPSIS." IMDRF PATIENT CODE E0743 CAPTURES THE REPORTABLE EVENT OF "RAPID DECLINE IN RESPIRATORY FUNCTION." IMDRF PATIENT CODE E0733 CAPTURES THE REPORTABLE EVENT OF "ASPIRATION PNEUMONIA." IMDRF IMPACT CODE F23 CAPTURES THE REPORTABLE EVENT OF "PATIENT WAS INTUBATED." IMDRF IMPACT CODE F17 CAPTURES THE REPORTABLE EVENT OF "PATIENT WAS SEDATED." IMDRF IMPACT CODE F0801 CAPTURES THE REPORTABLE EVENT OF "PATIENT ADMITTED TO THE INTENSIVE CARE UNIT (ICU)." IMDRF IMPACT CODE F2303 CAPTURES THE REPORTABLE EVENT OF "NORMAL SALINE BOLUS, PHENYLEPHRINE, AND ADENOSINE WERE ADMINISTERED TO MAINTAIN PERFUSION." IMDRF IMPACT CODE F14 CAPTURES THE REPORTABLE EVENT OF "THE PATIENT WAS TRANSFERRED TO HOSPICE CARE." IMDRF IMPACT CODE F02 CAPTURES THE REPORTABLE EVENT OF "THE PATIENT REPORTEDLY EXPIRED, BUT THE CAUSE OF DEATH WAS NOT REPORTED." Manufacturer Narrative: BLOCK B2: THE EXACT DATE OF THE PATIENT'S DEATH IS UNKNOWN. THE PROVIDED PATIENT'S DEATH DATE OF (B)(6) 2021, WAS CHOSEN AS BEST ESTIMATE BASED ON 30 DAYS FROM THE PROCEDURE DATE WHICH HAPPENED SOMETIME IN (B)(6) 2021. BLOCK B3: THE EXACT EVENT ONSET DATE IS UNKNOWN. THE PROVIDED EVENT DATE OF MARCH 14, 2021, WAS CHOSEN AS THE BEST ESTIMATE BASED ON THE PROCEDURE WHICH HAPPENED SOMETIME IN (B)(6) 2021 AND THE DAY OF ADVERSE EVENT REPORTED AT 13 DAYS (POD13) POST-PROCEDURE. BLOCKS D4, H4: THE COMPLAINANT WAS UNABLE TO REPORT THE UPN AND SERIAL NUMBER; THEREFORE, THE UNIQUE IDENTIFIER (UDI) #, MANUFACTURE DATE, AND EXPIRATION DATE ARE UNKNOWN. BLOCK H6: IMDRF PATIENT CODE E0306 CAPTURES THE REPORTABLE EVENT OF "THE PATIENT DEVELOPED SEPSIS." IMDRF PATIENT CODE E0743 CAPTURES THE REPORTABLE EVENT OF "RAPID DECLINE IN RESPIRATORY FUNCTION." IMDRF IMPACT CODE F2203 CAPTURES THE REPORTABLE EVENT OF "RENAL ENDOSCOPY." IMDRF IMPACT CODE F23 CAPTURES THE REPORTABLE EVENT OF "PATIENT WAS INTUBATED." IMDRF IMPACT CODE F0801 CAPTURES THE REPORTABLE EVENT OF "PATIENT ADMITTED TO THE INTENSIVE CARE UNIT (ICU)." IMDRF IMPACT CODE F2303 CAPTURES THE REPORTABLE EVENT OF "NORMAL SALINE BOLUS, PHENYLEPHRINE, AND ADENOSINE WERE ADMINISTERED TO MAINTAIN PERFUSION." IMDRF IMPACT CODE F14 CAPTURES THE REPORTABLE EVENT OF "THE PATIENT WAS TRANSFERRED TO HOSPICE CARE." IMDRF IMPACT CODE F02 CAPTURES THE REPORTABLE EVENT OF "THE PATIENT REPORTEDLY EXPIRED, BUT THE CAUSE OF DEATH WAS NOT REPORTED."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21703958&pc=GBO | 2124215-2025-18574 | 2021-03-14 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-03-26 04:00 | GBO | PERCUTANEOUS NEPHROSTOMY SET | Adverse Event Without Identified Device or Use Problem | Fatigue; Muscle Weakness; Pneumonia; Sepsis; Weight Changes; Respiratory Insufficiency; Unspecified Tissue Injury |  |  | 1 | Event Description: NOTE: THIS REPORT PERTAINS TO ONE OF TWO DEVICES USED FOR THE SAME PATIENT. REFER TO MANUFACTURER REPORT # 2124215-2025-18575 FOR THE ASSOCIATED DEVICE INFORMATION. IT WAS REPORTED TO BOSTON SCIENTIFIC CORPORATION THAT A PERCUTANEOUS NEPHROSTOMY SET WAS USED TO TREAT A LEFT NEPHROLITHIASIS DURING A LEFT PERCUTANEOUS NEPHROSTOLITHOTOMY (PCNL) OF STONE GREATER THAN 2 CM PROCEDURE PERFORMED SOMETIME IN (B)(6) 2021. TWO NAVIGUIDE DEVICES WERE UTILIZED AFTER DISCOVERING ANOTHER STONE DURING A RENAL ENDOSCOPY. THIRTEEN DAYS POST-PROCEDURE, THE PATIENT DEVELOPED SEPSIS AND WAS INTUBATED AND ADMITTED TO THE INTENSIVE CARE UNIT (ICU) DUE TO A RAPID DECLINE IN RESPIRATORY FUNCTION. NORMAL SALINE BOLUS, PHENYLEPHRINE, AND ADENOSINE WERE ADMINISTERED TO MAINTAIN PERFUSION. THE PATIENT, WHO WAS CACHECTIC, WAS TRANSFERRED TO HOSPICE CARE IN THE ICU. THE PATIENT REPORTEDLY EXPIRED, BUT THE CAUSE OF DEATH WAS NOT REPORTED. PER PHYSICIAN'S ASSESSMENT, THE EVENT WAS SERIOUS, WAS POSSIBLY RELATED TO THE DEVICE, AND POSSIBLY RELATED TO THE PROCEDURE.Event Description: NOTE: THIS REPORT PERTAINS TO ONE OF TWO DEVICES USED FOR THE SAME PATIENT. REFER TO MANUFACTURER REPORT # 2124215-2025-18575 FOR THE ASSOCIATED DEVICE INFORMATION. IT WAS REPORTED TO BOSTON SCIENTIFIC CORPORATION THAT A PERCUTANEOUS NEPHROSTOMY SET WAS USED TO TREAT A LEFT NEPHROLITHIASIS DURING A LEFT PERCUTANEOUS NEPHROSTOLITHOTOMY (PCNL) OF STONE GREATER THAN 2 CM PROCEDURE PERFORMED SOMETIME IN (B)(6) 2021. TWO NAVIGUIDE DEVICES WERE UTILIZED AFTER DISCOVERING ANOTHER STONE DURING A RENAL ENDOSCOPY. THIRTEEN DAYS POST-PROCEDURE, THE PATIENT DEVELOPED SEPSIS AND WAS INTUBATED AND ADMITTED TO THE INTENSIVE CARE UNIT (ICU) DUE TO A RAPID DECLINE IN RESPIRATORY FUNCTION. NORMAL SALINE BOLUS, PHENYLEPHRINE, AND ADENOSINE WERE ADMINISTERED TO MAINTAIN PERFUSION. THE PATIENT, WHO WAS CACHECTIC, WAS TRANSFERRED TO HOSPICE CARE IN THE ICU. THE PATIENT REPORTEDLY EXPIRED, BUT THE CAUSE OF DEATH WAS NOT REPORTED. PER PHYSICIAN'S ASSESSMENT, THE EVENT WAS SERIOUS, WAS POSSIBLY RELATED TO THE DEVICE, AND POSSIBLY RELATED TO THE PROCEDURE.Event Description: NOTE: THIS REPORT PERTAINS TO ONE OF TWO DEVICES USED FOR THE SAME PATIENT. REFER TO MANUFACTURER REPORT # 2124215-2025-18575 FOR THE ASSOCIATED DEVICE INFORMATION. IT WAS REPORTED TO BOSTON SCIENTIFIC CORPORATION THAT A PERCUTANEOUS NEPHROSTOMY SET WAS USED TO TREAT A LEFT NEPHROLITHIASIS DURING A LEFT PERCUTANEOUS NEPHROSTOLITHOTOMY (PCNL) OF STONE GREATER THAN 2 CM PROCEDURE PERFORMED SOMETIME IN (B)(6) 2021. TWO NAVIGUIDE DEVICES WERE UTILIZED AFTER DISCOVERING ANOTHER STONE DURING A RENAL ENDOSCOPY. THIRTEEN DAYS POST-PROCEDURE, THE PATIENT DEVELOPED SEPSIS AND WAS INTUBATED AND ADMITTED TO THE INTENSIVE CARE UNIT (ICU) DUE TO A RAPID DECLINE IN RESPIRATORY FUNCTION. NORMAL SALINE BOLUS, PHENYLEPHRINE, AND ADENOSINE WERE ADMINISTERED TO MAINTAIN PERFUSION. THE PATIENT, WHO WAS CACHECTIC, WAS TRANSFERRED TO HOSPICE CARE IN THE ICU. THE PATIENT REPORTEDLY EXPIRED, BUT THE CAUSE OF DEATH WAS NOT REPORTED. PER PHYSICIAN'S ASSESSMENT, THE EVENT WAS SERIOUS, WAS POSSIBLY RELATED TO THE DEVICE, AND POSSIBLY RELATED TO THE PROCEDURE. ***ADDITIONAL INFORMATION RECEIVED ON MAY 13, 2025*** ADDITIONAL CLINICAL DETAILS REVEALED THAT THE PATIENT WAS INTUBATED AND SEDATED IN THE ICU, RECEIVING VASOPRESSOR SUPPORT AND BROAD-SPECTRUM ANTIBIOTICS FOR SEPSIS, WITH A DIFFERENTIAL DIAGNOSIS OF ASPIRATION PNEUMONIA VERSUS A GENITOURINARY SOURCE FOLLOWING RECENT LEFT URETERAL STENT PLACEMENT. IN ALIGNMENT WITH THE FAMILY'S EXPRESSED WISHES AND THE PATIENT'S GOALS OF CARE, A PALLIATIVE EXTUBATION WAS PURSUED. THE HEALTHCARE PROVIDER REQUESTED COMFORT-FOCUSED CARE. THE PATIENT WAS SUBSEQUENTLY EXTUBATED IN ACCORDANCE WITH THE PALLIATIVE PLAN. Manufacturer Narrative: BLOCKS B5, H6 (PATIENT AND IMPACT CODES) AND H11 HAVE BEEN UPDATED BASED ON THE ADDITIONAL INFORMATION RECEIVED ON MAY 13, 2025. BLOCK B2: THE EXACT DATE OF THE PATIENT'S DEATH IS UNKNOWN. THE PROVIDED PATIENT'S DEATH DATE OF MARCH 31, 2021, WAS CHOSEN AS BEST ESTIMATE BASED ON 30 DAYS FROM THE PROCEDURE DATE WHICH HAPPENED SOMETIME IN MARCH 2021. BLOCK B3: THE EXACT EVENT ONSET DATE IS UNKNOWN. THE PROVIDED EVENT DATE OF MARCH 14, 2021, WAS CHOSEN AS THE BEST ESTIMATE BASED ON THE PROCEDURE WHICH HAPPENED SOMETIME IN MARCH 2021, AND THE DAY OF ADVERSE EVENT REPORTED AT 13 DAYS (POD13) POST-PROCEDURE. BLOCKS D4, H4: DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BASED ON THE NATURE OF THE INFORMATION PROVIDED TO BSC, IT IS NOT POSSIBLE TO PERFORM A GOOD FAITH EFFORT TO OBTAIN ADDITIONAL INFORMATION. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER PRODUCT SPECIFIC INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. BLOCK H6: IMDRF PATIENT CODE E0306 CAPTURES THE REPORTABLE EVENT OF "THE PATIENT DEVELOPED SEPSIS." IMDRF PATIENT CODE E0743 CAPTURES THE REPORTABLE EVENT OF "RAPID DECLINE IN RESPIRATORY FUNCTION." IMDRF PATIENT CODE E0733 CAPTURES THE REPORTABLE EVENT OF "ASPIRATION PNEUMONIA." IMDRF IMPACT CODE F23 CAPTURES THE REPORTABLE EVENT OF "PATIENT WAS INTUBATED." IMDRF IMPACT CODE F17 CAPTURES THE REPORTABLE EVENT OF "PATIENT WAS SEDATED." IMDRF IMPACT CODE F0801 CAPTURES THE REPORTABLE EVENT OF "PATIENT ADMITTED TO THE INTENSIVE CARE UNIT (ICU)." IMDRF IMPACT CODE F2303 CAPTURES THE REPORTABLE EVENT OF "NORMAL SALINE BOLUS, PHENYLEPHRINE, AND ADENOSINE WERE ADMINISTERED TO MAINTAIN PERFUSION." IMDRF IMPACT CODE F14 CAPTURES THE REPORTABLE EVENT OF "THE PATIENT WAS TRANSFERRED TO HOSPICE CARE." IMDRF IMPACT CODE F02 CAPTURES THE REPORTABLE EVENT OF "THE PATIENT REPORTEDLY EXPIRED, BUT THE CAUSE OF DEATH WAS NOT REPORTED." Manufacturer Narrative: BLOCK B2: THE EXACT DATE OF THE PATIENT'S DEATH IS UNKNOWN. THE PROVIDED PATIENT'S DEATH DATE OF (B)(6) 2021, WAS CHOSEN AS BEST ESTIMATE BASED ON 30 DAYS FROM THE PROCEDURE DATE WHICH HAPPENED SOMETIME IN (B)(6) 2021. BLOCK B3: THE EXACT EVENT ONSET DATE IS UNKNOWN. THE PROVIDED EVENT DATE OF MARCH 14, 2021, WAS CHOSEN AS THE BEST ESTIMATE BASED ON THE PROCEDURE WHICH HAPPENED SOMETIME IN (B)(6) 2021 AND THE DAY OF ADVERSE EVENT REPORTED AT 13 DAYS (POD13) POST-PROCEDURE. BLOCKS D4, H4: THE COMPLAINANT WAS UNABLE TO REPORT THE UPN AND SERIAL NUMBER; THEREFORE, THE UNIQUE IDENTIFIER (UDI) #, MANUFACTURE DATE, AND EXPIRATION DATE ARE UNKNOWN. BLOCK H6: IMDRF PATIENT CODE E0306 CAPTURES THE REPORTABLE EVENT OF "THE PATIENT DEVELOPED SEPSIS." IMDRF PATIENT CODE E0743 CAPTURES THE REPORTABLE EVENT OF "RAPID DECLINE IN RESPIRATORY FUNCTION." IMDRF IMPACT CODE F2203 CAPTURES THE REPORTABLE EVENT OF "RENAL ENDOSCOPY." IMDRF IMPACT CODE F23 CAPTURES THE REPORTABLE EVENT OF "PATIENT WAS INTUBATED." IMDRF IMPACT CODE F0801 CAPTURES THE REPORTABLE EVENT OF "PATIENT ADMITTED TO THE INTENSIVE CARE UNIT (ICU)." IMDRF IMPACT CODE F2303 CAPTURES THE REPORTABLE EVENT OF "NORMAL SALINE BOLUS, PHENYLEPHRINE, AND ADENOSINE WERE ADMINISTERED TO MAINTAIN PERFUSION." IMDRF IMPACT CODE F14 CAPTURES THE REPORTABLE EVENT OF "THE PATIENT WAS TRANSFERRED TO HOSPICE CARE." IMDRF IMPACT CODE F02 CAPTURES THE REPORTABLE EVENT OF "THE PATIENT REPORTEDLY EXPIRED, BUT THE CAUSE OF DEATH WAS NOT REPORTED." Manufacturer Narrative: BLOCK H6 PATIENT CODES AND IMPACT CODES CORRECTED. BLOCK H11 ADDITIONAL MFR NARRATIVE CORRECTED. BLOCK B2: THE EXACT DATE OF THE PATIENT'S DEATH IS UNKNOWN. THE PROVIDED PATIENT'S DEATH DATE OF MARCH 31, 2021, WAS CHOSEN AS BEST ESTIMATE BASED ON 30 DAYS FROM THE PROCEDURE DATE WHICH HAPPENED SOMETIME IN (B)(6) 2021. BLOCK B3: THE EXACT EVENT ONSET DATE IS UNKNOWN. THE PROVIDED EVENT DATE OF MARCH 14, 2021, WAS CHOSEN AS THE BEST ESTIMATE BASED ON THE PROCEDURE WHICH HAPPENED SOMETIME IN (B)(6) 2021 AND THE DAY OF ADVERSE EVENT REPORTED AT 13 DAYS (POD13) POST-PROCEDURE. BLOCKS D4, H4: DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER PRODUCT SPECIFIC INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. BLOCK H6: IMDRF PATIENT CODE E0306 CAPTURES THE REPORTABLE EVENT OF "THE PATIENT DEVELOPED SEPSIS." IMDRF PATIENT CODE E0743 CAPTURES THE REPORTABLE EVENT OF "RAPID DECLINE IN RESPIRATORY FUNCTION." IMDRF IMPACT CODE F23 CAPTURES THE REPORTABLE EVENT OF "PATIENT WAS INTUBATED." IMDRF IMPACT CODE F0801 CAPTURES THE REPORTABLE EVENT OF "PATIENT ADMITTED TO THE INTENSIVE CARE UNIT (ICU)." IMDRF IMPACT CODE F2303 CAPTURES THE REPORTABLE EVENT OF "NORMAL SALINE BOLUS, PHENYLEPHRINE, AND ADENOSINE WERE ADMINISTERED TO MAINTAIN PERFUSION." IMDRF IMPACT CODE F14 CAPTURES THE REPORTABLE EVENT OF "THE PATIENT WAS TRANSFERRED TO HOSPICE CARE." IMDRF IMPACT CODE F02 CAPTURES THE REPORTABLE EVENT OF "THE PATIENT REPORTEDLY EXPIRED, BUT THE CAUSE OF DEATH WAS NOT REPORTED."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21703989&pc=NIP | 9681442-2025-00095 | 2025-03-14 04:00 | Death | ANGIOMED GMBH & CO. MEDIZINTECHNIK KG | 2025-03-26 04:00 | NIP | LIFESTENT 5F VASCULAR STENT | Break; Difficult to Remove; Detachment of Device or Device Component | Hemorrhage/Bleeding; Multiple Organ Failure | P070014 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE LITIGATION PROCESS, THAT DURING A STENT PLACEMENT PROCEDURE, THE STENT RELEASE SYSTEM (CATHETER) BROKE AND DETACHED, THEN BECAME STUCK IN THE PATIENT. THE HCP WAS UNABLE TO RETRIEVE THE DETACHED SEGMENTS, A SURGICAL PROCEDURE WAS REQUIRED. REPORTEDLY, HIGH BLOOD LOSS WAS DOCUMENTED. FORTY HOURS AFTER THE INCIDENT, THE PATIENT EXPIRED DUE TO MULTI-ORGAN FAILURE. THE DEVICE IS CURRENTLY UNAVAILABLE FOR RETURN AS IT HAS BEEN RETAINED BY THE LOCAL AUTHORITIES FOR INVESTIGATION FOLLOWING THE DECLARATION OF AN UNNATURAL DEATH.Event Description: IT WAS REPORTED THROUGH THE LITIGATION PROCESS, THAT DURING A STENT PLACEMENT PROCEDURE, THE STENT RELEASE SYSTEM (CATHETER) BROKE AND DETACHED, THEN BECAME STUCK IN THE PATIENT. THE HCP WAS UNABLE TO RETRIEVE THE DETACHED SEGMENTS, A SURGICAL PROCEDURE WAS REQUIRED. REPORTEDLY, HIGH BLOOD LOSS WAS DOCUMENTED. FORTY HOURS AFTER THE INCIDENT, THE PATIENT EXPIRED DUE TO MULTI-ORGAN FAILURE. THE DEVICE IS CURRENTLY UNAVAILABLE FOR RETURN AS IT HAS BEEN RETAINED BY THE LOCAL AUTHORITIES FOR INVESTIGATION FOLLOWING THE DECLARATION OF AN UNNATURAL DEATH. Manufacturer Narrative: H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE LIFESTENT VASCULAR STENT THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE LIFESTENT VASCULAR STENT ARE IDENTIFIED IN D2 AND G4. H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10:D4 (EXPIRY DATE: 07/2026). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. Manufacturer Narrative: H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE LIFESTENT VASCULAR STENT THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE LIFESTENT VASCULAR STENT ARE IDENTIFIED IN D2 AND G4. H10: MANUFACTURING REVIEW: BASED ON THE INFORMATION AVAILABLE IT IS NOT REASONABLY SUGGESTED THAT A MANUFACTURING PROCESS MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED ISSUE. HOWEVER, THE LOT HISTORY RECORDS OF THIS LOT WERE REVIEWED WITH SPECIAL ATTENTION TO THE MANUFACTURING AND INSPECTION OF THIS PRODUCT AND THE PRODUCT WAS FOUND TO HAVE MET THE SPECIFICATION PRIOR TO SHIPMENT. INVESTIGATION SUMMARY: THE VASCULAR STENT SYSTEM USED BY THE CUSTOMER IS NOT AVAILABLE FOR EVALUATION. NO IMAGES WERE PROVIDED. BASED ON THE INFORMATION AVAILABLE, THE INVESTIGATION HAS BEEN CLOSED WITHOUT A DEFINITIVE CONCLUSION. A DEFINITIVE CAUSE FOR THE ALLEGED DEVICE DEFICIENCY COULD NOT BE DETERMINED. LABELING REVIEW: RELEVANT LABELING APPLICABLE FOR THIS PRODUCT WAS REVIEWED. POTENTIAL COMPLICATIONS AND ADVERSE EVENTS, WHICH MAY OCCUR DURING USE OF THE LIFESTENT¿ 5F VASCULAR STENT SYSTEM ARE ADDRESSED BY THE INSTRUCTIONS FOR USE, WHICH INCLUDE BUT ARE NOT LIMITED TO OPEN SURGICAL INTERVENTION. THE CORRECT USE OF THE DEVICE IN PARTICULAR PROPER STENT DEPLOYMENT WAS FOUND TO BE ADDRESSED. H10: G3. H11: H6 (METHOD). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21704038&pc=LWS | 9614453-2025-01204 | 2025-02-03 05:00 | Death | MEDTRONIC EUROPE SARL | 2025-03-26 04:00 | LWS | EVERA MRI XT DR SURESCAN | Adverse Event Without Identified Device or Use Problem | Septic Shock; Cardiogenic Shock | P980016 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO SEPTIC SHOCK AND CARDIOGENIC SHOCK. THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) SYSTEM WAS EXPLANTED. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE DEVICE WAS RETURNED AND ANALYZED. RETURNED PRODUCT ANALYSIS WAS PERFORMED AND NO ANOMALIES WERE FOUND. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21704049&pc=DTB | 2649622-2025-08344 | 2025-02-03 05:00 | Death | MPRI | 2025-03-26 04:00 | DTB | CAPSUREFIX NOVUS MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Septic Shock; Cardiogenic Shock | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THAT THE PATIENT PASSED AWAY DUE TO SEPTIC SHOCK AND CARDIOGENIC SHOCK. THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) SYSTEM WAS EXPLANTED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE PROXIMAL PORTION OF THE LEAD WAS RETURNED, ANALYZED, AND NO ANOMALIES WERE FOUND. VISUAL ANALYSIS OF THE LEAD INDICATED APPARENT EXPLANT DAMAGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21704071&pc=NVY | 2649622-2025-08345 | 2025-02-03 05:00 | Death | MPRI | 2025-03-26 04:00 | NVY | SPRINT QUATTRO SECURE MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Septic Shock; Cardiogenic Shock | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THAT THE PATIENT PASSED AWAY DUE TO SEPTIC SHOCK AND CARDIOGENIC SHOCK. THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) SYSTEM WAS EXPLANTED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE PROXIMAL PORTION OF THE LEAD WAS RETURNED, ANALYZED, AND NO ANOMALIES WERE FOUND. VISUAL ANALYSIS OF THE LEAD INDICATED APPARENT EXPLANT DAMAGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21704293&pc=MSD | 3011642792-2025-00019 | 2024-12-24 05:00 | Death | MERIT MEDICAL SYSTEMS INC - MEXICO | 2025-03-26 04:00 | MSD | DIALYSIS.DURAMAX.CATHETER.CHRONIC HEMODIALYSIS CATHETER.15.5 | Difficult to Remove; Material Separation; Separation Failure; Activation, Positioning or Separation Problem | Air Embolism; Airway Obstruction; Bacterial Infection; Cardiac Arrest; Hemorrhage/Bleeding; Hypoxia; Pulmonary Dysfunction; Twitching; Brain Injury; Ventilator Dependent; Obstruction/Occlusion; Low Oxygen Saturation; Decreased Respiratory Rate; Convulsion/Seizure; Epistaxis | K101843 |  | 1 | Event Description: MW5166325 THE ACCOUNT ALLEGES THAT A 55-YEAR-OLD MALE PATIENT WAS ADMITTED ON (B)(6) 2024, "STATUS POST" [S/P] SEPSIS DUE TO PNEUMONIA. ON (B)(6) 2024, THE PATIENT ARRIVED IN IR [INTERVENTIONAL RADIOLOGY] FOR A TUNNELED HEMODIALYSIS CATHETER PLACEMENT. PER RADIOLOGY NOTE: "THE PATIENT HAD 02 DESATURATIONS IMMEDIATELY AFTER PLACEMENT OF A TUNNELED HD CATHETER". THE TEAR-AWAY PORTION OF THE SHEATH HAD ALLEGEDLY MALFUNCTIONED, AND THE PHYSICIAN HAD TO REMOVE THE HD CATHETER, RESULTING IN A SUSPECTED AIR EMBOLUS. A RAPID RESPONSE WAS CALLED. THE PATIENT WAS PLACED IN THE LEFT LATERAL POSITION (LLP), WHERE SUPPLEMENTAL OXYGEN WAS ADMINISTERED. THE PHYSICIAN SUSPECTED A COMPONENT OF AN AIR EMBOLISM VERSUS OBSTRUCTIVE AIRWAY OR A COMBINATION OF BOTH. EVENTUALLY, THE PATIENT'S OXYGEN SATURATIONS IMPROVED WITH THE NON-REBREATHER MASK [NRM]. THE PATIENT WAS TRANSFERRED BACK TO THE ICU ON 100% OXYGEN AND STILL REQUIRED INTUBATION TUBE PLACEMENT DUE TO A COMPROMISED AIRWAY FROM ATTEMPTS AT PLACING A NASAL TRUMPET TO ASSIST WITH THE BAG-VALVE-MASK RESPIRATIONS. THE NASAL TRUMPET PLACED IN THE RIGHT NOSTRIL DURING THE PROCEDURE LED TO EPISTAXIS. CXR DEMONSTRATED "DIFFUSE BILATERAL AIRSPACE DISEASE". AT 11:15, A CODE BLUE WAS CALLED, AND THE PHYSICIANS RESPONDED PROMPTLY. ACLS PROTOCOLS WERE UNDERWAY WHEN THE PHYSICIANS ARRIVED. THE CRITICAL CARE TEAM WAS IN THE ROOM PERFORMING BRONCHOSCOPY WITH CONCERNS FOR UPPER AIRWAY OBSTRUCTION. PEA (PULSELESS ELECTRICAL ACTIVITY) EVENTUALLY OCCURRED. THE PATIENT WAS ADMINISTERED EPINEPHRINE X3, SODIUM BICARBONATE X1, CALCIUM CHLORIDE X1; ROSC (RETURN OF SPONTANEOUS CIRCULATION) WAS ACHIEVED. THE CHAPLAIN WAS PRESENT AT 12:16. BRONCHIAL EXAMINATION WAS NORMAL. NO ENDOBRONCHIAL NODULARITY OR MASSES WERE IDENTIFIED. (B)(6) 2024, THE PATIENT DEVELOPED FACIAL TWITCHING AND WAS BLEEDING FROM THE ORAL CAVITY, REQUIRING SUCTIONING. THE PATIENT REMAINS ON FULL VENT SUPPORT AND IS OFF PRESSORS. PATIENT IS ON CONTINUOUS RENAL REPLACEMENT THERAPY [CRRT]. LOWER RESPIRATORY CULTURES WITH NORMAL RESPIRATORY FLORA, AND A FEW GRAM-NEGATIVE BACILLI WERE ALSO NOTED. (B)(6) - CRX (CHEST X-RAY) DEMONSTRATED IMPROVING BILATERAL INTERSTITIAL ALVEOLAR INFILTRATES. EEG STUDY DEMONSTRATED RECURRENT FOCAL LEFT HEMISPHERIC MOTOR SEIZURES AND DIFFUSE CEREBRAL DYSFUNCTION. NEUROLOGY WAS CONSULTED, AND THE PATIENT IS NOW ON CONTINUOUS EEG, AND PHENOBARBITAL WAS ADMINISTERED. PACKED RED BLOOD CELLS [PRBC] WERE TRANSFUSED. DECEMBER 29TH - NURSING REPORT STATES: WITH THE WEANING OFF THE PROPOFOL, THE PATIENT APPEARED TO DEVELOP TWITCHING AND REPETITIVE MOVEMENTS OF FACIAL MUSCLES. RESPIRATORY CULTURES GROWING ACINETOBACTER BAUMANNII AND PSEUDOMONAS WERE NOTED. EEG WITH RECURRENT FOCAL LEFT HEMISPHERIC MOTOR SEIZURES AND DIFFUSE CEREBRAL DYSFUNCTION NOTED. (B)(6) - NEURO DEPT. PACKED THE PATIENT'S NOSE, HEMOSTASIS ACHIEVED. PACKED RED BLOOD CELLS [PRBC] TRANSFUSED. A DIC (DISSEMINATED INTRAVASCULAR COAGULATION) PANEL WAS ORDERED. (B)(6) 2024, THE PATIENT REMAINS INTUBATED ON VENT SUPPORT WITH CONTINUED SEIZURE ACTIVITY REQUIRING BLOOD TRANSFUSIONS EVERY OTHER DAY. THE FAMILY DECIDED TO PURSUE COMFORT MEASURES ONLY [CMO]. PATIENT'S TIME OF DEATH 11:34. Manufacturer Narrative: THE SUSPECT MEDICAL DEVICE WAS NOT RETURNED FOR EVALUATION. THE COMPLAINT IS CONFIRMED BUT THE ROOT CAUSE COULD NOT BE DETERMINED. THE DEVICE HISTORY RECORD WAS REVIEWED, AND NO EXCEPTION DOCUMENTS WERE FOUND. A SEARCH OF THE COMPLAINT DATABASE WAS PERFORMED AND NO SIMILAR COMPLAINTS FOR THIS LOT NUMBER WERE IDENTIFIED. SHOULD THE DEVICE BE RETURNED LATER, THE INVESTIGATION WILL BE REOPENED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21686206&pc=OZD | 1220648-2025-26926 | 2025-02-27 05:00 | Death | ABIOMED, INC. | 2025-03-25 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hematoma | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WAS ON IMPELLA CP SUPPORT AND DURING SUPPORT, THE PATIENT DEVELOPED A HEMATOMA AT THE ACCESS SITE. FEMSTOP WAS PLACE, PACKED RED BLOOD CELLS AND PLATELETS WERE GIVEN AND HEPARIN WAS WITHHELD. SUPPORT CONTINUED BUT THE PATIENT REMAINED CRITICAL. HOWEVER, CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21686534&pc=OZD | 1220648-2025-26930 | 2025-02-26 05:00 | Death | ABIOMED, INC. | 2025-03-25 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Ischemia | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A 68 YEAR FEMALE PATIENT HAD LOST PULSES IN THE LOWER EXTREMITY FOLLOWING IMPELLA CP IMPLANT. THE PATIENT'S FOOT WAS COLD TO THE TOUCH AND FULL STRENGTH FLOLAN WAS ADDED TO ASSIST WITH THE CONDITION. BILATERAL PERFUSION CATHETERS WERE PLACED AND SUPPORT CONTINUED WITH THE IMPLANTED PUMP. FIVE DAYS LATER, AN EMERGENT CT SCAN SHOWED THAT THE PATIENT ALSO SUFFERED A BRAIN BLEED. THE FAMILY DECIDED TO WITHDRAW CARE AND THE PATIENT PASSED AWAY. THERE IS NO ENOUGH EVIDENCE TO DISASSOCIATE THE USE OF THE IMPELLA WITH THE PATIENT'S PASSING. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMART ASSIST SYSTEM STATES THE FOLLOWING: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21686547&pc=LWR | 2025587-2025-02084 | 2025-02-27 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-25 04:00 | LWR | FREESTYLE AORTIC ROOT BIOPROSTHESIS | Insufficient Information | Insufficient Information | P970031 |  | 1 | Event Description: MEDTRONIC RECEIVED INFORMATION THAT 10 YEARS AND 6 MONTHS POST IMPLANT OF THIS AORTIC BIOPROSTHETIC VALVE, THE PATIENT IS BEING EVALUATED FOR A TRANSCATHETER VALVE-IN-VALVE REPLACEMENT. THE REASON FOR EVALUATION WAS NOT REPORTED. AN INTERVENTION WAS SCHEDULED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. SUBSEQUENTLY, 3 DAYS LATER, THE PATIENT DIED. IT WAS REPORTED THAT THE PHYSICIAN DOES NOT BELIEVE THE MEDTRONIC PRODUCT CAUSED OR CONTRIBUTED TO THE DEATH. THE CAUSE OF DEATH WAS NOT RECEIVED. THEREFORE, RELATEDNESS OF THE DEATH TO THE VALVE COULD NOT BE EXCLUDED. IT IS UNKNOWN WHETHER AN AUTOPSY WAS PERFORMED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21686732&pc=MIH | 9612164-2025-01520 | 2025-03-05 05:00 | Death | MEDTRONIC IRELAND | 2025-03-25 04:00 | MIH | ENDURANT IIS BIFURCATED STENT GRAFT | Malposition of Device; Material Deformation | Ischemia; Obstruction/Occlusion | P100021 |  | 1 | Event Description: AN ENDURANT IIS BIFURCATE STENT GRAFT SYSTEM WAS IMPLANTED IN THE ENDOVASCULAR TREATMENT OF BILATERAL COMMON ILIAC ANEURYSMS. DURING THE INDEX PROCEDURE THE DISTAL AORTA WAS REINFORCED BY USING TWO 8- OR 9-MM NON-MDT BALLOONS AT THE COMMON AND EXTERNAL ILIAC ARTERIES AFTER DEPLOYING THE LIMB EXTENSIONS. THE PHYSICIAN ALSO USED TWO 9 MM NON-MDT BILIARY STENTS AND THEN USED TWO NON-MDT BALLOONS TO REINFORCE THE DISTAL AORTA. THIS WAS DONE BECAUSE THE DISTAL AORTA WAS ABOUT 18MM IN DIAMETER. IT WAS NOTED THAT GATE DEPLOYMENT WAS CROSSED DUE TO WIRES CROSSING. THE RIGHT ILIAC EXTENSION ON THE RIGHT ILIAC DID NOT DEPLOY AT MAX OVERLAP (50 MM) AND WAS DEPLOYED ABOUT 10 MM BELOW THE FLOW DIVIDER. IT WAS REPORTED THAT LATER THAT EVENING, INTERVENTION WAS PERFORMED AS IT WAS REPORTED THE LIMB HAD SHUT DOWN. BOTH LIMBS WERE IDENTICAL AND IT WAS NOT POSSIBLE TO IDENTIFY WHICH SERIAL NUMBER IS FOR THE LEFT VS RIGHT ILIAC. DURING THE INTERVENTION, ACCESS WAS ESTABLISHED AND AN ANGIOGRAM WAS DONE. A NARROWING OF THE CONTRALATERAL GATE COULD BE VISUALIZED. THREE NON-MDT STENTS WERE IMPLANTED AT THE FLOW DIVIDER AND INTO THE BILATERAL ILIAC LIMBS. A FINAL ANGIOGRAM SHOWED IMPROVEMENT. THE FOLLOWING DAY THE PATIENT EXPIRED. PER THE PHYSICIAN THE CAUSE OF THE DEFORMATION AND SUBSEQUENT DEATH WAS ANATOMY. IT WAS SAID THE IMA WAS OCCLUDED BY THE BIFURCATE GRAFT. BOTH HYPOGASTRIC ARTERIES WERE SEVERELY DISEASED AND ALMOST OCCLUDED. THE LIMB EXTENSIONS WERE LANDED IN THE EXTERNAL ILIAC ARTERIES. IT IS BELIEVED POTENTIAL OVERSIZING AND POTENTIAL ISCHEMIC ISSUES COULD HAVE CONTRIBUTED. Manufacturer Narrative: FILM ANALYSIS REVIEW THE REPORTED DEFORMATION OF THE STENT GRAFT COULD NOT BE FULLY APPRECIATED ON THE IMAGES PROVIDED DUE TO THE QUALITY OF THE IMAGES RECEIVED. THE IMAGES RECEIVED DID NOT ALLOW FOR REVIEW OF THE OCCLUSION OF PATIENT VESSELS. POST-IMPLANT CT'S WERE NOT AVAILABLE FOR REVIEW AND THE STENT GRAFT IN-VIVO CONFIGURATION COULD NOT BE EVALUATED. THE PATIENT'S ANATOMY AT IMPLANT WAS NOTED AS BEING NARROWED IN THE DISTAL AORTIC AREA AT 18 MM AND SO IT IS LIKELY THAT THE PATIENT'S ANATOMY AND THE NOTED POTENTIAL OVERSIZING AND ISCHEMIC ISSUES CONTRIBUTED TO THE REPORTED EVENTS. ANALYSIS OF THE RETURNED IMAGES DID NOT REVEAL ANY OBVIOUS OUT OF SPECIFICATION STENT GRAFT INTEGRITY ISSUES. CONTINUATION OF D10: SECTION D INFORMATION REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER RELEVANT DEVICE(S) ARE: PRODUCT ID: ETLW1610C146E, SERIAL/LOT #: (B)(6), UBD: 29-OCT-2026, UDI#: (B)(4); PRODUCT ID: ETLW1610C146E, SERIAL/LOT #: (B)(6), UBD: 02-MAY-2026, UDI#: (B)(4). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21686778&pc=OZD | 1220648-2025-26933 | 2025-02-24 05:00 | Death | ABIOMED, INC. | 2025-03-25 04:00 | OZD | IMPELLA CP | Restricted Flow rate | Hemorrhage/Bleeding; Ventricular Fibrillation | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WITH ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS INITIALLY IMPLANTED WITH AN IMPELLA CP FOLLOWED BY RIGHT-SIDE SUPPORT VIA AN IMPELLA RP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS). PRIOR TO THE IMPELLAS BEING IMPLANTED, THE PATIENT WAS UNDERGOING A TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE AND EXPERIENCED A DISSECTION OF THE RIGHT CORONARY ARTERY. THE PATIENT RAPIDLY DECOMPENSATED, VENTRICULAR FIBRILLATION ARRESTED REQUIRING AT LEAST 20 SHOCKS. THE VALVE WAS DEPLOYED IN THE DESCENDING AORTA AND COVERED WITH A STENT. THE IMPELLA CP DEVICE WAS THEN IMPLANTED VIA THE RIGHT FEMORAL ARTERY. A STENT WAS PLACED IN THE RIGHT CORONARY ARTERY BUT NO FLOW. THE PATIENT WAS TAKEN TO THE UNIT ON IMPELLA CP MCS. LATER THIS DAY, BLEEDING FROM ALL ACCESS SITES, MOUTH AND NOSE WAS NOTED. ONE UNIT OF RED BLOOD CELLS WAS ADMINISTERED. THE PATIENT HAD DECREASED FLOWS NOTED AND ECHOCARDIOGRAM SHOWED RIGHT VENTRICULAR FAILURE. THE PATIENT WENT INTO VENTRICULAR FIBRILLATION ARREST REQUIRING THREE SHOCKS. THE DECISION WAS MADE TO PLACE AN IMPELLA RP DEVICE VIA THE RIGHT FEMORAL VEIN, WHICH WAS SUCCESSFULLY COMPLETED. THE IMPELLA CP SUPPORT WAS ABLE TO BE INCREASED WITHOUT SUCTION. ACCORDING TO THE REPORT, THE PATIENT REMAINED CRITICALLY ILL, AND THE FOLLOWING DAY IT WAS NOTED THE PATIENT WAS STILL BLEEDING FOR ALL SITES AND BEING TRANSFUSED. TWO DAYS LATER, THE PATIENT HAD A TEMPORARY ATRIAL PACEMAKER LEAD PLACED AND FIVE DAYS THEREAFTER THE PATIENT'S OUTCOME AT EXPLANT WAS NOTED AS "EXPIRED." NO FURTHER DETAILS SURROUNDING THE DEMISE OUTCOME WERE MADE AVAILABLE. THE CAUSE OF DEATH WAS NOT PROVIDED. THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA CP DEVICE AS AN ASSOCIATED FACTOR TO THE OUTCOME (DEMISE) GIVEN SITE BLEEDING, THE PERIOD OF TIME WITH DECREASE FLOWS, AND CARDIAC ARRHYTHMIA REQUIRING CARDIOVERSION WHICH MAY HAVE ADDED, TO WHAT APPEARED TO HAVE BEEN, AN ALREADY COMPLEX SITUATION. Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING, LOW PUMP FLOW, AND VF/VT/AF/AT HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. CLINICAL MENTIONS BLEEDING WAS FROM MULTIPLE ACCESS SITES INCLUDING IMPELLA ACCESS SITE. NO OTHER ADDITIONAL INFORMATION RELATED TO BLEEDING WAS PROVIDED. THE ROOT CAUSE OF THE ACCESS SITE BLEEDING - MAJOR WAS MOST LIKELY PATIENT CONDITION RELATED AS EVIDENCED BY THE MULTIPLE-SITE BLEEDING INFORMATION. AS THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE VT/VF WAS NOT DETERMINED. NO PRODUCT WAS RETURNED FOR ANALYSIS. THE DATA LOGS RETURNED DID NOT COVER THE ENTIRE CP SUPPORT DURATION. THE PUMP WAS EXPLANTED ON 4/3 AND LOGS WERE ONLY PROVIDED FROM IMPLANT TO 3/28. THE LOGS DO NOT SHOW ANY MOTOR CURRENT SPIKES AND CLINICAL DID NOT REPORT ANY POSITIONING ISSUES. THERE WERE SUCTION ALARMS WHICH RESOLVED AND THE LOGS CONFIRMED LOW FLOW FOR MATCHING P-LEVELS. THE LOW FLOWS RESOLVED AFTER PUMP TURNED TO P7. THE ROOT CAUSE OF THE LOW PUMP FLOW WAS NOT DETERMINED AS NO PRODUCT WAS RETURNED FOR ANALYSIS AND LIMITED CLINICAL INFORMATION RELATED TO FAILURE WAS PROVIDED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. THIS IS ONE OF TWO MEDICAL DEVICE REPORTS ASSOCIATED WITH THIS EVENT. REFER TO INITIAL MEDICAL DEVICE REPORT FOR IMPELLA RP SERIAL NUMBER (B)(6) WITH DATE OF EVENT 25-FEB-2025 AND IDENTIFIER ENDING IN (B)(6). AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM: SECTION: GENERAL PATIENT CARE CONSIDERATIONS: "ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA." SECTION: ENTERING CARDIAC OUTPUT: USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: "REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE." SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): "ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21686906&pc=OZD | 1220648-2025-26934 | 2025-02-26 05:00 | Death | ABIOMED, INC. | 2025-03-25 04:00 | OZD | IMPELLA 5.5 | Infusion or Flow Problem | Insufficient Information | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THE PATIENT WAS ON IMPELLA 5.5 SUPPORT AND DURING SUPPORT, THERE WAS FLOW SATURATION. THERE WERE INCREASED MOTOR CURRENTS. SUPPORT CONTINUED FOR TWO DAYS AT WHICH TIME THE CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE INVESTIGATION OF SATURATED FLOW HAS BEEN COMPLETED. THE PUMP WAS RETURNED CUT AT CATHETER. UPON VISUAL INSPECTION, DRIED BLOOD WAS PRESENT ON IMPELLER AND WITHIN OUTFLOW CAGE. BIOMATERIAL WAS PRESENT WITHIN THE PURGE GAP. PUMP TEARDOWN WAS PERFORMED AND INSIDE OF MOTOR HOUSING DID NOT SHOW ANY BIOMATERIAL OR ABNORMALITIES WITHIN THE MOTOR. DATA LOGS WERE REVIEWED AND LT LOG OF CASE SHOWS PUMP FLOWS BECOMING SATURATED AROUND DAY 5 OF SUPPORT AFTER PUMP EXPERIENCED AN INSTANCE OF HIGH PURGE PRESSURE. A SLIGHT TREND UP IN MOTOR CURRENT IS OBSERVED WITH SATURATION OF FLOWS. NO MOTOR CURRENT SPIKES OBSERVED TO INDICATE BIOMATERIAL INGESTION. ¿PLACEMENT SIGNAL LOW¿ ALARMS PRESENT WITH SATURATED FLOWS. CLINICAL DETAILS NOTED THAT THE WHEN SATURATED FLOWS WERE OBSERVED, INVERTED AO/LV WAVEFORMS WERE NOTED BY CLINICAL TEAM. IT WAS NOTED THAT PATIENT HAD STOPPED SYSTEMIC HEPARIN DUE TO A RETROPERITONEAL BLEED. PATIENT REMAINED ON SUPPORT FOR AN ADDITIONAL 2 DAYS. THE ROOT CAUSE OF THE SATURATED FLOW CANNOT BE DEFINITIVELY DETERMINED AS THE PUMP WAS RETURNED CUT AND WAS UNABLE TO BE RUN IN ATTEMPT TO REPRODUCE SATURATED FLOWS. Manufacturer Narrative: THE IMPELLA DEVICE WAS RETURNED, AND AN EVALUATION IS UNDERWAY. UPON COMPLETION OF OUR ANALYSIS, A SUPPLEMENTAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21687193&pc=OZD | 1220648-2025-26935 | 2025-02-27 05:00 | Death | ABIOMED, INC. | 2025-03-25 04:00 | OZD | IMPELLA CP | Restricted Flow rate; Deformation Due to Compressive Stress | Insufficient Information | P140003 |  | 1 | Event Description: THE COMPLAINANT HAD AN IMPELLA CP PUMP PLACED TO SUPPORT A PATIENT ADMITTED WITH ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK. THE PUMP WAS EXPLANTED AFTER SEVEN HOURS DUE TO A KINKED CANNULA. THE PUMP WAS EXPLANTED AND REPLACED BY A NEW IMPELLA CP. THE PATIENT DID EXPIRE ON THE SECOND PUMP. HOWEVER, THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE FIRST IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE INVESTIGATION HAS BEEN COMPLETED. KINKED CANNULA DUE TO AORTIC ARCHE AND CAUSED LOW FLOW. THE PUMP WAS REPLACED. LOW FLOW WAS ADDED SINCE THE KINKED CANNULA WHICH CAUSED LOW FLOW PER CLINICAL DESCRIPTION. THE DATA LOGS WERE NOT RETURNED FOR REVIEW. VISUAL INSPECTION FOUND A KINK ON CANNULA ABOUT 15 MM FROM THE OF CAGE AND CANNULA BONDING SITE. NO OTHER ISSUES WERE FOUND ON THE REST OF THE PUMP. THE ROOT CAUSE OF LOW FLOW WAS MOST LIKELY KINKED CANNULA SINCE THE CLINICAL TEAM CONFIRMED THE KINK CAUSED LOW FLOW. THE ROOT CAUSE OF KINKED CANNULA WAS MOST LIKELY PATIENT CONDITION RELATED SINCE THE CLINICAL TEAM CONFIRMED THE KINK WAS DUE TO THE PATIENT HAS SHORT AORTIC ARCHE. H.6 MEDICAL DEVICE PROBLEM CODE ADDED DUE TO INVESTIGATION RESULTS. B.2 REQUIRED INTERVENTION WAS INADVERTENTLY OMITTED FROM INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26935. Manufacturer Narrative: THE IMPELLA PUMP AND PURGE CASSETTE WERE RETURNED, AND AN EVALUATION IS UNDERWAY. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21687225&pc=NVZ | 3008973940-2025-02160 | 2025-01-03 05:00 | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-03-25 04:00 | NVZ | AZURE XT DR MRI SURESCAN | Appropriate Term/Code Not Available | Bacterial Infection | P980035 |  | 1 | Event Description: IT WAS REPORTED BY A PATIENT REPRESENTATIVE THAT THE PATIENT DIED. "THE OLD DEVICE HAD BACTERIA ON IT THAT SPREAD THROUGH THEIR BODY BUT WHEN THEY PUT THE NEW DEVICE IN IT WAS TOO LATE". Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21687226&pc=OZD | 1220648-2025-26936 | 2025-02-27 05:00 | Death | ABIOMED, INC. | 2025-03-25 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A 52-YEAR-OLD MALE PATIENT WITH AN IMPELLA CP FOR A FALL PROMPTED BY ACUTE MYOCARDIAL INFARCTION WHERE HE HAD A VISIBLE HEMATOMA AND UNKNOWN NEUROLOGICAL STATUS. IT WAS REPORTED THAT THE DISTAL PULSES IN THE PATIENT'S LOWER EXTREMITY WAS LOST FOLLOWING IMPELLA CP IMPLANT. THE PHYSICIAN WAS MADE AWARE AND REVASCULARIZATION WAS PERFORMED. HOWEVER, THE FAMILY ULTIMATELY MADE THE PATIENT DO NOT RESUSCITATE (DNR) AS A RESULT OF THE OVERALL PROGNOSIS. THE PATIENT PASSED AWAY. THERE IS NOT ENOUGH EVIDENCE TO DISASSOCIATE THE PATIENT'S PASSING AND THE USE OF THE DEVICE. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH RISK PCI STATES THE FOLLOWING: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): "ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)".
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21687260&pc=NPT | 2025587-2025-02087 | 2024-10-08 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-25 04:00 | NPT | MEDTRONIC TRANSCATHETER VALVE | Inadequacy of Device Shape and/or Size; Patient Device Interaction Problem | Unspecified Heart Problem | P130021 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING THE IMPACT OF PROSTHESIS-PATIENT MISMATCH (PPM) AFTER TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) IN ASIAN PATIENTS. THE STUDY POPULATION CONSISTED OF 7,072 PATIENTS WHO UNDERWENT TAVR WITH EITHER A MEDTRONIC COREVALVE/EVOLUT R/EVOLUT PRO VALVE (N = 1,647) OR A NON-MEDTRONIC SAPIEN XT/SAPIEN 3 VALVE (N = 5 ,425). THE AUTHORS CLASSIFIED PPM AS MODERATE OR SEVERE BASED ON THE INDEXED EFFECTIVE ORIFICE AREA OF THE IMPLANTED VALVE. AMONG ALL PATIENTS, THE FOLLOWING ADVERSE OUTCOMES OCCURRED: PPM (MODERATE OR SEVERE), BLEEDING (LIFE-THREATENING OR MAJOR), MAJOR VASCULAR COMPLICATIONS, DISABLING STROKE, CORONARY ARTERY OBSTRUCTION, NEW ATRIAL FIBRILLATION, NEED FOR PACEMAKER IMPLANTATION, PARAVALVULAR LEAK (MILD TO SEVERE), ELEVATED MEAN AORTIC GRADIENTS (= 20 MMHG OR = 20 MMHG), AND ACUTE KIDNEY INJURY. ADDITIONALLY, THE AUTHORS FOUND THAT ¿SEVERE PPM, BUT NOT MODERATE PPM, WAS INDEPENDENTLY ASSOCIATED WITH AN INCREASED RISK OF 3-YEAR ALL-CAUSE MORTALITY, CARDIOVASCULAR MORTALITY, AND HEART FAILURE HOSPITALIZATION.¿ Manufacturer Narrative: CITATION: ISHIZU K, KITANO K, YAMAMOTO K, ET AL. UPDATE ON PROSTHESIS-PATIENT MISMATCH FOLLOWING TRANSCATHETER AORTIC VALVE REPLACEMENT IN ASIAN PATIENTS. JACC ASIA. 2024;4(11):793-806. PUBLISHED 2024 OCT 8. DOI:10.1016/J.JACASI.2024.08.010. EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT AND DATE OF DEATH. EARLIEST APPROVED MEDTRONIC TAVR PRODUCT USED FOR PRODUCT CODE AND PMA#. NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21687266&pc=PNJ | 9612164-2025-01522 | 2025-01-03 05:00 | Death | MEDTRONIC IRELAND | 2025-03-25 04:00 | PNJ | MICRA¿ AV2 | Appropriate Term/Code Not Available | Bacterial Infection | P150033 |  | 1 | Event Description: IT WAS REPORTED BY A PATIENT REPRESENTATIVE THAT THE PATIENT DIED. "THE OLD DEVICE HAD BACTERIA ON IT THAT SPREAD THROUGH HIS BODY BUT WHEN THEY PUT THE NEW DEVICE IN IT WAS TOO LATE". Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21687282&pc=DTB | 2649622-2025-08113 | 2025-01-03 05:00 | Death | MPRI | 2025-03-25 04:00 | DTB | CAPSURE SENSE LEAD MRI SURESCAN | Appropriate Term/Code Not Available | Bacterial Infection | P830061 |  | 1 | Event Description: IT WAS REPORTED BY A PATIENT REPRESENTATIVE THAT THE PATIENT DIED. "THE OLD DEVICE HAD BACTERIA ON IT THAT SPREAD THROUGH HIS BODY BUT WHEN THEY PUT THE NEW DEVICE IN IT WAS TOO LATE". Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21687375&pc=FJS | 1282497-2025-00189 | 2025-02-04 05:00 | Death | COVIDIEN LLC | 2025-03-25 04:00 | FJS | PERITONEAL CATHETER | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE PROSPECTIVE STUDY, A SINGLE-CENTER RETROSPECTIVE ANALYSIS OF 246 PEDIATRIC PATIENTS WAS INCLUDED AFTER THE IMPLANTATION OF A CURVED-END TENCKHOF CATHETER (TC) FOR PERITONEAL DIALYSIS (PD) BETWEEN 2002 AND 2022. A TENCKHOF CATHETER WITH A CURVED END WAS INSERTED LAPAROSCOPICALLY OR VIA A MINI LAPAROTOMY. THE AGE GROUP BREAKDOWN WAS 53 NEWBORNS (0¿28 DAYS), 32 INFANTS (29 DAYS TO 12 MONTHS), 80 TODDLERS (1¿3 YEARS), 44 SCHOOL-AGED CHILDREN (4¿10 YEARS), AND 37 ADOLESCENTS (10¿19 YEARS). THE IMPLANTATION SITES VARIED: LEFT LOWER ABDOMINAL QUADRANT (144, 48%), RIGHT LOWER QUADRANT (98, 32%), MIDDLE ABDOMINAL LINE (5, 1.8%), AND 52 (17%) WERE UNRECORDED. FIFTY-SIX DEATHS OCCURRED WITHIN 61 MONTHS. OVERALL, 18.7% OF PATIENTS EXPERIENCED MORTALITY ACROSS ALL GROUPS. THE MEDIAN CATHETER SURVIVAL TIME WAS 0.6 YEARS. MORTALITY RATES FOR CHILDREN REQUIRING RENAL REPLACEMENT THERAPY WERE 30 TIMES GREATER THAN THOSE OF HEALTHY PEERS.Event Description: ACCORDING TO THE LITERATURE, A SINGLE-CENTER RETROSPECTIVE ANALYSIS OF 246 PEDIATRIC PATIENTS WAS INCLUDED AFTER THE IMPLANTATION OF A CURVED-END TC (TENCKHOFF CATHETER) FOR PD (PERITONEAL DIALYSIS), BETWEEN 2002 AND 2022. A TENCKHOFF CATHETER WITH A CURVED END (COVIDIEN; MANSFELD, MA, USA) WAS INSERTED LAPAROSCOPICALLY OR VIA A MINI LAPAROTOMY. THE AGE GROUP BREAKDOWN WAS 53 NEWBORNS (0¿28 DAYS), 32 INFANTS (29 DAYS TO 12 MONTHS), 80 TODDLERS (1¿3 YEARS), 44 SCHOOL-AGED CHILDREN (4¿10 YEARS), AND 37 ADOLESCENTS (10¿19 YEARS). THE IMPLANTATION SITES VARIED: LEFT LOWER ABDOMINAL QUADRANT (144, 48%), RIGHT LOWER QUADRANT (98, 32%), MIDDLE ABDOMINAL LINE (5, 1.8%), AND 52 (17%) WERE UNRECORDED. FIFTY-SIX DEATHS OCCURRED WITHIN 61 MONTHS. OVERALL, 18.7% OF PATIENTS EXPERIENCED MORTALITY ACROSS ALL GROUPS. THE MEDIAN CATHETER SURVIVAL TIME WAS 0.6 YEARS. MORTALITY RATES FOR CHILDREN REQUIRING RENAL REPLACEMENT THERAPY REMAIN 30 TIMES GREATER THAN THOSE OF HEALTHY PEERS. Manufacturer Narrative: D10 - CONCOMITANT PRODUCT: UNK PD, PERITONEAL CATHETER, (LOT: UNKNOWN); UNK PD, PERITONEAL CATHETER, (LOT: UNKNOWN) AFTZOGLOU, MICHAIL. ¿SURGICAL RISK STRATIFICATION AND OUTCOME ANALYSIS OF TENCKHOF CATHETER IMPLANTATIONS IN PAEDIATRIC PATIENTS: A SINGLE-CENTRE EXPERIENCE.¿ EUROPEAN JOURNAL OF PEDIATRICS, VOL. 184, NO. 172, 2025, HTTPS://DOI.ORG/10.1007/S00431-025-06006-X. PAGES: 1-11. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTION: B5, E4(EMAIL) ADDITIONAL INFORMATION: G3. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21687376&pc=NVN | 2649622-2025-08117 | 2025-01-03 05:00 | Death | MPRI | 2025-03-25 04:00 | NVN | SELECTSECURE MRI SURESCAN | Appropriate Term/Code Not Available | Bacterial Infection | P030036 |  | 1 | Event Description: IT WAS REPORTED BY A PATIENT REPRESENTATIVE THAT THE PATIENT DIED. "THE OLD DEVICE HAD BACTERIA ON IT THAT SPREAD THROUGH THEIR BODY BUT WHEN THEY PUT THE NEW DEVICE IN IT WAS TOO LATE". Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21687425&pc=DSQ | 2916596-2025-01665 | 2025-03-04 05:00 | Death | THORATEC CORPORATION | 2025-03-25 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Ischemia Stroke | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY AFTER AN ISCHEMIC STROKE. THE PUMP WAS WORKING AS INTENDED DURING THE TIME OF SUPPORT. IT WAS NOTED THIS WAS NOT DEVICE OR THERAPY RELATED. IT WAS NOTED THAT THE DEVICE OPERATED AS INTENDED. THE DEVICE WOULD NOT BE RETURNED FOR EVALUATION. THERE WERE NO REPORTED CHANGES TO ANTICOAGULATION STATUS. THERE WERE NO RECENT CHANGES TO PUMP PARAMETERS. IT WAS NOT REPORTED WHAT THE PATIENT'S SYMPTOMS WERE. A THROMBECTOMY WAS PERFORMED AND CEREBRAL BLEEDING OCCURRED LATER ON. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF STROKE COULD NOT BE CONCLUSIVELY DETERMINED. STROKE IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. STROKE HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21687566&pc=DSQ | 2916596-2025-01932 | 2025-03-13 04:00 | Death | THORATEC CORPORATION | 2025-03-25 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTI ORGAN FAILURE. THE DEATH WAS RELATED TO THE PATIENT'S POOR HEALTH CONDITION AND WAS NOT DEVICE OR THERAPY RELATED. THE DEVICE WAS WORKING AS INTENDED DURING TIME OF SUPPORT. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RESPIRATORY FAILURE, RIGHT HEART FAILURE, RENAL FAILURE, HEPATIC DYSFUNCTION) AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21687671&pc=NGV | 2135147-2025-01539 | 2021-07-01 04:00 | Death | ABBOTT MEDICAL | 2025-03-25 04:00 | NGV | AMPLATZER AMULET | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P200049 |  | 1 | Event Description: THE ARTICLE, "CLINICAL VALUE OF CT-BASED 3D COMPUTATIONAL MODELING IN LEFT ATRIAL APPENDAGE OCCLUSION: AN IN-DEPTH ANALYSIS OF THE PRECISE LAAO STUDY', WAS REVIEWED. THE ARTICLE PRESENTED A PROSPECTIVE, MULTICENTER STUDY TO COMPARE THE OPERATOR'S DEGREE OF CERTAINTY IN DEVICE SIZE SELECTION AFTER STANDARD OF CARE (SOC) ASSESSMENT (CERTAINTY SOC) WITH THE DEGREE OF CERTAINTY AFTER FEOPS CONSULTATION (CERTAINTY FEOPS) AND TO INVESTIGATE THE SUBJECTIVE VALUE OF FEOPS ACCORDING TO THE IMPLANTING TEAMS. DEVICES INCLUDED IN THE STUDY WERE SOLELY AMPLATZER AMULET. THE ARTICLE CONCLUDED THAT COMPUTED TOMOGRAPHY (CT)-BASED COMPUTATIONAL MODELING USING FEOPS HAS A POSITIVE DECISIONAL IMPACT IN LAAO, INDUCING A CHANGE IN AMULET SIZE SELECTION IN ALMOST ONE OUT OF THREE PATIENTS AND INCREASING PROCEDURAL CONFIDENCE. [THE PRIMARY AND CORRESPONDING AUTHOR WAS IAN BUYSSCHAERT, DEPARTMENT OF CARDIOLOGY, AZ SINT-JAN BRUGGE AV, BRUGGE, BELGIUM, WITH CORRESPONDING EMAIL: IAN.BUYSSCHAERT@AZSINTJAN.BE]. THE TIME FRAME OF THE STUDY WAS FROM JULY 2021 TO APRIL 2023. A TOTAL OF 102 PATIENTS WERE INCLUDED IN THE STUDY, OF WHICH ALL RECEIVED AN ABBOTT AMPLATZER AMULET DEVICE. THE AVERAGE AGE WAS 76.2 YEARS AND THE MAJORITY GENDER WAS MALE. COMORBIDITIES INCLUDED CONGESTIVE HEART FAILURE, HYPERTENSION, DIABETES MELLITUS, PRIOR STROKE, PRIOR TRANSIENT ISCHEMIC ATTACK, VASCULAR DISEASE, PERMANENT ATRIAL FIBRILLATION, CEREBRAL BLEEDING, RECURRENT BLEEDING. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF AMPLATZER AMULET WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING CONGESTIVE HEART FAILURE, HYPERTENSION, DIABETES MELLITUS, PRIOR STROKE, PRIOR TRANSIENT ISCHEMIC ATTACK, VASCULAR DISEASE, PERMANENT ATRIAL FIBRILLATION, CEREBRAL BLEEDING, RECURRENT BLEEDING. COMPLICATIONS REPORTED INCLUDED DEATH, BLEEDING, GASTRO-INTESTINAL BLEEDING, ANEMIA, MYOCARDIAL INFARCTION, RESIDUAL SHUNT. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. B2: DEATH DATE IS ESTIMATED. B3: DATE OF EVENT IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE: CLINICAL VALUE OF CT-BASED 3D COMPUTATIONAL MODELING IN LEFT ATRIAL APPENDAGE OCCLUSION: AN IN-DEPTH ANALYSIS OF THE PRECISE LAAO STUDY. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. B2: DEATH DATE IS ESTIMATED. B3: DATE OF EVENT IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: CLINICAL VALUE OF CT-BASED 3D COMPUTATIONAL MODELING IN LEFT ATRIAL. APPENDAGE OCCLUSION: AN IN-DEPTH ANALYSIS OF THE PRECISE LAAO STUDY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21687757&pc=GDW | 3005075853-2025-02131 | 2025-02-25 05:00 | Death | ETHICON ENDO-SURGERY, LLC. | 2025-03-25 04:00 | GDW | ECHELON 3000 60MM STANDARD | Failure to Fire; Difficult to Open or Close; Adverse Event Without Identified Device or Use Problem; Activation Problem | Air Embolism; Cardiac Arrest | K213633 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A GUN SHOT TRAUMA PROCEDURE EARLY THIS MORNING THAT AN UNKNOWN ISSUE OCCURRED WITH THE DEVICE AND THE PATIENT PASSED OF AN EMBOLUS. NO FURTHER INFORMATION AVAILABLE AT THIS TIME. Manufacturer Narrative: (B)(4). DATE SENT: 4/17/2025. D4: BATCH # C9DV27. INVESTIGATION SUMMARY: THE PRODUCT WAS RETURNED TO ETHICON ENDO-SURGERY FOR EVALUATION. VISUAL INSPECTION AND FUNCTIONAL TESTING WERE CONDUCTED ON THE RETURNED DEVICE. VISUAL ANALYSIS OF THE RETURNED SAMPLE DETERMINED THAT ONE ECH60S DEVICE WAS RETURNED WITH ONE GST60B RELOAD LOADED ON THE DEVICE. THE RELOAD WAS RECEIVED PARTIALLY FIRED 1/3 AND WITH DAMAGE ON RELOAD DECK. THE DAMAGE TO THE RELOAD IS CONSISTENT WITH THE DEVICE BEING CLAMPED OVER A HARD OBJECT. UPON VISUAL INSPECTION, IT WAS NOTED THAT THE JOINT COVER WAS DAMAGED. IT IS POSSIBLE THAT THE DEVICE WAS ATTENDED TO BE PULLED OUT FROM A TROCAR IN THE ARTICULATED POSITION, RESULTING IN THE EDGE OF THE TROCAR DAMAGING THE JOINT COVER. IN ADDITION, THE ANVIL BENT WAS NOTED UPWARDS. IT IS POSSIBLE THAT THE DEVICE WAS CLAMPED OVER AN EXCESS OF TISSUE CAUSING THE ANVIL TO BEND AND FOR THE FIRING STROKE AND THE STAPLE FORM TO BE INCOMPLETE. THE MANUAL OVERRIDE DOOR OUT OF POSITION; THE OVERRIDE LEVER WAS UP WHICH DENOTES THAT THE KNIFE WAS MANUALLY RETURNED TO HOME POSITION. IT SHOULD BE NOTED AFTER THE MANUAL OVERRIDE SYSTEM IS USED THE INSTRUMENT IS DISABLED AND CANNOT BE ARTICULATED OR USED FOR ANY SUBSEQUENT FIRINGS. THE BAILOUT SYSTEM WAS RESET AND THEN, IT WAS TESTED FOR FUNCTIONALITY IN THE STRAIGHT POSITION WITH A TEST RELOAD AND ACHIEVED ITS COMPLETE FIRING SEQUENCE WITHOUT ANY DIFFICULTIES. THE STAPLE LINE AND CUT LINE WERE COMPLETE AND THE STAPLES MEET THE STAPLE RELEASE CRITERIA. HOWEVER, THE CUT WAS NOTED TO BE JAGGED DUE TO A DAMAGED KNIFE. ONE POSSIBLE CAUSE FOR THIS TYPE OF DAMAGE TO THE KNIFE IS WHEN THE DEVICE IS FIRED OVER AN ALREADY EXISTING STAPLE LINE, HARD OBJECT, OR THICKER TISSUE THAN INDICATED. REPEATEDLY FIRING ACROSS EXISTING STAPLE LINES CAN ALSO REDUCE THE ABILITY TO CUT CLEANLY. TO MITIGATE THE POTENTIAL FOR STAPLES GETTING INTO THE RELOAD AND INTERFERING WITH THE KNIFE PATH DURING DEVICE FIRING, PRIOR TO RELOADING THE DEVICE, RINSE THE ANVIL AND RELOAD JAW IN STERILE SOLUTION AND THEN WIPE THE ANVIL AND RELOAD JAW TO CLEAN ANY FORMED BUT UNUSED STAPLES FROM THE DEVICE. ADDITIONALLY, PROPER CARE SHOULD BE TAKEN WHEN PLACING THE DEVICE ON THE TISSUE TO BE STAPLED, TO ENSURE THAT NO HARD OBSTRUCTION SUCH AS A CLIP IS INCLUDED WITH THE TISSUE INSIDE THE JAWS. PLEASE REFERENCE THE INSTRUCTION FOR USE FOR MORE INFORMATION. THE EVENT DESCRIBED OF WOULD NOT OPEN AND MANUAL OVERRIDE WOULD NOT WORK COULD NOT BE CONFIRMED AS THE DEVICE OPENED AND CLOSED WITHOUT ANY DIFFICULTIES NOTED AND THE MANUAL OVERRIDE WAS TOTALLY FUNCTIONAL. THE DEVICE EVENT LOG WAS REVIEWED. THE LOG INDICATES THAT THE DEVICE HAD A HIGH FORCE TO FIRE FOR ONE OF THE CLINICAL FIRINGS, INDICATING THAT THE DEVICE WAS FIRING OVER AN OBSTRUCTION, OR TOO THICK OF TISSUE. IN ADDITION, THE LOG INDICATES THAT ARTICULATION WAS ATTEMPTED WHILE THE DEVICE WAS CLOSED. ARTICULATION FUNCTION IS BLOCKED WHEN THE DEVICE IS CLOSED TO PREVENT INADVERTENT PATIENT INJURY. WHEN THE HOME BUTTON IS PRESSED WHILE THE JAWS ARE CLOSED THE DEVICE WILL ATTEMPT TO RETRACT THE KNIFE. WHEN ANY OF THE OTHER ARTICULATION BUTTONS ARE PRESSED WHILE THE JAWS ARE CLOSED THE DEVICE WILL PROVIDE HAPTIC FEEDBACK INDICATING THAT ARTICULATION FUNCTION IS BLOCKED. DEVICE HISTORY REVIEW: A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH NUMBER C9DV27, AND NO NON-CONFORMANCES RELATED TO THE REPORTED COMPLAINT CONDITION WERE IDENTIFIED. Manufacturer Narrative: (B)(4). DATE SENT: 3/25/2024. D4: UNK. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: THE EVENT DEVICE IS THE ONE WITH THE MANUAL OVERRIDE ENGAGED. PATIENT DEATH (B)(6) 2025. PATIENT ARRIVED AT FACILITY LATE (B)(6) 2025, AND PASSED (B)(6) 2025 EARLY MORNING. WHAT WAS THE UNKNOWN ISSUE WITH THE ETHICON DEVICE? THE DEVICE STOPPED FIRING IN THE MIDDLE OF THE FIRE AND THE MANUAL OVERRIDE DID NOT RELEASE THE TISSUE. WHAT WAS THE ANATOMICAL STRUCTURE THE DEVICE WAS BEING FIRED UPON? THE SUPERIOR PULMONARY ARTERY ON THE LEFT. WAS THE SECOND DEVICE USED FOR THIS PROCEDURE? IF SO WAS THERE ANY ISSUE WITH THE SECOND DEVICE? THE SECOND DEVICE WAS NOT USED. PLEASE PROVIDE A TIMELINE OF EVENTS. THE PATIENT WAS IN THE OR. MULTIPLE FIRES OF THE STAPLER WERE USED TO RESECT THE LEFT UPPER LOBE OF THE LUNG. THE HILUM WAS CONTROLLED WITH A CLAMP AND THE STAPLER WAS USED WITH PERI STRIPS ACROSS THE VESSEL. DURING THE MIDPOINT OF THE FIRE, THE STAPLER LOCKED UP AND THE MANUAL OVERRIDE WAS NOT EFFECTIVE AT RELEASING THE TISSUE. ONCE ANOTHER CLAMP WAS PLACED PROXIMAL, WE HAD TO CUT THE STAPLER OUT OF THE PATIENT. THE CLAMP THEN SLIPPED OFF DURING THIS PROCESS AND THE PATIENT SUFFERED A CARDIAC ARREST AND WAS UNABLE TO RECOVER. WHAT WAS THE COD? UNKNOWN BUT LIKELY ACUTE COR PULMONALE AND AIR EMBOLISM. IS THERE AN AUTOPSY REPORT AVAILABLE? UNKNOWN. PLEASE PROVIDE MORE INFORMATION REGARDING THE EMBOLISM. WAS THE EMBOLISM AIR OR BLOOD? AIR. DOES THE SURGEON BELIEVE THAT THE DIFFICULTIES OF THE DEVICE RESULTED IN THE EMBOLUS THAT RESULTED IN THE PATIENT DEMISE? YES. NOT THE CAUSE OF DEATH BUT LIKELY CONTRIBUTED. IS THE SURGEON INTERESTED IN SPEAKING WITH ETHICON MEDICAL AND ENGINEERING STAFF? NO. THE DEVICE/PHOTO UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21687785&pc=PID | 3005099803-2025-01243 | 2006-11-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-03-25 04:00 | PID | RIGIFLEX II | Adverse Event Without Identified Device or Use Problem | Perforation; Sepsis |  |  | 1 | Event Description: BOSTON SCIENTIFIC BECAME AWARE OF AN EVENT INVOLVING RIGIFLEX II BALLOON THROUGH THE ARTICLE TITLED, "SAFETY OF PNEUMATIC DILATION IN OLDER ADULTS WITH ACHALASIA: AN INTERNATIONAL MULTICENTER CROSS-SECTIONAL STUDY", BY NIR BAR ET. AL. PER THE ARTICLE, BETWEEN NOVEMBER 2006 TO OCTOBER 2020 THE PATIENT EXPERIENCED PERFORATION WHICH RESULTED IN DEATH FROM SEPSIS 3 MONTHS AFTER THE PERFORATION WAS SEALED USING AN OVER THE SCOPE CLIP AND A LONG HOSPITALIZATION. PLEASE SEE THE REFERENCED ARTICLE FOR FULL DETAILS AND FULL LISTING OF PHYSICIANS AND HEALTHCARE FACILITIES. Manufacturer Narrative: BLOCK B3: THE EXACT DATE OF EVENT WAS NOT REPORTED. THE RETROSPECTIVE STUDY DATE NOVEMBER 1, 2006, IS USED FOR THE ESTIMATED DATE OF EVENT BETWEEN NOVEMBER 2006 AND OCTOBER 2020. BLOCK D4, H4: THE SUSPECT DEVICE LOT NUMBER AND UPN WERE NOT REPORTED. THEREFORE, THE UNIQUE IDENTIFIER (UDI) NUMBER, MANUFACTURE AND EXPIRATION DATES ARE UNKNOWN. BLOCK G2: LITERATURE SOURCE: NIR BAR, ET AL., "SAFETY OF PNEUMATIC DILATION IN OLDER ADULTS WITH ACHALASIA: AN INTERNATIONAL MULTICENTER CROSS-SECTIONAL STUDY" GASTROENTEROLOGY AND HEPATOLOGY DEPARTMENT, TEL AVIV MEDICAL CENTER, TEL AVIV FACULTY OF MEDICINE, TEL AVIV UNIVERSITY, TEL AVIV 6423906, HTTPS://DOI.ORG/10.3390/JCM12206682. BLOCK H6: IMDRF PATIENT CODE E2114 CAPTURES THE REPORTABLE EVENT OF AN ESOPHAGEAL PERFORATION. IMDRF PATIENT CODE E0306 CAPTURES THE REPORTABLE EVENT OF SEPSIS. IMDRF IMPACT CODE F02 CAPTURES THE REPORTABLE EVENT OF DEATH. IMDRF IMPACT CODE F2301 CAPTURES THE REPORTABLE EVENT OF AN ADDITIONAL DEVICE REQUIRED TO SEAL THE PERFORATION. IMDRF IMPACT CODE F08 CAPTURES THE REPORTABLE EVENT OF HOSPITALIZATION DUE TO PERFORATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21688067&pc=QTC | 1417592-2025-00141 | 2023-02-01 05:00 | Death | MEDLINE INDUSTRIES LP | 2025-03-25 04:00 | QTC | DBQ-BAR,BED ASSIST | Insufficient Information | Laceration(s) |  |  | 1 | Event Description: "ENTANGLED WITH THE PRODUCT. THIS LED TO SEVERE LACERATIONS AND BLOOD LOSS." Manufacturer Narrative: IT WAS REPORTED THAT THE USER "BECAME ENTANGLED WITH THE PRODUCT". "THIS LED TO SEVERE LACERATIONS AND BLOOD LOSS. SHE WAS RUSHED TO THE HOSPITAL AND LATER PASSED AWAY". IT HAS BEEN DETERMINED THAT THE REPORTED EVENT CAUSED OR CONTRIBUTED TO SERIOUS INJURY, THEREFORE, THIS MEDWATCH IS BEING FILED. IF ANY FURTHER RELEVANT INFORMATION IS IDENTIFIED OR OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21688542&pc=OUT | 2029214-2025-00766 |  | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-03-25 04:00 | OUT | PIPELINE | Material Deformation; Physical Resistance/Sticking | Obstruction/Occlusion |  |  | 1 | Event Description: MEDTRONIC RECEIVED A REPORT THAT A PIPELINE BECAME DENSE DUE TO THE PUSH DURING STENT DEPLOYMENT. AFTER SURGERY, THERE WAS AN OCCLUSION AT THE STENT SITE. IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE REPORTED DEVICE AND ACCESSORY DEVICES WERE PREPARED AS INDICATED IN THE INSTRUCTION FOR USE (IFU). THE PATIENT WAS BEING TREATED FOR AN UNRUPTURED ANEURYSM WITH A SACCULAR MORPHOLOGY.Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED VESSEL TORTUOSITY WAS MODERATE. THE PATIENT PASSES AWAY IN THE HOSPITAL. THE PATIENT DEATH RELATION TO MEDTRONIC DEVICE OR THERAPY IS UNDENIABLE. LEADING UP TO THE PATIENT DEATH THERE WAS UNRUPTURED CEREBRAL ANEURYSM. THERE WAS NO RESISTANCE IN THE CATHETER. CONTINUOUS SALINE FLUSH ADMINISTERED AND MAINTAINED AS INDICATED IN THE INSTRUCTIONS FOR USE (IFU). THERE WAS NO DAMAGE OBSERVED TO THE PIPELINE PUSHWIRE OR CATHETER. THE PATIENT UNDERWENT THROMBECTOMY THERAPY TO TO RESOLVE THE OCCLUSION. DUAL ANTIPLATELET TREATMENT WAS ADMINISTERED. ANCILLARY DEVICES: FG15150-0615-1S CATHETER. Manufacturer Narrative: B3: EVENT DATE IS NOT KNOWN. PLEASE SEE B5 FOR APPROXIMATE DATE RANGE, IF APPLICABLE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21688605&pc=NPT | 2025587-2025-02097 | 2024-02-20 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-25 04:00 | NPT | EVOLUT FX VALVE | Adverse Event Without Identified Device or Use Problem; Insufficient Information | Insufficient Information; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING THE IMPLANT OF A TRANSCATHETER VALVE, THE PATIENT DIED. THE CAUSE OF DEATH WAS UNKNOWN. PER THE PHYSICIAN, THE DEATH WAS NOT RELATED TO THE VALVE OR VALVE IMPLANT PROCEDURE.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT FOLLOWING THE IMPLANT OF THE TRANSCATHETER VALVE AND "FRACTURE TREATMENT", THE PATIENT WAS SUBSEQUENTLY HOSPITALIZED. THE PATIENT'S ABILITY TO COMPLETE ACTIVITIES OF DAILY LIVING DECREASED, AND THE PATIENT SUBSEQUENTLY DIED. PER THE PHYSICIAN, THE CAUSE OF DEATH WAS UNKNOWN BUT IT WAS LIKELY THAT THE PATIENT DIED OF OLD AGE.Event Description: IT WAS REPORTED THAT FOLLOWING THE IMPLANT OF A TRANSCATHETER VALVE, THE PATIENT DIED. THE CAUSE OF DEATH WAS UNKNOWN. PER THE PHYSICIAN, THE DEATH WAS NOT RELATED TO THE VALVE OR VALVE IMPLANT PROCEDURE. ADDITIONAL INFORMATION WAS RECEIVED THAT FOLLOWING THE IMPLANT OF THE TRANSCATHETER VALVE AND "FRACTURE TREATMENT", THE PATIENT WAS SUBSEQUENTLY HOSPITALIZED. THE PATIENT'S ABILITY TO COMPLETE ACTIVITIES OF DAILY LIVING DECREASED, AND THE PATIENT SUBSEQUENTLY DIED. PER THE PHYSICIAN, THE CAUSE OF DEATH WAS UNKNOWN BUT IT WAS LIKELY THAT THE PATIENT DIED OF OLD AGE. ADDITIONAL INFORMATION WAS RECEIVED TO CLARIFY THE PATIENT WAS HOSPITALIZED DUE TO A DECLINE IN ACTIVITIES OF DAILY LIVING ASSOCIATED WITH ADVANCED AGE. INFORMATION WAS INITIALLY REPORTED, BUT NOT CAPTURED IN THE ABOVE EVENT NARRATIVE: THE PATIENT DIED APPROXIMATELY SIX MONTHS FOLLOWING THE VALVE IMPLANT PROCEDURE. Manufacturer Narrative: UPDATED DATA: B5. A THIRD PARAGRAPH WAS ADDED. H6. CODES WERE UPDATED. CORRECTED DATA; E. FACILITY FULL STREET ADDRESS: THE INFORMATION WAS AVAILABLE AT THE TIME OF INITIAL REPORT SUBMISSION, BUT WAS ERRONEOUSLY OMITTED FROM THE INITIAL AND SUPPLEMENTAL 001 REPORTS. D. SUSPECT MEDICAL DEVICE FULL UDI# THE INFORMATION WAS AVAILABLE AT THE TIME OF SUPPLEMENTAL REPORT 001 SUBMISSION, BUT WAS ERRONEOUSLY OMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. D4. H4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21688693&pc=NPT | 2135147-2025-01541 | 2025-02-28 05:00 | Death | ABBOTT MEDICAL | 2025-03-25 04:00 | NPT | NAVITOR VISION | Migration or Expulsion of Device | Low Blood Pressure/ Hypotension | P190023 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 25MM NAVITOR VALVE WAS SELECTED FOR IMPLANT USING A SMALL FLEXNAV DELIVERY SYSTEM (DS). THE ANATOMICAL MEASUREMENTS OF THE PATIENT WERE: MINIMUM DIAMETER OF 20.3MM, MAX DIAMETER OF 25.2MM, AVERAGE DIAMETER OF 22.8MM, DERIVED AREA OF 22.8MM, DERIVED PERIMETER OF 23.1MM, AREA OF 408.9MM^2, AND PERIMETER OF 72.6MM. A PRE-DILATION WAS PERFORMED WITH A 20MM WIDE X 45MM LONG NON-ABBOTT BALLOON. THE DEVICE WAS SUCCESSFULLY IMPLANTED AT A DEPTH OF NCC ~1-2MM AND LCC~4-5MM. POST-DEPLOYMENT OF THE VALVE, THE NOSE CONE OF THE DELIVERY SYSTEM GOT STUCK ON WHAT APPEARED TO BE THE BOTTOM STRUTS OF THE VALVE ON THE NCC SIDE. MULTIPLE ATTEMPTS WERE MADE TO PUSH THE DELIVERY SYSTEM BACK INTO LEFT VENTRICLE. THE PHYSICIAN ALSO ATTEMPTED TO SNARE THE NOSECONE TO GET IT CENTRALIZED. ADDITIONAL MANIPULATION AND REMOVING TENSION WAS DONE BUT IT WAS DIFFICULT FOR THE PHYSICIAN TO MOVE THE NOSE CONE PAST THE VALVE. THE VALVE POPPED UP TOWARD THE AORTA AND A SECOND NON-ABBOTT VALVE WAS IMPLANTED VALVE-IN-VALVE. AFTER DEPLOYMENT OF THE SECOND VALVE, THE PATIENTS BLOOD PRESSURE CRASHED. CPR WAS PERFORMED BUT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS NOTED TO BE SURE TO ISSUES WITH THE LEFT VENTRICLE.Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 25MM NAVITOR VALVE WAS SELECTED FOR IMPLANT USING A SMALL FLEXNAV DELIVERY SYSTEM (DS). THE ANATOMICAL MEASUREMENTS OF THE PATIENT WERE: MINIMUM DIAMETER OF 20.3MM, MAX DIAMETER OF 25.2MM, AVERAGE DIAMETER OF 22.8MM, DERIVED AREA OF 22.8MM, DERIVED PERIMETER OF 23.1MM, AREA OF 408.9MM^2, AND PERIMETER OF 72.6MM. A PRE-DILATION WAS PERFORMED WITH A 20MM WIDE X 45MM LONG NON-ABBOTT BALLOON. THE DEVICE WAS SUCCESSFULLY IMPLANTED AT A DEPTH OF NCC ~1-2MM AND LCC~4-5MM. POST-DEPLOYMENT OF THE VALVE, THE NOSE CONE OF THE DELIVERY SYSTEM GOT STUCK ON WHAT APPEARED TO BE THE BOTTOM STRUTS OF THE VALVE ON THE NCC SIDE. MULTIPLE ATTEMPTS WERE MADE TO PUSH THE DELIVERY SYSTEM BACK INTO LEFT VENTRICLE. THE PHYSICIAN ALSO ATTEMPTED TO SNARE THE NOSECONE TO GET IT CENTRALIZED. ADDITIONAL MANIPULATION AND REMOVING TENSION WAS DONE BUT IT WAS DIFFICULT FOR THE PHYSICIAN TO MOVE THE NOSE CONE PAST THE VALVE. THE VALVE POPPED UP TOWARD THE AORTA AND A SECOND NON-ABBOTT VALVE WAS IMPLANTED VALVE-IN-VALVE. AFTER DEPLOYMENT OF THE SECOND VALVE, THE PATIENTS BLOOD PRESSURE CRASHED. CPR WAS PERFORMED BUT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS NOTED TO BE SURE TO ISSUES WITH THE LEFT VENTRICLE. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. AN EVENT OF MIGRATION, HYPOTENSION, AND DEATH WAS REPORTED. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED, AND THE PRODUCT MET ALL SPECIFICATIONS. THE EVENT AND PROVIDED IMAGING WAS REVIEWED BY AN ABBOTT DIRECTOR OF MEDICAL AFFAIRS. BASED ON ALL AVAILABLE INFORMATION, THE REPORTED MIGRATION APPEARS TO BE RELATED TO PROCEDURAL CONDITIONS (INTERACTION BETWEEN VALVE AND NOSE CONE DURING DELIVERY SYSTEM REMOVAL). THE REPORTED HYPOTENSION AND DEATH ALSO APPEAR TO BE RELATED TO A COMBINATION OF PROCEDURAL CONDITIONS AND PATIENT CONDITION (CORONARY OBSTRUCTION, SEVERE LEFT VENTRICLE IMPAIRMENT). THE REPORTED UNEXPECTED MEDICAL INTERVENTION AND SURGERY WERE THE RESULTS OF CASE-SPECIFIC CIRCUMSTANCES. CODES REMOVED: TYPE OF INVESTIGATION: 4118 TYPE OF INVESTIGATION NOT YET DETERMINED; INVESTIGATION FINDINGS: INVESTIGATION CONCLUSIONS: 11 CONCLUSION NOT YET AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21688986&pc=CAF | 2518422-2025-038523 | 2023-12-21 05:00 | Death | RESPIRONICS, INC. | 2025-03-25 04:00 | CAF | I-NEB CF UK UNIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information; No Clinical Signs, Symptoms or Conditions | K102454 |  | 1 | Event Description: THE MANUFACTURER BECAME AWARE OF AN ALLEGATION OF I-NEB CF UK UNIT ALLEGES DEATH - PATIENT NO LONGER USING INEB DEVICE. THERE IS NO ALLEGATION OF SERIOUS OR PERMANENT HARM OR INJURY. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER¿S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH AN I-NEB CF UK UNIT DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE DEVICE WAS RETURNED AS NO LONGER NEEDED AND WAS SCRAPPED WITHOUT TESTING. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED. IN THIS REPORT, BOX D, G, H HAS BEEN CORRECTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21689392&pc=FKX | 0008030665-2025-00697 | 2025-03-16 04:00 | Death | ERIKA DE REYNOSA, S.A. DE C.V. | 2025-03-25 04:00 | FKX | LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | Adverse Event Without Identified Device or Use Problem | Chronic Obstructive Pulmonary Disease (COPD) | K222318 |  | 1 | Event Description: IT WAS REPORTED THAT A PERITONEAL DIALYSIS (PD) PATIENT PASSED AWAY. A PD REGISTERED NURSE (PDRN) CONTACTED FRESENIUS TO DEACTIVATE THE PATIENT'S ACCOUNT IN ORDER FOR THE PATIENT'S FAMILY TO NOT RECEIVE ADDITIONAL SUPPLIES. ADDITIONAL INFORMATION WAS OBTAINED THROUGH FOLLOW-UP WITH THE PATIENT¿S PDRN. THE PATIENT WAS FOUND DECEASED ON (B)(6) 2025. THE PATIENT WAS STILL CONNECTED TO THE LIBERTY SELECT CYCLER WHEN FOUND. HOWEVER, THE LAST COMPLETED TREATMENT WAS RECORDED AS (B)(6) 2025. THE PATIENT¿S RECORDS DOCUMENTED THAT A FULL TREATMENT WAS COMPLETED WITHOUT ANY ISSUES. THE CAUSE OF DEATH WAS DOCUMENTED AS CHRONIC OBSTRUCTIVE PULMONARY DISEASE (COPD). NO ADDITIONAL INFORMATION WAS PROVIDED.Event Description: IT WAS REPORTED THAT A PERITONEAL DIALYSIS (PD) PATIENT PASSED AWAY. A PD REGISTERED NURSE (PDRN) CONTACTED FRESENIUS TO DEACTIVATE THE PATIENT'S ACCOUNT IN ORDER FOR THE PATIENT'S FAMILY TO NOT RECEIVE ADDITIONAL SUPPLIES. ADDITIONAL INFORMATION WAS OBTAINED THROUGH FOLLOW-UP WITH THE PATIENT¿S PDRN. THE PATIENT WAS FOUND DECEASED ON (B)(6) 2025. THE PATIENT WAS STILL CONNECTED TO THE LIBERTY SELECT CYCLER WHEN FOUND. HOWEVER, THE LAST COMPLETED TREATMENT WAS RECORDED AS (B)(6) 2025. THE PATIENT¿S RECORDS DOCUMENTED THAT A FULL TREATMENT WAS COMPLETED WITHOUT ANY ISSUES. THE CAUSE OF DEATH WAS DOCUMENTED AS CHRONIC OBSTRUCTIVE PULMONARY DISEASE (COPD). NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: ADDITIONAL INFORMATION: D9, G4, H3 PLANT INVESTIGATION: THE ACTUAL DEVICE WAS RETURNED TO THE MANUFACTURER FOR PHYSICAL EVALUATION. AN EXTERNAL VISUAL INSPECTION WAS PERFORMED ON THE CYCLER AND SHOWED SIGNS OF PHYSICAL DAMAGE DUE TO DISPLAY MISALIGNED. ALTHOUGH THERE WAS NO EVIDENCE OF DRIED FLUID PRESENT WITHIN THE CASSETTE COMPARTMENT, THERE WERE NO BURRS OR SHARP EDGES IN THE CASSETTE AREA THAT COULD HAVE PUNCTURED THE CASSETTE MEMBRANE. THERE WERE VISUAL INDICATIONS OF PARTICULATES WITHIN THE CASSETTE AREA. A SIMULATED TREATMENT USING (AS-RECEIVED) TREATMENT SETTINGS WITH REDUCED DWELL TIMES WAS PERFORMED AND COMPLETED WITHOUT FAILURES. NO FLUID LEAKS IN THE TEST CASSETTE DURING THE TREATMENT TEST. THE CYCLER UNDERWENT AND PASSED A SYSTEM AIR LEAK TEST, VALVE ACTUATION TEST, AND TECH PUMPS TEST. THERE WERE NO VISUAL DISCREPANCIES ENCOUNTERED DURING THE INTERNAL INSPECTION. A REVIEW OF THE DEVICE MANUFACTURING RECORDS WAS CONDUCTED BY THE MANUFACTURER. THERE WERE NO DEVIATIONS OR NON-CONFORMANCES DURING THE MANUFACTURING PROCESS. IN ADDITION, A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND VERIFIED THAT THE RESULTS OF THE IN-PROGRESS AND FINAL QUALITY CONTROL (QC) TESTING MET ALL REQUIREMENTS. UPON COMPLETION OF THE EVALUATION, THERE WERE NO MALFUNCTIONS THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE CYCLER PERFORMED AS DESIGNED AND AN ASSOCIATED CAUSE COULD NOT BE DETERMINED. Manufacturer Narrative: CLINICAL INVESTIGATION: THERE IS A TEMPORAL RELATIONSHIP BETWEEN PERITONEAL DIALYSIS UTILIZING THE LIBERTY SELECT CYCLER AND THE PATIENT EVENT OF DEATH. HOWEVER, THERE IS NO DOCUMENTATION OF A CAUSAL RELATIONSHIP BETWEEN THE PATIENT DEATH AND USE OF THE LIBERTY SELECT CYCLER. THE PATIENT WAS FOUND DECEASED ON THE CYCLER, SO IT IS UNKNOWN AT WHAT TIME THE PATIENT EXPIRED. HOWEVER, THE PATIENT DID COMPLETE A FULL TREATMENT PRIOR TO THE DEATH. THE PATIENT¿S CAUSE OF DEATH IS DOCUMENTED AS COPD. THIS COMBINATION OF COPD AND CKD IS INDEPENDENTLY ASSOCIATED WITH A HIGHER PREVALENCE OF OTHER COMORBIDITIES (ESPECIALLY CARDIOVASCULAR) AND INCREASED MORTALITY. BASED ON THE AVAILABLE INFORMATION AND NO ALLEGATION OR EVIDENCE OF A MALFUNCTION OR DEFICIENCY, THE LIBERTY SELECT CYCLER CAN BE EXCLUDED AS THE CAUSE OF THE PATIENT¿S DEATH. THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21689558&pc=OZD | 1220648-2025-26943 | 2025-04-15 04:00 | Death | ABIOMED, INC. | 2025-03-25 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A 76 YEAR OLD FEMALE PATIENT WHO HAD A CP PUMP FOR SUPPORT FOR HIGH RISK PERCUTANEOUS CARDIAC INTERVENTION (HRPCI). THE PATIENT WAS NOTED TO HAVE EXPIRED AS A DO NOT RESUSCITATE (DNR). ABIOMED'S LONG-TERM OUTCOME AND QUALITY INDICATOR (LOQI) DATABASE LATER REPORTED THAT AN ADVERSE EVENT OF VENTRICULAR TACHYCARDIA (VT) WITH "POSSIBLE" RELATIONSHIP TO THE PUMP USE. NO ADDITIONAL DETAILS WERE PROVIDED. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMART ASSIST SYSTEM STATES THE FOLLOWING: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21689831&pc=DSQ | 2916596-2025-01776 | 2025-02-06 05:00 | Death | THORATEC CORPORATION | 2025-03-25 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO VASOPLEGIA AND RIGHT VENTRICULAR (RV) FAILURE. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED. THE DEVICE OPERATED AS EXPECTED.Event Description: RIGHT VENTRICULAR (RV) FAILURE DID NOT EXIST PRE-LEFT VENTRICULAR ASSIST DEVICE (LVAD) IMPLANT. A DEVICE RELATED ISSUE WAS NOT CONSIDERED TO CAUSE OR CONTRIBUTE TO RV FAILURE. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. IT WAS COMMUNICATED THAT THE PUMP WOULD NOT BE RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING RIGHT HEART FAILURE AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. ADDITIONALLY, SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿RIGHT HEART FAILURE¿) STATES THAT PATIENTS MAY DEVELOP RIGHT VENTRICULAR FAILURE DURING OR SHORTLY AFTER IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS. THIS SECTION (UNDER ¿CAUTION!¿) ALSO EXPLAINS THAT RIGHT VENTRICULAR DYSFUNCTION, ESPECIALLY WHEN COMBINED WITH PULMONARY VASCULAR RESISTANCE, MAY LIMIT THE EFFECTIVENESS OF THE LVAS DUE TO REDUCED FILLING OF THE PUMP. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21689851&pc=DSQ | 2916596-2025-01802 | 2025-03-07 05:00 | Death | THORATEC CORPORATION | 2025-03-25 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE PATIENT OUTCOME, THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. THE OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND THE REPORTED PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICLE ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21689860&pc=DSI | 1218950-2025-000185 | 2025-02-27 05:00 | Death | PHILIPS MEDICAL SYSTEMS | 2025-03-25 04:00 | DSI | MX40 PATIENT WEARABLE MONITOR | Defective Alarm | Insufficient Information | K113125 |  | 1 | Event Description: IT WAS REPORTED THE DEVICE DID NOT ALARM FOR AN OXYGEN DESATURATION AND THE PATIENT PASSED AWAY. IT WAS INDICATED THE NURSES THOUGHT THE TELEMETRY DEVICE DID NOT ALARM FOR LOW SPO2; HOWEVER, IT WAS ALSO INDICATED THERE WAS AN ALARM PRESENT FOR AN EXTENDED PERIOD OF TIME. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTION WAS REPORTED; THEREFORE, GOOD FAITH EFFORTS ARE BEING PERFORMED.Event Description: IT WAS INDICATED THE NURSES THOUGHT THE TELEMETRY DEVICE DID NOT ALARM FOR LOW SPO2; HOWEVER, IT WAS ALSO INDICATED THERE WAS AN ALARM PRESENT FOR AN EXTENDED PERIOD OF TIME. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTION WAS REPORTED. THE DEVICE WAS IN USE MONITORING A PATIENT AT THE TIME OF THE EVENT. AN ADVERSE EVENT INVOLVING A PATIENT OR USER WAS REPORTED. Manufacturer Narrative: ATTEMPTS WERE MADE TO OBTAIN THE DEVICE CONTEXT OF USE, EVALUATION, REPAIR, AND OPERATIONAL STATUS WITH NO ADDITIONAL INFORMATION FROM THE CUSTOMER. IT WAS FURTHER CLARIFIED THAT THE CUSTOMER CANCELLED THE WORK ORDER. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED IF NEW INFORMATION IS OBTAINED. Manufacturer Narrative: PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE. REPORTING ADDRESS POSTAL: (B)(6).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21689918&pc=DSQ | 2916596-2025-01639 | 2025-01-24 05:00 | Death | THORATEC CORPORATION | 2025-03-25 04:00 | DSQ | HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM, JP | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Insufficient Information; No Clinical Signs, Symptoms or Conditions | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH/ DETAILS LEADING UP TO DEATH WAS UNKNOWN. IT WAS UNKNOWN IF THE DEATH WAS CONSIDERED TO BE DEVICE RELATED OR IF THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: THIS EVENT WAS DETERMINED TO BE SUPPLEMENTAL INFORMATION TO MANUFACTURER REPORT NUMBER 2916596-2025-00767. A FINAL REPORT WILL BE SUBMITTED UNDER THAT EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21689986&pc=OZD | 1220648-2025-26944 | 2025-02-27 05:00 | Death | ABIOMED, INC. | 2025-03-25 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemolysis; Hemorrhage/Bleeding; Perforation of Vessels | P140003 |  | 1 | Event Description: THE COMPLAINANT HAD A IMPELLA CP PUMP PLACED TO SUPPORT A PATIENT ADMITTED IN WITH ACUTE MYOCARDIAL INFARCTION AND CARDIOGENIC SHOCK. THE PUMP WAS SUPPORTING WHEN THERE WAS A BLEED AT THE ACCESS SITE THAT WAS DUE TO A VASCULAR DAMAGE DURING INSERTION. THE SURGEON DID VASCULAR REPAIR. THE TEAM ALSO NOTED SIGNS OF HEMOLYSIS. THERE WAS VERY LITTLE URINE OUTPUT AND POSITIVE HEMOLYSIS LABS. THE PUMP REMAINS ON BUT THE TEAM WAS WEANING AND PER THE LOCAL TEAM THE PATIENT EXPIRED. THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21690231&pc=BSZ | 2112667-2025-02474 | 2025-02-06 05:00 | Death | DATEX-OHMEDA, INC. | 2025-03-25 04:00 | BSZ | AISYS | No Flow | Cardiac Arrest | K172045 |  | 1 | Event Description: THE HOSPITAL REPORTED DIFFICULTY VENTILATING A PATIENT CONNECTED TO AN AISYS. ALLEGEDLY, THE UNIT STOPPED VENTILATING. IT WAS REPORTED THAT THE PATIENT DIED. GE HEALTHCARE WILL SUBMIT A FOLLOW-UP REPORT WHEN THE INVESTIGATION HAS BEEN COMPLETED. Manufacturer Narrative: GE HEALTHCARE INVESTIGATION INTO THE REPORTED OCCURRENCE IS ONGOING. A FOLLOW-UP REPORT WILL BE ISSUED WHEN THE INVESTIGATION HAS BEEN COMPLETED. BLOCK A5 AND A6: NO INFORMATION PROVIDED. D4 PRIMARY UDI NUMBER: THERE IS NO UDI AVAILABLE AS THE DEVICE WAS MANUFACTURED PRIOR TO UDI COMPLIANCE REQUIREMENTS. LEGAL MANUFACTURER: HCS MADISON - 3030 OHMEDA DR, USA MADISON, WI 53718. Manufacturer Narrative: GE HEALTHCARE'S (GEHC) INVESTIGATION HAS BEEN COMPLETED AND THE ROOT CAUSE COULD NOT BE DETERMINED. THE GEHCÂ'S FIELD ENGINEER COMPLETED A REVIEW OF THE SYSTEM LOGS AND DETERMINED THERE WAS NO MALFUNCTION OF THE GEHC DEVICE. THEREFORE, THE ROOT CAUSE OF THE PATIENT DEATH IS UNDETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21690288&pc=LOX | 2124215-2025-16925 | 2025-03-02 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-03-25 04:00 | LOX | NC EMERGE | Defective Device | Low Blood Pressure/ Hypotension; Perforation of Vessels; Loss of consciousness |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. THE PATIENT WAS PRESENTED WITH COMPLAINTS OF CHEST PAIN. AN ELECTROCARDIOGRAM (ECG) REVEALED DIFFUSE ST-SEGMENT DEPRESSIONS ALONG WITH ST ELEVATIONS IN THE AVR LEAD AND A LARGE MYOCARDIAL INFARCTION (MI WAS DIAGNOSED. THE PATIENT WAS IMMEDIATELY TRANSFERRED TO THE CATHETERIZATION ROOM FOR FURTHER EVALUATION AND TREATMENT. DURING CORONARY ANGIOGRAPHY, SIGNIFICANT STENOSIS OF THE LEFT MAIN ARTERY (LM) WAS IDENTIFIED, ACCOMPANIED BY HEAVY CALCIFICATION IN THE DISTAL SEGMENT OF THE ARTERY, WHICH COMPLICATED THE INTERVENTION. WIRES WERE ADVANCED INTO THE LEFT ANTERIOR DESCENDING (LAD) AND LEFT CIRCUMFLEX (LCX) ARTERIES TO FACILITATE TREATMENT. FOLLOWING THIS, BALLOON DILATION WAS PERFORMED USING A 2.5 NC EMERGE BALLOON CATHETER TO OPEN THE NARROWED ARTERIES, AFTER WHICH A 4.5 X 15 NON-BOSTON SCIENTIFIC STENT WAS IMPLANTED. HOWEVER, FURTHER DILATION WITH A 5.00MM X 8MM NC EMERGE BALLOON WAS NECESSARY TO EXPAND THE STENT IN A CALCIFIED AREA THAT DID NOT OPEN PROPERLY. FOLLOWING DILATATION, A SIGNIFICANT PERFORATION OF THE ARTERY OCCURRED, LEADING TO BLOOD EXTRAVASATION INTO THE PERICARDIUM. THIS RESULTED IN AN ACUTE DETERIORATION OF THE PATIENT'S CONDITION, LOST CONSCIOUSNESS, AND BLOOD PRESSURE DROPPED SIGNIFICANTLY, PROMPTING THE ACTIVATION OF A CODE BLUE. DESPITE IMMEDIATE RESUSCITATIVE EFFORTS, INCLUDING ATTEMPTS TO PROVIDE MECHANICAL CIRCULATORY SUPPORT VIA EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO)AND INTRA-AORTIC BALLOON PUMP (IABP), THESE INTERVENTIONS FAILED DUE TO THE LACK OF ACCESS. THE PATIENTS PERIPHERAL ARTERIES WERE SEVERELY CALCIFIED, PREVENTING EFFECTIVE ACCESS FOR THE DEVICES AND EVENTUALLY, THE PATIENT DIED. ACCORDING TO THE HOSPITAL, THE BALLOON LENGTH SHOULD HAVE BEEN 8MM, BUT THE BALLOON MEASURED LENGTH WAS16MM ON FILM.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21690323&pc=DSQ | 2916596-2025-01652 | 2024-04-04 04:00 | Death | THORATEC CORPORATION | 2025-03-25 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Ischemia; Necrosis; Renal Failure; Septic Shock; Thrombosis/Thrombus; Drug Resistant Bacterial Infection | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLES ¿ALPHA-GAL-KO XENOTHYMOKIDNEY FUNCTION IN A LIVING HUMAN RECIPIENT WITH A DURABLE MECHANICAL CIRCULATORY SUPPORT DEVICE¿ AND ¿ELECTROLYTE HOMEOSTASIS IN A LIVING HUMAN RECIPIENT WITH AN ALPHA-GAL-KO XENOTHYMOKIDNEY¿ THAT HEARTMATE 3 (HM3) LEFT VENTRICULAR ASSIST DEVICE (LVAD) IMPLANT MAY BE ASSOCIATED WITH KIDNEY FAILURE, RESPIRATORY FAILURE, VANCOMYCIN-RESISTANT ENTEROCOCCI (VRE) BACTEREMIA, PSEUDOMONAS AERUGINOSA (P. AERUGINOSA) BACTERIURIA, CANDIDA FUNGEMIA, SEPTIC SHOCK, ACUTE TUBULAR INJURY (ATI), CORTICAL NECROSIS, INTERSTITIAL HEMORRHAGE, THROMBUS, HYPOTENSION, AND DEATH. AN ALPHA-GAL-KO XENOTHYMOKIDNEY WAS TRANSPLANTED INTO A LIVING HUMAN 8 DAYS FOLLOWING HM3 IMPLANTATION. IMMUNOSUPPRESSION INCLUDED THYMOGLOBULIN, BELATACEPT, RITUXIMAB, CORTICOSTEROIDS, MYCOPHENOLATE MOFETIL, TACROLIMUS, PEGCETACOPLAN, AND ECULIZUMAB. PLASMA EXCHANGE WAS PERFORMED ON DAYS -4 AND -2 PRETRANSPLANT AND ON POST-OPERATIVE DAYS (POD) 3, 5, 7, 26, 35, 36, 38, AND 40. URINE OUTPUT, SERUM CREATININE (SCR), 6-HOUR CREATININE CLEARANCE (6CRCL), AND ESTIMATED GLOMERULAR FILTRATION RATE (EGFR) BY CYSTATIN C WERE MONITORED. BIOPSY WAS PERFORMED ON POD 0, 13, 32, 42, AND 47. SCR DECREASED FROM 1.52MG/DL (POST HEMODIALYSIS) TO 0.55MG/DL ON POD 1 WHILE THE PATIENT WAS ON VASOPRESSOR SUPPORT. BY POD 2, SCR INCREASED TO A VALUE OF 1.55MG/DL SECONDARY TO AN APPARENT MISMATCH BETWEEN OPTIMAL LVAD HEMODYNAMICS AND XENOGRAFT PERFUSION REQUIREMENTS. DURING THIS PERIOD 6CRCL DECREASED FROM 75.8ML/MIN TO 12ML/MIN. SCR AND 6CRCL NORMALIZED AFTER ALTERING LVAD SETTINGS AND VASOPRESSOR SUPPORT TO CREATE PULSATILE FLOW WITH A GOAL OF SYSTOLIC PRESSURES =100MMHG. FROM POD 7 TO 30 THE AVERAGE SCR AND 6CRCL WAS 0.73MG/DL AND 50.38ML/MIN, RESPECTIVELY. EGFR CLOSELY APPROXIMATED THE 6CRCL AT 43.33ML/MIN/BSA. BIOPSY ON POD 13 DID NOT REVEAL GLOMERULITIS, TUBULITIS, ARTERITIS, OR THROMBOTIC MICROANGIOPATHY (TMA) ON H&E STAINING. XENOGRAFT FUNCTION BEGAN TO DECLINE BY POD31 AS BACTEREMIA AND A DRUG FEVER MADE IT DIFFICULT TO MAINTAIN HEMODYNAMIC GOALS AND VASOPRESSOR SUPPORT HAD TO BE INCREASED, WITH SCR PEAKING AT 3.69MG/DL AND 6CRCL LOW OF 1.10ML/MIN. BIOPSY ON POD 32 SHOWED SIGNIFICANT ACUTE TUBULAR INJURY WITHOUT EVIDENCE OF REJECTION, POSSIBLY A RESULT OF A MISMATCH BETWEEN OPTIMAL LVAD HEMODYNAMICS AND XENOGRAFT REQUIREMENTS IN THE SETTING OF BACTEREMIA AND INCREASED VASOPRESSOR SUPPORT. BIOPSY ON POD 42 SHOWED PROGRESSION TO CORTICAL NECROSIS AND INTERSTITIAL HEMORRHAGE WITHOUT EVIDENCE OF IMMUNOLOGIC REJECTION. CRRT WAS INITIATED ON POD 45 WITH EXPLANT OF XENOGRAFT ON POD 47. ADDITIONALLY, ROUTINE SERUM ELECTROLYTE MEASUREMENTS WERE PERFORMED. POTASSIUM, MAGNESIUM, PHOSPHATE, AND CALCIUM WERE REPLETED WHEN CLINICALLY INDICATED. FROM XENOGRAFT REPERFUSION UNTIL CRRT INITIATION ON POD 45, AVERAGE SERUM ELECTROLYTE LEVELS WERE AS FOLLOWS: SODIUM 143 MMOL/L (RANGE: 135-151), POTASSIUM 4.02 MMOL/L (RANGE: 3.0-5.9), MAGNESIUM 2.10 MG/DL (RANGE: 1.4-3.2), PHOSPHATE 4.21 MMOL/L (RANGE: 2.2-6.9), CHLORIDE 102 MG/DL (RANGE: 89-112), AND CALCIUM 8.76 MG/DL (RANGE: 7.2-10.1). TO ADDRESS HYPOKALEMIA DUE TO DIURESIS WITH NA-CL LOOP COTRANSPORTER INHIBITORS, AN AVERAGE OF 103 MEQ/DAY OF POTASSIUM CHLORIDE WAS ADMINISTERED. ADDITIONALLY, AN AVERAGE OF 2 G/DAY OF MAGNESIUM SULFATE WAS GIVEN DURING THE SAME PERIOD. SERUM PHOSPHORUS LEVELS BEGAN TO RISE ON POD 29. HYPOCALCEMIA WAS OBSERVED ON POD 35, AND FROM THAT POINT UNTIL THE INITIATION OF CRRT, AN AVERAGE OF 1 G OF CALCIUM GLUCONATE WAS ADMINISTERED DAILY. IN A CRITICALLY ILL PATIENT, SERUM SODIUM, MAGNESIUM, AND CHLORIDE REMAINED WITHIN NORMAL LIMITS THROUGHOUT THE DURATION OF XENOGRAFT IMPLANTATION. FREQUENT POTASSIUM REPLETION WAS NECESSARY IN SETTING OF AGGRESSIVE DIURESIS. SERUM PHOSPHATE AND CALCIUM WERE NORMAL UNTIL XENOGRAFT FUNCTION DETERIORATED. STABLE KIDNEY FUNCTION WAS OBSERVED THE FIRST MONTH POST-PROCEDURE BUT DETERIORATED IN THE SETTING OF DIFFICULTY MEETING XENOGRAFT BLOOD PRESSURE REQUIREMENTS WITH LVAD SUPPORT IN THE CONTEXT OF PROFOUND SYSTEMIC VASODILATION AND RISING VASOPRESSOR REQUIREMENTS. THE POST-OPERATIVE COURSE WAS ALSO COMPLICATED BY MULTIPLE INFECTIONS INCLUDING VRE BACTEREMIA, NOROVIRUS, ORAL HERPES SIMPLEX VIRUS 1 (HSV1), AND PSEUDOMONAS AERUGINOSA (P. AERUGINOSA) BACTERIURIA, AS WELL AS PRESSURE ULCERS AND DECONDITIONING, RESPIRATORY FAILURE REQUIRING A TRACHEOSTOMY, AND THE EVENTUAL XENOGRAFT FAILURE, WHICH WAS DETERMINED TO LIKELY BE DUE TO A COMBINATION OF THE CHRONIC ISCHEMIC INJURY AND ANTIBODY-MEDIATED REJECTION (AMR). POST-EXPLANTATION, THE PATIENT DEVELOPED SEPTIC SHOCK IN THE SETTING OF CANDIDA FUNGEMIA. LVAD SUPPORT WAS WITHDRAWN ON (B)(6) 2024, AND THE PATIENT PASSED AWAY.Event Description: IT WAS ADDITIONALLY REPORTED FROM THE RESEARCH ARTICLE "ALPHA-GAL-KO XENOTHYMOKIDNEY RESPONSE TO ENDOGENOUS ARGININE VASOPRESSIN IN A LIVING HUMAN RECIPIENT WITH A DURABLE MECHANICAL CIRCULATORY SUPPORT DEVICE" THAT THIS SINGLE CENTER CASE STUDY EVALUATED A HIGHLY SENSITIZED 54-YEAR-OLD FEMALE, CALCULATED PANEL REACTIVE ANTIBODY (CPRA) OF 93%, WITH MULTIPLE CO-MORBIDITIES INCLUDING END-STAGE HEART AND KIDNEY DISEASE WHO WAS DETERMINED TO NOT BE A HEART KIDNEY TRANSPLANT CANDIDATE AND UNDERWENT XENOTRANSPLANT WITH AN ALPHA-GAL-KO XENOTHYMOKIDNEY 8 DAYS FOLLOWING IMPLANTATION OF A HEARTMATE 3 LEFT VENTRICULAR ASSIST DEVICE (LVAD). PERIODS OF HYPERNATREMIA PRE-CRRT INITIATION CORRESPONDED TO LIMITED FREE WATER INTAKE AND ONGOING DIURETIC USE FOR WORSENING VOLUME STATUS AND CORRELATE WITH RISING COPEPTIN LEVELS. Manufacturer Narrative: SECTION D6A: DATE OF IMPLANT CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. NO PRODUCT WAS RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, REV. C AND THE HEARTMATE 3 PATIENT HANDBOOK, REV. D ARE CURRENTLY AVAILABLE. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING INFECTION (LOCAL, DRIVELINE, PUMP POCKET), SEPSIS, MULTI ORGAN FAILURE (INCLUDING RESPIRATORY FAILURE AND RENAL DYSFUNCTION), THROMBOEMBOLISM, BLEEDING, AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 5 OF THE IFU, "SURGICAL PROCEDURES", INSTRUCTS TO INSPECT THE VENTRICULAR CHAMBER FOR MURAL THROMBI AND CROSSING TRABECULAE FOLLOWING REMOVAL OF THE CORE AND TO ADDRESS ONE OR BOTH, AS NEEDED. THIS SECTION ALSO INSTRUCTS TO INSPECT THE VENTRICLE AND REMOVE ANY PREVIOUSLY FORMED CLOTS THAT MAY CAUSE EMBOLISM OR ANY TRABECULAE THAT MAY IMPEDE FLOW. SECTION 6, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿ANTICOAGULATION¿), PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) RANGE, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. THIS SECTION ALSO LISTS INFECTION AND THROMBOEMBOLISM AS POTENTIAL LATE POSTIMPLANT COMPLICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: H6 - ADDITIONAL HEALTH EFFECT - CLINICAL CODES: 2484 - RESPIRATORY FAILURE, 2068 - SEPTIC SHOCK, 2419 - FUNGAL INFECTION, 2120 - URINARY TRACT INFECTION, 4846 - BACTEREMIA, 1914 - HYPOTENSION, 2127 - VASODILATION, 4825 - HYPOKALEMIA, 4823 - HYPOCALCEMIA, 2248 - VIRAL INFECTION, 2326 - PRESSURE SORES. ALJABBAN, I. ET AL. (2025). ALPHA-GAL-KO XENOTHYMOKIDNEY FUNCTION IN A LIVING HUMAN RECIPIENT WITH A DURABLE MECHANICAL CIRCULATORY SUPPORT DEVICE. AMERICAN JOURNAL OF TRANSPLANTATION, 25(1). HTTPS://DOI.ORG/10.1016/J.AJT.2024.12.019. ALJABBAN, I., ET AL. (2025). ELECTROLYTE HOMEOSTASIS IN A LIVING HUMAN RECIPIENT WITH AN ALPHA-GAL-KO XENOTHYMOKIDNEY. AMERICAN JOURNAL OF TRANSPLANTATION, 25(1), S88. HTTPS://DOI.ORG/10.1016/J.AJT.2024.12.168. NEW YORK UNIVERSITY LANGONE TRANSPLANT INSTITUTE, NEW YORK, NY. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21690519&pc=LWP | 2017865-2025-38475 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-03-25 04:00 | LWP | ASSURITY MRI | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P140033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED WITHOUT CLEAR INFORMATION AS WHETHER THE ADVISORY PULSE GENERATOR CAUSED OR CONTRIBUTED TO THE DEATH. NO FURTHER INFORMATION WAS AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21690939&pc=OZD | 1220648-2025-26949 | 2025-02-28 05:00 | Death | ABIOMED, INC. | 2025-03-25 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Retroperitoneal Hemorrhage | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS ON IMPELLA CP SUPPORT AND THE PUMP WAS EXPLANTED. TWO HOURS AFTER THE EXPLANT, THE PATIENT STARTED TO DECOMPENSATE. FIVE UNITS OF BLOOD PRODUCTS WERE GIVEN. THE PHYSICIAN SUSPECTED A RETROPERITONEAL BLEED FOLLOWING IMPELLA REMOVAL. THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21691701&pc=PRL | 2017233-2025-05978 | 2025-02-24 05:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-03-25 04:00 | PRL | GORE® VIABAHN® VBX BALLOON EXPANDABLE ENDOPROSTHESIS | Adverse Event Without Identified Device or Use Problem | Calcium Deposits/Calcification; Hemorrhage/Bleeding; Rupture | P160021 |  | 1 | Event Description: THIS PROCEDURE WAS INITIALLY PLANNED FOR A COVERED ENDOVASCULAR RECONSTRUCTION OF AORTIC BIFURCATION (CERAB) TO ADDRESS HEAVILY CALCIFIED AORTA AND COMMON ILIAC ARTERIES. THE PROCEDURE BEGAN WITH A RIGHT FEMORAL ENDARTERECTOMY, FOLLOWED BY WIRE ACCESS BEING ACHIEVED BILATERALLY. AN 11MM X 59MM GORE® VIABAHN® VBX BALLOON EXPANDABLE ENDOPROSTHESIS (VBX DEVICE S/N (B)(6) WAS ADVANCED TO THE DISTAL AORTA, AND THE BALLOON WAS INFLATED TO RATED BURST PRESSURE. THE VBX DEVICE DELIVERY SYSTEM CATHETER WAS WITHDRAWN SUCCESSFULLY. SUBSEQUENTLY, A 16MM ARMADA BALLOON WAS ADVANCED INTO THE 11MM X 59MM VBX DEVICE AND INFLATED FOR POST DILATION TO 16MM. AT APPROXIMATELY NOMINAL PRESSURE THE ARMADA BALLOON RUPTURED. AT THIS TIME, A FLUSH CATHETER WAS INTRODUCED TO TAKE IMAGES AND REVEALED AN AORTIC RUPTURE AT THE MID-LEVEL OF THE VBX DEVICE. TO MANAGE THE LEAKAGE AND RUPTURE, PHYSICIAN DECIDED TO USE A 20MM GORE® EXCLUDER® CONFORMABLE AAA ENDOPROSTHESIS (AORTIC CUFF) TO COVER THE AREA OF LEAKAGE AND RUPTURE. THE PHYSICIAN THEN ATTEMPTED TO INTRODUCE THE CUFF THROUGH A GORE® 15FR DRYSEAL FLEX INTRODUCER SHEATH ON THE RIGHT, BUT ADVANCEMENT WAS HINDERED DUE TO CALCIFIED ILIAC. IN THE PROCESS OF CREATING A BIGGER HOLE IN THE FEMORAL EXCESSIVE BLEEDING OCCURRED EXTENDING DOWN TO THE SUPERFICIAL FEMORAL ARTERY (SFA). THE PHYSICIAN DETERMINED THAT HE NEEDED TO PLACE A 7MM GORE® VIABAHN® ENDOPROSTHESIS (VIABAHN DEVICE) AT THE PROXIMAL SFA TO CONTROL THE BLEEDING. AT THIS TIME, A 7MM X 5CM GORE® VIABAHN® ENDOPROSTHESIS (VIABAHN DEVICE SERIAL NUMBER (B)(6) WAS PLACED NEAR THE ORIGIN OF THE SFA/PROFUNDA. HOWEVER, IT DID NOT QUITE REACH THE ORIGIN, SO AN ADDITIONAL 7MM X 7.5CM GORE® VIABAHN® ENDOPROSTHESIS (VIABAHN DEVICE SERIAL NUMBER (B)(6) WAS POSITIONED HIGHER INSIDE THE 7MM X 5CM VIABAHN DEVICE AND AT THE SFA/PROFUNDA ORIGIN, SUCCESSFULLY CONTROLLING THE BLEEDING. NEXT A 11MM X29MM GORE® VIABAHN® VBX BALLOON EXPANDABLE ENDOPROSTHESIS (VBX DEVICE SERIAL NUMBER (B)(6) WAS INTRODUCED IN THE DISTAL AORTA WITH A PLAN TO POST DILATE TO 16MM AT THE LEVEL OF THE RUPTURE. THE 11MM X 29MM VBX DEVICE WAS INFLATED TO NOMINAL PRESSURE AND HELD IN PLACE BY AN 8MM BALLOON (UNKNOWN BRAND) FROM THE LEFT SIDE. A 16MM BALLOON WAS INTRODUCED AND INFLATED TO POST DILATE THE VBX DEVICE TO 16MM. DESPITE THESE EFFORTS, SUBSEQUENT IMAGING SHOWED CONTINUED LEAKAGE. IT WAS DETERMINED THAT THE LEAK MAY BE COMING FROM ABOVE DUE TO LACK OF SEAL. TO ADDRESS THIS, A 11MM X 29MM GORE® VIABAHN® VBX BALLOON EXPANDABLE ENDOPROSTHESIS (VBX DEVICE S/N (B)(6) WAS INTRODUCED IN THE DISTAL AORTA AND EXTENDED UPWARDS TO ACHIEVE A BETTER SEAL. IT WAS POST DILATED TO 16MM AND IMAGING CONTINUED TO SHOW LEAKAGE. THE PHYSICIAN THEN ATTEMPTED AN 18MM ARMADA BALLOON AT THE TOP TO TRY AND GET IT TO SEAL; HOWEVER, THE BALLOON RUPTURED. AS THE PATIENT¿S BLOOD PRESSURE DROPPED AND STOMACH BECAME DISTENDED, THE PHYSICIAN CALLED IN ANOTHER PHYSICIAN TO PERFORM AN OPEN SURGERY TO STOP THE INTERNAL BLEEDING. UNFORTUNATELY, THE INTERNAL BLEEDING COULD NOT BE CONTROLLED AND THE PATIENT EXPIRED ON THE OPERATING TABLE. ADDITIONALLY, IT WAS REPORTED THAT DURING THE PROCEDURE TWO 9MM X 29MM VBX DEVICES (SERIAL NUMBERS (B)(6) AND TWO 9MM X 39MM VBX DEVICES (SERIAL NUMBERS (B)(6) WERE USED AND IMPLANTED IN THE COMMON ILIAC ARTERIES. THE PHYSICIAN SUSPECTS THAT THE ISSUE WAS CAUSED BY SEVERE CALCIFICATION WHICH LED TO AORTIC RUPTURE FOLLOWING VBX DEVICE POST DILATION. THE CAUSE WAS BELIEVED TO BE DUE TO SHARP CALCIUM WHICH PERFORATED THROUGH EXPANDED POLYTETRAFLUOROETHYLENE (EPTFE) MATERIAL AND RUPTURED THE BALLOONS. THE PHYSICIAN WAS UNSURE OF THE EXACT CAUSE OF DEATH BUT SUSPECTS IT COULD BE RELATED TO BLOOD LOSS. Manufacturer Narrative: SECTION H6 UPDATED TO REFLECT COMPLETION OF INVESTIGATION. A REVIEW OF MANUFACTURING RECORDS VERIFIED THAT THE LOT INVOLVED IN THIS COMPLAINT MET ALL PRE-RELEASE SPECIFICATIONS. THE DEVICE REMAINS IMPLANTED AND WAS, THEREFORE, NOT AVAILABLE FOR ANALYSIS. NO CLINICAL IMAGES ENABLING DIRECT ASSESSMENT OF PRODUCT PERFORMANCE WERE RETURNED FOR EVALUATION. CAUSE OF THE REPORTED EVENT CANNOT BE ESTABLISHED BASED ON THE INFORMATION REPORTED TO GORE. FURTHER INFORMATION REGARDING THIS EVENT WAS REQUESTED BY GORE, BUT NO FURTHER INFORMATION HAS BEEN REPORTED, THEREFORE THIS INVESTIGATION IS CONSIDERED COMPLETE. Manufacturer Narrative: CBAS® HEPARIN SURFACE INCORPORATES CARMEDA HEPARIN MANUFACTURED FROM HEPARIN SODIUM API, WHICH IS COVALENTLY BOUND TO THE DEVICE SURFACE AND IS ESSENTIALLY NON-ELUTING. H6 - CODE B20: DEVICE REMAINS IMPLANTED, AND NO IMAGES WERE PROVIDED; THEREFORE, DIRECT PRODUCT ANALYSIS WAS NOT POSSIBLE. H6 - CODE C21: RESULTS PENDING COMPLETION OF INVESTIGATION. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21691777&pc=LWP | 2017865-2025-38492 | 2025-03-06 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-03-25 04:00 | LWP | ENDURITY MRI PACEMAKER, SR | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P140033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. IT WAS REPORTED THAT THE CAUSE OF DEATH IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21691797&pc=NVN | 2017865-2025-38493 | 2025-03-06 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-03-25 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. IT WAS REPORTED THAT THE CAUSE OF DEATH IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21692812&pc=LKK | 2182207-2025-00807 | 2025-03-07 05:00 | Death | MEDTRONIC NEUROMODULATION | 2025-03-25 04:00 | LKK | IMPLANTABLE DRUG INFUSION PUMP | Adverse Event Without Identified Device or Use Problem | Cognitive Changes |  |  | 1 | Event Description: INFORMATION WAS RECEIVED FROM A FOREIGN HEALTHCARE PROVIDER (HCP) VIA A DISTRIBUTER (DIST) REGARDING A PATIENT RECEIVING GABALON OF AN UNKNOWN DRUG CONCENTRATION AT AN UNKNOWN DOSE RATE VIA IMPLANTED INFUSION PUMP. IT WAS REPORTED THAT A REFILL WAS PERFORMED AT ANOTHER HOSPITAL; HOSPITAL WHERE THE PATIENT VISITS NORMALLY). THE PATIENT WAS TRANSPORTED DUE TO IMPAIRED CONSCIOUSNESS. INTRATHECAL BACLOFEN WAS BEING FOLLOWED UP AT ANOTHER HOSPITAL SO THE SITUATION WAS UNKNOWN. THE OUTCOME OF THE ADVERSE EVENT WAS NOT RECOVERED. THE CAUSAL RELATIONSHIP OF IMPAIRED CONSCIOUSNESS WAS UNRELATED TO DRUG, BUT UNKNOWN IF RELATED TO THE PUMP, CATHETER, PROGRAMMING, OR PROCEDURE.Event Description: ADDITIONAL INFORMATION RECEIVED FROM A FOREIGN HEALTH CARE PROVIDER (FOR, HCP) VIA A DISTRIBUTOR (DIST) INDICATED THAT THE SEVERITY CLASSIFICATION WAS DEATH. ADDITIONAL INFORMATION RECEIVED FROM A FOREIGN HEALTH CARE PROVIDER (FOR, HCP) VIA A DISTRIBUTOR (DIST) INDICATED THAT THE DETAILS REGARDING THE MODEL NUMBER, SERIAL NUMBER, AND IMPLANT DATE OF THE PUMP WERE NOT KNOWN. REGARDING THE PATIENT REQUIRING HOSPITALIZATION, IT WAS STATED THAT THE SITUATION WAS THAT THE PATIENT WAS TREATED BY THE DOCTOR IN THE EMERGENCY ROOM. SINCE THEN, HE HAD NOT VISITED THE HOSPITAL AND DID NOT TAKE CONSULTATION AT THE HOSPITAL. IT WAS STATED THAT SINCE APRIL, IT WAS DIFFICULT TO OBTAIN INFORMATION DUE TO THE DOCTOR'S TRANSFER. IT WAS STATED THAT THE PRODUCT HAD NOT BEEN REMOVED AND WOULD NOT BE RETURNED BECAUSE IT WAS STILL IN USE.Event Description: ADDITIONAL INFORMATION RECEIVED FROM A FOREIGN HEALTH CARE PROVIDER (FOR, HCP) VIA A DISTRIBUTOR (DIST) INDICATED THAT THE PATIENT WAS TREATED BY A DOCTOR IN THE EMERGENCY ROOM. SINCE THEN, THE PATIENT HAD NOT VISITED THE HOSPITAL, AND IN APRIL, THE DOCTOR WAS TRANSFERRED TO ANOTHER HOSPITAL SO IT WAS NOT POSSIBLE TO INVESTIGATE. IT WAS ALSO STATED THAT IT WAS JUST THAT THE PATIENT WAS DEALT WITH IN THE EMERGENCY ROOM AND THE PATIENT DID NOT COME TO THE HOSPITAL AFTER THAT, AND THAT THE PATIENT DID NOT DIE. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID NEU_UNKNOWN_CATH LOT# UNKNOWN. PRODUCT TYPE CATHETER. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: H1: DUE TO ADDITIONAL INFORMATION RECEIVED, THE TYPE OF REPORTABLE EVENT IS NO LONGER DEATH, BUT SERIOUS INJURY. H6: DUE TO ADDITIONAL INFORMATION RECEIVED, ALL CODES INDICATING DEATH AND HOSPITALIZATION NO LONGER APPLY TO THIS EVENT AND WERE REMOVED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: B2: FOLLOW-UP IS BEING CONDUCTED FOR THE DATE OF DEATH. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21692843&pc=OZD | 1220648-2025-26951 | 2025-02-27 05:00 | Death | ABIOMED, INC. | 2025-03-25 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Renal Failure; Gastrointestinal Hemorrhage | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A 64 YEAR OLD MALE PATIENT DEVELOPED A GASTROINTESTINAL BLEED DURING IMPELLA 5.5 SUPPORT. HEPARIN WAS BEING RUN THROUGH THE PURGE LINE AT THE TIME OF THE BLEED AND ANTICOAGULANTS WERE HALTED TO MANAGE THE CONDITION. SUPPORT CONTINUED UNINTERRUPTED. NINE DAYS LATER, THE FAMILY MADE THE DECISION TO WITHDRAW CARE AND THE PATIENT PASSED AWAY. THERE IS NOT ENOUGH EVIDENCE TO DISASSOCIATE THE DEATH WITH THE USE OF THE IMPELLA.Event Description: IT WAS LATER REPORTED VIA ABIOMED¿S LONG-TERM OUTCOME AND QUALITY INDICATOR (LOQI) IMPELLA REGISTRY THAT THE PATIENT ENDURED AN ACUTE KIDNEY INJURY (AKI) DURING IMPELLA SUPPORT. THE AKI WAS BELIEVED TO HAVE A POSSIBLE RELATIONSHIP TO THE IMPELLA DEVICE. CONTINUOUS RENAL REPLACEMENT THERAPY WAS INITIATED TO TREAT THE CONDITION. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK STATES THE FOLLOWING: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿. Manufacturer Narrative: THE INVESTIGATION OF VASCULAR DAMAGE - PERIPHERAL VESSEL HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. E4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26951 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. G1 REPORTING CONTACT FACILITY NAME AND FAX NUMBER WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26951. H6 HEALTH EFFECT - IMPACT CODE 4641 WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26951. Manufacturer Narrative: THE PRODUCT INVOLVED IN THE REPORT HAS NOT BEEN RETURNED. A REVIEW OF THE DEVICE HISTORY RECORD IS IN-PROGRESS. ALL INFORMATION REASONABLY KNOWN AS OF 25 MAR 2025 HAS BEEN INCLUDED IN THE HEALTH AUTHORITY REPORT. SHOULD ADDITIONAL INFORMATION BE OBTAINED A FOLLOW-UP HEALTH AUTHORITY REPORT WILL BE PROVIDED. THE INFORMATION PROVIDED BY AIRLIFE. REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO AIRLIFE. AIRLIFE HAS NO INDEPENDENT KNOWLEDGE OF THE EVENT REPORTED BUT IS RELAYING THE INFORMATION THAT WAS PROVIDED BY THE USER FACILITY WHERE THE INCIDENT OCCURRED. THIS PRODUCT INCIDENT IS DOCUMENTED IN THE AIRLIFE COMPLAINT DATABASE AND IS IDENTIFIED AS COMPLAINT (B)(4). THIS INFORMATION IS SUBMITTED PURSUANT TO 21CFR803, IN COMPLIANCE WITH THE MEDICAL DEVICE REPORTING REQUIREMENT AND SHOULD NOT BE CONSIDERED TO BE AN ADMISSION THAT A AIRLIFE PRODUCT IS DEFECTIVE OR CAUSES SERIOUS INJURY. THE COMPLAINT OF "SPONTANEOUS XXX ETCO2 MODULE ON SCREEN" REGARDING PARTS 301-5107ET AND 301-5707F-SE WAS NOT CONFIRMED. THE ROOT CAUSE WAS NOT DETERMINED BUT COULD POSSIBLY BE A RESULT OF USER ERROR. A RISK ASSESSMENT WAS PERFORMED AND THE ULTIMATE RISK WAS DETERMINED TO BE LOW WHICH DOES NOT REQUIRE ESCALATION TO THE CAPA REVIEW BOARD. THERE HAVE BEEN 9 OTHER COMPLAINTS REGARDING THE SAME PART AND A SIMILAR ISSUE WITHIN THE 24 MONTHS PRECEDING THIS REPORTED EVENT. A RESOLUTION LETTER WAS SENT TO THE CUSTOMER. COMPLAINTS WILL CONTINUE TO BE MONITORED FOR POSSIBLE TRENDS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21692911&pc=OZD | 1220648-2025-26954 | 2025-02-28 05:00 | Death | ABIOMED, INC. | 2025-03-25 04:00 | OZD | IMPELLA CP | Failure to Pump | Ischemia | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A 77-YEAR-OLD FEMALE PATIENT ON IMPELLA CP SUPPORTDURING SUPPORT, THERE WAS SLOW DISTAL PERFUSION NOTED VIA ANGIOGRAM FROM IMPELLA SIDE PORT. THE PATIENT DEVELOPED ISCHEMIA WITH A COLD MOTTLED FOOT. ADDITIONALLY, A PUMP STOP OCCURRED. THERE WERE LOW PURGE PRESSURES WITH HIGH MOTOR CURRENT PRIOR TO THE PUMP STOP. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE INVESTIGATION OF LIMB ISCHEMIA AND PUMP STOP HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. E4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26954 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. G1 REPORTING CONTACT FACILITY NAME AND FAX NUMBER WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26954. H6 COMPONENT CODE 3038 WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26954. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMART ASSIST SYSTEM STATES THE FOLLOWING: SECTION: TROUBLESHOOTING THE PURGE SYSTEM ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP.¿ SECTION: IMPELLA STOPPED ¿IF THE IMPELLA CATHETER HAS STOPPED SUDDENLY: 1. TRY TO RESTART THE CATHETER AT PREVIOUS P-LEVEL. 2. IF THE IMPELLA DOES NOT RESTART, TRY TO RESTART AT P-2. 3. IF THE IMPELLA DOES NOT RESTART OR STOPS AGAIN, WAIT 1 MINUTE AND TRY TO RESTART AGAIN. 4. IF THE IMPELLA RESTARTS, WEAN DOWN TO P-2 AS THE PATIENT CAN TOLERATE. UNDER THESE CIRCUMSTANCES, CATHETER FUNCTION IS NOT RELIABLE AND THE IMPELLA MAY STOP AGAIN. 5. IF THE IMPELLA DOES NOT RESTART, REMOVE THE IMPELLA FROM THE VENTRICLE AS SOON AS POSSIBLE TO AVOID AORTIC INSUFFICIENCY.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21693862&pc=CCK | 1314417-2025-00015 | 2024-08-24 04:00 | Death | BOUND TREE MEDICAL, LLC | 2025-03-25 04:00 | CCK | CURAPLEX | Display or Visual Feedback Problem; Output Problem | Respiratory Failure | K063813 |  | 1 | Event Description: ETCO2 PLUGGED INTO MONITOR, ATTACHED TO A PATIENT, MONITOR (LP15) NOTED TO SPONTANEOUS XXX ETCO2 MODULE ON SCREEN. ETCO2 REPLACED AND REATTACHED TO SAME MONITOR - SAME RESULT. ADDITIONAL MONITOR OBTAINED, SECOND ETCO2 PROBE ATTACHED AND ATTACHED TO ETT - MONITOR SUCCESSFULLY DISPLAYED ETCO2 WITHOUT ISSUE. PATIENT PRONOUNCED DEAD AFTER FURTHER RESUSCITATION EFFORTS. OTHER: (B)(6). TWO SEPERATE DEVICES WERE USED IN RELATION TO THIS EVENT, THE SECOND REPORT NUMBER IS 1314417-2025-00016.Event Description: ETCO2 PLUGGED INTO MONITOR, ATTACHED TO A PATIENT, MONITOR (LP15) NOTED TO SPONTANEOUS XXX ETCO2 MODULE ON SCREEN. ETCO2 REPLACED AND REATTACHED TO SAME MONITOR - SAME RESULT. ADDITIONAL MONITOR OBTAINED, SECOND ETCO2 PROBE ATTACHED AND ATTACHED TO ETT - MONITOR SUCCESSFULLY DISPLAYED ETCO2 WITHOUT ISSUE. PATIENT PRONOUNCED DEAD AFTER FURTHER RESUSCITATION EFFORTS. OTHER: (B)(6). TWO SEPARATE DEVICES WERE USED IN RELATION TO THIS EVENT, THE SECOND REPORT NUMBER IS 1314417-2025-00016. Manufacturer Narrative: THE PRODUCT INVOLVED IN THE REPORT HAS NOT BEEN RETURNED. A REVIEW OF THE DEVICE HISTORY RECORD IS IN-PROGRESS. ALL INFORMATION REASONABLY KNOWN AS OF 25 MAR 2025 HAS BEEN INCLUDED IN THE HEALTH AUTHORITY REPORT. SHOULD ADDITIONAL INFORMATION BE OBTAINED A FOLLOW-UP HEALTH AUTHORITY REPORT WILL BE PROVIDED. THE INFORMATION PROVIDED BY AIRLIFE. REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO AIRLIFE. AIRLIFE HAS NO INDEPENDENT KNOWLEDGE OF THE EVENT REPORTED BUT IS RELAYING THE INFORMATION THAT WAS PROVIDED BY THE USER FACILITY WHERE THE INCIDENT OCCURRED. THIS PRODUCT INCIDENT IS DOCUMENTED IN THE AIRLIFE COMPLAINT DATABASE AND IS IDENTIFIED AS COMPLAINT - (B)(4). THIS INFORMATION IS SUBMITTED PURSUANT TO 21CFR803, IN COMPLIANCE WITH THE MEDICAL DEVICE REPORTING REQUIREMENT AND SHOULD NOT BE CONSIDERED TO BE AN ADMISSION THAT A AIRLIFE PRODUCT IS DEFECTIVE OR CAUSES SERIOUS INJURY. Manufacturer Narrative: THE PRODUCT INVOLVED IN THE REPORT HAS NOT BEEN RETURNED. A REVIEW OF THE DEVICE HISTORY RECORD IS IN-PROGRESS. ALL INFORMATION REASONABLY KNOWN AS OF 25 MAR 2025 HAS BEEN INCLUDED IN THE HEALTH AUTHORITY REPORT. SHOULD ADDITIONAL INFORMATION BE OBTAINED A FOLLOW-UP HEALTH AUTHORITY REPORT WILL BE PROVIDED. THE INFORMATION PROVIDED BY AIRLIFE. REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO AIRLIFE. AIRLIFE HAS NO INDEPENDENT KNOWLEDGE OF THE EVENT REPORTED BUT IS RELAYING THE INFORMATION THAT WAS PROVIDED BY THE USER FACILITY WHERE THE INCIDENT OCCURRED. THIS PRODUCT INCIDENT IS DOCUMENTED IN THE AIRLIFE COMPLAINT DATABASE AND IS IDENTIFIED AS COMPLAINT (B)(4). THIS INFORMATION IS SUBMITTED PURSUANT TO 21CFR803, IN COMPLIANCE WITH THE MEDICAL DEVICE REPORTING REQUIREMENT AND SHOULD NOT BE CONSIDERED TO BE AN ADMISSION THAT A AIRLIFE PRODUCT IS DEFECTIVE OR CAUSES SERIOUS INJURY. THE COMPLAINT OF "SPONTANEOUS XXX ETCO2 MODULE ON SCREEN" REGARDING PARTS 301-5107ET AND 301-5707F-SE WAS NOT CONFIRMED. THE ROOT CAUSE WAS NOT DETERMINED BUT COULD POSSIBLY BE A RESULT OF USER ERROR. A RISK ASSESSMENT WAS PERFORMED AND THE ULTIMATE RISK WAS DETERMINED TO BE LOW WHICH DOES NOT REQUIRE ESCALATION TO THE CAPA REVIEW BOARD. THERE HAVE BEEN 9 OTHER COMPLAINTS REGARDING THE SAME PART AND A SIMILAR ISSUE WITHIN THE 24 MONTHS PRECEDING THIS REPORTED EVENT. A RESOLUTION LETTER WAS SENT TO THE CUSTOMER. COMPLAINTS WILL CONTINUE TO BE MONITORED FOR POSSIBLE TRENDS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21693863&pc=CCK | 1314417-2025-00016 | 2024-08-24 04:00 | Death | BOUND TREE MEDICAL, LLC | 2025-03-25 04:00 | CCK | CURAPLEX | Display or Visual Feedback Problem; Output Problem | Respiratory Failure | K063813 |  | 1 | Event Description: ETCO2 PLUGGED INTO MONITOR, ATTACHED TO A PATIENT, MONITOR (LP15) NOTED TO SPONTANEOUS XXX ETCO2 MODULE ON SCREEN. ETCO2 REPLACED AND REATTACHED TO SAME MONITOR - SAME RESULT. ADDITIONAL MONITOR OBTAINED, SECOND ETCO2 PROBE ATTACHED AND ATTACHED TO ETT - MONITOR SUCCESSFULLY DISPLAYED ETCO2 WITHOUT ISSUE. PATIENT PRONOUNCED DEAD AFTER FURTHER RESUSCITATION EFFORTS. OTHER: (B)(6). TWO SEPERATE DEVICES WERE USED IN RELATION TO THIS EVENT, THE FIRST REPORT NUMBER IS 1314417-2025-00015.Event Description: ETCO2 PLUGGED INTO MONITOR, ATTACHED TO A PATIENT, MONITOR (LP15) NOTED TO SPONTANEOUS XXX ETCO2 MODULE ON SCREEN. ETCO2 REPLACED AND REATTACHED TO SAME MONITOR - SAME RESULT. ADDITIONAL MONITOR OBTAINED, SECOND ETCO2 PROBE ATTACHED AND ATTACHED TO ETT - MONITOR SUCCESSFULLY DISPLAYED ETCO2 WITHOUT ISSUE. PATIENT PRONOUNCED DEAD AFTER FURTHER RESUSCITATION EFFORTS. OTHER: (B)(6). TWO SEPARATE DEVICES WERE USED IN RELATION TO THIS EVENT, THE FIRST REPORT NUMBER IS 1314417-2025-00015. Manufacturer Narrative: THE PRODUCT INVOLVED IN THE REPORT HAS NOT BEEN RETURNED. A REVIEW OF THE DEVICE HISTORY RECORD IS IN-PROGRESS. ALL INFORMATION REASONABLY KNOWN AS OF 25 MAR 2025 HAS BEEN INCLUDED IN THE HEALTH AUTHORITY REPORT. SHOULD ADDITIONAL INFORMATION BE OBTAINED A FOLLOW-UP HEALTH AUTHORITY REPORT WILL BE PROVIDED. THE INFORMATION PROVIDED BY AIRLIFE. REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO AIRLIFE. AIRLIFE HAS NO INDEPENDENT KNOWLEDGE OF THE EVENT REPORTED BUT IS RELAYING THE INFORMATION THAT WAS PROVIDED BY THE USER FACILITY WHERE THE INCIDENT OCCURRED. THIS PRODUCT INCIDENT IS DOCUMENTED IN THE AIRLIFE COMPLAINT DATABASE AND IS IDENTIFIED AS COMPLAINT (B)(4). THIS INFORMATION IS SUBMITTED PURSUANT TO 21CFR803, IN COMPLIANCE WITH THE MEDICAL DEVICE REPORTING REQUIREMENT AND SHOULD NOT BE CONSIDERED TO BE AN ADMISSION THAT A AIRLIFE PRODUCT IS DEFECTIVE OR CAUSES SERIOUS INJURY. THE COMPLAINT OF "SPONTANEOUS XXX ETCO2 MODULE ON SCREEN" REGARDING PARTS 301-5107ET AND 301-5707F-SE WAS NOT CONFIRMED. THE ROOT CAUSE WAS NOT DETERMINED BUT COULD POSSIBLY BE A RESULT OF USER ERROR. A RISK ASSESSMENT WAS PERFORMED AND THE ULTIMATE RISK WAS DETERMINED TO BE LOW WHICH DOES NOT REQUIRE ESCALATION TO THE CAPA REVIEW BOARD. THERE HAVE BEEN 9 OTHER COMPLAINTS REGARDING THE SAME PART AND A SIMILAR ISSUE WITHIN THE 24 MONTHS PRECEDING THIS REPORTED EVENT. A RESOLUTION LETTER WAS SENT TO THE CUSTOMER. COMPLAINTS WILL CONTINUE TO BE MONITORED FOR POSSIBLE TRENDS. Manufacturer Narrative: THE PRODUCT INVOLVED IN THE REPORT HAS NOT BEEN RETURNED. A REVIEW OF THE DEVICE HISTORY RECORD IS IN-PROGRESS. ALL INFORMATION REASONABLY KNOWN AS OF 25 MAR 2025 HAS BEEN INCLUDED IN THE HEALTH AUTHORITY REPORT. SHOULD ADDITIONAL INFORMATION BE OBTAINED A FOLLOW-UP HEALTH AUTHORITY REPORT WILL BE PROVIDED. THE INFORMATION PROVIDED BY AIRLIFE. REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO AIRLIFE. AIRLIFE HAS NO INDEPENDENT KNOWLEDGE OF THE EVENT REPORTED BUT IS RELAYING THE INFORMATION THAT WAS PROVIDED BY THE USER FACILITY WHERE THE INCIDENT OCCURRED. THIS PRODUCT INCIDENT IS DOCUMENTED IN THE AIRLIFE COMPLAINT DATABASE AND IS IDENTIFIED AS COMPLAINT - (B)(4). THIS INFORMATION IS SUBMITTED PURSUANT TO 21CFR803, IN COMPLIANCE WITH THE MEDICAL DEVICE REPORTING REQUIREMENT AND SHOULD NOT BE CONSIDERED TO BE AN ADMISSION THAT A AIRLIFE PRODUCT IS DEFECTIVE OR CAUSES SERIOUS INJURY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21694039&pc=LWP | 2017865-2025-38526 | 2025-03-06 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-03-25 04:00 | LWP | ASSURITY MRI | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P140033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD DECEASED DUE TO AN UNKNOWN CAUSE. NO ADDITIONAL INFORMATION WAS REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21694095&pc=NVN | 2017865-2025-38527 | 2025-03-06 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-03-25 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD DECEASED DUE TO AN UNKNOWN CAUSE. NO ADDITIONAL INFORMATION WAS REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21694099&pc=NVN | 2017865-2025-38528 | 2025-03-06 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-03-25 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD DECEASED DUE TO AN UNKNOWN CAUSE. NO ADDITIONAL INFORMATION WAS REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21694266&pc=DYB | 2015691-2025-02258 | 2025-02-28 05:00 | Death | EDWARDS LIFESCIENCES | 2025-03-25 04:00 | DYB | EDWARDS ESHEATH+ INTRODUCER SET | Peeled/Delaminated; Material Puncture/Hole; Physical Resistance/Sticking | Great Vessel Perforation | K200258 |  | 1 | Event Description: AS REPORTED BY THE FIELD CLINICAL SPECIALIST (FCS), DURING A LEFT-SIDE TRANSFEMORAL TAVR (TRANSCATHETER AORTIC VALVE REPLACEMENT) PROCEDURE WITH A 29MM SAPIEN 3 ULTRA RESILIA (S3UR) VALVE IN THE NATIVE AORTIC POSITION, THE S3UR VALVE (LOADED ON THE 29MM COMMANDER DELIVERY SYSTEM (DS)) WAS UNABLE TO ADVANCE THROUGH THE 16FR ESHEATH+ AND BECAME STUCK AT THE AORTO-ILIAC BIFURCATION. A TEAR WAS OBSERVED IN THE ESHEATH+ SEAM AND THE ESHEATH+ HAD PERFORATED THE ABDOMINAL AORTA. IT WAS NOTED THAT THE ESHEATH+ WAS NOT SUTURED IN PLACE. THE ESHEATH+ MAY HAVE ALSO BEEN TWISTED AT SOME POINT DURING ADVANCEMENT. THE OPERATING TEAM PANNED DOWN WHEN RESISTANCE WAS FELT WITH THE ESHEATH+, BUT BY THEN THE S3UR VALVE HAD PUNCTURED THE ABDOMINAL AORTA. IN RESPONSE, THE S3UR VALVE AND DS WERE RETRACTED BACK INTO THE ESHEATH+. DURING NAVIGATION OF THE SYSTEM FOR REMOVAL OUT OF THE PATIENT, THE S3UR VALVE EXITED THE DISTAL TIP OF THE ESHEATH+ AND LODGED INTO THE PATIENT'S RIGHT FEMORAL ARTERY. IT IS HYPOTHESIZED THAT THE ESHEATH+ MAY HAVE BEGUN TO SPLIT IN THE RIGHT EXTERNAL ILIAC. A SURGICAL CUT-DOWN WAS PERFORMED IN ORDER TO REMOVE THE DEVICES. A CODA BALLOON WAS UTILIZED AND STENT GRAFTS WERE IMPLANTED TO STOP THE BLEEDING. CHEST COMPRESSIONS, EPINEPHRINE AND A BLOOD TRANSFUSION WERE GIVEN TO RESTORE THE PATIENT'S BLOOD PRESSURE (BP). ONCE THE PATIENT'S BP WAS STABILIZED, A SECOND 29MM SAPIEN 3 ULTRA RESILIA VALVE WAS PREPPED. PRIOR TO THE SECOND VALVE BEING ADVANCED THROUGH THE SECOND 16FR ESHEATH+ ON THE PATIENT'S LEFT SIDE, THE PATIENT BECAME UNSTABLE AGAIN. NEVERTHELESS, THE SECOND VALVE WAS ADVANCED AND SUCCESSFULLY DEPLOYED. HOWEVER, THE PATIENT WAS UNABLE TO RECOVER AND EXPIRED.Event Description: OBSERVATION OF THE RETURNED DEVICE PRIOR TO DECONTAMINATION REVEALED TWO LINER STRANDS NEAR THE DISTAL STRAIN RELIEF, THE FIRST APPROXIMATELY 9IN ADJACENT AND THE SECOND APPROXIMATELY 9.5IN ADJACENT; THE LENGTH OF EACH STRAND WAS 0.25IN AND 1IN, RESPECTIVELY. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED DUE TO PRE-DECONTAMINATION FINDINGS OF THE RETURNED DEVICE. SECTION B4, B5, D9, G3, G6, H2, H3, H6: DEVICE CODE(S) AND TYPE OF INVESTIGATION HAVE BEEN UPDATED. ADDITIONS TO SECTIONS B5, D9, H6: DEVICE CODE(S) AND TYPE OF INVESTIGATION. CORRECTION TO SECTION H3. THE INVESTIGATION IS ONGOING. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED DUE TO ENGINEERING EVALUATION FINDINGS. SECTIONS B4, G3, G6, H2, H6: COMPONENT CODE, TYPE OF INVESTIGATION, INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAVE BEEN UPDATED. ADDITIONS TO SECTIONS H6: COMPONENT CODE AND TYPE OF INVESTIGATION. CORRECTIONS TO SECTIONS H6: INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS. THE EVENT(S) REPORTED IS/ARE ANTICIPATED IN THE RISK MANAGEMENT DOCUMENTATION FOR TRANSCATHETER HEART VALVE PROCEDURES. A PREVIOUS INVESTIGATION INTO THIS TYPE OF EVENT IS CAPTURED IN AN EDWARDS LIFESCIENCES TECHNICAL SUMMARY AND APPLIES TO THIS COMPLAINT. ADDITIONAL ASSESSMENT OF THE FAILURE MODE(S) IS NOT REQUIRED AT THIS TIME. THE ESHEATH+ DEVICE WAS RETURNED FOR EVALUATION, AND THE FOLLOWING OBSERVATIONS WERE MADE: SHEATH SHAFT CURVATURE, LINER LIFTED 3.25IN FROM DISTAL STRAIN RELIEF WITH REMAINING LINER UNEXPANDED, LINER PUNCTURE ON THE FULL LENGTH OF THE SHEATH SHAFT, FIRST LINER STRAND STARTING AT 9IN FROM DISTAL STRAIN RELIEF (0.25IN LONG), SECOND LINER STRAND STARTING AT 9.5IN FROM DISTAL STRAIN RELIEF (1IN LONG), DISTAL TIP OPENED AS DESIGNED, C-MARKER PRESENT, HDPE FOLD STARTING AT 8.75IN FROM DISTAL STRAIN RELIEF (1.5IN LONG), AND HDPE DAMAGE ALONG THE LINER PUNCTURE. PER THE TECHNICAL SUMMARY, THE INSTRUCTIONS FOR USE (IFU), TRAINING MANUALS AND CURRENT RISK MITIGATIONS HAVE BEEN REVIEWED. NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. IMAGING WAS PROVIDED AND REVEALED TORTUOSITY AND CALCIFICATION IN THE PATIENT'S LEFT ACCESS VESSEL. IN THIS CASE, THE COMPLAINTS OF INABILITY TO ADVANCE THROUGH SHEATH, LINER PUNCTURED AND SHEATH LINER STRAND ARE CONFIRMED BASED ON THE RETURNED DEVICE EVALUATION AND PROVIDED IMAGERY. THE DEVICE EVALUATION REVEALED A LINER PUNCTURE AND LINER STRANDS. PER THE IFU, ''INSERT THE SHEATH WITH THE EDWARDS LOGO FACING UPWARD SO THE SEAM REMAINS DOWN TO ENSURE PROPER SHEATH PERFORMANCE'' AND ''SUTURE SHEATH IN PLACE''. PER THE TRAINING MANUAL, ''FOLLOWING PROPER VALVE CRIMPING TECHNIQUE AND ENSURE VALVE IS DELIVERED AS STRAIGHT AS POSSIBLE''. IF THE SHEATH WAS NOT SUTURED, THIS MAY LEAD TO INSTABILITY DURING DELIVERY SYSTEM ADVANCEMENT AND RESULT IN NON-COAXIAL ALIGNMENT BETWEEN THE DEVICES. ADDITIONALLY, IF THE SHEATH WAS TORQUED DURING PROCEDURE, THIS MAY FURTHER CONTRIBUTE TO NON-COAXIAL ADVANCEMENT. PER IMAGERY REVIEW, TORTUOSITY WAS OBSERVED IN THE PATIENT'S LEFT ACCESS VESSEL. TORTUOSITY CAN PROMOTE SUB-OPTIMAL ANGLES THAT CAN CONTRIBUTE TO NON-COAXIAL ALIGNMENT DURING DELIVERY SYSTEM ADVANCEMENT. NON-COAXIAL ADVANCEMENT MAY CAUSE RESISTANCE THROUGH INCREASED INTERACTIONS BETWEEN THE SYSTEM AND SHEATH. HIGH PUSH FORCE USED TO OVERCOME ANY RESISTANCE WHEN NAVIGATING CHALLENGING ANATOMY COUPLED WITH NON-COAXIAL ADVANCEMENT TRAJECTORIES CAN LEAD TO LINER PUNCTURE AND LINER STRAND DUE TO DIRECT INTERACTION WITH CRIMPED VALVE (E.G. CATCHING OF BENT STRUT). AS SUCH, AVAILABLE INFORMATION SUGGESTS THAT USE ERROR (SHEATH NOT SUTURED), PATIENT FACTORS (TORTUOSITY) AND/OR PROCEDURAL FACTORS (DEVICE MANIPULATION) MAY HAVE CONTRIBUTED TO THE COMPLAINT EVENTS. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. SINCE NO EDWARDS PRODUCT DEFECTS OR LABELING DEFICIENCIES WERE IDENTIFIED, NO CORRECTIVE OR PREVENTATIVE ACTION IS REQUIRED. Manufacturer Narrative: THE INVESTIGATION IS ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21694285&pc=NPT | 2015691-2025-02259 | 2025-02-28 05:00 | Death | EDWARDS LIFESCIENCES | 2025-03-25 04:00 | NPT | EDWARDS COMMANDER DELIVERY SYSTEM | Device Dislodged or Dislocated; Insufficient Information | Great Vessel Perforation | P140031 |  | 1 | Event Description: AS REPORTED BY THE FIELD CLINICAL SPECIALIST (FCS), DURING A LEFT-SIDE TRANSFEMORAL TAVR (TRANSCATHETER AORTIC VALVE REPLACEMENT) PROCEDURE WITH A 29MM SAPIEN 3 ULTRA RESILIA (S3UR) VALVE IN THE NATIVE AORTIC POSITION, THE S3UR VALVE (LOADED ON THE 29MM COMMANDER DELIVERY SYSTEM (DS)) WAS UNABLE TO ADVANCE THROUGH THE 16FR ESHEATH+, AND BECAME STUCK AT THE AORTO-ILIAC BIFURCATION. A TEAR WAS OBSERVED IN THE ESHEATH+ SEAM AND THE ESHEATH+ HAD PERFORATED THE ABDOMINAL AORTA. IT WAS NOTED THAT THE ESHEATH+ WAS NOT SUTURED IN PLACE. THE ESHEATH+ MAY HAVE ALSO BEEN TWISTED AT SOME POINT DURING ADVANCEMENT. THE OPERATING TEAM PANNED-DOWN WHEN RESISTANCE WAS FELT WITH THE ESHEATH+, BUT BY THEN THE S3UR VALVE HAD PUNCTURED THE ABDOMINAL AORTA. IN RESPONSE, THE S3UR VALVE AND DS WERE RETRACTED BACK INTO THE ESHEATH+. DURING NAVIGATION OF THE SYSTEM FOR REMOVAL OUT OF THE PATIENT, THE S3UR VALVE EXITED THE DISTAL TIP OF THE ESHEATH+ AND LODGED INTO THE PATIENT'S RIGHT FEMORAL ARTERY. IT IS HYPOTHESIZED THAT THE ESHEATH+ MAY HAVE BEGUN TO SPLIT IN THE RIGHT EXTERNAL ILIAC. A SURGICAL CUT-DOWN WAS PERFORMED IN ORDER TO REMOVE THE DEVICES. A CODA BALLOON WAS UTILIZED AND STENT GRAFTS WERE IMPLANTED TO STOP THE BLEEDING. CHEST COMPRESSIONS, EPINEPHRINE AND A BLOOD TRANSFUSION WERE GIVEN TO RESTORE THE PATIENT'S BLOOD PRESSURE (BP). ONCE THE PATIENT'S BP WAS STABILIZED, A SECOND 29MM SAPIEN 3 ULTRA RESILIA VALVE WAS PREPPED. PRIOR TO THE SECOND VALVE BEING ADVANCED THROUGH THE SECOND 16FR ESHEATH+ ON THE PATIENT'S LEFT SIDE, THE PATIENT BECAME UNSTABLE AGAIN. NEVERTHELESS, THE SECOND VALVE WAS ADVANCED AND SUCCESSFULLY DEPLOYED. HOWEVER, THE PATIENT WAS UNABLE TO RECOVER AND EXPIRED. Manufacturer Narrative: THE INVESTIGATION IS ONGOING. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED DUE TO ADMINISTRATIVE REVIEW. SECTIONS B1, B4, G3, G6, H2, H6: HEALTH EFFECT - IMPACT CODE AND DEVICE CODE(S) HAVE BEEN UPDATED. ADDITION TO SECTION H6: HEALTH EFFECT - IMPACT CODE. CORRECTIONS TO SECTIONS B1 AND H6: DEVICE CODE(S). THE INVESTIGATION IS ONGOING. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED DUE TO ENGINEERING EVALUATION FINDINGS. SECTIONS B4, D9, G3, G6, H2, H3, H6: TYPE OF INVESTIGATION, INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAVE BEEN UPDATED. ADDITIONS TO SECTION H6: TYPE OF INVESTIGATION. CORRECTIONS TO SECTIONS H6: INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS. PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. ACCORDING TO THE DEVICE TRAINING MANUALS, RISK FACTORS FOR AORTIC DISSECTION, CARDIAC/AORTIC HEMATOMA, OR ANNULAR RUPTURE DURING THE TVR PROCEDURE INCLUDE SIGNIFICANT VALVE OVER SIZING, SEVERELY OBLITERATED SINUSES OF VALSALVA, PORCELAIN AORTA AND/OR PRESENCE OF BULKY CALCIFICATION, AND NARROW CALCIFIED SINOTUBULAR JUNCTION. IN ADDITION, ADVANCED AGE, FEMALE GENDER, SMALL BODY WEIGHT, AND STEROID DEPENDENCY CAN ALSO BE CONTRIBUTING FACTORS. THE SAPIEN TRANSCATHETER HEART VALVE (ALL MODELS) RELIES ON CALCIUM TO SECURELY ANCHOR IN THE LANDING ZONE. DESPITE THIS BENEFICIAL ASPECT OF CALCIUM, BULKY CALCIUM CAN INCREASE THE RISK OF CALCIFIC NODULE DISPLACEMENT INTO THE VASCULATURE, WHICH CAN LEAD TO VASCULAR INJURY. AT TIMES, THE EXTENT AND DISTRIBUTION OF CALCIUM CAN IMPAIR EASE OF DELIVERY OF THE VALVE, CORRECT POSITIONING OF THE VALVE, DEPLOYMENT OF THE VALVE, AND PROCEDURAL SUCCESS. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THE PERFORATION OF THE PATIENT'S ABDOMINAL AORTA. INVESTIGATION RESULTS ARE INCONCLUSIVE DUE TO THE LIMITED INFORMATION PROVIDED; HOWEVER, PROCEDURAL FACTORS LIKELY CONTRIBUTED TO THE ABDOMINAL AORTIC PERFORATION DUE TO TWISTING OF THE 16FR ESHEATH+ DURING ADVANCEMENT AND THE ESHEATH+ NOT BEING SUTURED IN PLACE PRIOR TO SYSTEM ADVANCEMENT. PATIENT FACTORS (ADVANCED AGE OF 81 YEARS AND TORTUOUS ANATOMY) MAY HAVE ALSO CONTRIBUTED TO THE AORTIC PERFORATION. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE EVENT(S) REPORTED IS/ARE ANTICIPATED IN THE RISK MANAGEMENT DOCUMENTATION FOR TRANSCATHETER HEART VALVE PROCEDURES. ADDITIONAL ASSESSMENT OF THE FAILURE MODE(S) IS NOT REQUIRED AT THIS TIME. THE DELIVERY SYSTEM DEVICE WAS RETURNED FOR EVALUATION, AND THE FOLLOWING OBSERVATIONS WERE MADE: SLIGHT CURVATURE ON THE FLEX SHAFT WITH NO OTHER ABNORMALITIES, AND NO TRANSCATHETER HEART VALVE (THV) RETURNED. THE INSTRUCTIONS FOR USE (IFU), TRAINING MANUALS AND CURRENT RISK MITIGATIONS HAVE BEEN REVIEWED. NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. IMAGING WAS PROVIDED AND REVEALED THE RIGHT ACCESS VESSEL WAS ADEQUATELY SIZED (GREATER THAN OR EQUAL TO 6.0MM) TO ACCOMMODATE A 16F ESHEATH+, AND THERE WAS MINIMAL TORTUOSITY WITH LITTLE TO NO CALCIFICATION PRESENT THE ACCESS VESSEL. IN THIS CASE, THE COMPLAINT OF THV DISLODGED OFF BALLOON WAS CONFIRMED BASED ON THE PROVIDED PROCEDURAL IMAGERY AND EVALUATION OF THE RETURNED DEVICES. THE ASSOCIATED ESHEATH+ HAD A DAMAGED LINER AND THE VALVE WAS PUNCTURING OUT OF THE ESHEATH+, INTERACTING WITH THE VASCULATURE. IT IS LIKELY THAT HIGH PULL FORCE WAS APPLIED IN AN ATTEMPT TO WITHDRAW THE VALVE, WHICH WAS STUCK IN THE VASCULATURE. THIS CONDITION OF HIGH PULL FORCE WHILE THE VALVE WAS INTERACTING WITH THE AORTA MAY HAVE CONTRIBUTED TO THE VALVE DISLODGING OFF OF THE DELIVERY SYSTEM AND REMAINING IN THE PATIENT VASCULATURE AS THE DELIVERY SYSTEM WAS BEING WITHDRAWN THROUGH THE ESHEATH+. AS SUCH, AVAILABLE INFORMATION SUGGESTS THAT PROCEDURAL FACTORS (VALVE INTERACTED WITH VASCULATURE AND HIGH PULL FORCE) MAY HAVE CONTRIBUTED TO THE REPORTED DISLODGMENT. HOWEVER, A DEFINITIVE ROOT CAUSE IS UNABLE TO BE DETERMINED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21675109&pc=OZP | 2032227-2025-141339 | 2024-10-21 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-03-24 04:00 | OZP | 670G INSULIN PUMP MMT-1780KL | Adverse Event Without Identified Device or Use Problem; Appropriate Term/Code Not Available | Hyperglycemia; Hypoglycemia | P150001 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2024. THE LAST KNOWN PRODUCT FOR THIS CUSTOMER ARE AS FOLLOWS: MMT-1780KL, MMT-332A, UNOMEDICAL. TROUBLESHOOTING WAS PERFORMED. THE CUSTOMER WAS NOT HOSPITALIZED. THE CAUSE OF DEATH WAS DUE TO DIABETES COMPLICATIONS. THE PUMP WAS WORN AT THE TIME OF PASSING. THE CUSTOMER WAS NOT USING SENSOR. NO PRODUCT RETURN IS REQUIRED FOR MMT-1780KL, MMT-332A AND UNOMEDICAL. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21675403&pc=BZD | 2518422-2025-038171 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-03-24 04:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer; Unspecified Respiratory Problem; Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, RESPIRATORY TRACT IRRITATION, LUNG DISEASE CANCER, AND OTHER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21675816&pc=BZD | 2518422-2025-103316 | 2025-02-10 05:00 | Death | RESPIRONICS, INC. | 2025-03-24 04:00 | BZD | REP DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS INFORMED THAT A PATIENT USING A REP DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THE PATIENT'S HUSBAND INQUIRED ABOUT HOW TO DISPOSE OF THE DEVICE. HE WAS ADVISED TO CONTACT THE MEDICAL COMPANY THAT PROVIDED IT, AS THEY COULD ASSIST IN PROPERLY DISPOSING OF THE DEVICE. WHEN THE HUSBAND ASKED IF HE COULD SIMPLY THROW IT AWAY, HE WAS RECOMMENDED TO REACH OUT TO THE DOCTOR TO OBTAIN THE NAME OF THE DME PROVIDER, WHO COULD HELP WITH THE DISPOSAL PROCESS. THE HUSBAND THEN EXPRESSED CONCERNS ABOUT THE TIME INVOLVED IN DRIVING AROUND AND MENTIONED THAT HE MIGHT JUST DISPOSE OF THE DEVICE HIMSELF. NO ADDITIONAL DETAILS WERE PROVIDED, AND THERE ARE NO ALLEGATIONS SUGGESTING THAT THE DEVICE CONTRIBUTED TO OR CAUSED THE PATIENT'S DEATH. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21675862&pc=OZD | 1220648-2025-26903 | 2025-02-24 05:00 | Death | ABIOMED, INC. | 2025-03-24 04:00 | OZD | IMPELLA CP | Increase in Pressure | Ischemia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WITH ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT AND EXPERIENCED HIGH PURGE PRESSURE FOLLOWED BY LIMB ISCHEMIA WITH A DEMISE OUTCOME. THE PATIENT UNDERWENT ELECTIVE TRANSCATHETER AORTIC VALVE REPLACEMENT. WHILE STILL IN THE PROCEDURE ROOM WHILE ATTEMPTING TO EXTUBATE, THE PATIENT STARTED HAVING RHYTHM CHANGES AND DECREASING SATS EVENTUALLY GOING INTO VENTRICULAR FIBRILLATION AND VENTRICULAR TACHYCARDIA REQUIRING NUMEROUS DEFIBRILLATIONS. CORONARY ANGIOGRAM SHOWED CORONARIES WERE NOT COMPROMISED. TRANSESOPHAGEAL ECHOCARDIOGRAM SHOWED A SEVERELY DILATED RIGHT VENTRICLE. THEREFORE, THE DECISION WAS MADE TO PLACE RIGHT-SIDED SUPPORT AS WELL. THE PATIENT WAS PLACED ON AN IMPELLA CP AND IMPELLA RP FLEX DEVICES (BIPELLA SUPPORT) WITH A "VERY" POOR PROGNOSIS. APPROXIMATELY FOUR HOURS INTO SUPPORT, THE IMPELLA CP DEVICE EXPERIENCED HIGH PURGE PRESSURE. TISSUE PLASMINOGEN ACTIVATOR WAS INITIATED IN RESPONSE TO THE PURGE PRESSURE HIGH ALARM. IT WAS ADDITIONALLY NOTED THAT THE PATIENT HAD NO DISTAL FLOW IN THEIR LOWER LEFT EXTREMITY (IMPELLA CP LEG). THE PERSISTENT CONDITION WAS MONITORED BY THE STAFF. AFTER TWO DAYS OF SUPPORT, THE PATIENT WAS NO LONGER RESPONDING TO THE MAXIMUM DOSES OF VASOPRESSORS AND INOTROPES. THE PATIENT PASSED AWAY ON SUPPORT. MEDICAL SAFETY REVIEW OF THE EVENT NOTED THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA CP DEVICE AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME (DEMISE). Manufacturer Narrative: THE INVESTIGATION OF HIGH PURGE PRESSURE AND LIMB ISCHEMIA HAS BEEN COMPLETED. THE IMPELLA CP WAS RETURNED AND EVALUATED. NO KINKS WERE FOUND ON THE PUMP CATHETER OR PURGE CASSETTE TUBING. BIOMATERIAL WAS FOUND IN THE PURGE GAP. HIGH PURGE PRESSURE WAS REPRODUCED WHILE PURGING THE PUMP IN THE LAB. NO BLOOD INGRESS INTO THE CATHETER PURGE LINE WAS OBSERVED. THE HIGH PURGE PRESSURE RESOLVED AFTER CUTTING THE CATHETER JUST ABOVE THE MOTOR HOUSING. DATA LOGS SHOWED THAT INITIAL PURGE PRESSURE WAS 700-800 MMHG AND PURGE FLOW WAS BELOW FOUR MILLILITERS PER HOUR. ABOUT FOUR HOURS INTO THE CASE, THE PURGE PRESSURE ROSE FROM ~700 TO GREATER THAN 1000 MMHG WITHIN ABOUT 30 MINUTES. THE HIGH PURGE PRESSURE DID NOT RESOLVE BY THE END OF THE CASE, WHICH WAS 43 HOURS TOTAL. NO SPIKES IN MOTOR CURRENT WERE SEEN, BUT A SLIGHT INCREASE THROUGHOUT THE CASE WAS SEEN. THE ROOT CAUSE OF THE HIGH PURGE PRESSURE WAS MOST LIKELY BIOMATERIAL BUILDUP SINCE THE BIOMATERIAL WAS FOUND ON THE PURGE GAP OF THE RETURNED PUMP, HIGH PURGE PRESSURE WAS REPRODUCED, AND DATA LOGS SHOWED A GRADUAL RISE IN PURGE PRESSURE. AS ALL THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE LIMB ISCHEMIA WAS NOT DETERMINED. Manufacturer Narrative: THE IMPELLA PUMP, PURGE CASSETTE AND DATA STICK WERE RETURNED, AND AN EVALUATION IS UNDERWAY. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21676049&pc=PNJ | 9612164-2025-01478 | 2025-02-25 05:00 | Death | MEDTRONIC IRELAND | 2025-03-24 04:00 | PNJ | MICRA¿ AV2 | Use of Device Problem; Positioning Problem; High Capture Threshold | Anemia; Cardiac Arrest; Shock; Cardiac Tamponade | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT INTRA OPERATIVELY DURING THE PLACEMENT OF THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG), THERE WAS SOME RESISTANCE IN PASSING THE TRICUSPID VALVE. WHEN THE PLACEMENT POSITION WAS REACHED AND ATTEMPTED TO BE PLACED, HIGH THRESHOLDS WERE NOTED. ON THE THIRD ATTEMPT, SUITABLE THRESHOLDS WERE OBTAINED WITH RELATIVELY LOW IMPEDANCE. THE DEVICE WAS EVENTUALLY PLACED. POST OPERATIVELY, THE PATIENT WAS ANEMIC AND WERE IN A STATE OF SHOCK. CARDIAC TAMPONADE AND RETROPERITONEAL HEMORRHAGE WERE NOTED. LATER, IT WAS REPORTED THAT THE PATIENT PASSED AWAY THE NEXT DAY DUE TO HEMORRHAGIC SHOCK.Event Description: IT WAS FURTHER REPORTED THAT DURING THE IMPLANT PROCEDURE OF THE LEADLESS IPG, THE CATHETER SLIPPED/DISLODGEMENT TO THE HINGE SIDE WHEN MAKING A GOOSENECK. IT WAS ALSO REPORTED THAT POST DEVICE PLACEMENT, PATIENT'S ANEMIA PROGRESSED LEADING TO CARDIOPULMONARY ARREST AND THE PATIENT WAS MOVED TO INTENSIVE CARE UNIT (ICU). IT WAS REPORTED THAT THE BLEEDING WAS DUE TO THE IMPACT OF CARDIOPULMONARY RESUSCITATION. Manufacturer Narrative: SECTION D REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [MICRA] PRODUCT ID [MC2AVR1-DELSYS] (SERIAL: [(B)(6)]). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21676857&pc=NPT | 2025587-2025-02054 | 2024-01-31 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-24 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Dysphagia/ Odynophagia; Muscle Weakness; Pleural Effusion; Respiratory Insufficiency; Swelling/ Edema | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT HAD DIFFICULTY AWAKING AFTER THE PROCEDURE. PROLONGED VENTILATOR SUPPORT WAS REQUIRED, A TRACHEOSTOMY WAS PERFORMED, AND RESPIRATORY SUPPORT WAS PROVIDED USING CONTINUOUS POSITIVE AIRWAY PRESSURE MODE. IT WAS NOTED THAT THE PRE-EXISTING AORTIC STENOSIS WAS RELIEVED AND THE PATIENT HAD GOOD CARDIAC FUNCTION; HOWEVER, THE PATIENT HAD DIFFICULTY SWALLOWING AND REQUIRED TUBE FEEDING WITH TOTAL PARENTERAL NUTRITION. DISUSE AND MALN UTRITION CONTRIBUTED TO THE DEVELOPMENT OF PLURAL EFFUSION AND PERIPHERAL EDEMA, ACCORDING TO THE PHYSICIAN. OVER TIME, THE PATIENT GRADUALLY WEAKENED. THE PATIENT PASSED AWAY DUE TO HEART-RELATED CAUSES APPROXIMATELY 4 MONTHS AFTER THE PROCEDURE. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21676884&pc=DWF | 1317056-2025-00077 | 2025-02-26 05:00 | Death | ANGIODYNAMICS, INC | 2025-03-24 04:00 | DWF | ANGIOVAC | Adverse Event Without Identified Device or Use Problem | Pulmonary Embolism | K190594 |  | 1 | Event Description: AN ANGIODYNAMICS TERRITORY MANAGER REPORTED A PATIENT DEATH THAT OCCURRED DURING USE OF AN ANGIOVAC C180 WITH OBTURATOR PG AND CIRCUIT WITH BUBBLE TRAP. THE FOLLOWING WAS REPORTED: "PATIENT HAD HISTORY OF CANCER 2018 (CERVICAL CANCER), ATRIAL FIBRILLATION, COPD, CTEPH, AND DVT. PATIENT HAD BLOOD IN URINE IN (B)(6) AND WAS TAKEN OFF XARELTO. INSTRUCTED TO BEGIN XARELTO AGAIN AFTER 5 DAYS. PATIENT FAILED TO BEGIN TAKING XARELTO AGAIN AS INSTRUCTED. YESTERDAY PATIENT PRESENTED AT ED AT (B)(6) WITH SUSPECTED PULMONARY EMBOLISM. PATIENT WAS TRANSFERRED TO DEACONESS GATEWAY FOR TREATMENT. CT SCAN SHOWED SMALL SADDLE EMBOLI AS WELL AS EMBOLI IN LEFT PULMONARY ARTERY THAT WASN'T DEEMED CLINICALLY SIGNIFICANT PER DR. (B)(6) . TEE SHOWED LARGE RA MASS (GREATER THAN 4CM). MASS APPEARED TO BE BROAD BASED, CONCENTRIC, AND LACKING JAGGED EDGES. SMALLER CLOT (2CM) APPEARED TO BE PRESENT AT THE IVC/RA JUNCTION. PATIENT HAD BEEN INTUBATED DURING THE TEE AND HAD REMAINED INTUBATED SINCE. DR. (B)(6), CT SURGEON, SUSPECTED CARDIAC TUMOR BASED ON CONVERSATIONS IN THE CONTROL ROOM AND HE MENTIONED THE RISK OF EMBOLIZATION AND MASSIVE PE TO DR. (B)(6). PRIOR TO CASE BEGINNING, DR.(B)(6) MET WITH ADAM DAVIS AND I TO DISCUSS SPECIFICS OF THE CASE. WE DISCUSSED RIJ CANNULA/LFV RETURN AND RFV CANNULA/LFV RETURN WITH DR. (B)(6); HE CHOSE RFV CANNULA/LFV RETURN APPROACH FOR THIS CASE. DEBULKING OF THE RIGHT ATRIAL CLOT/MASS WAS THE GOAL OF THE CASE. PATIENT WAS ALREADY INTUBATED AND RIGHT ATRIAL MASS WAS CONFIRMED TO BE PRESENT PRIOR TO CASE. WE GAINED ACCESS USING A 26F GORE DRY SEAL ON THE RIGHT AND A 17F MEDTRONIC ARTERIAL RETURN CATHETER ON THE LEFT. PREP OF THE ANGIOVAC DEVICE WAS WITHOUT EVENT AND PER ANGIODYNAMICS IFU. DR. (B)(6) GAINED ACCESS WITH AN AMPLATZ SUPER STIFF WIRE; WIRE WAS IN PLACE AND ANGIOVAC CANNULA AND OBTURATOR WERE LOADED ONTO THE WIRE AND PLACED IN THE RIGHT ATRIUM. ONCE DR. (B)(6) WAS HAPPY WITH THE POSITION OF THE ANGIOVAC CANNULA, THE WIRE AND OBTURATOR WERE REMOVED. FUNNEL WAS THEN EXPOSED. AT THIS POINT DR. (B)(6) ASKED THE PERFUSION TEAM TO COME ON SUCTION. THEY DID AND FLOW VARIED BETWEEN 2.5 TO 5L/MIN. ON SEVERAL OCCASIONS THE FUNNEL ENGAGED MATERIAL/ATRIAL WALL AND FLOW WAS COMPLETELY BLOCKED. THE CANNULA WOULD BE REPOSITIONED EACH TIME AND FLOW WAS RESTORED. THE RIGHT ATRIAL MASS WAS ENGAGED AND A LARGE AMOUNT OF MATERIAL WAS REMOVED (BRIGHT WHITE MATERIAL) AND LODGED IN THE BUBBLE TRAP. IT WAS NOTICED ON TEE THAT THE MAJORITY OF THE MATERIAL WAS NO LONGER IN THE RA, WITH ONLY A SINGLE PIECE OF MATERIAL REMAINING. THE REMAINING MATERIAL WAS ROUGHLY 3CM LONG AND EXTREMELY PEDUNCULATED AND STALK LIKE. AT THIS POINT A LONG PAUSE FOLLOWED IN CARDIAC FUNCTION, FOLLOWED BY CARDIAC ARREST. HE CALLED TO INITIATE CPR AND RESUSCITATION EFFORTS BEGAN IN THE FORM OF CHEST COMPRESSIONS . THE ANGIOVAC DEVICE AND MEDTRONIC ARTERIAL RETURN SHEATH WERE BOTH REMOVED. CPR WAS MAINTAINED FOR 14 MINUTES WITH BRIEF PAUSES TO CHECK FOR RETURN OF FUNCTION. DR. (B)(6) BELIEVED THAT MATERIAL HAD EMBOLIZED FROM THE RA TO THE RIGHT PULMONARY ARTERY. HE CALLED FOR THE INARI REPRESENTATIVE. A PIGTAIL CATHETER WAS PLACED IN THE MAIN PA AND A FILLING DEFECT WAS NOTICED IN THE LEFT PULMONARY ARTERIES. DR.(B)(6) PLACED A FLOWTRIEVER IN THE LEFT PULMONARY ARTERY AND MADE 7 TO 8 PULLS IN AN ATTEMPT TO TRY AND EVACUATE THE MATERIAL. HE DID NOT ASPIRATE ANY MATERIAL FROM THE PULMONARY ARTERIES AND THE DECISION WAS MADE AFTER 27 MINUTES TO GO SPEAK WITH THE FAMILY AND CEASE RESUSCITATION EFFORTS." DR. (B)(6) IS A VERY EXPERIENCED ANGIOVAC USER AND FELT THIS WAS THE BEST OPTION TO ATTEMPT TO TREAT A VERY, CHRONICALLY ILL PATIENT AND DOES NOT BELIEVE THERE WAS ANY SORT OF DEVICE MALFUNCTION THAT CONTRIBUTED TO THIS EVENT. THIS MEDICAL OPINION WAS DISCUSSED AND AGREED UPON WITH SEVERAL PHYSICIANS WITHIN THIS ACCOUNT. Manufacturer Narrative: THE DEVICE IS NOT AVAILABLE TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AN INVESTIGATION INTO THE ROOT CAUSE FOR THE EVENT IS CURRENTLY IN PROGRESS. THE RESULTS OF THE INVESTIGATION WILL BE SENT VIA A FOLLOW UP MEDWATCH. REFERENCE (B)(4). Manufacturer Narrative: THE CUSTOMER'S REPORTED COMPLAINT DESCRIPTION OF PATIENT THROMBOSIS (PULMONARY EMBOLISM - PE) THAT RESULTED IN CARDIAC ARREST CANNOT BE CONFIRMED GIVEN THE PATIENT CENTRIC NATURE OF THE PATIENT SERIOUS ADVERSE EVENT (SAE). NO ANGIOVAC CANNULA/CIRCUIT WAS RETURNED FOR EVALUATION AS THERE WAS NO REPORTED DEVICE MALFUNCTION OR PERFORMANCE ISSUE DURING THE PROCEDURE. AS NOTED IN THE EVENT DESCRIPTION: "THE PHYSICIAN IS A VERY EXPERIENCED ANGIOVAC USER AND FELT THIS WAS THE BEST OPTION TO ATTEMPT TO TREAT A VERY, CHRONICALLY ILL PATIENT AND DOES NOT BELIEVE THERE WAS ANY SORT OF DEVICE MALFUNCTION THAT CONTRIBUTED TO THIS EVENT [PATIENT PE, CARDIAC ARREST AND EXPIRATION]. THIS MEDICAL OPINION WAS DISCUSSED AND AGREED UPON WITH SEVERAL PHYSICIANS WITHIN THIS ACCOUNT." A DEVICE HISTORY RECORD (DHR) REVIEW OF THE PACKAGING/ASSEMBLY LOTS REVEALED NO QUALITY RELATED ISSUES OR MANUFACTURING DEFICIENCIES AT THE TIME OF MANUFACTURE. THERE HAVE BEEN NO OTHER REPORTED COMPLAINTS FOR THE INDICATED PACKAGING/ASSEMBLY LOTS FOR ANY DEVICE FAILURE MODES OR SIMILAR PATIENT THROMBOSIS/PE ISSUE. LABELING REVIEW: THE ANGIOVAC CANNULA SAMPLE WAS NOT RETURNED FOR EVALUATION SINCE THERE WAS NO REPORTED DEVICE MALFUNCTION DURING THE PROCEDURE. IT CANNOT BE DETERMINED IF DEVICE WAS USED IN ACCORDANCE WITH ITS LABELING. DIRECTIONS FOR USE (DFU) THAT IS PROVIDED WITH THIS DEVICE CONTAINS THE FOLLOWING STATEMENTS: WARNINGS: - SELECTION OF THE PATIENT AS A CANDIDATE FOR USE WITH THIS DEVICE AND FOR SUCH PROCEDURES AS IT IS INTENDED IS THE PHYSICIANS' SOLE RESPONSIBILITY. THE OUTCOME IS DEPENDENT ON MANY VARIABLES INCLUDING, PATIENT PATHOLOGY, SURGICAL PROCEDURE, AND PERFUSION PROCEDURE/TECHNIQUE. THE BENEFITS OF USE OF THIS DEVICE MUST BE WEIGHED AGAINST THE RISKS INCLUDING RISKS OF SYSTEMIC ANTICOAGULATION AND MUST BE ASSESSED BY THE PRESCRIBING PHYSICIAN. - AS WITH ALL MEDICAL DEVICES, THIS DEVICE AND ANCILLARY EQUIPMENT ARE TO BE USED BY TRAINED PHYSICIANS ONLY. SPECIFICALLY, THIS DEVICE IS TO BE USED ONLY BY MEDICAL PERSONNEL EXPERIENCED WITH INITIATING AND MONITORING EXTRACORPOREAL BYPASS AND BY PHYSICIANS TRAINED AND EXPERIENCED USING SURGICAL AND/OR PERCUTANEOUS (SELDINGER) VASCULAR ACCESS TECHNIQUES. ADVERSE EVENTS: THIS DEVICE, AS DO ALL EXTRACORPOREAL BLOOD VESSEL DEVICES, HAS POSSIBLE SIDE EFFECTS, WHICH INCLUDE BUT ARE NOT LIMITED TO INFECTIONS, BLOOD LOSS, THROMBUS FORMATION, EMBOLIC EVENTS, VESSEL, VENTRICULAR OR VALVULAR DAMAGE AND COMPLICATIONS OF PERCUTANEOUS OR SURGICAL INSERTION TECHNIQUES. THESE MAY OCCUR IF THE INSTRUCTIONS FOR USE ARE NOT FOLLOWED. POSSIBLE COMPLICATIONS INCLUDE THOSE NORMALLY ASSOCIATED WITH LARGE BORE SURGICAL AND/OR PERCUTANEOUS VESSEL CATHETERIZATION/CANNULATION, ANTICOAGULATION, EXTRACORPOREAL CIRCULATION AND APPLICATION OF INTRAVASCULAR INTRODUCER SYSTEMS WHICH INCLUDE BUT ARE NOT LIMITED TO: - DISTAL EMBOLIZATION OF THROMBUS. - PULMONARY EMBOLISM. - ARRHYTHMIAS. - DEATH. A REVIEW OF THE ANGIODYNAMICS COMPLAINT SYSTEM NOTED NO ADVERSE TRENDS FOR THIS COMPLAINT TYPE AND PRODUCT FAMILY. THIS TYPE OF COMPLAINT WILL CONTINUE TO BE MONITORED FOR TRENDS. REFERENCE (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21676892&pc=OZD | 1220648-2025-26907 | 2025-02-24 05:00 | Death | ABIOMED, INC. | 2025-03-24 04:00 | OZD | AUTOMATED IMPELLA CONTROLLER | Application Program Problem | Cardiac Arrest | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT (UNKNOWN AGE AND ETHNICITY) WITH AN EJECTION FRACTION OF 20% (DECREASE FROM 45% ONE DAY PRIOR) IN CARDIOGENIC SHOCK FROM AN ACUTE MYOCARDIAL INFARCTION WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. THEREAFTER, IT WAS NOTED THE AUTOMATED IMPELLA CONTROLLER (AIC) SCREEN WAS BLACK WHILE PATIENT WAS BEING ACTIVELY SUPPORTED. ASSISTANCE WAS CALLED AND THE TEAM WAS INSTRUCTED TO TURN THE AIC BACK ON USING THE POWER BUTTON. UPON POWER UP, ¿CONTROLLER FAILURE¿ AND ¿UNEXPECTED CONSOLE SHUTDOWN¿ ALARMS PRESENTED. THE TEAM THEN ADVISED TO DO AN AIC-TO-AIC TRANSFER, AND THE CASE COMPLETED WITHOUT ISSUE ON SECONDARY CONTROLLER. THE PATIENT WAS SENT TO THE UNIT ON BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) FOR RECOVERY AND SUBSEQUENT TRANSFER. HOWEVER, WITHIN THREE HOURS OF START OF SUPPORT, WHILE ON THE UNIT, THE PATIENT EXPIRED. THE PATIENT WAS NOTED TO BE DOING FINE ON BIPAP ON ARRIVAL TO UNIT; HOWEVER, THE PATIENT'S OXYGEN SATURATION STARTED DROPPING. DURING INTUBATION THE PATIENT WAS GIVEN A LARGE AMOUNT OF SEDATIVES. THE PATIENT'S HEART STOPPED, AND CARDIOPULMONARY RESUSCITATION WAS STARTED; THE PATIENT CODED FOR APPROXIMATELY 40 MINUTES UNTIL TIME OF DEMISE WAS CALLED. Manufacturer Narrative: THE INVESTIGATION OF AUTOMATED IMPELLA CONTROLLER (AIC) - CONTROLLER FAILURE (RED ALARM) HAS BEEN COMPLETED. THE CONSOLE LOGS FROM THE REPORTED DAY OF EVENT SHOW THAT ABOUT HALF AN HOUR AFTER PUMP START, IMC LOG ENTRIES END ABRUPTLY, AND THEN AIC IS POWER-CYCLED, WITH THE FIRST BOOT-UP UNSUCCESSFUL AND SECOND BOOT-UP CONCLUDES WITH CONTROLLER FAILURE ALARM AND UNEXPECTED CONTROLLER SHUTDOWN. ALTHOUGH THE PUMP WAS AUTOMATICALLY RESTARTED WITHIN 30 SECONDS, THE PURGE SYSTEM WAS NO LONGER OPERATIONAL. FOUR MINUTES LATER THE PUMP WAS TRANSFERRED TO A BACKUP CONSOLE, WHERE THE SUPPORT CONTINUED WITHOUT ANY ISSUES FOR ABOUT 2.5 HOURS. CONSOLE WAS TESTED AND CONTROLLER FAILURE ALARM WAS REPRODUCED. REPLACING PURGE PRESSURE SENSOR DID NOT RESOLVE THE ISSUE, HOWEVER WHEN PUD BOARD WAS REPLACED, THE ALARM RESOLVED EVEN WITH THE ORIGINAL PURGE PRESSURE SENSOR. CONSOLE SHUTDOWN/RESTART WAS NOT REPRODUCED DURING TESTING. REPORTED CONTROLLER FAILURE ALARM WAS CAUSED BY MALFUNCTION OF PURGE PRESSURE MEASUREMENT SYSTEM, DUE TO PUD BOARD DEFECT. E4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26907 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. Manufacturer Narrative: THE AUTOMATED IMPELLA CONTROLLER DEVICE WAS RETURNED, AND AN EVALUATION IS UNDERWAY. UPON COMPLETION OF OUR ANALYSIS, A SUPPLEMENTAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21677005&pc=OZD | 1220648-2025-26909 | 2025-02-26 05:00 | Death | ABIOMED, INC. | 2025-03-24 04:00 | OZD | IMPELLA CP | Restricted Flow rate | Ischemia; Renal Failure; Tachycardia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THE PATIENT WAS ON IMPELLA CP SUPPORT AND DURING SUPPORT, THE PATIENT HAD ACUTE KIDNEY INJURY (AKI). THE PATIENT'S URINE WAS WINE COLOR. THE PATIENT'S CR WAS 1.18 ON ADMISSION AND WAS 2.18 WHILE ON SUPPORT. DIALYSIS WAS STARTED. THE RELATIONSHIP BETWEEN THE AKI AND THE IMPELLA IS UNKNOWN. ADDITIONALLY, THERE WAS LOW PUMP FLOW. ONE UNIT OF PACKED RED BLOOD CELLS WERE GIVEN TO THE PATIENT AND THE SUCTION EVENTS RESOLVED. FURTHERMORE, THE PATIENT HAD LIMB ISCHEMIA. A FEM BYPASS WAS PERFORMED. DISTAL LIMB STILL COLD, ISCHEMIC AND MOTTLED. THE PUMP WAS ATTEMPTED TO BE WEANED BUT THE PATIENT WENT INTO VENTRICULAR TACHYCARDIA. THE PATIENT WAS SHOCKED MANY TIMES. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21677013&pc=QOC | MW5168015 | 2024-08-09 04:00 | Death | ABIOMED, INC. | 2025-03-24 04:00 | QOC | IMPELLA SMART ASSIST.5 |  | Insufficient Information |  |  | 1 | Event Description: IMPELLA PLACED AND DAD NEVER RECOVERED AND GOT EXTREMELY WORSE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21677026&pc=BZD | MW5168016 | 2025-02-22 05:00 | Death | DEVILBISS HEALTHCARE LLC /3B MEDICAL INC. | 2025-03-24 04:00 | BZD | DEVILBISS INTELLIPAP |  | Dyspnea |  |  | 1 | Event Description: ON (B)(6) 2025, MY MOTHER WAS PUT ON HOSPICE. SHE WAS DISCHARGED FROM THE HOSPITAL AND WENT HOME AS BIPAP DEPENDENT. MY MOTHER HAD COPD AND THOUGH SHE WAS PUT ON HOSPICE SHE WAS NO WHERE NEAR READY TO PASS. MY MOTHER ATE DINNER ON FRIDAY NIGHT, SHE GOT UP BY HERSELF, USED THE RESTROOM AND WAS FULLY ALLERT. SHE WENT TO SLEEP THAT NIGHT AT 1045 PM AND WOKE UP SOMETIME BETWEEN 5:30 AND 6 AM. AT THIS TIME MY MOTHER SAT UP IN HER BED AND WHEN DOING SO HER BIPAP TUBE CAME COMPLETELY OFF OF HER MACHINE. ONCE THIS HAPPENED, SHE WAS UNABLE TO REATTACH IT AND I WAS TOO LATE GETTING TO HER. THE MASK WAS STILL FULLY ATTACHED TO HER FACE BUT THE END THAT ATTACHES TO THE BIPAP MACHINE CAME COMPLETELY OFF. THE MACHINE HAS NO EJECTION PORT, NO ALARM AND JUST SLIDES RIGHT OFF THE MACHINE. AT THE TIME I GOT TO MY MOTHER SHE COULD NOT CATCH HER BREATH AND DIED AT 6:42 AM. I TRIED REATTACHING BUT BECAUSE OF THIS MACHINES FAILURE MY MOTHER SUFFERED AND DIED A HORRIBLE DEATH. SOMETHING NEEDS TO BE DONE TO MAKE THESE MACHINES SAFE FOR PEOPLE THAT ARE DEPENDENT UPON THEM. THE SIMPLE SLIDING ON AND OFF IS NOT SAFE AND KILLS PEOPLE WHO ARE DEPENDENT UPON THESE MACHINES. THE MACHINE SHE WAS USING: DEVILBISS INTELLIPAP, SN (B)(6).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21677039&pc=NIQ | 2024168-2025-03019 | 2025-03-10 04:00 | Death | ABBOTT VASCULAR INC. | 2025-03-24 04:00 | NIQ | XIENCE SKYPOINT¿ | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Thrombosis/Thrombus | P110019 |  | 1 | Event Description: IT WAS REPORTED THE PROCEDURE WAS TO TREAT LESIONS IN THE LEFT ANTERIOR DESCENDING (LAD) AND LEFT CIRCUMFLEX (LCX) ARTERIES. A 3.50X48MM AND A 3.50X38MM XIENCE SKYPOINT STENTS WERE IMPLANTED WITHOUT ISSUE. AT THE END THE PROCEDURE, THE PATIENT WAS AWAKE, ALERT, AND IN RECOVERY WITH FAMILY. HOWEVER, LATER THAT DAY THE PATIENT CODED AND WAS BROUGHT BACK TO THE CATH LAB WHERE ANGIOGRAPHY SHOWED THROMBUS IN THE LEFT MAIN (LM). A GUIDE WIRE WAS PLACED AND THROMBECTOMY WAS ATTEMPTED HOWEVER THE PATIENT EXPIRED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. PRODUCTION RECORD AND CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A QUERY OF THE COMPLAINT HANDLING DATABASE FOR THE REPORTED LOT REVEALED THERE IS NO INDICATION OF A LOT SPECIFIC ISSUE. THE REPORTED PATIENT EFFECTS OF DEATH AND THROMBOSIS ARE LISTED IN THE XIENCE SKYPOINT EVEROLIMUS ELUTING CORONARY STENT SYSTEMS INSTRUCTIONS FOR USE AS KNOWN PATIENT EFFECTS OF CORONARY STENTING PROCEDURES. A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECTS, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED; HOWEVER, THE SUBSEQUENT TREATMENT APPEARS TO BE RELATED TO THE OPERATIONAL CONTEXT OF THE PROCEDURE. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING. THE ADDITIONAL DEVICE REFERENCED IN B5 IS FILED UNDER A SEPARATE MEDWATCH REPORT NUMBER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21677040&pc=NIQ | 2024168-2025-03020 | 2025-03-10 04:00 | Death | ABBOTT VASCULAR INC. | 2025-03-24 04:00 | NIQ | XIENCE SKYPOINT¿ | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Thrombosis/Thrombus | P110019 |  | 1 | Event Description: IT WAS REPORTED THE PROCEDURE WAS TO TREAT LESIONS IN THE LEFT ANTERIOR DESCENDING (LAD) AND LEFT CIRCUMFLEX (LCX) ARTERIES. A 3.50X48MM AND A 3.50X38MM XIENCE SKYPOINT STENTS WERE IMPLANTED WITHOUT ISSUE. AT THE END THE PROCEDURE, THE PATIENT WAS AWAKE, ALERT, AND IN RECOVERY WITH FAMILY. HOWEVER, LATER THAT DAY THE PATIENT CODED AND WAS BROUGHT BACK TO THE CATH LAB WHERE ANGIOGRAPHY SHOWED THROMBUS IN THE LEFT MAIN (LM). A GUIDE WIRE WAS PLACED AND THROMBECTOMY WAS ATTEMPTED HOWEVER THE PATIENT EXPIRED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. PRODUCTION RECORD AND CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A QUERY OF THE COMPLAINT HANDLING DATABASE FOR THE REPORTED LOT REVEALED THERE IS NO INDICATION OF A LOT SPECIFIC ISSUE. THE REPORTED PATIENT EFFECTS OF DEATH AND THROMBOSIS ARE LISTED IN THE XIENCE SKYPOINT EVEROLIMUS ELUTING CORONARY STENT SYSTEMS INSTRUCTIONS FOR USE AS A KNOWN PATIENT EFFECTS OF CORONARY STENTING PROCEDURES. A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECTS, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED; HOWEVER, THE SUBSEQUENT TREATMENT APPEARS TO BE RELATED TO THE OPERATIONAL CONTEXT OF THE PROCEDURE. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING. THE ADDITIONAL DEVICE REFERENCED IN B5 IS FILED UNDER A SEPARATE MEDWATCH REPORT NUMBER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21677042&pc=DYB | MW5168017 | 2025-02-04 05:00 | Death | ABBOTT MEDICAL | 2025-03-24 04:00 | DYB | ABBOTT SL1 SHEATH |  | Great Vessel Perforation |  |  | 1 | Event Description: IT WAS REPORTED BY THE CALLER THAT WHILE PERFORMING A PROCEDURE IN THE ANIMAL LAB THEY PERFORATED THE AORTA WITH THE SL1 SHEATH AND THE ANIMAL PASSED AWAY. BWI PRODUCTS IN USE WERE: CARTO 3 SYSTEM SER# (B)(6), STSF DF CURVE. CATALOG AND LOT #'S ARE UNKNOWN. THIS REPORT REFLECTS INFORMATION RECEIVED BY FDA IN THE FORM OF A NOTIFICATION PER 803.22 (B)(2).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21677259&pc=DSQ | 3007042319-2025-00681 | 2024-12-22 05:00 | Death | HEARTWARE, INC. | 2025-03-24 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT, WITH A MEDICAL HISTORY OF CONGESTIVE HEART FAILURE, WAS PLACED ON HOSPICE CARE DUE TO HEALTH ISSUES NOT RELATED TO THE VENTRICULAR ASSIST DEVICE (VAD) AND SUBSEQUENTLY PASSED AWAY. THE TREATING PHYSICIAN DETERMINED THAT THE EVENT WAS NOT RELATED TO THE DEVICE OR PROCEDURE, AND THERE WAS NO DEVICE MALFUNCTION. THE CAUSE OF DEATH WAS NOT DISCLOSED. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: THE VENTRICULAR ASSIST DEVICE (VAD) (B)(6) WAS NOT RETURNED FOR EVALUATION AS IT REMAINED IMPLANTED. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE MANUFACTURING DOCUMENTATION WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT WAS PLACED ON HOSPICE CARE DUE TO HEALTH ISSUES NOT RELATED TO THE VAD AND SUBSEQUENTLY EXPIRED. BASED ON THE LIMITED INFORMATION AVAILABLE, THE DEVICE MAY HAVE CAUSE OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, DEATH IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. THERE WAS NO EVIDENCE THAT THE PATIENT HAD A HISTORY OF SIMILAR ADVERSE EVENTS. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21677687&pc=DSQ | 2916596-2025-01638 | 2024-04-04 04:00 | Death | THORATEC CORPORATION | 2025-03-24 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Fall; Hip Fracture; Hyponatremia | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT SUBMITTED LOG FILES APPEARED TO CAPTURE A DRIVELINE DISCONNECTION ON (B)(6) 2024. THE DISCONNECTION WAS DUE TO A WITHDRAWAL OF CARE AT THE REQUEST OF THE PATIENT'S FAMILY. THE PATIENT HAD MULTIPLE READMISSIONS FOR MULTIPLE CO-MORBIDITIES INCLUDING NON-SMALL CELL LUNG ADENOCARCINOMA, CHRONIC HYPONATREMIA, NAUSEA, VOMITING, FAILURE TO THRIVE/MALNUTRITION, AND DEHYDRATION, AND A RECENT FALL THAT CAUSED A HIP FRACTURE (BELIEVED TO BE DUE TO DEHYDRATION/MALNUTRITION). THE PATIENT WAS NOT A SURGICAL, RADIATION, OR CHEMOTHERAPY CANDIDATE. THE PATIENT WAS ON ANTIEMETICS AND HAD A GASTROJEJUNOSTOMY (GJ) TUBE FOR BOLUS FEEDINGS, AND LATER CONTINUOUS FEEDINGS, ALONG WITH FREE WATER BOLUSES, AND THE PATIENT FREQUENTLY REFUSED FEEDING AND FLUIDS DUE TO DISCOMFORT. ON ADMISSION THE PATIENT WAS GIVEN INTRAVENOUS (IV) FLUIDS BUT WOULD CONTINUE TO REFUSE ANY ORAL FLUIDS OR ANY OTHER ENTERAL INTAKE. THE PATIENT PASSED AWAY ON (B)(6) 2024. THE OFFICIAL CAUSE OF DEATH WAS UNKNOWN, AS THE PATIENT DIED AT HOME. THE DEATH WAS NOT CONSIDERED DEVICE RELATED, AND THE DEVICE WAS NOTED TO HAVE OPERATED AS EXPECTED. NO AUTOPSY WAS PERFORMED, AND THE PUMP WAS NOT RETURNED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, ALSO ADDRESSES ALL PUMP PARAMETERS. SECTION 4 OF THE IFU, ¿SYSTEM MONITOR¿, DESCRIBES THE PUMP FLOW DISPLAY AND THE HAZARD ALARMS AND EXPLAINS THAT CHANGES IN PATIENT CONDITION CAN RESULT IN LOW FLOW. PER DESIGN, WHEN THE ESTIMATED FLOW VALUE IS CALCULATED AT LESS THAN 2.5 LITERS PER MINUTE (LPM), A LOW FLOW STATUS IS POSTED TO THE LOG FILE. IF THE FLOW REMAINS BELOW 2.5 LPM FOR 10 SECONDS, A LOW FLOW HAZARD ALARM IS TRIGGERED. IN REFERENCE TO PULSATILITY INDEX (PI), SECTION 1 AND SECTION 4 OF THE IFU STATE THAT ¿PI CALCULATION REPRESENTS CARDIAC PULSATILITY AND PI VALUES TYPICALLY RANGE FROM 1 TO 10. IN GENERAL, THE MAGNITUDE OF THE PI VALUE IS RELATED TO THE AMOUNT OF ASSISTANCE PROVIDED BY THE PUMP. HIGHER VALUES INDICATE MORE VENTRICULAR FILLING AND HIGHER PULSATILITY (I.E. THE PUMP IS PROVIDING LESS SUPPORT TO THE LEFT VENTRICLE). LOWER VALUES INDICATE LESS VENTRICULAR FILLING AND LOWER PULSATILITY (I.E. THE PUMP IS PROVIDING GREATER SUPPORT AND FURTHER UNLOADING THE VENTRICLE)¿. SECTION 5 OF THE PATIENT HANDBOOK, ¿ALARMS AND TROUBLESHOOTING¿, AND SECTION 7 OF THE IFU, ¿ALARMS AND TROUBLESHOOTING¿, PROVIDES INFORMATION ON ALL SYSTEM ALARM CONDITIONS AS WELL AS THE APPROPRIATE ACTIONS ASSOCIATED WITH EACH CONDITION. FURTHERMORE, SECTION 8 OF THE PATIENT HANDBOOK, ¿HANDLING EMERGENCIES¿, ALSO PROVIDES EXAMPLES OF EMERGENCIES AND THE PROPER ACTIONS TO TAKE IN THE EVENT AN EMERGENCY OCCURS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21677963&pc=OZD | 1220648-2025-26912 | 2025-02-16 05:00 | Death | ABIOMED, INC. | 2025-03-24 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT WHILE A PATIENT WAS ON IMPELLA CP SUPPORT THEY SUFFERED A CEREBRAL HEMORRHAGE AND DIED WITH NO SIGNS OF IMPROVEMENT IN BILATERAL HEART FAILURE. THE DOCTORS COMMENTED THAT THE CAUSE OF DEATH WAS NOT RELATED TO THE IMPELLA AND THE IMPELLA FUNCTIONED NORMALLY WITHOUT ANY ISSUES. HOWEVER, THE DEVICE COULD NOT CONCLUSIVELY BE DISASSOCIATED FROM THE OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21678412&pc=LWP | 2124215-2025-17893 | 2017-01-27 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-03-24 04:00 | LWP | ACCOLADE MRI DR | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P150012 |  | 1 | Event Description: IT WAS REPORTED BY THE PATIENT'S SPOUSE THAT THE PATIENT HAD PASSED AWAY ON THE DATE THE PACEMAKER WAS IMPLANTED IN 2017. THE DEVICE WAS SUSPECTED TO BE BURIED WITH THE PATIENT AS IT WAS NOT RETURNED FOR ANALYSIS. Manufacturer Narrative: IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21678554&pc=DSQ | 2916596-2025-01612 | 2025-01-01 05:00 | Death | THORATEC CORPORATION | 2025-03-24 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED IN THE RESEARCH ARTICLE ¿GENOMICS OF ADVANCED HEART FAILURE IN NON-ISCHEMIC DILATED CARDIOMYOPATHY¿ COMPARISON OF CLINICAL PRESENTATIONS, CO-MORBIDITIES, AND OUTCOMES AFTER LVAD PLACEMENT IN PATIENTS WITH NON-ISCHEMIC CARDIOMYOPATHY (NICMP) AFTER LVAD PLACEMENT WAS DONE. GENETIC TESTING THROUGH A 157 GENE ARRYTHMIA AND CARDIOMYOPATHY COMPREHENSIVE PANEL WAS DONE TO ASSESS PATIENTS WITH IDENTIFIABLE PATHOGENIC MUTATIONS AND THOSE WITHOUT. PRIMARY ENDPOINT WAS MORTALITY. SECONDARY OUTCOMES INCLUDED RV FAILURE, HEART FAILURE EXACERBATIONS, VENTRICULAR ARRYTHMIAS. OVERALL SURVIVAL WAS COMPARED USING COX-PROPORTIONAL HAZARDS REGRESSION AND SECONDARY OUTCOMES WERE ANALYZED USING MULTIVARIABLE LOGISTIC REGRESSION MODELS ADJUSTED FOR AGE, SEX, AND INTERMACS CLASSIFICATION. THE STUDY INCLUDED 75 PATIENTS; 13 PATIENTS HAD PATHOGENIC MUTATIONS (46.2%TITINMUTATIONS). FOR ALL PATIENTS, THE MEAN AGE WAS 49.7YEARS, 32% FEMALE, 40% REQUIRED INOTROPE THERAPY PRIOR TO ADMISSION, 59% WERE EITHER INTERMACS I OR II ON PRESENTATION, 48% HAD HEARTMATE3 IMPLANTED, AND 43% HAD HEARTWARE IMPLANTED. THE PRIMARY ENDPOINT AT MEDIAN FOLLOW-UP OF 3.8 YEARS WAS MORTALITY( 7.7% FOR THE IDENTIFIED PATHOGENIC COHORT VS 19.4%, ADJUSTED HR 0.43, 95% CI [0.06 3.33],P=0.42).SECONDARY OUTCOMES 1-YER POST LVAD PLACEMENT INCLUDED HEART FAILURE EXACERBATIONS (7.7% VS 26.7%, ADJUSTED HR 0.84[0.66-1.08], P=0.18), VENTRICULAR ARRYTHMIAS (23.1%VS15%, ADJUSTEDHR2.16[0.46 10.11],P=0.33), STROKE OCCURRENCE OVER 1YEA R(23.1% VS 6.67%, ADJUSTED HR 4.54 [0.81-25.5], P=0.09), AND OCCURRENCE OF RV FAILURE (38.5% VS 45.9%, ADJUSTED HR 0.74[0.22-2.55],P=0.64).OTHER OUTCOMES INCLUDED OCCURRENCE OF HEART TRANSPLANTATION (38.5%VS. 19.4%;ADJUSTEDHR2.8 [0.75-10.5],P=0.13)OR LVAD EXPLANTATION (0% VS. 4.8%,P=0.975).CONCLUSIONS: IN THIS SMALL, SINGLE-CENTER, RETROSPECTIVE STUDY, THERE WAS NO STATISTICALLY SIGNIFICANT DIFFERENCE IN OUTCOMES FOR PATIENTS WITH IDENTIFIED PATHOGENIC MUTATIONS IN THE NICMP POPULATION IN THE 1-YEAR POST LVAD PLACEMENT. Manufacturer Narrative: SECTION E1: REPORTER EMAIL WAS NOT AVAILABLE. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM AND THE REPORTED PATIENT DEATHS COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE HEARTMATE 3 DEVICE SERIAL NUMBERS, AS WELL AS OTHER SPECIFIC CASE/PATIENT INFORMATION, ARE NOT AVAILABLE. THE SERIAL NUMBERS OR OTHER IDENTIFYING INFORMATION OF THE PRODUCTS WERE NOT REPORTED AND WERE NOT ABLE TO BE DETERMINED DURING THE INVESTIGATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. CHAPTER 1, ¿INTRODUCTION, LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. DATE OF EVENT IS LISTED AS 01JAN2025, THE DATE OF PUBLICATION, AS NO IMPLANT INFORMATION WAS GIVEN. AJIT K KODURI, CLAIRE SCOTT, RIYAD Y KHERALLAH, HANNAH SMATI, AJITH NAIR, GENOMICS OF ADVANCED HEART FAILURE IN NON-ISCHEMIC DILATED CARDIOMYOPATHY, JOURNAL OF CARDIAC FAILURE, VOLUME 31, ISSUE 1, 2025, PAGE 224, ISSN 1071-9164, HTTPS://DOI.ORG/10.1016/J.CARDFAIL.2024.10.114. BAYLOR COLLEGE OF MEDICINE, HOUSTON, TX. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21678760&pc=MKJ | 3030677-2025-000942 | 2025-03-20 04:00 | Death | PHILIPS NORTH AMERICA LLC | 2025-03-24 04:00 | MKJ | FR3 DEFIB, ECG UK ENGLISH (T) EXCHANGE | Failure to Deliver Shock/Stimulation | Cardiac Arrest | P160028 |  | 1 | Event Description: IT HAS BEEN REPORTED THAT THE DEVICE'S ANALYSIS OF THE RHYTHM WAS DELAYED DUE TO EXCESSIVE ARTIFACTS DETECTED. PATIENT IS DECEASED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21678944&pc=DSI | 1218950-2025-000182 | 2025-02-16 05:00 | Death | PHILIPS MEDICAL SYSTEMS | 2025-03-24 04:00 | DSI | MX40 1.4 GHZ SMART HOPPING | No Audible Alarm | Insufficient Information | K113125 |  | 1 | Event Description: IT WAS REPORTED ALARMS WERE NOT GENERATED FOR AN INCIDENT WITH A TELEMETRY PATIENT WHO PASSED AWAY. THE TELEMETRY PATIENT WAS CHECKED BY THE INCOMING NURSE AS THE CENTRAL STATION WAS DISPLAYING A QUESTION MARK IN THE PATIENT SECTOR AND FOUND UNRESPONSIVE WITH NO PULSE. CPR WAS INITIATED. THE STAFF WAS NOT SURE WHAT HAPPENED BUT IT WAS DETERMINED THE PATIENT'S TELEMETRY DEVICE WAS NEVER ASSIGNED AT THE CENTRAL STATION; HOWEVER, THIS WAS NOT THE CASE, AND THE PATIENT WAS ADMITTED. THE CUSTOMER INDICATED THE PATIENT PASSED AWAY DUE TO UNKNOWN CAUSES A FEW DAYS LATER. THE AUDIT LOGS REVEALED ALARMS HAD BEEN GENERATED AND ACKNOWLEDGED. Manufacturer Narrative: THE EVENT TOOK PLACE ON FEBRUARY 16, 2025 BETWEEN 12:53 AND 18:45 IN BED (B)(4). DURING THE INVESTIGATION, A TECHNICAL AND CLINICAL ANALYSIS OF THE LOG DATA PROVIDED WAS PERFORMED. THE FINDINGS ARE BELOW. THE INVESTIGATION FOUND THAT ON (B)(6) 2025 AT 22:09:13, THE PATIENT WAS ADMITTED. FROM ADMISSION TO ON (B)(6) 2025 2:06:35 (3 HOURS 57 MIN 22 SECONDS), THERE WAS EITHER NO ECG ALARMING POSSIBLE OR NO PHYSIOLOGICAL MONITORING AT ALL POSSIBLE FOR A TOTAL OF 1 HOUR 34 MIN AND 18 SECONDS, DUE TO: ECG LEADS OFF, LEADSET UNPLUGGED, NO DATA TELE CONDITIONS. DURING THIS PERIOD OF TIME THERE WERE ALSO LA [LEFT ARM] LEAD OFF INOPS WHICH LIMITS ARRHYTHMIA ANALYSIS TO JUST ONE LEAD FOR A TOTAL OF 27 MINUTES AND 33 SECONDS. ANALYSIS AND ALARMS OCCURRED BETWEEN ON (B)(6) 2025 2:06:35 AND ON (B)(6) 2025 12:51:53 (10 HOURS 45 MINUTES 16 SECONDS). STARTING ON (B)(6) 2025 12:51:53 UNTIL ON (B)(6) 2025 18:45:54 (5 HOURS 54 MINUTES AND 1 SECOND), THERE WAS EITHER NO ECG ALARMING POSSIBLE OR NO PHYSIOLOGICAL MONITORING AT ALL POSSIBLE FOR A TOTAL OF 3 MINUTES AND 52 SECONDS, DUE TO: ECG LEADS OFF, LEADSET UNPLUGGED (3 MINUTES 49 SECOND EVENT ¿ THIS LEADS TO AN ARRYTHMIA RELEARN) CONDITIONS. FOLLOWING THESE EVENTS, THERE WERE LA LEAD OFF INOPS WHICH LIMITS ARRHYTHMIA ANALYSIS TO JUST ONE LEAD FOR A TOTAL OF 1 HOUR 51 MINUTES AND 48 SECONDS. THE LA LEAD OFF INOP WAS CORRECTED AT 15:57 AND MONITORING WITH ALL AVAILABLE ECG LEADS CONTINUED UNTIL THE V LEAD OFF ALARM GENERATED AT 18:45:46. STARTING ON (B)(6) 2025 18:51:08, THERE WAS A 2 HOUR 20 MINUTES AND 28 SECONDS ECG LEADS OFF CONDITION (NO ECG ALARMING POSSIBLE). THE ECG LEADS OFF ENDED DUE TO NO DATA TELE AT 21:11:15 AND THEN THE PATIENT WAS TRANSFERRED TO A DIFFERENT ROOM AT 21:18 ON (B)(6) 2025. THEY WOULD HAVE HAD A ¿?¿ IN THE SECTOR FOR HR WITH A FLAT WAVEFORM WHILE THE ECG LEADS OFF WAS PRESENT. ONCE THE DEVICE LOST CONNECTION TO THE PIC IX, THERE IS NO ¿?¿. THE SECTOR IS JUST BLANK. THE INVESTIGATION FOUND THAT THERE WAS EITHER NO ECG ALARMING POSSIBLE, NO PHYSIOLOGICAL MONITORING AT ALL POSSIBLE, OR LIMITED MONITORING DUE TO ONGOING ECG LEAD OFF ISSUES DURING THE REPORTED PERIOD OF TIME. FROM THE AUDIT LOG, ECG MONITORING WAS AVAILABLE FROM 15:57 WHEN THE LA LEAD OFF ENDED UNTIL!! ECG LEADS OFF GENERATED AT 18:50:58. THIS ALARM SOUNDED UNTIL IT WAS ACKNOWLEDGED AT 19:07:58 AND CONTINUED TO REMIND EVERY THREE MINUTES UNTIL 21:08:44. THIS ALARM WAS ACKNOWLEDGED AGAIN AT 19:42:01. THE REPORTED PROBLEM IS NOT CONFIRMED. BASED ON THIS INFORMATION, THERE DOES NOT APPEAR TO BE ANY DEVICE MALFUNCTION; HOWEVER, A COMBINATION OF USE ERROR, ALARM MANAGEMENT, AND CLINICAL WORKFLOW DURING USE OF THE DEVICE MAY HAVE BEEN FACTORS IN THE PATIENT'S OUTCOME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21678981&pc=MNS | 2518422-2025-103340 | 2025-03-19 04:00 | Death | RESPIRONICS, INC. | 2025-03-24 04:00 | MNS | DREAMSTATION | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K102465 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION STATING A PATIENT SADLY PASSED AWAY. THE PATIENT WAS USING A BIPAP AVAPS30 DEVICE, ALTHOUGH NO INFORMATION WAS PROVIDED TO VERIFY IF THE DEVICE WAS IN USE WHEN THE PATIENT PASSED. THE REPORTED EVENT MAKES NO ALLEGATION OF ANY SPECIFIC DEVICE MALFUNCTION, USE ERROR, FAILURE, LABELING, OR OTHER DEFICIENCY THAT IS RELEVANT TO THE REPORTED EVENT. ADDITIONALLY, THERE IS NO OTHER CLINICAL INFORMATION PROVIDED TO INDICATE ANY CAUSAL RELATIONSHIP BETWEEN THE DEVICE AND THE REPORTED DEATH. ALSO, PER REPORT OF THE EVENT, "THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH." THEREFORE, BASED ON AVAILABLE INFORMATION THE REPORTED DEATH IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE AND DID NOT CAUSE OR CONTRIBUTE TO A SERIOUS INJURY OR DEATH. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM IF ANY MALFUNCTION OCCURRED. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21679023&pc=DSQ | 2916596-2025-01610 | 2025-03-03 05:00 | Death | THORATEC CORPORATION | 2025-03-24 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO INTRACRANIAL BLEED (ICB) AFTER A SEVERE FALL FROM STAIRS. THE ICB WAS REPORTED NOT TO BE RELATED TO THE PUMP OR ITS THERAPY. NO DEVICE AND/OR ANTICOAGULATION ISSUES WERE MENTIONED PRIOR TO THE OUTCOME. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED FALL, INTRACRANIAL BLEEDING, AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. G IS CURRENTLY AVAILABLE. THE IFU LISTS POTENTIAL ADVERSE EVENTS, INCLUDING BLEEDING, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. THE IFU ALSO OUTLINES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING THE INTERNATIONAL NORMALIZED RATIO RANGE FOR PATIENTS USING THE HEARTMATE 3 LVAS, AS WELL AS THE SUGGESTED ANTICOAGULATION MODIFICATIONS IN THE EVENT THAT THERE IS A RISK OF BLEEDING. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21679377&pc=DSQ | 2916596-2025-01605 | 2023-07-01 04:00 | Death | THORATEC CORPORATION | 2025-03-24 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: THE RETROSPECTIVE STUDY ENTITLED ¿REEVALUATING HOW WE VIEW RENAL DYSFUNCTION DURING LEFT VENTRICULAR ASSIST DEVICE CONSIDERATION¿ SOUGHT HOW TO VIEW COMORBIDITIES LIKE RENAL DISEASE WHEN CONSIDERING PATIENT¿S FOR LEFT VENTRICULAR ASSIST DEVICE, LVAD, BY ASSESSING SURVIVAL RATES WHEN CONSIDERING THEIR ESTIMATED GLOMERULAR FILTRATION RATE, EGFR, AT TIME OF IMPLANTATION. THIS SINGLE CENTER, RETROSPECTIVE STUDY REVIEWED 177 PATIENTS IMPLANTED WITH HEARTMATE 3, HM3, LVADS BETWEEN 1/1/2021 AND 7/1/2023, OF WHICH WERE THEN SORTED INTO 3 GROUPS: GROUP A CONSISTED OF 73 PATIENTS WITH EGFR RATES OF = 60, GROUP B CONSISTED OF 64 PATIENTS WITH EGFR RATES OF 60 > AND > 30, AND GROUP C CONSISTED OF 40 PATIENTS WITH EGFR SCORE = 30 OR ON DIALYSIS. EACH GROUPS SURVIVAL AT 6,12 AND 24 MONTHS WAS EVALUATED AND COMPARED THEIR DEVELOPMENT OR REDEVELOPMENT OF A DIALYSIS REQUIREMENT POST IMPLANT WITH A CHI-SQUARED TEST. 11 GROUP C PATIENTS WERE ON DIALYSIS AT THE TIME OF IMPLANTATION, (1 HEMODIALYSIS, 10 CONTINUOUS RENAL REPLACEMENT THERAPY, CRRT) WITH THREE REMAINING ON DIALYSIS INDEFINITELY. THE MEDIAN LENGTH OF DIALYSIS REQUIREMENT FOR THESE ELEVEN PATIENTS AFTER LVAD IMPLANTATION WAS 33 DAYS. GROUP C PATIENTS HAD A SEEMINGLY LOWER 6-MONTH SURVIVAL RATES; HOWEVER, THEY HAD A SIMILAR 12 AND 24-MONTH SURVIVAL COMPARED TO HIGHER EGFR GROUPS (FIGURE, TABLE). THE PERCENTAGE OF PATIENTS WHO DEVELOPED OR REDEVELOPED A DIALYSIS REQUIREMENT IN EACH GROUP WAS 12.3%, 12.5%, AND 18.9%, RESPECTIVELY (P=0.59).
CONCLUSION: THIS STUDY DEMONSTRATED THAT AN EGFR =30 OR BEING ON DIALYSIS AT IMPLANTATION IS LIKELY INSUFFICIENT TO PREDICT A POOR OUTCOME POST-LVAD IN THOSE IMPLANTED WITH AN HM3. Manufacturer Narrative: SECTION B3 DATE CORRECTION. MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. NO PRODUCT WAS EVALUATED UNDER THIS COMPLAINT. THE HEARTMATE 3 DEVICE SERIAL NUMBERS, AS WELL AS OTHER SPECIFIC CASE/PATIENT INFORMATION, ARE NOT AVAILABLE. THE SERIAL NUMBER OR OTHER IDENTIFYING INFORMATION OF THE PRODUCT WAS NOT REPORTED AND WAS NOT ABLE TO BE DETERMINED DURING THE INVESTIGATION. THE HEARTMATE 3 LVAS IFU, REV. C IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21679380&pc=NPT | 2025587-2025-02071 | 2025-03-04 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-24 04:00 | NPT | EVOLUT FX PLUS VALVE | Perivalvular Leak; Patient Device Interaction Problem | Rupture; Valvular Insufficiency/ Regurgitation | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING THE IMPLANT OF A TRANSCATHETER AORTIC VALVE, A POST-IMPLANT BALLOON AORTIC VALVULOPLASTY (BAV) WAS PERFORMED. DURING THE POST-IMPLANT BAV, AN ANNULAR RUPTURE OCCURRED. THE PATIENT WAS PLACED ON CARDIOPULMONARY BYPASS, AND AN AORTIC ROOT REPLACEMENT WAS PERFORMED. THE VALVE WAS SUBSEQUENTLY EXPLANTED, AND THE PATIENT EXPIRED BEFORE BEING TAKEN OFF OF BYPASS. PERTHE PHYSICIAN, THE CAUSE OF DEATH WAS DUE TO THE ANNULAR RUPTURE.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT DURING THE VALVE IMPLANT, THE POST IMPLANT BALLOON AORTIC VALVULOPLASTY (BAV) WAS PERFORMED DUE TO MODERATE PARAVALVULAR LEAK (PVL). THE POST IMPLANT BAV WAS PERFORMED WITH A NON-MEDTRONIC (TRUE) 21 MILLIMETER (MM) SEMI-COMPLAINT BALLOON. ONE INFLATION WAS PERFORMED USING A DEFLATOR DEVICE. Manufacturer Narrative: CORRECTED DATA: H6 - CORRECTED TO INCLUDE IME CODE E0621 AND ANNEX A CODE A050405 ADDITIONAL CODES - CORRECTED TO INCLUDE IMG COMPONENT CODE G04027. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5 CONTINUATION OF D10: PRODUCT ID 21 MM TRUE BALLOON; PRODUCT LOT/SERIAL NUMBER NA; PRODUCT TYPE: DILATION BALLOON. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID {D-EVPROP2329US}; PRODUCT LOT/SERIAL NUMBER {UNKNOWN}; PRODUCT TYPE: {0195-HEARTVALVES}; IMPLANT DATE {N/A}; EXPLANT DATE {N/A}; PRODUCT ID {UNKNOWN}; PRODUCT LOT/SERIAL NUMBER {UNKNOWN}; PRODUCT TYPE: {BALLOON AORTIC VALVULOPLASTY}; IMPLANT DATE {N/A}; EXPLANT DATE {N/A}. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21679543&pc=GCX | 2020550-2025-00491 | 2025-03-19 04:00 | Death | KARL STORZ SE & CO. KG | 2025-03-24 04:00 | GCX | S-PILOT | Power Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  |  | 1 | Event Description: IT WAS REPORTED THAT, THE INITIAL CHECKS WERE DONE, THE S-PILOT WAS WORKING, INTRA-OPERATIVELY POWER COULD NOT BE ESTABLISHED. A PROLONGED MORCELLATION TIME WAS INDICATED. THE WHOLE PROCEDURE WAS PROLONGED MORE THAN 60 MINUTES. AS CONSEQUENCE, FURTHER INTERVENTIONS WERE REQUIRED (CT SCAN, X-RAY AND ITU ADMISSION) AND THEREFORE ADDITIONAL HOSPITALIZATION NECESSARY. IN ADDITION, ON 2025-03-19, THE INFORMATION WAS RECEIVED THAT THE PATIENT HAS PASSED AWAY. Manufacturer Narrative: THE AFFECTED DEVICE HAS BEEN REQUESTED FOR INVESTIGATION BY THE MANUFACTURER. DEVICE WAS NOT RETURNED AT THE TIME OF MDR SUBMISSION. THE EVENT IS FILED UNDER INTERNAL KARL STORZ COMPLAINT ID: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21679979&pc=NPT | 2025587-2025-02074 | 2025-02-26 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-24 04:00 | NPT | EVOLUT FX PLUS VALVE | Device Dislodged or Dislocated; Activation Failure; Patient Device Interaction Problem | Low Blood Pressure/ Hypotension; Perforation | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT PRIOR TO A PROCEDURE INVOLVING THE TRANSCATHETER AORTIC VALVE IN A PATIENT WITH A SEVERELY CALCIFIED AORTIC VALVE, A HORIZONTAL AORTIC ROOT, AND A DILATED ASCENDING AORTA, A PRE-IMPLANT BALLOON AORTIC VALVULOPLASTY WAS PERFORMED USING AN 18 MM NON-MEDTRONIC BALLOON DUE TO CALCIFICATION. THE FIRST VALVE WAS IMPLANTED, BUT WAS UNDER-EXPANDED. A POST-DILATION WAS PERFORMED WITH A 21 MM NON-MEDTRONIC BALLOON. UPON REMOVAL, THE BALLOON CAUGHT ON THE VALVE, CAUSING IT TO DISLODGE INTO THE ASCENDING AORTA AT THE LEVEL OF THE SINOTUBULAR JUNCTION. A SECOND VALVE WAS INSERTED INTO THE PATIENT AND PLACED THROUGH THE DISLODGED VALVE, BUT IT PUSHED THE INFLOW OF THE FIRST VALVE INTO THE AORTA, RESULTING IN A PERFORATION. SUBSEQUENTLY, THE PATIENT'S BLOOD PRESSURE DROPPED. THE SECOND VALVE AND DELIVERY CATHETER SYSTEM WERE WITHDRAWN FROM THE PATIENT. SURGICAL INTERVENTION WAS PERFORMED. THE DISLODGED VALVE WAS EXPLANTED FROM THE PATIENT, A SURGICAL AORTIC VALVE REPLACEMENT WAS PERFORMED, AND THE PERFORATED ASCENDING AORTA WAS REPAIRED. DESPITE THESE EFFORTS, THE PATIENT PASSED AWAY ON THE OPERATING ROOM TABLE.Event Description: IT WAS REPORTED THAT PRIOR TO A PROCEDURE INVOLVING THE TRANSCATHETER AORTIC VALVE IN A PATIENT WITH A SEVERELY CALCIFIED AORTIC VALVE, A HORIZONTAL AORTIC ROOT, AND A DILATED ASCENDING AORTA, A PRE-IMPLANT BALLOON AORTIC VALVULOPLASTY WAS PERFORMED USING AN 18 MM NON-MEDTRONIC BALLOON DUE TO CALCIFICATION. THE FIRST VALVE WAS IMPLANTED BUT WAS UNDER-EXPANDED. A POST-DILATION WAS PERFORMED WITH A 21 MM NON-MEDTRONIC BALLOON. UPON REMOVAL, THE BALLOON CAUGHT ON THE VALVE, CAUSING IT TO DISLODGE INTO THE ASCENDING AORTA AT THE LEVEL OF THE SINOTUBULAR JUNCTION. A SECOND VALVE WAS INSERTED INTO THE PATIENT AND PLACED THROUGH THE DISLODGED VALVE, BUT IT PUSHED THE INFLOW OF THE FIRST VALVE INTO THE AORTA, RESULTING IN A PERFORATION. SUBSEQUENTLY, THE PATIENT'S BLOOD PRESSURE DROPPED. THE SECOND VALVE AND DELIVERY CATHETER SYSTEM WERE WITHDRAWN FROM THE PATIENT. SURGICAL INTERVENTION WAS PERFORMED. THE DISLODGED VALVE WAS EXPLANTED FROM THE PATIENT, A SURGICAL AORTIC VALVE REPLACEMENT WAS PERFORMED, AND THE PERFORATED ASCENDING AORTA WAS REPAIRED. DESPITE THESE EFFORTS, THE PATIENT PASSED AWAY ON THE OPERATING ROOM TABLE. ADDITIONAL INFORMATION WAS RECEIVED THAT THE PATIENT WAS RAPID PACED DURING POST-IMPLANT DILATION. THE PATIENT HAD A SEVERELY CALCIFIED AORTIC VALVE, HORIZONTAL AORTIC ROOT ANGLE AND DILATED ASCENDING AORTA, WHICH CONTRIBUTED TO THE DISLODGEMENT. THE DEPLOYMENT STARTING POINT WAS AT THE BOTTOM OF THE PIGTAIL CATHETER. PRIOR TO VALVE DISLODGEMENT, THE IMPLANT DEPTH WAS 4 MM ON THE NON-CORONARY CUSP (NCC) AND 5 MM ON THE LEFT CORONARY CUSP (LCC). A MEDTRONIC (CONFIDA) GUIDEWIRE WAS USED DURING THE PROCEDURE. PER THE PHYSICIAN, THE CAUSE OF DEATH WAS PERFORATION OF THE ASCENDING AORTA WHICH WAS NOT RELATED TO THE FIRST DELIVERY CATHETER SYSTEM (DCS), HOWEVER, THE INFLOW CROWNS OF THE DISLODGED VALVE WERE BELIEVED TO CAUSE THE PERFORATION WHEN TRYING TO PASS THE SECOND VALVE THROUGH THE FIRST. THE PHYSICIAN INDICATED THAT THE DIFFICULT ANATOMY PLAYED A FACTOR IN CONTRIBUTING TO THE PATIENT¿S DEATH. Manufacturer Narrative: OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: PRODUCT ID GWBC30 (LOT: UNKNOWN); PRODUCT TYPE: 0193-GUIDEWIRE; IMPLANT DATE N/A; EXPLANT DATE N/A UPDATED: A4, B5, H11. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID D-EVOLUTFX-2329 (LOT: 0012610943); PRODUCT TYPE: 0195-HEART VALVES; PRODUCT ID EVFXPLUS-26 (SERIAL: (B)(6); PRODUCT TYPE: 0195-HEART VALVES; PRODUCT ID D-EVOLUTFX-2329 (0012610948); PRODUCT TYPE: 0195-HEART VALVES; MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21680030&pc=NPT | 2025587-2025-02075 | 2025-02-26 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-24 04:00 | NPT | EVOLUT FX DCS | Patient Device Interaction Problem | Low Blood Pressure/ Hypotension; Perforation | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT PRIOR TO A PROCEDURE INVOLVING THE TRANSCATHETER AORTIC VALVE IN A PATIENT WITH A SEVERELY CALCIFIED AORTIC VALVE, A HORIZONTAL AORTIC ROOT, AND A DILATED ASCENDING AORTA, A PRE-IMPLANT BALLOON AORTIC VALVULOPLASTY WAS PERFORMED USING AN 18 MM NON-MEDTRONIC BALLOON DUE TO CALCIFICATION. THE FIRST VALVE WAS IMPLANTED, BUT WAS UNDER-EXPANDED. A POST-DILATION WAS PERFORMED WITH A 21 MM NON-MEDTRONIC BALLOON. UPON REMOVAL, THE BALLOON CAUGHT ON THE VALVE, CAUSING IT TO DISLODGE INTO THE ASCENDING AORTA AT THE LEVEL OF THE SINOTUBULAR JUNCTION. A SECOND VALVE WAS INSERTED INTO THE PATIENT AND PLACED THROUGH THE DISLODGED VALVE, BUT IT PUSHED THE INFLOW OF THE FIRST VALVE INTO THE AORTA, RESULTING IN A PERFORATION. SUBSEQUENTLY, THE PATIENT'S BLOOD PRESSURE DROPPED. THE SECOND VALVE AND DELIVERY CATHETER SYSTEM WERE WITHDRAWN FROM THE PATIENT. SURGICAL INTERVENTION WAS PERFORMED. THE DISLODGED VALVE WAS EXPLANTED FROM THE PATIENT, A SURGICAL AORTIC VALVE REPLACEMENT WAS PERFORMED, AND THE PERFORATED ASCENDING AORTA WAS REPAIRED. DESPITE THESE EFFORTS, THE PATIENT PASSED AWAY ON THE OPERATING ROOM TABLE.Event Description: IT WAS REPORTED THAT PRIOR TO A PROCEDURE INVOLVING THE TRANSCATHETER AORTIC VALVE IN A PATIENT WITH A SEVERELY CALCIFIED AORTIC VALVE, A HORIZONTAL AORTIC ROOT, AND A DILATED ASCENDING AORTA, A PRE-IMPLANT BALLOON AORTIC VALVULOPLASTY WAS PERFORMED USING AN 18 MM NON-MEDTRONIC BALLOON DUE TO CALCIFICATION. THE FIRST VALVE WAS IMPLANTED, BUT WAS UNDER-EXPANDED. A POST-DILATION WAS PERFORMED WITH A 21 MM NON-MEDTRONIC BALLOON. UPON REMOVAL, THE BALLOON CAUGHT ON THE VALVE, CAUSING IT TO DISLODGE INTO THE ASCENDING AORTA AT THE LEVEL OF THE SINOTUBULAR JUNCTION. A SECOND VALVE WAS INSERTED INTO THE PATIENT AND PLACED THROUGH THE DISLODGED VALVE, BUT IT PUSHED THE INFLOW OF THE FIRST VALVE INTO THE AORTA, RESULTING IN A PERFORATION. SUBSEQUENTLY, THE PATIENT'S BLOOD PRESSURE DROPPED. THE SECOND VALVE AND DELIVERY CATHETER SYSTEM WERE WITHDRAWN FROM THE PATIENT. SURGICAL INTERVENTION WAS PERFORMED. THE DISLODGED VALVE WAS EXPLANTED FROM THE PATIENT, A SURGICAL AORTIC VALVE REPLACEMENT WAS PERFORMED, AND THE PERFORATED ASCENDING AORTA WAS REPAIRED. DESPITE THESE EFFORTS, THE PATIENT PASSED AWAY ON THE OPERATING ROOM TABLE. ADDITIONAL INFORMATION WAS RECEIVED THAT THE PATIENT WAS RAPID PACED DURING POST-IMPLANT DILATION. THE PATIENT HAD A SEVERELY CALCIFIED AORTIC VALVE, HORIZONTAL AORTIC ROOT ANGLE AND DILATED ASCENDING AORTA, WHICH CONTRIBUTED TO THE DISLODGEMENT. THE DEPLOYMENT STARTING POINT WAS AT THE BOTTOM OF THE PIGTAIL CATHETER. PRIOR TO VALVE DISLODGEMENT, THE IMPLANT DEPTH WAS 4 MM ON THE NON-CORONARY CUSP (NCC) AND 5 MM ON THE LEFT CORONARY CUSP (LCC). A MEDTRONIC (CONFIDA) GUIDEWIRE WAS USED DURING THE PROCEDURE. PER THE PHYSICIAN, THE CAUSE OF DEATH WAS PERFORATION OF THE ASCENDING AORTA WHICH WAS NOT RELATED TO THE FIRST DELIVERY CATHETER SYSTEM (DCS), HOWEVER, THE INFLOW CROWNS OF THE DISLODGED VALVE WERE BELIEVED TO CAUSE THE PERFORATION WHEN TRYING TO PASS THE SECOND VALVE THROUGH THE FIRST. THE PHYSICIAN INDICATED THAT THE DIFFICULT ANATOMY PLAYED A FACTOR IN CONTRIBUTING TO THE PATIENT¿S DEATH. Manufacturer Narrative: OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: PRODUCT ID: GWBC30 (LOT: UNKNOWN); PRODUCT TYPE: 0193-GUIDEWIRE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID D-EVOLUTFX-2329 (LOT: 0012610943); PRODUCT TYPE: 0195-HEART VALVES; PRODUCT ID EVFXPLUS-26 (SERIAL: (B)(6)); PRODUCT TYPE: 0195-HEART VALVES; PRODUCT ID EVFXPLUS-26 ((B)(6)); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE (B)(6) 2025; EXPLANT DATE (B)(6) 2025. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21681340&pc=FKX | 0008030665-2025-00687 | 2025-03-06 05:00 | Death | ERIKA DE REYNOSA, S.A. DE C.V. | 2025-03-24 04:00 | FKX | LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | Adverse Event Without Identified Device or Use Problem | Peritonitis | K222318 |  | 1 | Event Description: IT WAS REPORTED THAT THIS PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON CONTINUOUS CYCLIC PERITONEAL DIALYSIS [CC(PD)] FOR RENAL REPLACEMENT THERAPY (RRT) EXPIRED ON 06/MAR/2025 REPORTEDLY DUE TO PERITONITIS AND A CLOSTRIDIUM DIFFICILE (C-DIFF) INFECTION. SUBSEQUENT ATTEMPTS TO OBTAIN ADDITIONAL CLINICAL INFORMATION (E.G., PATIENT DEMOGRAPHICS, HOSPITAL RECORDS, TREATMENT RECORDS, ESRD DEATH NOTIFICATION, DEATH CERTIFICATE, AUTOPSY REPORT, DISCHARGE SUMMARY) WERE UNSUCCESSFUL.Event Description: IT WAS REPORTED THAT THIS PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON CONTINUOUS CYCLIC PERITONEAL DIALYSIS [CC(PD)] FOR RENAL REPLACEMENT THERAPY (RRT) EXPIRED ON (B)(6) 2025 REPORTEDLY DUE TO PERITONITIS AND A CLOSTRIDIUM DIFFICILE (C-DIFF) INFECTION. SUBSEQUENT ATTEMPTS TO OBTAIN ADDITIONAL CLINICAL INFORMATION (E.G., PATIENT DEMOGRAPHICS, HOSPITAL RECORDS, TREATMENT RECORDS, ESRD DEATH NOTIFICATION, DEATH CERTIFICATE, AUTOPSY REPORT, DISCHARGE SUMMARY) WERE UNSUCCESSFUL. Manufacturer Narrative: THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL REVIEW: IT IS UNKNOWN IF A TEMPORAL RELATIONSHIP EXISTS BETWEEN CCPD THERAPY UTILIZING THE LIBERTY SELECT CYCLER AND THE SERIOUS ADVERSE EVENTS OF PERITONITIS, A C-DIFF INFECTION, AND DEATH. THE ETIOLOGY OF THE SERIOUS ADVERSE EVENTS IS UNKNOWN; THEREFORE, CAUSALITY COULD NOT BE ESTABLISHED. ACCORDING TO THE INITIAL REPORTING, THE PATIENT¿S DEATH WAS ATTRIBUTED TO THE PERITONITIS AND C-DIFF INFECTIONS. IT SHOULD BE NOTED THAT LIMITED CLINICAL FOLLOW-UP INFORMATION PRECLUDED A MORE COMPREHENSIVE CLINICAL ASSESSMENT. THOSE INDIVIDUALS UNDERGOING PD FOR RRT (MANUAL OR CYCLER BASED) ARE AT HIGH RISK FOR INFECTIONS OF THE PERITONEUM. ADDITIONALLY, RISK FACTORS ASSOCIATED WITH THE DEVELOPMENT OF C. DIFFICILE INCLUDE LONGER DURATIONS OF ANTIBIOTIC ADMINISTRATION REQUIRED FOR ESRD PATIENTS TO CLEAR THE INFECTION. LASTLY, THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. BASED ON THE INFORMATION AVAILABLE, THE LIBERTY SELECT CYCLER AND LIBERTY CYCLER SET CANNOT BE DISASSOCIATED FROM HAVING A POSSIBLE CAUSAL OR CONTRIBUTORY ROLE IN THE SERIOUS ADVERSE EVENTS. THERE IS CURRENTLY NO ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) DEFICIENCY OR MALFUNCTION CAUSED AND/OR CONTRIBUTED TO THE SERIOUS ADVERSE EVENTS. HOWEVER, GIVEN THE LACK OF PATIENT DATA, TREATMENT DATA, CLINICAL FOLLOW-UP, DISCHARGE SUMMARY, DEATH CERTIFICATE, ESRD DEATH NOTIFICATION AND NO MANUFACTURER EVALUATION OF THE PATIENT¿S DEVICE, THIS CLINICAL INVESTIGATION CANNOT DISASSOCIATE THE PATIENT¿S LIBERTY SELECT CYCLER FROM PLAYING A POSSIBLE ROLE IN THE SERIOUS ADVERSE EVENTS. Manufacturer Narrative: CORRECTION PROVIDED IN G4. ADDITIONAL INFORMATION PROVIDED IN D9 AND H3. PLANT INVESTIGATION: A VISUAL INSPECTION OF THE RETURNED CYCLER EXTERIOR SHOWED NO SIGNS OF PHYSICAL DAMAGE. THERE WERE VISUAL INDICATIONS OF DRIED FLUID WITHIN THE CASSETTE COMPARTMENT ON THE PUMP. THERE WERE NO VISUAL INDICATIONS OF PARTICULATES WITHIN THE CASSETTE AREA. THERE WERE NO BURRS OR SHARP EDGES IN CASSETTE AREA THAT MAY HAVE PUNCTURED A CASSETTE MEMBRANE. AN (AS RECEIVED WITH REDUCED DWELL) SIMULATED TREATMENT WAS PERFORMED AND COMPLETED. VALVE ACTUATION TEST PASSED. SYSTEM AIR LEAK TEST PASSED. AN INTERNAL VISUAL INSPECTION OF THE RETURNED CYCLER WAS PERFORMED. THERE WERE VISUAL INDICATIONS OF DRIED FLUID ON THE BOTTOM COVER BEHIND THE FRONT PANEL ASSEMBLY. THE CAUSE OF THE OBSERVED DRIED FLUID COULD NOT BE DETERMINED. THERE WERE NO DISCREPANCIES ENCOUNTERED WITH THE MUSHROOM HEADS. A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND VERIFIED THAT THE RESULTS OF THE IN-PROGRESS AND FINAL QUALITY CONTROL (QC) TESTING MET ALL REQUIREMENTS. UPON COMPLETION OF THE EVALUATION, THERE WERE NO MALFUNCTIONS THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE CYCLER PERFORMED AS DESIGNED AND AN ASSOCIATED CAUSE COULD NOT BE DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21681389&pc=GCX | 9610617-2025-00491 | 2025-03-19 04:00 | Death | KARL STORZ SE & CO. KG | 2025-03-24 04:00 | GCX | S-PILOT | Power Problem; Unexpected Shutdown | Pulmonary Edema; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | EXEMPT |  | 1 | Event Description: IT WAS REPORTED THAT "THE INITIAL CHECKS WERE DONE, THE S-PILOT WAS WORKING, INTRA-OPERATIVELY POWER COULD NOT BE ESTABLISHED. A PROLONGED MORCELLATION TIME WAS INDICATED. THE WHOLE PROCEDURE WAS PROLONGED MORE THAN 60 MINUTES. AS CONSEQUENCE, FURTHER INTERVENTIONS WERE REQUIRED (CT SCAN, X-RAY AND ITU ADMISSION) AND THEREFORE ADDITIONAL HOSPITALIZATION NECESSARY. IN ADDITION, ON (B)(6) 2025 THE INFORMATION WAS RECEIVED THAT THE PATIENT HAS PASSED AWAY.Event Description: CONFIRMATION WAS RECEIVED ON APRIL 11, 2025, THAT THE KARL STORZ S-PILOT DEVICE DID NOT HAVE ANY INTERVENTION WITH THE REPORTED PATIENT EVENT. Manufacturer Narrative: THE AFFECTED DEVICE HAS BEEN REQUESTED FOR INVESTIGATION BY THE MANUFACTURER. DEVICE WAS NOT RETURNED AT THE TIME OF MDR SUBMISSION. THE EVENT IS FILED UNDER INTERNAL KARL STORZ COMPLAINT ID: (B)(4). Manufacturer Narrative: AN INCORRECT ADDRESS WAS PROVIDED IN SECTION G1 IN THE INITIAL REPORT. THIS SUPPLEMENTAL REPORT IS TO PROVIDE THE CORRECT CONTACT INFORMATION FOR THIS SECTION. ADDITIONAL INFORMATION WAS ADDED IN SECTION B5 - CONFIRMATION WAS RECEIVED ON APRIL 11, 2025, THAT THE KARL STORZ S-PILOT DEVICE DID NOT HAVE ANY INTERVENTION WITH THE REPORTED PATIENT EVENT. DEVICE EVALUATION: THE DEVICE WAS SENT TO THE MANUFACTURER FOR INSPECTION. DURING THE MANUFACTURER'S TECHNICAL INSPECTION, THE ERROR DESCRIPTION PROVIDED BY THE CUSTOMER "SHOWING NO POWER - DURING USE" WAS CONFIRMED. BASED ON THE DEFECTS FOUND DURING THE TECHNICAL INSPECTION IT IS CONCLUDED THAT THE MOST LIKELY CAUSE FOR THE DESCRIBED ERROR IS "OVERCURRENT". SUBSEQUENTLY, THE INTEGRATED FUSE APPLIED ON THE MAINBOARD BLOWN. THE IS STATING THAT THE PRODUCT CAN FAIL DURING USE. THE ABOVE INVESTIGATIONS DID NOT REVEAL A ROOT CAUSE RELATED TO THE LABELLING, THE DEVICE, OR ITS HISTORY. ALL PRODUCTION-RELATED QUALITY CHECKS WERE PASSED, AND NO NON-CONFORMITIES WERE IDENTIFIED IN THE DEVICE'S MANUFACTURING RECORDS. THE LABELLING WAS FOUND TO CONTAIN ADEQUATE INSTRUCTIONS AND WARNINGS. THE COMPLAINT HISTORY DID NOT REVEAL ANY PICKUP IN SIMILAR COMPLAINTS; NO TREND WAS IDENTIFIED. THE EVENT IS FILED UNDER INTERNAL KARL STORZ COMPLAINT ID: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21682609&pc=OZD | 1220648-2025-26920 | 2025-02-26 05:00 | Death | ABIOMED, INC. | 2025-03-24 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT HAD PLACED A IMPELLA CP PUMP FOR THE SUPPORT OF A PATIENT ADMITTED IN WITH ST-SEGMENT ELEVATION MYOCARDIAL INFARCTION AND CODING DURING CARDIAC PERCUTANEOUS CORONARY INTERVENTION. THE CP WAS PLACED AND SUPPORTED FOR 51 HOURS TIL PUMP REMOVED. THE ACCESS SITE WAS OOZING AND BLOOD PRODUCTS WERE GIVEN, MANUAL HOLD AND HEPARIN TITRATION. IT WAS REPORTED THAT CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21684245&pc=QZK | 2017233-2025-05968 | 2025-02-28 05:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-03-24 04:00 | QZK | GORE® EXCLUDER® THORACOABDOMINAL BRANCH ENDOPROSTHESIS | Adverse Event Without Identified Device or Use Problem | Coagulation Disorder | P230023 |  | 1 | Event Description: THE FOLLOWING INFORMATION WAS REPORTED: ON (B)(6) 2025, THE PATIENT UNDERWENT TREATMENT OF AN AORTIC ANEURYSM USING A GORE® EXCLUDER® THORACOABDOMINAL BRANCH ENDOPROSTHESIS (TAMBE). IT WAS REPORTED THAT THE RIGHT RENAL ARTERY WAS CANNULATED AND STENTED, THEN THE LEFT RENAL ARTERY, THEN THE SUPERIOR MESENTERIC ARTERY (SMA) AND LAST THE CELIAC ARTERY. THE PROCEDURE FOR THE TAMBE TOOK APPROXIMATELY THREE HOURS. UPON THE FINAL COMPLETION ANGIOGRAM, IT WAS NOTED THAT THE RENAL ARTERIES DID NOT HAVE FLOW. THE PHYSICIAN CANNULATED THE RENAL ARTERIES AND DISCOVERED BOTH RENAL ARTERIES HAD THROMBUS PRESENT. THE PHYSICIAN PREFORMED A THROMBECTOMY WITH PENUMBRA WHERE SHE REESTABLISHED FLOW. INTRAVASCULAR ULTRASOUND (IVUS) WAS USED TO VISUALIZE THE RENAL ARTERIES TO MAKE SURE THE STENTS WHERE NOT CRUSHED. EVERYTHING LOOKED APPROPRIATE VIA IVUS. DURING THIS TIME, THE PHYSICIAN DISCOVERED THAT ANESTHESIA HAD BEEN STRUGGLING KEEPING THE ACTIVATED CLOTTING TIME (ACT) THERAPEUTIC. WITH ACT¿S BEING BELOW 200. THE PATIENT WAS SWITCHED OVER TO ARGATROBAN, AT THAT POINT THE PATIENT¿S ACT WENT TO THERAPEUTIC LEVELS. THE PATIENT¿S LABS WERE SENT IN TO DETERMINE IF THE PATIENT HAD HEPARIN-INDUCED THROMBOCYTOPENIA (HITT), THAT CAME BACK NEGATIVE. AS THE SHEATH WAS COMING OUT OF THE GROIN, THE PHYSICIAN NOTICED NO PULSES IN THE ACCESSED LEG. UPON FURTHER INVESTIGATION IT WAS DETERMINED THE LEFT LEG HAD THROMBOSED AS WELL. THE PHYSICIAN PREFORMED A CUT DOWN ON THE ACCESS AND DID A THROMBECTOMY WHERE FLOW WAS REESTABLISHED. ON MARCH 04, 2025, THE FIELD SALES ASSOCIATE WAS CALLED BY THE PHYSICIAN AND INFORMED THAT THE PATIENT WAS EXPERIENCING ABDOMINAL PAIN AND HAD A LACTIC ACID OF OVER 6, PATIENT RECEIVED A CTA AND DISCOVERED ALL 4 VISCERAL VESSELS WHERE THROMBOSED. THE PHYSICIAN TOOK THE PATIENT BACK TO THE OPERATING ROOM WHERE A THROMBECTOMY WAS PERFORMED OF THE CELIAC ARTERY AND SMA. FLOW WAS ESTABLISHED HOWEVER BRANCH VESSELS HAD THROMBUS STILL PRESENT. THE PATIENT ULTIMATELY WENT INTO RENAL AND LIVER FAILURE AND PASSED AWAY ON (B)(6)2025. Manufacturer Narrative: PATIENT MEDICATIONS: ALBUTEROL HFA, NORVASC, ELIQUIS, ASPIRIN, ACETAMINOPHEN, ALBUMIN HUMAN 5% INFUSION, CATHFLO ACTIVASE, ARGATROBAN, ARTIFICIAL TEARS OPHTHALMIC SOLUTION, PULMICORT, CALCIUM CHLORIDE, COREG, MEFOXIN, DECADRON, FOLIC-ACID, FENTANYL, APRESOLINE, DILAUDID, OMNIPAQUE, AND TRANDATE. THE GORE® EXCLUDER® THORACOABDOMINAL BRANCH ENDOPROSTHESIS (TAMBE) INSTRUCTIONS FOR USE (IFU) STATES THAT ADVERSE EVENTS THAT MAY OCCUR AND/OR REQUIRE INTERVENTION INCLUDE BUT ARE NOT LIMITED TO ARTERIAL OR VENOUS THROMBOSIS AND OCCLUSION OF DEVICE OR NATIVE VESSEL. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. Manufacturer Narrative: CORRECTION SECTION D. Manufacturer Narrative: CORRECTION TO H1, FURTHER DETERMINATION, THIS WILL BE DEEMED A REPORTABLE DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21684364&pc=FKX | 0008030665-2025-00688 | 2025-03-06 05:00 | Death | ERIKA DE REYNOSA, S.A. DE C.V. | 2025-03-24 04:00 | FKX | LIBERTY CYCLER SET, SINGLE CONN./EXT. DL | Adverse Event Without Identified Device or Use Problem | Peritonitis | K173718 |  | 1 | Event Description: IT WAS REPORTED THAT THIS PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON CONTINUOUS CYCLIC PERITONEAL DIALYSIS [CC(PD)] FOR RENAL REPLACEMENT THERAPY (RRT) EXPIRED ON (B)(6) 2025 REPORTEDLY DUE TO PERITONITIS AND A CLOSTRIDIUM DIFFICILE (C-DIFF) INFECTION. SUBSEQUENT ATTEMPTS TO OBTAIN ADDITIONAL CLINICAL INFORMATION (E.G., PATIENT DEMOGRAPHICS, HOSPITAL RECORDS, TREATMENT RECORDS, ESRD DEATH NOTIFICATION, DEATH CERTIFICATE, AUTOPSY REPORT, DISCHARGE SUMMARY) WERE UNSUCCESSFUL.Event Description: IT WAS REPORTED THAT THIS PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON CONTINUOUS CYCLIC PERITONEAL DIALYSIS [CC(PD)] FOR RENAL REPLACEMENT THERAPY (RRT) EXPIRED ON (B)(6) 2025 REPORTEDLY DUE TO PERITONITIS AND A CLOSTRIDIUM DIFFICILE (C-DIFF) INFECTION. SUBSEQUENT ATTEMPTS TO OBTAIN ADDITIONAL CLINICAL INFORMATION (E.G., PATIENT DEMOGRAPHICS, HOSPITAL RECORDS, TREATMENT RECORDS, ESRD DEATH NOTIFICATION, DEATH CERTIFICATE, AUTOPSY REPORT, DISCHARGE SUMMARY) WERE UNSUCCESSFUL. Manufacturer Narrative: THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL REVIEW: IT IS UNKNOWN IF A TEMPORAL RELATIONSHIP EXISTS BETWEEN CCPD THERAPY UTILIZING THE LIBERTY SELECT CYCLER AND THE SERIOUS ADVERSE EVENTS OF PERITONITIS, A C-DIFF INFECTION, AND DEATH. THE ETIOLOGY OF THE SERIOUS ADVERSE EVENTS IS UNKNOWN; THEREFORE, CAUSALITY COULD NOT BE ESTABLISHED. ACCORDING TO THE INITIAL REPORTING, THE PATIENT¿S DEATH WAS ATTRIBUTED TO THE PERITONITIS AND C-DIFF INFECTIONS. IT SHOULD BE NOTED THAT LIMITED CLINICAL FOLLOW-UP INFORMATION PRECLUDED A MORE COMPREHENSIVE CLINICAL ASSESSMENT. THOSE INDIVIDUALS UNDERGOING PD FOR RRT (MANUAL OR CYCLER BASED) ARE AT HIGH RISK FOR INFECTIONS OF THE PERITONEUM. ADDITIONALLY, RISK FACTORS ASSOCIATED WITH THE DEVELOPMENT OF C. DIFFICILE INCLUDE LONGER DURATIONS OF ANTIBIOTIC ADMINISTRATION REQUIRED FOR ESRD PATIENTS TO CLEAR THE INFECTION. LASTLY, THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. BASED ON THE INFORMATION AVAILABLE, THE LIBERTY SELECT CYCLER AND LIBERTY CYCLER SET CANNOT BE DISASSOCIATED FROM HAVING A POSSIBLE CAUSAL OR CONTRIBUTORY ROLE IN THE SERIOUS ADVERSE EVENTS. THERE IS CURRENTLY NO ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) DEFICIENCY OR MALFUNCTION CAUSED AND/OR CONTRIBUTED TO THE SERIOUS ADVERSE EVENTS. HOWEVER, GIVEN THE LACK OF PATIENT DATA, TREATMENT DATA, CLINICAL FOLLOW-UP, DISCHARGE SUMMARY, DEATH CERTIFICATE, ESRD DEATH NOTIFICATION AND NO MANUFACTURER EVALUATION OF THE PATIENT¿S DEVICE, THIS CLINICAL INVESTIGATION CANNOT DISASSOCIATE THE PATIENT¿S LIBERTY SELECT CYCLER FROM PLAYING A POSSIBLE ROLE IN THE SERIOUS ADVERSE EVENTS. Manufacturer Narrative: PLANT INVESTIGATION: THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER AND THE LOT NUMBER WAS NOT PROVIDED. A MANUFACTURING REVIEW WAS PERFORMED ON THE PRODUCTS SHIPPED TO THE PATIENT FOR THE THREE (3) MONTH TIME FRAME WHICH IMMEDIATELY PRECEDED THE EVENT OCCURRENCE DATE. THIS REVIEW INCLUDED THE LOT NUMBERS FOR ALL FRESENIUS LIBERTY CYCLER SETS SHIPPED TO THIS ACCOUNT WITHIN THE SELECTED TIME FRAME. THE ENTIRE SET OF LOTS HAVE BEEN SOLD AND DISTRIBUTED. THERE WERE NO NON-CONFORMANCES OR ABNORMALITIES IDENTIFIED DURING THE MANUFACTURING PROCESS WHICH COULD BE ASSOCIATED WITH THE REPORTED EVENT. AN INVESTIGATION OF THE DEVICE HISTORY RECORDS (DHR) WAS CONDUCTED AND CONFIRMED THAT THE RESULTS OF THE IN-PROGRESS AND FINAL QUALITY CONTROL (QC) TESTING MET ALL REQUIREMENTS. THE PRODUCT LOTS INVOLVED MET ALL SPECIFICATIONS FOR RELEASE. A REVIEW OF THE DHR DID NOT REVEAL A PROBABLE CAUSE FOR THE CUSTOMER COMPLAINT. AS A PHYSICAL EVALUATION COULD NOT BE PERFORMED, A DEFINITIVE CONCLUSION REGARDING THE REPORTED INCIDENT COULD NOT BE REACHED AND A CAUSE COULD NOT BE CONFIRMED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21684795&pc=LPB | 2029046-2025-00895 | 2024-03-04 05:00 | Death | BIOSENSE WEBSTER INC | 2025-03-24 04:00 | LPB | THERMOCOOL SMARTTOUCH SF | Adverse Event Without Identified Device or Use Problem | Fistula | P030031 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A LEFT ATRIAL TACHYCARDIA ABLATION WITH A THERMOCOOL SMARTTOUCH SF AND A NGEN GENERATOR AND THE PATIENT EXPERIENCED ESOPHAGEAL FISTULA THAT RESULTED IN DEATH. ROUGHLY ONE MONTH AFTER AN ATRIAL FIBRILLATION PROCEDURE WAS PERFORMED, AN ESOPHAGEAL FISTULA WAS NOTICED. THE PATIENT CALLED THE OFFICE AFTER FEELING CHEST PAINS ROUGHLY ONE MONTHS AFTER THE PROCEDURE. THE PATIENT PASSED AWAY ON THEIR WAY TO THE EMERGENCY ROOM (ER). IT WAS UNKNOWN HOW THE INJURY WAS CONFIRMED OR IF ANY MEDICAL INTERVENTION WAS PROVIDED FOR THE PATIENT. THERE WERE NO COMPLICATIONS DURING THE INITIAL PROCEDURE. THEY BELIEVE THE INJURY WAS CAUSED BY BURNING AT HIGH POWER IN THE POSTERIOR WALL USING A RADIO FREQUENCY CATHETER. IT WAS REPORTED THAT THEY WERE "NO LONGER A CLINICAL AT THIS ACCOUNT AND IS NOW DOING CLINICAL THERAPY ADVANCEMENT". OCTARAY CATHETER, SOUNDSTAR CATHETER AND A THERMOCOOL SMARTTOUCH SF CATHETER WERE USED DURING THE PROCEDURE. THE PROCEDURE WAS PERFORMED ROUGHLY 12 MONTHS AGO, THE EVENT WAS REPORTED NOW BECAUSE THEY DID NOT KNOW IT WAS PREVIOUSLY REPORTED AND THEY JUST FOUND OUT ABOUT THE EVENT. THE EVENT WAS JUST FOUND OUT AND IT WAS REPORTED THAT THE INFORMATION ABOUT THIS EVENT WAS GOING TO BE SPARCE SINCE THEY FOUND OUT SO LATE. Manufacturer Narrative: DEVICE INVESTIGATION DETAILS: AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. AN EVALUATION OF THE MANUFACTURING RECORD COULD NOT BE PERFORMED AS THE REQUIRED PRODUCT IDENTIFICATION NUMBER WAS NOT PROVIDED TO COMPLETE THE EVALUATION. HOWEVER, IF THE PRODUCT OR PRODUCT ID NUMBERS ARE RECEIVED AT A LATER DATE, THE INVESTIGATION WILL BE UPDATED AS APPLICABLE. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. NOTE FOR FIELD D4: UDI: AS THE LOT NUMBER FOR THE DEVICE INVOLVED IN THE EVENT WAS NOT PROVIDED, THE FULL UDI IS CURRENTLY NOT AVAILABLE. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY BIOSENSE WEBSTER INC., OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, BIOSENSE WEBSTER INC., OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. MANUFACTURER'S REF. # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21684817&pc=LPB | 2029046-2025-00896 | 2024-03-04 05:00 | Death | BIOSENSE WEBSTER INC | 2025-03-24 04:00 | LPB | NGEN GENERATOR | Adverse Event Without Identified Device or Use Problem | Fistula | P030031 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A LEFT ATRIAL TACHYCARDIA ABLATION WITH A THERMOCOOL SMARTTOUCH SF AND A NGEN GENERATOR AND THE PATIENT EXPERIENCED ESOPHAGEAL FISTULA THAT RESULTED IN DEATH. ROUGHLY ONE MONTH AFTER AN ATRIAL FIBRILLATION PROCEDURE WAS PERFORMED, AN ESOPHAGEAL FISTULA WAS NOTICED. THE PATIENT CALLED THE OFFICE AFTER FEELING CHEST PAINS ROUGHLY ONE MONTHS AFTER THE PROCEDURE. THE PATIENT PASSED AWAY ON THEIR WAY TO THE EMERGENCY ROOM (ER). IT WAS UNKNOWN HOW THE INJURY WAS CONFIRMED OR IF ANY MEDICAL INTERVENTION WAS PROVIDED FOR THE PATIENT. THERE WERE NO COMPLICATIONS DURING THE INITIAL PROCEDURE. THEY BELIEVE THE INJURY WAS CAUSED BY BURNING AT HIGH POWER IN THE POSTERIOR WALL USING A RADIO FREQUENCY CATHETER. IT WAS REPORTED THAT THEY WERE "NO LONGER A CLINICAL AT THIS ACCOUNT AND IS NOW DOING CLINICAL THERAPY ADVANCEMENT". OCTARAY CATHETER, SOUNDSTAR CATHETER AND A THERMOCOOL SMARTTOUCH SF CATHETER WERE USED DURING THE PROCEDURE. THE PROCEDURE WAS PERFORMED ROUGHLY 12 MONTHS AGO, THE EVENT WAS REPORTED NOW BECAUSE THEY DID NOT KNOW IT WAS PREVIOUSLY REPORTED AND THEY JUST FOUND OUT ABOUT THE EVENT. THE EVENT WAS JUST FOUND OUT AND IT WAS REPORTED THAT THE INFORMATION ABOUT THIS EVENT WAS GOING TO BE SPARCE SINCE THEY FOUND OUT SO LATE. Manufacturer Narrative: IT WAS REPORTED THAT A PATIENT UNDERWENT A LEFT ATRIAL TACHYCARDIA ABLATION WITH A THERMOCOOL SMARTTOUCH SF AND A NGEN GENERATOR AND THE PATIENT EXPERIENCED ESOPHAGEAL FISTULA THAT RESULTED IN DEATH. ROUGHLY ONE MONTH AFTER AN ATRIAL FIBRILLATION PROCEDURE WAS PERFORMED, AN ESOPHAGEAL FISTULA WAS NOTICED. THE PATIENT CALLED THE OFFICE AFTER FEELING CHEST PAINS ROUGHLY ONE MONTHS AFTER THE PROCEDURE. THE PATIENT PASSED AWAY ON THEIR WAY TO THE EMERGENCY ROOM (ER). IT WAS UNKNOWN HOW THE INJURY WAS CONFIRMED OR IF ANY MEDICAL INTERVENTION WAS PROVIDED FOR THE PATIENT. THERE WERE NO COMPLICATIONS DURING THE INITIAL PROCEDURE. THEY BELIEVE THE INJURY WAS CAUSED BY BURNING AT HIGH POWER IN THE POSTERIOR WALL USING A RADIO FREQUENCY CATHETER. IT WAS REPORTED THAT THEY WERE "NO LONGER A CLINICAL AT THIS ACCOUNT AND IS NOW DOING CLINICAL THERAPY ADVANCEMENT". DEVICE INVESTIGATION DETAILS: TECHNICAL SERVICES PERFORMED A REMOTE EVALUATION OF THE DEVICE, HOWEVER, THE SERVICE WAS DECLINED. THE CUSTOMER STATED THE SYSTEM HAS BEEN FUNCTIONAL AND IS DECLINING SERVICE. A DEVICE HISTORY RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE NUMBER, AND NO INTERNAL ACTIONS RELATED TO THE REPORTED COMPLAINT CONDITION WERE IDENTIFIED. ALL DEVICES ARE MANUFACTURED, INSPECTED, AND RELEASED TO APPROVED SPECIFICATIONS AS PART OF JOHNSON & JOHNSON MEDTECH'S QUALITY PROCESS. IF ADDITIONAL INFORMATION IS RECEIVED REGARDING THIS EVENT, A SUPPLEMENTAL 3500A REPORT WILL BE SUBMITTED TO THE FDA. MANUFACTURER'S REF. # (B)(4). Manufacturer Narrative: THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY BIOSENSE WEBSTER INC., OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, BIOSENSE WEBSTER INC., OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. NOTE: PER FDA REQUEST, MDR SUBMISSIONS FOR THE NGEN RF GENERATOR ARE TO BE REPORTED WITH PMA DETAILS OF THE CATHETER USED ALONG WITH THE NGEN RF GENERATOR. MANUFACTURER'S REF. # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21684833&pc=EOQ | 2955842-2025-04259 | 2025-02-04 05:00 | Death | INTUITIVE SURGICAL, INC | 2025-03-24 04:00 | EOQ | ION | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | K182188 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNEVENTFUL ION ENDOBRONCHIAL LUNG BIOPSY WITH MEDIASTINAL STAGING PROCEDURE. THE PHYSICIAN REPORTED THAT THERE WERE NO INTRA-PROCEDURAL COMPLICATIONS OR MALFUNCTIONS WITH THE ION SYSTEM. APPROXIMATELY 30 MINUTES POST-PROCEDURE, THE PATIENT SHOWED ACUTE STROKE SYMPTOMS OF LEFT-SIDED WEAKNESS AND GAZE DEVIATION TREATED WITH TISSUE PLASMINOGEN ACTIVATOR (TPA) THAT RESULTED IN EPISODES OF HEMOPTYSIS. THE PHYSICIAN STATED THAT THE PATIENT SUFFERED AN ISCHEMIC STROKE; UNSPECIFIED HEAD IMAGES DID NOT SHOW ANY ACUTE CHANGES. THE PATIENT EXPIRED SHORTLY THEREAFTER. THE CAUSE OF THE STROKE WAS REPORTED TO BE RELATED TO THE PATIENT¿S MEDICAL HISTORY, INCLUDING SUBSTANTIAL CARDIOVASCULAR DISEASE WITH MULTIPLE STENTS AND MULTIPLE STROKES TREATED WITH ASPIRIN AND PLAVIX THERAPY. THE PLAVIX HAD BEEN HELD FOR 5 DAYS PRIOR TO THE PROCEDURE, BUT THE ASPIRIN THERAPY WAS CONTINUED. Manufacturer Narrative: A REVIEW OF THE SYSTEM LOGS FOR THE REPORTED PROCEDURE DATE SHOWED THERE WERE NO RELATED SYSTEM ERRORS DURING THE PROCEDURE THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT AN ELDERLY PATIENT WITH MULTIPLE MEDICAL CO-MORBIDITIES INCLUDING MULTIPLE PRIOR STROKES, CARDIOVASCULAR DISEASE REQUIRING MULTIPLE VASCULAR STENTS UNDERWENT AN ION LUNG BIOPSY. PLAVIX WAS HELD FOR 5 DAYS IN PREPARATION FOR THE BIOPSY. THE PROCEDURE WAS COMPLETED WITHOUT ISSUE AND THE PATIENT WAS TRANSPORTED TO THE PACU. ABOUT 30 MINUTES LATER THE PATIENT DEVELOPED A LEFT SIDED WEAKNESS AND GAZE DEVIATION PROMPTING TPA TO TREAT AN ACUTE STROKE. WORKUP WAS CONSISTENT WITH AN ISCHEMIC STROKE AND THE PATIENT DIED LATER THE SAME DAY. BASED ON THE AVAILABLE DATA THE EVENT WAS PROCEDURE RELATED IN A PATIENT WITH MULTIPLE PRIOR STROKES WHOSE PLAVIX HAD BEEN STOPPED FOR THE BIOPSY AND WAS NOT DEVICE RELATED. BRONCHOSCOPY AND BIOPSY IS A MINIMALLY INVASIVE PROCEDURE WITH A LOW COMPLICATION RATE AND RARELY ASSOCIATED WITH FATAL COMPLICATIONS. IN A MULTICENTER PROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES THE TOTAL NUMBER OF ANY COMPLICATIONS WAS REPORTED TO BE 227 (1.08%) WITH 5 (0.02%) TRANSIENT ISCHEMIC ATTACKS AND 4 (0.02%) TOTAL DEATHS BUT NO STROKES. ANOTHER MULTICENTER STUDY REPORTED 13 (2.2%) COMPLICATIONS WITH NO STROKES AND NO DEATHS ASSOCIATED WITH 581 CASES. A RECENT PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 BRONCHOSCOPIES REPORTED 1 ASSOCIATED DEATH (0.08%) AND NO STROKES. ANOTHER PROSPECTIVE SERIES OF 415 ION PROCEDURES REPORTED 1 CVA (0.2%). THERE WAS NO ALLEGATION OF A MALFUNCTION OF THE ION SYSTEM, INSTRUMENT OR ACCESSORIES ASSOCIATED WITH THE REPORTED COMPLICATION. THE RISK OF STROKE ASSOCIATED WITH GENERAL ANESTHESIA HAS BEEN REPORTED TO VARY FROM 0.1-1.9% IN NON-CARDIAC, NON-NEUROLOGIC, AND NON-MAJOR SURGERY AND AS HIGH AS 10% IN THE SETTING OF BRAIN OR HIGH-RISK CARDIOVASCULAR SURGERY. FACCIOLONGO N, PATELLI M, GASPARINI S, ET AL. INCIDENCE OF COMPLICATIONS IN BRONCHOSCOPY. MULTICENTRE PROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES. MONALDI ARCHIVES FOR CHEST DISEASE. 2009;71(1). OST DE, ERNST A, LEI X, ET AL. DIAGNOSTIC YIELD AND COMPLICATIONS OF BRONCHOSCOPY FOR PERIPHERAL LUNG LESIONS. RESULTS OF THE AQUIRE REGISTRY. AM J RESPIR CRIT CARE MED. 2016;193(1):68-77. FOLCH EE, PRITCHETT MA, NEAD MA, ET AL. ELECTROMAGNETIC NAVIGATION BRONCHOSCOPY FOR PERIPHERAL PULMONARY LESIONS: ONE-YEAR RESULTS OF THE PROSPECTIVE, MULTICENTER NAVIGATE STUDY. JOURNAL OF THORACIC ONCOLOGY. 2019. BROWNLEE AR, WATSON JJJ, AKHMEROV A, ET AL. ROBOTIC NAVIGATIONAL BRONCHOSCOPY IN A THORACIC SURGICAL PRACTICE: LEVERAGING TECHNOLOGY IN THE MANAGEMENT OF PULMONARY NODULES. JTCVS. 2023. BATEMAN BT, SCHUMACHER HC, WANG S, SHAEFI S, BERMAN MF. PERIOPERATIVE ACUTE ISCHEMIC STROKE IN NONCARDIAC AND NONVASCULAR SURGERY: INCIDENCE, RISK FACTORS, AND OUTCOMES. ANESTHESIOLOGY. 2009. SELIM M. PERIOPERATIVE STROKE. NEW ENGLAND JOURNAL OF MEDICINE. 2007. SECTION A2: AGE REPORTED AS "IN THEIR 80'S."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21673431&pc=BZD | 2518422-2025-103309 | 2025-03-10 04:00 | Death | RESPIRONICS, INC. | 2025-03-23 04:00 | BZD | DREAMSTATION | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A CUSTOMER REQUESTING INFORMATION ON RETURNING A DEVICE DUE TO A PATIENT PASSING. THERE IS AN ALLEGATION THAT THE DEVICE IS NOT WORKING CORRECTLY. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21673459&pc=BZD | 2518422-2025-103313 | 2022-04-03 04:00 | Death | RESPIRONICS, INC. | 2025-03-23 04:00 | BZD | DREAMSTATION | Degraded | Unspecified Kidney or Urinary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE DEVICE WAS NOT WORKING CORRECTLY. THE PATIENT HAD KIDNEY FAILURE AND PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21673468&pc=BZD | 2518422-2025-103315 | 2024-07-25 04:00 | Death | RESPIRONICS, INC. | 2025-03-23 04:00 | BZD | DS2ADV AUTO CPAP W/HUMID+HT CELL/BT | Adverse Event Without Identified Device or Use Problem | Unspecified Heart Problem | K200480 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A DREAMSTATION 2 DEVICE. THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WAS IN THE HOSPITAL AND PASSED AWAY FROM A HEART ISSUE. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21673668&pc=LWR | 3004478276-2025-00129 |  | Death | CORCYM CANADA CORP. | 2025-03-23 04:00 | LWR | PERCEVAL SUTURELESS AORTIC HEART VALVE | Insufficient Information | Insufficient Information | P150011 |  | 1 | Event Description: RECEIVED THE FOLLOWING INFORMATION FROM PATIENT TRACKING SYSTEM. A PERCEVAL VALVE PVS23 WAS IMPLANTED IN A PATIENT ON (B)(6) 2024. THE PATIENT PASSED AWAY AT UNKNOWN DATE DUE TO UNKNOWN REASON. NO FURTHER INFORMATION IS AVAILABLE AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21674522&pc=LJT | 1317056-2025-00080 | 2023-02-18 05:00 | Death | ANGIODYNAMICS, INC | 2025-03-23 04:00 | LJT | BIOFLO VORTEX | Adverse Event Without Identified Device or Use Problem | Muscle Weakness; Sepsis | K190559 |  | 1 | Event Description: ON OR ABOUT (B)(6) 2023, DECEDENT G.W. UNDERWENT PLACEMENT OF AN ANGIODYNAMICS SMARTPORT+ PRODUCT, DEVICE IDENTIFICATION NUMBER (B)(4), AND CATALOGUE NUMBER H787CT80LTBAVI0. THE DEVICE WAS IMPLANTED BY DR. MARSHALL L. DINES, M.D., AT ADVOCATE (B)(6, ILLINOIS, FOR THE PURPOSE OF PROVIDING CHEMOTHERAPY. ON OR ABOUT (B)(6) 2023, DECEDENT G.W. PRESENTED HERSELF TO THE EMERGENCY DEPARTMENT OF ADVOCATE (B)(6) HOSPITAL ILLINOIS WITH A CHIEF COMPLAINT OF WEAKNESS, WHERE SHE WAS DIAGNOSED WITH SEPSIS. DECEDENT G.W. WAS PLACED IN THE ICU ON (B)(6) 2023 AND WAS PRONOUNCED DECEASED THE SAME DAY. Manufacturer Narrative: THE CUSTOMER'S REPORTED COMPLAINT DESCRIPTION OF PATIENT SEPSIS INFECTION CANNOT BE CONFIRMED GIVEN THE PATIENT CENTRIC NATURE OF THIS SERIOUS ADVERSE EVENT (SAE). NO PORT DEVICE WAS RETURNED FOR EVALUATION SINCE THERE WAS NO REPORT OF PORT DEVICE MALFUNCTION, DAMAGE OR PERFORMANCE ISSUE DURING USE. A DEVICE HISTORY RECORD (DHR) REVIEW OF THE SHIP HISTORY REPORT (SHR) PACKAGING LOTS REVEALED NO QUALITY RELATED ISSUES OR MANUFACTURING DEFICIENCIES AT THE TIME OF MANUFACTURE. STERILIZATION LOAD RELEASE RECORDS WERE REVIEWED FOR THE PACKAGING LOTS IN QUESTION; NO ISSUES WERE OBSERVED. LABELING REVIEW: THE DIRECTION FOR USE (DFU) PROVIDED WITH THE PORT DEVICE CONTAINS THE FOLLOWING STATEMENTS: CONTRAINDICATIONS: PRESENCE OF INFECTION, BACTEREMIA, OR SEPTICEMIA. WARNINGS: THE DEVICE IS TO BE IMPLANTED, USED, MAINTAINED AND REMOVED IN ACCORDANCE WITH CURRENT EVIDENCE-BASED GUIDANCE FOR INFECTION PREVENTION FROM AGENCIES/ORGANIZATIONS SUCH AS CENTERS OF DISEASE CONTROL (CDC) OR THE WORLD HEALTH ORGANIZATION. IT IS RECOMMENDED THAT INSTITUTIONAL POLICIES AND PROCEDURES ARE UPDATED PERIODICALLY TO REFLECT CURRENT EVIDENCE-BASED GUIDELINES FOR INFECTION PREVENTION. CONTAMINATION OF THE DEVICE MAY LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: CAREFULLY READ AND FOLLOW ALL INSTRUCTIONS PRIOR TO USE AFTER IMPLANTATION OR USAGE OF THE PORT, THE SYSTEM SHOULD BE FLUSHED TO CLEAR INFUSATES AND/OR BLOOD COMPONENTS IN ORDER TO MAINTAIN PATENCY. IF MORE THAN ONE DRUG IS TO BE ADMINISTERED, BETWEEN THE INDIVIDUAL DRUG APPLICATIONS, FLUSH THE SYSTEM WITH 5 TO 10 ML NORMAL SALINE FOR INJECTION TO PREVENT DRUG INTERACTION. POTENTIAL COMPLICATIONS: BACTEREMIA, CATHETER THROMBOSIS, DEATH, INFLAMMATION, INFECTION, THROMBOEMBOLISM, THROMBOPHLEBITIS, TUNNEL INFECTION, VASCULAR THROMBOSIS. A REVIEW OF THE ANGIODYNAMICS COMPLAINT SYSTEM NOTED NO ADVERSE TRENDS FOR THIS COMPLAINT TYPE AND PRODUCT FAMILY. THIS TYPE OF COMPLAINT CONTINUES TO BE MONITORED FOR TRENDS. REFERENCE (B)(4). Manufacturer Narrative: THE DEVICE IS NOT AVAILABLE TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AN INVESTIGATION INTO THE ROOT CAUSE FOR THE EVENT IS CURRENTLY IN PROGRESS. THE RESULTS OF THE INVESTIGATION WILL BE SENT VIA A FOLLOW UP MEDWATCH. REFERENCE (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21674595&pc=NWX | 3003775186-2025-00069 | 2025-03-03 05:00 | Death | ORBUSNEICH MEDICAL (SHENZHEN) CO., LTD | 2025-03-23 04:00 | NWX | SCOREFLEX NC SCORING PTCA CATHETER | Adverse Event Without Identified Device or Use Problem | Thrombosis/Thrombus | P200041 |  | 1 | Event Description: IT WAS REPORTED THAT THE PROCEDURE WAS PERCUTANEOUS CORONARY INTERVENTION (PCI) ON OSTIAL LEFT ANTERIOR DESCENDING (LAD)/D LEFT MAIN (LM) ARTERY. UPON GETTING ON THE TABLE, THE PATIENT BEGAN TO HAVE CONSTRICTIVE PERICARDITIS (CP). WHEN THE PHYSICIAN WAS CANULATING VESSEL WITH A NON-ABBOTT GUIDE CATHETER, THE PATIENT BEGAN TO HAVE ST ELEVATIONS. THE VESSEL WAS WIRED WITH CONTINUING ST ELEVATIONS, AND THE PHYSICIAN REQUESTED A 2.0X12 COMPLIANT BALLOON. HOWEVER, A 4.0X15 SCOREFLEX BALLOON WAS USED AND INFLATED TO 12 ATMOSPHERES. TWO INFLATIONS WERE COMPLETED, AND NO INITIAL ANGIOGRAM WAS TAKEN PRIOR TO INFLATION. ANGIOGRAPHY WAS PERFORMED AFTER TWO INFLATIONS AND ALL VESSELS WERE SHUT DOWN EXCEPT FOR LM AND DIAGONAL BRANCH. PATIENT WENT INTO VENTRICULAR FIBRILLATION AND SHOCKED AT 200J. A NON-ABBOTT BALLOON WAS USED DOWN THE VESSEL. THE PHYSICIAN ATTEMPTED TO WIRE THE CIRCUMFLEX BUT WAS UNABLE TO. THE PATIENT LOST PULSE AND BEGAN TO CODE. IT WAS UNABLE TO REGAIN FLOW TO VESSEL AND OBTAIN RETURN OF SPONTANEOUS CIRCULATION (ROSC). THE PATIENT EXPIRED. IN THE OPINION OF THE PHYSICIAN, THE CAUSE OF THE DEATH WAS THROMBUS IN THE LM THAT EMBOLIZED WHEN THE VESSEL WAS CANNULATED. NO RETURNED DEVICE OR NO OTHER INFORMATION WAS PROVIDED FOR ANALYSIS. RECEIVED THE FOLLOWING ADDITIONAL INFORMATION: LOT NUMBER CANNOT BE OBTAINED. THERE WAS NOTHING WRONG WITH THE SCOREFLEX NC WHEN IT WAS USED. MD STATES THAT BALLOON COULD HAVE CONTRIBUTED TO THROMBUS EMBOLIZING, BUT HE STATES THAT ANY BALLOON WOULD HAVE PROGRESSED THE SITUATION. PT HAD ST ELEVATIONS, PRIOR TO ENGAGING THE VESSEL, ST ELEVATIONS DOES MEAN THAT THE PATIENT WAS BEGINNING TO INFARCT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21671434&pc=DSQ | 2916596-2025-01629 | 2022-01-01 05:00 | Death | THORATEC CORPORATION | 2025-03-22 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE 'IMPACT OF PREOPERATIVE IMPELLA® ON LEFT VENTRICULAR ASSIST DEVICE PATIENTS¿ OUTCOMES' THAT THIS STUDY AIMED TO COMPARE THE MORTALITY AND POST IMPLANTATION COMPLICATIONS OF LVAD IMPLANTATIONS, BOTH SCHEDULED AND EMERGENT DUE TO CS, TREATED WITH IMPELLA® OR OTHER CIRCULATORY SUPPORT DEVICES. THE RETROSPECTIVE STUDY WAS CONDUCTED FROM 2010 TO 2022 AND INCLUDED 119 PATIENTS WITH 45% (N=54) USING A HEARTMATE 3. THEY WERE CATEGORIZED INTO THREE GROUPS BASED ON THEIR INTERMACS PROFILES: THE ¿IMPELLA®¿ GROUP (=3) EITHER USED IMPELLA® EXCLUSIVELY OR WERE WEANED OFF IT LESS THAN 8 DAYS BEFORE IMPLANTATION; THE ¿ECLS/NO DEVICE¿ GROUP (= 3) RECEIVED OTHER CIRCULATORY SUPPORTS; AND THE ¿INTERMACS>3¿ GROUP, AN AMBULATORY COHORT SCHEDULED FOR LVAD IMPLANTATION. IN THE IMPELLA® GROUP, THERE WAS A SIGNIFICANT REDUCTION IN THE INCIDENCE OF POSTOPERATIVE RIGHT VENTRICULAR FAILURE (P=0.024), SEVERE RIGHT VENTRICULAR FAILURE REQUIRING CIRCULATORY SUPPORT (P=0.004), AND MASSIVE BLOOD TRANSFUSIONS (P=0.011). PATIENTS IN THE IMPELLA® GROUP EXHIBITED BETTER ONE-YEAR SURVIVAL COMPARED TO THOSE IN THE ECLS/NO DEVICE GROUP (LOG RANK; P=0.040) AND SIMILAR TO THOSE IN THE INTERMACS>3 PATIENT COHORT. THE USE OF IMPELLA® DURING CS AS A BRIDGE TO LVAD IMPLANTATION SIGNIFICANTLY REDUCES POSTOPERATIVE RIGHT VENTRICULAR DYSFUNCTIONS AND MAY IMPROVE THE ONE-YEAR PROGNOSIS. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEMS (LVAS) AND THE REPORTED EVENTS COULD NOT CONCLUSIVELY BE ESTABLISHED THROUGH THIS EVALUATION. THE SERIAL NUMBERS OF THE HEARTMATE 3 LEFT VENTRICULAR SYSTEMS IN USE WERE NOT PROVIDED. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THIS DOCUMENT LISTS THE ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SECTIONS A AND D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. IMPROVING HEMODYNAMICS IN THE ISCHEMIC HEART, INSERM U1034, PESSAC, FRANCE; SERVICE D¿ANESTHÉSIE-RÉANIMATION CARDIOVASCULAIRE, CHU HAUT-LÉVÊQUE, PESSAC, FRANCE; SERVICE DE TRAITEMENT DE L¿INSUFFISANCE CARDIAQUE, CHU HAUT-LÉVÊQUE, PESSAC, FRANCE; SERVICE DE CHIRURGIE CARDIOVASCULAIRE, CHU HAUT-LÉVÊQUE, PESSAC, FRANCE. BEURTON, A., LAGARDERE, Y., NUBRET, K., FAURE, M., CUKO, B., PERNOT, M., IMBAULT, J., & OUATTARA, A. (N.D.). IMPACT OF PREOPERATIVE IMPELLA® ON LEFT VENTRICULAR ASSIST DEVICE PATIENTS¿ OUTCOMES. ARCHIVES OF CARDIOVASCULAR DISEASES, VOLUME 118, ISSUE 1, SUPPLEMENT, 2025, PAGE S151, ISSN 1875-2136. HTTPS://DOI.ORG/10.1016/J.ACVD.2024.10.289. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21671456&pc=MVK | 3008642652-2025-02995 | 2025-01-24 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-03-22 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED SIX INAPPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 21:30:00 ON (B)(6) 2025. AT 03:04:56 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH MOTION ARTIFACT. AT 03:05:32, THE PATIENT RECEIVED THE FIRST INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON-LIFE-SUSTAINING RHYTHM, WITH MOTION ARTIFACT AND ELECTRODE LEAD FALLOFF. AT 03:12:23, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH MOTION/TACTILE ARTIFACT. AT 03:13:51, THE PATIENT RECEIVED THE SECOND INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH MOTION/TACTILE ARTIFACT. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON-LIFE-SUSTAINING RHYTHM, WITH MOTION/TACTILE ARTIFACT. AT 03:19:32, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH MOTION/TACTILE ARTIFACT. AT 03:20:05, THE PATIENT RECEIVED THE THIRD INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH TACTILE ARTIFACT. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON-LIFE-SUSTAINING RHYTHM, WITH TACTILE ARTIFACT. AT 03:20:30, THE PATIENT RECEIVED THE FOURTH INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH TACTILE ARTIFACT. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON-LIFE-SUSTAINING RHYTHM, WITH MOTION ARTIFACT. AT03:44:55, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH CPR/MOTION ARTIFACT. AT 03:46:02, THE PATIENT RECEIVED THE FIFTH INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON-LIFE-SUSTAINING RHYTHM, WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT. AT 03:46:24, THE PATIENT RECEIVED THE SIXTH INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON-LIFE-SUSTAINING RHYTHM, WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT. THE DEVICE WAS SHUT DOWN AT 04:08:56 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21671458&pc=MVK | 3008642652-2025-02996 | 2025-02-23 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-03-22 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED TWO INAPPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 18:57:29 ON (B)(6) 2025. AT 19:55:55, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE. AT 19:57:04, THE PATIENT RECEIVED THE FIRST INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON LIFE-SUSTAINING RHYTHM. AT 20:02:17, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE. AT 20:03:22, THE PATIENT RECEIVED THE SECOND INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON LIFE-SUSTAINING RHYTHM. THE ELECTRODE BELT WAS DISCONNECTED AT 00:00:23 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21671459&pc=MVK | 3008642652-2025-02997 | 2025-03-02 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-03-22 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED AN INAPPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 23:33:10 ON (B)(6) 2025. AT 20:01:20 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE. AT 20:11:39, THE PATIENT RECEIVED THE INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON-LIFE-SUSTAINING RHYTHM, WITH MOTION ARTIFACT AND ELECTRODE LEAD FALL OFF. THE DEVICE WAS SHUT DOWN AT 20:18:03 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21671642&pc=LWS | 2124215-2025-17740 | 2025-02-25 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-03-22 04:00 | LWS | EMBLEM MRI S-ICD | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P110042 |  | 1 | Event Description: IT WAS REPORTED THAT THIS PATIENT PASSED AWAY ON THE (B)(6) 2025, WHICH WAS THE DAY OF IMPLANT OF THE SUBCUTANEOUS IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (S-ICD) SYSTEM. MEDICAL/SURGICAL INTERVENTIONS WERE REPORTED TO BE CARDIOPULMONARY RESUSCITATION (CPR), A TEMPORARY PACEMAKER AND IMPELLA. NO DEVICE ISSUES WERE REPORTED TO BE SUSPECTED, AND NO OTHER DETAILS WERE PROVIDED.Event Description: IT WAS REPORTED THAT THIS PATIENT PASSED AWAY ON THE (B)(6) 2025, WHICH WAS THE DAY OF IMPLANT OF THE SUBCUTANEOUS IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (S-ICD) SYSTEM. MEDICAL/SURGICAL INTERVENTIONS WERE REPORTED TO BE CARDIOPULMONARY RESUSCITATION (CPR), A TEMPORARY PACEMAKER AND IMPELLA. NO DEVICE ISSUES WERE REPORTED TO BE SUSPECTED, AND NO OTHER DETAILS WERE PROVIDED. IT WAS ADDITIONALLY REPORTED THAT AN AUTOPSY WAS NOT PERFORMED, AND THE BODY WAS RELEASED BY THE CORONER. THE CAUSE OF DEATH WAS NOT DETERMINED. Manufacturer Narrative: REFER TO SECTION B.5. FOR ADDITIONAL EVENT DESCRIPTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21660399&pc=NKM | 2135147-2025-01470 | 2025-02-24 05:00 | Death | ABBOTT MEDICAL | 2025-03-21 04:00 | NKM | MITRACLIP G4 CLIP DELIVERY SYSTEM (MDR) | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Pain; Pericardial Effusion | P100009 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A MITRACLIP PROCEDURE TO TREAT FUNCTIONAL MITRAL REGURGITATION (MR) WITH A GRADE OF 4. ON (B)(6) 2025, THE PATIENT DEVELOPED A PERICARDIAL EFFUSION, ALONG WITH ABDOMINAL PAIN, FOLLOWED BY CARDIAC ARREST. CPR WAS PERFORMED. ON (B)(6) 2025, THE PATIENT EXPIRED. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR EVALUATION. A REVIEW OF THE LOT HISTORY RECORD IDENTIFIED NO MANUFACTURING NONCONFORMITIES ISSUED TO THE REPORTED LOT THAT WOULD HAVE CONTRIBUTED TO THE REPORTED EVENT. BASED ON THE INFORMATION REVIEWED, A CAUSE FOR THE REPORTED PAIN, CARDIAC ARREST AND PERICARDIAL EFFUSION RESULTING IN DEATH CANNOT BE DETERMINED. DEATH, PERICARDIAL EFFUSION, PAIN AND CARDIAC ARREST ARE KNOWN POSSIBLE COMPLICATIONS ASSOCIATED WITH MITRACLIP PROCEDURES. UNEXPECTED MEDICAL INTERVENTION WAS A RESULT OF CASE-SPECIFIC CIRCUMSTANCES. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21660800&pc=KPR | 3003768251-2025-000042 | 2025-02-25 05:00 | Death | PHILIPS MEDICAL SYSTEMS DMC GMBH | 2025-03-21 04:00 | KPR | RADIOGRAPHY 7300 C | Adverse Event Without Identified Device or Use Problem | Fall; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K233662 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A PATIENT EXAMINATION, AN OLD AGE PATIENT FELL OFF FROM THE TABLE WHILE THE OPERATOR WAS IN THE CONTROL ROOM. SIX DAYS LATER, THE PATIENT PASSED AWAY IN THE HOSPITAL, THOUGH IT'S UNCLEAR WHETHER THE FALL CONTRIBUTED TO THE DEATH. THE FALL OCCURRED ONLY ONCE AND WAS PROMPTLY FOLLOWED UP WITH A CT SCAN FOR INJURIES. THERE WAS NO IMPACT TO THE USER. THE PATIENT IMPACT IS DEATH. Manufacturer Narrative: STILL UNDER INVESTIGATION. Manufacturer Narrative: THE RADIOGRAPHY 7300 C IS USED IN GENERAL RADIOGRAPHIC EXAMINATIONS AND APPLICATIONS. IT IS SUITABLE FOR ALL ROUTINE RADIOGRAPHY EXAMINATIONS, INCLUDING SPECIALIST AREAS LIKE INTENSIVE CARE, TRAUMA OR PEDIATRIC WORK, EXCLUDING FLUOROSCOPY, ANGIOGRAPHY AND MAMMOGRAPHY. STATEMENT OF THE REPORTED PROBLEM: PHILIPS RECEIVED FEEDBACK THROUGH CFMS (CUSTOMER FEEDBACK MANAGEMENT SYSTEM) THAT A PATIENT FELL FROM THE TABLE OF THE RADIOGRAPHY 7300 C SYSTEM DIED 6 DAYS LATER. THE SEQUENCE OF EVENTS IS PROVIDED BELOW. ACCORDING TO THE HEALTH CARE PROVIDERS AT THE CUSTOMER LOCATION, ON (B)(6) 2025, 01:15 PM A FEMALE PATIENT, AGE 89 YEARS, UNDERWENT RADIOLOGICAL EXAMINATION WITH THE RADIOGRAPHY 7300 C SYSTEM. RADIOGRAPHS OF THE LUMBAR SPINE WERE REQUESTED FOR VERTEBRAL FRACTURES AS THE PATIENT WAS EXPERIENCING INCREASED PAIN IN THE LUMBAR SPINE AND PELVIS. PATIENT WAS STABLE BUT LIMITED IN COOPERATION. A ROLL BOARD WAS USED TO TRANSFER THE PATIENT FROM THE BED ONTO THE X-RAY TABLE, WITH NO CONFIRMATION OF ANY RESTRAINTS OR HAND GRIPS BEING USED. JUST AS THE OPERATOR STEPPED INTO THE CONTROL ROOM TO START THE SCAN, THE PATIENT TURNED WHILE ON THE TABLE AND FELL OFF THE TABLE ONTO THE FLOOR. SHORTLY AFTER THE FALL, THE PATIENT WAS STABLE AND CONSCIOUS. THE PATIENT WAS EXAMINED FOR INJURIES DUE TO THE FALL, AND A CT SCAN OF THE HEAD WAS PERFORMED THAT SHOWED GENERALIZED INVOLUTION AND PATCHY, INHOMOGENEOUS MARGIN DEGENERATION, AND PREVIOUSLY KNOWN FALX CALCIFICATIONS. THE CT SHOWED NO EVIDENCE OF A CALVARIAL FRACTURE, NO BLEEDING, NO CONTUSION, AND NO ACUTE CSF CIRCULATION DISTURBANCE. AFTER THE FALL, THE CASE WAS ALSO PRESENTED TO THE SURGEON ON DUTY AND THE PATIENT UNDERWENT ASSESSMENT IN THE EMERGENCY ROOM, BUT NO FURTHER DETAILS REGARDING THE TREATMENT WERE PROVIDED BY THE HEALTH CARE PROVIDERS AT THE CUSTOMER LOCATION. AFTER SIX DAYS, ON (B)(6) 2025, THE PATIENT DIED. THE DEATH CERTIFICATE AND OFFICIAL CAUSE OF DEATH WERE NOT PROVIDED DUE TO DATA PROTECTION REGULATIONS IN GERMANY. ANALYSIS PERFORMED: A SERVICEMAX CASE WAS NOT CREATED FOR THIS CASE AS THERE WAS NO ALLEGATION OF PRODUCT MALFUNCTION FROM THE CUSTOMER. NO FIELD SERVICE ENGINEER VISITED THE SITE TO EVALUATE THE SYSTEM. HOWEVER, COMPLAINT WAS ESCALATED FOR A TECHNICAL INVESTIGATION TO EVALUATE THE SYSTEM PERFORMANCE AT THE TIME OF THE INCIDENT ((B)(6) 2025). THE SYSTEM LOG FILES FROM THE DATE OF INCIDENT (B)(6) 2025 WERE RETRIEVED AND EVALUATED BY SYSTEM ENGINEER. THE RESULTS OF LOG FILE ANALYSIS ARE DOCUMENTED IN THE TECHNICAL INVESTIGATION CHILD RECORD- (B)(4). SUMMARY OF ANALYSIS OF LOG FILE RECEIVED (I.E. ST.VINZENZ-KRANKENHAUSHANAU_RADIOGRAPHYDR7300C_RAUMB_20250417_1345_SYSLOG.ZIP): 1) AROUND THE EVENT TIME I.E. 25TH OF FEBRUARY 2025, 01:15 PM, IT WAS OBSERVED THERE ARE NO ERROR LOGS REPORTED ON THE TABLE OR ANY GEOMETRY COMPONENTS. 2) THERE IS NO TABLE MOVEMENT INITIATED DURING THIS TIME. 3) THERE IS NO EVIDENCE OF ANY EXPOSURE INITIATED BY THE OPERATOR DURING THIS TIME. THE INSTRUCTIONS FOR USE (IFU, 3000 140 63981) FOR THE RADIOGRAPHY 7300 C DEVICE CONTAINS INFORMATION ABOUT THE USE OF HAND GRIPS (P286) FOR PATIENT SUPPORT, SIDE BARS (P291) AND ADJUSTABLE STRAPS (P293) FOR PATIENT SUPPORT. TWO TABLE TYPES CAN BE USED WITH THE RADIOGRAPHY 7300 C DEVICE, THE HEIGHT ADJUSTABLE TABLE (TH2) AND THE SINGLE SIDE SUSPENDED TABLE (TH-S). ADJUSTABLE STRAPS AND A SIDE BAR CAN BE USED WITH THE TH-S, WHILE HAND GRIPS CAN BE USED WITH THE TH2 TO HELP STABILIZE THE PATIENT. THE INSTRUCTIONS FOR USE (IFU, 3000 140 63981) FOR THE RADIOGRAPHY 7300 C DEVICE WARNS THAT THE PATIENT SHOULD BE SECURED AND UNDER CONSTANT OBSERVATION WHILE ON THE TABLE (P143) AND ADVISES THAT THE PATIENT USE HAND GRIPS AND/OR STRAPS IF NECESSARY (P144). FURTHER INVESTIGATION REVEALED THAT THE CUSTOMER DID NOT PURCHASE HAND GRIPS FOR TABLE TH-2 FROM PHILIPS OR FROM OTHER VENDORS. CUSTOMER DID NOT USE ANY MATTRESS AT THE TIME OF THE INCIDENT AS CONFIRMED BY THE FIELD SERVICE ENGINEER ON 02 MAY 2025. AS NO DEVICE FAULTS WERE REPORTED BY THE CUSTOMER AND NO DEVICE ERRORS WERE FOUND DURING SUBSEQUENT LOG FILE ANALYSIS, EVIDENCE DOES NOT EXIST TO SUGGEST THAT A DEVICE MALFUNCTION CAUSED OR CONTRIBUTED TO INJURY OR DEATH OF THE PATIENT. Manufacturer Narrative: ADDITIONAL NARRATIVE IS BASED ON THE RELEASED HEALTH HAZARD EVALUATION (HHE) OUTCOME AND PRODUCTION AND POST-MARKET RISK REVIEW (PPMRR) FOR THIS COMPLAINT: INVESTIGATION CONCLUSION: AS NO DEVICE FAULTS WERE REPORTED BY THE CUSTOMER AND NO DEVICE ERRORS WERE FOUND DURING SUBSEQUENT LOG FILE ANALYSIS, EVIDENCE DOES NOT EXIST TO SUGGEST THAT A DEVICE MALFUNCTION CAUSED OR CONTRIBUTED TO INJURY OR DEATH OF THE PATIENT. BASED ON THE HEALTH HAZARD EVALUATION IT HAS BEEN CONCLUDED THAT THE HAZARD WAS NOT A RESULT OF FAULTY EQUIPMENT BUT USER ERROR. HEALTH HAZARD EVALUATION (HHE) OUTCOME: SUMMARY OF HEALTH RISK: THE HAZARD MECHANICAL ENERGY/GRAVITY HAD BEEN IDENTIFIED IN THE RISK MANAGEMENT MATRIX [4] AT A LOWER SEVERITY LEVEL (S3). THE HAZARD WAS NOT A RESULT OF FAULTY EQUIPMENT BUT USER ERROR. THE IFU [1] CLEARLY STATES: WARNING "SAFETY PRECAUTIONS": SECURE THE PATIENT FROM FALLING OFF. TO KEEP A PATIENT ON THE TABLE UNDER CONSTANT OBSERVATION, IN PARTICULAR IF NARCOTIZED OR RESTIVE. WITH ALL MOVEMENTS MAKE SURE THAT YOU DO NOT RUN OVER ANY CABLES, PASS OVER ANYONE'S FEET OR COLLIDE WITH ANY OBSTACLES. THE BENEFITS OF THE DEVICE CONTINUE TO OUTWEIGH THE RISKS WITH RESPECT TO ITS INTENDED USE. THE DEVICE MAY CONTINUE TO BE USED AND / OR DISTRIBUTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21661147&pc=NIO | 9681442-2025-00089 | 2021-04-05 04:00 | Death | ANGIOMED GMBH & CO. MEDIZINTECHNIK KG | 2025-03-21 04:00 | NIO | E-LUMINEXX VASCULAR STENT | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P080007 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESULTS OF THE CLINICAL TRIAL THAT APPROXIMATELY FOUR YEARS EIGHT MONTHS AND TEN DAYS POST A STENT PLACEMENT ON THE LEFT LEG FOR THE DE-NOVO STENOSIS ON THE EXTERNAL ILIAC ARTERY, THE SUBJECT HAD AN ADVERSE EVENT OF DEATH. THE PRIMARY CAUSE OF DEATH WAS NOT PROVIDED. REPORTEDLY, THE ADVERSE EVENT DEATH WAS POSSIBLY RELATED TO THE STUDY DEVICE AND NOT RELATED TO THE PROCEDURE.Event Description: IT WAS REPORTED THROUGH THE RESULTS OF THE CLINICAL TRIAL THAT APPROXIMATELY FOUR YEARS, EIGHT MONTHS AND TEN DAYS POST A STENT PLACEMENT ON THE LEFT LEG FOR THE DE-NOVO STENOSIS ON THE EXTERNAL ILIAC ARTERY, THE SUBJECT HAD AN ADVERSE EVENT OF DEATH. THE PRIMARY CAUSE OF DEATH WAS NOT PROVIDED. REPORTEDLY, THE ADVERSE EVENT DEATH WAS POSSIBLY RELATED TO THE STUDY DEVICE AND NOT RELATED TO THE PROCEDURE. Manufacturer Narrative: H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. Manufacturer Narrative: H11: MANUFACTURING REVIEW: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. INVESTIGATION SUMMARY: THE STENT REMAINS IMPLANTED. X-RAY IMAGES WERE NOT PROVIDED FOR EVALUATION. THE EVALUATION OF THE MATERIAL AVAILABLE IS SUBJECT TO THE RETROSPECTIVE ASSESSMENT AS PART OF POST-MARKET CLINICAL FOLLOW-UP (PMCF) ACTIVITIES. THE DEATH OF THE PATIENT WAS DOCUMENTED NOT BEING RELATED TO THE PLACED STENT OR THE STENT PLACEMENT PROCEDURE. BASED ON THE INFORMATION AVAILABLE, THE INVESTIGATION IS CLOSED WITH INCONCLUSIVE RESULT. BASED ON INFORMATION AVAILABLE, THERE IS NO INDICATION THAT THE DEATH OF THE PATIENT MIGHT BE RELATED TO THE PLACED STENT OR THE STENT PLACEMENT PROCEDURE. LABELING REVIEW: RELEVANT LABELING SUPPLIED WITH THIS PRODUCT WAS REVIEWED. POTENTIAL COMPLICATIONS AND ADVERSE EVENTS WHICH MAY OCCUR DURING USE OF THE DEVICE WERE FOUND TO BE REFERENCED IN THE INSTRUCTIONS FOR USE. USERS AND/OR PATIENTS SHOULD REPORT ANY SERIOUS INCIDENT THAT HAS OCCURRED IN RELATION TO THE DEVICE TO THE MANUFACTURER AND TO EITHER THE EUROPEAN UNION NATIONAL COMPETENT AUTHORITY OR TO THE REGULATORY AUTHORITY OF THE COUNTRY IN WHICH THE USER AND/OR PATIENT IS ESTABLISHED. IN THIS CASE THE DEATH OF THE PATIENT WAS DOCUMENTED NOT BEING RELATED TO THE PLACED STENT OR THE STENT PLACEMENT PROCEDURE. G3, H6 (METHOD). SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21661157&pc=NIO | 9681442-2025-00090 | 2017-07-07 04:00 | Death | ANGIOMED GMBH & CO. MEDIZINTECHNIK KG | 2025-03-21 04:00 | NIO | E-LUMINEXX VASCULAR STENT | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P080007 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESULTS OF THE CLINICAL TRIAL THAT APPROXIMATELY THREE YEARS NINE MONTHS AND TWENTY SIX DAYS POST A STENT PLACEMENT ON THE LEFT LEG FOR THE DE-NOVO STENOSIS ON THE COMMON AND EXTERNAL ILIAC ARTERY, THE SUBJECT HAD AN ADVERSE EVENT OF DEATH. THE PRIMARY CAUSE OF DEATH WAS NOT PROVIDED. REPORTEDLY, THE ADVERSE EVENT DEATH WAS POSSIBLY RELATED TO THE STUDY DEVICE AND NOT RELATED TO THE PROCEDURE.Event Description: IT WAS REPORTED THROUGH THE RESULTS OF THE CLINICAL TRIAL THAT APPROXIMATELY THREE YEARS, NINE MONTHS AND TWENTY SIX DAYS POST A STENT PLACEMENT ON THE LEFT LEG FOR THE DE-NOVO STENOSIS ON THE COMMON AND EXTERNAL ILIAC ARTERY, THE SUBJECT HAD AN ADVERSE EVENT OF DEATH. THE PRIMARY CAUSE OF DEATH WAS NOT PROVIDED. REPORTEDLY, THE ADVERSE EVENT DEATH WAS POSSIBLY RELATED TO THE STUDY DEVICE AND NOT RELATED TO THE PROCEDURE. Manufacturer Narrative: H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. Manufacturer Narrative: H11: MANUFACTURING REVIEW: BASED ON THE EVENT INFORMATION PROVIDED THERE IS NO INDICATION THAT A MANUFACTURING PROCESS MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED ISSUE. THEREFORE, A DHR REVIEW IS CONSIDERED TO BE NOT REQUIRED. INVESTIGATION SUMMARY: THE STENT REMAINS IMPLANTED. X-RAY IMAGES WERE NOT PROVIDED FOR EVALUATION. THE EVALUATION OF THE MATERIAL AVAILABLE IS SUBJECT TO THE RETROSPECTIVE ASSESSMENT AS PART OF POST-MARKET CLINICAL FOLLOW-UP (PMCF) ACTIVITIES. THE DEATH OF THE PATIENT WAS DOCUMENTED NOT BEING RELATED TO THE PLACED STENT OR THE STENT PLACEMENT PROCEDURE. BASED ON THE INFORMATION AVAILABLE, THE INVESTIGATION IS CLOSED WITH INCONCLUSIVE RESULT. BASED ON INFORMATION AVAILABLE, THERE IS NO INDICATION THAT THE DEATH OF THE PATIENT MIGHT BE RELATED TO THE PLACED STENT OR THE STENT PLACEMENT PROCEDURE. LABELING REVIEW: RELEVANT LABELING SUPPLIED WITH THIS PRODUCT WAS REVIEWED. POTENTIAL COMPLICATIONS AND ADVERSE EVENTS WHICH MAY OCCUR DURING USE OF THE DEVICE WERE FOUND TO BE REFERENCED IN THE INSTRUCTIONS FOR USE. USERS AND/OR PATIENTS SHOULD REPORT ANY SERIOUS INCIDENT THAT HAS OCCURRED IN RELATION TO THE DEVICE TO THE MANUFACTURER AND TO EITHER THE EUROPEAN UNION NATIONAL COMPETENT AUTHORITY OR TO THE REGULATORY AUTHORITY OF THE COUNTRY IN WHICH THE USER AND/OR PATIENT IS ESTABLISHED. IN THIS CASE THE DEATH OF THE PATIENT WAS DOCUMENTED NOT BEING RELATED TO THE PLACED STENT OR THE STENT PLACEMENT PROCEDURE. D4 (MEDICAL DEVICE EXPIRATION DATE), G3, H6 (METHOD). SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21661235&pc=NIO | 9681442-2025-00092 | 2019-11-03 04:00 | Death | ANGIOMED GMBH & CO. MEDIZINTECHNIK KG | 2025-03-21 04:00 | NIO | E-LUMINEXX VASCULAR STENT | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P080007 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESULTS OF THE CLINICAL TRIAL THAT APPROXIMATELY TWO YEARS FOUR MONTHS AND THIRTEEN DAYS POST A STENT PLACEMENT ON THE LEFT LEG FOR THE DE-NOVO STENOSIS ON THE EXTERNAL ILIAC ARTERY, THE SUBJECT HAD AN ADVERSE EVENT OF DEATH. THE PRIMARY CAUSE OF DEATH WAS NOT PROVIDED. REPORTEDLY, THE ADVERSE EVENT DEATH WAS POSSIBLY RELATED TO THE STUDY DEVICE AND NOT RELATED TO THE PROCEDURE.Event Description: IT WAS REPORTED THROUGH THE RESULTS OF THE CLINICAL TRIAL THAT APPROXIMATELY TWO YEARS, FOUR MONTHS AND THIRTEEN DAYS POST A STENT PLACEMENT ON THE LEFT LEG FOR THE DE-NOVO STENOSIS ON THE EXTERNAL ILIAC ARTERY, THE SUBJECT HAD AN ADVERSE EVENT OF DEATH. THE PRIMARY CAUSE OF DEATH WAS NOT PROVIDED. REPORTEDLY, THE ADVERSE EVENT DEATH WAS POSSIBLY RELATED TO THE STUDY DEVICE AND NOT RELATED TO THE PROCEDURE. Manufacturer Narrative: H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. Manufacturer Narrative: H11: MANUFACTURING REVIEW: BASED ON THE EVENT INFORMATION PROVIDED THERE IS NO INDICATION THAT A MANUFACTURING PROCESS MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED ISSUE. THEREFORE, A DHR REVIEW IS CONSIDERED TO BE NOT REQUIRED. INVESTIGATION SUMMARY: THE STENT REMAINS IMPLANTED. X-RAY IMAGES WERE NOT PROVIDED FOR EVALUATION. THE EVALUATION OF THE MATERIAL AVAILABLE IS SUBJECT TO THE RETROSPECTIVE ASSESSMENT AS PART OF POST-MARKET CLINICAL FOLLOW-UP (PMCF) ACTIVITIES. THE DEATH OF THE PATIENT WAS DOCUMENTED NOT BEING RELATED TO THE PLACED STENT OR THE STENT PLACEMENT PROCEDURE. BASED ON THE INFORMATION AVAILABLE, THE INVESTIGATION IS CLOSED WITH INCONCLUSIVE RESULT. BASED ON INFORMATION AVAILABLE, THERE IS NO INDICATION THAT THE DEATH OF THE PATIENT MIGHT BE RELATED TO THE PLACED STENT OR THE STENT PLACEMENT PROCEDURE. LABELING REVIEW: RELEVANT LABELING SUPPLIED WITH THIS PRODUCT WAS REVIEWED. POTENTIAL COMPLICATIONS AND ADVERSE EVENTS WHICH MAY OCCUR DURING USE OF THE DEVICE WERE FOUND TO BE REFERENCED IN THE INSTRUCTIONS FOR USE. USERS AND/OR PATIENTS SHOULD REPORT ANY SERIOUS INCIDENT THAT HAS OCCURRED IN RELATION TO THE DEVICE TO THE MANUFACTURER AND TO EITHER THE EUROPEAN UNION NATIONAL COMPETENT AUTHORITY OR TO THE REGULATORY AUTHORITY OF THE COUNTRY IN WHICH THE USER AND/OR PATIENT IS ESTABLISHED. IN THIS CASE THE DEATH OF THE PATIENT WAS DOCUMENTED NOT BEING RELATED TO THE PLACED STENT OR THE STENT PLACEMENT PROCEDURE. D4 (MEDICAL DEVICE EXPIRATION DATE), G3, H6 (METHOD). SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21661236&pc=NIO | 9681442-2025-00091 | 2021-01-05 05:00 | Death | ANGIOMED GMBH & CO. MEDIZINTECHNIK KG | 2025-03-21 04:00 | NIO | E-LUMINEXX VASCULAR STENT | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P080007 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESULTS OF THE CLINICAL TRIAL THAT APPROXIMATELY FIVE YEARS THREE MONTHS AND TWENTY SEVEN DAYS POST A STENT PLACEMENT ON THE RIGHT LEG FOR THE DE-NOVO STENOSIS ON THE COMMON AND EXTERNAL ILIAC ARTERY, THE SUBJECT HAD AN ADVERSE EVENT OF DEATH. THE PRIMARY CAUSE OF DEATH WAS NOT PROVIDED. REPORTEDLY, THE ADVERSE EVENT DEATH WAS POSSIBLY RELATED TO THE STUDY DEVICE AND NOT RELATED TO THE PROCEDURE.Event Description: IT WAS REPORTED THROUGH THE RESULTS OF THE CLINICAL TRIAL THAT APPROXIMATELY FIVE YEARS, THREE MONTHS AND TWENTY SEVEN DAYS POST A STENT PLACEMENT ON THE RIGHT LEG FOR THE DE-NOVO STENOSIS ON THE COMMON AND EXTERNAL ILIAC ARTERY, THE SUBJECT HAD AN ADVERSE EVENT OF DEATH. THE PRIMARY CAUSE OF DEATH WAS NOT PROVIDED. REPORTEDLY, THE ADVERSE EVENT DEATH WAS POSSIBLY RELATED TO THE STUDY DEVICE AND NOT RELATED TO THE PROCEDURE. Manufacturer Narrative: H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE E-LUMINEXX VASCULAR STENT PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE E-LUMINEXX VASCULAR STENT PRODUCTS ARE IDENTIFIED IN D2 AND G4. H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. Manufacturer Narrative: H11: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE E-LUMINEXX VASCULAR STENT PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE E-LUMINEXX VASCULAR STENT PRODUCTS ARE IDENTIFIED IN D2 AND G4. MANUFACTURING REVIEW: BASED ON THE EVENT INFORMATION PROVIDED THERE IS NO INDICATION THAT A MANUFACTURING PROCESS MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED ISSUE. THEREFORE, A DHR REVIEW IS CONSIDERED TO BE NOT REQUIRED. INVESTIGATION SUMMARY: THE STENT REMAINS IMPLANTED. X-RAY IMAGES WERE NOT PROVIDED FOR EVALUATION. THE EVALUATION OF THE MATERIAL AVAILABLE IS SUBJECT TO THE RETROSPECTIVE ASSESSMENT AS PART OF POST-MARKET CLINICAL FOLLOW-UP (PMCF) ACTIVITIES. THE DEATH OF THE PATIENT WAS DOCUMENTED NOT BEING RELATED TO THE PLACED STENT OR THE STENT PLACEMENT PROCEDURE. BASED ON THE INFORMATION AVAILABLE, THE INVESTIGATION IS CLOSED WITH INCONCLUSIVE RESULT. BASED ON INFORMATION AVAILABLE, THERE IS NO INDICATION THAT THE DEATH OF THE PATIENT MIGHT BE RELATED TO THE PLACED STENT OR THE STENT PLACEMENT PROCEDURE. LABELING REVIEW: RELEVANT LABELING SUPPLIED WITH THIS PRODUCT WAS REVIEWED. POTENTIAL COMPLICATIONS AND ADVERSE EVENTS WHICH MAY OCCUR DURING USE OF THE DEVICE WERE FOUND TO BE REFERENCED IN THE INSTRUCTIONS FOR USE. USERS AND/OR PATIENTS SHOULD REPORT ANY SERIOUS INCIDENT THAT HAS OCCURRED IN RELATION TO THE DEVICE TO THE MANUFACTURER AND TO EITHER THE EUROPEAN UNION NATIONAL COMPETENT AUTHORITY OR TO THE REGULATORY AUTHORITY OF THE COUNTRY IN WHICH THE USER AND/OR PATIENT IS ESTABLISHED. IN THIS CASE THE DEATH OF THE PATIENT WAS DOCUMENTED NOT BEING RELATED TO THE PLACED STENT OR THE STENT PLACEMENT PROCEDURE. D4 (MEDICAL DEVICE EXPIRATION DATE), G3, H6 (METHOD). SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21661687&pc=NIK | 3004209178-2025-05043 | 2007-04-19 04:00 | Death | MEDTRONIC MED REL MEDTRONIC PUERTO RICO | 2025-03-21 04:00 | NIK | CONCERTO | Failure to Deliver Shock/Stimulation | Insufficient Information | P010031 |  | 1 | Event Description: IT WAS REPORTED APPROXIMATELY EIGHTEEN YEARS AFTER THE PATIENT EXPIRED THAT THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) WAS DEFECTIVE AND DID NOT SHOCK THE PATIENT. ADDITIONAL INFORMATION WAS REQUESTED BUT NOT OBTAINED. Manufacturer Narrative: CORRECTION: IMF (ANNEX F) HEALTH IMPACT CODES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: 419488 LEAD IMPLANTED (B)(6) 2006. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21661760&pc=MIH | 9612164-2025-01445 |  | Death | MEDTRONIC IRELAND | 2025-03-21 04:00 | MIH | VALIANT CAPTIVIA - FF | Leak/Splash; Adverse Event Without Identified Device or Use Problem | Chest Pain; Pneumonia; Bowel Perforation | P100040 |  | 1 | Event Description: VALIANT CAPTIVIA STENT GRAFTS WERE IMPLANTED DURING THE ENDOVASCULAR TREATMENT OF A THORACIC AORTIC ANEURYSM . ADDITIONALLY, AN ENDURANT II STENT GRAFT SYSTEM WAS IMPLANTED THE SAME DAY. APPROX. 1 YEAR AND 2 MONTHS POST INDEX PROCEDURE, A VALIANT NAVION STENT GRAFT WAS IMPLANTED ALONG WITH AN ADDITIONAL ENDURANT LIMB. IT WAS REPORTED ON AN UNKNOWN DATE, THE PATIENT PRESENTED EMERGENTLY EXPERIENCING CHEST AND BACK PAIN. CT IMAGING REVEALED A TYPE IA ENDOLEAK, A TYPE IB ENDOLEAK, AND A TYPE III ENDOLEAK, THOUGH NO RUPTURE WAS OBSERVED. THE PATIENT UNDERWENT RELINING WITH A NON-MDT THORACIC GRAFT. THE RELINING OF THE GRAFTS WAS SUCCESSFUL, BUT MULTIPLE COMPLICATIONS AROSE DURING THE PROCEDURE, INCLUDING A LACERATED BOWEL DUE TO ANESTHESIA-RELATED ISSUES, AND PNEUMONIA. THESE COMPLICATIONS ULTIM ATELY LED TO THE PATIENT¿S DEATH 10 DAYS LATER. PER THE PHYSICIAN THE CAUSE OF THE ENDOLEAKS AND GRAFT ISSUES WAS UNKNOWN. AN AUTOPSY REVEALED SIX HOLES IN THE IN THE OUTERMOST VALIANT CAPTIVIA STENT GRAFT WHICH APPEARED TO BE CAUSED BY THE INNER FREE FLOW GRAFT¿S BARE METAL RUBBING THROUGH THE OUTER GRAFT. THE CAUSE OF DEATH WAS DUE TO MULTIPLE COMPLICATIONS FROM THE PROCEDURE INCLUDING THE LACERATED BOWEL, ANESTHESIA-RELATED DIFFICULTIES, AND PNEUMONIA. Manufacturer Narrative: PRODUCT ANALYSIS #(B)(4):ONE (1) IMAGE OF THE EXPLANT WAS RECEIVED. A TEAR IS VISIBLE TO THE GRAFT FABRIC. A BARE STENT STRUT IS VISIBLE PROTRUDING THROUGH THE FABRIC MATERIAL. THE REPORTED ENDOLEAKS WERE CONFIRMED THROUGH IMAGE REVIEW. CONTINUATION OF D10: SECTION D INFORMATION REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER RELEVANT DEVICE(S) ARE: PRODUCT ID: VAMF3838C100TU, SERIAL/LOT #: (B)(6) , UBD: 14-DEC-2019, UDI#: (B)(4). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: H6; ADDITIONAL CODES ADDED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATE TO H6 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21661774&pc=MIH | 9612164-2025-01446 |  | Death | MEDTRONIC IRELAND | 2025-03-21 04:00 | MIH | VALIANT NAVION | Leak/Splash; Adverse Event Without Identified Device or Use Problem | Chest Pain; Pneumonia; Bowel Perforation | P100040 |  | 1 | Event Description: VALIANT CAPTIVIA STENT GRAFTS WERE IMPLANTED DURING THE ENDOVASCULAR TREATMENT OF A THORACIC AORTIC ANEURYSM. ADDITIONALLY, AN ENDURANT II STENT GRAFT SYSTEM WAS IMPLANTED THE SAME DAY. APPROX. 1 YEAR AND 2 MONTHS POST INDEX PROCEDURE, A VALIANT NAVION STENT GRAFT WAS IMPLANTED ALONG WITH AN ADDITIONAL ENDURANT LIMB. IT WAS REPORTED ON AN UNKNOWN DATE, THE PATIENT PRESENTED EMERGENTLY EXPERIENCING CHEST AND BACK PAIN. CT IMAGING REVEALED A TYPE IA ENDOLEAK, A TYPE IB ENDOLEAK, AND A TYPE III ENDOLEAK, THOUGH NO RUPTURE WAS OBSERVED. ON (B)(6) 2025, THE PATIENT UNDERWENT RELINING WITH A NON-MDT THORACIC GRAFT. THE RELINING OF THE GRAFTS WAS SUCCESSFUL, BUT MULTIPLE COMPLICATIONS AROSE DURING THE PROCEDURE, INCLUDING A LACERATED BOWEL DUE TO ANESTHESIA-RELATED ISSUES, AND PNEUMONIA. THESE COMPLICATIONS ULTIMATELY LED TO THE PATIENT¿S DEATH 10 DAYS LATER. PER THE PHYSICIAN THE CAUSE OF THE ENDOLEAKS AND GRAFT ISSUES WAS UNKNOWN. AN AUTOPSY REVEALED SIX HOLES IN THE IN THE OUTERMOST VALIANT CAPTIVIA STENT GRAFT WHICH APPEARED TO BE CAUSED BY THE INNER FREE FLOW GRAFT¿S BARE METAL RUBBING THROUGH THE OUTER GRAFT. THE CAUSE OF DEATH WAS DUE TO MULTIPLE COMPLICATIONS FROM THE PROCEDURE INCLUDING THE LACERATED BOWEL, ANESTHESIA-RELATED DIFFICULTIES, AND PNEUMONIA. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: H6; ADDITIONAL CODES ADDED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATE TO H6 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21661780&pc=MIH | 9612164-2025-01447 |  | Death | MEDTRONIC IRELAND | 2025-03-21 04:00 | MIH | ENDURANT II BIFURCATED STENT GRAFT | Leak/Splash; Adverse Event Without Identified Device or Use Problem | Chest Pain; Pneumonia; Bowel Perforation | P100021 |  | 1 | Event Description: VALIANT CAPTIVIA STENT GRAFTS WERE IMPLANTED DURING THE ENDOVASCULAR TREATMENT OF A THORACIC AORTIC ANEURYSM. ADDITIONALLY, AN ENDURANT II STENT GRAFT SYSTEM WAS IMPLANTED THE SAME DAY. APPROX. 1 YEAR AND 2 MONTHS POST INDEX PROCEDURE, A VALIANT NAVION STENT GRAFT WAS IMPLANTED ALONG WITH AN ADDITIONAL ENDURANT LIMB . IT WAS REPORTED ON AN UNKNOWN DATE, THE PATIENT PRESENTED EMERGENTLY EXPERIENCING CHEST AND BACK PAIN. CT IMAGING REVEALED A TYPE IA ENDOLEAK, A TYPE IB ENDOLEAK, AND A TYPE III ENDOLEAK, THOUGH NO RUPTURE WAS OBSERVED. THE PATIENT UNDERWENT RELINING WITH A NON-MDT THORACIC GRAFT. THE RELINING OF THE GRAFTS WAS SUCCESSFUL, BUT MULTIPLE COMPLICATIONS AROSE DURING THE PROCEDURE, INCLUDING A LACERATED BOWEL DUE TO ANESTHESIA-RELATED ISSUES, AND PNEUMONIA. THESE COMPLICATIONS ULTIMATELY LED TO THE PATIENT¿S DEATH 10 DAYS LATER. PER THE PHYSICIAN THE CAUSE OF THE ENDOLEAKS AND GRAFT ISSUES WAS UNKNOWN. AN AUTOPSY REVEALED SIX HOLES IN THE IN THE OUTERMOST VALIANT CAPTIVIA STENT GRAFT WHICH APPEARED TO BE CAUSED BY THE INNER FREE FLOW GRAFT¿S BARE METAL RUBBING THROUGH THE OUTER GRAFT. THE CAUSE OF DEATH WAS DUE TO MULTIPLE COMPLICATIONS FROM THE PROCEDURE INCLUDING THE LACERATED BOWEL, ANESTHESIA-RELATED DIFFICULTIES, AND PNEUMONIA. Manufacturer Narrative: CONTINUATION OF D10: SECTION D INFORMATION REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER RELEVANT DEVICE(S) ARE: PRODUCT ID: P RODUCT ID: ETLW1620C124E, SERIAL/LOT #: (B)(6) UBD: 09-JAN-2020, UDI#: (B)(4) ; PRODUCT ID: ETLW1613C156E, SERIAL/LOT #:(B)(6), UBD: 09-JAN-2020, UDI#: (B)(4); PRODUCT ID: ETLW1610C124E, SERIAL/LOT #: (B)(6)+, UBD: 10-OCT-2020, UDI#: (B)(4). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: H6; ADDITIONAL CODES ADDED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21661853&pc=LWR | 9612164-2025-01451 | 2025-02-03 05:00 | Death | MEDTRONIC MEXICO | 2025-03-21 04:00 | LWR | AVALUS AORTIC TISSUE VALVE | Insufficient Information | Insufficient Information | P170006 |  | 1 | Event Description: MEDTRONIC RECEIVED INFORMATION FROM A PATIENT REPRESENTATIVE THAT 9 YEARS AND 11 MONTHS POST IMPLANT OF THIS 27MM AORTIC BIOPROSTHETIC VALVE, THE PATIENT DIED AT HOME. THE CAUSE OF DEATH WAS NOT RECEIVED. THEREFORE, RELATEDNESS OF THE DEATH TO THE VALVE COULD NOT BE EXCLUDED. IT WAS NOTED THAT THE PATIENT HAD NOT BEEN FEELING VERY WELL DAYS PRIOR TO THEIR DEATH. IT IS UNKNOWN WHETHER AN AUTOPSY WAS PERFORMED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21662123&pc=MHX | 9610816-2025-000235 | 2025-01-26 05:00 | Death | PHILIPS MEDIZIN SYSTEME BÖBLINGEN GMBH | 2025-03-21 04:00 | MHX | INTELLIVUE X3 | No Audible Alarm | Bradycardia | K171801 |  | 1 | Event Description: IT WAS REPORTED AFTER A RECENT NEW DEVICE INSTALLATION, THERE WAS AN INTENSIVE CARE UNIT (ICU) PATIENT WHO PASSED AWAY. THE PATIENT WAS IN BRADYCARDIA AND LAYING DOWN IN THEIR BED AWAITING TRANSFER TO THE WARD BUT STAFF LATER FOUND THEM DECEASED. THE CUSTOMER INDICATED THERE ARE MORE ALARMS WITH THE NEW MONITORS, ALARMS WERE MISSED OR IGNORED BY STAFF, WHICH MAY HAVE CONTRIBUTED TO THE REPORTED EVENT. IN ADDITION, THE CUSTOMER INDICATED THERE WERE ISSUES WITH THE WAY OF WORKING IN THE ICU AS WELL. THE CUSTOMER INVESTIGATED THE LOGS, REVEALING TECHNICAL INOP ALARMS FOR ECG PATCH AND SPO2 CONNECTIONS; HOWEVER, IT WAS FOUND THE DEVICES BEHAVED PER SPECIFICATION. Manufacturer Narrative: REPORTING INSTITUTION PHONE#: (B)(6). REPORTER PHONE#: (B)(6). A PHILIPS FIELD SERVICE ENGINEER (FSE) WAS NOTIFIED THAT A PATIENT WAS IN BRADYCARDIA AND LAYING DOWN IN THEIR BED AWAITING TRANSFER TO THE WARD BUT STAFF LATER FOUND THEM DECEASED. THE CUSTOMER INDICATED THERE ARE MORE ALARMS WITH THE NEW MONITORS, ALARMS WERE MISSED OR IGNORED BY STAFF, WHICH MAY HAVE CONTRIBUTED TO THE REPORTED EVENT. IN ADDITION, THE CUSTOMER INDICATED THERE WERE ISSUES WITH THE WAY OF WORKING IN THE INTENSIVE CARE UNIT (ICU) AS WELL. THE CUSTOMER INVESTIGATED THE LOGS, REVEALING TECHNICAL INOP ALARMS FOR ECG PATCH AND SPO2 CONNECTIONS; HOWEVER, IT WAS FOUND THE DEVICES BEHAVED PER SPECIFICATION. ADDITIONALLY, SOME PATIENT DATA WAS NOT AVAILABLE TO SEND TO PHILIPS AS THE PATIENT WAS NOT ADMITTED TO THE SYSTEM. ADMISSION TO THE SYSTEM IS A NEW WORKFLOW FOR THE CUSTOMER WITH THE NEW MONITORS. A PHILIPS REMOTE SERVICE ENGINEER (RSE) SUPPORTED THE CUSTOMER TO COLLECT THE LOGS. THE EVENT TOOK PLACE IN ROOM 9 ON (B)(6) 2025 AT 1930-2255 HOURS. THE CUSTOMER VERIFIED THAT A 3-LEAD ECG CABLE WAS IN USE. THE DATA WAS SENT TO A PHILIPS PRODUCT SUPPORT ENGINEER (PSE) AND A PHILIPS CLINICAL SPECIALIST (PCS) FOR REVIEW AND A SUMMARY OF THE FINDINGS IS DOCUMENTED BELOW. THE LOGS SHOW THAT THE MONITOR WAS TAKEN OUT OF STANDBY AT 19:35:56 (B)(6) 2025 19:35:56 (B)(6), RUM 9 STANDBY AVSLUTAS MT34463 EXITING STANDBY. AT 19:35:52, ECG LEADS OFF IS ANNOUNCED. (B)(6) 2025 19:35:57 (B)(6), RUM 9 EKG LÖSA ELEKTR SKAPADES 19:35:52. PIC IX: KGAIXAKUTC1 LOGGED INOP. AT 19:37:00, HEART RATE ALARMS ARE SWITCHED OFF AND PULSE BECOMES THE ALARM SOURCE. THE ECG/ARRHYTHMIA ALARMS ARE CHANGED FROM ON TO OFF. FROM THE IFU, ¿SELECTING PULSE AS THE ACTIVE ALARM SOURCE FOR HR/PULSE SWITCHES OFF THE ARRHYTHMIA ALARMS, INCLUDING ASYSTOLE, VFIB AND VTACH ALARMS, AND THE HEART RATE ALARMS. THIS IS INDICATED BY THE MESSAGE ECG/ARRH ALARMSOFF (UNLESS THIS HAS BEEN CONFIGURED OFF FOR YOUR MONITOR), AND THE CROSSED-OUT ALARM SYMBOL BESIDE THE ECG HEART RATE NUMERIC. THE MESSAGE ECG/ARRH ALARMSOFF CAN BE CONFIGURED OFF, OR TO SWITCH TO A YELLOW (MEDIUM SEVERITY) INOP AFTER A FIXED NUMBER OF HOURS. HIGH AND LOW PULSE RATE AND EXTREME BRADYCARDIA AND EXTREME TACHYCARDIA ALARMS FROM PULSE ARE ACTIVE.¿ IF THE ALARM SOURCE IS SET TO AUTO, THIS SWITCHING OFF MAY BE AUTOMATICALLY FOLLOWING A MANUAL REQUEST TO TURN ON. WHEN THE ALARM SOURCE IS SET TO AUTO, THE MONITOR WILL USE THE HEART RATE FROM THE ECG MEASUREMENT AS THE ALARM SOURCE WHENEVER THE ECG MEASUREMENT IS SWITCHED ON AND AT LEAST ONE ECG LEAD CAN BE MEASURED WITHOUT AN INOP CONDITION. SEE DATA SET B FOR MANY EXAMPLES OF ALARM ON/OFF ENTRIES. AT 19:37:07. **SPO2 86<90 WAS GENERATED. AT 19:37:16, **PULSE 45<50 WAS GENERATED. 19:37:27, ALARMS WERE ACKNOWLEDGED AT MT34463 BEDSIDE MONITOR. **SPO2 86<90 ENDED AT 19:37:57 AND THEN GENERATED AGAIN AT 19:38:22. **PULSE 40<50 ENDED AT 19:38:02. ** PULSE 35<50 GENERATED AT 19:38:21. THESE ALARMS ENDED AT 19:38:29 WHEN THE ***BRADY/PULSE ALARM 26<40 GENERATED. THE ***BRADY/PULSE ALARM ENDED WITH A MAXIMUM VIOLATION OF 25<40 AT 19:38:38. THE ECG/ARRHYTHMIA ALARMS WERE SWITCHED ON AND OFF SEVERAL TIMES BETWEEN 19:37 AND 22:46:30. AT 19:41:32, **PULS 46<50 GENERATED AND ENDED AT 19:42:21 WITH A MAXIMUM OF VIOLATION OF 35. **PULSE 46<50 GENERATED AT 19:43:09 AND ENDED AT 19:44:02. AT 19:44:24, **PULS 42<50 GENERATED. AT 19:44:40, ECG LEAD OFF ENDED AND THE ECG/ARRHYTHMIA ALARM SWITCHED BACK ON. THIS RESULTED IN THE **PULSE ALARM ENDING AND THE SOUNDING OF THE *HR 48<50 ALARM AT 19:44:46. THE AUDIT LOG CONTINUES TO SHOW *HR ALARMS GENERATING AND ENDING UNTIL 20:10:31 WHEN THE ***XBRADY 39<40 GENERATED. THE ALARM ENDED AT 20:10:47 WITH A MAXIMUM VIOLATION OF 36 <40. SPO2 PULSE ? AND SPO2 NO PULSE INOPS ARE NOTED. THE PATIENT CONTINUED TO HAVE LOW HR AND XBRADY ALARMS AT 20:46:06, CANNOT ANALYSIS ECG GENERATED. AT 21:24:59, !!ECG LEADS OFF INOP GENERATED. A DEVICE PROBLEM IS NOT CONFIRMED. THE INVESTIGATION DETERMINED THAT THE DEVICE OPERATED AS EXPECTED. A CUSTOMER RESPONSE LETTER IS BEING PROVIDED TO THE CUSTOMER TO RESOLVE THE ISSUE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21662146&pc=DTB | 3008973940-2025-02059 | 2025-02-26 05:00 | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-03-21 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Device Dislodged or Dislocated | Cardiac Tamponade; Tricuspid Valve Insufficiency/ Regurgitation; Insufficient Information | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT POST-OPERATIVELY AFTER THE IMPLANT OF THE RIGHT VENTRICULAR (RV) LEAD, THE LEAD WAS DISLODGED AND WAS FLOATING AROUND THE RV. THE DISLODGED LEAD WAS CAUSING ECTOPY. PATIENT HAD EXTREMELY DILATED HEART AND TORRENTIAL TRICUSPID REGURGITATION. THE LEAD AND THE IMPLANTABLE PULSE GENERATOR (IPG) WAS DECIDED TO BE REPLACED WITH A LEADLESS IPG. POST-OPERATIVELY AFTER THE LEADLESS IPG IMPLANTATION, THE PATIENT HAD DEVELOPED CARDIAC TAMPONADE. PERICARDIOCENTESIS WAS PERFORMED AND BLOOD WAS ASPIRATED. THE PATIENT DECEASED. THE EXACT CAUSE OF THE TAMPONADE HAS NOT BEEN CONFIRMED AND THE PHYSICIAN SUSPECTED DUE TO THE RV LEAD. THE CORONER CONFIRMED THAT TWO TINES OF THE DEVICE HAD PERFORATED CAUSING THE TAMPONADE. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21662202&pc=KFM | 21662202 | 2025-02-04 05:00 | Death | THORATEC CORPORATION | 2025-03-21 04:00 | KFM | THORATEC/ABBOTT CENTRIMAG BLOOD PUMP |  |  |  |  | 1
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21662232&pc=NVZ | 2182208-2025-01420 | 2025-02-26 05:00 | Death | MEDTRONIC, INC. | 2025-03-21 04:00 | NVZ | AZURE S SR MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P980035 |  | 1 | Event Description: IT WAS REPORTED THAT POST-OPERATIVELY AFTER THE IMPLANT OF THE RIGHT VENTRICULAR (RV) LEAD, THE LEAD WAS DISLODGED AND WAS FLOATING AROUND THE RV. THE DISLODGED LEAD WAS CAUSING ECTOPY. PATIENT HAD EXTREMELY DILATED HEART AND TORRENTIAL TRICUSPID REGURGITATION. THE LEAD AND THE IMPLANTABLE PULSE GENERATOR (IPG) WAS DECIDED TO BE REPLACED WITH A LEADLESS IPG. POST-OPERATIVELY AFTER THE LEADLESS IPG IMPLANTATION, THE PATIENT HAD DEVELOPED CARDIAC TAMPONADE. PERICARDIOCENTESIS WAS PERFORMED AND BLOOD WAS ASPIRATED. THE PATIENT DECEASED. THE EXACT CAUSE OF THE TAMPONADE HAS NOT BEEN CONFIRMED AND THE PHYSICIAN SUSPECTED DUE TO THE RV LEAD. THE CORONER CONFIRMED THAT TWO TINES OF THE DEVICE HAD PERFORATED CAUSING THE TAMPONADE. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTION H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21662300&pc=PNJ | 9612164-2025-01459 | 2025-02-26 05:00 | Death | MEDTRONIC IRELAND | 2025-03-21 04:00 | PNJ | MICRA¿ VR2 | Adverse Event Without Identified Device or Use Problem | Cardiac Tamponade; Cardiac Perforation; Insufficient Information | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT POST-OPERATIVELY AFTER THE IMPLANT OF THE RIGHT VENTRICULAR (RV) LEAD, THE LEAD WAS DISLODGED AND WAS FLOATING AROUND THE RV. THE DISLODGED LEAD WAS CAUSING ECTOPY. PATIENT HAD EXTREMELY DILATED HEART AND TORRENTIAL TRICUSPID REGURGITATION. THE LEAD AND THE IMPLANTABLE PULSE GENERATOR (IPG) WAS DECIDED TO BE REPLACED WITH A LEADLESS IPG. POST-OPERATIVELY AFTER THE LEADLESS IPG IMPLANTATION, THE PATIENT HAD DEVELOPED CARDIAC TAMPONADE. PERICARDIOCENTESIS WAS PERFORMED AND BLOOD WAS ASPIRATED. THE PATIENT DECEASED. THE EXACT CAUSE OF THE TAMPONADE HAS NOT BEEN CONFIRMED AND THE PHYSICIAN SUSPECTED DUE TO THE RV LEAD. THE CORONER CONFIRMED THAT TWO TINES OF THE DEVICE HAD PERFORATED CAUSING THE TAMPONADE. Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME: [MICRA], PRODUCT ID: [MC2VR01-DELSYS], (SERIAL: (B)(6). D9: NO. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21662346&pc=OZD | 1220648-2025-26868 | 2025-02-23 05:00 | Death | ABIOMED, INC. | 2025-03-21 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Pain; Sepsis | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT AFTER REACHING FOR THE BEDSIDE TABLE DURING IMPELLA 5.5 SUPPORT, THE PATIENT DEVELOPED INCREASING PAIN AND SWELLING AT THEIR ACCESS SITE. A LIDOCAINE PATCH WAS APPLIED AND OXYCONTIN WAS GIVEN TO MANAGE THE PAIN. SUPPORT CONTINUED UNINTERRUPTED WITH THE IMPLANTED PUMP. ROUGHLY TEN DAYS INTO SUPPORT, THE PATIENT DEVELOPED AN INFECTION OF UNKNOWN ORIGIN. ANTIBIOTICS WERE GIVEN, BUT THE PATIENT BECAME SEPTIC AND BEGAN TO DECOMPENSATE. THE DECISION WAS MADE TO WITHDRAW CARE AND THE PATIENT PASSED AWAY. THUS, THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: TABLE 3.2 IMPELLA CATHETER COMPONENTS ¿THE INFUSION FILTER PREVENTS BACTERIAL CONTAMINATION AND AIR FROM ENTERING THE PURGE LUMEN.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿ ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ Manufacturer Narrative: THE INVESTIGATION OF PAIN AT INSERTION SITE AND INFECTION/SEPSIS HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21662421&pc=NVZ | 3008973940-2025-02060 | 2025-02-24 05:00 | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-03-21 04:00 | NVZ | AZURE XT DR MRI SURESCAN | Pacing Problem | Bradycardia; Presyncope | P980035 |  | 1 | Event Description: IT WAS REPORTED BY THE PATIENT REPRESENTATIVE THAT THE PATIENT ALMOST FAINTED AND EXPERIENCED BRADYCARDIA AND IT WAS IDENTIFIED THAT THE IMPLANTABLE PULSE GENERATOR (IPG) WAS FUNCTIONING BELOW THE SET PROGRAMMED RATE. IT WAS FURTHER REPORTED THAT THE PATIENT WAS DECEASED.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT IS ALIVE WAS NOT DECEASED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTION TO A2. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21662656&pc=KFM | 3013876692-2025-00020 | 2025-01-29 05:00 | Death | MAQUET CARDIOPULMONARY GMBH | 2025-03-21 04:00 | KFM | HEART LUNG MACHINE | Electrical /Electronic Property Problem | Adult Respiratory Distress Syndrome |  |  | 1 | Event Description: ORIGINALLY IT WAS REPORTED TO THE FACTORY- THE EVENT OCCURRED IN INDIA. IT WAS REPORTED THAT THE ERROR MESSAGE ¿LAMPTEST¿ WAS DISPLAYED ON THE ROTAFLOW CONSOLE. IT WAS NOTICED DURING PATIENT TREATMENT AND THE DEVICE WAS IMMEDIATELY EXCHANGED WITH A BACKUP DEVICE. DURING SERVICE THE ERROR MESSAGE ¿HEAD ERROR¿ OCCURRED ON THE ROTAFLOW CONSOLE. NO HARM TO ANY PERSON HAS BEEN REPORTED. DUE TO THE REPORTED EXCHANGE OF THE DEVICE DURING PATIENT TREATMENT AND DUE TO THAT THE REPORTED FAILURE "HEAD ERROR" COULD LEAD TO THE PUMP STOP A REPORT IS REQUIRED. COMPLAINT ID: (B)(4). NEW INFORMATION WAS THEN PROVIDED BY THE SSU DATED ON (B)(6) 2025 THAT THE PATIENT PASSED AWAY. COMPLAINT ID: (B)(4). THIS FACILITATED AN INITIAL MDR TO BE FILED BY THE SSU.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21662706&pc=MHY | 3004209178-2025-05048 | 2024-12-24 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-03-21 04:00 | MHY | PERCEPT | Adverse Event Without Identified Device or Use Problem | Dysphagia/ Odynophagia; Unspecified Infection; Dysphasia; Shaking/Tremors; Lethargy; Convulsion/Seizure; Unspecified Nervous System Problem; Skin Inflammation/ Irritation; Insufficient Information | P960009 |  | 1 | Event Description: INFORMATION WAS RECEIVED FROM A HEALTHCARE PROVIDER (HCP) WHO REPORTED A PATIENT WHO WAS IMPLANTED FOR ESSENTIAL TREMOR FELL AND HIT THEIR HEAD ON (B)(6) AFTER BEING CONFUSED AND EXPERIENCING HALLUCINATIONS RESULTING IN PRESENTATION TO THE EMERGENCY ROOM. MRI WAS PERFORMED WHICH SHOWED NO ACUTE FINDINGS AND MENTAL STATUS HAD IMPROVED AT THAT TIME. THE PATIENT WAS WORKING ON DISCHARGE BACK TO THE FACILITY THEY LIVED IN BUT DEVELOPED FEVER AND HYPOXIA THE EVENING OF (B)(6) AFTER WHICH WAS FOUND TO BE COVID POSITIVE. THE PATIENT DEVELOPED ACUTE RESPIRATORY DISTRESS/FAILURE FROM COVID WHICH WORSENED TO PNEUMONIA REQUIRING INTUBATION AND IV ANTIBIOTIC TREATMENT FOR FIVE DAYS. THE PATIENT WAS EXTUBATED AND EVENTUALLY DISCHARGED ON (B)(6). AT THAT TIME THE PATIENT WAS NOTED TO HAVE SIGNIFICANT WORSENING OF FACIAL DYSTONIA AS WELL AS PERSISTENT DYSTONIA IN THE NECK AND UPPER EXTREMITIES. BLOOD WORK WAS PERFORMED WITH NORMAL RESULTS. THE PATIENT WAS SEEN BY NEUROLOGY, BUT THEY WERE UNABLE TO CONTROL HIS SYMPTOMS EVEN AFTER THE PATIENT¿S ATIVAN WAS INCREASED TO 1 MG EVERY FROM HOURS FROM EVERY 6 AND THE ADDITION OF KEPPRA FOR SUSPECTED SEIZURE. SINCE DISCHARGE THE PATIENT WAS NEARLY APHASIC AND HAD NOT SPOKEN ANY WORDS. WITH NO IMPROVEMENT THE PATIENT WAS TRANSFERRED TO ANOTHER HEALTHCARE FACILITY. UPON ARRIVAL AT THE NEW HEALTHCARE FACILITY ON (B)(6) 2025 THE PATIENT WAS LETHARGIC, SLEEPY, WOULD NOT OPEN HIS EYES, AND WAS UNABLE TO FOLLOW ANY COMMANDS. IN ADDITION, THE PATIENT HAD CONSTANT RHYTHMIC MOVEMENTS OF THEIR MOUTH AND BLINKING OF THE EYES. LUNG SOUNDS NOTED RHONCHOROUS BREATH SOUNDS BILATERALLY. THE PATIENT HAD A PEG TUBE PLACED THREE DAYS PRIOR DUE TO DYSPHAGIA AND PLACED ON LIQUID NUTRITION REPLACEMENT. ALL OTHER SYSTEMS WERE NORMAL UPON EXAM. FOLLOWING INITIAL ADMISSION TO THE MOST RECENT FACILITY, A CHEST X-RAY WAS OBTAINED WHICH SHOWED A FEW AREAS OF POSSIBLE INFECTION WITH THE PATIENT EXPERIENCING COUGHING EPISODES WITH THICK SECRETIONS AND WAS STARTED ON AN ANTIBIOTIC REGIMEN ALONG WITH NEBULIZER TREATMENT. NEUROLOGY WAS CONSULTED WHICH DOCUMENTED THE PATIENT HAD A PAST MEDICAL HISTORY OF CHRONIC MOVEMENT DISORDER WITH OROFACIAL DYSTONIA FOR WHICH THEY RECEIVED PERIODIC BOTOX (THE LAST INJECTION WAS (B)(6) 2024). ORAL MEDICATIONS HAD ALSO BEEN USED HISTORICALLY INCLUDING ATIVAN AND ARTANE 2MG 3 TIMES A DAY. THE PATIENT ALSO HAD A HISTORY OF ESSENTIAL TREMOR AND WAS ONPROPRANOLOL 60 MG TWICE A DAY, BUT IT WAS UNCLEAR IF THIS HAD BEEN CONTINUED WHEN THE PATIENT WAS INTUBATED AND HOSPITALIZED MOST RECENTLY. A DEEP BRAIN STIMULATOR WAS IMPLANTED YEARS PRIOR AS WELL. IT WAS NOTED THE PATIENT ALSO HAD COGNITIVE DEFICITS WHICH HAD REMAINED FAIRLY MILD LEADING UP TO THEIR MOST RECENT ADMISSION. FOLLOWING A THOROUGH EVALUATION THE NEUROLOGIST NOTED THE PATIENT HAD SEVERE DEFICITS WHICH HAD BEEN PRESENT SINCE BEING EXTUBATED WITH SEVERE PERSISTENT DYSTONIA CONSISTENT WITH STATUS DYSTONICUS OVER THE PAST TWO WEEKS. IN ADDITION, NEUROLOGY¿S ASSESSMENT LED TO THE ADDITION OF BACLOFEN AND BENADRYL ALONG WITH THE PATIENT¿S PREVIOUS MEDICATIONS. THE NEUROLOGIST NOTED THEY DIDN¿T THINK THE PATIENT WAS HAVING SEIZURES AND CHOSE TO HOLD THE KEPPRA. THE PATIENT¿S SPOUSE NOTED THE PATIENT DID WELL AFTER NEUROSTIMULATOR CHANGE ON (B)(6) 2024 WITH NO NEW CHANGES IN THEIR SETTINGS. THE PATIENT¿S SPOUSE WAS CONFIDENT THE DEEP BRAIN STIMULATOR WAS ON AT THIS TIME. THE NEUROLOGIST MENTIONED A FORMAL INTERROGATION OF THE DEVICE COULD BE PERFORMED. LAB RESULTS FROM (B)(6) SHOWED HIGH PC02 V-1, PO2 V-1, HC03, BE V1, NEUTROPHILS, AND MPV. THE PATIENT WAS REEVALUATED ON (B)(6) WHERE DUE TO THE PATIENT BEING REFRACTORY TO TREATMENT AND REQUIRING SIGNIFICANT SEDATION TO GET ANY RELIEF, WAS THUS TRANSFERRED TO NEUROCRITICAL CARE AND INTUBATED. IN SPITE OF THESE INTERVENTIONS OVER THE NEXT SEVERAL DAYS THE PATIENT HAD NO SIGNIFICANT IMPROVEMENT IN REGARD TO THEIR DYSTONIA, BUT RATHER THEIR DYSTONIC MOVEMENTS ONLY CONTINUED TO WORSEN. DISCUSSION TOOK PLACE WITH THE PATIENT¿S FAMILY WHO ELEVATED TO TRANSITION THE PATIENT TO COMFORT CARE ON (B)(6) 2025 AN WAS EXTUBATED. ADJUSTMENTS WERE MADE TO THE PATIENT¿S MEDICATIONS TO DIMINISH BREAKTHROUGH DYSTONIA AND DISCOMFORT. THE PATIENT PASSED AWAY ON (B)(6) 2025.Event Description: ADDITIONAL INFORMATION WAS RECEIVED FROM A HEALTHCARE PROFESSIONAL (HCP). THE HCP REPORTED THAT THE REPLACEMENT ON (B)(6) 2024 WAS THE ONLY PROCEDURE/INTERVENTION DONE AT THAT TIME. THE INS REPLACEMENT WAS PERFORMED UNDER GENERAL ANESTHESIA (WITH PROPOFOL) AND AN LMA (LARYNGEAL MASK AIRWAYS) WAS USED. THERE WERE NO NOTED COMPLICATIONS PER THE ANESTHESIA OR PACU (POST-ANESTHESIA CARE UNIT) RECORDS. PATIENT VITAL SIGNS REMAINED STABLE THROUGHOUT. THE PATIENT'S INDICATION FOR USE WAS FOR CENTRAL TREMORS ONLY. NO AUTOPSY RESULTS ARE AVAILABLE AT THIS TIME.Event Description: ***THIS EVENT IS NO LONGER A REPORTABLE EVENT. MDR DECISION CORRECTED TO NOT REPORTABLE. NO ADDITIONAL SUPPLEMENTAL MDRS ARE REQUIRED UNLESS ADDITIONAL INFORMATION RECEIVED MAKES THE EVENT REPORTABLE. DUE TO GCH FUNCTIONALITY, THE COMPLAINT FLAG CANNOT BE FLIPPED TO 'NO' AS A REGULATORY REPORT EXISTS IN THIS PE*** Manufacturer Narrative: H10: REVIEW OF THIS MDR AND/OR ADDITIONAL INFORMATION RECEIVED SHOWS THAT THERE IS NO INFORMATION TO REASONABLY SUGGEST THAT THE DEVICE IN THIS REPORT MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH OR SERIOUS INJURY OR THAT THE DEVICE IN THIS REPORT HAS MALFUNCTIONED. THEREFORE, THIS EVENT DOES NOT MEET THE REPORTING REQUIREMENTS STIPULATED IN 21 CFR 803. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SECTION D10 REFERENCES: PRODUCT ID: 3708660, SERIAL#: (B)(6). PRODUCT TYPE EXTENSION PRODUCT ID 3708660, SERIAL#: (B)(6), PRODUCT TYPE EXTENSION H3. ANALYSIS OF THE NEUROSTIMULATOR (S/N (B)(6)) FOUND NO SIGNIFICANT ANOMALY. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21662730&pc=BZD | 2518422-2025-103145 | 2025-01-24 05:00 | Death | RESPIRONICS, INC. | 2025-03-21 04:00 | BZD | DS2ADV AUTO CPAP W/HUMID+HT CELL/BT | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED A CALL ON DREANSTATION REPLACEMENT DEVICE, FROM THE PATIENT'S PARTNER CALLED ON HIS BEHALF. THE PATIENT PASSED AWAY AND SHE WAS NOT SURE WHAT TO DO WITH HIS REPLACEMENT DEVICE. SHE WANTED TO RETURN THE DEVICE IF SHE COULD. SHE DID NOT HAVE EMAIL ACCESS. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THERE WAS NO REPORT OF MEDICAL INTERVENTION. REPORTABLE, ALTHOUGH THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO THE DEATH, CAUSE OF DEATH IS UNKNOWN. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21662849&pc=NIP | 2024168-2025-02978 | 2025-02-05 05:00 | Death | ABBOTT VASCULAR | 2025-03-21 04:00 | NIP | SUPERA | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Perforation of Vessels | P120020 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT DIED POST SUPERA STENT IMPLANT DUE TO UNCONTROLLED BLEEDING FROM A PERFORATION. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORDS FOR THIS PRODUCT COULD NOT BE REVIEWED AND A SIMILAR COMPLAINT QUERY COULD NOT BE PERFORMED BECAUSE THE PART NUMBER AND LOT NUMBER WAS NOT REPORTED, AND THE DEVICE WAS NOT RETURNED. THE REPORTED PATIENT EFFECTS OF DEATH, PERFORATION AND HEMORRHAGIC EVENT, ARE LISTED IN THE SUPERA PERIPHERAL STENT SYSTEMS INSTRUCTIONS FOR USE, AS POTENTIAL ADVERSE EVENTS. A CONCLUSIVE CAUSE FOR THE REPORTED DEATH, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED. A CONCLUSIVE CAUSE FOR THE REPORTED HEMORRHAGE/BLOOD LOSS/BLEEDING AND PERFORATION OF VESSELS, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED. BASED ON THE REPORTED INFORMATION AND RESULTS OF THE COMPLAINT INVESTIGATION THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO THE DESIGN, MANUFACTURE, OR LABELING OF THE DEVICE. B2: DATE OF DEATH ESTIMATED AS (B)(6) 2025. B3: DATE OF PROCEDURE ESTIMATED AS (B)(6) 2025. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. D6A: DATE OF IMPLANT ESTIMATED AS (B)(6) 2025.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21662892&pc=BZD | 2518422-2025-103254 | 2025-02-19 05:00 | Death | RESPIRONICS, INC. | 2025-03-21 04:00 | BZD | REMSTAR AUTO A-FLEX | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K091319 |  | 1 | Event Description: THE MANUFACTURER HAS BEEN INFORMED OF THE RECENT PASSING OF A PATIENT WHO WAS USING A REMSTAR AUTO A-FLEX DEVICE. ACCORDING TO THE PATIENT'S WIFE, HER HUSBAND RECENTLY PASSED AWAY, AND SHE WAS UNABLE TO LOCATE THE DEVICE. SHE HAS INDICATED THAT SHE WILL RETURN IT IF SHE COMES ACROSS IT WHILE SORTING THROUGH HIS BELONGINGS. NO FURTHER DETAILS HAVE BEEN PROVIDED, AND THERE ARE NO ALLEGATIONS SUGGESTING THAT THE DEVICE HAD ANY ROLE IN OR CONTRIBUTED TO THE PATIENT'S DEATH. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21662909&pc=MCX | 2124215-2025-17315 | 2025-03-05 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-03-21 04:00 | MCX | ROTAPRO? | Detachment of Device or Device Component; Adverse Event Without Identified Device or Use Problem; Output Problem; Intermittent Loss of Power; Program or Algorithm Execution Failure | Ischemia; Foreign Body In Patient | P900056 |  | 1 | Event Description: IT WAS REPORTED THAT THE BURR FRACTURED AND THE PATIENT EXPIRED. A 1.50MM ROTAPRO AND A ROTAWIRE DRIVE WERE SELECTED FOR USE IN THE MODERATELY TORTUOUS AND SEVERELY CALCIFIED MID CIRCUMFLEX ARTERY. DURING THE PROCEDURE, AFTER THE ROTAWIRE DRIVE CROSSED THE LESION, ADVANCEMENT WAS NOT SUCCESSFUL. FURTHERMORE, A 1.25X20 NON-BOSTON SCIENTIFIC DEVICE DID CROSS. THE ROTAWIRE DRIVE WAS REPLACED AND A 1.50MM ROTAPRO BURR WAS ADVANCED. WHILE BURRING SEVERE ANGULATION OF THE CX OSTIUM LESION, THE ROTAPRO BURR SLIGHTLY DECELERATED BY APPROXIMATELY 10,000RPM AND STALLED. A LIGHT PULL ON THE ROTAPRO BURR WAS ATTEMPTED AND IT FRACTURED AT THE HYPOTUBE PORTION. DYNAGLIDE MODE WAS NOT SUCCESSFULLY USED. THE PROCEDURE WAS CANCELLED. AT AN UNKNOWN TIME, THE PATIENT EXPIRED.Event Description: IT WAS REPORTED THAT THE BURR FRACTURED AND THE PATIENT EXPIRED. A 1.50MM ROTAPRO AND A ROTAWIRE DRIVE WERE SELECTED FOR USE IN THE MODERATELY TORTUOUS AND SEVERELY CALCIFIED MID CIRCUMFLEX ARTERY. DURING THE PROCEDURE, AFTER THE ROTAWIRE DRIVE CROSSED THE LESION, ADVANCEMENT WAS NOT SUCCESSFUL. FURTHERMORE, A 1.25X20 NON-BOSTON SCIENTIFIC DEVICE DID CROSS. THE ROTAWIRE DRIVE WAS REPLACED AND A 1.50MM ROTAPRO BURR WAS ADVANCED. WHILE BURRING SEVERE ANGULATION OF THE CX OSTIUM LESION, THE ROTAPRO BURR SLIGHTLY DECELERATED BY APPROXIMATELY 10,000RPM AND STALLED. A LIGHT PULL ON THE ROTAPRO BURR WAS ATTEMPTED AND IT FRACTURED AT THE HYPOTUBE PORTION. DYNAGLIDE MODE WAS NOT SUCCESSFULLY USED. THE PROCEDURE WAS CANCELLED. AT AN UNKNOWN TIME, THE PATIENT EXPIRED. IT WAS FURTHER REPORTED THAT THE ROTAPRO BURR WAS NOT FRACTURED; RATHER, THE END OF THE ROTAPRO WAS FRACTURED/AVULSED, LEAVING THE LOWER PORTION OF THE ROTAPRO AND THE TIP OF THE ROTAWIRE IN THE CORONARY ARTERY, WHILE THE REST OF THE ROTAPRO WAS ABLE TO BE TAKEN OUT. THE TIP OF THE ROTAWIRE WAS BROKEN OFF AT THE SITE WHERE THE ROTAPRO FRACTURED/AVULSED. THE FRACTURED ROTAPRO BURR BECAME LODGED IN THE VESSEL, LEADING TO OCCLUSION OF THE CIRCUMFLEX (CX) ARTERY, WHICH RESULTED IN ACUTE MYOCARDIAL INFARCTION AND WAS LISTED AS THE OFFICIAL CAUSE OF PATIENT DEATH. MULTIPLE RETRIEVAL ATTEMPTS WERE MADE, INCLUDING TRYING 3-4 VARIATIONS OF SNARES, ATTEMPTING TO WRAP A WIRE AROUND THE BROKEN DEVICE, AND USING AN 8F GUIDEZILLA TO FACILITATE ACCESS. ADDITIONALLY, EFFORTS WERE MADE TO WIRE THE LAD AND USE A LOW-PRESSURE BALLOON IN AN ATTEMPT TO FREE UP THE BROKEN BURR AND ALLOW FOR SNARE REMOVAL. DESPITE THESE MEASURES, THE DEVICE COULD NOT BE RETRIEVED, AND THE PATIENT WENT INTO VENTRICULAR FIBRILLATION (VF) DUE TO ISCHEMIA FROM THE OCCLUDED CX. AN AUTOPSY WAS DECLINED AND NOT PERFORMED. THE PHYSICIAN STATED THAT HAD THE BURR NOT FRACTURED AND GOTTEN STUCK, THE PATIENT WOULD NOT HAVE GONE INTO VF.Event Description: IT WAS REPORTED THAT THE BURR FRACTURED AND THE PATIENT EXPIRED. A 1.50MM ROTAPRO AND A ROTAWIRE DRIVE WERE SELECTED FOR USE IN THE MODERATELY TORTUOUS AND SEVERELY CALCIFIED MID CIRCUMFLEX ARTERY. DURING THE PROCEDURE, AFTER THE ROTAWIRE DRIVE CROSSED THE LESION, ADVANCEMENT WAS NOT SUCCESSFUL. FURTHERMORE, A 1.25X20 NON-BOSTON SCIENTIFIC DEVICE DID CROSS. THE ROTAWIRE DRIVE WAS REPLACED AND A 1.50MM ROTAPRO BURR WAS ADVANCED. WHILE BURRING SEVERE ANGULATION OF THE CX OSTIUM LESION, THE ROTAPRO BURR SLIGHTLY DECELERATED BY APPROXIMATELY 10,000RPM AND STALLED. A LIGHT PULL ON THE ROTAPRO BURR WAS ATTEMPTED AND IT FRACTURED AT THE HYPOTUBE PORTION. DYNAGLIDE MODE WAS NOT SUCCESSFULLY USED. THE PROCEDURE WAS CANCELLED. AT AN UNKNOWN TIME, THE PATIENT EXPIRED. IT WAS FURTHER REPORTED THAT THE ROTAPRO BURR WAS NOT FRACTURED; RATHER, THE END OF THE ROTAPRO WAS FRACTURED/EVOLSED, LEAVING THE LOWER PORTION OF THE ROTAPRO AND THE TIP OF THE ROTAWIRE IN THE CORONARY ARTERY, WHILE THE REST OF THE ROTAPRO WAS ABLE TO BE TAKEN OUT. THE TIP OF THE ROTAWIRE WAS BROKEN OFF AT THE SITE WHERE THE ROTAPRO FRACTURED/EVOLSED. THE FRACTURED ROTAPRO BURR BECAME LODGED IN THE VESSEL, LEADING TO OCCLUSION OF THE CIRCUMFLEX (CX) ARTERY, WHICH RESULTED IN ACUTE MYOCARDIAL INFARCTION AND WAS LISTED AS THE OFFICIAL CAUSE OF PATIENT DEATH. MULTIPLE RETRIEVAL ATTEMPTS WERE MADE, INCLUDING TRYING 3-4 VARIATIONS OF SNARES, ATTEMPTING TO WRAP A WIRE AROUND THE BROKEN DEVICE, AND USING AN 8F GUIDEZILLA TO FACILITATE ACCESS. ADDITIONALLY, EFFORTS WERE MADE TO WIRE THE LAD AND USE A LOW-PRESSURE BALLOON IN AN ATTEMPT TO FREE UP THE BROKEN BURR AND ALLOW FOR SNARE REMOVAL. DESPITE THESE MEASURES, THE DEVICE COULD NOT BE RETRIEVED, AND THE PATIENT WENT INTO VENTRICULAR FIBRILLATION (VF) DUE TO ISCHEMIA FROM THE OCCLUDED CX. AN AUTOPSY WAS DECLINED AND NOT PERFORMED. THE PHYSICIAN STATED THAT HAD THE BURR NOT FRACTURED AND GOTTEN STUCK, THE PATIENT WOULD NOT HAVE GONE INTO VF. Manufacturer Narrative: GOOD FAITH EFFORT ATTEMPTS WERE MADE TO TRY AND RETRIEVE ADDITIONAL DETAILS REGARDING THE REPORTED EVENT, BUT FURTHER INFORMATION WAS UNABLE TO BE OBTAINED. Manufacturer Narrative: H6 - PATIENT CODE: UPDATED. H6 - IMPACT CODE: UPDATED. DEVICE EVALUATED BY MANUFACTURER. THE COMPLAINT DEVICE WAS RECEIVED FOR PRODUCT ANALYSIS. A VISUAL INSPECTION OF THE DEVICE FOUND THAT THE COIL WAS KINKED, STRETCHED AND DETACHED. THE TOTAL LENGTH OF THE COIL RETURNED WAS 141.5CM. THE BURR FAILED TO RETURN FOR FURTHER ANALYSIS. THE DAMAGE PRESENT DIRECTLY IMPACTS ROTATIONAL OUTPUT; HOWEVER, BASED ON THE REPORTED EVENTS, FURTHER TESTING WAS COMPLETED TO RULE OUT THE CAUSE OF REPORTED SPEED AND STALL ISSUE. AFTER DESTRUCTIVE TESTING WAS PERFORMED, IT WAS FOUND THAT THE DRIVE SHAFT (TURBINE) WAS CORRODED, INDICATING THE PRESENCE OF DRIED SALINE WITHIN THE DEVICE. DURING ANALYSIS, THE DYNAGLIDE MODE CHANGE BUTTON, THE DYNAGLIDE MOMENTARY ON BUTTON, AND THE ABLATION BUTTON WERE TESTED AND FOUND TO RESPOND WHEN PRESSED WITH NO RESISTANCE OR ISSUES. WHEN THE ROTAPRO ADVANCER WAS CONNECTED TO THE ROTAPRO CONSOLE AND THE ABLATION BUTTON WAS PRESSED, THE DEVICE STALLED AND DID NOT RUN. IN ORDER TO DETERMINE THE REASON FOR THE DEVICE STALL, THE ADVANCER WAS DISMANTLED AND THE COMPONENTS INSIDE THE ADVANCER WERE INSPECTED. A PHOTO OF A KINKED, STRETCHED, AND DETACHED COIL WAS ATTACHED TO THE COMPLAINT RECORD AND WAS ABLE TO CONFIRM THE REPORTED DETACHMENT. ADDITIONAL PHOTOS RECEIVED THROUGH GOOD FAITH EFFORT, SHOWED THE SAME COIL DAMAGE, A PORTION OF THE SHEATH AND COIL WITH NO DAMAGE NOTED, AND A VIDEO PROVIDED SHOWING THE SEPARATION OF THE COIL WITHIN THE BODY. NO OTHER ISSUES WERE IDENTIFIED DURING THE PRODUCT ANALYSIS. AS THERE ARE NO INDICATIONS OF A DEVICE MALFUNCTION OR DEFECT PRIOR TO ATTEMPTED USE WITHIN THE PROCEDURE, IT WAS CONSIDERED LIKELY THAT THE REPORTED DECELERATION, STALL, DETACHMENT, AND FOLLOWING CASCADE OF EVENTS WERE ATTRIBUTABLE TO PROCEDURAL FACTORS. BASED ON THE TARGET LESION CHARACTERISTICS REPORTED, IT WAS UNABLE TO BE DETERMINED IF THE SEVERITY OF THE TARGET LESION WAS A POTENTIAL CONTRIBUTING FACTOR TO THE REPORTED EVENTS. Manufacturer Narrative: H6 - PATIENT CODE: UPDATED. H6 - IMPACT CODE: UPDATED. THE DEVICE WAS NOT RETURNED FOR ANALYSIS. HOWEVER, A PHOTO OF A KINKED, STRETCHED, AND DETACHED COIL WAS ATTACHED TO THE COMPLAINT RECORD AND WAS ABLE TO CONFIRM THE REPORTED DETACHMENT. AN ADDITIONAL PHOTO OF THE DAMAGED PHOTO AND A VIDEO SHOWING THE SEPARATION OF THE COIL WITHIN THE BODY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21663006&pc=MCX | 2124215-2025-17314 | 2025-03-05 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-03-21 04:00 | MCX | ROTAWIRE DRIVE | Detachment of Device or Device Component; Difficult to Advance | Ischemia; No Clinical Signs, Symptoms or Conditions | P900056 |  | 1 | Event Description: IT WAS REPORTED THAT THE BURR FRACTURED AND THE PATIENT EXPIRED. A 1.50MM ROTAPRO AND A ROTAWIRE DRIVE WERE SELECTED FOR USE IN THE MODERATELY TORTUOUS AND SEVERELY CALCIFIED MID CIRCUMFLEX ARTERY. DURING THE PROCEDURE, AFTER THE ROTAWIRE DRIVE CROSSED THE LESION, ADVANCEMENT WAS NOT SUCCESSFUL. FURTHERMORE, A 1.25X20 NON-BOSTON SCIENTIFIC DEVICE DID CROSS. THE ROTAWIRE DRIVE WAS REPLACED AND A 1.50MM ROTAPRO BURR WAS ADVANCED. WHILE BURRING SEVERE ANGULATION OF THE CX OSTIUM LESION, THE ROTAPRO BURR SLIGHTLY DECELERATED BY APPROXIMATELY 10,000RPM AND STALLED. A LIGHT PULL ON THE ROTAPRO BURR WAS ATTEMPTED AND IT FRACTURED AT THE HYPOTUBE PORTION. DYNAGLIDE MODE WAS NOT SUCCESSFULLY USED. THE PROCEDURE WAS CANCELLED. AT AN UNKNOWN TIME, THE PATIENT EXPIRED.Event Description: IT WAS REPORTED THAT THE BURR FRACTURED AND THE PATIENT EXPIRED. A 1.50MM ROTAPRO AND A ROTAWIRE DRIVE WERE SELECTED FOR USE IN THE MODERATELY TORTUOUS AND SEVERELY CALCIFIED MID CIRCUMFLEX ARTERY. DURING THE PROCEDURE, AFTER THE ROTAWIRE DRIVE CROSSED THE LESION, ADVANCEMENT WAS NOT SUCCESSFUL. FURTHERMORE, A 1.25X20 NON-BOSTON SCIENTIFIC DEVICE DID CROSS. THE ROTAWIRE DRIVE WAS REPLACED AND A 1.50MM ROTAPRO BURR WAS ADVANCED. WHILE BURRING SEVERE ANGULATION OF THE CX OSTIUM LESION, THE ROTAPRO BURR SLIGHTLY DECELERATED BY APPROXIMATELY 10,000RPM AND STALLED. A LIGHT PULL ON THE ROTAPRO BURR WAS ATTEMPTED AND IT FRACTURED AT THE HYPOTUBE PORTION. DYNAGLIDE MODE WAS NOT SUCCESSFULLY USED. THE PROCEDURE WAS CANCELLED. AT AN UNKNOWN TIME, THE PATIENT EXPIRED. IT WAS FURTHER REPORTED THAT THE ROTAPRO BURR WAS NOT FRACTURED; RATHER, THE END OF THE ROTAPRO WAS FRACTURED/EVOLSED, LEAVING THE LOWER PORTION OF THE ROTAPRO AND THE TIP OF THE ROTAWIRE IN THE CORONARY ARTERY, WHILE THE REST OF THE ROTAPRO WAS ABLE TO BE TAKEN OUT. THE TIP OF THE ROTAWIRE WAS BROKEN OFF AT THE SITE WHERE THE ROTAPRO FRACTURED/EVOLSED. THE FRACTURED ROTAPRO BURR BECAME LODGED IN THE VESSEL, LEADING TO OCCLUSION OF THE CIRCUMFLEX (CX) ARTERY, WHICH RESULTED IN ACUTE MYOCARDIAL INFARCTION AND WAS LISTED AS THE OFFICIAL CAUSE OF PATIENT DEATH. MULTIPLE RETRIEVAL ATTEMPTS WERE MADE, INCLUDING TRYING 3-4 VARIATIONS OF SNARES, ATTEMPTING TO WRAP A WIRE AROUND THE BROKEN DEVICE, AND USING AN 8F GUIDEZILLA TO FACILITATE ACCESS. ADDITIONALLY, EFFORTS WERE MADE TO WIRE THE LAD AND USE A LOW-PRESSURE BALLOON IN AN ATTEMPT TO FREE UP THE BROKEN BURR AND ALLOW FOR SNARE REMOVAL. DESPITE THESE MEASURES, THE DEVICE COULD NOT BE RETRIEVED, AND THE PATIENT WENT INTO VENTRICULAR FIBRILLATION (VF) DUE TO ISCHEMIA FROM THE OCCLUDED CX. AN AUTOPSY WAS DECLINED AND NOT PERFORMED. THE PHYSICIAN STATED THAT HAD THE BURR NOT FRACTURED AND GOTTEN STUCK, THE PATIENT WOULD NOT HAVE GONE INTO VF. Manufacturer Narrative: GOOD FAITH EFFORT ATTEMPTS WERE MADE TO TRY AND RETRIEVE ADDITIONAL DETAILS REGARDING THE REPORTED EVENT, BUT FURTHER INFORMATION WAS UNABLE TO BE OBTAINED. Manufacturer Narrative: REMOVED DEVICE CODE - DIFFICULT TO ADVANCE. H6 - COMPONENT CODE: CORRECTED. H6 - DEVICE CODE: CORRECTED. H6: UPDATED PATIENT AND IMPACT CODE. THE COMPLAINT DEVICE WAS NOT RETURNED BY THE CUSTOMER. HOWEVER, A MEDIA ATTACHED TO THE PARENT RECORD SHOWS A PORTION OF THE HYPOTUBE BUT DOES NOT SHOW ANY EVIDENCE OF THE REPORTED ISSUE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21663094&pc=OZD | 1220648-2025-26869 | 2025-02-24 05:00 | Death | ABIOMED, INC. | 2025-03-21 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia | P140003 |  | 1 | Event Description: THE COMPLAINANT HAD A IMPELLA CP PUMP PLACED VIA THE 14FR SHEATH. THE TEAM HAD PLACED THE PUMP FOR THE SUPPORT OF A PATIENT ADMITTED IN WITH ACUTE MYOCARDIAL INFARCTION AND CARDIOGENIC SHOCK. THE TEAM DID NOT EXPLANT/PEEL AWAY THE 14FR SHEATH BUT RATHER LEFT IN THE FEMORAL. THE LEG WAS NOTED TO BE WARM AND PERFUSED. ADDITIONAL DETAILS NOTED THE PATIENT DIED DUE TO MULTI-ORGAN FAILURE. THE LEG WAS WEAK BUT PERFUSED ALL THE TIME. GIVEN THE SLIGHT REDUCTION IN PERFUSION THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA DEVICE AS AN ASSOCIATED FACTOR TO THE EVENT OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM, SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21663371&pc=NPV | 2025587-2025-02024 | 2025-02-24 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-21 04:00 | NPV | HARMONY TRANSCATHETER PULMONARY VALVE | Patient Device Interaction Problem; Migration | Atrial Fibrillation; Atrial Flutter; Bradycardia; Cardiac Arrest; Non specific EKG/ECG Changes; Tachycardia; Ventricular Fibrillation | P200046 |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING THE IMPLANT OF THE BIOPROSTHETIC TRANSCATHETER PULMONARY VALVE NON-SUSTAINED VENTRICULAR TACHYCARDIA (NSVT) OCCURRED. VENTRICULAR BIGEMY ALSO WAS NOTED. A CHEST X-RAY CONFIRMED A STABLE VALVE POSITION. RADIOGRAPHY REVIEW NOTED DEVICE MIGRATION. HOSPITALIZATION WAS PROLONGED AND TELEMETRY MONITORING OCCURRED FOR 3 DAYS FOLLOWING THE VALVE IMPLANT. AN INTRAVENOUS ANTI-ARRHYTHMIC AND ORAL BETA BLOCKERS WERE ADMINISTERED. THE PHYSICIAN INDICATED THE NSVT WAS POSSIBLY RELATED TO THE IMPLANT PROCEDURE AND THE VALVE. THE NSVT WAS REPORTED AS UNRESOLVED. HOLTER MONITORING OCCURRED ON DAYS 4-8 FOLLOWING THE VALVE IMPLANT. EIGHT DAYS FOLLOWING THE VALVE IMPLANT HOSPITALIZATION WAS REQUIRED FOR VENTRICULAR FIBRILLATION (VF) WHICH LED TO CARDIAC ARREST. CARDIOPULMONARY RESUSCITATION (CPR) AND DEFIBRILLATION WERE REQUIRED. HOWEVER, THE PATIENT DIED. THE DEATH WAS REPORTED AS SUDDEN CARDIAC DEATH. THE PHYSICIAN INDICATED THE VF PROBABLY RELATED THE VALVE IMPLANT PROCEDURE AND VALVE.Event Description: ADDITIONAL INFORMATION RECEIVED INDICATED THAT PRIOR TO IMPLANT OF THE VALVE, THE PATIENT WAS TAKING THE ANTI-ARRHYTHMIC METOPROLOL. NO PACEMAKER WAS PRESENT AT BASELINE. ALTHOUGH, THE IMPLANT RADIOGRAPHY NOTED MIGRATION, THE IMPLANT PHYSICIAN REPORTED THAT THE VALVE WAS IMPLANTED IN THE DESIRED LOCATION AND THE DISCHARGE RADIOGRAPHY DID NOT INDICATE MIGRATION. A PRIVATE AUTOPSY WAS PERFORMED BUT THE RESULTS WERE NOT MADE AVAILABLE.Event Description: ADDITIONAL INFORMATION NOTED THAT HOSPITALIZATION FOR THE VENTRICULAR FIBRILLATION OCCURRED 6 DAYS FOLLOWING THE VALVE IMPLANT. UNSPECIFIED MEDICATION WAS ALSO PROVIDED WITH THE VENTRICULAR FIBRILLATION EVENT.Event Description: ADDITIONAL INFORMATION PERTAINING TO THE HOLTER MONITORING WHICH STARTED 4 DAYS FOLLOWING THE VALVE IMPLANT UNTIL THE PATIENT¿S DEATH WERE RECEIVED. ON THE SIXTH DAY FOLLOWING THE VALVE IMPLANT, SINUS RHYTHM WAS NOTED WITH AVERAGE HEART RATE IN THE 50-60 BEATS PER MINUTE RANGE (BPM). OCCASIONAL VENTRICULAR ECTOPY, <5%, WAS OBSERVED AND MOSTLY IN ISOLATION AND COUPLETS. SUDDEN ONSET OF RAPID MULTIFORM VENTRICULAR TACHYCARDIA FOLLOWED BY VENTRICULAR FIBRILLATION OCCURRED. SINUS TACHYCARDIA VERSUS AN ATRIAL TACHYCARDIA WITH 2:1 CONDUCTION WAS SEEN THEREAFTER. SIGNIFICANT ST CHANGES WERE APPARENT. A SLOW WIDE COMPLEX RHYTHM AT A RATE OF 26 BPM. AN ESTIMATED 27 MINUTES ELAPSED BEFORE SIGNIFICANT BASELINE ARTIFACT WAS SEEN, WHICH MAY HAVE CORRELATED TO THE COMMENCEMENT OF RESUSCITATIVE EFFORTS. THE PHYSICIAN INDICATED THE IMPLANT DEPTH HAD CONTRIBUTED TO THE ARRHYTHMIA.Event Description: ADDITIONAL INFORMATION RECEIVED INDICATED THAT THE CHEST X-RAY THAT WAS PERFORMED CONFIRMED A STABLE VALVE POSITION.Event Description: IT WAS REPORTED THAT FOLLOWING THE IMPLANT OF THE BIOPROSTHETIC TRANSCATHETER PULMONARY VALVE NON-SUSTAINED VENTRICULAR TACHYCARDIA (NSVT) OCCURRED. VENTRICULAR BIGEMY ALSO WAS NOTED. A CHEST X-RAY CONFIRMED A STABLE VALVE POSITION. RADIOGRAPHY REVIEW NOTED DEVICE MIGRATION. HOSPITALIZATION WAS PROLONGED AND TELEMETRY MONITORING OCCURRED FOR 3 DAYS FOLLOWING THE VALVE IMPLANT. AN INTRAVENOUS ANTI-ARRHYTHMIC AND ORAL BETA BLOCKERS WERE ADMINISTERED. THE PHYSICIAN INDICATED THE NSVT WAS POSSIBLY RELATED TO THE IMPLANT PROCEDURE AND THE VALVE. THE NSVT WAS REPORTED AS UNRESOLVED. HOLTER MONITORING OCCURRED ON DAYS 4-8 FOLLOWING THE VALVE IMPLANT. EIGHT DAYS FOLLOWING THE VALVE IMPLANT HOSPITALIZATION WAS REQUIRED FOR VENTRICULAR FIBRILLATION (VF) WHICH LED TO CARDIAC ARREST. CARDIOPULMONARY RESUSCITATION (CPR) AND DEFIBRILLATION WERE REQUIRED. HOWEVER, THE PATIENT DIED. THE DEATH WAS REPORTED AS SUDDEN CARDIAC DEATH. THE PHYSICIAN INDICATED THE VF PROBABLY RELATED THE VALVE IMPLANT PROCEDURE AND VALVE. ADDITIONAL INFORMATION RECEIVED INDICATED THAT PRIOR TO IMPLANT OF THE VALVE, THE PATIENT WAS TAKING THE ANTI-ARRHYTHMIC METOPROLOL. NO PACEMAKER WAS PRESENT AT BASELINE. ALTHOUGH, THE IMPLANT RADIOGRAPHY NOTED MIGRATION, THE IMPLANT PHYSICIAN REPORTED THAT THE VALVE WAS IMPLANTED IN THE DESIRED LOCATION AND THE DISCHARGE RADIOGRAPHY DID NOT INDICATE MIGRATION. A PRIVATE AUTOPSY WAS PERFORMED BUT THE RESULTS WERE NOT MADE AVAILABLE. ADDITIONAL INFORMATION NOTED THAT HOSPITALIZATION FOR THE VENTRICULAR FIBRILLATION OCCURRED 6 DAYS FOLLOWING THE VALVE IMPLANT. UNSPECIFIED MEDICATION WAS ALSO PROVIDED WITH THE VENTRICULAR FIBRILLATION EVENT. ADDITIONAL INFORMATION PERTAINING TO THE HOLTER MONITORING FROM 4 DAYS FOLLOWING THE VALVE IMPLANT UNTIL THE PATIENT¿S DEATH WERE RECEIVED. ON THE SIXTH DAY FOLLOWING THE VALVE IMPLANT, SINUS RHYTHM WAS NOTED WITH AVERAGE HEART RATE IN THE 50-60 BEATS PER MINUTE RANGE (BPM). OCCASIONAL VENTRICULAR ECTOPY, 5%, WAS OBSERVED AND MOSTLY IN ISOLATION AND COUPLETS. SUDDEN ONSET OF RAPID MULTIFORM VENTRICULAR TACHYCARDIA FOLLOWED BY VENTRICULAR FIBRILLATION OCCURRED. SINUS TACHYCARDIA VERSUS AN ATRIAL TACHYCARDIA WITH 2:1 CONDUCTION WAS SEEN THEREAFTER. SIGNIFICANT ST CHANGES WERE APPARENT. A SLOW WIDE COMPLEX RHYTHM AT A RATE OF 26 BPM. AN ESTIMATED 27 MINUTES ELAPSED BEFORE SIGNIFICANT BASELINE ARTIFACT WAS SEEN, WHICH MAY HAVE CORRELATED TO THE COMMENCEMENT OF RESUSCITATIVE EFFORTS. THE PHYSICIAN INDICATED THE IMPLANT DEPTH HAD CONTRIBUTED TO THE ARRHYTHMIA. Manufacturer Narrative: UPDATED DATA: H6. CONCLUSION: AS THE EVENT REPORTED AN ADVERSE EVENT IN A PROCEDURE INVOLVING A MEDTRONIC DEVICE AN INVESTIGATION WAS REQUIRED. AS THE DEVICE WAS STILL IN USE AT THE TIME THIS INVESTIGATION WAS COMPLETED, NO DEVICE ANALYSIS COULD BE PERFORMED. A COMPREHENSIVE REVIEW OF THE DEVICE HISTORY RECORDS FOR THE DEVICE ASSOCIATED WITH THIS REPORTED EVENT WAS PERFORMED. IN THIS INSTANCE NO MANUFACTURING ISSUES OR SIGNALS, THAT MAY HAVE BEEN CAUSATIVE IN CONSIDERING THIS ISSUE WERE IDENTIFIED. THE REPORTED EVENT IS UNCONFIRMED AS THE DEVICE MET SPECIFICATIONS FOR THE REPORTED ISSUE. IT IS NOT POSSIBLE TO DETERMINE A CAUSE OF THE EVENT FROM THE INFORMATION PROVIDED. THE PRODUCT DESIGN OR MANUFACTURING PROCESSES OF THE MEDTRONIC DEVICE HAVE NOT BEEN IDENTIFIED AS THE CAUSE OF THE EVENT. THEREFORE, THE CAUSE OF THE EVENT IS UNDETERMINED. A MEDICAL SAFETY ASSESSMENT WAS PERFORMED FOR THIS EVENT: ADVERSE EVENTS/HARMS, PATIENT OUTCOME, AND TREATMENTS FOR THIS EVENT INCLUDE: NON-SUSTAINED VENTRICULAR TACHYCARDIA, VENTRICULAR BIGEMINY, TREATED WITH MEDICATIONS, AND CONTINUED MONITORING. VALVE MIGRATION WITH NO INTERVENTION. VENTRICULAR FIBRILLATION WITH SUBSEQUENT CARDIAC ARREST TREATED WITH CARDIOPULMONARY RESUSCITATION BUT THAT RESULTED IN DEATH. POTENTIAL CONTRIBUTING FACTORS (IF APPLICABLE): NONE PROVIDED. RELATEDNESS ASSESSMENT & RATIONALE: BASED ON THE TIMING OF THE EVENTS, INFORMATION RECEIVED, AND MEDICAL EXPERTISE, THE MIGRATION WAS LIKELY RELATED TO THE VALVE AS IT DID NOT REMAIN IN THE INTENDED LOCATION. THE ARRHYTHMIAS AND DEATH WERE ALSO LIKELY RELATED TO THE VALVE. PRODUCT LABELING WAS REVIEWED. THE PRIMARY HARMS APPEAR TO BE DOCUMENTED IN THE PRODUCT LABELING. RISK MANAGEMENT FILES (RMF) WERE REVIEWED. THE PRIMARY HARMS AND SEVERITIES APPEAR TO BE SUFFICIENTLY DOCUMENTED IN THE RMF. RECOMMENDED ACTION(S): BASED ON THE MEDICAL SAFETY ASSESSMENT OF THIS EVENT, NO FURTHER ACTION IS RECOMMENDED. THE HARMONY INSTRUCTIONS FOR USE (IFU) LISTS CONDUCTION DISTURBANCE, MIGRATION, AND DEATH AS POTENTIAL RISKS ASSOCIATED WITH THE TREATMENT PROCEDURES. THE IFU IS DEVELOPED FROM THE PRODUCT RISK DOCUMENTATION AS IS THE PRODUCT LABELLED ADVERSE EVENTS DOCUMENT. THERE IS NO IDENTIFIED INADEQUACY WITH THE IFU IN RELATION TO THIS EVENT THIS EVENT WILL BE INCLUDED IN MONITORING AND TRENDING. CONDUCTION DISTURBANCES, SUCH AS VENTRICULAR FIBRILLATION, ATRIAL FIBRILLATION, BRADYCARDIA, AND VENTRICULAR TACHYCARDIA, ARE KNOWN POTENTIAL ADVERSE EFFECTS PER THE DEVICE IFU. FACTORS THAT MAY IMPACT THE DEVELOPMENT OF CONDUCTION DISTURBANCES INCLUDE DEPTH OF IMPLANT, BASELINE CONDUCTION DEFECTS, AND ANATOMICAL CONSIDERATIONS. THE PHYSICIAN INDICATED THE IMPLANT DEPTH MAY HAVE BEEN A CONTRIBUTING FACTOR TO THE ARRHYTHMIA; HOWEVER, A CONCLUSIVE CAUSE OF THE REPORTED EVENTS COULD NOT BE DETERMINED. MIGRATION IS A KNOWN POTENTIAL ADVERSE EFFECT PER DEVICE IFU AND CAN BE CAUSED BY A VARIETY OF FACTORS, INCLUDING VALVE POSITIONING, PATIENT ANATOMY, CALCIFICATION LEVEL, SPARSELY CALCIFIED NATIVE ANATOMY PROVIDING INSUFFICIENT ANCHORING, ASYMMETRICAL ROOT CALCIFICATION, UNDER EXPANSION OF THE VALVE AND OTHERS. NO ROOT CAUSE CAN BE DETERMINED BASED ON THE INFORMATION PROVIDED. CARDIAC ARREST IS LISTED UNDER THE POTENTIAL ADVERSE EVENTS PER THE DEVICE IFU AND CAN BE RELATED TO SEVERAL FACTORS (EX. PROCEDURE, PATIENT COMORBIDITIES, ETC.). A RELATIONSHIP OF THE CARDIAC ARREST TO THE DEVICE OR PROCEDURE COULD NOT BE DETERMINED WITH THE LIMITED INFORMATION AVAILABLE. PATIENT DEATH IS A POTENTIAL RISK ASSOCIATED WITH THE IMPLANTATION OF A BIOPROSTHETIC TRANSCATHETER PULMONARY VALVE PER THE DEVICE IFU. A PROCEDURE- OR VALVE-RELATED DEATH IS AN INHERENT RISK WHEN THE PATIENT CONDITION IS SUCH THAT THE VALVE IS NEEDED TO SUSTAIN CARDIAC FUNCTION AND IT CAN OCCUR DESPITE AN IDEAL IMPLANT PROCEDURE OR DEVICE FUNCTIONALITY. BASED ON THE MEDICAL SAFETY ASSESSMENT, THE DEATH WAS LIKELY RELATED TO THE VALVE, BUT A CONCLUSIVE CAUSE COULD NOT BE DETERMINED. THE COMPLAINT INVESTIGATION DETERMINED THAT THIS EVENT IS FORESEEN IN RISK MANAGEMENT AND IS INCLUDED IN MONITORING/TRENDING. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: IMAGING REVIEW: PRE-PROCEDURAL SCREENING REPORT REVIEWED. NO LANDING ZONE PREDICTED. THE INFLOW OVERSIZING CRITERIA NOT MET. A PRE-IMPLANT ECHOCARDIOGRAM, (B)(6) 2025, NOTED A NORMAL LEFT VENTRICLE (LV) SIZE AND FUNCTION, TRIVIAL MITRAL REGURGITATION (MR), AND NO AORTIC INSUFFICIENCY. THE RIGHT VENTRICULE (RV) WAS MODERATELY DILATED WITH NORMAL APPEARING FUNCTION. THE DOPPLER WAVEFORM DENSITY WAS CONSISTENT WITH MODERATE PULMONARY REGURGITATION. TRIVIAL TRICUSPID REGURGITATION IDENTIFIED. THERE WAS NO OBVIOUS ARRYTHMIA ON THE ECHOCARDIOGRAM IMAGING CLIPS. THE IMPLANT INTRACARDIAC ECHOCARDIOGRAPHY/TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) (B)(6) 2025 NOTED SEVERE PULMONARY REGURGITATION (PR) AT PRE-DEPLOYMENT OF THE TRANSCATHETER VALVE. POST VALVE DEPLOYMENT, THE VALVE APPEARED TO ROCK IN THE RIGHT VENTRICULAR OUTFLOW TRACT (RVOT). THE TRANSCATHETER LEAFLET MOTION APPEARED NORMAL. SOME ECTOPY WAS NOTED. THERE WAS TURBULENCE ACROSS THE VALVE, NO SIGNIFICANT REGURGITATION NOTED. THE PEAK VELOCITY MEASURED ~1.8 METERS PER SECOND (M/SEC). MILD TRICUSPID REGURGITATION NOTED. FURTHER, THE IMPLANT ANGIOGRAPHY (B)(6) 2025 NOTED THE TRANSCATHETER PULMONARY VALVE WAS UNCOVERED, SUPRA-ANNULARLY WITH POOR APPOSITION. THE TRANSCATHETER VALVE WAS PULLED LOWER, ONE STRUT BELOW ANNULUS, WITH STILL POOR APPOSITION. POST-RELEASE OF THE TRANSCATHETER PULMONARY VALVE, RIGHT VENTRICULAR ANGIOGRAM (RV-GRAM NOTED ANNULAR POSITION WITH STILL SOME LACK OF APPOSITION. THERE WAS SOME ROCKING NOTED (BEST OBSERVED IN THE MOVING IMAGES). THE PULMONARY ANGIOGRAM (PA-GRAM) NOTED SOME LEAK INTO THE SINUS, BUT NO LARGE CENTRAL OR PARAVALVULAR LEAK OBSERVED. THE DISCHARGE ECHOCARDIOGRAM, (B)(6) 2025, NOTED FREQUENT ECTOPY. THE LV WAS NORMAL IN SIZE AND FUNCTION. THERE WAS TRIVIAL MR AND NO AORTIC INSUFFICIENCY. THE RV RIGHT VENTRICLE WAS MODERATELY DILATED WITH A DECREASE IN FUNCTION. THE ECTOPY MADE IT DIFFICULT TO DETERMINE THE SEVERITY. THERE WAS MILD TR WITH AN ESTIMATED RIGHT VENTRICULAR SYSTOLIC PRESSURE (RVSP) OF 21 MILLIMETERS OF MERCURY (MM HG) + CENTRAL VENOUS PRESSURE (CVP). THE VALVE APPEARED IN THE RVOT. MILD PR AND POSSIBLE MILD PARAVALVULAR LEAK (PVL). THE INFLOW/OUTFLOW JETS WERE NOT WELL VISUALIZED, AND A SMALL AMOUNT OF FLOW WAS NOTED ADJACENT TO VALVE HOUSING. THE PEAK VELOCITY ACROSS THE TRANSCATHETER VALVE WAS 2.27 M/S WITH A PEAK GRADIENT OF 21 MMHG. AN ECHOCARDIOGRAM PERFORMED 6 DAYS FOLLOWING THE VALVE IMPLANT (B)(6) 2025 NOTED A NORMAL LV SIZE. THERE WAS A DECREASED SYSTOLIC FUNCTION. THE CALCULATED EJECTION FRACTION (EF) MEASURED 38-49%. THERE WAS TRIVIAL MR AND NO AORTIC INSUFFICIENCY. THE RV WAS MODERATELY DILATED WITH DECREASED FUNCTION. TRIVIAL TR WAS OBSERVED. THE VALVE APPEARED SEATED IN THE ANNULAR POSITION WITH THE INFLOW OBSERVED IN THE RVOT. TRIVIAL PR NOTED. THE PEAK VELOCITY ACROSS THE TRANSCATHETER PULMONARY VALVE MEASURED 2.33 M/S WITH A PEAK GRADIENT OF 22 MMHG. THE ECHOCARDIOGRAMS CONCLUDED THE TRANSCATHETER PULMONARY VALVE APPEARED IMPLANTED IN AN ANNULAR POSITION WITH A ROCKING MOTION WITH NO OBVIOUS VALVE DYSFUNCTION, NO SIGNIFICANT PR, PVL, OR OBSTRUCTION. POST-IMPLANT, ECTOPY WAS NOTED, HOWEVER, THE ECHOCARDIOGRAM WAS NOT A DIAGNOSTIC TOOL FOR THE PATIENT RHYTHM. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID {{HARMONY-DCS}}; PRODUCT LOT/SERIAL NUMBER (B)(6); PRODUCT TYPE: {{DELIVERY CATHETER SYSTEM}}; IMPLANT DATE {{NA}}; EXPLANT DATE {{NA}}. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: A1 A2 B5 H6 - ANNEX B. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21663523&pc=NPT | 2015691-2025-02129 | 2024-10-12 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-21 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Insufficient Information | Insufficient Information | P140031 |  | 1 | Event Description: THIS COMPLAINT WAS OPENED FROM THV/TVT REGISTRY ALTERNATIVE SUMMARY REPORTING ADVERSE EVENT SUBMISSION FOR EVENTS RECEIVED BY EW DURING QUARTER 1 2025 FOR THE SAPIEN 3, SAPIEN 3 ULTRA AND SAPIEN 3 ULTRA RESILIA VALVE. MULTIPLE EVENTS WERE IDENTIFIED TO HAVE OCCURRED MORE THAN 1 YEAR POST PROCEDURE. READMISSION (VALVE RELATED) - DEATH EVENT 459 DAYS POST IMPLANT OF A 26MM SAPIEN 3 ULTRA VALVE FOR AN 87-YEAR-OLD MALE. Manufacturer Narrative: THE EVENT DOCUMENTED IN THIS COMPLAINT WAS RECEIVED BY EW FROM THE TRANSCATHETER VALVE THERAPY (TVT) REGISTRY DURING QUARTER 1 OF 2025 FOR THE SAPIEN XT, SAPIEN 3, SAPIEN 3 ULTRA AND SAPIEN 3 ULTRA RESILIA VALVE. THIS EVENT WAS IDENTIFIED TO HAVE OCCURRED MORE THAN 1-YEAR POST-INDEX PROCEDURE. THEREFORE, THE EVENT DOES NOT MEET FDA'S SUMMARY REPORTING EXEMPTION (E2016006) CRITERIA FOR TVT REGISTRY ADVERSE EVENTS. THE DEVICE IDENTIFICATION (DI) NUMBER FOR A 26MM EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE IS (B)(4). PER THE INSTRUCTIONS FOR USE (IFU), NON-EMERGENT REOPERATION, EMERGENCY CARDIAC SURGERY, REOPERATION, DEVICE MIGRATION OR MALPOSITION REQUIRING INTERVENTION AND DEVICE EXPLANT ARE LISTED AS POTENTIAL RISKS ASSOCIATED WITH THE DEVICE AND TRANSCATHETER VALVE REPLACEMENT PROCEDURE. THE IFU CAUTIONS THAT LONG-TERM DURABILITY HAS NOT BEEN ESTABLISHED FOR THE VALVE. REGULAR MEDICAL FOLLOW-UP IS ADVISED TO EVALUATE VALVE PERFORMANCE. ACCELERATED DETERIORATION OF THE VALVE DUE TO CALCIFIC DEGENERATION MAY OCCUR IN CHILDREN, ADOLESCENTS, OR YOUNG ADULTS AND IN PATIENTS WITH AN ALTERED CALCIUM METABOLISM. VALVE RECIPIENTS SHOULD BE MAINTAINED ON ANTICOAGULANT/ANTIPLATELET THERAPY, EXCEPT WHEN CONTRAINDICATED, AS DETERMINED BY THEIR PHYSICIAN. THIS DEVICE HAS NOT BEEN TESTED FOR USE WITHOUT ANTICOAGULATION. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. IT IS POSSIBLE THAT THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, HOWEVER GIVEN THE LIMITED INFORMATION AVAILABLE FURTHER ASSESSMENT OF THIS ADVERSE EVENT IS NOT POSSIBLE AT THIS TIME. VALVE RELATED READMISSIONS MAY BE DUE TO VALVE DYSFUNCTION (E.G., SEVERE, OR CLINICALLY SIGNIFICANT PARAVALVULAR LEAK, CENTRAL LEAK, SIGNIFICANT MALPOSITION, EARLY OR LATE ENDOCARDITIS, THROMBOSIS, STRUCTURAL VALVE DETERIORATION OR DEGENERATION) AND/OR OTHER REASONS UNRELATED TO THE EDWARDS DEVICES (E.G., TREATMENT OF OTHER PATIENT COMORBIDITIES). DURING THE MANUFACTURING PROCESS, ALL SAPIEN VALVES (ALL MODELS) ARE 100% VISUALLY INSPECTED FOR DEFECTS AND 100% TESTED UNDER PHYSIOLOGICAL BACKPRESSURE CONDITIONS PRIOR TO RELEASE FOR DISTRIBUTION. THEREFORE, IT IS HIGHLY UNLIKELY THAT A MANUFACTURING DEFECT OR DEVICE MALFUNCTION CONTRIBUTED TO THE EVENT. IN THIS CASE, READMISSION OCCURRED 459 DAYS POST-IMPLANT WITH NO INDICATION OF INTERVENTION. IT IS UNKNOWN IF THE EVENT WAS RELATED TO THE EDWARD'S DEVICES OR PATIENT CO-MORBIDITIES. NO ADDITIONAL INFORMATION WAS PROVIDED, AND NO ADDITIONAL INVESTIGATION IS POSSIBLE. DUE TO THE LIMITED INFORMATION AVAILABLE, THE DEFINITIVE CAUSE OR CONTRIBUTING FACTORS FOR THIS EVENT CANNOT BE DETERMINED. THE AVAILABLE INFORMATION DOES NOT REASONABLY SUGGEST THERE WAS A MALFUNCTION OF THE EDWARDS DEVICE OR THAT THE USE OR MISS-USE OF THE DEVICE CAUSED OR CONTRIBUTED TO THE EVENT. THERE IS NO ALLEGATION OF DEFICIENCIES RELATED TO THE IDENTITY, QUALITY, DURABILITY, RELIABILITY, SAFETY, LABELING EFFECTIVENESS, OR PERFORMANCE OF THIS EDWARD DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21663659&pc=DSR | 3007972010-2025-00021 | 2025-02-21 05:00 | Death | CVRX, INC | 2025-03-21 04:00 | DSR | BAROSTIM NEO2 | Adverse Event Without Identified Device or Use Problem | Respiratory Arrest; Insufficient Information | P180050 |  | 1 | Event Description: A BAROSTIM SYSTEM WAS IMPLANTED ON (B)(6) 2025. DAYS AFTER IMPLANTATION, THE PATIENT EXPIRED.Event Description: A BAROSTIM SYSTEM WAS IMPLANTED ON (B)(6) 2025. ON (B)(6) 2025, THE PATIENT EXPIRED. PER THE OPINING OF THE PHYSICIAN, THE CAUSE OF THE PATIENT'S DEATH WAS A RESPIRATORY ARREST UNRELATED TO BAROSTIM. Manufacturer Narrative: UPDATED FIELDS: B4, G3, G6, H2, H6, H11. CVRX ID# (B)(4). Manufacturer Narrative: UPDATED FIELDS: A3A, A6, B2, B4, B5, G3, G6, H2, H6, H11. CVRX ID # (B)(4). Manufacturer Narrative: THE RESULTS OF THE INVESTIGATION ARE INCONCLUSIVE AT THIS TIME, AND THE REPORTED DEVICE WAS NOT RETURNED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED EVENT COULD NOT BE CONCLUSIVELY DETERMINED. THE DEVICE HISTORY AND STERILIZATION RECORD FOR THIS DEVICE SERIAL NUMBER HAS BEEN REVIEWED. NO ISSUES OR DISCREPANCIES WERE NOTED DURING THIS REVIEW THAT WOULD HAVE CONTRIBUTED TO THE REPORTED EVENT. THE DEVICE MET MATERIAL, ASSEMBLY, AND QUALITY CONTROL REQUIREMENTS. CVRX ID# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21663902&pc=OZD | 1220648-2025-26871 | 2025-02-22 05:00 | Death | ABIOMED, INC. | 2025-03-21 04:00 | OZD | IMPELLA CP | Malposition of Device; Material Split, Cut or Torn | Cardiac Arrest | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WITH CARDIOMYOPATHY WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS) AND THE DEVICE BECAME MALPOSITIONED REQUIRING REPOSITIONING IN WHICH THE STERILE SLEEVE WAS DAMAGED IN THE PROCESS. THE PATIENT WOULD SUBSEQUENTLY EXPIRE. THE PATIENT CAME AS AN OUTPATIENT FOR AN ENDOSCOPY. DURING RECOVERY THE PATIENT WENT INTO PULSELESS ELECTRICAL ACTIVITY ARREST AND RECEIVED THREE ROUNDS OF CARDIOPULMONARY RESUSCITATION WITH RETURN OF SPONTANEOUS CIRCULATION. THE PATIENT WAS CANNULATED FOR EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) AND IMPELLA IMPLANTED FOR VENT. APPROXIMATELY TWO DAYS AFTER START OF THE SUPPORT AN IMPELLA IN AORTA ALARM TRIGGERED. ECHOCARDIOGRAM CONFIRMED, AND THE PHYSICIAN WAS ABLE TO REPOSITION THE PUMP BACK ACROSS THE AORTIC VALVE INTO THE LEFT VENTRICLE MEASURING 5.0 CM. IN THE PROCESS, THE STERILE SLEEVE WAS DAMAGED (NO ACTION WAS TAKEN REGARDING THE DAMAGED SLEEVE PER THE REPORT). THE PATIENT CONTINUED ON SUPPORT, BUT PROGNOSIS WAS "POOR." SIX DAYS LATER, THE PATIENT HAD A FLAT ELECTROENCEPHALOGRAPHY AND PUPILS WERE NON-REACTIVE. THE PATIENT WAS PLANNED TO UNDERGO BRAIN DEATH TESTING. TWO DAYS LATER, HOWEVER, IT WAS NOTED THE PATIENT QUICKLY DECOMPENSATED DESPITE BEING MAXED ON PRESSERS AND IMPELLA MCS, CODED AND WAS UNABLE TO BE REVIVED AFTER CARDIOPULMONARY RESUSCITATION. MEDICAL SAFETY REVIEW OF THE EVENT NOTED THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR TO THE EVENT GIVEN THE MALPOSITION OF THE DEVICE. THE PUMP MOVED, WAS DISPLACED FROM ITS INTENDED POSITION AND CAUSE IS UNKNOWN. Manufacturer Narrative: THE INVESTIGATION OF POSITIONING ISSUES AND PRODUCT DAMAGE (SLEEVE) HAS BEEN COMPLETED. THE PUMP WAS NOT RETURNED FOR INVESTIGATION. AS PER THE CLINICAL INFORMATION PROVIDED, THE STERILE SLEEVE WAS DAMAGED WHILE REPOSITIONING THE PUMP. THE CAUSE OF THE PRODUCT DAMAGE (SLEEVE) ISSUE WAS NOT DETERMINED SINCE PRODUCT/IMAGES WERE NOT RETURNED. AS PER THE CLINICAL INFORMATION PROVIDED, THE PATIENT WAS ROLLED OVER WHICH LED TO THE PUMP GETTING PULLED INTO THE AORTA. THEREFORE, THE ROOT CAUSE OF THE POSITIONING ISSUE WAS MOST LIKELY PATIENT MOVEMENT. E4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26871 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. H6 MEDICAL DEVICE PROBLEM CODE 4008 WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26871. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21664029&pc=BZD | 2518422-2025-103269 | 2025-03-17 04:00 | Death | RESPIRONICS, INC. | 2025-03-21 04:00 | BZD | DREAMSTATION CPAP PRO | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER HAS BEEN NOTIFIED OF THE RECENT PASSING OF A PATIENT WHO WAS USING A DREAMSTATION CPAP PRO DEVICE. THE PATIENT'S DAUGHTER HAS REQUESTED INFORMATION REGARDING THE RECALL AND WHETHER THE RECALLED DEVICE COULD BE LINKED TO HER MOTHER'S CANCER DIAGNOSIS. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21664457&pc=OZD | 1220648-2025-26875 | 2025-01-10 05:00 | Death | ABIOMED, INC. | 2025-03-21 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemolysis | P140003 |  | 1 | Event Description: A NOTIFICATION WAS RECEIVED VIA ABIOMED¿S LONG-TERM OUTCOME AND QUALITY INDICATOR IMPELLA REGISTRY REGARDING A PATIENT WHO EXPERIENCED ENDURED HEMOLYSIS WITH A "POSSIBLE" RELATIONSHIP TO THE IMPELLA DEVICE USED. THE PATIENT WAS A HEART TRANSPLANT REJECTION AND WOULD SUBSEQUENTLY EXPIRED. THE PATIENT WAS ON IMPELLA CP MECHANICAL CIRCULATORY SUPPORT PRIOR FOR HEART KIDNEY TRANSPLANT EVALUATION. THAT IMPELLA WAS SUCCESSFULLY WEANED. HOWEVER, THE PATIENT HAS SINCE DECOMPENSATED REQUIRING INCREASED INOTROPES/PRESSORS. THE PATIENT HAD A PULSELESS ELECTRICAL ACTIVITY ARREST AT THE OUTSIDE HOSPITAL REQUIRING VENOARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION (VA ECMO) AND IMPELLA CP PLACEMENT (ECPELLA). THE PATIENT WAS ALSO STARTED ON CONTINUOUS RENAL REPLACEMENT THERAPY. ON DAY 8 OF ECPELLA SUPPORT, THE PATIENT WAS TRANSITIONED TO BIPELLA SUPPORT AND VA ECMO WAS REMOVED. ADDITIONALLY, ON THIS DAY PER THE REGISTRY, APPROXIMATELY EIGHT DAYS AFTER THE START OF THE IMPELLA CP SUPPORT, DURING THE NIGHT, THE PATIENT HAD IMPELLA RP SUCTION ALARM AND WAS DECREASED TO PERFORMANCE LEVEL P-7. PER NOTES, THE IMPELLA CP ALSO DECREASED TO P-7 FOR HEMOLYSIS. THE PATIENT WAS TRANSFUSED WITH PACKED RED BLOOD CELLS DUE TO INCREASING LACTATE AND ESCALATING NOREPINEPHRINE. THERE WAS NO BLEEDING NOTED BUT MONITORED FOR SIGNS. PER THE REGISTRY, IT WAS REPORTED THAT THE PATIENT/EVENT DID NOT RESOLVE. THE PATIENT WOULD SUBSEQUENTLY EXPIRE FIVE DAYS LATER. IT WAS NOTED THE PATIENT WAS NOT IMPROVING IN HEART AND RENAL FAILURE AND THE PATIENT ASKED FOR WITHDRAWAL OF LIFE-SUSTAINING TREATMENT. ALTHOUGH CARE WAS WITHDRAWN, THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA CP DEVICE AS AN ASSOCIATED FACTOR GIVEN THE UNRESOLVED HEMOLYSIS EVENT. Manufacturer Narrative: THE PUMP WAS NOT RETURNED FOR INVESTIGATION. HOWEVER, DATA LOGS WERE RECEIVED AND ANALYZED. HEMOLYSIS WAS REPORTED DURING IMPELLA CP SUPPORT. CLINICAL INFORMATION MENTIONS BRIEF SUCTION ON THE IMPELLA CP AND IMPELLA RP FLEX PUMPS WHICH ULTIMATELY WAS RESOLVED. BLOOD WAS GIVEN. DATA LOG SHOWED BRIEF SUCTION ALARM. NO POSITIONING ISSUE OCCURRED. NO PLACEMENT SIGNAL TREND OBSERVED. NO MOTOR CURRENT SPIKE OUTSIDE OF P-LEVEL CHANGES OBSERVED. THE CAUSE OF THE HEMOLYSIS ISSUE COULD NOT BE DETERMINED SINCE THE COMPLAINT DEVICE NOT RETURNED FOR ANALYSIS AND INSUFFICIENT CLINICAL DATA WAS PROVIDED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM, SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER: ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE: ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS: ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION: ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21664875&pc=MNS | 2518422-2025-103286 | 2025-03-19 04:00 | Death | RESPIRONICS, INC. | 2025-03-21 04:00 | MNS | RENTAL, DREAMSTATION ST30, GB | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K102465 |  | 1 | Event Description: IT WAS REPORTED THAT THE RENTAL DREAMSTATION ST30 DEVICE WAS BEING RETURNED DUE TO THE USER PASSING AWAY. THERE WAS NO ALLEGATION OF MALFUNCTION OR THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH. THERE WERE NO REPORTS OF MEDICAL INTERVENTION. ADDITIONAL INFORMATION IS BEING REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21664879&pc=MNS | 2518422-2025-103280 | 2025-03-18 04:00 | Death | RESPIRONICS, INC. | 2025-03-21 04:00 | MNS | RENTAL, DREAMSTATION ST30, GB | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K102465 |  | 1 | Event Description: IT WAS REPORTED THAT THE RENTAL DREAMSTATION ST30 DEVICE WAS BEING RETURNED DUE TO THE USER PASSING AWAY. THERE IS NO ALLEGATION OF MALFUNCTION, OR THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH. THERE WERE NO REPORTS OF MEDICAL INTERVENTION. ADDITIONAL INFORMATION IS BEING REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.Event Description: IT WAS REPORTED THAT THE RENTAL DREAMSTATION ST30 DEVICE WAS BEING RETURNED DUE TO THE USER PASSING AWAY. THERE IS NO ALLEGATION OF MALFUNCTION OR THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH. THERE WERE NO REPORTS OF MEDICAL INTERVENTION. ADDITIONAL INFORMATION IS BEING REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER. DURING EVALUATION, NO ERROR WAS FOUND IN THE DEVICE LOG, AND THE DEVICE WAS AWAITING SCRAPPING CONFIRMATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. Manufacturer Narrative: IT WAS REPORTED THAT THE RENTAL DREAMSTATION ST30 DEVICE WAS BEING RETURNED DUE TO THE USER PASSING AWAY. THERE IS NO ALLEGATION OF MALFUNCTION OR THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH. THERE WERE NO REPORTS OF MEDICAL INTERVENTION. ADDITIONAL INFORMATION IS BEING REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER. DURING EVALUATION, NO ERROR WAS FOUND IN THE DEVICE LOG, AND THE DEVICE WAS AWAITING SCRAPPING CONFIRMATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21664996&pc=MVK | 3008642652-2025-02991 | 2025-02-24 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-03-21 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Signal Artifact/Noise | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED A NON-LIFEVEST DEFIBRILLATION. THE DEVICE WAS STARTED UP AT 11:22:09 ON (B)(6) 2025. AT 11:38:47 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF WITH VARYING AMPLITUDES AND MOTION ARTIFACT. THE DEVICE PROPERLY DETECTED VF. VARYING AMPLITUDES AND MOTION ARTIFACT PREVENTED THE LIFEVEST FROM TREATING THE PATIENT. AT 11:44:15, THE PATIENT RECEIVED A NON-LIFEVEST DEFIBRILLATION. RHYTHM AT THE TIME OF THE NON-LIFEVEST DEFIBRILLATION WAS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY. POST SHOCK RHYTHM WAS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY. THE DEVICE WAS SHUT DOWN AT 11:48:15 ON (B)(6) 2025. Manufacturer Narrative: THE ELECTRODE BELT AND MONITOR HAVE BEEN RETURNED AND THE EVALUATION IS UNDERWAY. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21665004&pc=NPT | 2015691-2025-02149 | 2024-09-09 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-21 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Insufficient Information | Insufficient Information | P140031 |  | 1 | Event Description: THIS COMPLAINT WAS OPENED FROM THV/TVT REGISTRY ALTERNATIVE SUMMARY REPORTING ADVERSE EVENT SUBMISSION FOR EVENTS RECEIVED BY EW DURING QUARTER 1 2025 FOR THE SAPIEN 3, SAPIEN 3 ULTRA AND SAPIEN 3 ULTRA RESILIA VALVE. MULTIPLE EVENTS WERE IDENTIFIED TO HAVE OCCURRED MORE THAN 1 YEAR POST PROCEDURE. READMISSION (VALVE RELATED) - DEATH EVENT 411 DAYS POST IMPLANT OF A 29MM SAPIEN 3 ULTRA RESILIA VALVE FOR AN 69-YEAR-OLD MALE. Manufacturer Narrative: THE EVENT DOCUMENTED IN THIS COMPLAINT WAS RECEIVED BY EW FROM THE TRANSCATHETER VALVE THERAPY (TVT) REGISTRY DURING QUARTER 1 OF 2025 FOR THE SAPIEN XT, SAPIEN 3, SAPIEN 3 ULTRA AND SAPIEN 3 ULTRA RESILIA VALVE. THIS EVENT WAS IDENTIFIED TO HAVE OCCURRED MORE THAN 1-YEAR POST-INDEX PROCEDURE. THEREFORE, THE EVENT DOES NOT MEET FDA'S SUMMARY REPORTING EXEMPTION (E2016006) CRITERIA FOR TVT REGISTRY ADVERSE EVENTS. THE DEVICE IDENTIFICATION (DI) NUMBER FOR A 29MM EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IS (B)(4). PER THE INSTRUCTIONS FOR USE (IFU), NON-EMERGENT REOPERATION, EMERGENCY CARDIAC SURGERY, REOPERATION, DEVICE MIGRATION OR MALPOSITION REQUIRING INTERVENTION AND DEVICE EXPLANT ARE LISTED AS POTENTIAL RISKS ASSOCIATED WITH THE DEVICE AND TRANSCATHETER VALVE REPLACEMENT PROCEDURE. THE IFU CAUTIONS THAT LONG-TERM DURABILITY HAS NOT BEEN ESTABLISHED FOR THE VALVE. REGULAR MEDICAL FOLLOW-UP IS ADVISED TO EVALUATE VALVE PERFORMANCE. ACCELERATED DETERIORATION OF THE VALVE DUE TO CALCIFIC DEGENERATION MAY OCCUR IN CHILDREN, ADOLESCENTS, OR YOUNG ADULTS AND IN PATIENTS WITH AN ALTERED CALCIUM METABOLISM. VALVE RECIPIENTS SHOULD BE MAINTAINED ON ANTICOAGULANT/ANTIPLATELET THERAPY, EXCEPT WHEN CONTRAINDICATED, AS DETERMINED BY THEIR PHYSICIAN. THIS DEVICE HAS NOT BEEN TESTED FOR USE WITHOUT ANTICOAGULATION. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. IT IS POSSIBLE THAT THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, HOWEVER GIVEN THE LIMITED INFORMATION AVAILABLE FURTHER ASSESSMENT OF THIS ADVERSE EVENT IS NOT POSSIBLE AT THIS TIME. VALVE RELATED READMISSIONS MAY BE DUE TO VALVE DYSFUNCTION (E.G., SEVERE, OR CLINICALLY SIGNIFICANT PARAVALVULAR LEAK, CENTRAL LEAK, SIGNIFICANT MALPOSITION, EARLY OR LATE ENDOCARDITIS, THROMBOSIS, STRUCTURAL VALVE DETERIORATION OR DEGENERATION) AND/OR OTHER REASONS UNRELATED TO THE EDWARDS DEVICES (E.G., TREATMENT OF OTHER PATIENT COMORBIDITIES). DURING THE MANUFACTURING PROCESS, ALL SAPIEN VALVES (ALL MODELS) ARE 100% VISUALLY INSPECTED FOR DEFECTS AND 100% TESTED UNDER PHYSIOLOGICAL BACKPRESSURE CONDITIONS PRIOR TO RELEASE FOR DISTRIBUTION. THEREFORE, IT IS HIGHLY UNLIKELY THAT A MANUFACTURING DEFECT OR DEVICE MALFUNCTION CONTRIBUTED TO THE EVENT. IN THIS CASE, READMISSION OCCURRED 411 DAYS POST-IMPLANT WITH NO INDICATION OF INTERVENTION. IT IS UNKNOWN IF THE EVENT WAS RELATED TO THE EDWARD'S DEVICES OR PATIENT CO-MORBIDITIES. NO ADDITIONAL INFORMATION WAS PROVIDED, AND NO ADDITIONAL INVESTIGATION IS POSSIBLE. DUE TO THE LIMITED INFORMATION AVAILABLE, THE DEFINITIVE CAUSE OR CONTRIBUTING FACTORS FOR THIS EVENT CANNOT BE DETERMINED. THE AVAILABLE INFORMATION DOES NOT REASONABLY SUGGEST THERE WAS A MALFUNCTION OF THE EDWARDS DEVICE OR THAT THE USE OR MISS-USE OF THE DEVICE CAUSED OR CONTRIBUTED TO THE EVENT. THERE IS NO ALLEGATION OF DEFICIENCIES RELATED TO THE IDENTITY, QUALITY, DURABILITY, RELIABILITY, SAFETY, LABELING EFFECTIVENESS, OR PERFORMANCE OF THIS EDWARD DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21665242&pc=LWS | 2124215-2025-17594 | 2025-03-06 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-03-21 04:00 | LWS | EMBLEM S-ICD | Fracture; Failure to Convert Rhythm; Low impedance; Defective Device | No Clinical Signs, Symptoms or Conditions | P110042/S077 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE IMPLANT PROCEDURE OF A NEW SUBCUTANEOUS IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (S-ICD) DEVICE, THE PHYSICIAN PERFORMED STANDARD SHOCK TESTING. A VENTRICULAR FIBRILLATION (VF) WAS TRIGGERED, PROVIDING A 65J SHOCK DELIVERY. THE DEVICE DETECTED THE VF, HOWEVER, THE ARRHYTHMIA WAS UNABLE TO BE CONVERTED. EXTERNAL DEFIBRILLATION WAS NEEDED TO RETURN TO NORMAL SINUS RHYTHM, WITH A SHOCK IMPEDANCE OF LESS THAN 30 OHMS. THE PHYSICIAN ELECTED SUSPECTED POTENTIAL ELECTRODE DAMAGE, AND IT WAS SUBSEQUENTLY EXPLANTED. THE S-ICD DEVICE WAS ALSO REMOVED PRIOR TO WOUND CLOSURE, AND THE POCKET WAS CLOSED. AT THIS TIME, THIS PATIENT REMAINS HOSPITALIZED PENDING A TRANSVENOUS SYSTEM IMPLANT PROCEDURE, AND NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS ELECTRODE IS EXPECTED TO BE RETURNED FOR ANALYSIS.Event Description: IT WAS REPORTED THAT DURING THE IMPLANT PROCEDURE OF A NEW SUBCUTANEOUS IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (S-ICD) DEVICE, THE PHYSICIAN PERFORMED STANDARD SHOCK TESTING. A VENTRICULAR FIBRILLATION (VF) WAS TRIGGERED, PROVIDING A 65J SHOCK DELIVERY. THE DEVICE DETECTED THE VF, HOWEVER, THE ARRHYTHMIA WAS UNABLE TO BE CONVERTED. EXTERNAL DEFIBRILLATION WAS NEEDED TO RETURN TO NORMAL SINUS RHYTHM, WITH A SHOCK IMPEDANCE OF LESS THAN 30 OHMS. THE PHYSICIAN ELECTED SUSPECTED POTENTIAL ELECTRODE DAMAGE, AND IT WAS SUBSEQUENTLY EXPLANTED. THE S-ICD DEVICE WAS ALSO REMOVED PRIOR TO WOUND CLOSURE, AND THE POCKET WAS CLOSED. UNFORTUNATELY, THE PATIENT PASSED AWAY SEVERAL DAYS AFTER THIS PROCEDURE. THIS ELECTRODE HAS BEEN RECEIVED FOR ANALYSIS AND IS PENDING THE INVESTIGATION CONCLUSION. ONCE THE ANALYSIS IS COMPLETED, THIS RECORD WILL BE UPDATED WITH RESULTS. Manufacturer Narrative: AT THIS TIME, NO FURTHER INFORMATION IS AVAILABLE. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE IN THE FUTURE, WE WILL PROVIDE A SUPPLEMENTAL REPORT. Manufacturer Narrative: THIS REPORT IS BEING SENT TO PROVIDE NEW INFORMATION IN SECTIONS B2 (OUTCOMES ATTRIB TO ADV EVENT), B5 (EVENT DESCRIPTION), D6B (EXPLANT DATE), H1 (TYPE OF REPORTABLE EVENT), AND H6 (IMPACT CODES).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21665254&pc=OUT | 2029214-2025-00733 | 2024-10-01 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-03-21 04:00 | OUT | PIPELINE FLEX W/SHIELD TECHNOLOGY | Adverse Event Without Identified Device or Use Problem | Ischemia Stroke | P100018 |  | 1 | Event Description: GOERTZ, L., HOHENSTATT, S., VOLLHERBST, D. F., PFLAEGING, M., GRONEMANN, C., SIEBERT, E., ZOPFS, D., PENNIG, L., KOTTLORS, J., SCHLAMANN, M., BOHNER, G., DORN, F., LIEBIG, T., MÖHLENBRUCH, M., & KABBASCH, C. (2024). MULTICENTER EXPERIENCE WITH THE PIPELINE FLEX AND VANTAGE WITH SHIELD TECHNOLOGY FOR INTRACRANIAL ANEURYSM TREATMENT. AJNR. AMERICAN JOURNAL OF NEURORADIOLOGY, 45(10), 1488¿1494. HTTPS://DOI.ORG/10.3174/AJNR.A8352. MEDTRONIC REVIEW OF THE LITERATURE ARTICLE FOUND A MULTICENTER RETROSPECTIVE REVIEW OF 141 PATIENTS WHO HAD PIPELINE EMBOLIZATION DE VICES IMPLANTED TO TREAT 147 ANEURYSMS BETWEEN 2017 AND 2023. IN 112 CASES AT 4 CENTERS BETWEEN 2020 AND 2023, PIPELINE VANTAGE STENTS WERE IMPLANTED AND IN 32 CASES IN 3 CENTERS BETWEEN 2017 AND 2023 PIPELINE FLEX WITH SHIELD STENTS WERE IMPLANTED. NAVIEN 5F, PHENOM 21, PHENOM 27, AND REBAR 27 CATHETERS WERE USED IN SOME CASES AS WELL AS OTHER MANUFACTURERS' CATHETER. IT WAS NOT SPECIFIED WHICH CATHETER(S) WERE USED IN EACH PROCEDURE. NO DEVICE RELATED ISSUES OR COMPLICATIONS WERE REPORTED IN THE ARTICLE ASSOCIATED WITH THE CATHETER. ALL CASES WERE TECHNICALLY SUCCESSFUL, HOWEVER, IT WAS NOTED THAT MORE THAN ONE PIPELINE WAS IMPLANTED IN UNSPECIFIED "SELECT CASES" WITH A DYSPLASTIC PARENT VESSEL, FOR INCOMPLETE WALL APPOSITION, OR FOR FULL ANEURYSM NECK COVERAGE; IT WAS NOT SPECIFIED IN HOW MANY CASES THIS MAY HAVE OCCURRED. IT WAS ADDITIONALLY MENTIONED THAT BALLOON ANGIOPLASTY WAS USED FOR IMPROVED APPOSITION IN 20 PROCEDURES. 5 PATIENT DEATHS WERE REPORTED IN THE ARTICLE. HOWEVER, 3 OF THE DEATHS WERE REPORTED TO BE DUE TO PATIENT CONDITION OF SUBARACHNOID HEMORRHAGE (SAH) AND NOT TREATMENT RELATED. 2 OTHER PATIENTS, BOTH TREATED WITH PIPELINE FLEX WITH SHIELD STENTS FOR BASILAR ANEURYSMS EXPERIENCED IN-STENT OCCLUSION. ONE PATIENT EXPERIENCED OCCLUSION ON POST-OPERATIVE DAY 2 DESPITE ADHERENCE TO DUAL ANTIPLATELET TREATMENT (DAPT). THE OTHER PATIENT EXPERIENCED OCCLUSION ON POST-OPERATIVE DAY 3 DUE TO MISSED DAPT. BOTH PATIENTS DIED OF BRAINSTEM INFARCTION. Manufacturer Narrative: ADVERSE EVENTS THAT DID NOT HAVE OUTCOME OF MORTALITY WILL BE REPORTED SEPARATELY. A2. REPORTED PATIENT AGE IS THE MEAN AGE FOR ALL PATIENTS INCLUDED IN THE ARTICLE STUDY GROUP. A3. REPORTED PATIENT SEX IS REPRESENTATIVE OF THE MAJORITY OF PATIENT INCLUDED IN THE ARTICLE STUDY GROUP. B3. THE ARTICLE WAS PUBLISH OCTOBER 2024. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21665275&pc=OZD | 1220648-2025-26878 | 2025-02-19 05:00 | Death | ABIOMED, INC. | 2025-03-21 04:00 | OZD | IMPELLA 5.5 | Increase in Pressure | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA 5.5 DEVICE FOR MECHANICAL CIRCULATORY SUPPORT, AND APPROXIMATELY THREE DAYS INTO SUPPORT, BLEEDING DEVELOPED FROM ARTIFICIAL VASCULAR ANASTOMOSIS. A BLOOD TRANSFUSION OF AN UNKNOWN QUANTITY WAS PERFORMED, AND ANTICOAGULATION WAS PUT ON HOLD IN RESPONSE TO THE BLEED. THIS RESULTED IN AN INCREASE IN PURGE PRESSURE, AND DUE TO CONCERNS OF A POTENTIAL PUMP STOP, THE DECISION WAS MADE TO REPLACE THE PUMP WHICH WAS SUCCESSFULLY COMPLETED. HOWEVER, THE PATIENT WOULD EXPIRE 23 DAYS LATER. IT WAS NOTED HEART FAILURE PERSISTED AND THE PATIENT EXPIRED. THE RELATIONSHIP BETWEEN THE BLEEDING AND IMPELLA 5.5 DEVICE WAS NOTED TO BE UNKNOWN. MEDICAL SAFETY REVIEW NOTED THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA 5.5 PUMP 1 AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME (DEMISE). BLEEDING DEVELOPED FROM THE ARTIFICIAL VASCULAR ANASTOMOSIS. A BLOOD TRANSFUSION OF UNKNOWN QUANTITY WAS PERFORMED AND ANTICOAGULATION WITHHELD. Manufacturer Narrative: THE INVESTIGATION OF HIGH PURGE PRESSURE AND VASCULAR DAMAGE - PERIPHERAL VESSEL HAS BEEN COMPLETED. HIGH PURGE PRESSURE: CLINICAL DETAILS REPORT THAT ON (B)(6) 2025, AN INCREASE IN PURGE PRESSURE WAS NOTED, THOUGH NOT HIGH ENOUGH TO TRIGGER AN ALARM. EARLIER THE SAME DAY, ANTICOAGULATION WAS STOPPED TO TROUBLESHOOT BLEEDING. DATA LOG REVIEW CONFIRMED THE RISE IN PURGE PRESSURE. THE PURGE PRESSURE ROSE OVER ~2 MINUTES TO 850 MMHG. THERE WERE NO MOTOR CURRENT SPIKES LEADING UP TO THE EVENT, BUT MOTOR CURRENT MODULATION DID DAMPEN. THE PURGE PRESSURE REMAINED ELEVATED THROUGH (B)(6) 2025, AND DURING THIS TIME THERE WAS AN INCREASED JAGGED AMPLITUDE IN THE PURGE PRESSURE. ON THE (B)(6) 2025, IT DROPPED TO PREVIOUS LEVELS WITHIN A FEW MINUTES. THE ONLY OTHER PUMP METRIC THAT SHOWED ANY CHANGE AT THE SAME TIME WAS A SLIGHT STEP UP IN THE PS. RETURNED PRODUCT WAS INVESTIGATED AND THERE WERE NO KINKS IN THE CATHETER, BIOMATERIAL IN THE PURGE GAP, OR ANY OTHER VISUAL ABNORMALITIES. THE PUMP AND PURGE CASSETTE WERE CONNECTED TO THE CONSOLE, SUCCESSFULLY PRIMED AND RUN FOR OVER AN HOUR WITHOUT HIGH PURGE PRESSURE. THAT BEING SAID, PUMP FLOWS WERE SATURATED FOR THE FIRST HALF OF THE HOUR, BEFORE DROPPING WITHIN SPECIFICATION ON THEIR OWN. THE PATTERNS NOTED FROM DATA LOG REVIEW AND RETURNED PRODUCT COULD BE EVIDENCE OF SEVERAL CAUSES FOR HIGH PURGE PRESSURE. THEREFORE, BASED ON THE INCONCLUSIVE EVIDENCE FROM DATA LOG AND PRODUCT REVIEW, THE ROOT CAUSE OF THE HIGH PURGE PRESSURE COULD NOT BE DETERMINED. VASCULAR DAMAGE: IT WAS REPORTED THAT ON THE 4TH DAY OF SUPPORT BLEEDING WAS NOTED FROM THE VASCULAR GRAFT. THE BLEED HAD TO BE TREATED SURGICALLY BY OPEN HEMOSTASIS AND TRANSFUSIONS. DATA LOGS ARE NOT RELEVANT TO THE COMPLICATION. RETURNED PRODUCT WAS INVESTIGATED AND THERE WERE NO VISUAL ABNORMALITIES NOTED. SINCE LIMITED CLINICAL DETAILS WERE PROVIDED AND RETURNED PRODUCT DID NOT EXHIBIT ANY ABNORMALITIES, THE ROOT CAUSE OF THE VASCULAR DAMAGE COULD NOT BE DETERMINED. H6 HEALTH EFFECT - IMPACT CODE 1802 DEATH WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26878. Manufacturer Narrative: THE IMPELLA DEVICE WAS RETURNED, AND AN EVALUATION IS UNDERWAY. UPON COMPLETION OF THE ANALYSIS, A SUPPLEMENTAL REPORT WILL BE FILED. SECTION: WARNINGS & CAUTIONS: CAUTIONS ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH KNOWN OR SUSPECTED UNREPAIRED ABDOMINAL AORTIC ANEURYSM OR SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM OR DISSECTION OF THE ASCENDING, TRANSVERSE, OR DESCENDING AORTA.¿ CAUTION: ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS SECTION: PRE-SUPPORT EVALUATION SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER SECTION: DIRECT AORTIC INSERTION SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP. CLINICIANS SHOULD MONITOR MOTOR CURRENT AND CONSIDER REPLACING THE IMPELLA CATHETER WHENEVER A RISE IN MOTOR CURRENT IS SEEN.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21665344&pc=DSQ | 2916596-2025-01650 | 2025-02-12 05:00 | Death | THORATEC CORPORATION | 2025-03-21 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2025 DUE TO A CEREBROVASCULAR ACCIDENT (CVA). Manufacturer Narrative: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF STROKE COULD NOT BE CONCLUSIVELY DETERMINED. STROKE IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. STROKE HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: A4: PATIENT WEIGHT NOT PROVIDED, REQUEST FOR THIS INFORMATION WAS MADE. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21665597&pc=MCX | 3004742232-2025-00060 | 2025-02-24 05:00 | Death | CARDIOVASCULAR SYSTEMS INCORPORATED | 2025-03-21 04:00 | MCX | DIAMONDBACK CORONARY | Adverse Event Without Identified Device or Use Problem | Angina; Cardiac Arrest; Perforation of Vessels; Obstruction/Occlusion | P130005 |  | 1 | Event Description: IT WAS REPORTED THAT THE TARGET LESION WAS IN LEFT MAIN CORONARY ARTERY (LM) WHICH WAS 95% STENOSED, HEAVILY CALCIFIED AND NON-TORTUOUS. ACCESS WAS GAINED THROUGH THE LEFT FEMORAL ARTERY. THE HEART PUMP COULD NOT BE ADVANCED THROUGH THE LEFT ILIAC DUE TO HEAVY CALCIFICATION. THE PATIENT HAD PORCELAIN AORTA; THE OUTLINE OF AORTA WAS VISIBLE IN ANGIOGRAM WITHOUT CONTRAST DUE TO FULL CALCIFICATION. INTRAVASCULAR LITHOTRIPSY (IVL) WAS PERFORMED IN THE LEFT ILIAC. THE HEART PUMP SHEATH WAS ADVANCED AND THE HEART PUMP WAS INSERTED. THE CIRCUMFLEX ARTERY (CFX) WAS WIRED BUT THE LEFT ANTERIOR DESCENDING ARTERY (LAD) COULD NOT BE WIRED DUE TO CALCIUM BLOCKAGE. INTRAVASCULAR ULTRASOUND (IVUS) WAS USED DOWN THE CFX AND LM. A NON-ABBOTT DUAL-LUMEN CATHETER WAS USED TO ASSIST WITH WIRING THE LAD. THE DUAL-LUMEN CATHETER WAS REMOVED. IVUS WAS READVANCED TO THE LAD BUT WAS UNSUCCESSFUL. THE VIPERWIRE ADVANCE CORONARY GUIDE WIRE WITH FLEX TIP WAS IN PLACE. THE DIAMONDBACK 360 PRECISION CORONARY ORBITAL ATHERECTOMY DEVICE (OAD) WAS ADVANCED TO THE LAD. ONE LOW-SPEED TREATMENT AND THREE HIGH-SPEED TREATMENTS WERE PERFORMED. THE PATIENT EXPERIENCED CHEST PAIN. THE OAD WAS REMOVED. ANGIOGRAPHIC IMAGING SHOWED PERFORATION IN LM OSTIAL LAD. TO RESOLVE THE PERFORATION, THE PHYSICIAN TRIED TO PLACE A NON-ABBOTT COVERED STENT IN LAD BUT WAS UNSUCCESSFUL DUE TO THE STENT REPEATEDLY GOING THROUGH THE PERFORATION. THE STENT ENDED UP BEING PLACED IN DISTAL LM INTO THE LEFT CFX. THE PHYSICIAN THOUGHT THE PERFORATION WAS RESOLVED, BUT THERE WAS NO FLOW TO THE LAD. THE PATIENT LOST PULSE. CARDIOPULMONARY RESUSCITATION (CPR) PROTOCOL WAS ADMINISTERED. EPINEPHRINE AND OTHER STANDARD MEDICATIONS WERE GIVEN. PERICARDIOCENTESIS WAS PERFORMED. CPR CONTINUED. PATIENT REGAINED PULSE FOR A WHILE. A TEMPORARY PACEMAKER WAS INSERTED. THE HEART PUMP WAS STILL IN PLACE. CPR WAS CONTINUED. THE PATIENT LOST PULSE AGAIN. AFTER A BRIEF TIME, PATIENT EXPIRATION WAS PRONOUNCED. SOMETIME DURING THE CODE, THE PATIENT ALSO RECEIVED BLOOD TRANSFUSION AND WAS INTUBATED DURING CPR. THE PHYSICIAN INDICATED HE HAS NEVER SEEN AN OAD PERFORATE. THE PRIMARY CAUSE OF DEATH WAS PERFORATION BY OAD. THERE IS NO ALLEGATION OF MALFUNCTION AGAINST THE VIPERWIRE.Event Description: IT WAS REPORTED THAT THE PATIENT WAS A CORONARY ARTERY BYPASS GRAFTING (CABG) TURNDOWN WHO WANTED SOMETHING DONE AND DID NOT WANT TO BE SENT HOME AND MEDICALLY MANAGED. THE TARGET LESION WAS IN LEFT MAIN CORONARY ARTERY (LM) WHICH WAS 95% STENOSED, HEAVILY CALCIFIED AND NON-TORTUOUS. ACCESS WAS GAINED THROUGH THE LEFT FEMORAL ARTERY. THE HEART PUMP COULD NOT BE ADVANCED THROUGH THE LEFT ILIAC DUE TO HEAVY CALCIFICATION. THE PATIENT HAD PORCELAIN AORTA; THE OUTLINE OF AORTA WAS VISIBLE IN ANGIOGRAM WITHOUT CONTRAST DUE TO FULL CALCIFICATION. INTRAVASCULAR LITHOTRIPSY (IVL) WAS PERFORMED IN THE LEFT ILIAC. THE HEART PUMP SHEATH WAS ADVANCED AND THE HEART PUMP WAS INSERTED. THE CIRCUMFLEX ARTERY (CFX) WAS WIRED BUT THE LEFT ANTERIOR DESCENDING ARTERY (LAD) COULD NOT BE WIRED DUE TO CALCIUM BLOCKAGE. INTRAVASCULAR ULTRASOUND (IVUS) WAS USED DOWN THE CFX AND LM. A NON-ABBOTT DUAL-LUMEN CATHETER WAS USED TO ASSIST WITH WIRING THE LAD. THE DUAL-LUMEN CATHETER WAS REMOVED. IVUS WAS READVANCED TO THE LAD BUT WAS UNSUCCESSFUL. THE VIPERWIRE ADVANCE CORONARY GUIDE WIRE WITH FLEX TIP WAS IN PLACE. THE DIAMONDBACK 360 PRECISION CORONARY ORBITAL ATHERECTOMY DEVICE (OAD) WAS ADVANCED TO THE LAD. ONE LOW-SPEED TREATMENT AND THREE HIGH-SPEED TREATMENTS WERE PERFORMED. THE PATIENT EXPERIENCED CHEST PAIN. THE OAD WAS REMOVED. ANGIOGRAPHIC IMAGING SHOWED PERFORATION IN LM OSTIAL LAD. TO RESOLVE THE PERFORATION, THE PHYSICIAN TRIED TO PLACE A NON-ABBOTT COVERED STENT IN LAD BUT WAS UNSUCCESSFUL DUE TO THE STENT REPEATEDLY GOING THROUGH THE PERFORATION. THE STENT ENDED UP BEING PLACED IN DISTAL LM INTO THE LEFT CFX. THE PHYSICIAN THOUGHT THE PERFORATION WAS RESOLVED, BUT THERE WAS NO FLOW TO THE LAD. THE PATIENT LOST PULSE. CARDIOPULMONARY RESUSCITATION (CPR) PROTOCOL WAS ADMINISTERED. EPINEPHRINE AND OTHER STANDARD MEDICATIONS WERE GIVEN. PERICARDIOCENTESIS WAS PERFORMED. CPR CONTINUED. PATIENT REGAINED PULSE FOR A WHILE. A TEMPORARY PACEMAKER WAS INSERTED. THE HEART PUMP WAS STILL IN PLACE. CPR WAS CONTINUED. THE PATIENT LOST PULSE AGAIN. AFTER A BRIEF TIME, PATIENT EXPIRATION WAS PRONOUNCED. SOMETIME DURING THE CODE, THE PATIENT ALSO RECEIVED BLOOD TRANSFUSION AND WAS INTUBATED DURING CPR. THE PHYSICIAN INDICATED HE HAS NEVER SEEN AN OAD PERFORATE. THE PRIMARY CAUSE OF DEATH WAS PERFORATION BY OAD. THERE IS NO ALLEGATION OF MALFUNCTION AGAINST THE VIPERWIRE. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. PRODUCTION RECORD AND CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A QUERY OF THE COMPLAINT HANDLING DATABASE FOR THE REPORTED LOT REVEALED THERE IS NO INDICATION OF A LOT-SPECIFIC ISSUE. BASED ON THE INFORMATION RECEIVED, THE INVESTIGATION DETERMINED THAT THE REPORTED PATIENT PERFORATION OF VESSELS, CARDIAC ARREST, ANGINA, VESSEL OCCLUSION, AND RELATED MEDICATION, SURGICAL INTERVENTION AND DEATH APPEARS TO BE RELATED TO OPERATIONAL CONTEXT. IN THIS CASE IT IS LIKELY THAT THE REPORTED PATIENT PERFORATION OF VESSELS, CARDIAC ARREST, ANGINA, VESSEL OCCLUSION, AND RELATED MEDICATION, SURGICAL INTERVENTION AND DEATH ARE A RESULT OF THE PATIENT¿S DISEASE SEVERITY COMBINED WITH USE TECHNIQUES EMPLOYED; HOWEVER, SINCE THE DEVICE WAS NOT RETURNED THIS COULD NOT BE CONFIRMED. BASED ON THE RESULTS OF THE COMPLAINT INVESTIGATION, THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO THE DESIGN, MANUFACTURE, OR LABELING OF THE DEVICE. UPDATED: G3, G6, H2, H6. H6: CODES 4118, 3233, AND 11 NO LONGER APPLY. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21665765&pc=OZD | 1220648-2025-26879 | 2025-02-25 05:00 | Death | ABIOMED, INC. | 2025-03-21 04:00 | OZD | IMPELLA | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Perforation of Vessels | P140003 |  | 1 | Event Description: A LITERATURE ARTICLE FOUND THAT A PATIENT WAS ON IMPELLA SUPPORT AND DURING SUPPORT, WHEN THE PATIENT ATTEMPTED TO WEAN FROM CARDIOPULMONARY BYPASS (CPB) AFTER RESUMING PERCUTANEOUS CARDIOPULMONARY SUPPORT, A LEFT VENTRICULAR FREE WALL RUPTURE WAS OBSERVED. THE PATIENT CONTINUED TO BLEED AFTERWARDS AND UNDERWENT SURGICAL TREATMENT (LEFT VENTRICULAR RECONSTRUCTION UNDER CARDIAC ARREST). THE DETAILS OF THE LEFT VENTRICULAR RUPTURE AND ITS RELATIONSHIP TO IMPELLA WERE UNKNOWN. THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. A.1, A2, A.5 ARE UNKNOWN. D.4 WITHOUT EXACT MODEL NUMBER, THE CATALOG NUMBER, SERIAL NUMBER, LOT NUMBER, EXPIRATION DATE AND UNIQUE IDENTIFIER (UDI) #, D.6A AND D.6B IMPLANT AND EXPLANT DATES, AND H.4 DEVICE MANUFACTURE DATE IS UNKNOWN. E.1 NAME PREFIX/TITLE, FIRST NAME AND LAST NAME ARE UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21666021&pc=NPT | 2025587-2025-02048 | 2025-01-15 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-21 04:00 | NPT | EVOLUT FX VALVE | Patient Device Interaction Problem | Renal Failure; Multiple Organ Failure; Heart Failure/Congestive Heart Failure | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT AN UNSPECIFIED ADVERSE EVENT OCCURRED THREE DAYS AFTER THE IMPLANT OF THE EVOLUT FX TRANSCATHETER AORTIC VALVE. APPROXIMATELY THREE MONTHS LATER, THE PATIENT DIED DUE TO A COMBINATION OF MULTIPLE ORGAN FAILURE, RENAL FAILURE AND HEART FAILURE. ACCORDING TO THE REPORTER, THERE WAS NO CAUSAL RELATIONSHIP BETWEEN THE DEATH AND THE VALVE OR THE VALVE IMPLANT PROCEDURE. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: D4., H4., H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21667394&pc=NIK | 2017865-2025-28083 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-03-21 04:00 | NIK | GALLANT ¿ HF | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Discomfort | P030054 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-28084. RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-28085. IT WAS REPORTED THAT THE PATIENT HAS DECEASED. THE PATIENT EXPERIENCED A SUDDEN SHOCK IN THE CHEST AND LOW BLOOD PRESSURE. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE RELATED. THE CAUSE OF DEATH WAS UNKNOWN. NO ADDITIONAL INFORMATION WAS REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21667482&pc=NIK | 2017865-2025-28084 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-03-21 04:00 | NIK | QUARTET | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Discomfort | P030054 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-28083. RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-28085. IT WAS REPORTED THAT THE PATIENT HAS DECEASED. THE PATIENT EXPERIENCED A SUDDEN SHOCK IN THE CHEST AND LOW BLOOD PRESSURE. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE RELATED. THE CAUSE OF DEATH WAS UNKNOWN. NO ADDITIONAL INFORMATION WAS REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21667486&pc=LWS | 2017865-2025-28085 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-03-21 04:00 | LWS | DURATA STS OPTIM ACTIVE FIXATION | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Discomfort | P950022 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-28083. RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-28084. IT WAS REPORTED THAT THE PATIENT HAS DECEASED. THE PATIENT EXPERIENCED A SUDDEN SHOCK IN THE CHEST AND LOW BLOOD PRESSURE. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE RELATED. THE CAUSE OF DEATH WAS UNKNOWN. NO ADDITIONAL INFORMATION WAS REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21667673&pc=DSQ | 2916596-2025-01814 | 2020-01-01 05:00 | Death | THORATEC CORPORATION | 2025-03-21 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IN THE RESEARCH ARTICLE TITLED, ¿EVALUATION OF DRIVELINE INFECTIONS IN DIABETIC PATIENTS WITH LEFT VENTRICULAR ASSIST DEVICES¿, LEFT VENTRICULAR ASSIST DEVICES (LVADS) ARE CRITICAL IN ADVANCED HEART FAILURE MANAGEMENT, YET DRIVELINE INFECTIONS REMAIN A SIGNIFICANT COMPLICATION. THIS STUDY AIMED TO EVALUATE THE MICROBIOLOGICAL PROFILE AND CLINICAL OUTCOMES OF DRIVELINE INFECTIONS IN DIABETIC AND NONDIABETIC LVAD PATIENTS. A RETROSPECTIVE STUDY BETWEEN JANUARY 2020 AND DECEMBER 2024 IN DR. SIYAMI ERSEK THORACIC AND CARDIOVASCULAR SURGERY TRAINING AND RESEARCH HOSPITAL, IN THE CARDIOVASCULAR SURGERY CLINIC, WITH PATIENTS WHO WERE HOSPITALIZED WITH A DIAGNOSIS OF LVAD DRIVELINE INFECTION OR WHO WERE FOUND TO HAVE DRIVELINE INFECTION DURING HOSPITALIZATION. DEMOGRAPHIC INFORMATION SUCH AS AGE, HEIGHT AND WEIGHT OF THE PATIENTS WERE NOTED, BODY MASS INDEX (BMI) VALUES WERE CALCULATED FOR EACH PATIENT, AND THOSE WITH A BMI OF 30 KG/M2 AND ABOVE WERE INCLUDED IN THE OBESITY GROUP. THE PRESENCE OF DIABETES MELLITUS (DM) AS A COMORBID DISEASE WAS NOTED. DURING THE EVALUATION OF CULTURE RESULTS, PATIENTS WERE DIVIDED INTO "DIABETIC" AND "NON-DIABETIC" GROUPS AND DIFFERENCES WERE EVALUATED. THE ETIOLOGY OF HEART FAILURE WAS DEFINED AS "ISCHEMIC" IN PATIENTS WITH ISCHEMIC CORONARY ARTERY DISEASE, AND THE REMAINING PATIENTS WERE INCLUDED IN THE "NON-ISCHEMIC" GROUP. HEART MATE 2 OR HEART MATE 3 MODEL LVAD DEVICES WERE USED IN THE OPERATION AND APPROPRIATE POSTOPERATIVE WOUND CARE WAS PERFORMED IN EACH PATIENT. THE DIAGNOSIS OF DRIVELINE INFECTION WAS MADE ACCORDING TO THE CRITERIA OF THE INTERNATIONAL HEART AND LUNG TRANSPLANT CONSENSUS BY EVALUATING PURULENT DISCHARGE, PAIN, TENDERNESS, FEVER, LABORATORY DATA AND CULTURE RESULTS. ALL PATIENTS WERE TREATED WITH APPROPRIATE WOUND CARE AND INTRAVENOUS ANTIBIOTIC THERAPY FROM THE TIME OF DIAGNOSIS. GRAM POSITIVE BACTERIA WERE THE MOST COMMON MICROORGANISM GROUP IN BOTH GROUPS WITH 53.2% (N=33) IN THE DIABETIC GROUP AND 63.6% (N=28) IN THE NON-DIABETIC GROUP (P=0.285). STAPHYLOCOCCUS SPECIES, ESPECIALLY NON-AUREUS SPECIES, WERE FREQUENTLY ISOLATED IN BOTH GROUPS AND THE DIFFERENCE BETWEEN THE GROUPS WAS NOT SIGNIFICANT (DIABETIC 24.2%, NON-DIABETIC 29.5%; P=0.538). S. AUREUS WAS ISOLATED 12.9% IN THE DIABETIC GROUP AND 25.0% IN THE NON-DIABETIC GROUP, BUT THE DIFFERENCE WAS NOT STATISTICALLY SIGNIFICANT (P=0.110). METHICILLIN-RESISTANT S. AUREUS (MRSA) WAS RARE IN BOTH GROUPS AND WAS ISOLATED IN ONLY THREE CASES (2.8%) (P=1.000). GRAM NEGATIVE BACTERIA WERE OBSERVED IN 38.7% (N=24) OF DIABETIC PATIENTS AND 31.8% (N=14) OF NON-DIABETIC PATIENTS (P=0.466). PSEUDOMONAS, KLEBSIELLA AND ACHROMOBACTER SPECIES WERE AMONG THE MOST FREQUENTLY ISOLATED GRAM-NEGATIVE BACTERIA. HOWEVER, NO STATISTICALLY SIGNIFICANT DIFFERENCE WAS FOUND BETWEEN THE GROUPS IN TERMS OF THE DISTRIBUTION OF THESE SPECIES (P>0.05 FOR ALL). THE RATE OF FUNGAL GROWTH WAS 8.1% IN DIABETICS AND 4.5% IN NON-DIABETICS. ALL FUNGAL MICROORGANISMS ISOLATED WERE CANDIDA SPECIES AND THERE WAS NO SIGNIFICANT DIFFERENCE BETWEEN THE GROUPS (P=0.697). IN CONCLUSION, NO SIGNIFICANT DIFFERENCE WAS FOUND BETWEEN DIABETIC AND NON-DIABETIC GROUPS IN TERMS OF MICROBIAL GROWTH AND ISOLATED MICROORGANISM SPECIES. IN ADDITION, NO STATISTICAL SIGNIFICANCE WAS OBSERVED IN RECURRENCE RATES, BACTEREMIA OR MORTALITY RATES BETWEEN DIABETIC AND NON-DIABETIC GROUPS. CONCLUSION: THIS STUDY COMPARED THE MICROBIOLOGICAL PROFILES OF DRIVELINE INFECTIONS AND THEIR EFFECTS ON MORTALITY IN DIABETIC AND NON-DIABETIC PATIENTS USING LVADS. GRAM-POSITIVE BACTERIA WERE THE MOST FREQUENTLY ISOLATED INFECTIOUS AGENTS IN BOTH GROUPS. THERE WAS NO SIGNIFICANT DIFFERENCE IN THE DISTRIBUTION OF INFECTIOUS AGENTS, RECURRENCE RATES AND MORTALITY BETWEEN DIABETIC AND NON-DIABETIC PATIENTS. HOWEVER, AGE WAS FOUND TO BE AN INDEPENDENT RISK FACTOR FOR MORTALITY. THE STUDY FINDINGS SUGGEST THAT THERE IS NO NEED FOR DIFFERENT STRATEGIES FOR DIABETIC AND NON-DIABETIC PATIENTS IN THE MANAGEMENT AND PROPHYLAXIS OF INFECTIONS. HOWEVER, PROSPECTIVE STUDIES WITH LARGER SAMPLE GROUPS ARE NECESSARY TO CONFIRM THESE RESULTS. FIVE PERCENT OF PATIENTS WITH DRIVELINE INFECTIONS RECEIVED TRANSPLANTS. Manufacturer Narrative: SECTION B2: DATE OF DEATH ESTIMATED TO DATE OF EVENT. SECTION E1: CUSTOMER SITE CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 DEVICES AND THE REPORTED OUTCOMES COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. A TOTAL OF 40 PATIENTS WERE INCLUDED IN THIS STUDY (20 HEARTMATE II AND 20 HEARTMATE 3 PATIENTS). OF THE 40 PATIENTS, 24 PATIENTS EXPIRED. THE HEARTMATE 3 DEVICE SERIAL NUMBERS, AS WELL AS OTHER SPECIFIC CASE/PATIENT INFORMATION, ARE NOT AVAILABLE. THE SERIAL NUMBER OR OTHER IDENTIFYING INFORMATION OF THE PRODUCT WAS NOT REPORTED AND WAS NOT ABLE TO BE DETERMINED DURING THE INVESTIGATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SECTION A AND D - SPECIFIC PATIENT AND DEVICE INFORMATION WAS UNKNOWN. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. DEPARTMENT OF CARDIOVASCULAR SURGERY, DR. SIYAMI ERSEK THORACIC AND CARDIOVASCULAR SURGERY TRAINING AND RESEARCH HOSPITAL, ISTANBUL, TÜRKIYE. JOURNAL OF THE CARDIOVASCULAR THORACIC ANAESTHESIA AND INTENSIVE CARE SOCIETY. GKDA DERG 2024;30(4):145-150 DOI: 10.14744/GKDAD.2024.72324.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21667683&pc=OZD | 1220648-2025-26883 | 2025-02-25 05:00 | Death | ABIOMED, INC. | 2025-03-21 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Hemolysis | P140003 |  | 1 | Event Description: IT WAS REPORTED VIA ABIOMED¿S LONG-TERM OUTCOME AND QUALITY INDICATOR (LOQI) IMPELLA REGISTRY THAT A 57 YEAR OLD MALE PATIENT ENDURED HEMOLYSIS DURING IMPELLA 5.5 SUPPORT. THE CONDITION WAS BELIEVED TO HAVE A POSSIBLE RELATIONSHIP TO THE IMPELLA DEVICE. HEMOGLOBIN LEVELS WERE DROPPING AND THREE UNITS OF PACKED RED BLOOD CELLS WERE GIVEN TO TREAT THE CONDITION. AFTER THIRTY SIX DAYS OF SUPPORT, THE PATIENT'S FAMILY MADE THE DECISION TO WITHDRAW CARE AND THE PATIENT PASSED AWAY. THERE IS NOT ENOUGH EVIDENCE TO DISASSOCIATE THE USE OF THE IMPELLA WITH THE PATIENT'S PASSING. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMARTASSIST SYSTEM STATES THE FOLLOWING: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS ¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES ¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21667987&pc=QYX | 3007208829-2025-00070 | 2025-01-31 05:00 | Death | IRHYTHM TECHNOLOGIES, INC | 2025-03-21 04:00 | QYX | ZIO AT | Insufficient Information; Inaccurate Information | Unspecified Heart Problem | K240029 |  | 1 | Event Description: THE PATIENT EXPIRED DURING THEIR PRESCRIBED ZIO AT WEAR PERIOD. IRHYTHM ATTEMPTED TO GATHER MORE INFORMATION ABOUT THE CAUSE OF DEATH FROM THE ACCOUNT, BUT NO FURTHER DETAILS WERE PROVIDED.Event Description: THE PATIENT EXPERIENCED AN ARRHYTHMIA THAT MET MEDICAL DOCTOR NOTIFICATION (MDN) REQUIREMENTS THAT WAS NOT COMMUNICATED DURING THE WEAR PERIOD. THE INVESTIGATION REVEALED THAT A PRELIMINARY ECG INTERPRETATION PROVIDED TO THE PHYSICIAN WAS MISCLASSIFIED. FOLLOWING THE WEAR PERIOD AND WHILE COMPILING THE FINAL REPORT, THE INTERPRETATION WAS AMENDED. THE HEALTHCARE PROVIDER (HCP) WAS NOTIFIED, AND IRHYTHM LEARNED THAT THE PATIENT HAD EXPIRED. THE PATIENT¿S CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE INVESTIGATION IS STILL ONGOING, AND A FOLLOW-UP REPORT WILL BE SUBMITTED UPON ITS COMPLETION. Manufacturer Narrative: THIS SUPPLEMENTAL REPORT IS BEING SUBMITTED TO CORRECT THE UDI FORMAT AND TO PROVIDE ADDITIONAL EVENT DETAILS AS WELL AS THE FINDINGS FROM THE COMPLAINT INVESTIGATION. A REVIEW OF THIS COMPLAINT REVEALED THAT THE PATIENT WAS PRESCRIBED A ZIO AT DEVICE FOR A 14-DAY PRESCRIBED WEAR PERIOD. THE ZIO AT DEVICE WAS RETURNED TO IRHYTHM ON DAY 14, AND THE CLINICAL DATA WAS DOWNLOADED. ON DAY 2 OF WEAR, THE ZIO AT DEVICE RECORDED AN ARRHYTHMIA (VT) THAT SHOULD HAVE PROMPTED NOTIFICATION TO THE PROVIDER. THE ZIO AT RECORDED AND TRANSMITTED THE ARRHYTHMIA; HOWEVER, THE QUALITY CARE TECHNICIAN (QCT) CLASSIFIED IT AS SVT, WHICH DID NOT MEET NOTIFICATION CRITERIA. THE ZIO AT DEVICE RECORDED ANOTHER EPISODE OF VT 5.5 HOURS LATER, WHICH DID NOT TERMINATE AND DEGENERATED INTO VENTRICULAR FIBRILLATION (VF) AND THEN ASYSTOLE. THE ZIO AT RECORDED AND TRANSMITTED THE ARRYTHMIAS; HOWEVER, THE REPORTING TIMELINESS FOR ZIO AT IS NOT CONSISTENT WITH LIFE-THREATENING ARRYTHMIAS. THE PATIENT SUBSEQUENTLY PASSED AWAY. SINCE IT CANNOT BE EXCLUDED THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE PATIENT'S DEATH. THIS REPORT IS BEING SUBMITTED IN AN ABUNDANCE OF CAUTION. THE ZIO AT CLINICAL REFERENCE MANUAL STATES IN THE ¿INDICATIONS FOR USE¿ SECTION THAT: ¿IT IS NOT INTENDED FOR USE ON CRITICAL CARE PATIENTS.¿ THERE IS ALSO A WARNING THAT: ¿THIS DEVICE IS NOT INTENDED FOR USE IN CRITICAL CARE PATIENTS BECAUSE THE REPORTING TIMELINESS IS NOT CONSISTENT WITH LIFE-THREATENING ARRHYTHMIAS SUCH AS VENTRICULAR FIBRILLATION.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21668152&pc=NPW | 2015691-2025-02170 | 2025-02-21 05:00 | Death | EDWARDS LIFESCIENCES | 2025-03-21 04:00 | NPW | EVOQUE TRICUSPID VALVE REPLACEMENT SYSTEM | Malposition of Device | Hemolysis; Cardiovascular Insufficiency | P230013 |  | 1 | Event Description: AS REPORTED BY THE EDWARDS LIFESCIENCES AFFILIATE IN GERMANY WHERE, AFTER DEPLOYMENT, THE VALVE REMAINED IN POSITION; HOWEVER, THE SEPTAL PART OF THE VALVE STARTED TO ROCK. ONE ANTERIO-SEPTAL PVL GRADE WAS OBSERVED TO BE MODERATE TO SEVERE WITH HIGH VELOCITY. THE PATIENT GOT HEMOLYSIS AND WAS IN THE ICU. IT WAS NOT POSSIBLE TO BE TREATED SURGICALLY. THE PATIENT PASSED AWAY DUE TO RV FAILURE. THE VALVE WAS FURTHER MIGRATED WITH TWO DIFFERENT JETS OF PVL AND SOME ROCKING MOTION. THE PATIENT NEVER LEFT THE ICU. AT THE START OF PROCEDURE, DURING WIRE PLACEMENT, THE PATIENT'S BLOOD PRESSURE DROPPED. THE HEART TEAM RAISED THE PRESSURE WHICH HAD A CONSEQUENCE THAT PULMONIC PRESSURES (RESISTANCE) WENT UP AND LEFT VENTRICLE (LV) FUNCTION WENT DOWN/STARTED TO FAIL. A SPECIALIZED ANESTHETIC TEAM JOINED THE PROCEDURE AND NEEDED ONE HOUR TO STABILIZE THE PATIENT BY ADMINISTERING FLUIDS TO INCREASE THE VOLUME. AFTER THE PATIENT WAS STABLE, THE PHYSICIAN DECIDED TO CONTINUE/START THE PROCEDURE. THE PROCEDURE WAS COMPLETED WITHOUT ANY NEGATIVE EVENTS. SUCCESSFUL LEAFLET ENGAGEMENT WAS ACHIEVED FOR ALL ANCHORS, UNTIL THE START OF EXPANSION, WITH BLUE KNOB, TO THE FINAL VENTRICULAR EXPANSION. THE BLUE KNOB WAS ACTIVATED UNTIL 1/3 OF THE INNER CAPSULE WAS RETRACTED. AFTER SOME SECONDS REACHING THIS EXPANSION STAGE, IT WAS OBSERVED A JUMP IN POSITION IN TEE (MPR) WHERE THE SEPTAL AND POSTERIOR PART OF THE VALVE WAS HIGHER THAN BEFORE (MOVED UP SUDDENLY). CHECKING THE FLUORO AGAIN INDICATED THAT INNER CAPSULE HAD RETRACTED ALREADY MORE THAN 2/3. FURTHERMORE, THE TEE IMAGING GOT WORSE (A LOT OF SHADOWING). MPR AND 2D MID-, DEEP- AND TRANS-GASTRIC VIEWS WERE USED TO CONFIRM LEAFLET ENGAGEMENT (ALL ANCHORS STILL GOOD) AND ANCHOR TO HINGE POINT DISTANCE. SEPTAL AND POSTERIOR ANCHORS WERE LOWERED. AFTER THAT THE PROCEDURE CONTINUED FOLLOWING PROCEDURE STEPS TO BRING ANCHORS AS HIGH AS POSSIBLE PARALLEL TO THE ANNULUS. AFTER DEPLOYMENT, THE VALVE REMAINED IN POSITION; HOWEVER, THE SEPTAL PART OF THE VALVE STARTED TO ROCK. FIRST WITH A HIGHER AMPLITUDE, SYNCHRONIZED WITH MANY EXTRA SYSTOLES AND HEART RATE OF 115 BPM AND THEN WITH A LOWER AMPLITUDE. LEAFLET ENGAGEMENT AT THE SEPTAL ANCHORS WAS DOUBLE-CHECKED AND CONFIRMED. THE FINAL DISTANCE FROM VALVE TO ANNULUS WAS BETWEEN 0-5MM. ONE ANTERIO-SEPTAL PVL GRADE WAS OBSERVED TO BE MODERATE TO SEVERE WITH HIGH VELOCITY AND NO TRANSVALVULAR BACKFLOW. THE PATIENT GOT HEMOLYSIS AND WAS IN THE ICU. IT WAS NOT POSSIBLE TO BE TREATED SURGICALLY. THE PATIENT PASSED AWAY DUE TO RV FAILURE. THE VALVE WAS FURTHER MIGRATED WITH TWO DIFFERENT JETS OF PVL AND SOME ROCKING MOTION. THE PATIENT NEVER LEFT THE ICU. Manufacturer Narrative: THE FOLLOWING SECTIONS WERE UPDATED/CORRECTED/ADDED: B4, G3, G6, H2, H6 AND H11. ADDITIONAL TYPE OF INVESTIGATION CODES THAT WERE UNABLE TO BE ADDED IN H6: LABELLING REVIEW: ANALYSIS OF IMAGES. INVESTIGATION CONCLUSION CODE THAT WAS UNABLE TO BE ADDED IN H6: ADVERSE EVENT RELATED TO PATIENT ANATOMY AND/OR CONDITION. THE COMPLAINT FOR VALVE DEPLOYED TOO VENTRICULAR WAS CONFIRMED WITH OBJECTIVE EVIDENCE. THE IMAGING REVIEW REVEALED THAT THE EVOQUE VALVE APPEARS TO BE DEPLOYED LOWER AT THE ANTEROLATERAL ASPECT OF THE VENTRICULAR. IN ADDITION, IT WAS NOTED THAT THERE WAS EVIDENCE OF VALVE INSTABILITY, WITH ROCKING MOTION OF THE POSTEROSEPTAL/SEPTAL ASPECTS. THE DEVICE POSITIONING CONTRIBUTED TO PVL OBSERVED. AT THE START OF THE PROCEDURE (WIRE PLACEMENT), PRIOR TO DELIVERY SYSTEM INSERTION, IT WAS STATED THAT THE PATIENT EXPERIENCED HEART FAILURE. A SPECIALIZED ANESTHETISTS THAT JOINED THE TEAM TO STABILIZE THE PATIENT ADMINISTERED FLUIDS TO INCREASE THE VOLUME. ALTHOUGH THE VOLUME INCREASE SHOWED MINIMAL CHANGE TO THE ANNULAR MEASUREMENTS, THE PATIENT'S CONDITION MAY ALSO BE RELATED TO THE INSTABILITY OF THE VALVE. NO MANUFACTURING NON-CONFORMITIES WERE IDENTIFIED FROM THE IMAGING EVALUATION. AVAILABLE INFORMATION SUGGESTS THAT PATIENT CONDITIONS (HEART FAILURE) AND PROCEDURE RELATED ISSUES MAY HAVE CONTRIBUTED TO THE REPORTED EVENT. SINCE NO MANUFACTURING NON-CONFORMANCES WERE FOUND AND NO LABELING, TRAINING, OR IFU DEFICIENCIES WERE IDENTIFIED, NO PREVENTATIVE OR CORRECTIVE ACTIONS WERE REQUIRED. THE DHR REVIEW WAS PERFORMED, AND BASED ON THIS REVIEW, THERE WERE NO NON-CONFORMANCES RELATED TO THE ISSUE OBSERVED AND THE DEVICE PASSED ALL MANUFACTURING SPECIFICATIONS. Manufacturer Narrative: THE FOLLOWING SECTIONS WERE UPDATED/CORRECTED/ADDED: B4, D9, G3, G6, H2, H6 AND H11. ADDITIONAL H6 TYPE OF INVESTIGATION CODES: 'ANALYSIS OF IMAGES' AND 'LABELLING REVIEW' ADDITIONAL H6 INVESTIGATION CONCLUSION CODE: 'ADVERSE EVENT RELATED TO PATIENT ANATOMY AND/OR CONDITION'. THE COMPLAINT FOR VALVE DEPLOYED TOO VENTRICULAR WAS CONFIRMED WITH OBJECTIVE EVIDENCE. NO MANUFACTURING NON-CONFORMITIES WERE IDENTIFIED FROM THE IMAGING EVALUATION. BASED ON THE DHR REVIEW, THERE WERE NO NON-CONFORMANCES RELATED TO THE ISSUE OBSERVED AND THE DEVICE PASSED ALL MANUFACTURING SPECIFICATIONS. SINCE NO EDWARDS DEFECT WAS IDENTIFIED, CORRECTIVE OR PREVENTATIVE ACTIONS ARE NOT REQUIRED. AVAILABLE INFORMATION SUGGESTS THAT PATIENT CONDITIONS (HEART FAILURE) AND PROCEDURE RELATED ISSUES MAY HAVE CONTRIBUTED TO THE REPORTED EVENT. Manufacturer Narrative: THE MANUFACTURER'S INVESTIGATION IS ONGOING AND A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21668408&pc=PAM | 3002773840-2025-00017 | 2025-02-17 05:00 | Death | BIOFIRE DIAGNOSTICS, LLC | 2025-03-21 04:00 | PAM | BIOFIRE® BLOOD CULTURE IDENTIFICATION 2 (BCID2) PANEL | False Negative Result | Insufficient Information | K193519 |  | 1 | Event Description: SUMMARY: (B)(6) REPORTED A POTENTIAL FALSE NEGATIVE METHICILLIN-RESISTANT STAPHYLOCOCCUS AUREUS (MRSA) RESULT ON THE BIOFIRE BLOOD CULTURE IDENTIFICATION 2 (BCID2) PANEL AFTER TESTING A PATIENT'S BLOOD CULTURE SAMPLE. THE CUSTOMER REPORTED THAT DUE TO THE BIOFIRE BCID2 PANEL RESULT, THE PATIENT'S TREATMENT WAS IMPACTED, RESULTING IN THEIR DEATH. BIOFIRE HAS REQUESTED FURTHER INFORMATION FROM THE CUSTOMER AND IS CURRENTLY INVESTIGATING THIS EVENT. NO REMEDIAL ACTION, CORRECTIVE ACTION, PREVENTIVE ACTION, OR FIELD SAFETY CORRECTIVE ACTION (FSCA) HAS BEEN DEEMED NECESSARY AT THIS TIME.Event Description: SUMMARY: (B)(6) REPORTED A POTENTIAL NEGATIVE METHICILLIN-RESISTANT STAPHYLOCOCCUS AUREUS (MRSA) RESULT ON THE BIOFIRE BLOOD CULTURE IDENTIFICATION 2 (BCID2) PANEL AFTER TESTING A PATIENT'S BLOOD CULTURE SAMPLE. THE CUSTOMER REPORTED THAT DUE TO THE BIOFIRE BCID2 PANEL RESULT, THE PATIENT'S TREATMENT WAS IMPACTED, CONTRIBUTING TO THEIR DEATH. BIOFIRE IS CURRENTLY INVESTIGATING THIS EVENT. NO REMEDIAL ACTION, CORRECTIVE ACTION, PREVENTIVE ACTION, OR FIELD SAFETY CORRECTIVE ACTION (FSCA) HAS BEEN DEEMED NECESSARY AT THIS TIME.Event Description: SUMMARY: (B)(6) HEALTHCARE REPORTED A POTENTIAL NEGATIVE METHICILLIN-RESISTANT STAPHYLOCOCCUS AUREUS (MRSA) RESULT ON THE BIOFIRE BLOOD CULTURE IDENTIFICATION 2 (BCID2) PANEL AFTER TESTING A PATIENT'S BLOOD CULTURE SAMPLE. THE CUSTOMER REPORTED THAT DUE TO THE BIOFIRE BCID2 PANEL RESULT, THE PATIENT'S TREATMENT WAS IMPACTED, CONTRIBUTING TO THEIR DEATH. UPON INVESTIGATION, NO MALFUNCTION OCCURRED, AND THE BIOFIRE BCID2 PANEL WAS WORKING WITHIN SPECIFICATION. THE INVESTIGATION CONCLUDED THAT THE MOST LIKELY CAUSE FOR THE DISCREPANT MRSA WAS DUE TO AN ORGANISM STRAIN THAT HAS REDUCED REACTIVITY WITH THE BIOFIRE BCID2 PANEL'S MREJ ASSAY. Manufacturer Narrative: INVESTIGATION: THE PATIENT WAS A 9-YEAR-OLD MALE. THE PATIENT PRESENTED WITH SIGNS AND SYMPTOMS OF TACHYPNEA AND TACHYCARDIA. THE CUSTOMER REPORTED THERE WAS ALSO CONCERN FOR SEPSIS WITH ELEVATED C-REACTIVE PROTEIN (CRP) WHEN BLOOD CULTURES WERE DRAWN. THE PATIENT WAS ALREADY PAN-CYTOPENIC AND NEUTROPENIC. WHEN ASKED WHAT THE UNDERLYING CAUSE OF THE PATIENT'S IMMUNOCOMPROMISED STATUS (PAN-CYTOPENIC AND NEUTROPENIC) WAS, THE CUSTOMER COMMUNICATED THE FOLLOWING: "MULTIFACTORIAL. THE PATIENT WAS PREVIOUSLY HEATHY BUT PRESENTED WITH [FEBRILE INFECTION-RELATED EPILEPSY SYNDROME] FIRES AND ALSO DEVELOPED PANCYTOPENIA (OF UNKNOWN ETIOLOGY DESPITE HEME/ONC INVOLVEMENT) AND [HEMOPHAGOCYTIC LYMPHOHISTIOCYTOSIS] HLH DURING THE PATIENT'S PROLONGED COURSE OF ILLNESS. THE PATIENT RECEIVED CYTOKINE DIRECTED THERAPY WITH BOTH IL-6 BLOCKER AND ANAKINRA AS WELL AS HIGH-DOSE STEROIDS FOR FIRES. ALSO RECEIVED IMMUNOSUPPRESSIVE THERAPY FOR HLH." WHEN ASKED IF THE PATIENT WAS DIAGNOSED WITH ANY ORGAN FAILURE AT THE TIME OF TESTING, SUCH AS SEPTIC SHOCK, THE NEED FOR MECHANICAL VENTILATION, OR ANY OTHER FORM OF ORGAN SUPPORT, THE CUSTOMER REPORTED THAT THE PATIENT WAS ALREADY VENTILATED DUE TO NEUROLOGICAL ISSUE AND RENAL FAILURES (UNKNOWN TIME POINT FOR VENTILATION AND RENAL FAILURE) AND THE PATIENT DEVELOPED HEPATIC FAILURE DURING ANTIBIOTIC TREATMENT REQUIRING CONTINUOUS RENAL REPLACEMENT THERAPY (CRRT). ON (B)(6) 2025, THE PATIENT'S BLOOD CULTURE SAMPLE WAS TESTED ON THE BIOFIRE BCID2 PANEL. THE BIOFIRE BCID2 PANEL REPORTED STAPHYLOCOCCUS SPP. AND S. AUREUS AS DETECTED. THE BIOFIRE BCID2 PANEL WAS POSITIVE FOR MECA/C BUT NEGATIVE FOR MREJ, RESULTING IN A NOT DETECTED RESULT FOR 'MECA/C AND MREJ (MRSA).' NOTE THAT THE BIOFIRE BCID2 PANEL WILL ONLY REPORT A DETECTED RESULT FOR 'MECA AND MREJ (MRSA)' IF S. AUREUS, MREJ, AND MECA/C ASSAYS ARE ALL POSITIVE. MICROSCAN SUSCEPTIBILITY TESTING SHOWED OXACILLIN (MIC= >2 UG/ML) RESISTANCE AND WAS CEFOXITIN SCREEN POSITIVE (MIC= >4 UG/ML). THE ABBOTT CLEARVIEW PENICILLIN-BINDING PROTEIN 2A (PBP2A) CULTURE COLONY TEST WAS POSITIVE FOR PBP2A. GRAM STAIN SHOWED GRAM-POSITIVE COCCI IN CLUSTERS. THE CUSTOMER REPORTED ANOTHER BLOOD CULTURE SAMPLE WAS TESTED ON THE BIOFIRE BCID2 PANEL AND WAS POSITIVE FOR MSSA. THE RUN FILE INFORMATION FOR THIS TEST WAS NOT PROVIDED. THE CUSTOMER REPORTED DUE TO THE BIOFIRE BCID2 PANEL RESULT, THE PATIENT'S ANTIBIOTIC TREATMENT WAS CHANGED TO MSSA DIRECTED THERAPY. THIS INAPPROPRIATE CHANGE IN TREATMENT OCCURRED FOR TWO DAYS. THE CUSTOMER STATED THAT THE FALSE NEGATIVE MRSA RESULT ON THE BIOFIRE BCID2 PANEL "DIRECTLY IMPACTED THE PATIENT'S TREATMENT, RESULTING IN THEIR DEATH." THE CUSTOMER COMMUNICATED THE FOLLOWING REGARDING THE PATIENT'S TREATMENT CHANGE: "WHEN THE PATIENT DETERIORATED, INITIALLY, THE PATIENT WAS ON VANCOMYCIN. HOWEVER, ONCE WE RECEIVED THE RESULT OF [BIOFIRE BCID2 PANEL] INDICATING MSSA, WE SWITCHED TO OXACILLIN FOR DIRECTED THERAPY (AND ADDED CLINDAMYCIN FOR TOXIN MEDIATED DISEASE AS HE ALSO DEVELOPED ERYTHRODERMA) GIVEN THE PATIENT ALSO DEVELOPED AKI (WANTED TO AVOID VANC) AS WELL AS RESPIRATORY FAILURE (DAPTO WAS NOT OPTIMAL). GIVEN PANCYTOPENIA, LINEZOLID WAS NOT HIGH ON CUSTOMER OPTIONS INITIALLY. ONCE CULTURE SHOWED MRSA, CUSTOMER SWITCHED OXACILLIN TO CEFTAROLINE FOR MRSA." THE ANALYSIS OF THE INITIAL BIOFIRE BCID2 PANEL RUN FILE PROVIDED BY THE CUSTOMER SHOWED ROBUST AMPLIFICATION SIGNATURES FOR BOTH MECA/C AND S. AUREUS ASSAYS, WHILE NO MELT/AMPLIFICATION SIGNATURE WAS SEEN FOR THE MREJ ASSAY. QUALITY CONTROL (QC) RECORDS FOR POUCH LOT# 3H7D24 (KIT LOT# 1977124) WERE ALSO REVIEWED. THIS POUCH LOT PASSED QC CRITERIA AND WAS FOUND WITHIN SPECIFICATIONS. NO RUN MALFUNCTION OCCURRED AT THE CUSTOMER SITE AND THE FILMARRAY INSTRUMENT (SERIAL NUMBER# (B)(6)) WAS WORKING WITHIN DESIGNED SPECIFICATIONS. IN-HOUSE INVESTIGATION: BIOFIRE RECEIVED ONE TUBE OF ISOLATE AND TWO TUBES OF BLOOD CULTURE ALIQUOTS FOR TESTING. THE TWO BLOOD CULTURE ALIQUOTS WERE TESTED ON THE BIOFIRE BCID2 PANEL. THE BIOFIRE BCID2 PANEL DETECTED S. AUREUS, MECA/C AND MREJ (MRSA) FOR BOTH ALIQUOTS. THE MECA/C AND S. AUREUS ASSAYS EACH HAD ROBUST SIGNATURES, WHILE A LATE SIGNATURE WAS OBSERVED FOR THE MREJ ASSAY. THE ISOLATE WAS TESTED TWICE ON THE BIOFIRE BCID2 PANEL AND S. AUREUS, MECA/C AND MREJ (MRSA) WAS DETECTED BOTH TIMES. AGAIN, THE MECA/C AND S. AUREUS ASSAYS EACH HAD ROBUST SIGNATURES, WHILE A LATE SIGNATURE WAS OBSERVED FOR THE MREJ ASSAY. GROWTH OF ONE MORPHOLOGY TYPE WAS SEEN FROM CULTURE. ISOLATES WERE TESTED USING VITEK® MS AND VITEK® 2, WITH BOTH PLATFORMS IDENTIFYING THE ORGANISM AS S. AUREUS, VITEK® MS REPORTING A HIGH CONFIDENCE. VITEK® 2 AST CONFIRMED MRSA WITH A CEFOXITIN POSITIVE CALL AND OXACILLIN RESISTANCE. BASED ON THE BIOFIRE BCID2 PANEL'S AMPLIFICATION SIGNATURE PATTERN OF S. AUREUS, MECA/C, AND MREJ ASSAYS, INITIAL FINDINGS INDICATED IT WAS LIKELY THAT THE PATIENT SAMPLE CONTAINED A MRSA STRAIN CARRYING AN MREJ VARIANT WITH REDUCED REACTIVITY ON THE BIOFIRE BCID2 PANEL. WHOLE GENOME SEQUENCING WAS PERFORMED TO PROVIDE ADDITIONAL INFORMATION ON THE SPECIFIC MREJ VARIANT. THE MRSA ISOLATE WAS GROWN TO A HIGH TITER IN LIQUID CULTURE TO PREPARE A GLYCEROL STOCK FOR WHOLE GENOME SEQUENCING. ANALYSIS OF TARGET SEQUENCES OBSERVED MULTIPLE MISMATCHES UNDER THE MREJ ASSAY'S PRIMER BINDING SITES. THE ANALYSIS FURTHER INDICATED THE ISOLATE RECOVERED FROM THE PATIENT'S BLOOD CULTURE IS LIKELY AN MREJ VARIANT THAT IS EXPECTED TO EITHER HAVE REDUCED OR NO REACTIVITY (NOT DETECTED) WITH THE BIOFIRE BCID2 PANEL'S MREJ ASSAY. HOWEVER, THE BLOOD CULTURE ALIQUOT AND ISOLATE PROVIDED BY THE CUSTOMER BOTH DETECTED MREJ BY THE BIOFIRE BCID2 PANEL DURING INTERNAL TESTING, SUGGESTING THE MREJ VARIANT CAN BE REPORTED AS DETECTED WHEN THE TARGET NUCLEIC ACID IS PRESENT AT A HIGH CONCENTRATION. ALTERNATIVELY, THE MREJ VARIANT MAY BE REPORTED AS NOT DETECTED WHEN THE TARGET IS NOT PRESENT IN THE BLOOD CULTURE AT HIGH ENOUGH CONCENTRATION, WHICH IS LIKELY THE REASON IT WAS NOT DETECTED AT THE CUSTOMER'S SITE. CONCLUSION: THE INVESTIGATION CONCLUDED THAT THE MOST LIKELY CAUSE FOR THE DISCREPANT MRSA RESULT WAS DUE TO A MREJ VARIANT THAT HAS REDUCED REACTIVITY WITH THE BIOFIRE BCID2 PANEL'S MREJ ASSAY. THE BIOFIRE BCID2 PANEL IS A MULTIPLEXED NUCLEIC ACID TEST INTENDED FOR THE SIMULTANEOUS QUALITATIVE DETECTION AND IDENTIFICATION OF MULTIPLE BACTERIAL AND YEAST NUCLEIC ACIDS AND SELECT GENETIC DETERMINANTS ASSOCIATED WITH ANTIMICROBIAL RESISTANCE. EACH AMR GENE RESULT IS ASSOCIATED WITH A SINGLE CORRESPONDING ASSAY EXCEPT FOR THE MECA/C AND MREJ RESULT, WHICH IS DEPENDENT ON BOTH THE MECA/C ASSAY AND THE MREJ ASSAY. IN ORDER FOR MECA/C AND MREJ (MRSA) TO BE REPORTED AS DETECTED, THE FILMARRAY MUST BE POSITIVE FOR ALL THREE ASSAYS: S. AUREUS, MECA/C, AND MREJ. MORE INFORMATION ON MECA/C AND MREJ ASSAY INTERPRETATION CAN BE FOUND IN TABLE 10 OF THE BIOFIRE BCID2 PANEL INSTRUCTION FOR USE (IFU) (WWW.ONLINE-IFU.COM/ITI0048). THE ANALYSIS OF THE INITIAL BIOFIRE BCID2 PANEL RUN FILE PROVIDED BY THE CUSTOMER SHOWED ROBUST AMPLIFICATION SIGNATURES FOR BOTH MECA/C AND S. AUREUS ASSAYS, WHILE NO MELT/AMPLIFICATION SIGNATURE WAS SEEN FOR THE MREJ ASSAY. HOWEVER, THE CUSTOMER'S SUBCULTURE AND SUSCEPTIBILITY TESTING REVEALED MRSA AND GRAM STAIN SHOWED GRAM-POSITIVE COCCI IN CLUSTERS. TO INVESTIGATE FURTHER, IN-HOUSE TESTING AT BIOFIRE WAS PERFORMED, AND THE CUSTOMER'S RESULT OF A NEGATIVE MREJ/NEGATIVE MRSA COULD NOT BE REPRODUCED. ALL BIOFIRE BCID2 PANEL TESTS RAN INTERNALLY WERE POSITIVE RESULTS FOR MECA/C, S. AUREUS, AND MREJ (MRSA). HOWEVER, IT IS NOTEWORTHY THAT THE MREJ SIGNATURES OBSERVED WERE LATER AND LESS ROBUST THAN THOSE SEEN FOR MECA/C AND S. AUREUS. ADDITIONALLY, ANALYSIS OF THE WHOLE GENOME SEQUENCE IDENTIFIED A MREJ VARIANT THAT WOULD BE EXPECTED HAVE REDUCED REACTIVITY WITH THE BIOFIRE BCID2 PANEL'S MREJ ASSAY. ACCORDING TO THE "METHICILLIN RESISTANCE GENES DETECTION BY THE BIOFIRE BLOOD CULTURE IDENTIFICATION 2 (BCID2) PANEL" TECHNICAL NOTE [BFR0002-4828], METHICILLIN RESISTANCE IN STAPHYLOCOCCUS SPECIES IS PRIMARILY CONFERRED WHEN THE MOBILE SCCMEC GENETIC CASSETTE, CARRYING EITHER OF THE RESISTANCE GENES MECA OR MECC, INTEGRATES INTO A SPECIFIC REGION OF THE STAPHYLOCOCCUS GENOME. IN S. AUREUS, THIS STAPHYLOCOCCAL CASSETTE CHROMOSOME MEC (SCCMEC) INSERTION CREATES A JUNCTION BETWEEN THE S. AUREUS GENOME AND THE SCCMEC CASSETTE THAT IS CALLED MREJ (SCCMEC RIGHT-EXTREMITY JUNCTION). THE BIOFIRE BCID2 PANEL'S MREJ ASSAY IS DESIGNED TO DETECT THIS SPECIFIC INTEGRATION EVENT IN THE S. AUREUS GENOME. THERE ARE A VARIETY OF DIFFERENT MREJ TYPES (I-XXI) CLASSIFIED BY THE POINT OF INSERTION OF SCCMEC INTO S. AUREUS. THE MREJ ASSAY IS PREDICTED TO HAVE IMPAIRED REACTIVITY OR BE NON-REACTIVE TO SOME OF THESE MREJ TYPES AND VARIANT SEQUENCES. THESE MREJ TYPES INCLUDE IX, XV, AND XVIII, AS WELL AS TYPES XIX AND XX (ASSOCIATED WITH METHICILLIN-SENSITIVE S. AUREUS; MSSA), AND MREJ SEQUENCES ANNOTATED FROM NON-AUREUS STAPHYLOCOCCUS SPECIES AND NON-STAPYLOCOCCI ORGANISMS. MORE INFORMATION CAN BE FOUND IN TABLE 118. RESULTS FOR MECA/C AND MREJ(MRSA) IN STAPHYLOCOCCUS AUREUS ISOLATES TESTED AND PREDICTED REACTIVITY FOR MREJ TYPES IN THE BIOFIRE BCID2 PANEL'S IFU. OVERALL, EACH AMR GENE ASSAY ON THE BIOFIRE BCID2 PANEL WAS FOUND TO DETECT THE MAJORITY OF AMR GENE TYPES BASED ON TESTING AND ANALYSIS OF AVAILABLE SEQUENCE DATA. HOWEVER, THE MREJ ASSAY MAY BE AFFECTED BY ORGANISM STRAINS THAT CONTAIN SEQUENCE VARIANTS OR REARRANGEMENTS IN THE GENE TARGETS OF THE ASSAY. THE BCID2 PANEL'S MREJ ASSAY IS DESIGNED TO TARGET THE MOST PREVALENT TYPES OF MREJS, BUT NOT ALL KNOWN TYPES. TYPES KNOWN TO HAVE IMPAIRED OR NO REACTIVITY ARE LISTED IN TABLE 118 OF THE BIOFIRE BCID2 PANEL'S IFU, HOWEVER, THE SPECIFIC MREJ VARIANT RECOVERED FROM THIS PATIENT BLOOD CULTURE WAS UNABLE TO BE DEFINITIVELY LINKED TO THOSE LISTED IN TABLE 118. HOWEVER, BASED ON THE SEQUENCE ALIGNMENT, THIS SEQUENCE IS NOT EXPECTED TO BE DETECTED BY THE MREJ ASSAY ON THE BCID2 PANEL, AND MAY BE A NOVEL MREJ TYPE. NOTE: PER THE "LIMITATIONS" SECTION OF THE BIOFIRE BCID2 PANEL IFU, ANTIMICROBIAL RESISTANCE CAN OCCUR VIA MULTIPLE MECHANISMS. A NOT DETECTED RESULT FOR A GENETIC MARKER OF ANTIMICROBIAL RESISTANCE DOES NOT INDICATE SUSCEPTIBILITY TO ASSOCIATED ANTIMICROBIAL DRUGS OR DRUG CLASSES. CULTURE IS REQUIRED TO OBTAIN ISOLATES FOR ANTIMICROBIAL SUSCEPTIBILITY TESTING, AND BIOFIRE BCID2 PANEL RESULTS SHOULD BE USED IN CONJUNCTION WITH CULTURE RESULTS FOR THE DETERMINATION OF SUSCEPTIBILITY OR RESISTANCE. ALL IDENTIFICATION RESULTS PROVIDED BY THE BIOFIRE BCID2 PANEL ARE INTENDED TO BE INTERPRETED IN CONJUNCTION WITH GRAM STAIN RESULTS. THE BIOFIRE BCID2 PANEL RESULTS MUST BE CORRELATED WITH THE CLINICAL HISTORY, EPIDEMIOLOGICAL DATA, AND OTHER DATA AVAILABLE TO THE CLINICIAN EVALUATING THE PATIENT. CLINICAL PERFORMANCE: CLINICAL PERFORMANCE CAN BE FOUND IN TABLE 65. BIOFIRE BCID2 PANEL CLINICAL PERFORMANCE SUMMARY, MECA/C AND MREJ (MRSA) AND TABLE 81. MECA/C AND MREJ (MRSA) PERFORMANCE TABLE (AS COMPARED TO PHENOTYPIC AST METHODS FOR METHICILLIN (OXACILLIN/CEFOXITIN) RESISTANCE ON CULTURED ISOLATE(S) FROM PROSPECTIVE PBC SPECIMENS) OF THE BIOFIRE BCID2 PANEL IFU (WWW.ONLINE-IFU.COM/ITI0048). Manufacturer Narrative: INVESTIGATION: (B)(6) REPORTED A POTENTIAL FALSE NEGATIVE METHICILLIN-RESISTANT STAPHYLOCOCCUS AUREUS (MRSA) RESULT ON THE BIOFIRE BLOOD CULTURE IDENTIFICATION 2 (BCID2) PANEL AFTER TESTING A PATIENT'S BLOOD CULTURE SAMPLE. THE CUSTOMER REPORTED THAT DUE TO THE BIOFIRE BCID2 PANEL RESULT, THE PATIENT'S TREATMENT WAS IMPACTED, RESULTING IN THEIR DEATH. BIOFIRE'S INVESTIGATION INTO THIS EVENT IS ONGOING AND FURTHER INFORMATION HAS BEEN REQUESTED FROM THE CUSTOMER. THE FULL INVESTIGATION AND ASSOCIATED CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. NO REMEDIAL ACTION, CORRECTIVE ACTION, PREVENTIVE ACTION, OR FSCA HAS BEEN DEEMED NECESSARY AT THIS TIME. CONCLUSION: N/A FOR INITIAL REPORT. Manufacturer Narrative: INVESTIGATION: THE PATIENT WAS A 9-YEAR-OLD MALE. THE PATIENT PRESENTED WITH SIGNS AND SYMPTOMS OF TACHYPNEA AND TACHYCARDIA. THE CUSTOMER REPORTED THERE WAS ALSO CONCERN FOR SEPSIS WITH ELEVATED C-REACTIVE PROTEIN (CRP) WHEN BLOOD CULTURES WERE DRAWN. THE PATIENT WAS ALREADY PAN-CYTOPENIC AND NEUTROPENIC. WHEN ASKED WHAT THE UNDERLYING CAUSE OF THE PATIENT'S IMMUNOCOMPROMISED STATUS (PAN-CYTOPENIC AND NEUTROPENIC) WAS, THE CUSTOMER COMMUNICATED THE FOLLOWING: "MULTIFACTORIAL. THE PATIENT WAS PREVIOUSLY HEATHY BUT PRESENTED WITH [FEBRILE INFECTION-RELATED EPILEPSY SYNDROME] FIRES AND ALSO DEVELOPED PANCYTOPENIA (OF UNKNOWN ETIOLOGY DESPITE HEME/ONC INVOLVEMENT) AND [HEMOPHAGOCYTIC LYMPHOHISTIOCYTOSIS] HLH DURING THE PATIENT'S PROLONGED COURSE OF ILLNESS. THE PATIENT RECEIVED CYTOKINE DIRECTED THERAPY WITH BOTH IL-6 BLOCKER AND ANAKINRA AS WELL AS HIGH-DOSE STEROIDS FOR FIRES. ALSO RECEIVED IMMUNOSUPPRESSIVE THERAPY FOR HLH." WHEN ASKED IF THE PATIENT WAS DIAGNOSED WITH ANY ORGAN FAILURE AT THE TIME OF TESTING, SUCH AS SEPTIC SHOCK, THE NEED FOR MECHANICAL VENTILATION, OR ANY OTHER FORM OF ORGAN SUPPORT, THE CUSTOMER REPORTED THAT THE PATIENT WAS ALREADY VENTILATED DUE TO NEUROLOGICAL ISSUE AND RENAL FAILURES (UNKNOWN TIME POINT FOR VENTILATION AND RENAL FAILURE) AND THE PATIENT DEVELOPED HEPATIC FAILURE DURING ANTIBIOTIC TREATMENT REQUIRING CONTINUOUS RENAL REPLACEMENT THERAPY (CRRT). ON (B)(6) 2025, THE PATIENT'S BLOOD CULTURE SAMPLE WAS TESTED ON THE BIOFIRE BCID2 PANEL. THE BIOFIRE BCID2 PANEL REPORTED STAPHYLOCOCCUS SPP. AND S. AUREUS AS DETECTED. THE BIOFIRE BCID2 PANEL WAS POSITIVE FOR MECA/C BUT NEGATIVE FOR MREJ, RESULTING IN A NOT DETECTED RESULT FOR 'MECA/C AND MREJ (MRSA).' NOTE THAT THE BIOFIRE BCID2 PANEL WILL ONLY REPORT A DETECTED RESULT FOR 'MECA AND MREJ (MRSA)' IF S. AUREUS, MREJ, AND MECA/C ASSAYS ARE ALL POSITIVE. MICROSCAN SUSCEPTIBILITY TESTING SHOWED OXACILLIN (MIC= >2 UG/ML) RESISTANCE AND WAS CEFOXITIN SCREEN POSITIVE (MIC= >4 UG/ML). THE ABBOTT CLEARVIEW PENICILLIN-BINDING PROTEIN 2A (PBP2A) CULTURE COLONY TEST WAS POSITIVE FOR PBP2A. GRAM STAIN SHOWED GRAM-POSITIVE COCCI IN CLUSTERS. THE CUSTOMER REPORTED ANOTHER BLOOD CULTURE SAMPLE WAS TESTED ON THE BIOFIRE BCID2 PANEL AND WAS POSITIVE FOR MSSA. THE RUN FILE INFORMATION FOR THIS TEST WAS NOT PROVIDED. THE CUSTOMER REPORTED DUE TO THE BIOFIRE BCID2 PANEL RESULT, THE PATIENT'S ANTIBIOTIC TREATMENT WAS CHANGED TO MSSA DIRECTED THERAPY. THIS INAPPROPRIATE CHANGE IN TREATMENT OCCURRED FOR TWO DAYS. THE CUSTOMER STATED THAT THE FALSE NEGATIVE MRSA RESULT ON THE BIOFIRE BCID2 PANEL "DIRECTLY IMPACTED THE PATIENT'S TREATMENT, RESULTING IN THEIR DEATH." THE CUSTOMER COMMUNICATED THE FOLLOWING REGARDING THE PATIENT'S TREATMENT CHANGE: "WHEN THE PATIENT DETERIORATED, INITIALLY, THE PATIENT WAS ON VANCOMYCIN. HOWEVER, ONCE WE RECEIVED THE RESULT OF [BIOFIRE BCID2 PANEL] INDICATING MSSA, WE SWITCHED TO OXACILLIN FOR DIRECTED THERAPY (AND ADDED CLINDAMYCIN FOR TOXIN MEDIATED DISEASE AS HE ALSO DEVELOPED ERYTHRODERMA) GIVEN THE PATIENT ALSO DEVELOPED AKI (WANTED TO AVOID VANC) AS WELL AS RESPIRATORY FAILURE (DAPTO WAS NOT OPTIMAL). GIVEN PANCYTOPENIA, LINEZOLID WAS NOT HIGH ON CUSTOMER OPTIONS INITIALLY. ONCE CULTURE SHOWED MRSA, CUSTOMER SWITCHED OXACILLIN TO CEFTAROLINE FOR MRSA." THE ANALYSIS OF THE INITIAL BIOFIRE BCID2 PANEL RUN FILE PROVIDED BY THE CUSTOMER SHOWED ROBUST AMPLIFICATION SIGNATURES FOR BOTH MECA/C AND S. AUREUS ASSAYS, WHILE NO MELT/AMPLIFICATION SIGNATURE WAS SEEN FOR THE MREJ ASSAY. IN-HOUSE INVESTIGATION: BIOFIRE RECEIVED ONE TUBE OF ISOLATE AND TWO TUBES OF BLOOD CULTURE ALIQUOTS FOR TESTING. THE TWO BLOOD CULTURE ALIQUOTS WERE TESTED ON THE BIOFIRE BCID2 PANEL. THE BIOFIRE BCID2 PANEL DETECTED S. AUREUS, MECA/C AND MREJ (MRSA) FOR BOTH ALIQUOTS. THE MECA/C AND S. AUREUS ASSAYS EACH HAD ROBUST SIGNATURES, WHILE A LATE SIGNATURE WAS OBSERVED FOR THE MREJ ASSAY. THE PATIENT'S BLOOD CULTURE WAS ALSO TESTED ON AN INTERNAL COMPARATOR PCR ASSAY THAT DIFFERENTIATES BETWEEN MREJA & MREJC. 2/2 REPLICATES CALLED POSITIVE FOR MREJA AND 2/2 REPLICATES CALLED POSITIVE FOR MREJC, EACH WITH A LATE SIGNATURE. ADDITIONALLY, THE ISOLATE WAS TESTED TWICE ON THE BIOFIRE BCID2 PANEL AND S. AUREUS, MECA/C AND MREJ (MRSA) WAS DETECTED BOTH TIMES. AGAIN, THE MECA/C AND S. AUREUS ASSAYS EACH HAD ROBUST SIGNATURES, WHILE A LATE SIGNATURE WAS OBSERVED FOR THE MREJ ASSAY. THE INTERNAL COMPARATOR TESTING FOR THE ISOLATE HAD 3/3 REPLICATES CALLED POSITIVE FOR MREJA AND 2/3 REPLICATES CALLED POSITIVE FOR MREJC, EACH WITH LATE SIGNATURES. GROWTH OF ONE MORPHOLOGY TYPE WAS SEEN FROM CULTURE. ISOLATES WERE TESTED USING VITEK® MS AND VITEK® 2, WITH BOTH PLATFORMS IDENTIFYING THE ORGANISM AS S. AUREUS, VITEK® MS REPORTING A HIGH CONFIDENCE. VITEK® 2 AST CONFIRMED MRSA WITH A CEFOXITIN POSITIVE CALL AND OXACILLIN RESISTANCE. BASED ON THE BIOFIRE BCID2 PANEL'S AMPLIFICATION SIGNATURE PATTERN OF S. AUREUS, MECA/C, AND MREJ ASSAYS, IT IS LIKELY THAT THE PATIENT SAMPLE CONTAINED A MRSA STRAIN CARRYING AN MREJ VARIANT WITH REDUCED REACTIVITY ON THE BIOFIRE BCID2 PANEL. WHOLE GENOME SEQUENCING RESULTS MAY PROVIDE ADDITIONAL SEQUENCE INFORMATION ON THE SPECIFIC MREJ VARIANT. FOR FURTHER INVESTIGATION, THE MRSA ISOLATE WAS ALSO GROWN TO A HIGH TITER IN LIQUID CULTURE IN ORDER TO PREPARE A GLYCEROL STOCK AND WAS SENT OUT FOR WHOLE GENOME SEQUENCING. QUALITY CONTROL (QC) RECORDS FOR POUCH LOT# 3H7D24 (KIT LOT# 1977124) WERE ALSO REVIEWED. THIS POUCH LOT PASSED QC CRITERIA AND WAS FOUND WITHIN SPECIFICATIONS. NO RUN MALFUNCTION OCCURRED AT THE CUSTOMER SITE AND THE FILMARRAY INSTRUMENT (SERIAL NUMBER# (B)(6)) WAS WORKING WITHIN DESIGNED SPECIFICATIONS. CONCLUSION: INVESTIGATION ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21669691&pc=DRB | 3007284006-2025-00050 | 2025-03-05 05:00 | Death | SPECTRANETICS CORPORATION | 2025-03-21 04:00 | DRB | LLD EZ LEAD LOCKING DEVICE | Adverse Event Without Identified Device or Use Problem | Cardiac Perforation; Pericardial Effusion | K142116 |  | 1 | Event Description: A LEAD EXTRACTION PROCEDURE COMMENCED TO REMOVE A RIGHT VENTRICULAR (RV) LEAD DUE TO CIED/POCKET INFECTION. A SPECTRANETICS LLD EZ LEAD LOCKING DEVICE (LLD EZ) WAS INSERTED INTO THE LEAD TO PROVIDE TRACTION. BEGINNING WITH A SPECTRANETICS 16F GLIDELIGHT LASER SHEATH, PROGRESS WAS MADE TO THE MID INNOMINATE VEIN. THEN, SWITCHING TO A SPECTRANETICS 13F TIGHTRAIL ROTATING DILATOR SHEATH, ADVANCEMENT WAS MADE INTO THE RIGHT VENTRICLE. WHILE PULLING WITH TRACTION, THE LEAD RELEASED, EXCEPT THE PATIENT¿S BLOOD PRESSURE DROPPED, AND A PERICARDIAL EFFUSION WAS DETECTED VIA TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE). RESCUE EFFORTS BEGAN, INCLUDING PERICARDIOCENTESIS AND STERNOTOMY. AN RV PERFORATION WAS DISCOVERED AND REPAIRED, BUT THE PATIENT¿S EJECTION FRACTION (EF) REMAINED LOW, AND WAS PLACED ON EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) TO ALLOW THE HEART RECOVER. HOWEVER, THE FOLLOWING DAY, THE DECISION WAS MADE TO WITHDRAW LIFE SUPPORT, AND THE PATIENT DID NOT SURVIVE. THIS REPORT CAPTURES THE LLD EZ DEVICE IN USE WHEN THE RV PERFORATION OCCURRED, REQUIRING INTERVENTION BUT RESULTING IN DEATH. THERE WAS NO ALLEGED MALFUNCTION OF ANY SPECTRANETICS DEVICES IN USE DURING THE PROCEDURE. Manufacturer Narrative: A2) PATIENT DATE OF BIRTH - NOT AVAILABLE FROM FACILITY. A4) PATIENT WEIGHT - NOT AVAILABLE FROM FACILITY. H3) THE LLD WAS DISCARDED, THUS NO RETURNED PRODUCT INVESTIGATION WAS PERFORMED. H6) CARDIAC PERFORATION AND DEATH ARE KNOWN RISKS OF COMPLICATION WITH USE OF THE LLD. SUBMISSION OF THIS REPORT DOES NOT, IN ITSELF, REPRESENT A CONCLUSION BY THE MANUFACTURER AND/OR AUTHORIZED REPRESENTATIVE OR THE NATIONAL COMPETENT AUTHORITY THAT THE CONTENT OF THIS REPORT IS COMPLETE OR ACCURATE, THAT THE MEDICAL DEVICE(S) LISTED FAILED IN ANY MANNER AND/OR THAT THE MEDICAL DEVICE(S) CAUSED OR CONTRIBUTED TO AN ALLEGED DEATH OR DETERIORATION IN THE STATE OF THE HEALTH OF ANY PERSON.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21670748&pc=NIK | 2017865-2025-28093 | 2025-03-02 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-03-21 04:00 | NIK | GALLANT HF | Under-Sensing | Arrhythmia | P030054 |  | 1 | Event Description: IT WAS REPORTED THAT INTERMITTENT VENTRICULAR UNDER SENSING WAS OBSERVED ON THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD). THE PATIENT'S CLINIC REVIEWED TRANSMISSIONS DATED (B)(6) 2025 SENT AT 2.00 AM. EPISODES ON (B)(6) 2025 FROM 11.36 PM TO 11.45 PM DISPLAYED NON-SUSTAINED VENTRICULAR TACHYCARDIA (NSVT) AND NON-SUSTAINED RIGHT VENTRICULAR OVERSENSING (NSRVO). EPISODES ON (B)(6) 2025 AT 2 AM SHOWED NO CARDIAC ACTIVITY. THE PATIENT'S FAMILY NOTIFIED THE CLINIC ON (B)(6) 2025 OF THE PATIENT'S DEATH. A REVIEW OF THE ELECTROCARDIOGRAMS (EGMS) BY TECHNICAL SERVICES REVEALED UNDER SENSING DUE TO THE DIFFERING AMPLITUDE OF R WAVES AND THE PROGRAMMED SETTINGS. DUE TO THIS UNDER SENSING RESULTING FROM VARYING R WAVE AMPLITUDES THE DEVICE KEPT GOING IN AND OUT OF THE EPISODE. THE DEVICE WAS PROGRAMMED WITH VFTA (VENTRICULAR FIBRILLATION THERAPY ASSISTANCE) ON, BUT IT DID NOT SENSE TWO CONSECUTIVE DISCRIMINATION SIGNALS. THERE WERE ALSO INSTANCES WHERE THE DEVICE BIVENTRICULAR PACED INTO REFRACTORY TISSUE WHICH APPEARED AS NON-CAPTURE BUT WAS DUE TO UNDER SENSING BEFORE BIVENTRICULAR PACING. THE PHYSICIAN BELIEVED THE PATIENT'S OTHER PREEXISTING MEDICAL CONDITIONS CONTRIBUTED TO HIS DEATH. THE PATIENT HAD CHRONIC KIDNEY FAILURE AND REFUSED DIALYSIS TREATMENT. INFORMATION RECEIVED NOTES THE DEATH OCCURRED AT THE PATIENT'S RESIDENCE AT NIGHT, THE PATIENT WAS NOTED TO PASS PEACEFULLY. THE PHYSICIAN BELIEVED THE SYSTEM WAS FUNCTIONING AS PROGRAMMED. THE PHYSICIAN DID NOT BELIEVE THERE WAS ANY MALFUNCTION OF ABBOTT PRODUCTS PRIOR TO, DURING, OR AFTER THE EVENT. THE SYSTEM REMAINED IMPLANTED AND WAS NOT EXPECTED TO BE RETURNED. Manufacturer Narrative: THE PROVIDED SESSION RECORDS WERE REVIEWED BY TECHNICAL SERVICES (TS) AND IT WAS DETERMINED THAT THE DEVICE WAS UNDER-SENSING R-WAVES. VENTRICULAR FIBRILLATION THERAPY ASSISTANCE (VFTA) WAS NOT TRIGGERED DUE TO THE DEVICE SENSING IN AND OUT OF THE EPISODE. THE DEVICE PERFORMED PER PROGRAMMED SETTINGS. A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED, AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21649989&pc=DSP | 2249723-2025-0001320 | 2025-03-04 05:00 | Death | DATASCOPE CORP. | 2025-03-20 04:00 | DSP | CARDIOSAVE HYBRID, TYPE B PLUG | Fluid/Blood Leak | Insufficient Information | K112372 |  | 1 | Event Description: IT WAS REPORTED THAT CARDIOSAVE INTRA-AORTIC BALLOON PUMP (IABP) HAD BLOOD BACK. THIS WAS THE SECOND UNIT USED DURING THE TREATMENT. THIS UNIT WAS USING AN ALREADY USED CATHETER FROM A PREVIOUS MACHINE WITH GAS LOSS AND SIGNS OF BLOOD BACK. PATIENT DIED WHILE ON THIS UNIT AT 7:30 AM ON THE NEXT DAY.Event Description: N/AEvent Description: N/A Manufacturer Narrative: UPDATED FIELDS - B4, E1(POSTAL CODE), G3, G6, H2, H11. CORRECTED FIELDS - D10. Manufacturer Narrative: A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF OUR INVESTIGATION. Manufacturer Narrative: UPDATED FIELDS - B2, B4, D9, E1(INITIAL REPORTER, TELEPHONE, EMAIL), G3, G6, H2, H3, H6 ( TYPE OF INVESTIGATION, INVESTIGATION FINDINGS, COMPONENT CODES, INVESTIGATION CONCLUSION), H11. A GETINGE FIELD SERVICE ENGINEER (FSE) EVALUATED THE UNIT AND REPLACED PIM (0997-00-1178) DUE TO BLOOD BACK. COMPLETED REPAIR SUCCESSFULLY. PRODUCT PASSED ALL FUNCTIONAL AND SAFETY TESTS PER MANUFACTURER SPECIFICATIONS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21649991&pc=DSP | 2249723-2025-0001319 | 2025-03-04 05:00 | Death | DATASCOPE CORP. | 2025-03-20 04:00 | DSP | CARDIOSAVE HYBRID, TYPE B PLUG | Fluid/Blood Leak; Gas/Air Leak | Insufficient Information | K112372 |  | 1 | Event Description: N/A.Event Description: AT 2 PM A CATHETER WAS INSERTED AND PATIENT WAS TRANSPORTED TO ICU AT 3 PM. IT WAS REPORTED THAT CARDIOSAVE INTRA-AORTIC BALLOON PUMP (IABP) EXPERIENCED GAS LOSS ALARM. HOWEVER IT WAS STILL USED ON THE PATIENT. THE CATHETER HAD A TEAR, FILLED WITH BLOOD AND CONDENSATION. HOSPITAL STAFF STILL KEPT USING THE EQUIPMENT ON THE PATIENT. LATER AT NIGHT THE UNIT WAS SWITCHED WITH ANOTHER CARDIOSAVE DEVICE AND USED THE SAME CATHETER IN THIS UNIT TO CONTINUE PUMPING. PATIENT EXPIRED AT 7:30 AM IN THE MORNING. Manufacturer Narrative: UPDATED FIELDS - B4, D9, E1 (INITIAL REPORTER, EMAIL AND TELEPHONE), G3, G6, H2, H3, H6 (TYPE OF INVESTIGATION, INVESTIGATION FINDINGS, COMPONENT CODES, INVESTIGATION CONCLUSIONS), H11. A GETINGE FIELD SERVICE ENGINEER (FSE) WAS DISPATCHED TO INVESTIGATE THE ISSUE. IN ORDER TO FIX THE ISSUE, THE FSE REPLACED THE PNEUMATIC MODULE ASSY (0997-00-1178). THE FSE THEN COMPLETED THE REPAIR SUCCESSFULLY. THE UNIT PASSED ALL FUNCTIONAL AND SAFETY TESTS PER MANUFACTURER SPECIFICATIONS. Manufacturer Narrative: A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF OUR INVESTIGATION. Manufacturer Narrative: UPDATED FIELDS - B2, B4, E1(EVENT SITE POSTAL CODE), G3, G6, H2, H6(INVESTIGATION FINDINGS), H11 CORRECTED FIELDS - H1. Manufacturer Narrative: UPDATED : B4, G3, G6, H2, H6(MEDICAL DEVICE ¿ PROBLEM CODE), H11. CORRECTED FIELDS - D10. Manufacturer Narrative: UPDATED DATA: B4, G3, G6, H1, H2,H11, EVENT SITE STATE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21650256&pc=DSP | 2248146-2025-0000158 |  | Death | DATASCOPE CORP. | 2025-03-20 04:00 | DSP | SENSATION PLUS 7.5FR. 40CC IAB & ACCESSORIES | Fluid/Blood Leak; Leak/Splash | Insufficient Information | K122628 |  | 1 | Event Description: IT WAS REPORTED THAT THE INTRA-AORTIC BALLOON (IAB) CATHETER EXHIBITED BLOOD BACKFLOW ON UNIT (B)(6). DESPITE THIS OBSERVATION, THE SAME CATHETER WAS SUBSEQUENTLY USED ON UNIT (B)(6). AFTER THE PATIENT WAS DECLARED DECEASED, ATTEMPTS WERE MADE TO REMOVE THE BALLOON; HOWEVER, REMOVAL WAS UNSUCCESSFUL DUE TO CLOT FORMATION. AS A RESULT, THE CATHETER HAD TO BE CUT FOR REMOVAL. Manufacturer Narrative: A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF OUR INVESTIGATION. Manufacturer Narrative: UPDATED FIELDS: B4, G3, G6, H2, H6 (TYPE OF INVESTIGATION, MEDICAL DEVICE ¿ PROBLEM CODE, INVESTIGATION FINDINGS, COMPONENT CODES, INVESTIGATION CONCLUSIONS), H11. CORRECTED FIELDS: D7A, D10. NO DECON OR VISUAL INSPECTION PERFORMED SINCE PRODUCT WAS NOT RETURNED AND COULD NOT BE EVALUATED. THEREFORE, WE WERE UNABLE TO CONFIRM OR DETERMINE A ROOT CAUSE FOR THE REPORTED FAILURE. THE FAILURE MODE IS ADDRESSED IN THE RISK FILE AND IS OPERATING WITHIN ITS RISK PROFILE. THE IFU ADDRESSES THE REPORTED FAILURE. THE INVESTIGATION DOES NOT INDICATE THAT THE DEVICE WAS INADVERTENTLY RELEASED AS NON-CONFORMING OR AN ADULTERATED PRODUCT OR WAS A COUNTERFEIT. EACH IAB MANUFACTURED IS 100% INSPECTED AND THE CONTROLS PUT IN PLACE DURING THE MANUFACTURING OF THE IAB WOULD CATCH A DEFECT AND NOT ALLOW NON-CONFORMING DEVICE TO BE RELEASED. THE TREND REVIEW DID NOT IDENTIFY AN ADVERSE TREND (INCREASE IN NUMBER OF COMPLAINTS OVER PAST THREE (3) MONTHS). BASED ON THE RATIONAL PROVIDED ABOVE, NO ESCALATION TO THE CAPA PROCESS IS REQUIRED. COMPLAINT ID: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21650353&pc=DTB | 2649622-2025-07729 | 2025-02-24 05:00 | Death | MPRI | 2025-03-20 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Under-Sensing | Insufficient Information | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE RIGHT ATRIAL (RA) LEAD EXHIBITED UNDERSENSING ON STORED ELECTROGRAMS (EGM) TRIGGERING DEVICE CLASSIFIED FALSE TERMINATION OF ATRIAL TACHYCARDIA/ATRIAL FIBRILLATION (AT/AF) EVENT(S) AT TIMES AND RESULTING IN FALSE TERMINATION OF AT/AF DETECTION. THE LEAD REMAINS IN PATIENT. IT WAS ALSO NOTED THAT THE PATIENT PASSED AWAY APPROXIMATELY TWO WEEKS AFTER THE REMOTE TRANSMISSION.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT'S DEATH WAS NOT LINKED TO THE PERFORMANCE OF THE RA LEAD. Manufacturer Narrative: CONTINUATION OF D10: 3387-40 DBS LEAD, 3387-40 DBS LEAD, 748251 NEURO EXTENSION, 748251 NEURO EXTENSION WAS IMPLANTED ON (B)(6) 2003. 64001 NEURO ADAPTOR, 37612 DBS IPG WAS IMPLANTED ON ( B)(6) -2016. 5076-58 LEAD WAS IMPLANTED ON (B)(6) -2019. 37602 DBS IPG WAS IMPLANTED ON (B)(6) -2021. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21650737&pc=MNS | 2518422-2025-037726 | 2024-07-23 04:00 | Death | RESPIRONICS, INC. | 2025-03-20 04:00 | MNS | DREAMSTATION ST30 | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K102465 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION ST30 DEVICE. THE PATIENT HAS PASSED AWAY. THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR EVALUATION. THE DEVICE OPERATED AS DESIGNED. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21651295&pc=MSD | 3006260740-2025-01655 | 2025-02-03 05:00 | Death | C.R. BARD, INC. (BASD) -3006260740 | 2025-03-20 04:00 | MSD | HEMOSTAR | Material Discolored; Fracture; Material Split, Cut or Torn; Air/Gas in Device | Air Embolism; Dyspnea | K051748 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT STARTED HEMODIALYSIS IN THE LEFT FOREARM UNDER THE INTERNAL FISTULA PUNCTURE, ULTRAFILTRATION FOR ABOUT TWO HOURS. THE NURSE WAS READY TO GIVE ROUTINE CARE TO THE PATIENT'S INTRACERVICAL DIALYSIS CATHETER, AND IMMEDIATELY FOUND THAT THE PATIENT'S RESPIRATION WAS SHORT OF BREATH. THE NURSE REPORTED TO THE DOCTOR IMMEDIATELY, PLACED THE PATIENT IN LYING POSITION, AND STOP ULTRAFILTRATION. IT WAS IDENTIFIED THAT THE LONG-TERM TUBE ARTERIAL END OF THE CLAMP, 1 CM BELOW THE CATHETER WAS RUPTURED, THE CATHETER CLAMP INTACT, WAS IN CLAMPED CLOSED STATE. REPORTEDLY, THE PATIENT EXPERIENCED CONFUSION, CALLED OUT, BILATERAL PUPIL DILATION, RESPIRATION 29-35 TIMES/MIN, HEART RATE 120-136 TIMES/MIN, OXYGEN SATURATION 65-69%, AND HEMODIALYSIS TREATMENT WAS TERMINATED. THE HCP ACTIVELY RESUSCITATED THE CAROTID ARTERY AFTER THE CTA REPORT SUGGESTS THAT SOME OF THE INTRACEREBRAL VEINS AND FASCIA GAS SHADOW IN THE SINUS, CONSIDERED THE POSSIBILITY OF AN AIR EMBOLISM. AFTER THAT, THE PATIENT WAS TRANSFERRED TO ICU FOR TREATMENT. FURTHERMORE, THE TREATMENT WAS STILL INEFFECTIVE, AND THE PATIENT WAS PRONOUNCED DEAD.Event Description: IT WAS REPORTED THAT THE PATIENT STARTED HEMODIALYSIS IN THE LEFT FOREARM UNDER THE INTERNAL FISTULA PUNCTURE, ULTRAFILTRATION FOR ABOUT TWO HOURS. THE NURSE WAS READY TO GIVE ROUTINE CARE TO THE PATIENT'S INTRACERVICAL DIALYSIS CATHETER, AND IMMEDIATELY FOUND THAT THE PATIENT'S RESPIRATION WAS SHORT OF BREATH. THE NURSE REPORTED TO THE DOCTOR IMMEDIATELY, PLACED THE PATIENT IN LYING POSITION, AND STOP ULTRAFILTRATION. IT WAS IDENTIFIED THAT THE LONG-TERM TUBE ARTERIAL END OF THE CLAMP, 1CM BELOW THE CATHETER WAS RUPTURED, THE CATHETER CLAMP INTACT, WAS IN CLAMPED CLOSED STATE. REPORTEDLY, THE PATIENT EXPERIENCED CONFUSION, CALLED OUT, BILATERAL PUPIL DILATION, RESPIRATION 29-35 TIMES / MIN, HEART RATE 120-136 TIMES / MIN, OXYGEN SATURATION 65-69%, AND HEMODIALYSIS TREATMENT WAS TERMINATED. THE HCP ACTIVELY RESUSCITATED THE CAROTID ARTERY AFTER THE CTA REPORT SUGGESTS THAT SOME OF THE INTRACEREBRAL VEINS AND FASCIA GAS SHADOW IN THE SINUS, CONSIDERED THE POSSIBILITY OF AN AIR EMBOLISM. AFTER THAT, THE PATIENT WAS TRANSFERRED TO ICU FOR TREATMENT. FURTHERMORE, THE TREATMENT WAS STILL INEFFECTIVE, AND THE PATIENT WAS PRONOUNCED DEAD. Manufacturer Narrative: H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WAS PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. Manufacturer Narrative: H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED, AND THIS LOT MET ALL RELEASE CRITERIA. INVESTIGATION SUMMARY: ONE ELECTRONIC PHOTO WAS REVIEWED. THE PHOTOGRAPH SHOWED AN INSERTED DIALYSIS CATHETER WITH EXTENSION LEGS AND SOME OF THE BIFURCATION SIGNIFICANTLY YELLOWED, AS WELL AS THE MARKER INSERTS ON THE CLAMPS. THE TWO HUBS WERE BEING HELD IN A GLOVED HAND, BEING SURROUNDED LOOSELY BY GAUZE. BOTH CLAMPS APPEARED TO BE CLOSED. THE RED (ARTERIAL) EXTENSION LEG OF THE CATHETER HAD A LARGE, EVEN, CLEAN PARTIAL DISCONTINUITY ON THE PROXIMAL SIDE OF THE CLOSED CLAMP THAT APPEARED TO BE A CUT, BASED ON ITS EVENNESS AND APPARENT SMOOTHNESS. THE CUT WAS ROUGHLY TRANSVERSE, THOUGH AT A SHALLOW ANGLE RELATIVE TO THE TUBING AXIS. NO BLEEDING/LEAKAGE FROM THE CATHETER WAS EVIDENT, BUT AIR WAS VISIBLE WITHIN THE EXTENSION LEG ON THE SIDE OF THE CUT. THEREFORE, BASED ON THE PHOTO/VIDEO REVIEW, THE REPORTED RUPTURE WILL BE MORE SPECIFICALLY DESCRIBED AS A CUT, AND CONFIRMED, AND THE PROBLEMS OF DISCOLORATION AND AIR WITHIN THE DEVICE CAN BE CONFIRMED ALSO BASED ON THE PROVIDED PHOTOS/VIDEO. A DEFINITIVE ROOT CAUSE FOR THE DISCOLORATION, CUT, AND AIR LEAK COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE A LABELING REVIEW IS NOT REQUIRED. THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21651311&pc=NPU | 2015691-2025-02089 | 2025-01-26 05:00 | Death | EDWARDS LIFESCIENCES | 2025-03-20 04:00 | NPU | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Perivalvular Leak | Hemolysis | P140031 |  | 1 | Event Description: AS REPORTED TO THE IMPLANT PATIENT REGISTRY (IPR), 2 DAYS POST-IMPLANT OF A 29 MM SAPIEN 3 ULTRA RESILIA VALVE IN THE AORTIC POSITION, THE PATIENT EXPIRED. AFTER A REVIEW OF MEDICAL RECORDS PROVIDED, IT WAS CONFIRMED THAT THE PATIENT HAD EXPIRED 2-DAYS POST INDEX TMVR PROCEDURE. THE CAUSE OF DEATH WAS CONFIRMED AS MULTIORGAN FAILURE. THE INFORMATION AVAILABLE DOES NOT SUGGEST THERE WAS A MALFUNCTION OF THE EDWARDS DEVICE. HOWEVER, THE PATIENT DID HAVE MILD-MODERATE PARAVALVULAR LEAK (PVL), WHICH WAS FELT TO BE ACCEPTABLE AT THE TIME OF IMPLANTATION. HOWEVER, THE PATIENT THEN DEVELOPED ACUTE TUBAL NECROSIS WHICH REQUIRED HEMODIALYSIS, HEMOLYSIS AND SEVERE METABOLIC ACIDOSIS. THE PVL WAS THEN CLOSED USING THE TOOTSIE ROLL TECHNIQUE WITH A 10MM X 5 CM VIABAHN COVERED STENT AND 12MM AVP II DEVICE. HOWEVER, THE PATIENT CONTINUED TO BE ACIDOTIC WITH HYPERKALEMIA THAT WAS REFRACTORY TO DIALYSIS. THEY WERE MADE A DO NOT RESUSCITATE (DNR) AND PASSED AWAY FROM MULTIORGAN FAILURE. Manufacturer Narrative: THE EVENT REPORTED IS AN ANTICIPATED IN THE RISK MANAGEMENT DOCUMENTATION FOR TRANSCATHETER HEART VALVE PROCEDURES. A PREVIOUS INVESTIGATION INTO THIS TYPE OF EVENT IS CAPTURED IN AN EDWARDS LIFESCIENCES TECHNICAL SUMMARY AND APPLIES TO THIS COMPLAINT. ADDITIONAL ASSESSMENT OF THE FAILURE MODE IS NOT REQUIRED AT THIS TIME. DUE TO THE UNAVAILABILITY OF THE COMPLAINT DEVICE, ENGINEERING WAS UNABLE TO PERFORM ANY VISUAL INSPECTION, FUNCTIONAL TESTING, OR DIMENSIONAL ANALYSIS. PER THE TECHNICAL SUMMARY: THE IFU, CURRENT RISK MITIGATIONS INCLUDING DESIGN AND MANUFACTURING CONTROLS, AND TRAINING MANUALS HAVE BEEN REVIEWED, AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. PER THE INSTRUCTIONS FOR USE (IFU), PARAVALVULAR LEAK (PVL) IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH BIOPROSTHETIC HEART VALVES AND THE TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. PARAVALVULAR LEAK REFERS TO BLOOD FLOWING THROUGH A CHANNEL BETWEEN THE STRUCTURE OF THE IMPLANTED VALVE AND THE LANDING ZONE, AS A RESULT OF A LACK OF APPROPRIATE SEALING OF THE VALVE TO THE TARGET SITE. SOME PVL IS NOT UNCOMMON POST-DEPLOYMENT. MANY CASES ARE MILD TO MODERATE, AND EITHER RESOLVE OVER TIME OR DO NOT CAUSE SYMPTOMS. OTHERS MAY BE MORE CLINICALLY SIGNIFICANT AND REQUIRE INTERVENTION. THE PATIENT SCREENING MANUAL AND THE PROCEDURE DIDACTIC IDENTIFY SEVERAL PROCEDURAL AND ANATOMICAL FACTORS WHICH COULD CONTRIBUTE TO PVL, INCLUDING DEVICE MALPOSITION, INACCURATE MEASUREMENT OF THE LANDING ZONE, UNEVEN DISTRIBUTION OF CALCIUM ON THE LANDING ZONE, BULKY OR SEVERE CALCIFICATION, AN ELLIPTICAL ANNULUS SHAPE, AND VALVE UNDER-SIZING. THE MECHANISM BEHIND WORSENING OR LATE PVL IS NOT WELL UNDERSTOOD BUT MAY BE RELATED TO CARDIAC REMODELING. HEMOLYSIS IS THE DESTRUCTION OF RED BLOOD CELLS. THERE ARE MULTIPLE EXTRINSIC AND INTRINSIC FACTORS INCLUDING USE OF EXTRACORPOREAL CIRCULATION, TRANSFUSION, INFECTION, TUMORS, AUTOIMMUNE DISORDERS, MEDICATION SIDE EFFECTS, A METABOLIC ABNORMALITY, AND RED BLOOD CELL MEMBRANE INSTABILITY. NORMAL RED BLOOD CELLS (ERYTHROCYTES) HAVE A LIFESPAN OF ABOUT 120 DAYS. AFTER THE LIFESPAN IS OVER THE RED BLOOD CELLS BREAK DOWN AND ARE REMOVED FROM THE CIRCULATION BY THE SPLEEN. IN SOME MEDICAL CONDITIONS, OR BECAUSE OF TAKING CERTAIN MEDICATIONS, THIS BREAKDOWN OF RED BLOOD CELLS IS INCREASED. MEDICATIONS THAT HAVE BEEN ASSOCIATED WITH HEMOLYSIS INCLUDE ACETAMINOPHEN, PENICILLIN, AND OTHER PAIN MEDICATIONS. RED CELLS MAY BE DESTROYED DUE TO DEFECTS IN THE CELLS THEMSELVES. ANOTHER CAUSE OF HEMOLYTIC ANEMIA IS BREAK DOWN DUE TO MECHANICAL DAMAGE, SUCH AS FROM ARTIFICIAL HEART VALVES OR HEART-LUNG BYPASS; OR THEY AS A RESULT OF TURBULENT BLOOD FLOW PATTERN AND ERYTHROCYTE DESTRUCTION CAUSED BY PVL OR CENTRAL REGURGITATION. EDWARDS EXTENSIVELY TRAINS PHYSICIANS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVES (ALL MODELS). THE DEVICE TRAINING MANUALS INSTRUCTS THE OPERATOR ON PROPER IMAGING SCREENING REQUIREMENTS, INCLUDING THE USE OF GOOD QUALITY ECHOCARDIOGRAPHY AND/OR COMPUTED TOMOGRAPHY (CT) TO APPROPRIATELY MEASURE THE LANDING ZONE, ASSESS THE CONTENT AND DISTRIBUTION OF CALCIUM, AND OTHER PATIENT ANATOMICAL FACTORS. MULTIPLE IMAGING MODALITIES SHOULD BE CONSIDERED DURING VALVE SIZE SELECTION. CONTRAINDICATIONS, IMPORTANT CONSIDERATIONS WHEN ASSESSING THE VALVE, AND CHOOSING THE PROPER TRANSCATHETER HEART VALVE ARE ALSO DISCUSSED. THE DEVICE TRAINING MANUALS ALSO INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING PATIENTS SCREENING AND PROCEDURAL CONSIDERATIONS. DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES ARE ALSO INCLUDED. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST/INDICATE THE MITRAL ANNULUS IS OVAL IN SHAPE, WHEREAS THE EDWARDS SAPIEN VALVE IS CIRCULAR. IMPLANTING THE EDWARDS SAPIEN VALVE INTO THE NATIVE MITRAL ANNULUS CAN LEAD TO PARAVALVULAR LEAK (PVL), PARTICULARLY IF THE MITRAL ANNULUS IS CALCIFIED ENOUGH TO ENSURE SECURE IMPLANTATION. THIS MISMATCH IN SHAPES AND THE PRESENCE OF CALCIFICATION CAN INCREASE THE RISK OF PVL AT DEPLOYMENT AND COULD HAVE CAUSED OR CONTRIBUTED TO THE PVL WITH SUBSEQUENT HEMOLYSIS/HEMOLYTIC ANEMIA. OVER TIME, PVL MAY RESOLVE AS THE VALVE BECOMES ENDOTHELIALIZED WITHIN THE PATIENT'S ANNULUS. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21651527&pc=OUT | 2029214-2025-00721 | 2024-10-15 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-03-20 04:00 | OUT | PIPELINE | Adverse Event Without Identified Device or Use Problem | Hemorrhagic Stroke; Ischemia Stroke |  |  | 1 | Event Description: ANALYZING THE SAFETY AND EFFICACY OF PIPELINE EMBOLIZATION DEVICE IN PEDIATRIC ANEURYSMS: INSIGHT FROM A MULTICENTER COHORT AND POOLED ANALYSIS. THE TIME FRAME OF THIS STUDY WAS: PATIENTS WERE TREATED WITH PED FROM NOVEMBER 2014 TO OCTOBER 2019. LITERATURE REVIEW INCLUDED CASES FROM 2007 TO 2023. MULTIPLE MANUFACTURER¿S DEVICES WERE USED IN THE STUDY POPULATION. THE FOLLOWING MEDTRONIC DEVICES WERE USED: PIPELINE EMBOLIZATION DEVICE (PED) CLASSIC AND PED FLEX, DEPLOYED THROUGH MARKSMAN AND PHENOM 27 MICROCATHETERS. DEATHS OCCURRED IN THE STUDY POPULATION. THE CAUSES OF DEATH WERE: ACUTE IN-STENT THROMBOSIS LEADING TO ISCHEMIC STROKES (2 CASES). HEMORRHAGIC STROKES FOLLOWING ANEURYSM RUPTURES (2 CASES). (PAGE 4) AMONG PATIENT ADVERSE EVENTS INCLUDED: ISCHEMIC STROKES DUE TO ACUTE IN-STENT THROMBOSIS (2 CASES). HEMORRHAGIC STROKES FOLLOWING ANEURYSM RUPTURES (3 CASES). SYMPTOMATIC PARENT ARTERY STENOSIS (1 CASE). ASYMPTOMATIC PARENT ARTERY STENOSIS (1 CASE). MAJOR MORBIDITY (2 CASES). NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID NV UNK PIPELINE (LOT: UNKNOWN); PRODUCT TYPE: ; IMPLANT DATE N/A; EXPLANT DATE N/A PRODUCT ID NV UNK PIPELINE (UNKNOWN); PRODUCT TYPE: ; IMPLANT DATE N/A; EXPLANT DATE N/A G2: CITATION: AUTHORS: ZHAO, Y., LU, J., LI, Z., CHEN, X., YANG, X., ZHAO, Y., WANG, S., HAO Q.. ANALYZING THE SAFETY AND EFCACY OF PIPELINE EMBOLIZATION DEVICE IN PEDIATRIC ANEURYSMS: INSIGHT FROM A MUL-TICENTER COHORT AND POOLED ANALYSIS. ACTA NEUROCHIRURGICA 166:412 2024. DOI: 10.1007/S00701-024-06306-5. EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21651736&pc=OZD | 1220648-2025-26842 | 2025-02-22 05:00 | Death | ABIOMED, INC. | 2025-03-20 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. THE PATIENT'S CONDITION POST IMPELLA WAS NOTED AS CRITICAL AND APPROXIMATELY ONE DAY LATER THE PATIENT UNDERWENT A FEMORAL CROSSOVER BYPASS DUE TO LIMB ISCHEMIA THAT WAS NOTED TO BE RELATED TO IMPELLA ACCESS. THE PATIENT WOULD LATER EXPIRE ON SUPPORT THAT EVENING. NO FURTHER DETAILS REGARDING THE EVENT WERE NOTED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM, SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ Manufacturer Narrative: THE INVESTIGATION OF LIMB ISCHEMIA HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21651981&pc=MVK | 3008642652-2025-02922 | 2025-02-20 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-03-20 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing; Incorrect Interpretation of Signal | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED THREE APPROPRIATE TREATMENTS AND FIVE INAPPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 14:16:38 ON (B)(6) 2025. AT 23:36:24, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS AF @ 140 BPM WITH NSVT. AT 23:37:02, THE PATIENT RECEIVED THE FIRST INAPPROPRIATE TREATMENT. AF WITH RVR, NSVT AND OVERSENSING OF CARDIAC ACTIVITY CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS AF @ 20 BPM WITH NSVT. POST SHOCK RHYTHM WAS AF @ 20 BPM WITH MOTION ARTIFACT. AT 23:46:07, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS AF WITH RVR @ 150 BPM WITH PVCS/NSVT. AT 23:46:45, THE PATIENT RECEIVED THE SECOND INAPPROPRIATE TREATMENT. AF WITH RVR, NSVT AND OVERSENSING OF CARDIAC ACTIVITY CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS NSVT. POST SHOCK RHYTHM WAS VT @ 180 BPM. AT 23:47:11, THE PATIENT RECEIVED THE FIRST APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 180 BPM. POST SHOCK RHYTHM WAS AF @ 130 BPM WITH PVCS/NSVT. AT 23:48:29, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VT @ 180 BPM. AT 23:49:17, THE PATIENT RECEIVED THE THIRD INAPPROPRIATE TREATMENT. AF WITH RVR, NSVT AND OVERSENSING OF CARDIAC ACTIVITY CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS NSVT. POST SHOCK RHYTHM WAS VT @ 180 BPM FOR 69 SECONDS SELF-CONVERTING TO AF @ 140 BPM WITH PVCS/NSVT. AT 23:50:44, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VT @ 180 BPM. AT 23:51:50, THE PATIENT RECEIVED THE SECOND APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 180 BPM. POST SHOCK RHYTHM WAS SINUS BRADYCARDIA @ 20 BPM. AT 23:59:30, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VT @ 210 BPM. AT 00:00:00 ON (B)(6) 2025, THE PATIENT RECEIVED THE THIRD APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 210 BPM. POST SHOCK RHYTHM WAS SINUS RHYTHM @ 60 BPM. AT 00:35:32, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AND TACTILE ARTIFACT. AT 00:36:39, THE PATIENT RECEIVED THE FOURTH INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH TACTILE ARTIFACT. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON LIFE-SUSTAINING RHYTHM, WITH INTERMITTENT CARDIAC ACTIVITY AND TACTILE ARTIFACT. AT 00:37:07, THE PATIENT RECEIVED THE FIFTH INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH TACTILE ARTIFACT. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON LIFE-SUSTAINING RHYTHM, WITH TACTILE ARTIFACT. THE DEVICE WAS SHUT DOWN AT 07:13:28 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21651991&pc=MVK | 3008642652-2025-02924 | 2025-02-06 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-03-20 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Failure to Convert Rhythm | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT'S DATE OF DEATH WAS NOT REPORTED. THE PATIENT RECEIVED TEN APPROPRIATE TREATMENTS ON THEIR LAST WEAR DATE. THE DEVICE WAS STARTED UP AT 04:19:55 ON (B)(6) 2025. AT 22:31:44 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS SINUS RHYTHM @ 80 BPM WITH PVCS TRANSITIONING TO VT @ 220 BPM. AT 22:32:19, THE PATIENT RECEIVED THE FIRST APPROPRIATE TREATMENT. RHYTHM THE TIME OF TREATMENT WAS VT @ 220 BPM. POST SHOCK RHYTHM WAS VT @ 180 BPM. AT 22:32:50, THE PATIENT RECEIVED THE SECOND APPROPRIATE TREATMENT. RHYTHM THE TIME OF TREATMENT WAS VT @ 220 BPM. POST SHOCK RHYTHM WAS VT @ 220 BPM. AT 22:33:24, THE PATIENT RECEIVED THE THIRD APPROPRIATE TREATMENT. RHYTHM THE TIME OF TREATMENT WAS VT @ 250 BPM. POST SHOCK RHYTHM WAS VT @ 220 BPM. AT 22:33:55, THE PATIENT RECEIVED THE FOURTH APPROPRIATE TREATMENT. RHYTHM THE TIME OF TREATMENT WAS VT @ 240 BPM. POST SHOCK RHYTHM WAS VT @ 220 BPM. AT 22:34:29, THE PATIENT RECEIVED THE FIFTH APPROPRIATE TREATMENT. RHYTHM THE TIME OF TREATMENT WAS VT @ 250 BPM. POST SHOCK RHYTHM WAS VF/FINE VF. AT 22:47:30, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS FINE VF. AT 22:48:35, THE PATIENT RECEIVED THE SIXTH APPROPRIATE TREATMENT. RHYTHM THE TIME OF TREATMENT WAS FINE VF. POST SHOCK RHYTHM WAS ASYSTOLE WITH CPR/MOTION ARTIFACT. POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY. AT 22:50:48, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS FINE VF. AT 22:52:05, THE PATIENT RECEIVED THE SEVENTH APPROPRIATE TREATMENT. RHYTHM THE TIME OF TREATMENT WAS FINE VF. POST SHOCK RHYTHM WAS ASYSTOLE WITH CPR/MOTION ARTIFACT. POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY. AT 22:53:07, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF. AT 22:55:13, THE PATIENT RECEIVED THE EIGHTH APPROPRIATE TREATMENT. RHYTHM THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY. POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY. AT 22:56:57, THE PATIENT RECEIVED THE NINTH APPROPRIATE TREATMENT. RHYTHM THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS IDIOVENTRICULAR RHYTHM @ 80 BPM WITH CPR/MOTION ARTIFACT. AT 22:58:00, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS FINE VF. AT 22:59:55, THE PATIENT RECEIVED THE TENTH APPROPRIATE TREATMENT. RHYTHM THE TIME OF TREATMENT WAS FINE VF. POST SHOCK RHYTHM WAS IDIOVENTRICULAR RHYTHM @ 70 BPM WITH CPR/MOTION ARTIFACT. THE DEVICE WAS SHUT DOWN AT 01:30:36 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21651995&pc=MVK | 3008642652-2025-02921 | 2025-02-24 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-03-20 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED EIGHT INAPPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 13:55:37 ON (B)(6) 2025. AT 18:43:13, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT. AT 18:45:34, THE PATIENT RECEIVED THE FIRST INAPPROPRIATE TREATMENT. CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM THE TIME OF TREATMENT WAS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON-LIFE-SUSTAINING RHYTHM, WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT. AT 18:46:59, THE PATIENT RECEIVED THE SECOND INAPPROPRIATE TREATMENT. CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM THE TIME OF TREATMENT WAS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON-LIFE-SUSTAINING RHYTHM, WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT. AT 18:48:54, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT. AT 18:50:33, THE PATIENT RECEIVED THE THIRD INAPPROPRIATE TREATMENT. CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM THE TIME OF TREATMENT WAS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON-LIFE-SUSTAINING RHYTHM, WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT. AT 18:51:48, THE PATIENT RECEIVED THE FOURTH INAPPROPRIATE TREATMENT. CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM THE TIME OF TREATMENT WAS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. AT 18:52:37, THE PATIENT RECEIVED THE FIFTH INAPPROPRIATE TREATMENT. CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM THE TIME OF TREATMENT WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. AT 18:53:09, THE PATIENT RECEIVED THE SIXTH INAPPROPRIATE TREATMENT. CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM THE TIME OF TREATMENT WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. AT 18:55:11, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT. AT 18:57:33, THE PATIENT RECEIVED THE SEVENTH INAPPROPRIATE TREATMENT. CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM THE TIME OF TREATMENT WAS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON-LIFE-SUSTAINING RHYTHM, WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT. AT 18:59:50, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT. AT 19:01:27, THE PATIENT RECEIVED THE EIGHTH INAPPROPRIATE TREATMENT. CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM THE TIME OF TREATMENT WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT. THE ELECTRODE BELT WAS DISCONNECTED AT 23:30:26 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21652004&pc=MVK | 3008642652-2025-02923 | 2025-02-20 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-03-20 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Signal Artifact/Noise | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE DEVICE WAS STARTED UP AT 08:44:42 ON (B)(6) 2025. AT 07:15:42 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS IDIOVENTRICULAR RHYTHM @ 20 BPM THAT DEGRADES TO VT/VF WITH LOW/VARYING AMPLITUDE. THE DEVICE PROPERLY DETECTED VT/VF. LOW/VARYING AMPLITUDES, CPR/MOTION ARTIFACT AND ELECTRODE LEAD FALLOFF PREVENTED THE LIFEVEST FROM TREATING THE PATIENT. THE ELECTRODE BELT WAS DISCONNECTED AT 09:03:08 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21652054&pc=OZD | 1220648-2025-26844 | 2025-02-22 05:00 | Death | ABIOMED, INC. | 2025-03-20 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Hemolysis | P140003 |  | 1 | Event Description: THE COMPLAINANT HAD A IMPELLA 5.5 PUMP PLACED AS CARE ESCALATION FROM THE IMPELLA CP. IT WAS REPORTED THAT THE PATIENTS HAD ON AND OFF CONCENTRATED AMBER/RED URINE OUTPUT. IT WAS FURTHER REPORTED THAT CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: H1 TYPE OF REPORT ERRONEOUSLY SELECTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26844. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21652055&pc=PYX | 1220648-2025-26843 | 2025-02-21 05:00 | Death | ABIOMED, INC. | 2025-03-20 04:00 | PYX | IMPELLA RP FLEX | Patient-Device Incompatibility |  | P170011 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT (POST CARDIOTHORACIC SURGERY) WAS ON THE UNIT, DECOMPENSATED AND WAS BROUGHT TO THE CATHETERIZATION LAB FOR AN IMPELLA RP FLEX IMPLANT. THE RIGHT VENTRICLE DECOMPENSATION WAS NOTED ON ECHOCARDIOGRAM. THE ARTERIAL LINE WAS PLACED WITH MEAN ARTERIAL PRESSURE IN THE 10'S. CARDIOPULMONARY RESUSCITATION WAS INITIATED. VENOUS ACCESS WAS OBTAINED. THE IMPELLA RP FLEX WAS PRIMED AND CONNECTED. HOWEVER, INSERTION WAS ABORTED DUE TO PATIENT ANATOMY REASON. THE PATIENT WAS PRONOUNCED AS DECEASED. NO FURTHER DETAILS SURROUNDING THE EVENT WERE NOTED, AND THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA RP FLEX AS AN ASSOCIATED FACTOR. Manufacturer Narrative: THE INVESTIGATION OF DELIVERY ISSUES HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. BASED ON THE CLINICAL DETAILS THE ROOT CAUSE OF DELIVERY ISSUE IS DETERMINED TO BE PATIENT CONDITION BECAUSE THE PUMP WAS UNABLE TO PASS THROUGH RIGHT VENTRICLE AND LACK OF FORWARD FLOW. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA RP FLEX WITH SMART ASSIST SYSTEM WITH AUTOMATED IMPELLA CONTROLLER. SECTION: WARNINGS & CAUTIONS: WARNINGS, SECTION: PRE-SUPPORT EVALUATION, SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER, SECTION: DIRECT AORTIC INSERTION, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: ¿TO REDUCE THE RISK OF CARDIAC INJURY (INCLUDING VENTRICULAR PERFORATION), PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY IN THE SETTINGS OF ACUTE INFARCTION WITH TISSUE NECROSIS.¿ ¿TO REDUCE THE RISK OF VASCULAR INJURY, PHYSICIANS SHOULD EXERCISE CAUTION WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX PERIPHERAL VASCULAR ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: UNREPAIRED ABDOMINAL AORTIC ANEURYSM, SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM, DISSECTION OF THE ASCENDING/ TRANSVERSE/DESCENDING AORTA, CHRONIC ANATOMICAL CHANGES IN THE RELATIONSHIP OF THE AORTA/AORTIC VALVE/VENTRICULAR ALIGNMENT, SIGNIFICANT MOBILE ATHEROMATOUS DISEASE IN THE THORACIC OR ABDOMINAL AORTA OR PERIPHERAL VESSELS.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING A RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY IN THE LEFT VENTRICLE AFTER CPR WITH ECHOCARDIOGRAPHY GUIDANCE. SECTION: WARNINGS & CAUTIONS: WARNINGS, SECTION: PRE-SUPPORT EVALUATION, SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER, SECTION: DIRECT AORTIC INSERTION, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ARRHYTHMIA, ATRIAL FIBRILLATION, BLEEDING, CARDIAC TAMPONADE, CARDIOGENIC SHOCK, DEATH, DEVICE MALFUNCTION, HEMOLYSIS, HEPATIC FAILURE, INSERTION SITE INFECTION, PERFORATION, PHLEGMASIA CERULEA DOLENS (A SEVERE FORM OF DEEP VENOUS THROMBOSIS), PULMONARY VALVE INSUFFICIENCY, RESPIRATORY DYSFUNCTION, SEPSIS, THROMBOCYTOPENIA, THROMBOTIC VASCULAR (NON-CENTRAL NERVOUS SYSTEM COMPLICATION, TRICUSPID VALVE INJURY, CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION), VENOUS THROMBOSIS, VENTRICULAR FIBRILLATION AND/OR TACHYCARDIA.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿DO NOT ADVANCE OR WITHDRAW THE IMPELLA RP FLEX WITH SMARTASSIST SYSTEM CATHETER AGAINST RESISTANCE WITHOUT USING FLUOROSCOPY TO DETERMINE THE CAUSE OF THE RESISTANCE. DOING SO COULD RESULT IN SEPARATION OF THE CATHETER OR GUIDEWIRE TIP, DAMAGE TO THE CATHETER OR VESSEL, OR PERFORATION.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21652858&pc=DTB | MW5167846 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-03-20 04:00 | DTB | PERMANENT PACEMAKER ELECTRODE |  | Insufficient Information |  |  | 1 | Event Description: PATIENT'S SON IN LAW CALLED TO ASK FOR INFORMATION ON THE DEFIBRILLATOR LEADING UP TO A PATIENT'S DEATH. HE SAID THEY GOT A CALL THAT SAID "THE LEAD LOOKS OFF HOW IS THE PATIENT DOING?" DATE OF DEATH: (B)(6) 2025. THIS REPORT REFLECTS INFORMATION RECEIVED BY FDA IN THE FORM OF A NOTIFICATION PER 803.22 (B)(2).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21652881&pc=NVY | MW5167847 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-03-20 04:00 | NVY | PERMANENT DEFIBRILLATOR ELECTRODES |  | Insufficient Information |  |  | 1 | Event Description: PATIENT'S SON IN LAW CALLED TO ASK FOR INFORMATION ON THE DEFIBRILLATOR LEADING UP TO A PATIENT'S DEATH. HE SAID THEY GOT A CALL THAT SAID "THE LEAD LOOKS OFF HOW IS THE PATIENT DOING?". DATE OF DEATH: (B)(6) 2025. THIS REPORT REFLECTS INFORMATION RECEIVED BY FDA IN THE FORM OF A NOTIFICATION PER 803.22 (B)(2).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21653196&pc=DSQ | 2916596-2025-01637 | 2025-02-28 05:00 | Death | THORATEC CORPORATION | 2025-03-20 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. THE ACCOUNT DECLINED TO PROVIDE PATIENT OUTCOME INFORMATION FOR THIS EVENT, DUE TO PATIENT PRIVACY LAWS. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND THE REPORTED PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. NO AUTOPSY WAS PERFORMED, AND THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. DUE TO PATIENT PRIVACY LAWS, THE MANAGING HOSPITAL WOULD NOT COMMUNICATE ANY ADDITIONAL INFORMATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. G IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21653281&pc=DSQ | 3007042319-2025-00669 | 2025-03-14 04:00 | Death | HEARTWARE, INC. | 2025-03-20 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Sepsis | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WITH A VENTRICULAR ASSIST DEVICE (VAD) WAS ON HOSPICE CARE AND DEVELOPED SEPSIS FROM A DRIVELINE INFECTION. THE PATIENT PROGRESSIVELY WORSENED, AND THE DEVICE WAS TURNED OFF BY A HOSPICE NURSE, AS INSTRUCTED BY THE IMPLANTING CENTER. THE PATIENT PASSED AWAY AT HOME. IT IS UNKNOWN IF THERE WERE ANY PATIENT-RELATED COMPLICATIONS OR IF THE EVENT CONTRIBUTED TO THE PATIENT'S DEATH. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: THE VENTRICULAR ASSIST DEVICE (VAD) (B)(6) WAS NOT RETURNED FOR EVALUATION AS IT REMAINS IN USE. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE MANUFACTURING DOCUMENTATION WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED ON HOSPICE CARE AFTER DEVELOPING SEPSIS FROM A DRIVELINE INFECTION AND PROGRESSIVE WORSENING. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, D RIVELINE INFECTION, SEPSIS AND DEATH ARE KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. THERE WAS NO EVIDENCE THAT THE PATIENT HAD A HISTORY OF SIMILAR ADVERSE EVENTS. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, PATIENT'S COMPLEX POST-OPERATIVE COURSE, INCLUDING CARE OF THE DRIVELINE EXIT SITE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21653536&pc=DTZ | 3013876692-2025-00019 | 2025-02-16 05:00 | Death | MAQUET CARDIOPULMONARY GMBH | 2025-03-20 04:00 | DTZ | TUBING SET | Failure to Prime; Incomplete or Inadequate Priming; Air/Gas in Device | Cardiac Arrest | K112360 |  | 1 | Event Description: ON MARCH 18, 2025 AN INITIAL MDR REPORT (8010762-2025-0000119) FOR A PRODUCT PROBLEM WAS SUBMITTED BY GETINGE MANUFACTURER- STATING "THE EVENT OCCURRED IN THE USA DURING PRIMING. FOLLOWING WAS REPORTED IN THE FDA IMPORTER REPORT # (B)(4). CVL (CARDIOVASCULAR LAB) STEMI (ST-SEGMENT ELEVATION MYOCARDIAL INFARCTION) TEAM CALLED IN TO HAVE ECMO (EXTRACORPOREAL MEMBRANE OXYGENATION) CANNULATION AND INITIATION STARTED ON A CCU (CORONARY CARE UNIT) PATIENT. THE PATIENT WAS BROUGHT DOWN TO CATHETER LAB AND ECLS (EXTRACORPOREAL LIFE SUPPORT) TEAM STARTED SETTING UP ECMO DEVICE WHILE THE PROVIDER AND CVL STAFF BEGAN GETTING ACCESSES FOR SHEATH CANNULATION. SHEATH CANNULATION WAS COMPLETED. THE ECMO DEVICE WAS FILLING AND CIRCULATING FOAM AND WAS UNABLE TO BE CONNECTED TO PATIENT. ECLS TEAM NEEDED TO GO SET UP A NEW KIT AND PUMP UPSTAIRS. THE PATIENT WAS CONNECTED TO A NEW ECMO SET 1 HOUR AFTER CATHETER LAB TEAM ACHIEVED CANNULATION OF BILATERAL FEMORAL AND RETROGRADE ACCESS. ADDITIONAL 15 MINUTES FROM THE TIME THE PATIENT WAS CONNECTED TO THE PUMP TO THE PATIENT RECEIVING OXYGENATED BLOOD WERE NEEDED. NO HARM TO ANY PERSON HAS BEEN REPORTED. DUE TO THE DELAY IN TREATMENT OF THE PATIENT, A REPORT IS REQUIRED. COMPLAINT ID# (B)(4). ADDITIONAL INFORMATION WAS RECEIVED ON 2025-03-18/19 THAT THE PATIENT PASSED AWAY 24 HOURS AFTER ON THE SECOND HLS SET CONNECTED DUE TO THE FAMILY DECIDED TO GIVE UP TREATMENT. THE PATIENT WAS A STEMI THAT CODED MULTIPLE TIMES. THE PHYSICIAN INITIATED THE ECMO THERAPY AFTER THE PATIENT WAS REVIVED FROM MULTIPLE CODES AND WAS UNSTABLE. IT WAS STATED THAT THE FIRST USED SET WAS NEVER CONNECTED TO THE PATIENT AND THE CUSTOMER STATED THAT THE MOST PROBABLE ROOT CAUSE WITH THE FIRST HLS SET WAS AN IN-EXPERIENCED NEW ECMO- SPECIALIST WHO PRIMED THE CIRCUIT AND DID NOT DE AIR THE CIRCUIT CORRECTLY CAUSING THE FOAM. THE SECOND HLS SET WHICH WAS THAN USED ON THE PATIENT HAD NO FAILURE. THE PATIENT PASSED AWAY ON (B)(6) 2025. BOTH SETS ARE NOT AVAILABLE FOR INVESTIGATION, AS THEY WERE DISCARDED BY THE CUSTOMER. FURTHER THE CUSTOMER CONFIRMED THAT THE PATIENT PASSED AWAY FROM HIS PRE-EXISTING CONDITIONS AND NOT DUE TO THE EVENT WITH THE FIRST HLS SET, WHICH CAUSED A DELAY IN TREATMENT.Event Description: COMPLAINT ID# (B)(4).Event Description: COMPLAINT ID# (B)(4). Manufacturer Narrative: RECEIVED A REQUEST FROM THE FDA ON APRIL 7, 2025 STATING-"THIS IS TO BRING TO YOUR ATTENTION THAT WE ARE IN RECEIPT OF REPORT NUMBER 3013876692-2025-00019-1 WITH THE REPORT DATE 03/20/2025. UPON REVIEW OF THE INFORMATION INCLUDED FOR THE REPORT, IT REFERENCES RELATED MEDICAL DEVICE REPORT (MDR) (B)(4) IDENTIFIED BY YOUR FIRM FOR THE SAME EVENT. HOWEVER, THIS INFORMATION WAS NOT INCLUDED IN THE CORRESPONDING BLOCK H10, RELATED REPORT NUMBER FIELD, OF MEDWATCH FORM 3500A." THIS SUBMISSION SERVES TO ADDRESS THIS ERROR NOTED ABOVE. Manufacturer Narrative: THE EVENT OCCURRED IN THE USA DURING PRIMING. FOLLOWING WAS REPORTED IN THE IMPORTER REPORT (B)(4). CVL (CARDIOVASCULAR LAB) STEMI (ST-SEGMENT ELEVATION MYOCARDIAL INFARCTION) TEAM CALLED IN TO HAVE ECMO (EXTRACORPOREAL MEMBRANE OXYGENATION) CANNULATION AND INITIATION STARTED ON A CCU (CORONARY CARE UNIT) PATIENT. THE PATIENT WAS BROUGHT DOWN TO CATHETER LAB AND ECLS (EXTRACORPOREAL LIFE SUPPORT) TEAM STARTED SETTING UP ECMO DEVICE WHILE THE PROVIDER AND CVL STAFF BEGAN GETTING ACCESSES FOR SHEATH CANNULATION. SHEATH CANNULATION WAS COMPLETED. THE ECMO DEVICE WAS FILLING AND CIRCULATING FOAM AND WAS UNABLE TO BE CONNECTED TO PATIENT. ECLS TEAM NEEDED TO GO SET UP A NEW KIT AND PUMP UPSTAIRS. THE PATIENT WAS CONNECTED TO A NEW ECMO SET 1 HOUR AFTER CATHETER LAB TEAM ACHIEVED CANNULATION OF BILATERAL FEMORAL AND RETROGRADE ACCESS. ADDITIONAL 15 MINUTES FROM THE TIME THE PATIENT WAS CONNECTED TO THE PUMP TO THE PATIENT RECEIVING OXYGENATED BLOOD WERE NEEDED. ADDITIONAL INFORMATION WAS RECEIVED ON 2025-03-18/19 THAT THE PATIENT PASSED AWAY 24 HOURS AFTER HE WAS ON THE SECOND HLS SET CONNECTED DUE TO THE FAMILY DECIDED TO GIVE UP TREATMENT. THE PATIENT WAS A STEMI THAT CODED MULTIPLE TIMES. THE PHYSICIAN INITIATED THE ECMO THERAPY AFTER THE PATIENT WAS REVIVED FROM MULTIPLE CODES AND WAS UNSTABLE. IT WAS STATED THAT THE FIRST USED SET WAS NEVER CONNECTED TO THE PATIENT AND THE CUSTOMER STATED THAT THE MOST PROBABLE ROOT CAUSE WITH THE FIRST HLS SET WAS AN IN-EXPERIENCED NEW ECMO- SPECIALIST WHO PRIMED THE CIRCUIT AND DID NOT DE AIR THE CIRCUIT CORRECTLY CAUSING THE FOAM. THE SECOND HLS SET WHICH WAS THAN USED ON THE PATIENT HAD NO FAILURE. THE PATIENT PASSED AWAY ON (B)(6) 2025. BOTH SETS ARE NOT AVAILABLE FOR INVESTIGATION, AS THEY WERE DISCARDED BY THE CUSTOMER. FURTHER THE CUSTOMER CONFIRMED THAT THE PATIENT PASSED AWAY FROM HIS PRE-EXISTING CONDITIONS AND NOT FROM THE FAILURE OF THE FIRST HLS SET, WHICH CAUSED A DELAY IN TREATMENT. A MEDICAL REVIEW WAS PERFORMED BY GETINGE MEDICAL AFFAIRS ON (B)(6) 2025 WITH FOLLOWING CONCLUSION: "ACCORDING TO THE COMPLAINT NARRATIVE, A STEMI (ST-SEGMENT ELEVATION MYOCARDIAL INFARCTION) TEAM WAS CALLED TO INITIATE ECMO (EXTRACORPOREAL MEMBRANE OXYGENATION) CANNULATION FOR A PATIENT IN THE CORONARY CARE UNIT (CCU). THE PATIENT WAS TRANSFERRED TO THE CATH LAB, WHERE THE EXTRACORPOREAL LIFE SUPPORT (ECLS) TEAM SET UP THE ECMO DEVICE WHILE THE PROVIDER AND CVL STAFF PERFORMED VASCULAR ACCESS FOR SHEATH CANNULATION. CANNULATION OF BILATERAL FEMORAL AND RETROGRADE ACCESS WAS COMPLETED SUCCESSFULLY. DURING PRIMING, THE ECMO CIRCUIT EXHIBITED EXCESSIVE FOAM, PREVENTING PATIENT CONNECTION AND DELAYING EXTRACORPOREAL PERFUSION. WITH NO BACKUP SET ON-SITE, THE ECLS TEAM RETRIEVED A REPLACEMENT CIRCUIT AND PUMP FROM ANOTHER LOCATION, DELAYING TREATMENT. THE PATIENT WAS CONNECTED TO THE NEW ECMO SET ONE HOUR AFTER CANNULATION, WITH AN ADDITIONAL 15-MINUTE DELAY BEFORE OXYGENATED BLOOD CIRCULATION BEGAN. THE INTENDED PROCEDURE WAS TO PRIME THE ECMO CIRCUIT WITHOUT VISIBLE BUBBLES FOR IMMEDIATE PATIENT CONNECTION AND SUPPORT. HOWEVER, FOAMING COMPROMISED THE CIRCUIT, DELAYING PATIENT CONNECTION BY OVER AN HOUR. THE LACK OF A BACKUP SET FURTHER EXTENDED THIS DELAY, POTENTIALLY IMPACTING PATIENT STABILITY. A STUDY PUBLISHED IN ACS OMEGA EXAMINED THE STABILITY OF LIQUID FILMS FORMED BY BUBBLES IN BOVINE SERUM ALBUMIN SOLUTIONS. THE RESEARCH HIGHLIGHTED THAT THE AMPHIPATHIC CHARACTERISTICS OF ALBUMIN CONTRIBUTE TO THE FORMATION AND STABILIZATION OF FOAM FILMS, EMPHASIZING THE ROLE OF ALBUMIN AT THE AIR-LIQUID INTERFACE. THIS SUGGESTS THAT THE EXCESSIVE FOAM OBSERVED DURING PRIMING MAY HAVE BEEN DUE TO THE DIRECT INCLUSION OF ALBUMIN INTO THE PRIMING SOLUTION, LEADING TO INCREASED BUBBLE FORMATION. THE PATIENT, A STEMI CASE, HAD SUFFERED MULTIPLE CARDIAC ARRESTS BEFORE ECMO INITIATION. AFTER RESUSCITATION AND STABILIZATION, THE PHYSICIAN INITIATED CARDIO-HELP THERAPY DUE TO CONTINUED INSTABILITY. THE PATIENT¿S FAMILY OPTED TO DISCONTINUE SUPPORT AFTER 24 HOURS. THE ATTENDING PHYSICIAN CONFIRMED THAT THE PATIENT¿S DEATH RESULTED FROM PRE-EXISTING CARDIAC CONDITIONS RATHER THAN EQUIPMENT FAILURE. ADDITIONALLY, FEEDBACK FROM THE ECMO MANAGER SUGGESTED THAT THE INITIAL CIRCUIT PRIMING ISSUE MAY HAVE RESULTED FROM IMPROPER DE-AIRING. FURTHERMORE, IT WAS CONFIRMED THAT THE PATIENT EXPIRED DUE TO PRE-EXISTING CONDITIONS RATHER THAN ANY FAILURE OF THE ECMO SYSTEM. IN CONCLUSION, THERE WAS NO EVIDENCE OF DEVICE MALFUNCTION IN THIS INCIDENT. HOWEVER, THE DELAY IN INITIATING EXTRACORPOREAL PERFUSION DUE TO INITIAL PRIMING ISSUES MAY HAVE CONTRIBUTED TO THE OVERALL PATIENT OUTCOME. THE EXCESSIVE FOAM OBSERVED DURING PRIMING, WHICH PREVENTED IMMEDIATE PATIENT CONNECTION, MAY HAVE BEEN INFLUENCED BY THE INCLUSION OF ALBUMIN IN THE PRIMING SOLUTION. RESEARCH HAS SHOWN THAT ALBUMIN'S AMPHIPATHIC PROPERTIES CONTRIBUTE TO FOAM STABILIZATION, POTENTIALLY LEADING TO INCREASED BUBBLE FORMATION. AS BOTH ECMO SETS INVOLVED IN THIS CASE WERE DISCARDED BY THE CUSTOMER, NO FURTHER INVESTIGATION CAN BE CONDUCTED, AND A DEFINITIVE ROOT CAUSE CANNOT BE DETERMINED." ACCORDING TO THE INSTRUCTION OF USE OF THE HLS SET (CHAPTER "SAFETY INSTRUCTIONS FOR THE OXYGENATOR") IT IS STATED THAT THE DE-AIRING MEMBRANE MUST BE OPEN THROUGHOUT PRIMING TO ENSURE SAFE DE-AIRING OF THE OXYGENATOR AND TO ENSURE THAT THE DE-AIRING MEMBRANE OF THE OXYGENATOR IS CLOSED WITH THE YELLOW PROTECTIVE CAP DURING OPERATION. THE FLOW/BUBBLE SENSOR MUST ALWAYS BE AFFIXED ON THE ARTERIAL SIDE OF THE SET IF YOU ARE USING THE CARDIOHELP-I. WITH AN ACTIVATED INTERVENTION, THE DETECTION OF BUBBLES BY THE FLOW/BUBBLE SENSOR TRIGGERS A PUMP STOP. IN CHAPTER "PRIMING THE SYSTEM" IT IS MENTIONED THAT PRE-PRIMING CAN CAUSE AIR BUBBLES TO FORM IN THE SYSTEM. REFERRING TO THE CHAPTER "START PERFUSION" THE CUSTOMER SHOULD ENSURE THAT THE TUBE SYSTEM HAS BEEN COMPLETELY DE-AIRED BEFORE STARTING THE PERFUSION. IN ORDER TO AVOID REOCCURRENCE OF THE REPORTED FAILURE, THE CUSTOMER WILL BE INFORMED BY THE GETINGE SALES AND SERVICE UNIT (SSU) TO FOLLOW THE CHAPTER PRIMING THE SYSTEM AND TO RE-TRAIN THE NEW MEDICAL STUFF. THE PRODUCTION RECORDS OF THE AFFECTED PRODUCT WERE REVIEWED ON (B)(6) 2025. ACCORDING TO THE FINAL TEST RESULTS, THE PRODUCT PASSED THE TESTS AS PER SPECIFICATIONS. PRODUCTION RELATED INFLUENCES ARE UNLIKELY. THE REVIEW OF SCRAP, REWORK, ENHANCEMENTS AND DESIGN CHANGES WERE REVIEWED AN NO ABNORMALITIES IN REGARD TO THE REPORTED FAILURE WERE FOUND. IN ADDITION, A REVIEW FOR POTENTIAL FIELD ACTIONS AND CAPAS RELATED TO THE FAILURE AND PRODUCT IN THIS COMPLAINT WAS PERFORMED. THIS COMPLAINT IS NOT IN SCOPE OF ANY ONGOING FIELD ACTIONS AND/OR CAPAS. THE REVIEW OF THE NON-CONFORMITIES HAS BEEN PERFORMED FOR THE REVIEWED TIME PERIOD. IT DOES NOT SHOW ANY NON-CONFORMITY IN REGARD TO THE REPORTED FAILURE. BASED ON THE RESULTS THE REPORTED FAILURE "AIR IN PRODUCT" COULD BE CONFIRMED BUT WAS NOT A PRODUCT RELATED MALFUNCTION. THE CUSTOMER WILL BE INFORMED ABOUT THE RESULTS BY THE GETINGE SALES AND SERVICE UNIT. THE OCCURRENCE RATE WAS CALCULATED FOR THE REPORTED ISSUE AND IT WAS DETERMINED THAT THIS IS NOT A SYSTEMIC ISSUE. THEREFORE, NO REMEDIAL ACTION IS REQUIRED. THE OCCURRENCE RATE RELATED TO THE REPORTED ISSUE IS CURRENTLY BEING MONITORED AS PART OF MAQUET CARDIOPULMONARY¿ S TRENDING PROGRAM AND ADDITIONAL INVESTIGATIONS OR CORRECTIONS WILL BE IMPLEMENTED IN CASE OF ADVERSE TRENDING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21653762&pc=DSQ | 2916596-2025-01972 | 2025-02-12 05:00 | Death | THORATEC CORPORATION | 2025-03-20 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Thrombosis/Thrombus | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT RETURNED TO OPERATION ROOM FOR TUMOR RESECTION FROM KIDNEY. THE PATIENT ALSO HAD A THROMBUS IN VENA CAVA INFERIOR WHICH WAS PLANNED TO BE REMOVED. WHEN REMOVING THROMBUS, THROMBUS MOBILIZED AND GOT STUCK IN PULMONARY ARTERY WHICH LED TO PATIENT EXPIRED IN OPERATION ROOM. PUMP PERFORMED AS INTENDED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS COULD NOT BE CONCLUSIVELY ESTABLISHED. IT WAS COMMUNICATED THAT THE PUMP OPERATED AS INTENDED AND WOULD NOT BE RETURNED FOR EVALUATION. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. B, IS CURRENTLY AVAILABLE. THE IFU LISTS POTENTIAL ADVERSE EVENTS, INCLUDING THROMBOEMBOLISM AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. THE IFU ALSO OUTLINES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING THE INTERNATIONAL NORMALIZED RATIO RANGE FOR PATIENTS USING THE HEARTMATE 3 LVAS, AS WELL AS THE SUGGESTED ANTICOAGULATION MODIFICATIONS. THIS SECTION ALSO LISTS THROMBOEMBOLISM AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21653878&pc=OBJ | 2124215-2025-17036 | 2025-03-17 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-03-20 04:00 | OBJ | OPTICROSS | Entrapment of Device; Adverse Event Without Identified Device or Use Problem | Ischemia |  |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT DEATH HAD OCCURRED. AN OPTICROSS IMAGING CATHETER WAS ADVANCED FOR INTRAVASCULAR ULTRASOUND (IVUS) EXAMINATION OF THE TARGET LESION. DURING THE PROCEDURE, THE CATHETER BECAME TANGLED WITH STENT STRUTS OR THE GUIDEWIRE AND WAS TRAPPED. IT TOOK MORE THAN THREE HOURS TO REMOVE AND COMPLETE THE PROCEDURE. THE PATIENT PASSED AWAY 1.5 HOURS LATER, PROBABLY DUE TO ISCHEMIA DURING THE PROCEDURE.Event Description: IT WAS REPORTED THAT PATIENT DEATH HAD OCCURRED. AN OPTICROSS IMAGING CATHETER WAS ADVANCED FOR INTRAVASCULAR ULTRASOUND (IVUS) EXAMINATION OF THE TARGET LESION. DURING THE PROCEDURE, THE CATHETER BECAME TANGLED WITH STENT STRUTS OR THE GUIDEWIRE AND WAS TRAPPED. IT TOOK MORE THAN THREE HOURS TO REMOVE AND COMPLETE THE PROCEDURE. THE PATIENT PASSED AWAY 1.5 HOURS LATER, PROBABLY DUE TO ISCHEMIA DURING THE PROCEDURE. IT WAS FURTHER REPORTED THAT 5 STENTS WERE USED IN THE PROCEDURE. THE STENT INVOLVED WITH THE OPTICROSS ENTANGLEMENT WAS PROPERLY SEATED BUT WAS NOT POST DILATED. ANOTHER OPTICROSS WAS USED AFTER RETRIEVING THE TANGLED DEVICES. AN AUTOPSY WAS NOT PERFORMED.Event Description: IT WAS REPORTED THAT PATIENT DEATH HAD OCCURRED. AN OPTICROSS IMAGING CATHETER WAS ADVANCED FOR INTRAVASCULAR ULTRASOUND (IVUS) EXAMINATION OF THE TARGET LESION. DURING THE PROCEDURE, THE CATHETER BECAME TANGLED WITH STENT STRUTS OR THE GUIDEWIRE AND WAS TRAPPED. IT TOOK MORE THAN THREE HOURS TO REMOVE AND COMPLETE THE PROCEDURE. THE PATIENT PASSED AWAY 1.5 HOURS LATER, PROBABLY DUE TO ISCHEMIA DURING THE PROCEDURE. IT WAS FURTHER REPORTED THAT 5 STENTS WERE USED IN THE PROCEDURE. THE STENT INVOLVED WITH THE OPTICROSS ENTANGLEMENT WAS PROPERLY SEATED BUT WAS NOT POST DILATED. ANOTHER OPTICROSS WAS USED AFTER RETRIEVING THE TANGLED DEVICES. AN AUTOPSY WAS NOT PERFORMED. Manufacturer Narrative: E1: INITIAL REPORTER ADDRESS: (B)(6). Manufacturer Narrative: E1: INITIAL REPORTER ADDRESS: (B)(6). Manufacturer Narrative: E1: INITIAL REPORTER ADDRESS 1 - (B)(6). THE DEVICE WAS RETURNED FOR ANALYSIS. VISUAL AND MICROSCOPIC INSPECTION REVEALED THE SHEATH DETACHED FROM THE TELESCOPE SECTION. THE DETACHED SECTION OF THE SHEATH, INCLUDING THE IMAGING WINDOW AND TIP, WERE NOT RETURNED FOR ANALYSIS. THE DRIVESHAFT WAS FOUND BROKEN AND CUT WITH A WINDUP AND A SECTION TWISTED. THE DEVICE WAS RETURNED WITH 2 NON-BOSTON SCIENTIFIC GUIDEWIRES, ONE WITH ADDED MATERIAL. MICROSCOPE INSPECTION REVEALED THE DRIVE CABLE WAS ALSO CUT. MEDIA PROVIDED BY THE CUSTOMER WAS INSPECTED. TWO PICTURES SHOWED THE DRIVE CABLE (DRIVESHAFT) BROKEN BUT WERE NOT CLEAR ENOUGH TO IDENTIFY ANY OTHER FAILURES. ONE PICTURE SHOWED THE DRIVE CABLE WITH A SECTION TWISTED. ANOTHER PICTURE SHOWED THE DISTAL SECTION OF THE IMAGING WINDOW DETACHED WITH A BEND, BUT THE PICTURE IS NOT CLEAR ENOUGH TO IDENTIFY TIP DAMAGE. TWO PICTURES CAPTURED FROM A MOBILE PHONE WERE PROVIDED OF STILL ANGIOGRAPHIC FRAMES. THE FIRST STILL IMAGE SHOWED THE GUIDE CATHETER IN THE AORTA, OUTSIDE THE LEFT MAIN STEM, THE OPTICROSS CATHETER, AND WHAT APPEARS TO BE A SNARE CATHETER. IT ALSO SHOWED WHAT APPEARS TO POSSIBLY BE A KINKED CATHETER RIGHT OUTSIDE THE LEFT MAIN STEM. THE SECOND PICTURE SHOWED WHAT SEEMS TO BE THE SAME STILL FRAME VISIBLE IN THE FIRST PICTURE AND SHOWED THE OPTICROSS CATHETER INSIDE THE STENT(S) IN THE LAD. ONLY THE FRAME OF THE STENTS IS VISIBLE. THE STENT(S) SEEM TO BE UNDEREXPANDED IN THE MID-DISTAL PART OF THE ARTERY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21653881&pc=NPT | 2015691-2025-02096 | 2025-02-24 05:00 | Death | EDWARDS LIFESCIENCES | 2025-03-20 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Rupture; Pericardial Effusion | P140031 |  | 1 | Event Description: AS REPORTED BY OUR AFFILIATES IN FRANCE, DURING A TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) PROCEDURE USING A 26MM SAPIEN 3 ULTRA VALVE VIA A TRANSFEMORAL APPROACH, A LOSS OF PACING CAPTURE OCCURRED AT THE END OF THE DEPLOYMENT WHEN THE VALVE WAS FULLY DEPLOYED. THIS CAUSED THE VALVE TO EMBOLIZE ABOVE THE ANNULUS, AT THE LEVEL OF THE SINUSES OF VALSALVA. A SECOND VALVE WAS THEN SUCCESSFULLY DEPLOYED AS A VALVE-IN-VALVE. UNFORTUNATELY, ONE DAY AFTER THE PROCEDURE, THE PATIENT DEVELOPED A PERICARDIAL EFFUSION AND PASSED AWAY. IT WAS BELIEVED THAT THE INITIAL VALVE CAUSED A RUPTURE IN THE AORTA. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), VALVE EMBOLIZATION IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. THERE ARE MULTIPLE PATIENT AND PROCEDURAL FACTORS THAT ALONE OR IN COMBINATION CAN CAUSE OR CONTRIBUTE TO EMBOLIZATION OF THE DEVICE, INCLUDING IMPROPER POSITIONING BEFORE DEPLOYMENT, POOR IMAGE INTENSIFIER ANGLE, POOR COAXIAL ALIGNMENT OF THE VALVE AND DELIVERY SYSTEM, SEVERE SEPTAL HYPERTROPHY, MINIMALLY OR BULKY/SEVERELY CALCIFIED LEAFLETS, LOSS OF PACING CAPTURE, A NARROW SINOTUBULAR JUNCTION (IN AORTIC POSITION PROCEDURES), RAPID DEPLOYMENT, THE RELEASE OF STORED TENSION DURING DEPLOYMENT, INACCURATE MEASUREMENT OF THE LANDING ZONE, AND MOVEMENT OF THE DELIVERY SYSTEM BY THE OPERATOR. THE IFU CAUTIONS THAT INCORRECT SIZING OF THE LANDING ZONE MAY LED TO PARAVALVULAR LEAK, MIGRATION, OR EMBOLIZATION. THE DEVICE TRAINING MANUALS INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVES (ALL MODELS). TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES. THE CORRECT ALIGNMENT AND POSITIONING OF THE DEVICE AT THE POINT OF DEPLOYMENT IS EMPHASIZED AS A KEY FACTOR IN THE PLACEMENT AND FIXATION OF THE DEVICE. OPERATORS ARE ALSO INSTRUCTED TO USE FLUOROSCOPY AS THE PRIMARY METHOD OF VISUALIZATION FOR POSITIONING AND DEPLOYMENT. IN PATIENTS WITH HIGH-RISK ANATOMICAL FEATURES FOR EMBOLIZATION (I.E., MINIMAL LEAFLET CALCIFICATION, SEVERE SEPTAL HYPERTROPHY), BALLOON VALVULOPLASTY MAY INDICATE POTENTIAL BALLOON MOVEMENT DURING VALVE DEPLOYMENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST/INDICATE THAT PATIENT FACTORS (LOSS OF PACING AND RUPTURE) CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21654024&pc=NPT | 2025587-2025-01998 | 2024-06-01 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-20 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Insufficient Information | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 10 YEARS, 9 MONTHS, AND 3 WEEKS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT PASSED AWAY SUDDENLY. THE CAUSE OF DEATH IS UNKNOWN. THE RELATIONSHIP BETWEEN THE DEATH AND THE DEVICE OR PROCEDURE IS UNKNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: D4., H4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21654377&pc=DQY | 2135147-2025-01440 | 2025-02-26 05:00 | Death | ABBOTT MEDICAL | 2025-03-20 04:00 | DQY | AMPLATZER TORQVUE 45X45 DELIVERY SHEATH | Difficult to Advance | Cardiac Arrest; Hematoma; Hemorrhage/Bleeding; Low Blood Pressure/ Hypotension | K163000 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, AN AMPLATZER AMULET LEFT ATRIAL APPENDAGE (LAA) OCCLUDER WAS CHOSEN FOR IMPLANTATION UTILIZING A 14F TORQVUE 45X45 DELIVERY SHEATH. DURING ADVANCEMENT OF THE DELIVERY SYSTEM, RESISTANCE WAS ENCOUNTERED AND ADVANCEMENT WAS STOPPED. ON FLUOROSCOPY, THE SHEATH WAS WRAPPED AROUND ITSELF AND KINKED SO THE DELIVERY SYSTEM WAS REMOVED AND REPLACED. A 18F GROIN SHEATH WAS PLACED AND THE REPLACEMENT 14F TORQVUE 45X45 DELIVERY SYSTEM WAS ADVANCED. THE DELIVERY SYSTEM EXPERIENCED RESISTANCE WHEN TRYING TO CROSS THE SEPTUM. THE DELIVERY SYSTEM WAS REMOVED AND THE SEPTUM WAS DILATED. THE DELIVERY SYSTEM WAS THEN ABLE TO CROSS THE SEPTUM AND ADVANCED TO THE LAA. BLOOD PRESSURE BECAME UNSTABLE SO THE CASE WAS ABORTED. NO EFFUSION OR GROIN COMPLICATIONS WERE OBSERVED. ONCE PATIENT WAS BROUGHT TO THE RECOVERY AREA BLOOD PRESSURE DROPPED. A HEMATOMA WAS NOTED IN GROIN AREA. MANUAL COMPRESSION WAS PERFORMED, AND THE PATIENT CODED AND DIED. IT IS BELIEVED THAT THE PATIENT HAD A BLEED AT THE GROIN SITE AND ENTERED CARDIOGENIC SHOCK DUE TO A LOW EJECTION FRACTION. THE CAUSE OF THE BLEED IS UNKNOWN AS AN ARTERIAL LINE HAD ALSO BEEN INSERTED. Manufacturer Narrative: THE DEVICE IS NOT RETURNING FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. AN EVENT OF DIFFICULT TO ADVANCE THROUGH PATIENT ANATOMY WAS REPORTED. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED AND THE PRODUCT MET ALL SPECIFICATIONS. BASED ON THE AVAILABLE INFORMATION, THE CAUSE OF THE REPORTED DIFFICULT TO ADVANCE THROUGH PATIENT ANATOMY COULD NOT BE DETERMINED. THE PATIENT EFFECTS OF CARDIAC ARREST, HEMORRHAGE, HEMATOMA, AND HYPOTENSION COULD NOT DETERMINED. DEATH WAS RELATED TO PATIENT'S COMORBIDITIES. THE REPORTED UNEXPECTED MEDICAL INTERVENTIONS WERE A RESULT OF CASE-SPECIFIC CIRCUMSTANCES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH REGARDS TO MANUFACTURE, DESIGN, OR LABELING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21654456&pc=DSQ | 2916596-2025-01539 | 2025-02-22 05:00 | Death | THORATEC CORPORATION | 2025-03-20 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Sepsis; Septic Shock; Skin Infection | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD AN ONGOING PSEUDOMONAS DRIVELINE INFECTION (PREVIOUSLY REPORTED UNDER MFR #2916596-2024-04397). THE PATIENT WAS ORIGINALLY ON INTRAVENOUS (IV) ANTIBIOTICS BUT WAS OFF OF THEM FOR TWO MONTHS BEFORE STARTING LONG TERM HOME ANTIBIOTICS, AS THEY DID NOT WISH TO HAVE THE PERIPHERALLY INSERTED CENTRAL CATHETER (PICC) LINE REPLACED AND HAD RECURRING INFECTION DURING THAT TIME. THE PATIENT LATER STOPPED HOME ANTIBIOTICS BY CHOICE AND HAD MORE INFECTIONS FOR 1 WEEK, BEFORE PRESENTING TO THE HOSPITAL AND BEING ADMITTED TO INPATIENT HOSPICE ON (B)(6) 2025. THE PATIENT WAS STARTED ON COMFORT CARE MEDICATIONS AND HAD AN ALTERED MENTAL STATUS DUE TO THE INFECTION. THE INFECTION PROGRESSED TO BACTEREMIA AND THEN SEPTIC SHOCK, AND THE PATIENT PASSED AWAY AT 03:03 ON (B)(6) 2025. IT WAS STATED THAT THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) FUNCTIONED AS INTENDED AT THE TIME OF DEATH. THE PUMP WAS NOT EXPLANTED AND NO AUTOPSY WAS PERFORMED. IT WAS STATED THAT THE DEATH WAS DEVICE OR THERAPY RELATED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF INFECTION COULD NOT BE CONCLUSIVELY DETERMINED. INFECTION IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3LEFT VENTRICULAR ASSIST SYSTEM. INFECTION HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: VOLUNTARY MEDWATCH REPORT #(B)(4) WAS RECEIVED 17MAR2025. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21654634&pc=OZD | 1220648-2025-26860 | 2025-02-23 05:00 | Death | ABIOMED, INC. | 2025-03-20 04:00 | OZD | IMPELLA CP | Positioning Failure | Cardiac Arrest | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT FOLLOWING IMPELLA CP IMPLANT, THE STAFF WAS UNABLE TO GET OPTIMAL POSITIONING FOR THE PUMP. DESPITE MULTIPLE ATTEMPTS TO REPOSITION, THE STAFF HAD TO SETTLE FOR A DEPTH OF 3.4 CENTIMETERS WITH THE DEVICE POINTING TOWARDS THE MITRAL VALVE. AS SUPPORT CONTINUED, THE PATIENT CODED AND PASSED AWAY. IT IS UNKNOWN IF THE POSITIONING ISSUE CONTRIBUTED TO THE PATIENT'S PASSING. Manufacturer Narrative: THE INVESTIGATION OF POSITIONING ISSUES HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. BASED ON THE CLINICAL DETAILS THE CAUSE OF THE POSITIONING ISSUE MOST LIKELY WAS USE RELATED SINCE THE IMPELLA STACKED AGAINST MITRAL VALVE. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21654652&pc=NPT | 2025587-2025-02000 | 2024-02-18 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-20 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Insufficient Information | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 2 MONTHS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. Manufacturer Narrative: UPDATED DATA: D4., H4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21654686&pc=NPT | 2025587-2025-02001 | 2024-06-03 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-20 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Arrhythmia | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 2 MONTHS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT SUDDENLY DIED. THE CAUSE OF DEATH IS THOUGHT TO BE ARRHYTHMIA. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21654703&pc=MHX | 3009156722-2025-00005 | 2025-01-29 05:00 | Death | SHENZHEN MINDRAY BIO-MEDICAL ELECTRONICS CO., LTD | 2025-03-20 04:00 | MHX | BENEVISION N17 PATIENT MONITOR | Adverse Event Without Identified Device or Use Problem; Output Problem; No Visual Prompts/Feedback | Insufficient Information | K213799 |  | 1 | Event Description: THE USER REPORTED BEING UNABLE TO SEE PATIENT DISCHARGE DATA FOR (B)(6) 2025, ON BENEVISION N17 (SERIAL NUMBER (B)(6). THE USER WAS ABLE TO SEE DATA ON (B)(6) 2025, AND (B)(6) 2025. THE PATIENT EXPIRED.Event Description: THE USER REPORTED BEING UNABLE TO SEE PATIENT DISCHARGE DATA FOR (B)(6) 2025, ON BENEVISION N17 (SERIAL NUMBER (B)(6)). THE USER WAS ABLE TO SEE DATA ON (B)(6) 2025, AND (B)(6) 2025. THE PATIENT EXPIRED DAYS AFTER THE (B)(6) 2025 (THE EXACT DATE IS UNKNOWN).Event Description: DURING AN INVESTIGATION FOLLOW-UP ON APRIL 7, 2025, THE CUSTOMER CLARIFIED THE INVESTIGATION REQUEST. NO ALLEGED MALFUNCTION WAS REPORTED. THE CUSTOMER IS REQUESTING THE KEYSTROKE FOR THE A7 ANESTHESIA MACHINE (SERIAL NUMBER (B)(6)) AND THE BENEVISION N17 PATIENT MONITOR (SERIAL NUMBER (B)(6)). THE REQUEST IS FOR RAW DATA FROM 09:00 AM TO 11:00 AM ON (B)(6) 2025. SOME VENTILATION CHANGES WERE MADE TO THE ANESTHESIA MACHINE AT 10:28. THEREFORE, THE CUSTOMER REQUESTED TO PROVIDE DETAILS ON THE KEYSTROKE ACTIVITY ON THE ANESTHESIA MACHINE AND BEDSIDE MONITOR 10 MINUTES BEFORE AND 10 MINUTES AFTER THAT TIME. THE PATIENT EXPIRED DAYS AFTER (B)(6) 2025 (THE EXACT DATE IS UNKNOWN). Manufacturer Narrative: A MINDRAY FIELD SERVICE REPRESENTATIVE COLLECTED THE PATIENT DATABASE, AND THE BENEVISION N17 MONITOR LOGS FOR INVESTIGATION. THE LOG ANALYSIS SHOWS THE FOLLOWING: THE N17 WAS SET TO STANDBY MODE AT 15:28 ON (B)(6) 2025 AND EXITED STANDBY MODE AT 7:39 ON (B)(6) 2025. IN STANDBY MODE, THE MONITOR STOPS ALL PARAMETER MEASUREMENTS AND DISABLES ALL ALARM INDICATIONS. IN CONCLUSION, THERE IS NO PATIENT DATA FOR (B)(6) 2025, BECAUSE THE N17 WAS IN STANDBY MODE. BENEVISION N17 PERFORMED ACCORDING TO SPECIFICATIONS, NO PRODUCT MALFUNCTION OCCURRED. Manufacturer Narrative: A MINDRAY FIELD SERVICE REPRESENTATIVE COLLECTED THE PATIENT DATABASE AND THE BENEVISION N17 MONITOR LOGS FOR REVIEW. MINDRAY PROVIDED THE BENEVISION N17 KEYSTROKE ACTIVITY FOR THE REQUIRED PERIOD. A MINDRAY SERVICE REPRESENTATIVE, UPON CLARIFYING THE REQUEST, RETURNED TO THE SITE TO COLLECT LOGS FROM THE ANESTHESIA UNIT. HOWEVER, SINCE THE ATTEMPT TO COLLECT DATA WAS ON (B)(6) 2025, THE DATA FOR THE OCCURRENCE DATE ((B)(6) 2025) IS NO LONGER AVAILABLE ON THE A-SERIES. NO DEVICE MALFUNCTION WAS REPORTED WITH THE MINDRAY DEVICES. UNITS ARE MONITORED PER SPECIFICATIONS. Manufacturer Narrative: A MINDRAY FIELD SERVICE REPRESENTATIVE COLLECTED THE PATIENT DATABASE, AND THE BENEVISION N17 MONITOR LOGS FOR INVESTIGATION. THE LOG ANALYSIS SHOWS THE FOLLOWING: THE N17 WAS SET TO STANDBY MODE AT 15:28 ON (B)(6) 2025 AND EXITED STANDBY MODE AT 7:39 ON (B)(6) 2025. IN STANDBY MODE, THE MONITOR STOPS ALL PARAMETER MEASUREMENTS AND DISABLES ALL ALARM INDICATIONS. IN CONCLUSION, THERE IS NO PATIENT DATA FOR (B)(6) 2025, BECAUSE THE N17 WAS IN STANDBY MODE. BENEVISION N17 PERFORMED ACCORDING TO SPECIFICATIONS, NO PRODUCT MALFUNCTION OCCURRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21654746&pc=MHX | 2221819-2025-00005 | 2025-01-29 05:00 | Death | SHENZHEN MINDRAY BIO-MEDICAL ELECTRONICS CO., LTD | 2025-03-20 04:00 | MHX | BENEVISION N17 PATIENT MONITOR | Adverse Event Without Identified Device or Use Problem; Output Problem; No Visual Prompts/Feedback | Insufficient Information |  |  | 1 | Event Description: THE USER REPORTED BEING UNABLE TO SEE PATIENT DISCHARGE DATA FOR (B)(6) 2025, ON BENEVISION N17 (SERIAL NUMBER (B)(6). THE USER WAS ABLE TO SEE DATA ON (B)(6) 2025, AND (B)(6) 2025. THE PATIENT EXPIRED.Event Description: THE USER REPORTED BEING UNABLE TO SEE PATIENT DISCHARGE DATA FOR (B)(6) 2025, ON BENEVISION N17 (SERIAL NUMBER (B)(6)). THE USER WAS ABLE TO SEE DATA ON (B)(6) 2025, AND (B)(6) 2025. THE PATIENT EXPIRED DAYS AFTER THE (B)(6) 2025 (THE EXACT DATE IS UNKNOWN).Event Description: DURING AN INVESTIGATION FOLLOW-UP ON APRIL 7, 2025, THE CUSTOMER CLARIFIED THE INVESTIGATION REQUEST. NO ALLEGED MALFUNCTION WAS REPORTED. THE CUSTOMER IS REQUESTING THE KEYSTROKE FOR THE A7 ANESTHESIA MACHINE (SERIAL NUMBER (B)(6)) AND THE BENEVISION N17 PATIENT MONITOR (SERIAL NUMBER (B)(6)). THE REQUEST IS FOR RAW DATA FROM 09:00 AM TO 11:00 AM ON (B)(6) 2025. SOME VENTILATION CHANGES WERE MADE TO THE ANESTHESIA MACHINE AT 10:28. THEREFORE, THE CUSTOMER REQUESTED TO PROVIDE DETAILS ON THE KEYSTROKE ACTIVITY ON THE ANESTHESIA MACHINE AND BEDSIDE MONITOR 10 MINUTES BEFORE AND 10 MINUTES AFTER THAT TIME. THE PATIENT EXPIRED DAYS AFTER (B)(6) 2025 (THE EXACT DATE IS UNKNOWN). Manufacturer Narrative: A MINDRAY FIELD SERVICE REPRESENTATIVE COLLECTED THE PATIENT DATABASE AND THE BENEVISION N17 MONITOR LOGS FOR INVESTIGATION. THE LOG ANALYSIS SHOWS THE FOLLOWING: THE N17 WAS SET TO STANDBY MODE AT 15:28 ON (B)(6) 2025 AND EXITED STANDBY MODE AT 7:39 ON (B)(6) 2025. IN STANDBY MODE, THE MONITOR STOPS ALL PARAMETER MEASUREMENTS AND DISABLES ALL ALARM INDICATIONS. IN CONCLUSION, THERE IS NO PATIENT DATA FOR (B)(6) 2025, BECAUSE THE N17 WAS IN STANDBY MODE. BENEVISION N17 PERFORMED ACCORDING TO SPECIFICATIONS, NO PRODUCT MALFUNCTION OCCURRED. Manufacturer Narrative: A MINDRAY FIELD SERVICE REPRESENTATIVE COLLECTED THE PATIENT DATABASE AND THE BENEVISION N17 MONITOR LOGS FOR REVIEW. MINDRAY PROVIDED THE BENEVISION N17 KEYSTROKE ACTIVITY FOR THE REQUIRED PERIOD. A MINDRAY SERVICE REPRESENTATIVE, UPON CLARIFYING THE REQUEST, RETURNED TO THE SITE TO COLLECT LOGS FROM THE ANESTHESIA UNIT. HOWEVER, SINCE THE ATTEMPT TO COLLECT DATA WAS ON (B)(6) 2025, THE DATA FOR THE OCCURRENCE DATE ((B)(6) 2025) IS NO LONGER AVAILABLE ON THE A-SERIES. NO DEVICE MALFUNCTION WAS REPORTED WITH THE MINDRAY DEVICES. UNITS ARE MONITORED PER SPECIFICATIONS. Manufacturer Narrative: A MINDRAY FIELD SERVICE REPRESENTATIVE COLLECTED THE PATIENT DATABASE AND THE BENEVISION N17 MONITOR LOGS FOR INVESTIGATION. THE LOG ANALYSIS SHOWS THE FOLLOWING: THE N17 WAS SET TO STANDBY MODE AT 15:28 ON (B)(6), 2025 AND EXITED STANDBY MODE AT 7:39 ON (B)(6) 2025. IN STANDBY MODE, THE MONITOR STOPS ALL PARAMETER MEASUREMENTS AND DISABLES ALL ALARM INDICATIONS. IN CONCLUSION, THERE IS NO PATIENT DATA FOR (B)(6) 2025, BECAUSE THE N17 WAS IN STANDBY MODE. BENEVISION N17 PERFORMED ACCORDING TO SPECIFICATIONS, NO PRODUCT MALFUNCTION OCCURRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21655003&pc=DSQ | 2916596-2025-01745 | 2025-01-21 05:00 | Death | THORATEC CORPORATION | 2025-03-20 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO INTRACRANIAL HEMORRHAGE. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. NO AUTOPSY WAS PERFORMED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF STROKE COULD NOT BE CONCLUSIVELY DETERMINED. STROKE IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. STROKE HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21655103&pc=GEX | 2020550-2025-00340 | 2025-02-19 05:00 | Death | KARL STORZ SE & CO. KG | 2025-03-20 04:00 | GEX | UNKNOWN | Insufficient Information | Insufficient Information |  |  | 1 | Event Description: ACCORDING TO THE CURRENTLY AVAILABLE INFORMATION IT WAS REPORTED THAT A PATIENT PASSED AWAY AFTER A HOLEP CASE. Manufacturer Narrative: APART FROM THE INFORMATION THAT THE PROCEDURE WAS A "HOLEP CASE" THERE IS NO FURTHER INFORMATION AVAILABLE. THEREFORE, IT COULD NOT BE DETERMINED WHICH PRODUCT WAS INVOLVED IN THIS EVENT. THE PRODUCT CODE "GEX" STATED IN SECTION D2 WAS DETERMINED AS THE MOST FITTING BASED ON THE LIMITED INFORMATION. THE PRODUCT CODE MAY CHANGE IF FURTHER INFORMATION BECOMES AVAILABLE. THE AFFECTED DEVICE HAS BEEN REQUESTED FOR INVESTIGATION BY THE MANUFACTURER. THE DEVICE WAS NOT YET RETURNED FOR INVESTIGATION. THE EVENT IS FILED UNDER INTERNAL KARL STORZ COMPLAINT ID: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21655140&pc=GEX | 9610617-2025-00340 | 2025-02-19 05:00 | Death | KARL STORZ SE & CO. KG | 2025-03-20 04:00 | GEX | UNKNOWN | Insufficient Information | Insufficient Information |  |  | 1 | Event Description: ACCORDING TO THE CURRENTLY AVAILABLE INFORMATION IT WAS REPORTED THAT A PATIENT PASSED AWAY AFTER A HOLEP CASE. Manufacturer Narrative: APART FROM THE INFORMATION THAT THE PROCEDURE WAS A "HOLEP CASE" THERE IS NO FURTHER INFORMATION AVAILABLE. THEREFORE, IT COULD NOT BE DETERMINED WHICH PRODUCT WAS INVOLVED IN THIS EVENT. THE PRODUCT CODE "GEX" STATED IN SECTION D2 WAS DETERMINED AS THE MOST FITTING BASED ON THE LIMITED INFORMATION. THE PRODUCT CODE MAY CHANGE IF FURTHER INFORMATION BECOMES AVAILABLE. THE AFFECTED DEVICE HAS BEEN REQUESTED FOR INVESTIGATION BY THE MANUFACTURER. THE DEVICE WAS NOT YET RETURNED FOR INVESTIGATION. THE EVENT IS FILED UNDER INTERNAL KARL STORZ COMPLAINT ID: (B)(4). Manufacturer Narrative: EVEN AFTER MULTIPLE REQUESTS, NO FURTHER INFORMATION ABOUT THE EVENT OR THE INVOLVED PRODUCT(S) WAS RECEIVED. THE ONLY AVAILABLE INFORMATION IS THE INITIAL DESCRIPTION, WHICH INDICATES THAT A PATIENT PASSED AWAY AFTER A HOLEP CASE. AS THERE IS NO INFORMATION REGARDING THE INVOLVED PRODUCT(S) AN INVESTIGATION AND ROOT CAUSE DETERMINATION IS NOT POSSIBLE. IF NEW OR ADDITIONAL RELEVANT INFORMATION WILL BE RECEIVED, THE INVESTIGATION WILL BE RE-OPENED ACCORDINGLY. THE EVENT IS FILED UNDER INTERNAL KARL STORZ COMPLAINT ID: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21655142&pc=NPT | 2015691-2025-02110 | 2025-02-24 05:00 | Death | EDWARDS LIFESCIENCES | 2025-03-20 04:00 | NPT | SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Rupture | P140031 |  | 1 | Event Description: AS REPORTED BY AN EDWARDS UNITED KINGDOM AFFILIATE, DURING A TRANSFEMORAL TAVR PROCEDURE WITH A 26MM SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE, THE VALVE WAS SUCCESSFULLY IMPLANTED WITH NOMINAL VOLUME MINUS 2 ML. HOWEVER, AN ANNULAR RUPTURE OCCURRED RESULTING IN THE PATIENT'S DEATH. THERE WAS NO DEVICE FAILURE. THE PATIENT WAS CONSIDERED EXTREMELY HIGH RISK AND THE SURGEONS HAD AGREED THEY WOULD NOT INTERVENE, SHOULD ANY COMPLICATIONS ARISE. THE ROOT CAUSE OF THE ANNULAR RUPTURE WAS FRAIL PATIENT ANATOMY (THE VALVE WAS UNDERSIZED TO MINIMIZE RISK). Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. ACCORDING TO THE DEVICE TRAINING MANUALS, RISK FACTORS FOR AORTIC DISSECTION, CARDIAC/AORTIC HEMATOMA, OR ANNULAR RUPTURE DURING THE TVR PROCEDURE INCLUDE SIGNIFICANT VALVE OVER SIZING, SEVERELY OBLITERATED SINUSES OF VALSALVA, PORCELAIN AORTA AND/OR PRESENCE OF BULKY CALCIFICATION, AND NARROW CALCIFIED SINOTUBULAR JUNCTION. IN ADDITION, ADVANCED AGE, FEMALE GENDER, SMALL BODY WEIGHT, AND STEROID DEPENDENCY CAN ALSO BE CONTRIBUTING FACTORS. THE SAPIEN TRANSCATHETER HEART VALVE (ALL MODELS) RELIES CALCIUM TO SECURELY ANCHOR IN THE LANDING ZONE. DESPITE THIS BENEFICIAL ASPECT OF CALCIUM, BULKY CALCIUM CAN INCREASE THE RISK OF CALCIFIC NODULE DISPLACEMENT INTO THE VASCULATURE, WHICH CAN LEAD TO VASCULAR INJURY. AT TIMES, THE EXTENT AND DISTRIBUTION OF CALCIUM CAN IMPAIR EASE OF DELIVERY OF THE VALVE, CORRECT POSITIONING OF THE VALVE, DEPLOYMENT OF THE VALVE, AND PROCEDURAL SUCCESS. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST PATIENT AND PROCEDURAL FACTORS (FRAIL PATIENT ANATOMY) CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED, AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21655707&pc=MVK | 3008642652-2025-02960 | 2025-02-04 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-03-20 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Failure to Convert Rhythm | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED AN APPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 07:31:03 ON (B)(6) 2025. AT 14:47:29, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF WITH CPR/MOTION ARTIFACT. AT 14:48:49, THE PATIENT RECEIVED THE APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS VT @ 130 BPM WITH MOTION ARTIFACT. THE ELECTRODE BELT WAS DISCONNECTED AT 14:49:06 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR HAS BEEN COMPLETED. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION. THE ELECTRODE BELT HAS NOT BEEN RECOVERED. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21656127&pc=GXI | 3006630150-2025-01664 | 2025-02-21 05:00 | Death | RELIEVANT MEDSYSTEMS, INC. | 2025-03-20 04:00 | GXI | INTRACEPT RF PROBE | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K180369 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD AN INTRACEPT PROCEDURE AND THE PROCEDURE WAS COMPLETED SUCCESSFULLY. HOWEVER, THE PATIENT HAD AN ISSUE WITH THE ANESTHESIA. THE PATIENT WAS TRANSPORTED TO THE EMERGENCY ROOM (ER) POST-PROCEDURE. THE PATIENT PASSED AWAY AFTER SOME ANESTHESIA CHALLENGES POST-PROCEDURE. THE PHYSICIAN ACKNOWLEDGED THAT THE PATIENTS DEATH WAS NOT A PROCEDURE OR DEVICE-RELATED COMPLICATION. NO DEVICES WILL BE RETURNED.Event Description: IT WAS REPORTED THAT THE PATIENT HAD AN INTRACEPT PROCEDURE AND THE PROCEDURE WAS COMPLETED SUCCESSFULLY. HOWEVER, THE PATIENT HAD AN ISSUE WITH THE ANESTHESIA. THE PATIENT WAS TRANSPORTED TO THE EMERGENCY ROOM (ER) POST-PROCEDURE. THE PATIENT PASSED AWAY AFTER SOME ANESTHESIA CHALLENGES POST-PROCEDURE. THE PHYSICIAN ACKNOWLEDGED THAT THE PATIENTS DEATH WAS NOT A PROCEDURE OR DEVICE-RELATED COMPLICATION. NO DEVICES WILL BE RETURNED.Event Description: IT WAS REPORTED THAT THE PATIENT HAD AN INTRACEPT PROCEDURE AND THE PROCEDURE WAS COMPLETED SUCCESSFULLY. HOWEVER, THE PATIENT HAD AN ISSUE WITH THE ANESTHESIA. THE PATIENT WAS TRANSPORTED TO THE EMERGENCY ROOM (ER) POST-PROCEDURE. THE PATIENT PASSED AWAY AFTER SOME ANESTHESIA CHALLENGES POST-PROCEDURE. THE PHYSICIAN ACKNOWLEDGED THAT THE PATIENT'S DEATH WAS NOT A PROCEDURE OR DEVICE-RELATED COMPLICATION. Manufacturer Narrative: CORRECTION TO INITIAL EMDR IN BLOCK D4.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21656629&pc=LLZ | 2183926-2025-00014 | 2025-02-08 05:00 | Death | MERGE HEALTHCARE | 2025-03-20 04:00 | LLZ | MERGE CARDIO | Adverse Event Without Identified Device or Use Problem | Unspecified Heart Problem; Insufficient Information | K192276 |  | 1 | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION FROM MERGE HEALTHCARE AND OTHER VENDORS SYSTEMS INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. MERGE CARDIO IS INTENDED TO ALLOW USERS TO REVIEW DIAGNOSTIC AND NON-DIAGNOSTIC QUALITY IMAGES, ANNOTATE STUDIES, PERFORM DIGITAL SUBTRACTION ON IMAGES, TO PERFORM QUANTITATIVE MEASUREMENTS ON IMAGES (INCLUDING BUT NOT LIMITED TO QUANTITATIVE CORONARY ANALYSIS, LEFT VENTRICULAR ANALYSIS, TIME, AREA, LENGTH, VELOCITY, ANGLE, VOLUME, AND VELOCITY-TIME INTEGRALS), TO GENERATE PHYSICIAN GENERATED CLINICAL REPORTS (VIA STRUCTURE REPORTING AND TEMPLATE BASED TOOLS), AND TO STORE THIS INFORMATION IN A DATABASE. ON 02/19/2025, A CUSTOMER CONTACTED MERGE HEALTHCARE FOR ASSISTANCE WITH MERGE CARDIO'S STUDY LIST FILTERS. MERGE TECHNICAL SUPPORT REMOTELY CONNECTED TO THE CUSTOMER'S SYSTEM TO ASSIST WITH FILTER QUESTIONS. THE CUSTOMER WAS ATTEMPTING TO LOCATE A ECHOCARDIOGRAM PATIENT STUDY PERFORMED ON (B)(6) 2025 AND WAS UNABLE TO CREATE A NEW FILTER TO LOCATE THE IMAGING STUDY. TECHNICAL SUPPORT PROVIDED ASSISTANCE TO ENABLE THE CUSTOMER TO CREATE THE STUDY LIST FILTER. MERGE TECHNICAL SUPPORT ASSISTED THE CUSTOMER WITH FINDING THE STUDY AND CONFIRMED THAT THE NEW STUDY LIST FILTER WORKED AS EXPECTED. THE CUSTOMER REPRESENTATIVE MENTIONED TO TECHNICAL SUPPORT THAT THE PATIENT HAD PASSED AWAY ON (B)(6) 2025. MERGE HEALTHCARE FOLLOWED UP WITH THE CUSTOMER TO GATHER MORE INFORMATION AND DETERMINED THAT THE PATIENT OUTCOME WAS DUE TO A PREEXISTING CONDITION. BASED ON THE AVAILABLE INFORMATION, MERGE CARDIO IS PERFORMING AS DESIGNED AND DID NOT CAUSE OR CONTRIBUTE TO THE PATIENT OUTCOME. IF NEW INFORMATION BECOMES AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21657364&pc=LWQ | 1649833-2025-00007 |  | Death | ON-X LIFE TECHNOLOGIES, INC. | 2025-03-20 04:00 | LWQ | ON-X AAP UNKNOWN CONFIGURATION | Insufficient Information | Atrial Fibrillation; Stroke/CVA; Dyspnea; Fever; Low Blood Pressure/ Hypotension; Unspecified Infection; Pain; Pleural Effusion; Pneumonia; Pericardial Effusion; Syncope/Fainting | P000037 |  | 1 | Event Description: ACCORDING TO PUBLICATION, ¿MIDTERM FOLLOW-UP OF COMPOSITE GRAFT REPLACEMENT OF THE AORTIC ROOT (30-YEAR EXPERIENCE)¿REMARKABLY SAFE, EFFECTIVE, AND DURABLE¿ SYEDA MANAHIL HAIDER JEOFFREY, MD, MOHAMMAD A. ZAFAR, MBBS, JUAN VELASCO, MD, AHAD KHATTAK, MD, HESHAM ELLAUZI, MB BCH, BAO (HONS), AFSHEEN NASIR, MD, ASANISH KALYANASUNDARAM, MD, BULAT A. ZIGANSHIN, MD, PHD, AND JOHN A. ELEFTERIADES, MD, PHD (HON), THIS PUBLICATION ¿MIDTERM FOLLOW-UP OF COMPOSITE GRAFT REPLACEMENT OF THE AORTIC ROOT (30-YEAR EXPERIENCE)¿REMARKABLY SAFE, EFFECTIVE, AND DURABLE¿ BY HAIDER-JEOFFREY, ET AL READ AT THE 103RD ANNUAL MEETING OF THE AMERICAN ASSOCIATION FOR THORACIC SURGERY, LOS ANGELES, CALIFORNIA, MAY 6-9, 2023 WAS DISCOVERED DURING THE MOST RECENT ON-X AAP EU MDR CYCLE. THIS PAPER REPORTS ON A RETROSPECTIVE STUDY FOR 564 PATIENTS THAT UNDERWENT CVG (COMPOSITE VALVED GRAFT) BY A SINGLE SURGEON (J.A.E) AT YALE UNIVERSITY SCHOOL OF MEDICINE BETWEEN JAN 1990 AND 2020. MECHANICAL AND TISSUE PROTHESIS WERE USED IN 427 (75.7%) AND 136 (24.2%) OF PATIENTS RESPECTIVELY. ALL PATIENTS WITH MECHANICAL VALVE RECEIVED ORAL WARFARIN AND LONG-TERM ANTICOAGULATION WAS MAINTAINED WITH A TARGET INR OF 2 TO 2.5 FOR ST. JUDE VALVES AND 1.8 TO 2.5 FOR ON-X VALVES. THE ENDPOINTS OF THE STUDY WERE OPERATIVE MORTALITY, ENDOCARDITIS, STROKE, BLEEDING, THROMBOEMBOLISM AND REOPERATION OF THE ROOT ALONG WITH OVERALL SURVIVAL. THE AUTHORS CONCLUDE THAT COMPOSITE GRAFT REPLACEMENT OF THE AORTIC ROOT IS ASSOCIATED WITH LOW OPERATIVE RISK AND PRODUCES EXCELLENT LONG-TERM SURVIVAL, LOW INCIDENCE OF LATE REOPERATION OR LATE AORTIC EVENTS, AND VANISHINGLY LOW INCIDENCE OF REOPERATION IN THE ROOT PORTION OF THE AORTA. THIS INVESTIGATION IS RELEGATED TO ONXAAP UNKNOWN CONFIGURATION FOR: RAPID ATRIAL FIBRILLATION/ARRHYTHMIA PERICARDIAL EFFUSION, INFECTION (STERNAL, GROIN, SAPHENOUS VEIN SITE), ALTERED MENTAL STATUS/SYNCOPE/VISUAL CHANGES, DYSPNEA, PAIN, STROKE/TIA/TE, UNKNOWN, HYPOTENSION, FEVER, PLEURAL EFFUSION, PNEUMONIA. Manufacturer Narrative: THIS INVESTIGATION IS CURRENTLY ONGOING. ANY ADDITIONAL INFORMATION WILL BE PROVIDED IN THE FOLLOW-UP REPORT. THIS REPORT IS BEING SUBMITTED AS REQUIRED BY FEDERAL REGULATIONS AND DOES NOT CONSTITUTE AN ADMISSION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. FURTHERMORE, THIS REPORT REFLECTS THE EVENT AS ALLEGED BY THE COMPLAINANT AND DOES NOT IMPLY THAT THE INFORMATION REPORTED TO ARTIVION ¿ FORMERLY CRYOLIFE/JOTEC IS ACCURATE OR HAS BEEN CONFIRMED BY ARTIVION ¿ FORMERLY CRYOLIFE/JOTEC. Manufacturer Narrative: ACCORDING TO PUBLICATION, ¿MIDTERM FOLLOW-UP OF COMPOSITE GRAFT REPLACEMENT OF THE AORTIC ROOT (30-YEAR EXPERIENCE)¿REMARKABLY SAFE, EFFECTIVE, AND DURABLE¿ SYEDA MANAHIL HAIDER JEOFFREY, MD, MOHAMMAD A. ZAFAR, MBBS, JUAN VELASCO, MD, AHAD KHATTAK, MD, HESHAM ELLAUZI, MB BCH, BAO (HONS), AFSHEEN NASIR, MD, ASANISH KALYANASUNDARAM, MD, BULAT A. ZIGANSHIN, MD, PHD, AND JOHN A. ELEFTERIADES, MD, PHD (HON), THIS PUBLICATION ¿MIDTERM FOLLOW-UP OF COMPOSITE GRAFT REPLACEMENT OF THE AORTIC ROOT (30-YEAR EXPERIENCE)¿REMARKABLY SAFE, EFFECTIVE, AND DURABLE¿ BY HAIDER-JEOFFREY, ET AL READ AT THE 103RD ANNUAL MEETING OF THE AMERICAN ASSOCIATION FOR THORACIC SURGERY, LOS ANGELES, CALIFORNIA, MAY 6-9, 2023 WAS DISCOVERED DURING THE MOST RECENT ON-X AAP EU MDR CYCLE. THIS PAPER REPORTS ON A RETROSPECTIVE STUDY FOR 564 PATIENTS THAT UNDERWENT CVG (COMPOSITE VALVED GRAFT) BY A SINGLE SURGEON (J.A.E) AT YALE UNIVERSITY SCHOOL OF MEDICINE BETWEEN JAN 1990 AND 2020. MECHANICAL AND TISSUE PROTHESIS WERE USED IN 427 (75.7%) AND 136 (24.2%) OF PATIENTS RESPECTIVELY. ALL PATIENTS WITH MECHANICAL VALVE RECEIVED ORAL WARFARIN AND LONG-TERM ANTICOAGULATION WAS MAINTAINED WITH A TARGET INR OF 2 TO 2.5 FOR ST. JUDE VALVES AND 1.8 TO 2.5 FOR ON-X VALVES. THE ENDPOINTS OF THE STUDY WERE OPERATIVE MORTALITY, ENDOCARDITIS, STROKE, BLEEDING, THROMBOEMBOLISM AND REOPERATION OF THE ROOT ALONG WITH OVERALL SURVIVAL. THE AUTHORS CONCLUDE THAT COMPOSITE GRAFT REPLACEMENT OF THE AORTIC ROOT IS ASSOCIATED WITH LOW OPERATIVE RISK AND PRODUCES EXCELLENT LONG-TERM SURVIVAL, LOW INCIDENCE OF LATE REOPERATION OR LATE AORTIC EVENTS, AND VANISHINGLY LOW INCIDENCE OF REOPERATION IN THE ROOT PORTION OF THE AORTA. THIS INVESTIGATION IS RELEGATED TO ONXAAP UNKNOWN CONFIGURATION FOR THE 12 EVENTS LISTED BELOW: RAPID ATRIAL FIBRILLATION/ARRHYTHMIA. PERICARDIAL EFFUSION. INFECTION (STERNAL, GROIN, SAPHENOUS VEIN SITE). ALTERED MENTAL STATUS/SYNCOPE/VISUAL CHANGES. DYSPNEA. PAIN. STROKE/TIA/TE. UNKNOWN. HYPOTENSION. FEVER. PLEURAL EFFUSION. PNEUMONIA. MULTIPLE ATTEMPTS FOR ADDITIONAL INFORMATION HAVE GONE UNMET. NO ADDITIONAL INFORMATION FORTHCOMING. A REVIEW OF MANUFACTURING RECORDS COULD NOT BE PERFORMED AS A DEFINITIVE SERIAL NUMBER AND DATE OF SURGERY WERE NOT PROVIDED BY THE COMPLAINANT. A REVIEW OF THE AVAILABLE INFORMATION WAS PERFORMED. THIS INVESTIGATION REVIEWS THE CLINICAL OUTCOMES REPORTED IN THE 2024 PUBLICATION TITLED ¿MIDTERM FOLLOW-UP OF COMPOSITE GRAFT REPLACEMENT OF THE AORTIC ROOT (30-YEAR EXPERIENCE)¿REMARKABLY SAFE, EFFECTIVE, AND DURABLE¿ BY HAIDER-JEOFFREY ET AL. THE STUDY PRESENTS A SINGLE-SURGEON, SINGLE-CENTER RETROSPECTIVE COHORT ANALYSIS CONDUCTED AT THE YALE UNIVERSITY SCHOOL OF MEDICINE. IT INCLUDES 564 ADULT PATIENTS WHO UNDERWENT COMPOSITE VALVED GRAFT (CVG) REPLACEMENT OF THE AORTIC ROOT BETWEEN 1990 AND 2020. THE MEAN AGE OF THE PATIENT COHORT WAS 55.6 ± 13.8 YEARS, WITH THOSE RECEIVING MECHANICAL VALVES AVERAGING 52.2 YEARS AND THOSE RECEIVING TISSUE VALVES AVERAGING 65.7 YEARS. THE STUDY POPULATION WAS PREDOMINANTLY MALE (84%). MECHANICAL PROSTHESES WERE USED IN 427 PATIENTS (75.7%), WHILE TISSUE PROSTHESES WERE IMPLANTED IN 136 PATIENTS (24.2%). PREFABRICATED ST. JUDE MEDICAL VALVED CONDUITS (AND, MORE RECENTLY, ON-X VALVED CONDUITS) WERE USED FOR MECHANICAL VALVE REPLACEMENTS. FOR THE TISSUE PROSTHESIS GROUP, A CARPENTIEREDWARDS, EDWARDS MAGNA, OR MAGNA EASE VALVE WAS HAND-SEWN INTO A HEMASHIELD GRAFT TO CREATE THE CONDUIT. OPERATIVE MORTALITY WAS REPORTED IN 12 PATIENTS (2.1%), WITH EIGHT DEATHS IN THE MECHANICAL GROUP AND FOUR IN THE TISSUE VALVE GROUP. OF THESE, FOUR EARLY DEATHS (33.3%) OCCURRED IN PATIENTS PRESENTING WITH ACUTE TYPE A AORTIC DISSECTION. DURING BOTH SHORT- AND LONG-TERM FOLLOW-UP, A TOTAL OF 92 PATIENTS DIED, COMPRISING 71 DEATHS IN THE MECHANICAL VALVE GROUP AND 21 IN THE TISSUE VALVE GROUP. LONG-TERM SURVIVAL FOR THE OVERALL COHORT WAS 91.3% AT 5 YEARS, 85.4% AT 10 YEARS, 76.9% AT 15 YEARS, AND 70.6% AT 20 YEARS. THE CUMULATIVE INCIDENCE OF REOPERATION ON THE AORTIC ROOT WAS REMARKABLY LOW, REPORTED AT 0.6%, 1.1%, AND 1.1% AT 5, 10, AND 15 YEARS, RESPECTIVELY. FOR DISTAL AORTIC SEGMENTS, THE CUMULATIVE INCIDENCE OF REINTERVENTION WAS SLIGHTLY HIGHER¿0.9% AT 5 YEARS, 2.8% AT 10 YEARS, AND 4.9% AT 15 YEARS. MAJOR LATE COMPLICATIONS, INCLUDING BLEEDING AND THROMBOEMBOLIC EVENTS, OCCURRED IN 3.1% OF PATIENTS AT 5 YEARS, 4.2% AT 10 YEARS, AND 5.3% AT 15 YEARS. LONG-TERM SURVIVAL OUTCOMES WERE NOT STATISTICALLY DIFFERENT BETWEEN THE MECHANICAL AND BIOPROSTHETIC VALVE RECIPIENTS. MECHANICAL VALVE PATIENTS HAD SURVIVAL RATES OF 97.2%, 93.8%, 89.4%, AND 79.1% AT 1, 5, 10, AND 15 YEARS, RESPECTIVELY, WHILE BIOPROSTHETIC VALVE PATIENTS HAD RATES OF 94.8%, 92.5%, 82.1%, AND 74.2% OVER THE SAME INTERVALS (LOG-RANK P = 0.1). SIMILARLY, NO SIGNIFICANT DIFFERENCE IN THE CUMULATIVE INCIDENCE OF BLEEDING OR THROMBOEMBOLIC EVENTS WAS OBSERVED BETWEEN THE TWO GROUPS. SIX PATIENTS UNDERWENT TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) DUE TO DEGENERATION OF THEIR BIOPROSTHETIC VALVES¿TWO FOR AORTIC STENOSIS AND FOUR FOR AORTIC INSUFFICIENCY. THESE INTERVENTIONS OCCURRED AT 10, 13, 14, 15, 15, AND 18 YEARS FOLLOWING THE INITIAL SURGERY. ACCORDING TO THE AUTHORS, COMPOSITE GRAFT REPLACEMENT OF THE AORTIC ROOT DEMONSTRATED LOW OPERATIVE RISK, EXCELLENT LONG-TERM SURVIVAL, MINIMAL INCIDENCE OF REOPERATION, AND LOW OCCURRENCE OF LATE AORTIC EVENTS. IMPORTANTLY, THE NEED FOR REOPERATION AT THE ROOT SEGMENT WAS EXCEEDINGLY RARE OVER THE STUDY¿S LONG-TERM FOLLOW-UP. ALL COMPLICATIONS AND MORTALITY EVENTS REPORTED ARE CONSISTENT WITH THE RECOGNIZED RISKS OUTLINED IN THE INSTRUCTIONS FOR USE (IFU) FOR THE ON-X VALVE. REFERENCES: ON-X ASCENDING AORTIC PROSTHESIS WITH THE VASCUTEK GELWEAVE VALSALVA GRAFT, INSTRUCTIONS FOR USE. THIS CLINICAL OUTCOME ANALYSIS IS BASED ON A COMPREHENSIVE 30-YEAR EXPERIENCE FROM A SINGLE SURGEON AT YALE UNIVERSITY SCHOOL OF MEDICINE INVOLVING 564 ADULT PATIENTS UNDERGOING COMPOSITE GRAFT AORTIC ROOT REPLACEMENT. WHILE ON-X VALVED CONDUITS WERE USED IN A PORTION OF THE MECHANICAL GROUP, THE PUBLICATION DOES NOT SPECIFY WHICH PATIENTS RECEIVED ON-X VALVES VERSUS ST. JUDE MEDICAL VALVES. THEREFORE, IT IS NOT POSSIBLE TO ASSOCIATE ANY SPECIFIC OUTCOMES, INCLUDING COMPLICATIONS OR MORTALITY, DIRECTLY WITH THE ON-X AORTIC AAP DEVICE. ACCORDINGLY, WE ARE UNABLE TO EVALUATE THE CONTRIBUTION, IF ANY, OF THE AAP DEVICE TO THE REPORTED CLINICAL EVENTS. NOTABLY, ALL OBSERVED ADVERSE EVENTS ARE INCLUDED AMONG THE KNOWN RISKS DOCUMENTED IN THE ON-X VALVE¿S INSTRUCTIONS FOR USE (IFU). NO FURTHER ACTION IS REQUIRED WITHOUT FURTHER INFORMATION. A COMPLAINT AND RECALL QUERY COULD NOT BE PERFORMED AS THE SERIAL NUMBER IS UNKNOWN. BASED ON THE AVAILABLE INFORMATION, A DEFINITIVE ROOT CAUSE FOR THIS EVENT CANNOT BE DETERMINED. THIS EVENT DOES NOT IDENTIFY ADDITIONAL HAZARDS OR MODIFY THE PROBABILITY AND SEVERITY OF EXISTING HAZARDS. THERE IS NO INDICATION THAT AN ERROR OR DEFICIENCY OCCURRED AT ARTIVION ¿ FORMERLY CRYOLIFE/JOTEC AND THE IFU ADEQUATELY COMMUNICATES RISK. THIS COMPLAINT WAS REVIEWED FOR A CAPA EVALUATION AND A CAPA IS NOT WARRANTED AT THIS TIME. ARTIVION WILL CONTINUE TO MONITOR SIMILAR COMPLAINTS TO DETERMINE IF ADDITIONAL ACTIONS ARE WARRANTED; HOWEVER, AT THIS TIME NO FURTHER ACTIONS ARE NECESSARY. THIS REPORT IS BEING SUBMITTED AS REQUIRED BY FEDERAL REGULATIONS AND DOES NOT CONSTITUTE AN ADMISSION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. FURTHERMORE, THIS REPORT REFLECTS THE EVENT AS ALLEGED BY THE COMPLAINANT AND DOES NOT IMPLY THAT THE INFORMATION REPORTED TO ARTIVION ¿ FORMERLY CRYOLIFE/JOTEC IS ACCURATE OR HAS BEEN CONFIRMED BY ARTIVION ¿ FORMERLY CRYOLIFE/JOTEC.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21657489&pc=BZD | 2518422-2025-037897 | 2022-08-19 04:00 | Death | RESPIRONICS, INC. | 2025-03-20 04:00 | BZD | PHILIPS CPAP DEVICE | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A PHILIPS CPAP DEVICE. THE PATIENT HAS PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21657525&pc=LWS | 3008973940-2025-02045 | 2025-02-21 05:00 | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-03-20 04:00 | LWS | SPRINT QUATTRO SECURE S MRI SURESCAN | Over-Sensing | Syncope/Fainting; Insufficient Information | P920015 |  | 1 | Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT HAS PASSED AWAY.Event Description: IT WAS REPORTED THAT TWO WEEKS POST IMPLANT, THE PATIENT EXPERIENCED A VENTRICULAR TACHYCARDIA (VT)/VENTRICULAR FIBRILLATION (VF) STORM, COLLAPSED AND CARDIOPULMONARY RESUSCITATION WAS PERFORMED. THE RIGHT VENTRICULAR (RV) LEAD TRIGGERED A LEAD INTEGRITY ALERT (LIA) FOR SENSING INTEGRITY COUNTER (SIC) AND HIGH RATE NON-SUSTAINED EPISODES. IT WAS NOTED THAT ALL LEAD CHARACTERISTICS (IMPEDANCE, THRESHOLD, AND WAVE AMPLITUDE) WERE NORMAL AND STABLE UP TO THE VT/VF STORM. THE LEAD REMAINS IN USE. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. Manufacturer Narrative: CORRECTION: G3 AND H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: B2: YEAR VALID MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE LEAD WAS NOT RETURNED FOR ANALYSIS, HOWEVER, PERFORMANCE DATA COLLECTED FROM THE DEVICE WAS RECEIVED AND ANALYZED. ANALYSIS OF THE DEVICE MEMORY INDICATED OVERSENSING ASSOCIATED WITH THE RIGHT VENTRICULAR LEAD. ANALYSIS OF THE DEVICE MEMORY ALSO INDICATED THE CRITERIA FOR THE RIGHT VENTRICULAR LEAD INTEGRITY ALERT WERE METN AND OVERSENSING DUE TO NON-PHYSIOLOGIC SIGNALS/SENSING INTEGRITY COUNTER. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTION: G3 FIELD FROM PREVIOUS FOLLOW-UP REPORT 001 WAS OMITTED. THE OMITTED DATE SHOULD HAVE BEEN 24-FEB-2025. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTION: G3 FIELD FROM PREVIOUS FOLLOW-UP REPORT 001 WAS OMITTED. THE OMITTED DATE SHOULD HAVE BEEN 24-FEB-2025. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21658059&pc=OZD | 1220648-2025-26862 | 2025-02-19 05:00 | Death | ABIOMED, INC. | 2025-03-20 04:00 | OZD | IMPELLA CP | Restricted Flow rate | Cardiac Arrest | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT AND EXPERIENCED LOW PUMP FLOW WHICH RESULTED IN REPLACEMENT OF THE DEVICE WITH A DEMISE OUTCOME. THE PATIENT CARDIAC ARRESTED ON THE GOLF COURSE. CARDIOPULMONARY RESUSCITATION WAS PERFORMED, AND THE PATIENT WAS SHOCKED TWICE IN FIELD, ARRIVED AT THE EMERGENCY DEPARTMENT AND WAS INTUBATED. ELECTROCARDIOGRAM SHOWED ANTERIOR ST-SEGMENT ELEVATION MYOCARDIAL INFARCTION. THE PATIENT WAS TAKEN TO THE CATHETERIZATION LAB FOR PERCUTANEOUS CORONARY INTERVENTION. POST LEFT ANTERIOR DESCENDING ARTERY STENT FOR 95% LESIONS, THE DIAGNOSTIC SHEATH WAS EXCHANGED FOR 14FR X 13 CM PEEL AWAY ABIOMED SHEATH. THE PIGTAIL CATHETER WAS ADVANCED TO LEFT VENTRICLE WITH 0.035J WIRE. THE GUIDEWIRE WAS EXCHANGED FOR 0.018 ABIOMED WIRE, PIGTAIL REMOVED AND IMPELLA ADVANCED TO LEFT VENTRICLE. CARDIOPULMONARY RESUSCITATION WAS BEING ADMINISTERED DURING THIS TIME. THE PUMP WAS TURNED ON BUT WAS ONLY ABLE GET FLOWS OF 2L AND WAS GETTING SUCTION AT PERFORMANCE LEVEL P-4. THE PATIENT REGAINED BLOOD PRESSURE; ON LEVOPHED AT MAX DOSE. HUNG EPINEPHRINE, PATIENT PRESSURE INCREASED TO AROUND 60 MEAN ARTERIAL PRESSURE. TROUBLESHOOTING LOW FLOWS AND SUCTION, THE POSITION WAS CHECKED UNDER FLUOROSCOPY, REPOSITIONED THE DEVICE WITH NO RESOLUTION OF FLOW. ECHOCARDIOGRAM WAS COMPLETED; IMAGES WERE CHALLENGING. HOWEVER, THE PHYSICIAN DID NOT FEEL THE RIGHT HEART WAS AN ISSUE. THERE WAS A SMALL EFFUSION. THE PHYSICIAN FELT ALL OPTIONS WERE EXHAUSTED AND THE MEAN ARTERIAL PRESSURE DROPPED TO 40. THEREFORE, THE PHYSICIAN DECIDED TO EXPLANT AND REPLACE THE IMPELLA CP DEVICE DUE TO LOW FLOWS. THE SECOND IMPELLA CP WAS INSERTED. THE PIGTAIL CATHETER WAS ADVANCED TO THE LEFT VENTRICLE, ABIOMED 0.018 WIRE ADVANCED, THE PIGTAIL REMOVED, AND THE NEW IMPELLA CP WAS ADVANCED AND PLACED. THE PUMP WAS TURNED ON AND THE SAME LOW PUMP FLOW ISSUE OCCURRED WITH FLOW. FLOWS WAS UNABLE TO GO ABOVE 1.5L, CPR WAS STARTED DUE TO MEAN ARTERIAL PRESSURE DROPPING INTO THE 30¿S. DOCTOR WENT TO SPEAK WITH FAMILY. CODE PROTOCOL FOLLOWED. THE PUMP WAS TURNED ON AND OFF MULTIPLE TIMES TO SEE THE UNDERLYING PRESSURE. AFTER SPEAKING WITH FAMILY, THE DECISION WAS MADE TO TURN ON PUMP AND THE PATIENT REMAINED IN PULSELESS ELECTRICAL ACTIVITY AND EXPIRED. THE TREATING PHYSICIAN COMMENTED THAT HE SUSPECTED THE PATIENT MIGHT HAVE HAD A PULMONARY EMBOLISM AND THAT THE TEAM WAS UNABLE TO GET ENOUGH FLOW FROM THE RIGHT HEART. MEDICAL SAFETY REVIEWED THE EVENT AND NOTED THERE IS NOT ENOUGH EVIDENCE, HOWEVER, TO EXCLUDE IMPELLA CP DEVICE PUMP 1 AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. THERE WAS LOW PUMP FLOW FOR REASON THAT IS UNKNOWN AND DELAY IN NEEDED THERAPY INITIATED, RESULTED WITH IMPELLA CP DEVICE PUMP 1. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THIS IS ONE OF TWO DEVICE MANUFACTURING REPORTS ASSOCIATED WITH THIS EVENT. REFER TO INITIAL MEDICAL DEVICE REPORT FOR THE IMPELLA CP DEVICE (PUMP 2) SERIAL NUMBER (B)(6) WITH DATE OF EVENT 19-FEB-2025 AND IDENTIFIER ENDING IN (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21658234&pc=NVY | 1028232-2025-01420 | 2025-03-18 04:00 | Death | BIOTRONIK SE & CO. KG | 2025-03-20 04:00 | NVY | LINOX SMART S 65 | Fracture; Over-Sensing; Inappropriate/Inadequate Shock/Stimulation | Shock from Patient Lead(s) | P980023 |  | 1 | Event Description: PATIENT UNDERWENT SURGERY FOR OVERSENSING DUE TO FRACTURED LEAD RESULTING IN INAPPROPRIATE SHOCKS. DURING LASER LEAD EXTRACTION, THE SVC WAS RIPPED AND THE PATIENT PASSED AWAY. SHOULD ADDITIONAL INFORMATION BE RECEIVED THIS FILE WILL BE UPDATED. Manufacturer Narrative: COMBINATION PRODUCT: YES. Manufacturer Narrative: COMBINATION PRODUCT: YES. THE DEVICE IS CURRENTLY NOT AVAILABLE FOR ANALYSIS. NO CONCLUSION CAN BE DRAWN AT THIS TIME. NO ADDITIONAL INFORMATION IS AVAILABLE AT THE MOMENT. THE FILE IS CLOSED. THE INVESTIGATION WILL BE RE-OPENED SHOULD ADDITIONAL DATA BECOME AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21658235&pc=NPT | 2015691-2025-02114 | 2025-02-24 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-20 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P140031 |  | 1 | Event Description: AS REPORTED FROM PATIENT SUPPORT, PSC RECEIVED A CALL TODAY FROM A CARE ADVOCATE. THE CARE ADVOCATE INFORMED US THAT APPROXIMATELY 3 MONTHS POST TRANSFEMORAL TAVR PROCEDURE WITH A 26 MM SAPIEN 3 ULTRA VALVE, THE PATIENT EXPIRED. ACCORDING TO THE CARE ADVOCATE, ONE MONTH PRIOR TO THE IMPLANTING PROCEDURE, THE PATIENT SUFFERED A HEART ATTACK AND WAS BEING TREATED AT THE HOSPITAL. THE PATIENT HAD RECOVERED, BUT THEIR MEDICAL TEAM DETERMINED THAT THEY WOULD NEED A TAVR, WHICH COULD ONLY BE PERFORMED AT A DIFFERENT HOSPITAL. THE PATIENT WAS DISCHARGED HOME FOR A FEW DAYS AND WAS THEN ADMITTED TO THE IMPLANTING HOSPITAL. AFTER THE PROCEDURE, THE PATIENT ''SUFFERED A CATASTROPHIC STROKE AND WAS NEVER ABLE TO FULLY RECOVER.'' THE PATIENT PASSED AWAY FIVE DAYS LATER. ACCORDING TO THE CARE ADVOCATE, THE PATIENT'S CAUSE OF DEATH ON THEIR CERTIFICATE STATES ''BRAIN DEATH, CEREBRAL HEMORRHAGE.'' THE CARE ADVOCATE DID NOT ASK TO SPEAK WITH OUR PHYSICIANS AND DID NOT MAKE ANY ALLEGATIONS REGARDING OUR DEVICES. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), PERMANENT OR TRANSIENT NEUROLOGICAL EVENTS SUCH AS TRANSIENT ISCHEMIC ATTACK (TIA) AND STROKE ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND THE USE OF THE EDWARDS TRANSCATHETER HEART VALVE (THV) DEVICES. ACCORDING TO THE LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EW, STROKE IS RECOGNIZED IN THE LITERATURE AS A WELL-KNOWN COMPLICATION IN A SMALL NUMBER OF PATIENTS UNDERGOING THV. RISK FACTORS CORRELATING WITH SEVERAL PATIENT CO-MORBIDITIES HAVE BEEN IDENTIFIED. ALTHOUGH IN MANY CASES THE ROOT CAUSE OF THE EVENT IS UNABLE TO BE DETERMINED, STROKES DURING THV ARE UNDOUBTEDLY MULTIFACTORIAL, THE DOMINANT ETIOLOGY LIKELY BEING INTRA PROCEDURE EMBOLIC EVENTS. A TRANSCRANIAL DOPPLER STUDY DURING THV DEMONSTRATED THAT MOST PROCEDURAL EMBOLIC EVENTS OCCURRED DURING BALLOON VALVULOPLASTY, MANIPULATION OF CATHETERS ACROSS THE AORTIC VALVE, AND VALVE IMPLANTATION. AN ANALYSIS IN PATIENTS UNDERGOING VALVE SURGERY REVEALED FOUR BASELINE CHARACTERISTICS AND TWO PROCEDURAL EVENTS THAT WERE ASSOCIATED WITH EARLY POST-PROCEDURE STROKE: FEMALE SEX, EF < 30%, DIABETES, AGE OLDER THAN 70 YEARS, BYPASS PROCEDURE TIME> 120 MIN, AND CALCIFICATION OF THE ASCENDING AORTA. PREDICTORS OF LATE STROKE HAVE INCLUDED FEMALE SEX, AGE OLDER THAN 75 YEARS, ATRIAL FIBRILLATION, AND A HISTORY OF OR CURRENT SMOKING. THERE WERE NO SIGNIFICANT DIFFERENCES IN THE FREQUENCY OF LATE STROKES BETWEEN THV AND AVR PATIENTS. AFTER THV, THERE APPEARS TO BE A MORE CONSIDERABLE PROPORTION OF EARLY STROKES OCCURRING < 24 H POST-PROCEDURE, BUT THV PATIENTS WITH MULTIPLE CO-MORBIDITIES ARE PROBABLY AT HIGHER RISK OF BOTH EARLY AND LATE STROKES. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST/INDICATE THE PATIENT'S AGE MAY HAVE CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED, AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21658354&pc=OZD | 1220648-2025-26864 | 2025-02-23 05:00 | Death | ABIOMED, INC. | 2025-03-20 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemolysis; Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THE PATIENT WAS ON IMPELLA CP SUPPORT AND DURING SUPPORT, THE PATIENT HAD OOZING AT THE ACCESS SITE. ANGLE MATCHING WAS PERFORMED, LEG IMMOBILIZER WAS IN PLACE. HEPARIN WAS CHANGED TO BICARB IN THE PURGE. THE OOZING SLOWED AND SUPPORT CONTINUED. THE PATIENT ALSO HAD HEMOLYSIS. IT WAS FURTHER REPORTED THAT THE PATIENT EXPIRED WHILE ON SUPPORT HOWEVER NO OTHER INFORMATION WAS PROVIDED. THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21658488&pc=OZD | 1220648-2025-26865 | 2025-02-23 05:00 | Death | ABIOMED, INC. | 2025-03-20 04:00 | OZD | IMPELLA 5.5 | Increase in Pressure; Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT THE PATIENT ON IMPELLA 5.5 SUPPORT DEVELOPED A HEMATOMA AT THE ACCESS SITE NOTED UPON ARRIVAL TO THE INTENSIVE CARE UNIT. SANDBAG WAS APPLIED AND BLOOD PRODUCTS WERE ADMINISTERED. ADDITIONALLY, DURING INITIATION OF GASTROINTESTINAL (GI) CAPSULE STUDY, THE PATIENT BEGAN SHOWING SIGNS CEREBROVASCULAR ACCIDENT (CVA). RIGHT SIDED DEFICITS. THE PATIENT IS OFF SYSTEMIC ANTICOAGULATION FOR FOUR DAYS DUE TO GI BLEED. SUCCESSFUL STENT OF LEFT MAIN CORONARY ARTERY WAS DONE. NEUROLOGICAL SYMPTOMS SLOWLY WERE IMPROVING. IT WAS FURTHER REPORTED THAT THE PATIENT PREVIOUS EMBOLIC CVA FROM (B)(6) DEVELOPED INTO A HEMORRHAGIC CVA EVENING OF (B)(6). THE PATIENT REMAINS OFF SYSTEMIC ANTICOAGULATION. SERIAL COMPUTED TOMOGRAPHY SCANS WILL BE DONE TO MONITOR DEVELOPMENT. NO INTERVENTION AVAILABLE AT THIS TIME PER NEUROLOGY. THE PATIENT REMAINS ON SUPPORT.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE IMPELLA 5.5 HAD PURGE PRESSURE AND PURGE SYSTEM BLOCKED ALARM WHILE IN USE WITH THE PATIENT. THE PATIENT HAS BEEN ON IMPELLA SUPPORT SINCE (B)(6) 2025 AS NO LONG-TERM ADVANCED THERAPY WAS PLANNED AT THIS POINT. THE PURGE CASSETTE AND TUBING WERE CHANGED WITHOUT RESOLUTION. TISSUE PLASMINOGEN ACTIVATOR (TPA) WAS ADDED TO THE PURGE.Event Description: THE DECISION WAS MADE TO WITHDRAW CARE AND THE PATIENT EXPIRED. IT IS UNKNOWN IF THE USE OF THE IMPELLA WAS RELATED TO PATIENT DEATH, THUS THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING, STROKE/CVA, VASCULAR DAMAGE - PERIPHERAL VESSEL, AND HIGH PURGE PRESSURE HAS BEEN COMPLETED. THE ROOT CAUSE OF THE HIGH PURGE PRESSURE ISSUE WAS MOST LIKELY BIOMATERIAL PER LOG ANALYSIS. AS THIS CLINICAL INFORMATION WAS NOT PROVIDED, THE TRUE ROOT CAUSE OF ACCESS SITE BLEEDING, STROKE/CVA, AND VASCULAR DAMAGE - PERIPHERAL VESSEL, COULD NOT BE DETERMINED. B2 REVISED TO REFLECT THE PATIENT OUTCOME OF DEATH. B5 REVISED TO REFLECT THE PATIENT OUTCOME OF DEATH. D6B EXPLANT DATE ADDED. H1 TYPE OF REPORTABLE EVENT REVISED TO REFLECT THE PATIENT OUTCOME OF DEATH. H6 HEALTH EFFECT - IMPACT CODE 1802 DEATH ADDED. Manufacturer Narrative: B.1 REVISED TO REFLECT BOTH ADVERSE EVENT AND PRODUCT PROBLEM DUE TO ADDITIONAL INFORMATION RECEIVED FROM THE CLINICAL SITE. B.5 ADDITIONAL INFORMATION WAS RECEIVED FROM THE CLINICAL SITE. H.6 ADDED DUE TO MEDICAL DEVICE PROBLEM CODE ADDITIONAL INFORMATION RECEIVED FROM THE CLINICAL SITE. E.4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26865 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21659846&pc=DTB | 2017865-2025-27990 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-03-20 04:00 | DTB | BIPOLAR VENTRICULAR ENDOCARDIAL LEAD | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Unspecified Infection; Vascular Dissection; Thrombosis/Thrombus | K900550 |  | 1 | Event Description: VOLUNTARY MEDWATCH FORMS MW5166910, MW5166911 RECEIVED STATE: IT WAS REPORTED THAT THE PATIENT IMPLANTED WITH THIS (B)(6) PACEMAKER AND FOUR NON-(B)(6) LEADS (TWO CAPPED/ABANDONED AND TWO ACTIVE) UNDERWENT A LEAD EXTRACTION PROCEDURE DUE TO VEGETATION ON THE LEADS. UPON REMOVAL OF THE ABANDONED RIGHT ATRIAL (RA) LEAD MODEL 1018T, THE PATIENT'S BLOOD PRESSURE DROPPED AND AN EMERGENCY THORACOTOMY WAS PERFORMED. THE PATIENT HAD A TEAR IN THE RIGHT ATRIUM AND SUPERIOR VENA CAVA (SVC). THE PATIENT DEVELOPED DISSEMINATED INTRAVASCULAR COAGULATION (DIC), THE DAMAGED AREAS OF THE HEART COULD NOT BE REPAIRED, AND THE PATIENT DIED ON THE OPERATING TABLE. THE EXPLANTED PACEMAKER WAS PLANNED TO BE RETURNED FOR DISPOSAL. THIS REPORT REFLECTS INFORMATION RECEIVED BY FDA IN THE FORM OF A NOTIFICATION PER 803.22 (B)(2). Manufacturer Narrative: RELATED VOLUNTARY MEDWATCH FORM RECEIVED: MW5166910, MW5166911. Manufacturer Narrative: CORRECTION: SECTION E2 - "HEALTH PROFESSIONAL?" SHOULD HAVE BEEN "NO INFORMATION" CORRECTION: CUSTOMER SECTION SHOULD HAVE BEEN LEFT BLANK AS THE ORIGINAL REPORT SOURCE IS NOT KNOWN AND THE INFORMATION WAS RECEIVED VIA THE FDA. NOTE: HOSPITAL FACILITY, PATIENT DETAILS AND SPECIFIC PRODUCT DETAILS WERE NOT PROVIDED. FURTHER INFORMATION WAS NOT AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21659869&pc=LYJ | 1644487-2025-00315 | 2024-09-04 04:00 | Death | CYBERONICS - HOUSTON | 2025-03-20 04:00 | LYJ | PULSE GEN MODEL 1000 | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P970003 |  | 1 | Event Description: IT WAS REPORTED THAT THAT THE PATIENT HAD PASSED AWAY, THE CAUSE OF DEATH WAS INDICATED TO BE DUE TO RESPIRATORY FAILURE. THE VNS DEVICE WAS NOT EXPLANTED. NO OTHER RELEVANT INFORMATION HAS BEEN RECEIVED TO DATE. Manufacturer Narrative: LIVANOVA USA, INC. SUBMITS THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION, BASED ON INFORMATION THAT LIVANOVA HAS OBTAINED, BUT MAY NOT HAVE BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION, OR A CONCLUSION BY FDA OR ANYONE ELSE, THAT THE DEVICE, LIVANOVA, OR LIVANOVA'S EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY; DEFECTS¿ OR MALFUNCTIONS¿. THESE WORDS ARE INCORPORATED INTO THE FDA 3500A MEDWATCH FORM BY THE FDA, AND LIVANOVA OBJECTS TO THEIR USE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21659908&pc=NVN | 2017865-2025-27998 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-03-20 04:00 | NVN | FAST-PASS ACE | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Unspecified Infection; Vascular Dissection; Thrombosis/Thrombus |  |  | 1 | Event Description: VOLUNTARY MEDWATCH FORMS MW5166910; MW5166911 RECEIVED STATE: IT WAS REPORTED THAT THE PATIENT IMPLANTED WITH THIS (B)(6) PACEMAKER AND FOUR NON-(B)(6) LEADS (TWO CAPPED/ABANDONED AND TWO ACTIVE) UNDERWENT A LEAD EXTRACTION PROCEDURE DUE TO VEGETATION ON THE LEADS. UPON REMOVAL OF THE ABANDONED RIGHT ATRIAL (RA) LEAD MODEL 1018T, THE PATIENT'S BLOOD PRESSURE DROPPED AND AN EMERGENCY THORACOTOMY WAS PERFORMED. THE PATIENT HAD A TEAR IN THE RIGHT ATRIUM AND SUPERIOR VENA CAVA (SVC). THE PATIENT DEVELOPED DISSEMINATED INTRAVASCULAR COAGULATION (DIC), THE DAMAGED AREAS OF THE HEART COULD NOT BE REPAIRED, AND THE PATIENT DIED ON THE OPERATING TABLE. THE EXPLANTED PACEMAKER WAS PLANNED TO BE RETURNED FOR DISPOSAL. Manufacturer Narrative: RELATED VOLUNTARY MEDWATCH FORM RECEIVED:MW5166910, MW5166911. Manufacturer Narrative: CORRECTION: SECTION E2: "HEALTH PROFESSIONAL?" SHOULD HAVE BEEN "NO INFORMATION". CORRECTION: CUSTOMER SECTION SHOULD HAVE BEEN LEFT BLANK AS THE ORIGINAL REPORT SOURCE IS NOT KNOWN AND THE INFORMATION WAS RECEIVED VIA THE FDA. NOTE: HOSPITAL FACILITY, PATIENT DETAILS AND SPECIFIC PRODUCT DETAILS WERE NOT PROVIDED. FURTHER INFORMATION WAS NOT AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21639303&pc=NPU | 2015691-2025-01832 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-19 04:00 | NPU | COMMANDER DELIVERY SYSTEM UNKNOWN | Adverse Event Without Identified Device or Use Problem | Cardiac Perforation | P140031 | E2016006 | 2 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR MITRAL DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 2 MITRAL CARDIAC PERFORATION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 60 - 79 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 2 FEMALE(S) AND 0 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 2 CARDIAC PERFORATION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IN THE MITRAL POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 0 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS COMMANDER DELIVERY SYSTEM ARE: (B)(4). PER THE INSTRUCTIONS FOR USE, CARDIOVASCULAR INJURIES SUCH AS PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, MYOCARDIAL OR VALVULAR STRUCTURES ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE AND MAY REQUIRE INTERVENTION. THERE ARE SEVERAL POTENTIAL ETIOLOGIES FOR CARDIAC PERFORATION DURING A TRANSCATHETER VALVE REPLACEMENT PROCEDURE INCLUDING PERFORATION BY THE GUIDEWIRE, THE DELIVERY SYSTEM, OR THE TRANSVENOUS PACER (TVP) LEAD. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE TRANSCATHETER HEART VALVES. TRAINING INCLUDES PROPER GUIDEWIRE POSITIONING, FIXATION OF THE TVP TO PREVENT VENTRICLE PERFORATION, AND CAREFUL MANIPULATION OF DEVICES. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21639304&pc=NPU | 2015691-2025-01826 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-19 04:00 | NPU | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Obstruction/Occlusion | P140031 | E2016006 | 1 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR MITRAL DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 1 MITRAL LEFT VENTRICULAR OUTFLOW TRACT OBSTRUCTION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 86 - 86 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 FEMALE(S) AND 0 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 1 LEFT VENTRICULAR OUTFLOW TRACT OBSTRUCTION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IN THE MITRAL POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 0 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE ARE: (B)(4). PER THE INSTRUCTIONS FOR USE, VALVE MALPOSITION REQUIRING INTERVENTION IS A POTENTIAL ADVERSE EVENT ASSOCIATED WITH THE USE OF TRANSCATHETER HEART VALVES. OBSTRUCTION OF THE LEFT VENTRICULAR OUTFLOW TRACT CAN BE CAUSED BY PATIENT FACTORS (ANTERIOR MITRAL LEAFLET PROTRUDING INTO THE LVOT, SEPTAL BULGE) OR PROCEDURAL FACTORS (POSITIONING OF THE VALVE FRAME WITHIN THE ANNULUS). DEPENDING ON THE DEGREE OF OBSTRUCTION, CARDIAC OUTPUT AND HEMODYNAMIC STABILITY MAY BE AFFECTED. INACCURATE DEPLOYMENT IS GENERALLY A RESULT OF USE ERROR OR A COMBINATION OR PATIENT AND PROCEDURAL FACTORS. IN SOME CASES, LVOT OBSTRUCTION COULD RESULT IN CLINICALLY SIGNIFICANT HEMODYNAMIC COMPROMISE THAT MAY REQUIRE EXPLANTATION OF THE THV WITH SURGICAL CORRECTION. THE THV TRAINING MANUALS INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN THV. TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS AND PROCTORED PROCEDURES. THE CORRECT ALIGNMENT AND POSITIONING OF THE DEVICE AT THE POINT OF DEPLOYMENT IS EMPHASIZED AS A KEY FACTOR TO THE PLACEMENT AND FIXATION OF THE DEVICE. OPERATORS ARE ALSO INSTRUCTED TO USE FLUOROSCOPY AS THE PRIMARY METHOD OF VISUALIZATION FOR POSITIONING AND DEPLOYMENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21639305&pc=DYB | 2015691-2025-01827 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-19 04:00 | DYB | EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | Adverse Event Without Identified Device or Use Problem | Retroperitoneal Hemorrhage | K200258 | E2016006 | 2 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR MITRAL DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 2 MITRAL BLEEDING - RETROPERITONEAL DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 76 - 80 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 FEMALE(S) AND 1 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 2 BLEEDING - RETROPERITONEAL DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IN THE MITRAL POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 2 DAY(S). DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). THE INSTRUCTIONS FOR USE (IFU) LIST HEMORRHAGE REQUIRING TRANSFUSION OR INTERVENTION AS A POTENTIAL RISK ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING POTENTIAL ACCESS COMPLICATIONS ASSOCIATED WITH STANDARD CARDIAC CATHETERIZATION. THE MAJORITY OF TRANSFEMORAL ACCESS RELATED BLEEDING COMPLICATIONS ARE RELATED TO DISEASED ILEO-FEMORAL VESSELS AND/OR PROCEDURAL TECHNIQUE DURING THE INSERTION OR REMOVAL OF THE SHEATH AND/OR DILATORS. THERE ARE CASES WHERE THE BLEEDING IS SIGNIFICANT AND MORE COMPLEX INTERVENTION OR TRANSFUSION IS REQUIRED TO TREAT/PREVENT A PERMANENT INJURY FROM OCCURRING. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE VALVE CAN BE DELIVERED TRANSFEMORALLY. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE; HOWEVER, A TRANSFUSION OF BLOOD WAS REPORTED IN THE REGISTRY. BY VARC DEFINITION, THE MEET CRITERIA FOR MAJOR VASCULAR COMPLICATION; THEREFORE, THE EVENT IS BEING REPORTED. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21639306&pc=NPU | 2015691-2025-01825 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-19 04:00 | NPU | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P140031 | E2016006 | 1 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR MITRAL DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 1 MITRAL DEVICE MIGRATION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 60 - 60 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 FEMALE(S) AND 0 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 1 DEVICE MIGRATION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IN THE MITRAL POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 0 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE ARE: (B)(4). PER THE INSTRUCTIONS FOR USE, VALVE MIGRATION REQUIRING INTERVENTION IS A POTENTIAL ADVERSE EVENT ASSOCIATED WITH THE USE OF TRANSCATHETER HEART VALVES. VALVE MIGRATION RESULTS WHEN FORCES ACTING ON THE THV OVERCOME THE STRENGTH OF ATTACHMENT OF THE VALVE TO THE SITE OF DEPLOYMENT. STENT VALVES ARE SUBJECTED TO ANTEGRADE EJECTION FORCES DURING SYSTOLE. INCORRECT BIOPROSTHETIC VALVE SIZING AND INCOMPLETE FRAME EXPANSION CAN CONTRIBUTE TO VALVE MIGRATION. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE THV. TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS AND PROCTORED PROCEDURES. THE CORRECT SIZING, ALIGNMENT AND POSITIONING OF THE DEVICE ARE EMPHASIZED AS KEY FACTORS TO THE PLACEMENT AND FIXATION OF THE DEVICE. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21639307&pc=DYB | 2015691-2025-01830 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-19 04:00 | DYB | EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | K200258 | E2016006 | 1 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR MITRAL DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 1 MITRAL BLEEDING - ACCESS SITE DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 72 - 72 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 FEMALE(S) AND 0 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 1 BLEEDING - ACCESS SITE DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IN THE MITRAL POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 1 DAY(S). DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). THE INSTRUCTIONS FOR USE (IFU) LIST BLEEDING REQUIRING TRANSFUSION OR INTERVENTION AS A POTENTIAL RISK ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING POTENTIAL ACCESS COMPLICATIONS ASSOCIATED WITH STANDARD CARDIAC CATHETERIZATION. THE MAJORITY OF TRANSFEMORAL ACCESS RELATED BLEEDING COMPLICATIONS ARE RELATED TO DISEASED ILEO-FEMORAL VESSELS AND/OR PROCEDURAL TECHNIQUE DURING THE INSERTION OR REMOVAL OF THE SHEATH AND/OR DILATORS. THERE ARE CASES WHERE THE BLEEDING IS SIGNIFICANT AND MORE COMPLEX INTERVENTION OR TRANSFUSION IS REQUIRED TO TREAT/PREVENT A PERMANENT INJURY FROM OCCURRING. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE VALVE CAN BE DELIVERED TRANSFEMORALLY. IN THIS CASE, SPECIFIC PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE; HOWEVER, A TRANSFUSION OF BLOOD WAS REPORTED IN THE REGISTRY. BY VARC DEFINITION, THE MEET CRITERIA FOR MAJOR VASCULAR COMPLICATION; THEREFORE, THE EVENT IS BEING REPORTED. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21639308&pc=DYB | 2015691-2025-01831 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-19 04:00 | DYB | EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | Adverse Event Without Identified Device or Use Problem | Hematoma | K200258 | E2016006 | 1 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR MITRAL DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 1 MITRAL BLEEDING - HEMATOMA AT ACCESS SITE DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 76 - 76 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 0 FEMALE(S) AND 1 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 1 BLEEDING - HEMATOMA AT ACCESS SITE DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IN THE MITRAL POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 1 DAY(S). DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). THE INSTRUCTIONS FOR USE (IFU) LIST HEMORRHAGE REQUIRING TRANSFUSION OR INTERVENTION AS A POTENTIAL RISK ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING POTENTIAL ACCESS COMPLICATIONS ASSOCIATED WITH STANDARD CARDIAC CATHETERIZATION. THE MAJORITY OF TRANSFEMORAL ACCESS RELATED BLEEDING COMPLICATIONS ARE RELATED TO DISEASED ILEO-FEMORAL VESSELS AND/OR PROCEDURAL TECHNIQUE DURING THE INSERTION OR REMOVAL OF THE SHEATH AND/OR DILATORS. THERE ARE CASES WHERE THE BLEEDING IS SIGNIFICANT AND MORE COMPLEX INTERVENTION OR TRANSFUSION IS REQUIRED TO TREAT/PREVENT A PERMANENT INJURY FROM OCCURRING. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE VALVE CAN BE DELIVERED TRANSFEMORALLY. IN THIS CASE, SPECIFIC PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE; HOWEVER, A TRANSFUSION OF BLOOD WAS REPORTED IN THE REGISTRY. BY VARC DEFINITION, THE MEET CRITERIA FOR MAJOR VASCULAR COMPLICATION; THEREFORE, THE EVENT IS BEING REPORTED. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21639309&pc=NPU | 2015691-2025-01823 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-19 04:00 | NPU | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Insufficient Information | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P140031 | E2016006 | 4 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR MITRAL DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 4 MITRAL CARDIAC SURGERY OR INTERVENTION - OTHER UNPLANNED DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 60 - 79 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 4 FEMALE(S) AND 0 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 4 CARDIAC SURGERY OR INTERVENTION - OTHER UNPLANNED DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IN THE MITRAL POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 4 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE ARE: (B)(4). THE INSTRUCTIONS FOR USE (IFU) LIST REOPERATION, INCLUDING EMERGENCY CARDIAC SURGERY, AS ADDITIONAL POTENTIAL RISKS ASSOCIATED WITH THE USE OF THE THV, DELIVERY SYSTEM, AND/OR ACCESSORIES. THIS EVENT TYPE (CARDIAC SURGERY OR INTERVENTION - OTHER UNPLANNED) IS DEFINED AS: THE PATIENT SUBSEQUENTLY UNDERWENT OTHER TYPES OF CARDIAC SURGERY OR CATH LAB INTERVENTIONS (NOT OTHERWISE SPECIFIED) THAT WAS UNPLANNED. THIS COULD INCLUDE REOPERATION FOR DAMAGE TO ADJACENT STRUCTURES, OR POTENTIALLY VALVE EXPLANT RELATED TO EMBOLIZATION, OR OTHER EVENT TYPES. CARDIOVASCULAR INJURY, INCLUDING PERFORATION OR DISSECTION OF VALVULAR STRUCTURES AND VALVE EMBOLIZATION REQUIRING INTERVENTION ARE KNOWN POTENTIAL COMPLICATIONS LISTED IN THE PRODUCT INSTRUCTIONS FOR USE (IFU). PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE TRANSCATHETER HEART VALVES (THV). TRAINING INCLUDES PROPER GUIDEWIRE POSITIONING, AND CAREFUL MANIPULATION OF DEVICES. PER THE THV TA TRAINING MANUALS, AFTER GAINING LV ACCESS WITH THE 0.035'' SOFT GUIDEWIRE, THE WIRE SHOULD BE JIGGLED UNDER TEE IMAGING OF THE MITRAL VALVE. AN INCREASE IN MR SUGGESTS WIRE ENTANGLEMENT IN THE MITRAL SUB-VALVULAR APPARATUS. IF ENTANGLEMENT IS SUSPECTED, THE GUIDEWIRE SHOULD BE COMPLETELY REMOVED FROM THE VENTRICLE; THE OPERATOR SHOULD CHANGE DIRECTION OF THE NEEDLE AND REINSERT THE GUIDEWIRE INTO THE VENTRICLE, CHECKING AGAIN FOR WIRE ENTANGLEMENT. THE TF TRAINING MANUAL ALSO PROVIDES GUIDANCE FOR VALVE CROSSING AND WIRE EXCHANGE: EXCHANGE 0.035'' AMPLATZ EXTRA-STIFF WIRE WITH PRE-SHAPED DISTAL END IN LEFT VENTRICLE. ENSURE THE GUIDE CATHETER IS ADVANCED UPON WIRE EXCHANGE TO ENSURE EXCHANGE WIRE DOES NOT GET CAUGHT IN THE MITRAL CHORDAE. THERE ARE MULTIPLE PATIENT AND PROCEDURAL FACTORS THAT ALONE OR IN COMBINATION CAN CAUSE OR CONTRIBUTE TO VENTRICULAR EMBOLIZATION, INCLUDING IMPROPER POSITIONING PRIOR TO DEPLOYMENT, POOR IMAGE INTENSIFIER ANGLE, POOR COAXIAL ALIGNMENT OF THE VALVE/DELIVERY SYSTEM, RAPID DEPLOYMENT, RELEASE OF STORED TENSION DURING DEPLOYMENT, AND MOVEMENT OF THE DELIVERY SYSTEM BY THE OPERATOR. THE THV TRAINING MANUALS INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN THV. TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS AND PROCTORED PROCEDURES. THE CORRECT ALIGNMENT AND POSITIONING OF THE DEVICE AT THE POINT OF DEPLOYMENT IS EMPHASIZED AS A KEY FACTOR TO THE PLACEMENT AND FIXATION OF THE DEVICE. OPERATORS ARE ALSO INSTRUCTED TO USE FLUOROSCOPY AS THE PRIMARY METHOD OF VISUALIZATION FOR POSITIONING AND DEPLOYMENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21639318&pc=NPU | 2015691-2025-01833 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-19 04:00 | NPU | COMMANDER DELIVERY SYSTEM UNKNOWN | Insufficient Information | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P140031 | E2016006 | 1 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR MITRAL DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 1 MITRAL TRANSSEPTAL COMPLICATION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 72 - 72 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 FEMALE(S) AND 0 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 1 TRANSSEPTAL COMPLICATION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IN THE MITRAL POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 0 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS COMMANDER DELIVERY SYSTEM ARE: (B)(4). SPECIFIC DETAILS ARE NOT AVAILABLE TO DETERMINE IF THE EVENT REPRESENTS INJURY TO THE CARDIAC STRUCTURES RELATED TO NAVIGATION THROUGH THE ATRIAL SEPTUM, NEED FOR DEPLOYMENT OF A SEPTAL CLOSURE DEVICE, OR SOMETHING ELSE. PER THE INSTRUCTIONS FOR USE (IFU) CARDIOVASCULAR INJURY, SUCH AS PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, MYOCARDIUM, OR VALVULAR STRUCTURES, IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH TRANSCATHETER HEART VALVES. TRANSEPTAL APPROACHES ALLOW TRANSFEMORAL (VENOUS) ACCESS TO THE MITRAL VALVE BY PUNCTURING THE ATRIAL SEPTUM AND ADVANCING THE DELIVERY SYSTEM THROUGH THE OPENING. USUALLY, THE SEPTOSTOMY CLOSES ON ITS OWN OVER TIME AND DOES NOT REQUIRE TREATMENT. IN SOME CASES, THE INTERVENTIONALIST WILL ELECT TO CLOSE THE CREATED ATRIAL SEPTAL DEFECT (ASD) DURING THE INITIAL PROCEDURE PERCUTANEOUSLY WITH A SEPTAL OCCLUDER. PERSISTENT IATROGENIC ATRIAL SEPTAL DEFECTS (IASD) ARE NOT UNCOMMON AND MAY BE ASSOCIATED WITH THE USE OF LARGER SHEATHS. THESE MAY REQUIRE PLACEMENT OF A CLOSURE DEVICE IN A REPEAT PROCEDURE. THE RELATIONSHIP OF THE DEVICE TO THE EVENT IS NOT KNOWN. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21639319&pc=DYB | 2015691-2025-01829 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-19 04:00 | DYB | EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | Insufficient Information | Unspecified Vascular Problem | K200258 | E2016006 | 2 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR MITRAL DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 2 MITRAL VASCULAR SURGERY OR INTERVENTION - UNPLANNED DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 60 - 65 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 2 FEMALE(S) AND 0 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 2 VASCULAR SURGERY OR INTERVENTION - UNPLANNED DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IN THE MITRAL POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 12 DAY(S). DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). ACCORDING TO THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR COMPLICATIONS, INCLUDING PERFORATION OR DISSECTION OF VESSELS WHICH MAY REQUIRE INTERVENTION, ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE TRANSFEMORAL TRANSCATHETER VALVE REPLACEMENT PROCEDURE. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, VASCULAR COMPLICATIONS ARE A WELL-RECOGNIZED COMPLICATION OF THE TRANSFEMORAL TRANSCATHETER VALVE REPLACEMENT PROCEDURE IN THIS ELDERLY POPULATION WITH MULTIPLE CO-MORBIDITIES. EDWARDS HAS REVIEWED MANY REPORTS, INCLUDING SCREENING DATA RECORDS AND SOURCE DOCUMENTATION OF VASCULAR COMPLICATIONS, AND HAS FOUND THAT THE ROOT CAUSE IS TYPICALLY RELATED TO A COMBINATION OF VESSEL SIZE, TORTUOSITY, AND CALCIFICATIONS. ALTHOUGH THE INCIDENCE IS DECREASING WITH SMALLER SHEATH/DELIVERY SYSTEM SIZES AND PHYSICIAN EXPERIENCE, THERE WILL CONTINUE TO BE CASES IN WHICH VASCULAR COMPLICATIONS WILL OCCUR. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE TRAINING MANUAL INSTRUCTS THE OPERATOR ON PROPER SHEATH INSERTION AND WITHDRAWAL TECHNIQUES, INCLUDING PRE-DILATING THE VESSEL WITH THE EDWARDS DILATORS, AS NEEDED. IT ALSO NOTES THAT CALCIFICATION MAY REDUCE LUMEN DIAMETER AND LIMIT OR PREVENT TRANSFEMORAL PASSAGE OF THE DEVICES. THE IFU CONTRAINDICATES PATIENTS WITH VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE VASCULAR INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE THV CAN BE DELIVERED THROUGH THE PLANNED ACCESS SITE. ASSESSMENT OF LOCATION AND AMOUNT OF CIRCUMFERENTIAL CALCIUM WILL AID IN DETERMINING AREAS OF REDUCED VESSEL DIAMETERS. THE OPERATORS ARE TRAINED TO MEASURE MINIMUM VESSEL DIAMETER TAKING CALCIUM INTO ACCOUNT. THE PHYSICIAN TRAINING MANUAL ALSO LISTS THE MINIMUM RECOMMENDED VESSEL SIZE FOR EACH SIZE DEVICE. DESPITE THE BEST SCREENING TOOLS, A SMALL PERCENTAGE OF PATIENTS WILL HAVE VESSELS THAT ARE NOT AMENABLE TO PERCUTANEOUS VALVE DELIVERY OR WHERE INCREASED RESISTANCE IS ENCOUNTERED DURING INSERTION OF DEVICES. IN MANY CASES, THE VESSEL MINIMUM LUMINAL DIAMETER (MLD) MAY BE BORDERLINE OR BELOW THE INDICATED SIZE. IN ADDITION, SIGNIFICANT CALCIFICATION AND/OR TORTUOSITY, NOT ALWAYS APPRECIABLE ON IMAGING, COULD BE CONTRIBUTING FACTORS TO THE EVENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21639320&pc=NPU | 2015691-2025-01824 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-19 04:00 | NPU | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Insufficient Information | Mitral Valve Insufficiency/ Regurgitation | P140031 | E2016006 | 4 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR MITRAL DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 4 MITRAL REINTERVENTION - MITRAL VALVE DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 42 - 83 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 4 FEMALE(S) AND 0 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 4 REINTERVENTION - MITRAL VALVE DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IN THE MITRAL POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 13.5 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE ARE: (B)(4). INTRA-PROCEDURAL AND IN-HOSPITAL MITRAL VALVE RE-INTERVENTION WILL TYPICALLY RESULT FROM VALVE MALPOSITION OR REGURGITATION (PVL OR CENTRAL, INCLUDING LEAFLET RESTRICTION). THESE CONDITIONS ARE KNOWN POTENTIAL RISKS ASSOCIATED WITH THE USE OF THE THV, DELIVERY SYSTEM, AND/OR ACCESSORIES. THERE ARE MULTIPLE PATIENT AND PROCEDURAL FACTORS THAT ALONE OR IN COMBINATION CAN CAUSE OR CONTRIBUTE TO VALVE MALPOSITION, INCLUDING IMPROPER POSITIONING PRIOR TO DEPLOYMENT, POOR IMAGE INTENSIFIER ANGLE, POOR COAXIAL ALIGNMENT OF THE VALVE AND DELIVERY SYSTEM, CHARACTERISTICS OF THE PRE-EXISTING VALVE OR RING, LOSS OF PACING CAPTURE, RAPID DEPLOYMENT, RELEASE OF STORED TENSION DURING DEPLOYMENT, AND MOVEMENT OF THE DELIVERY SYSTEM BY THE OPERATOR. VALVE MALPOSITION HAS THE POTENTIAL TO CONTRIBUTE TO SUBOPTIMAL COAPTATION OF THE THV VALVE LEAFLETS AND CAUSE CENTRAL INSUFFICIENCY; IT CAN LEAD TO MIGRATION OR EMBOLIZATION OF THE PROSTHESIS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE THV. TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS AND PROCTORED PROCEDURES. THE CORRECT ALIGNMENT AND POSITIONING OF THE DEVICE AT THE POINT OF DEPLOYMENT IS EMPHASIZED AS A KEY FACTOR TO THE PLACEMENT AND FIXATION OF THE DEVICE. OPERATORS ARE ALSO INSTRUCTED TO USE FLUOROSCOPY AS THE PRIMARY METHOD OF VISUALIZATION FOR POSITIONING AND DEPLOYMENT. MITRAL VALVE RE-INTERVENTION IN THE FOLLOW-UP PERIOD (IN THE ABSENCE OF PROSTHETIC CARDIAC VALVE THROMBOSIS) WILL TYPICALLY RESULT FROM ON-GOING OR WORSENING REGURGITATION, VALVE DEGENERATION RELATED TO THE FORMATION OF CALCIFICATION OR PANNUS, OR VALVE MIGRATION. THESE EVENTS ARE IDENTIFIED IN THE PRODUCT INSTRUCTIONS FOR USE (IFU) AS POTENTIAL RISKS ASSOCIATED WITH THE USE OF THE THV. REGURGITATION WHICH DEVELOPS PROGRESSIVELY OVER TIME CAN BE DUE TO SEVERAL ISSUES INCLUDING PATIENT RELATED FACTORS OR STRUCTURAL VALVE DETERIORATION, INCLUDING CALCIFICATION, NON-CALCIFIC DEGENERATION, LEAFLET THICKENING OR FIBROSIS, OR A COMBINATION OF THESE. REGURGITATION MAY ALSO DEVELOP PROGRESSIVELY IF HOST FIBROTIC TISSUE, OR PANNUS, GROWS ONTO THE BIOPROSTHETIC VALVE. PANNUS, A CAUSE OF NONSTRUCTURAL DYSFUNCTION, MAY INTERFERE WITH FUNCTIONALITY OF THE DEVICE BY RESTRICTING THE LEAFLET MOTION LEADING TO ABNORMAL COAPTATION. CALCIFICATION IS A WELL-RECOGNIZED FAILURE MODE OF BIOPROSTHETIC VALVES. THE MECHANISM OF CALCIFICATION OF BIOMATERIALS IS NOT COMPLETELY UNDERSTOOD BUT IS PROBABLY RELATED TO AN INABILITY OF THE NON-VIABLE CELLS TO MAINTAIN THEIR NORMALLY LOW INTRACELLULAR CONCENTRATION OF CALCIUM. ANOTHER POTENTIAL CONDITION THAT MAY LEAD TO RE-INTERVENTION IS PATIENT PROSTHESIS MISMATCH AND/OR UNDER-EXPANDED VALVES DUE TO THV RESTRICTION FROM THE EXISTING VALVE FRAME. INCORRECT BIOPROSTHETIC VALVE SIZING, AND INCOMPLETE FRAME EXPANSION, CAN CONTRIBUTE TO VALVE MIGRATION. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21639323&pc=DYB | 2015691-2025-01828 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-19 04:00 | DYB | EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | Adverse Event Without Identified Device or Use Problem | Unspecified Vascular Problem | K200258 | E2016006 | 2 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR MITRAL DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 2 MITRAL VASCULAR COMPLICATION - MAJOR DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 60 - 76 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 FEMALE(S) AND 1 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 2 VASCULAR COMPLICATION - MAJOR DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IN THE MITRAL POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 0.5 DAY(S). DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). ACCORDING TO THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR COMPLICATIONS, INCLUDING PERFORATION OR DISSECTION OF VESSELS WHICH MAY REQUIRE INTERVENTION, ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE TRANSFEMORAL TRANSCATHETER VALVE REPLACEMENT PROCEDURE. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, VASCULAR COMPLICATIONS ARE A WELL-RECOGNIZED COMPLICATION OF THE TRANSFEMORAL PROCEDURE IN THIS ELDERLY POPULATION WITH MULTIPLE CO-MORBIDITIES. EDWARDS HAS REVIEWED MANY REPORTS, INCLUDING SCREENING DATA RECORDS AND SOURCE DOCUMENTATION OF VASCULAR COMPLICATIONS, AND HAS FOUND THAT THE ROOT CAUSE IS TYPICALLY RELATED TO A COMBINATION OF VESSEL SIZE, TORTUOSITY, AND CALCIFICATIONS. ALTHOUGH THE INCIDENCE IS DECREASING WITH SMALLER SHEATH/DELIVERY SYSTEM SIZES AND PHYSICIAN EXPERIENCE, THERE WILL CONTINUE TO BE CASES IN WHICH VASCULAR COMPLICATIONS WILL OCCUR. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE TRAINING MANUAL INSTRUCTS THE OPERATOR ON PROPER SHEATH INSERTION AND WITHDRAWAL TECHNIQUES, INCLUDING PRE-DILATING THE VESSEL WITH THE EDWARDS DILATORS, AS NEEDED. IT ALSO NOTES THAT CALCIFICATION MAY REDUCE LUMEN DIAMETER AND LIMIT OR PREVENT TRANSFEMORAL PASSAGE OF THE DEVICES. THE IFU CONTRAINDICATES PATIENTS WITH VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE VASCULAR INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE THV CAN BE DELIVERED THROUGH THE PLANNED ACCESS SITE. ASSESSMENT OF LOCATION AND AMOUNT OF CIRCUMFERENTIAL CALCIUM WILL AID IN DETERMINING AREAS OF REDUCED VESSEL DIAMETERS. THE OPERATORS ARE TRAINED TO MEASURE MINIMUM VESSEL DIAMETER TAKING CALCIUM INTO ACCOUNT. THE PHYSICIAN TRAINING MANUAL ALSO LISTS THE MINIMUM RECOMMENDED VESSEL SIZE FOR EACH SIZE DEVICE. DESPITE THE BEST SCREENING TOOLS, A SMALL PERCENTAGE OF PATIENTS WILL HAVE VESSELS THAT ARE NOT AMENABLE TO PERCUTANEOUS VALVE DELIVERY OR WHERE INCREASED RESISTANCE IS ENCOUNTERED DURING INSERTION OF DEVICES. IN MANY CASES, THE VESSEL MINIMUM LUMINAL DIAMETER (MLD) MAY BE BORDERLINE OR BELOW THE INDICATED SIZE. IN ADDITION, SIGNIFICANT CALCIFICATION AND/OR TORTUOSITY, NOT ALWAYS APPRECIABLE ON IMAGING, COULD BE CONTRIBUTING FACTORS TO THE EVENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21639551&pc=LJT | 3006260740-2025-01604 | 2025-03-10 04:00 | Death | C.R. BARD, INC. (BASD) -3006260740 | 2025-03-19 04:00 | LJT | POWERPORT ISP M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | Fracture | Unspecified Infection | K063377 |  | 1 | Event Description: IT WAS REPORTED THROUGH LITIGATION PROCESS THAT SOMETIME AFTER PORT PLACEMENT PROCEDURE, THE PORT CATHETER FRACTURED AND PATIENT DEVELOPED WITH UNSPECIFIED INFECTION. REPORTEDLY, THE PATIENT EXPIRED, AND THE CAUSE OF DEATH WAS NOT PROVIDED. Manufacturer Narrative: H10: DATE OF DEATH FOR THIS PATIENT WAS NOT PROVIDED, DATE OF DEATH UPDATED AS (B)(6) 1900 FOR THE MDR SUBMISSION REQUIREMENT. H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21640329&pc=DTZ | 8010762-2025-0000123 | 2025-03-16 04:00 | Death | MAQUET CARDIOPULMONARY GMBH | 2025-03-19 04:00 | DTZ | TUBING SET | Fluid/Blood Leak; Leak/Splash | Cardiac Arrest | K112360 |  | 1 | Event Description: THE EVENT OCCURRED IN CZECHIA DURING TREATMENT. IT WAS REPORTED THAT THERE WAS A BLOOD LEAK DIRECTLY FROM THE OXYGENATOR. THE CUSTOMER MANUALLY COMPRESSED THE LEAKAGE POINT AND REPLACED THE HLS SET. THE PATIENT HAD TO BE RESUSCITATED AND SURVIVED. A BLOOD TRANSFUSION WAS REQUIRED. AS THE PATIENT NEEDED A BLOOD TRANSFUSION AND REQUIRED RESUSCITATION, A REPORT IS REQUIRED. COMPLAINT ID# (B)(4).Event Description: COMPLAINT ID# (B)(4).Event Description: COMPLAINT ID (B)(4). Manufacturer Narrative: THE EVENT OCCURRED IN CZECHIA DURING TREATMENT. IT WAS REPORTED THAT THERE WAS A BLOOD LEAK DIRECTLY FROM THE OXYGENATOR. THE CUSTOMER MANUALLY COMPRESSED THE LEAKAGE POINT AND REPLACED THE HLS SET. THE PATIENT HAD TO BE RESUSCITATED AND SURVIVED. A BLOOD TRANSFUSION WAS REQUIRED. AS THE PATIENT NEEDED A BLOOD TRANSFUSION, REQUIRED RESUSCITATION AND LATER PASSED AWAY, A REPORT IS REQUIRED. NEW INFORMATION WAS RECEIVED THAT THE PATIENT LOST 1000ML OF BLOOD. THE PATIENT UNDERWENT A TWO MINUTE CARDIOPULMONARY RESUSCITATION DURING THE ECMO CIRCUIT EXCHANGE DUE TO SEVERE CARDIAC DYSFUNCTION SUBSEQUENT TO EXTENSIVE MYOCARDIAL INFARCTION. NEW INFORMATION WAS RECEIVED ON 2025-04-01 THAT THE PATIENT PASSED DUE TO LATE COMPLICATIONS, DAYS AFTER WEANING FROM THE ECMO. THE AFFECTED PRODUCT WAS INVESTIGATED BY THE GETINGE LABORATORY ON 2025-05-15. THE REPORTED FAILURE COULD BE CONFIRMED, BUT THE EXACT ROOT CAUSE COULD NOT BE DETERMINED BUT MOST PROBABLY RELATED TO A LEAKAGE IN THE INTERNAL PRESSURE SENSOR. A MEDICAL REVIEW WAS PERFORMED BY GETINGE MEDICAL AFFAIRS ON 2025-05-28 WITH FOLLOWING CONCLUSION: "ACCORDING TO THE COMPLAINT NARRATIVE, THE INCIDENT INVOLVES A 54-YEAR-OLD MALE PATIENT ON VA-ECMO SUPPORT WITH A CARDIOHELP SYSTEM AND HLS SET ADVANCED 7.0, WHO EXPERIENCED A SUDDEN BLOOD LEAK FROM THE OXYGENATOR APPROXIMATELY THREE HOURS AFTER THERAPY INITIATION. THIS LED TO A SIGNIFICANT DROP IN ARTERIAL PRESSURE AND REQUIRED URGENT CIRCUIT REPLACEMENT. MANUAL COMPRESSION WAS APPLIED TO THE LEAK SITE, AND THE HLS SET WAS REPLACED. THE PATIENT UNDERWENT TWO MINUTES OF CPR DUE TO SEVERE CARDIAC DYSFUNCTION AND REQUIRED BLOOD TRANSFUSION. THE PATIENT INITIALLY SURVIVED BUT LATER EXPIRED FROM COMPLICATIONS AFTER ECMO WEANING. INVESTIGATION CONFIRMED THE COMPLAINT. BLOOD RESIDUES AND LEAKAGE WERE LOCALIZED TO THE INTERNAL PRESSURE SENSOR AREA. THE HLS SET WAS INSPECTED PRIOR TO USE WITH NO VISIBLE ISSUES, AND NO LEAKS WERE DETECTED DURING PRIMING. THE LEAK INVOLVED BLOOD ONLY. IMMEDIATE CLINICAL RESPONSE PREVENTED FURTHER HARM. IT IS IMPORTANT TO NOTE THAT DURING THE TECHNICAL INVESTIGATION, ONLY DROPLET FORMATION WAS OBSERVED AT THE BLOOD INLET COVER AS EVIDENCE OF LEAKAGE. HOWEVER, THE COMPLAINT NARRATIVE REPORTS AN ESTIMATED TOTAL BLOOD LOSS OF APPROXIMATELY 1000 ML, WITHOUT SPECIFYING THE EXACT TIME FRAME OVER WHICH THIS LOSS OCCURRED. THIS DISCREPANCY HIGHLIGHTS A GAP IN THE INFORMATION REGARDING THE PROGRESSION AND SEVERITY OF THE LEAKAGE. A PRODUCT MALFUNCTION WAS CONFIRMED THROUGH THE INVESTIGATION, SPECIFICALLY RELATED TO DEFECTS IN THE INTERNAL PRESSURE SENSOR. ACCORDING TO THE CUSTOMER¿S STATEMENT, ¿THE PATIENT INITIALLY SURVIVED BUT LATER EXPIRED FROM COMPLICATIONS AFTER ECMO WEANING.¿ THIS SUGGESTS THAT THE PATIENT¿S EVENTUAL EXPIRATION WAS LIKELY NOT DIRECTLY CAUSED BY THE HLS MALFUNCTION. NONETHELESS, THE URGENT NEED TO REPLACE THE HLS SET AND THE RESUSCITATION EFFORTS DESCRIBED DURING THIS PERIOD UNDERSCORE THE POTENTIAL FOR SERIOUS HARM HAD THE INTERVENTION BEEN DELAYED OR UNSUCCESSFUL. IN SUMMARY, WHILE THE HLS MALFUNCTION WAS VERIFIED BY INTERNAL GETINGE TESTING. MOREOVER, THE REPORTED EVENT LIKELY CONTRIBUTED TO THE CRITICAL CLINICAL OCCURRENCE, SPECIFICALLY CPR AND BLOOD TRANSFUSION. THE PATIENT¿S EXPIRATION APPEARS TO BE UNRELATED TO THE DEVICE FAILURE ITSELF, HAVING OCCURRED AN UNDEFINED PERIOD AFTER THE EVENT. IT IS ASSUMED THAT THE LOSS OF EXTRACORPOREAL SUPPORT (I.E., EXPERIENCED DURING PRODUCT EXCHANGE) LIKELY PRECIPITATED CPR. FURTHER, THE LOSS OF BLOOD VOLUME APPEARS TO BE DIRECTLY RELATED TO THE ADMINISTRATION OF BLOOD PRODUCTS." ACCORDING TO THE INSTRUCTION FOR USE (IFU, HLS SET ADVANCED 6.0/7.0, HIT SET ADVANCED 5.0/7.0) IT IS STATED IN CHAPTER ¿SAFETY INSTRUCTIONS FOR CENTRIFUGAL PUMP¿ THAT LEAKS OR SCRATCHING NOISES IN THE CENTRIFUGAL PUMP CAN BE A SIGN THAT IT IS MALFUNCTIONING. MALFUNCTIONING CAN LEAD TO INADEQUATE PATIENT SUPPORT. FURTHERMORE, IT IS MENTIONED TO ALWAYS KEEP A REPLACEMENT HLS SET FOR THOSE SITUATIONS. IN CHAPTER ¿PRIMING THE SYSTEM¿ IT IS WRITTEN NOT TO PRE-PRIME THE CIRCUIT TOO EARLY AS THIS COULD LEAD TO LEAKS AND TO NOT USE A DEVICE IF THERE ARE ANY LEAKS. THE PRODUCTION RECORDS OF THE AFFECTED PRODUCT WERE REVIEWED ON 2025-06-02. ACCORDING TO THE FINAL TEST RESULTS, THE PRODUCT PASSED THE TESTS AS PER SPECIFICATIONS. PRODUCTION RELATED INFLUENCES ARE UNLIKELY. THE REVIEW OF SCRAP, REWORK, ENHANCEMENTS AND DESIGN CHANGES WERE REVIEWED AND NO ABNORMALITIES IN REGARDS TO THE REPORTED FAILURE WERE FOUND. IN ADDITION, A REVIEW FOR POTENTIAL FIELD ACTIONS AND CAPAS RELATED TO THE FAILURE AND PRODUCT IN THIS COMPLAINT WAS PERFORMED. THIS COMPLAINT IS NOT IN SCOPE OF ANY ONGOING FIELD ACTIONS AND/OR CAPAS. THE REVIEW OF THE NON-CONFORMITIES HAS BEEN PERFORMED FOR THE REVIEWED TIME PERIOD. IT DOES NOT SHOW ANY NON-CONFORMITY IN REGARDS TO THE REPORTED FAILURE. BASED ON THE RESULTS THE REPORTED FAILURE "LEAKAGE ON THE OXYGENATOR" COULD BE CONFIRMED. THE PATIENT¿S EXPIRATION APPEARS TO BE UNRELATED TO THE DEVICE FAILURE ITSELF, HAVING OCCURRED AN UNDEFINED PERIOD AFTER THE EVENT. THE CUSTOMER WILL BE INFORMED ABOUT THE RESULTS BY THE GETINGE SALES AND SERVICE UNIT. THE OCCURRENCE RATE WAS CALCULATED FOR THE REPORTED ISSUE AND IT WAS DETERMINED THAT THIS IS NOT A SYSTEMIC ISSUE. THEREFORE, NO REMEDIAL ACTION IS REQUIRED. THIS COMPLAINT WAS FOUND IN THE DATABASE OF CUSTOMER COMPLAINTS FOR THE HLS SET WITH ARTICLE (B)(4) AS A SINGLE EVENT (TIMEFRAME FROM 2024-03-17 TILL 2025-03-17). ACCORDING TO THE RISK REVIEW THE REPORTED FAILURE IS BELOW THE ACCEPTANCE THRESHOLD ACCORDING TO THE RISK MANAGEMENT PLAN OF THE HLS SET. THE OCCURRENCE RATE RELATED TO THE REPORTED ISSUE IS CURRENTLY BEING MONITORED AS PART OF MAQUET CARDIOPULMONARY¿ S TRENDING PROGRAM AND ADDITIONAL INVESTIGATIONS OR CORRECTIONS WILL BE IMPLEMENTED IN CASE OF ADVERSE TRENDING. Manufacturer Narrative: THE PATIENT LOST 1000ML OF BLOOD. THE PATIENT UNDERWENT A TWO MINUTE CARDIOPULMONARY RESUSCITATION DURING THE ECMO CIRCUIT EXCHANGE DUE TO SEVERE CARDIAC DYSFUNCTION SUBSEQUENT TO EXTENSIVE MYOCARDIAL INFARCTION. NEW INFORMATION WAS RECEIVED ON 2025-04-01 THAT THE OUTCOME OF THE PATIENT WAS DEATH DUE TO LATE COMPLICATIONS, DAYS AFTER WEANING OFF FROM THE ECMO. A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. NOTE: THIS EVENT OCCURRED ON THE CZECHIA MARKET. IT IS A SIGNIFICANTLY SIMILAR DEVICE TO "HLS SET¿ WHICH IS SOLD IN THE USA UNDER PREMARKET SUBMISSION NUMBER K112360. FOR SECTION D4 ALL AVAILABLE IDENTIFYING INFORMATION IS FOR THE OUT OF THE US DEVICE, SUBJECTED TO THIS REPORT. THEREFORE, NO GUDID INFORMATION EXISTS. FURTHERMORE, UDI INFORMATION (PRIMARY DI NUMBER) PROVIDED IN D4 IS FOR HLS SET WITH CATALOG NUMBER 701069073. Manufacturer Narrative: A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. NOTE: NO UDI NUMBER HAS BEEN INCLUDED IN THE SECTION D4 UNIQUE DEVICE IDENTIFIER (UDI) SINCE THE LOT NUMBER OF THE AFFECTED SET WAS NOT PROVIDED. THEREFORE IT IS NOT POSSIBLE TO IDENTIFY THE SET LOT NUMBER OF THE DEVICE IN QUESTION AND THEREFORE ITS UDI NUMBER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21640421&pc=LJT | 3006260740-2025-01613 | 2025-03-12 04:00 | Death | C.R. BARD, INC. (BASD) -3006260740 | 2025-03-19 04:00 | LJT | POWERPORT M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | Adverse Event Without Identified Device or Use Problem | Thrombosis/Thrombus | K063377 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE LITIGATION PROCESS THAT SOMETIME AFTER PORT PLACEMENT PROCEDURE, THE PATIENT ALLEGEDLY DEVELOPED THROMBOSIS. REPORTEDLY, THE PATIENT EXPIRED AND THE CAUSE OF DEATH WAS NOT PROVIDED.Event Description: IT WAS REPORTED THROUGH THE LITIGATION PROCESS THAT SOMETIME AFTER PORT PLACEMENT PROCEDURE, THE PATIENT ALLEGEDLY DEVELOPED THROMBOSIS. REPORTEDLY, THE PATIENT EXPIRED AND THE CAUSE OF DEATH WAS NOT PROVIDED. Manufacturer Narrative: H10: DATE OF DEATH FOR THIS PATIENT WAS NOT PROVIDED, DATE OF DEATH UPDATED AS (B)(6) 1900 FOR THE MDR SUBMISSION REQUIREMENT. H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. Manufacturer Narrative: H11: DATE OF DEATH FOR THIS PATIENT WAS NOT PROVIDED, DATE OF DEATH UPDATED AS (B)(6) 1900 FOR THE MDR SUBMISSION REQUIREMENT. MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED, AND THIS LOT MET ALL RELEASE CRITERIA. INVESTIGATION SUMMARY: THE DEVICE WAS NOT RETURNED FOR EVALUATION. NO MEDICAL RECORDS WERE PROVIDED. THE REPORTED THROMBOSIS AND DEATH THEREFORE CANNOT BE CONFIRMED. NO CAUSAL RELATIONSHIP BETWEEN THE THROMBOSIS OR THE PATIENT DEATH HAS BEEN ESTABLISHED. BASED ON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE OF THE THROMBOSIS AND DEATH IS UNKNOWN. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. D4 (UNIQUE IDENTIFIER (UDI) #), G3, H4, H6 (COMPONENT, METHOD). SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21640938&pc=MVK | 3008642652-2025-02831 | 2025-03-10 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-03-19 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST.. THE PATIENT RECEIVED AN INAPPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 17:59:42 ON (B)(6) 2025. AT 06:52:55 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH CPR/MOTION ARTIFACT. AT 06:53:39, THE PATIENT RECEIVED THE INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON LIFE-SUSTAINING RHYTHM, WITH INTERMITTENT CARDIAC ACTIVITY. THE DEVICE WAS SHUT DOWN AT 07:02:41 ON 3/10/2025. Manufacturer Narrative: THE ELECTRODE BELT AND MONITOR HAVE BEEN RETURNED AND THE EVALUATION IS UNDERWAY. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21640950&pc=MVK | 3008642652-2025-02833 | 2025-02-21 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-03-19 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED AN APPROPRIATE TREATMENT AND AN INAPPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 19:25:02 ON (B)(6) 2025. AT 06:26:42 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF. AT 06:27:20, THE PATIENT RECEIVED THE APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS SINUS BRADYCARDIA @ 30 BPM WITH NSVT. AT 06:28:00, THE PATIENT RECEIVED THE INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS SINUS BRADYCARDIA @ 30 BPM WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS SINUS BRADYCARDIA @ 20 BPM WITH PVCS AND CPR/MOTION ARTIFACT. THE DEVICE WAS SHUT DOWN AT 06:28:24 ON 2/21/2025. Manufacturer Narrative: THE ELECTRODE BELT AND MONITOR HAVE BEEN RETURNED AND THE EVALUATION IS UNDERWAY. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21641315&pc=DWF | 9612164-2025-01394 | 2025-03-06 05:00 | Death | MEDTRONIC MEXICO S. DE R.L. DE CV | 2025-03-19 04:00 | DWF | BIO-MEDICUS NG ARTERIAL CANNULA | Crack; Fluid/Blood Leak | Hemorrhage/Bleeding; No Clinical Signs, Symptoms or Conditions | K142673 |  | 1 | Event Description: IT WAS REPORTED THAT DURING USE OF A BIO-MEDICUS CANNULA, THERE WAS A SUDDEN BREAK AT THE LATERAL BRANCH CIRCULATION. THE USE OF THE DEVICE WAS UNSPECIFIED. IT WAS UNKNOWN IF THERE WAS ANY PATIENT INVOLVEMENT OR COMPLICATIONS AS A RESULT OF THE EVENT. Manufacturer Narrative: ADDITIONAL INFO B5: MEDTRONIC RECEIVED ADDITIONAL INFORMATION THAT THE LEAK CAME FROM THE CANNULA. THE CANNULA WAS NOT HEATED OR COOLED PRIOR TO USE. THE AMOUNT OF BLOOD LOSS WAS TOO GREAT TO PROVIDE DATA. A TRANSFUSION WAS REQUIRED AS A RESULT OF THIS LEAK. THERE WERE 2 UNITS OF BLOOD (250-300ML) TRANSFUSED. MEDTRONIC RECEIVED ADDITIONAL INFORMATION THAT THE CAUSE OF DEATH AND IF PATIENT'S DEATH BELIEVED TO BE LINKED TO THE PERFORMANCE OF CANNULA ARE BOTH YET TO BE CLARIFIED. MEDTRONIC RECEIVED ADDITIONAL INFORMATION THAT THE CANNULA WAS IN USE FROM (B)(6) 2025 (25 DAYS). CORRECTION H6: UPDATED THE ANNEX A AND ANNEX E CODES CORRECTION B5: IT WAS REPORTED THAT THERE WAS A SUDDEN BREAK AT THE LATERAL BRANCH CIRCULATION INVOLVING THE BIO-MEDICUS CANNULA. THE CANNULA WAS NOT REPLACED TO COMPLETE THE PROCEDURE. THE DOCTOR ASKED THE FAMILY TO MAKE THE DECISION DIRECTLY. THE PATIENT PASSED AWAY. CORRECTION D4: UPDATED THE MODEL NUMBER, CATALOG NUMBER, LOT NUMBER, UDI, AND DEVICE EXPIRY DATE CORRECTION E2: UPDATED TO YES CORRECTION E3: OCCUPATION UPDATED TO PHYSICIAN. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION B5: MEDTRONIC RECEIVED ADDITIONAL INFORMATION THAT THE CUSTOMER USED A SYRINGE TO PLUG THE PORT WHEN IT BROKE. CORRECTION D3 (MFR NAME), D3.6 (POSTAL CODE), D4.5 (UNIQUE IDENTIFIER (UDI) #): THESE FIELDS HAVE BEEN UPDATED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21642012&pc=NPT | 2025587-2025-01980 | 2025-02-21 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-19 04:00 | NPT | EVOLUT FX PLUS VALVE | Activation Failure | Insufficient Information | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT PRIOR TO THE ATTEMPTED IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, A PRE-IMPLANT BALLOON AORTIC VALVULOP LASTLY WAS PERFORMED DUE TO A BICUSPID AORTIC VALVE. FOLLOWING INITIAL DEPLOYMENT, THE VALVE WAS RECAPTURED DUE TO UNDER-EXPANSION OF THE VALVE FRAME. THE VALVE WAS RE-DEPLOYED, BUT THE SAME ISSUE OCCURRED. A SECOND RECAPTURE WAS PERFORMED AND THE VALVE WAS WITHDRAWN FROM THE PATIENT. A SECOND BAV WAS PERFORMED USING A LARGER BALLOON. A NEW VALVE WAS OPENED; HOWEVER, PRIOR TO INSERTION OF THE SECOND VALVE, THE PATIENT BEGAN TO DECOMPENSATE. SUBSEQUENTLY, THE PATIENT PASSED AWAY. PER THE PHYSICIAN, THE CAUSE OF DEATH WAS RELATED TO THE PATIENT'S UNDERLYING MEDICAL HISTORY AND NOT THE DEVICE OR THE PROCEDURE.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT PER THE PHYSICIAN, THE CAUSE OF DEATH IS UNKNOWN. PER THE PHYSICIAN, THE PATIENT HAD A COMPLICATED MEDICAL HISTORY AND WAS VERY SICK PRIOR TO THE EMERGENT TRANSCATHETER BIOPROSTHETIC VALVE IMPLANT ATTEMPT, WHICH CONTRIBUTED TO PATIENT'S DEATH. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID {D-EVOLUTFX-34}; PRODUCT LOT/SERIAL NUMBER {UNKNOWN}; PRODUCT TYPE: {0195-HEART VALVES}; IMPLANT DATE {N/A}; EXPLANT DATE {N/A}. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21642284&pc=DWE | 9612164-2025-01398 | 2025-02-04 05:00 | Death | MEDTRONIC MEXICO | 2025-03-19 04:00 | DWE | CUSTOM TUBING PACK | Air/Gas in Device | Air Embolism; Hemorrhage/Bleeding | K171979 |  | 1 | Event Description: MEDTRONIC RECEIVED INFORMATION THAT DURING USE OF 3 CUSTOM TUBING PACKS, IT WAS REPORTED THE PLASTIC PUMP HOUSING THAT CONNECTS TO THE ABBOTT CENTROMAG VENOUS CANNULA CRACKED AND CAUSED BUBBLES TO ENTER THE CIRCUIT. CUSTOMER ATTEMPTED TO STOP EXSANGUINOUS DRAINAGE, AND PERFORMED CIRCUIT EXCHANGE, BUT PATIENT CODED DURING THIS EXCHANGE, CUSTOMER PLACED PATIENT BACK ON VENO-VENOUS EXTRACORPOREAL MEMBRANE OXYGENATION (VV ECMO) BUT THE BUBBLES WERE ALREADY IDENTIFIED IN THE CIRCUIT, WITH INABILITY TO GET BUBBLES OUT WITHOUT CAUSING ADDITIONAL HARM TO THE PATIENT. DESPITE ALL EFFORTS, THE PATIENT EXPIRED AT 00:05AM ON (B)(6)2025. THE DEVICE WAS REPLACED TO COMPLETE THE PROCEDURE. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21642293&pc=NVY | 2649622-2025-07653 | 2025-03-03 05:00 | Death | MPRI | 2025-03-19 04:00 | NVY | EPSILA EV¿ MRI SURESCAN¿ | Appropriate Term/Code Not Available | Pulmonary Embolism; Hypoxia; Thrombosis/Thrombus | P220012 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IS DECEASED. THE PATIENT DIED AFTER THE PLACEMENT OF AN EXTRAVASCULAR (EV) IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) SYSTEM. DEFIBRILLATION THRESHOLD (DFT) TESTING HAD BEEN COMPLETED, AND THE POCKET CLOSED WHEN ANESTHESIA NOTED THE PATIENT¿S OXYGEN LEVEL AND BLOOD PRESSURE TO BE DROPPING. IT WAS INDICATED THAT A CLOT HAD FORMED IN THE PATIENT¿S RIGHT VENTRICLE WHICH ELIMINATED BLOOD FLOW TO THE BODY. THE CAUSE OF THE CLOT WAS UNKNOWN. IT WAS INDICATED THAT NEITHER THE EV ICD NOR EV LEAD CAUSED THE PATIENT¿S DEATH.Event Description: IT WAS FURTHER REPORTED THAT DFT TESTING FAILED AT 30 AND 40 JOULES (J), AND EXTERNAL DEFIBRILLATION WAS UNSUCCESSFUL. CHEST COMPRESSIONS WERE PERFORMED, AND A 40J SHOCK IN REVERSE POLARITY SUCCESSFULLY CONVERTED THE RHYTHM. ALL EV LEAD MEASUREMENTS WERE NORMAL FOLLOWING CONVERSION, AND THE PROCEDURE WAS DEEMED A SUCCESSFUL IMPLANT. APPROXIMATELY 15 MINUTES LATER, ANESTHESIA NOTED THE PATIENT¿S VITALS DETERIORATING, AND RESUSCITATION WAS PERFORMED WHICH INCLUDED SHOCKS AND PACING FROM THE EV ICD SYSTEM. HEART CATHETERIZATION WAS PERFORMED BY AN INTERVENTIONALIST, AND NO ANOMALIES WERE FOUND IN THE ARTERIES. THE PATIENT RECEIVED EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO), AND AN ULTRASOUND ULTIMATELY FOUND THE MASSIVE CLOT. THE PHYSICIAN DETERMINED THE CAUSE OF DEATH TO BE PULMONARY EMBOLISM. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21642351&pc=DWE | 9612164-2025-01399 | 2025-02-04 05:00 | Death | MEDTRONIC MEXICO | 2025-03-19 04:00 | DWE | CUSTOM TUBING PACK | Air/Gas in Device | Air Embolism; Hemorrhage/Bleeding | K171979 |  | 1 | Event Description: MEDTRONIC RECEIVED INFORMATION THAT DURING USE OF 3 CUSTOM TUBING PACKS, IT WAS REPORTED THE PLASTIC PUMP HOUSING THAT CONNECTS TO THE ABBOTT CENTROMAG VENOUS CANNULA CRACKED AND CAUSED BUBBLES TO ENTER THE CIRCUIT. CUSTOMER ATTEMPTED TO STOP EXSANGUINOUS DRAINAGE, AND PERFORMED CIRCUIT EXCHANGE, BUT PATIENT CODED DURING THIS EXCHANGE, CUSTOMER PLACED PATIENT BACK ON VENO-VENOUS EXTRACORPOREAL MEMBRANE OXYGENATION (VV ECMO) BUT THE BUBBLES WERE ALREADY IDENTIFIED IN THE CIRCUIT, WITH INABILITY TO GET BUBBLES OUT WITHOUT CAUSING ADDITIONAL HARM TO THE PATIENT. DESPITE ALL EFFORTS, THE PATIENT EXPIRED AT 00:05AM ON (B)(6)2025. THE DEVICE WAS REPLACED TO COMPLETE THE PROCEDURE. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21642438&pc=LWS | 9614453-2025-01093 | 2025-03-03 05:00 | Death | MEDTRONIC EUROPE SARL | 2025-03-19 04:00 | LWS | AURORA EV-ICD¿ MRI SURESCAN¿ | Appropriate Term/Code Not Available | Pulmonary Embolism; Hypoxia; Thrombosis/Thrombus | P220012 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IS DECEASED. THE PATIENT DIED AFTER THE PLACEMENT OF AN EXTRAVASCULAR (EV) IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) SYSTEM. DEFIBRILLATION THRESHOLD (DFT) TESTING HAD BEEN COMPLETED, AND THE POCKET CLOSED WHEN ANESTHESIA NOTED THE PATIENT¿S OXYGEN LEVEL AND BLOOD PRESSURE TO BE DROPPING. IT WAS INDICATED THAT A CLOT HAD FORMED IN THE PATIENT¿S RIGHT VENTRICLE WHICH ELIMINATED BLOOD FLOW TO THE BODY. THE CAUSE OF THE CLOT WAS UNKNOWN. IT WAS INDICATED THAT NEITHER THE EV ICD NOR EV LEAD CAUSED THE PATIENT¿S DEATH.Event Description: IT WAS FURTHER REPORTED THAT DFT TESTING FAILED AT 30 AND 40 JOULES (J), AND EXTERNAL DEFIBRILLATION WAS UNSUCCESSFUL. CHEST COMPRESSIONS WERE PERFORMED, AND A 40J SHOCK IN REVERSE POLARITY SUCCESSFULLY CONVERTED THE RHYTHM. ALL EV LEAD MEASUREMENTS WERE NORMAL FOLLOWING CONVERSION, AND THE PROCEDURE WAS DEEMED A SUCCESSFUL IMPLANT. APPROXIMATELY 15 MINUTES LATER, ANESTHESIA NOTED THE PATIENT¿S VITALS DETERIORATING, AND RESUSCITATION WAS PERFORMED WHICH INCLUDED SHOCKS AND PACING FROM THE EV ICD SYSTEM. HEART CATHETERIZATION WAS PERFORMED BY AN INTERVENTIONALIST, AND NO ANOMALIES WERE FOUND IN THE ARTERIES. THE PATIENT RECEIVED EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO), AND AN ULTRASOUND ULTIMATELY FOUND THE MASSIVE CLOT. THE PHYSICIAN DETERMINED THE CAUSE OF DEATH TO BE PULMONARY EMBOLISM. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21642447&pc=DTZ | 3013876692-2025-00017 | 2025-03-16 04:00 | Death | MAQUET CARDIOPULMONARY GMBH | 2025-03-19 04:00 | DTZ | TUBING SET | Fluid/Blood Leak; Leak/Splash | Cardiac Arrest | K112360 |  | 1 | Event Description: THE EVENT OCCURRED IN CZECHIA DURING TREATMENT. IT WAS REPORTED THAT THERE WAS A BLOOD LEAK DIRECTLY FROM THE OXYGENATOR. THE CUSTOMER MANUALLY COMPRESSED THE LEAKAGE POINT AND REPLACED THE HLS SET. THE PATIENT HAD TO BE RESUSCITATED AND SURVIVED. A BLOOD TRANSFUSION WAS REQUIRED. AS THE PATIENT NEEDED A BLOOD TRANSFUSION AND REQUIRED RESUSCITATION, A REPORT IS REQUIRED. COMPLAINT ID# (B)(4).Event Description: COMPLAINT ID#: (B)(4).Event Description: COMPLAINT ID (B)(4). Manufacturer Narrative: THE EVENT OCCURRED IN CZECHIA DURING TREATMENT. IT WAS REPORTED THAT THERE WAS A BLOOD LEAK DIRECTLY FROM THE OXYGENATOR. THE CUSTOMER MANUALLY COMPRESSED THE LEAKAGE POINT AND REPLACED THE HLS SET. THE PATIENT HAD TO BE RESUSCITATED AND SURVIVED. A BLOOD TRANSFUSION WAS REQUIRED. AS THE PATIENT NEEDED A BLOOD TRANSFUSION, REQUIRED RESUSCITATION AND LATER PASSED AWAY, A REPORT IS REQUIRED. NEW INFORMATION WAS RECEIVED THAT THE PATIENT LOST 1000ML OF BLOOD. THE PATIENT UNDERWENT A TWO MINUTE CARDIOPULMONARY RESUSCITATION DURING THE ECMO CIRCUIT EXCHANGE DUE TO SEVERE CARDIAC DYSFUNCTION SUBSEQUENT TO EXTENSIVE MYOCARDIAL INFARCTION. NEW INFORMATION WAS RECEIVED ON 2025-04-01 THAT THE PATIENT PASSED DUE TO LATE COMPLICATIONS, DAYS AFTER WEANING FROM THE ECMO. THE AFFECTED PRODUCT WAS INVESTIGATED BY THE GETINGE LABORATORY ON 2025-05-15. THE REPORTED FAILURE COULD BE CONFIRMED, BUT THE EXACT ROOT CAUSE COULD NOT BE DETERMINED BUT MOST PROBABLY RELATED TO A LEAKAGE IN THE INTERNAL PRESSURE SENSOR. A MEDICAL REVIEW WAS PERFORMED BY GETINGE MEDICAL AFFAIRS ON 2025-05-28 WITH FOLLOWING CONCLUSION: "ACCORDING TO THE COMPLAINT NARRATIVE, THE INCIDENT INVOLVES A 54-YEAR-OLD MALE PATIENT ON VA-ECMO SUPPORT WITH A CARDIOHELP SYSTEM AND HLS SET ADVANCED 7.0, WHO EXPERIENCED A SUDDEN BLOOD LEAK FROM THE OXYGENATOR APPROXIMATELY THREE HOURS AFTER THERAPY INITIATION. THIS LED TO A SIGNIFICANT DROP IN ARTERIAL PRESSURE AND REQUIRED URGENT CIRCUIT REPLACEMENT. MANUAL COMPRESSION WAS APPLIED TO THE LEAK SITE, AND THE HLS SET WAS REPLACED. THE PATIENT UNDERWENT TWO MINUTES OF CPR DUE TO SEVERE CARDIAC DYSFUNCTION AND REQUIRED BLOOD TRANSFUSION. THE PATIENT INITIALLY SURVIVED BUT LATER EXPIRED FROM COMPLICATIONS AFTER ECMO WEANING. INVESTIGATION CONFIRMED THE COMPLAINT. BLOOD RESIDUES AND LEAKAGE WERE LOCALIZED TO THE INTERNAL PRESSURE SENSOR AREA. THE HLS SET WAS INSPECTED PRIOR TO USE WITH NO VISIBLE ISSUES, AND NO LEAKS WERE DETECTED DURING PRIMING. THE LEAK INVOLVED BLOOD ONLY. IMMEDIATE CLINICAL RESPONSE PREVENTED FURTHER HARM. IT IS IMPORTANT TO NOTE THAT DURING THE TECHNICAL INVESTIGATION, ONLY DROPLET FORMATION WAS OBSERVED AT THE BLOOD INLET COVER AS EVIDENCE OF LEAKAGE. HOWEVER, THE COMPLAINT NARRATIVE REPORTS AN ESTIMATED TOTAL BLOOD LOSS OF APPROXIMATELY 1000 ML, WITHOUT SPECIFYING THE EXACT TIME FRAME OVER WHICH THIS LOSS OCCURRED. THIS DISCREPANCY HIGHLIGHTS A GAP IN THE INFORMATION REGARDING THE PROGRESSION AND SEVERITY OF THE LEAKAGE. A PRODUCT MALFUNCTION WAS CONFIRMED THROUGH THE INVESTIGATION, SPECIFICALLY RELATED TO DEFECTS IN THE INTERNAL PRESSURE SENSOR. ACCORDING TO THE CUSTOMER¿S STATEMENT, ¿THE PATIENT INITIALLY SURVIVED BUT LATER EXPIRED FROM COMPLICATIONS AFTER ECMO WEANING.¿ THIS SUGGESTS THAT THE PATIENT¿S EVENTUAL EXPIRATION WAS LIKELY NOT DIRECTLY CAUSED BY THE HLS MALFUNCTION. NONETHELESS, THE URGENT NEED TO REPLACE THE HLS SET AND THE RESUSCITATION EFFORTS DESCRIBED DURING THIS PERIOD UNDERSCORE THE POTENTIAL FOR SERIOUS HARM HAD THE INTERVENTION BEEN DELAYED OR UNSUCCESSFUL. IN SUMMARY, WHILE THE HLS MALFUNCTION WAS VERIFIED BY INTERNAL GETINGE TESTING. MOREOVER, THE REPORTED EVENT LIKELY CONTRIBUTED TO THE CRITICAL CLINICAL OCCURRENCE, SPECIFICALLY CPR AND BLOOD TRANSFUSION. THE PATIENT¿S EXPIRATION APPEARS TO BE UNRELATED TO THE DEVICE FAILURE ITSELF, HAVING OCCURRED AN UNDEFINED PERIOD AFTER THE EVENT. IT IS ASSUMED THAT THE LOSS OF EXTRACORPOREAL SUPPORT (I.E., EXPERIENCED DURING PRODUCT EXCHANGE) LIKELY PRECIPITATED CPR. FURTHER, THE LOSS OF BLOOD VOLUME APPEARS TO BE DIRECTLY RELATED TO THE ADMINISTRATION OF BLOOD PRODUCTS." ACCORDING TO THE INSTRUCTION FOR USE (IFU, HLS SET ADVANCED 6.0/7.0, HIT SET ADVANCED 5.0/7.0) IT IS STATED IN CHAPTER ¿SAFETY INSTRUCTIONS FOR CENTRIFUGAL PUMP¿ THAT LEAKS OR SCRATCHING NOISES IN THE CENTRIFUGAL PUMP CAN BE A SIGN THAT IT IS MALFUNCTIONING. MALFUNCTIONING CAN LEAD TO INADEQUATE PATIENT SUPPORT. FURTHERMORE, IT IS MENTIONED TO ALWAYS KEEP A REPLACEMENT HLS SET FOR THOSE SITUATIONS. IN CHAPTER ¿PRIMING THE SYSTEM¿ IT IS WRITTEN NOT TO PRE-PRIME THE CIRCUIT TOO EARLY AS THIS COULD LEAD TO LEAKS AND TO NOT USE A DEVICE IF THERE ARE ANY LEAKS. THE PRODUCTION RECORDS OF THE AFFECTED PRODUCT WERE REVIEWED ON 2025-06-02. ACCORDING TO THE FINAL TEST RESULTS, THE PRODUCT PASSED THE TESTS AS PER SPECIFICATIONS. PRODUCTION RELATED INFLUENCES ARE UNLIKELY. THE REVIEW OF SCRAP, REWORK, ENHANCEMENTS AND DESIGN CHANGES WERE REVIEWED AND NO ABNORMALITIES IN REGARD TO THE REPORTED FAILURE WERE FOUND. IN ADDITION, A REVIEW FOR POTENTIAL FIELD ACTIONS AND CAPAS RELATED TO THE FAILURE AND PRODUCT IN THIS COMPLAINT WAS PERFORMED. THIS COMPLAINT IS NOT IN SCOPE OF ANY ONGOING FIELD ACTIONS AND/OR CAPAS. THE REVIEW OF THE NON-CONFORMITIES HAS BEEN PERFORMED FOR THE REVIEWED TIME PERIOD. IT DOES NOT SHOW ANY NON-CONFORMITY IN REGARD TO THE REPORTED FAILURE. BASED ON THE RESULTS THE REPORTED FAILURE "LEAKAGE ON THE OXYGENATOR" COULD BE CONFIRMED. THE PATIENT¿S EXPIRATION APPEARS TO BE UNRELATED TO THE DEVICE FAILURE ITSELF, HAVING OCCURRED AN UNDEFINED PERIOD AFTER THE EVENT. THE CUSTOMER WILL BE INFORMED ABOUT THE RESULTS BY THE GETINGE SALES AND SERVICE UNIT. THE OCCURRENCE RATE WAS CALCULATED FOR THE REPORTED ISSUE AND IT WAS DETERMINED THAT THIS IS NOT A SYSTEMIC ISSUE. THEREFORE, NO REMEDIAL ACTION IS REQUIRED. THIS COMPLAINT WAS FOUND IN THE DATABASE OF CUSTOMER COMPLAINTS FOR THE HLS SET WITH ARTICLE (B)(4) AS A SINGLE EVENT (TIMEFRAME FROM 2024-03-17 TILL 2025-03-17). ACCORDING TO THE RISK REVIEW THE REPORTED FAILURE IS BELOW THE ACCEPTANCE THRESHOLD ACCORDING TO THE RISK MANAGEMENT PLAN OF THE HLS SET. THE OCCURRENCE RATE RELATED TO THE REPORTED ISSUE IS CURRENTLY BEING MONITORED AS PART OF MAQUET CARDIOPULMONARY¿ S TRENDING PROGRAM AND ADDITIONAL INVESTIGATIONS OR CORRECTIONS WILL BE IMPLEMENTED IN CASE OF ADVERSE TRENDING. Manufacturer Narrative: A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. NOTE: NO UDI NUMBER HAS BEEN INCLUDED IN THE SECTION D4 UNIQUE DEVICE IDENTIFIER (UDI) SINCE THE LOT NUMBER OF THE AFFECTED SET WAS NOT PROVIDED. THEREFORE IT IS NOT POSSIBLE TO IDENTIFY THE SET LOT NUMBER OF THE DEVICE IN QUESTION AND THEREFORE ITS UDI NUMBER. Manufacturer Narrative: THE PATIENT LOST 1000ML OF BLOOD. THE PATIENT UNDERWENT A TWO MINUTE CARDIOPULMONARY RESUSCITATION DURING THE ECMO CIRCUIT EXCHANGE DUE TO SEVERE CARDIAC DYSFUNCTION SUBSEQUENT TO EXTENSIVE MYOCARDIAL INFARCTION. NEW INFORMATION WAS RECEIVED ON 2025-04-01 THAT THE OUTCOME OF THE PATIENT WAS DEATH DUE TO LATE COMPLICATIONS, DAYS AFTER WEANING OFF FROM THE ECMO. A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. NOTE: THIS EVENT OCCURRED ON THE CZECHIA MARKET. IT IS A SIGNIFICANTLY SIMILAR DEVICE TO "HLS SET¿ WHICH IS SOLD IN THE USA UNDER PREMARKET SUBMISSION NUMBER K112360. FOR SECTION D4 ALL AVAILABLE IDENTIFYING INFORMATION IS FOR THE OUT OF THE US DEVICE, SUBJECTED TO THIS REPORT. THEREFORE, NO GUDID INFORMATION EXISTS. FURTHERMORE, UDI INFORMATION (PRIMARY DI NUMBER) PROVIDED IN D4 IS FOR HLS SET WITH CATALOG NUMBER 701069073.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21642465&pc=DWE | 9612164-2025-01400 | 2025-02-04 05:00 | Death | MEDTRONIC MEXICO | 2025-03-19 04:00 | DWE | CUSTOM TUBING PACK | Air/Gas in Device | Air Embolism; Hemorrhage/Bleeding | K171979 |  | 1 | Event Description: MEDTRONIC RECEIVED INFORMATION THAT DURING USE OF 3 CUSTOM TUBING PACKS, IT WAS REPORTED THE PLASTIC PUMP HOUSING THAT CONNECTS TO THE ABBOTT CENTROMAG VENOUS CANNULA CRACKED AND CAUSED BUBBLES TO ENTER THE CIRCUIT. CUSTOMER ATTEMPTED TO STOP EXSANGUINOUS DRAINAGE, AND PERFORMED CIRCUIT EXCHANGE, BUT PATIENT CODED DURING THIS EXCHANGE, CUSTOMER PLACED PATIENT BACK ON VENO-VENOUS EXTRACORPOREAL MEMBRANE OXYGENATION (VV ECMO) BUT THE BUBBLES WERE ALREADY IDENTIFIED IN THE CIRCUIT, WITH INABILITY TO GET BUBBLES OUT WITHOUT CAUSING ADDITIONAL HARM TO THE PATIENT. DESPITE ALL EFFORTS, THE PATIENT EXPIRED AT 00:05AM ON (B)(6) 2025. THE DEVICE WAS REPLACED TO COMPLETE THE PROCEDURE. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21642469&pc=NGV | 2124215-2025-16790 | 2025-03-09 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-03-19 04:00 | NGV | WATCHMAN FLX PRO | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P130013 |  | 1 | Event Description: REPORTED VIA CLINICAL STUDY. IT WAS REPORTED THAT DEATH OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS PERFORMED AND A 31MM WATCHMAN FLX PRO CLOSURE DEVICE WAS IMPLANTED WITH COMPLETE DEVICE SEAL NOTED. THE PATIENT WAS DISCHARGED ON RIVAROXABAN. 366 DAYS POST INDEX PROCEDURE, THE PATIENT WAS FOUND TO BE EXPIRED. AT THE TIME OF THE EVENT, THE PATIENT WAS ON ASPIRIN AND CLOPIDOGREL. THE CAUSE OF DEATH REMAINS UNKNOWN, AND NO FURTHER DETAILS WERE MADE AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21642940&pc=DSQ | 2916596-2025-01498 | 2023-12-21 05:00 | Death | THORATEC CORPORATION | 2025-03-19 04:00 | DSQ | THORATEC® HEARTMATE®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure; Heart Failure/Congestive Heart Failure | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2023 DUE TO RIGHT VENTRICLE FAILURE (RVF) AND MULTISYSTEM ORGAN FAILURE. IT WAS A SLOW, LONG COURSE, AND WAS REFRACTORY TO DOBUTAMINE AND DIURETICS. THE PATIENT PASSED AWAY AT HOME. THE PUMP WORKED APPROPRIATELY. IT WAS UNCLEAR IF THE RVF WAS PREVALENT BEFORE THE DEVICE.Event Description: IT WAS ADDITIONALLY REPORTED THAT THE DEATH WAS NOT CONSIDERED DEVICE RELATED AND THE DEVICE OPERATED AS INTENDED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE II LVAS, SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. (B)(6) WAS NOT RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM INSTRUCTIONS FOR USE IS CURRENTLY AVAILABLE. SECTION 1, "INTRODUCTION," OUTLINES POTENTIAL ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM, INCLUDING DEATH AND MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21643003&pc=MOZ | MW5167799 | 2025-03-01 05:00 | Death | SEIKAGAKU CORPORATION | 2025-03-19 04:00 | MOZ | GEL-ONE PFS 30MG/3ML |  | Insufficient Information |  |  | 1 | Event Description: PT IS REPORTED AS HAVING PASSED AWAY ON (B)(6) 2025. THE CAUSE OF DEATH WAS NOT PROVIDED BY THE REPORTER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21643184&pc=NPT | 2015691-2025-02040 | 2025-02-11 05:00 | Death | EDWARDS LIFESCIENCES | 2025-03-19 04:00 | NPT | SAPIEN 3 TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P140031 |  | 1 | Event Description: AS REPORTED, IT WAS A CASE OF AN IMPLANT OF A 23MM SAPIEN 3 ULTRA VALVE INSIDE OF A DEGENERATED NON-EDWARDS VALVE, IN AORTIC POSITION BY TRANSFEMORAL APPROACH. DURING PROCEDURE, THE VALVE WAS IMPLANTED WITH GOOD RESULTS AND THE FEMORAL ACCESS WAS SUCCESSFULLY CLOSED. THE PATIENT OUTCOME WAS WITHOUT POST-PROCEDURAL COMPLICATIONS AND REMAINED ASYMPTOMATIC. UNFORTUNATELY, THE PATIENT PASSED AWAY THE DAY AFTER THE IMPLANT. THE CAUSE OF DEATH WAS AN AORTIC DISSECTION ORIGINATED FROM THE PROXIMAL CONTACT POINT OF THE PREVIOUS NON-EDWARDS VALVE ON THE AORTA. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. ACCORDING TO THE DEVICE TRAINING MANUALS, RISK FACTORS FOR AORTIC DISSECTION, CARDIAC/AORTIC HEMATOMA, OR ANNULAR RUPTURE DURING THE TVR PROCEDURE INCLUDE SIGNIFICANT VALVE OVER SIZING, SEVERELY OBLITERATED SINUSES OF VALSALVA, PORCELAIN AORTA AND/OR PRESENCE OF BULKY CALCIFICATION, AND NARROW CALCIFIED SINOTUBULAR JUNCTION. IN ADDITION, ADVANCED AGE, FEMALE GENDER, SMALL BODY WEIGHT, AND STEROID DEPENDENCY CAN ALSO BE CONTRIBUTING FACTORS. THE SAPIEN TRANSCATHETER HEART VALVE (ALL MODELS) RELIES CALCIUM TO SECURELY ANCHOR IN THE LANDING ZONE. DESPITE THIS BENEFICIAL ASPECT OF CALCIUM, BULKY CALCIUM CAN INCREASE THE RISK OF CALCIFIC NODULE DISPLACEMENT INTO THE VASCULATURE, WHICH CAN LEAD TO VASCULAR INJURY. AT TIMES, THE EXTENT AND DISTRIBUTION OF CALCIUM CAN IMPAIR EASE OF DELIVERY OF THE VALVE, CORRECT POSITIONING OF THE VALVE, DEPLOYMENT OF THE VALVE, AND PROCEDURAL SUCCESS. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST THE AORTIC DISSECTION ORIGINATED FROM THE PROXIMAL CONTACT POINT OF THE PREVIOUS NON-EDWARDS VALVE ON THE AORTA. HOWEVER, THE CAUSE OF THE EVENT IS UNKNOWN DUE TO THE LIMITED INFORMATION AVAILABLE. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21643512&pc=MIH | 2247858-2025-00085 | 2024-10-15 04:00 | Death | BOLTON MEDICAL, INC. | 2025-03-19 04:00 | MIH | TREO ABDOMINAL STENT-GRAFT SYSTEM | Insufficient Information; Device Stenosis | Stenosis; Restenosis | P190015 |  | 1 | Event Description: "ON 28FEB2025, THE CORE LAB SENT NOTIFICATION OF SIGNIFICANT FINDING FOR SUBJECT (B)(6) OF STENOSIS IN THE PROXIMAL GRAFT FOR 3-YEAR IMAGING DATED (B)(6) 2024 AND UNSCHEDULED IMAGING DATED (B)(6) 2024. THE UNSCHEDULED IMAGING WAS OBTAINED DURING THE SAME HOSPITALIZATION FOR THE INITIAL >5MM INCREASE NOTED ON THEIR LATE 2-YEAR IMAGING." PATIENT OUTCOME: "THE SUBJECT WAS HOSPITALIZED AGAIN AND DEATH OCCURRED ON (B)(6) 2024 DUE TO SEPTIC SHOCK UNRELATED TO THE STUDY."Event Description: "ON 28FEB2025, THE CORE LAB SENT NOTIFICATION OF SIGNIFICANT FINDING FOR SUBJECT (B)(6) OF STENOSIS IN THE PROXIMAL GRAFT FOR 3-YEAR IMAGING DATED (B)(6) 2024 AND UNSCHEDULED IMAGING DATED (B)(6) 2024. THE UNSCHEDULED IMAGING WAS OBTAINED DURING THE SAME HOSPITALIZATION FOR THE INITIAL >5MM INCREASE NOTED ON THEIR LATE 2-YEAR IMAGING THAT RESULTED IN A RE-INTERVENTION. (COMPLAINT (B)(4)). BOTH CT REPORTS NOTE THROMBUS ALONG THE POSTERIOR WALL OF THE ABDOMINAL AORTA PROPAGATING INTO THE PROXIMAL LEFT RENAL ARTERY RESULTING IN 80% STENOSIS WITH THE (B)(6) 2024 CT REPORT FURTHER SPECIFYING THE THROMBUS IN THE STENT GRAFT. ON 06MAR2025, THE INVESTIGATOR RE-REVIEWED THE IMAGING STATING "THERE IS NOT SIGNIFICANT STENOSIS IN THE GRAFT. IT APPEARS WIDELY PATENT. I SEE NO ISSUES WITH THE DEVICE. NONE OF THIS IS DEVICE RELATED". THEREFORE, THERE WAS NO ADVERSE EVENT THAT OCCURRED." PATIENT OUTCOME: "THE SUBJECT WAS DISCHARGED (B)(6) 2024 AFTER THEIR RE-INTERVENTION HOSPITALIZATION RESULTED IN A PROLONGATION DUE TO THE SUBJECT DEVELOPING ACUTE CHOLECYSTITIS. ON 06MAR2025, THE INVESTIGATOR CONFIRMED THAT THE CHOLECYSTITIS WAS LIKELY NOT REINTERVENTION RELATED. THE SUBJECT WAS HOSPITALIZED AGAIN DUE TO COMPLICATIONS FROM HIS CHOLECYSTITIS RESULTING IN SEPTIC SHOCK AND THEN NON-STUDY RELATED DEATH OCCURRED ON (B)(6) 2024." Manufacturer Narrative: THIS COMPLAINT WAS INVOLVED WITH THREE DEVICES. DEVICE 1 IS BEING REPORTED UNDER MDR-2247858-2025-00083, DEVICE 2 IS BEING REPORTED UNDER MDR-2247858-2025-00084, AND DEVICE 3 IS BEING REPORTED UNDER MDR-. BOLTON MEDICAL, INC. (D/B/A TERUMO AORTIC), HEREIN KNOWN AS THE "COMPANY", IS SUBMITTING THIS REPORT PURSUANT TO 21 CFR PART 803, HAS MADE REASONABLE EFFORTS TO OBTAIN COMPLETE INFORMATION, AND HAS PROVIDED AS MUCH AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT IS BASED ON INFORMATION OBTAINED BY THE COMPANY, WHICH MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE REGULATIONS. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN, UNAVAILABLE, OR NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY ANYONE THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE CERTAIN REGULATIONS, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. Manufacturer Narrative: THIS COMPLAINT WAS INVOLVED WITH THREE DEVICES. DEVICE 1 IS BEING REPORTED UNDER MDR-2247858-2025-00083, DEVICE 2 IS BEING REPORTED UNDER MDR-2247858-2025-00084, AND DEVICE 3 IS BEING REPORTED UNDER MDR-2247858-2025-00085. BOLTON MEDICAL, INC. (D/B/A TERUMO AORTIC), HEREIN KNOWN AS THE "COMPANY", IS SUBMITTING THIS REPORT PURSUANT TO 21 CFR PART 803, HAS MADE REASONABLE EFFORTS TO OBTAIN COMPLETE INFORMATION, AND HAS PROVIDED AS MUCH AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT IS BASED ON INFORMATION OBTAINED BY THE COMPANY, WHICH MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE REGULATIONS. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN, UNAVAILABLE, OR NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY ANYONE THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE CERTAIN REGULATIONS, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21643520&pc=MIH | 2247858-2025-00083 | 2024-10-15 04:00 | Death | BOLTON MEDICAL, INC. | 2025-03-19 04:00 | MIH | TREO ABDOMINAL STENT-GRAFT SYSTEM | Insufficient Information; Device Stenosis | Stenosis; Restenosis | P190015 |  | 1 | Event Description: "ON (B)(6) 2025, THE CORE LAB SENT NOTIFICATION OF SIGNIFICANT FINDING FOR SUBJECT (B)(6) OF STENOSIS IN THE PROXIMAL GRAFT FOR 3-YEAR IMAGING DATED (B)(6) 2024 AND UNSCHEDULED IMAGING DATED (B)(6) 2024. THE UNSCHEDULED IMAGING WAS OBTAINED DURING THE SAME HOSPITALIZATION FOR THE INITIAL >5MM INCREASE NOTED ON THEIR LATE 2-YEAR IMAGING." PATIENT OUTCOME: "THE SUBJECT WAS HOSPITALIZED AGAIN AND DEATH OCCURRED ON (B)(6) 2024 DUE TO SEPTIC SHOCK UNRELATED TO THE STUDY."Event Description: "ON 28FEB2025, THE CORE LAB SENT NOTIFICATION OF SIGNIFICANT FINDING FOR SUBJECT (B)(6) OF STENOSIS IN THE PROXIMAL GRAFT FOR 3-YEAR IMAGING DATED (B)(6) 2024 AND UNSCHEDULED IMAGING DATED (B)(6) 2024. THE UNSCHEDULED IMAGING WAS OBTAINED DURING THE SAME HOSPITALIZATION FOR THE INITIAL >5MM INCREASE NOTED ON THEIR LATE 2-YEAR IMAGING THAT RESULTED IN A RE-INTERVENTION. (COMPLAINT (B)(4)). BOTH CT REPORTS NOTE THROMBUS ALONG THE POSTERIOR WALL OF THE ABDOMINAL AORTA PROPAGATING INTO THE PROXIMAL LEFT RENAL ARTERY RESULTING IN 80% STENOSIS WITH THE (B)(6) 2024 CT REPORT FURTHER SPECIFYING THE THROMBUS IN THE STENT GRAFT. ON 06MAR2025, THE INVESTIGATOR RE-REVIEWED THE IMAGING STATING "THERE IS NOT SIGNIFICANT STENOSIS IN THE GRAFT. IT APPEARS WIDELY PATENT. I SEE NO ISSUES WITH THE DEVICE. NONE OF THIS IS DEVICE RELATED". THEREFORE, THERE WAS NO ADVERSE EVENT THAT OCCURRED." PATIENT OUTCOME: "THE SUBJECT WAS DISCHARGED (B)(6) 2024 AFTER THEIR RE-INTERVENTION HOSPITALIZATION RESULTED IN A PROLONGATION DUE TO THE SUBJECT DEVELOPING ACUTE CHOLECYSTITIS. ON 06MAR2025, THE INVESTIGATOR CONFIRMED THAT THE CHOLECYSTITIS WAS LIKELY NOT REINTERVENTION RELATED. THE SUBJECT WAS HOSPITALIZED AGAIN DUE TO COMPLICATIONS FROM HIS CHOLECYSTITIS RESULTING IN SEPTIC SHOCK AND THEN NON-STUDY RELATED DEATH OCCURRED ON (B)(6) 2024." Manufacturer Narrative: THIS COMPLAINT WAS INVOLVED WITH THREE DEVICES. DEVICE 1 IS BEING REPORTED UNDER MDR-2247858-2025-00083, DEVICE 2 IS BEING REPORTED UNDER MDR-2247858-2025-00084, AND DEVICE 3 IS BEING REPORTED UNDER MDR-2247858-2025-00085. BOLTON MEDICAL, INC. (D/B/A TERUMO AORTIC), HEREIN KNOWN AS THE "COMPANY", IS SUBMITTING THIS REPORT PURSUANT TO 21 CFR PART 803, HAS MADE REASONABLE EFFORTS TO OBTAIN COMPLETE INFORMATION, AND HAS PROVIDED AS MUCH AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT IS BASED ON INFORMATION OBTAINED BY THE COMPANY, WHICH MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE REGULATIONS. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN, UNAVAILABLE, OR NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY ANYONE THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE CERTAIN REGULATIONS, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. Manufacturer Narrative: THIS COMPLAINT WAS INVOLVED WITH THREE DEVICES. DEVICE 1 IS BEING REPORTED UNDER MDR-, DEVICE 2 IS BEING REPORTED UNDER MDR-2247858-2025-00084, AND DEVICE 3 IS BEING REPORTED UNDER MDR-2247858-2025-00085. BOLTON MEDICAL, INC. (D/B/A TERUMO AORTIC), HEREIN KNOWN AS THE "COMPANY", IS SUBMITTING THIS REPORT PURSUANT TO 21 CFR PART 803, HAS MADE REASONABLE EFFORTS TO OBTAIN COMPLETE INFORMATION, AND HAS PROVIDED AS MUCH AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT IS BASED ON INFORMATION OBTAINED BY THE COMPANY, WHICH MAY HAVE NOT BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE REGULATIONS. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN, UNAVAILABLE, OR NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY ANYONE THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE CERTAIN REGULATIONS, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21643571&pc=MIH | 2247858-2025-00084 | 2024-10-15 04:00 | Death | BOLTON MEDICAL, INC. | 2025-03-19 04:00 | MIH | TREO ABDOMINAL STENT-GRAFT SYSTEM | Insufficient Information; Device Stenosis | Stenosis; Restenosis | P190015 |  | 1 | Event Description: "ON 28FEB2025, THE CORE LAB SENT NOTIFICATION OF SIGNIFICANT FINDING FOR SUBJECT (B)(6) OF STENOSIS IN THE PROXIMAL GRAFT FOR 3-YEAR IMAGING DATED (B)(6) 2024 AND UNSCHEDULED IMAGING DATED (B)(6) 2024. THE UNSCHEDULED IMAGING WAS OBTAINED DURING THE SAME HOSPITALIZATION FOR THE INITIAL >5MM INCREASE NOTED ON THEIR LATE 2-YEAR IMAGING." PATIENT OUTCOME: "THE SUBJECT WAS HOSPITALIZED AGAIN AND DEATH OCCURRED ON (B)(6) 2024 DUE TO SEPTIC SHOCK UNRELATED TO THE STUDY.Event Description: "ON 28FEB2025, THE CORE LAB SENT NOTIFICATION OF SIGNIFICANT FINDING FOR SUBJECT TPA-057-003 OF STENOSIS IN THE PROXIMAL GRAFT FOR 3-YEAR IMAGING DATED (B)(6) 2024 AND UNSCHEDULED IMAGING DATED (B)(6)2024. THE UNSCHEDULED IMAGING WAS OBTAINED DURING THE SAME HOSPITALIZATION FOR THE INITIAL >5MM INCREASE NOTED ON THEIR LATE 2-YEAR IMAGING THAT RESULTED IN A RE-INTERVENTION. (SEE COMPLAINT (B)(4)). BOTH CT REPORTS NOTE THROMBUS ALONG THE POSTERIOR WALL OF THE ABDOMINAL AORTA PROPAGATING INTO THE PROXIMAL LEFT RENAL ARTERY RESULTING IN 80% STENOSIS WITH THE (B)(6) 2024 CT REPORT FURTHER SPECIFYING THE THROMBUS IN THE STENT GRAFT. ON 06MAR2025, THE INVESTIGATOR RE-REVIEWED THE IMAGING STATING "THERE IS NOT SIGNIFICANT STENOSIS IN THE GRAFT. IT APPEARS WIDELY PATENT. I SEE NO ISSUES WITH THE DEVICE. NONE OF THIS IS DEVICE RELATED". THEREFORE, THERE WAS NO ADVERSE EVENT THAT OCCURRED." PATIENT OUTCOME: "THE SUBJECT WAS DISCHARGED (B)(6) 2024 AFTER THEIR RE-INTERVENTION HOSPITALIZATION RESULTED IN A PROLONGATION DUE TO THE SUBJECT DEVELOPING ACUTE CHOLECYSTITIS. ON (B)(6) 2025, THE INVESTIGATOR CONFIRMED THAT THE CHOLECYSTITIS WAS LIKELY NOT REINTERVENTION RELATED. THE SUBJECT WAS HOSPITALIZED AGAIN DUE TO COMPLICATIONS FROM HIS CHOLECYSTITIS RESULTING IN SEPTIC SHOCK AND THEN NON-STUDY RELATED DEATH OCCURRED ON (B)(6) 2024." Manufacturer Narrative: THIS COMPLAINT WAS INVOLVED WITH THREE DEVICES. DEVICE 1 IS BEING REPORTED UNDER MDR-2247858-2025-00083, DEVICE 2 IS BEING REPORTED UNDER MDR-2247858-2025-00084, AND DEVICE 3 IS BEING REPORTED UNDER MDR-2247858-2025-00085. BOLTON MEDICAL, INC. (D/B/A TERUMO AORTIC), HEREIN KNOWN AS THE "COMPANY", IS SUBMITTING THIS REPORT PURSUANT TO 21 CFR PART 803, HAS MADE REASONABLE EFFORTS TO OBTAIN COMPLETE INFORMATION, AND HAS PROVIDED AS MUCH AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT IS BASED ON INFORMATION OBTAINED BY THE COMPANY, WHICH MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE REGULATIONS. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN, UNAVAILABLE, OR NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY ANYONE THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE CERTAIN REGULATIONS, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. Manufacturer Narrative: THIS COMPLAINT WAS INVOLVED WITH THREE DEVICES. DEVICE 1 IS BEING REPORTED UNDER MDR-2247858-2025-00083, DEVICE 2 IS BEING REPORTED UNDER MDR- AND DEVICE 3 IS BEING REPORTED UNDER MDR-2247858-2025-00085. BOLTON MEDICAL, INC. (D/B/A TERUMO AORTIC), HEREIN KNOWN AS THE "COMPANY", IS SUBMITTING THIS REPORT PURSUANT TO 21 CFR PART 803, HAS MADE REASONABLE EFFORTS TO OBTAIN COMPLETE INFORMATION, AND HAS PROVIDED AS MUCH AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT IS BASED ON INFORMATION OBTAINED BY THE COMPANY, WHICH MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE REGULATIONS. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN, UNAVAILABLE, OR NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY ANYONE THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE CERTAIN REGULATIONS, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21643768&pc=OZD | 1220648-2025-26823 | 2025-02-20 05:00 | Death | ABIOMED, INC. | 2025-03-19 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Ischemia | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS) EXPERIENCED BLEEDING WITH HEPARIN DISCONTINUATION AND LIMB ISCHEMIA. HOWEVER, THE PATIENT WOULD SUBSEQUENTLY EXPIRE. THE PATIENT HAD A POSITIVE ELECTROCARDIOGRAM FOR INFERIOR LATERAL ST-SEGMENT ELEVATION MYOCARDIAL INFARCTION. AND WAS ELECTIVELY INTUBATED BEFORE BEING TAKEN TO THE CATHETERIZATION LAB. THE LEFT DOMINANT CIRCUMFLEX PROXIMALLY OCCLUDED AND THERE WAS SEVERE MITRAL REGURGITATION NOTED ON VENTRICULOGRAPHY. THE DECISION WAS MADE TO PLACE PATIENT ON IMPELLA CP MCS AND CONSULT CARDIOTHORACIC SURGERY FOR SURGICAL INTERVENTION. THE IMPELLA WAS SUCCESSFULLY IMPLANTED, AND THE PATIENT WAS TRANSPORTED TO THE UNIT. THREE DAYS LATER THE PATIENT UNDERWENT MITRAL VALVE REPLACEMENT AND CORONARY ARTERY BYPASS GRAFTING SURGERY. THE PATIENT WAS AT THIS TIME PLACE ON VENOARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION (VA ECMO). NOW DAY SIX OF IMPELLA MCS THE PATIENT WAS STARTED ON CONTINUOUS RENAL REPLACEMENT THERAPY (CRRT) WITH GOAL REMOVAL. ADDITIONALLY, IT WAS NOTED THE PATIENT DEVELOPED A GASTROINTESTINAL BLEED AND THE PURGE SOLUTION WAS SWITCHED TO SODIUM BICARBONATE. THREE UNITS OF PLATELETS WERE ADMINISTERED. TWO DAYS LATER THE PATIENT WAS ADMINISTERED TWO UNITS OF PACKED RED BLOOD CELLS FOR LOW HEMOGLOBIN. PER THE REPORT IT WAS STATED THE STAFF DOES NOT FEEL IT IS IMPELLA RELATED (HOWEVER UNCONFIRMED). THE PATIENT WAS HAVING PERSISTENT MELENA AND ALSO NOW NET NEGATIVE WITH CRRT REMOVAL, UNABLE TO WEAN FLOWS ON ECMO. THE FOLLOWING DAY, THE RIGHT LEG HAD LOSS OF PULSES. A PLAN WAS NOTED TO REMOVE THE IMPELLA CP WITH HOPES TO IMPROVE FLOW. PERSISTENT LOW HEMOGLOBIN NOTED WITH UNITS TRANSFUSED WAS NOTED. HOWEVER, THE PATIENT WOULD EXPIRE LATER IN THE DAY; CARE WAS WITHDRAWN. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY¿ (INCLUDING VENTRICULAR PERFORATION).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21643783&pc=DRE | 2124215-2025-17006 | 2025-02-27 05:00 | Death | BAYLIS MEDICALE CIE INC | 2025-03-19 04:00 | DRE | VERSACROSS CONNECT LAAC ACCESS SOLUTION | Adverse Event Without Identified Device or Use Problem | Air Embolism; Cardiac Arrest; Stroke/CVA; Non specific EKG/ECG Changes; Heart Block | K220414 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WAS AN AIR EMBOLISM AND THE PATIENT DIED. DURING A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE, A VERSACROSS CONNECT LAAC WAS SELECTED FOR USE. THE PHYSICIAN POSITIONED THE WATCHMAN TRUSTEER ACCESS SYSTEM (WAS) AND VERSACROSS CONNECT DILATOR OVER THE PIGTAIL WIRE THROUGH THE SEPTUM INTO THE LEFT ATRIUM (LA) OF THE PATIENT. UPON REMOVAL OF DILATOR AND PREPARING FOR A 6FR PIGTAIL CATHETER, THE PATIENT COUGHED, AND AIR WAS SUCTIONED INTO THE LA EITHER THROUGH THE WAS OR PIGTAIL CATHETER. AS PER THE PHYSICIAN, THE BLUE VALVE FOR THE STEERABLE SHEATH DOES NOT CLOSE COMPLETELY. AN AIR EMBOLISM WAS SEEN IN THE LA AND THE PHYSICIAN ASPIRATED THE AIR BUBBLES OUT FROM THE PATIENT. THE PHYSICIAN CONTINUED THE PROCEDURE WITH THE HOPE THAT ANY REMAINING AIR COULD BE TRAPPED BY IMPLANTING THE CLOSURE DEVICE. THE WDS WAS INSERTED, AND THE CLOSURE DEVICE WAS SUCCESSFULLY IMPLANTED IN THE LAA OF THE PATIENT. SHORTLY AFTER THE CLOSURE DEVICE HAD BEEN RELEASED THE PATIENT BEGAN TO EXPERIENCE AN ST SEGMENT ELEVATION WHICH LED TO COMPLETE HEART BLOCK. THE PATIENT RECEIVED CPR AND IV MEDICATION BEFORE THEY BECAME STABLE. THE PHYSICIAN BELIEVES THAT THE AIR EMBOLISM REACHED THE PATIENT'S BRAIN. PATIENT WAS ADMITTED TO INPATIENT ICU AND MONITORED. THE PATIENT DID NOT RECOVER FROM THIS EVENT AND THE PATIENT DIED. THE DEVICE IS NOT EXPECTED TO BE RETURNED FOR ANALYSIS (DISPOSED). THERE WAS NO DIFFICULTY WITH THE TRANSSEPTAL ACCESS. ACCORDING TO THE PHYSICIAN AND OBSERVATION UPON REMOVAL OF THE DILATOR (WIRE STILL INSIDE THE SHEATH) DESPITE THE PHYSICIAN TIGHTENING THE BLUE VALVE THERE WAS STILL BLEEDING BACK. THIS WAS A DEEP SEDATION (NO INTUBATION) CASE AND WHILE PREPARING TO INTRODUCE THE PIGTAIL 6F CATHETER THE PATIENT COUGHED, AND AIR WAS SUCKED INTO THE SHEATH AND DIRECTLY ENTERED THE LA.Event Description: IT WAS REPORTED THAT THERE WAS AN AIR EMBOLISM AND THE PATIENT DIED. DURING A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE, A VERSACROSS CONNECT LAAC WAS SELECTED FOR USE. THE PHYSICIAN POSITIONED THE WATCHMAN TRUSTEER ACCESS SYSTEM (WAS) AND VERSACROSS CONNECT DILATOR OVER THE PIGTAIL WIRE THROUGH THE SEPTUM INTO THE LEFT ATRIUM (LA) OF THE PATIENT. UPON REMOVAL OF DILATOR AND PREPARING FOR A 6FR PIGTAIL CATHETER, THE PATIENT COUGHED, AND AIR WAS SUCTIONED INTO THE LA EITHER THROUGH THE WAS OR PIGTAIL CATHETER. AS PER THE PHYSICIAN, THE BLUE VALVE FOR THE STEERABLE SHEATH DOES NOT CLOSE COMPLETELY. AN AIR EMBOLISM WAS SEEN IN THE LA AND THE PHYSICIAN ASPIRATED THE AIR BUBBLES OUT FROM THE PATIENT. THE PHYSICIAN CONTINUED THE PROCEDURE WITH THE HOPE THAT ANY REMAINING AIR COULD BE TRAPPED BY IMPLANTING THE CLOSURE DEVICE. THE WDS WAS INSERTED, AND THE CLOSURE DEVICE WAS SUCCESSFULLY IMPLANTED IN THE LAA OF THE PATIENT. SHORTLY AFTER THE CLOSURE DEVICE HAD BEEN RELEASED THE PATIENT BEGAN TO EXPERIENCE AN ST SEGMENT ELEVATION WHICH LED TO COMPLETE HEART BLOCK. THE PATIENT RECEIVED CPR AND IV MEDICATION BEFORE THEY BECAME STABLE. THE PHYSICIAN BELIEVES THAT THE AIR EMBOLISM REACHED THE PATIENT'S BRAIN. PATIENT WAS ADMITTED TO INPATIENT ICU AND MONITORED. THE PATIENT DID NOT RECOVER FROM THIS EVENT AND THE PATIENT DIED. THE DEVICE IS NOT EXPECTED TO BE RETURNED FOR ANALYSIS (DISPOSED). THERE WAS NO DIFFICULTY WITH THE TRANSSEPTAL ACCESS. ACCORDING TO THE PHYSICIAN AND OBSERVATION UPON REMOVAL OF THE DILATOR (WIRE STILL INSIDE THE SHEATH) DESPITE THE PHYSICIAN TIGHTENING THE BLUE VALVE THERE WAS STILL BLEEDING BACK. THIS WAS A DEEP SEDATION (NO INTUBATION) CASE AND WHILE PREPARING TO INTRODUCE THE PIGTAIL 6F CATHETER THE PATIENT COUGHED, AND AIR WAS SUCKED INTO THE SHEATH AND DIRECTLY ENTERED THE LA. IT WAS FURTHER CONFIRMED THAT THE PATIENT WAS DIAGNOSED WITH A CEREBRAL VASCULAR ACCIDENT CAUSED BY AN AIR EMBOLISM. Manufacturer Narrative: DUE TO ADDITIONAL INFORMATION RECEIVED, THIS SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR THE UPDATED FIELDS: DESCRIBE EVENT OR PROBLEM (B5), IN SECTION B - ADVERSE EVENT/PRODUCT PROB, AND PATIENT CODES (F10 AND H6), IN SECTIONS F - FOR USE BY UF/IMPORTER AND H - DEVICE MANUFACTURERS ONLY. IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION, BUT IT WAS NOT AVAILABLE. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION. Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION, BUT IT WAS NOT AVAILABLE. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21644138&pc=MIH | 9612164-2025-01411 | 2025-02-23 05:00 | Death | MEDTRONIC IRELAND | 2025-03-19 04:00 | MIH | ENDURANT II BIFURCATED STENT GRAFT | Adverse Event Without Identified Device or Use Problem | Abdominal Pain; Cardiac Arrest; Low Blood Pressure/ Hypotension | P100021 |  | 1 | Event Description: AN ENDURANT IIS STENT GRAFT SYSTEM WAS IMPLANTED DURING THE ENDOVASCULAR TREATMENT OF AN 80MM RUPTURED AAA. IT WAS REPORTED THAT THE PATIENT PRESENTED EMERGENTLY WITH ABDOMINAL PAIN AND WAS DIAGNOSED WITH A RUPTURED AAA. THE PHYSICIAN IMPLANTED FOUR ENDURANT II STENT GRAFTS HOWEVER AFTER THE STENTS WERE IMPLANTED, THE PATIENT'S BLOOD PRESSURE DROPPED AND HEART STOPPED BEATING. DEFIBRILLATION AND CARDIOPULMONARY RESUSCITATION FAILED AND THE PATIENT DIED DURING THE RUPTURED AAA REPAIR. PER THE PHYSICIAN THE CAUSE OF THE EVENT IS UNDETERMINED. IT COULD NOT BE CONFIRMED IF THE CAUSE OF DEATH WAS RELATED TO THE DEVICE OR PROCEDURE. NO ADDITIONAL CLINICAL SEQUELAE WERE PROVIDED AND THE PATIENT IS EXPIRED. Manufacturer Narrative: FILM ANALYSIS THE REPORTED EVENTS COULD NOT BE ASSESSED ON THE FILMS PROVIDED. ANALYSIS OF THE RETURNED STILLS DID NOT REVEAL ANY OUT OF SPECIFICATION STENT GRAFT INTEGRITY ISSUES WITH THE IMPLANTED ENDURANT STENT GRAFT SYSTEM. THE CAUSE OF THE EVENTS COULD NOT BE DETERMINED. CONTINUATION OF D10: SECTION D INFORMATION REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER RELEVANT DEVICE(S) ARE: PRODUCT ID: ETLW1616C124EE, SERIAL/LOT #: (B)(6), UBD: 30-AUG-2026, UDI#: (B)(4); PRODUCT ID: ETLW1624C93EE, SERIAL/LOT #: (B)(6), UBD: 14-JUN-2026, UDI#: (B)(4); PRODUCT ID: ETLW1624C124EE, SERIAL/LOT #: (B)(6), UBD: 24-JUN-2026, UDI#: (B)(4). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTION TO H6. ANNEX E CODE E1002 REMOVED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21644583&pc=DQY | 2124215-2025-16477 | 2019-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-03-19 04:00 | DQY | MAMBA | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: BOSTON SCIENTIFIC PERFORMED A RETROSPECTIVE DATA ANALYSIS ON MAMBA USING THE PINC AI HEALTHCARE DATABASE FROM PREMIER. PATIENTS ARE IDENTIFIED THROUGH USE OF MAMBA AS IDENTIFIED IN CHARGE MASTER BILLING. ADVERSE EVENTS WERE IDENTIFIED THROUGH THE RESPECTIVE CPT/ICD-10 CODING. THE PROCEDURES WERE PERFORMED FROM JANUARY 2019 TO DECEMBER 2023. THESE ANALYSES ARE BEING PERFORMED AS A SPECIFIC ACTIVITY TO ASSESS SAFETY AND PERFORMANCE RATES ASSOCIATED WITH THE SUBJECT DEVICE. MAMBA OUTCOMES WERE ASSESSED AGAINST ASAHI CORSAIR PRO, CARAVEL MICROCATHETER, TELEFLEX TURNPIKE CATHETER, AND TERUMO FINECROSS MG CORONARY MICRO-GUIDE CATHETER. RATES OF THE ADVERSE EVENTS INCLUDING DEATH, VESSEL TRAUMA, MYOCARDIAL INFARCTION (MI), STROKE, ABRUPT CLOSURE, BLEEDING, EMBOLISM, CARDIAC TAMPONADE, PERICARDIAL EFFUSION, ARRYTHMIA, AND ACUTE RENAL FAILURE ARE REPORTED BELOW. A TOTAL OF 2,272 PATIENTS WHO UNDERWENT A PROCEDURE USING A MAMBA MICROCATHETER DEVICE WERE INCLUDED IN THE STUDY. THE FOLLOWING IS A SUMMARY OF THOSE RESULTS OVER 5 YEARS (2019 -2023): MORTALITY RATES FOR MAMBA CASES: 14 OUT OF 2,272 PATIENTS RESULTING IN A RATE OF 0.62%, COMPARED TO THE COMPETITOR RATE OF 0.24%-1.41%. MORTALITY RATES FOR MAMBA FALL WITHIN THE RANGE OF OR LOWER THAN THOSE OF THE COMPETITORS IN YEARS 2020-2023. RATES FOR MORTALITY ARE THEREFORE COMPARABLE TO OR ARE LOWER THAN THOSE OF THE COMPETITIVE PRODUCTS AND ARE THEREFORE ACCEPTABLE. Manufacturer Narrative: B2 DATE OF DEATH AND B3 DATE OF EVENT: DATA WAS PROVIDED FOR EVENTS RECORDED (B)(6) 2019 -(B)(6) 2023. (B)(6) 2019 SELECTED AS THE EARLIEST POSSIBLE DATE OF EVENT AND DATE OF DEATH. DATA OBTAINED FOR THIS STUDY IS DE-IDENTIFIED BEFORE BEING ACCESSED BY BOSTON SCIENTIFIC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO OUR ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE STUDY DATA DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21644648&pc=DSQ | 2916596-2025-01427 | 2025-02-16 05:00 | Death | THORATEC CORPORATION | 2025-03-19 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Muscle Weakness; Obstruction/Occlusion | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD A REVERSAL OF SEDATION ON (B)(6)2025 AFTER LEFT VENTRICULAR ASSIST DEVICE (LVAD) IMPLANTATION. THE PROVIDERS NOTED A LEFT GAZE DEVIATION AND RIGHT SIGHTED WEAKNESS. THE COMPUTED TOMOGRAPHY (CT) SHOWED LEFT MIDDLE CEREBRAL ARTERY (MCA) INFARCTION, AND THE COMPUTED TOMOGRAPHIC ANGIOGRAPHY (CTA) SHOWED COMPLETE OCCLUSION OF LEFT M1. A FULL NEUROLOGIC EXAM WAS UNABLE TO BE DONE DUE TO THEIR CONDITION. THE PATIENT WAS NOT A CANDIDATE FOR TNKASE (TENECTEPLASE) DUE TO SURGERY AND LARGE INFARCT SIZE. A REPEAT CT ON (B)(6) 2025 SHOWED COMPLETE EFFACEMENT OF THE LEFT VENTRICLE AND MIDLINE SHIFT. THE PATIENT'S FAMILY DECLINED DECOMPRESSIVE HEMICRANIECTOMY AS IT WAS NOT IN LINE WITH THEIR GOALS OF CARE. THE PATIENT HAD A PERCUTANEOUS VENTRICULAR ASSIST DEVICE (IMPELLA) PLACEMENT PRIOR TO THE IMPLANT THAT HAD COMPLICATIONS WHICH WAS THE REASON FOR URGENT LVAD PLACEMENT WHICH HAD PREVIOUSLY BEEN SCHEDULED FOR A FEW DAYS LATER. IT WAS ELECTED TO TRANSITION TO COMFORT CARE DUE TO DEVASTATING CEREBROVASCULAR ACCIDENT (CVA). THE PATIENT PASSED PEACEFULLY ON (B)(6) 2025. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF STROKE COULD NOT BE CONCLUSIVELY DETERMINED. STROKE IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. STROKE HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21644847&pc=OZD | 1220648-2025-26827 | 2025-01-26 05:00 | Death | ABIOMED, INC. | 2025-03-19 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: IT WAS REPORTED VIA ABIOMED¿S LONG-TERM OUTCOME AND QUALITY INDICATOR (LOQI) IMPELLA REGISTRY THAT A PATIENT EXPERIENCED A GASTROINTESTINAL BLEED (GI) AND EXPIRED ON IMPELLA 5.5 SUPPORT. BOTH THE GI BLEED AND DEATH WERE BELIEVED TO BE POSSIBLY RELATED TO THE USE OF THE IMPELLA. NO ADDITIONAL DETAILS WERE PROVIDED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21644868&pc=DSQ | 2916596-2025-01720 | 2025-03-05 05:00 | Death | THORATEC CORPORATION | 2025-03-19 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO AN UNKNOWN REASON. WHETHER THE PATIENT OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED. THE PUMP WAS NOT EXPLANTED. NO FURTHER INFORMATION WAS PROVIDED PER HOSPITAL POLICY. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU). SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21644953&pc=KDI | 3003464075-2025-00055 | 2025-02-22 05:00 | Death | NXSTAGE MEDICAL, INC. | 2025-03-19 04:00 | KDI | NXSTAGE SYSTEM ONE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | K170469 |  | 1 | Event Description: A REPORT WAS RECEIVED ON 24 FEB 2025 FROM THE HOME THERAPY NURSE (HTN) OF A 62-YEAR-OLD MALE PATIENT APPROVED FOR PERFORMING SOLO HEMODIALYSIS WITH A MEDICAL HISTORY INCLUDING END STAGE RENAL DISEASE, WHO STATED THE PATIENT EXPIRED DURING A HOME HEMODIALYSIS TREATMENT ON (B)(6) 2025. ADDITIONAL INFORMATION WAS RECEIVED ON 25 FEB 2025 FROM THE HOME THERAPY NURSE (HTN) WHO STATED THE PATIENT WAS FOUND UNRESPONSIVE. EMERGENCY MEDICAL SERVICES (EMS) WAS DIALED AND UPON ARRIVAL, THE PATIENT WAS PRONOUNCED DECEASED. PER THE HTN, THE CAUSE OF THE EVENT WAS DUE TO THEIR UNDERLYING CARDIAC AND PULMONARY ISSUES, UNRELATED TO NXSTAGE THERAPY. Manufacturer Narrative: THE DEVICE WAS NOT RECEIVED FOR EVALUATION. A REVIEW OF THE DEVICE HISTORY RECORD (DHR) WAS CONDUCTED WHICH CONFIRMED THAT THE DEVICE MET ALL QUALITY CRITERIA AND MANUFACTURING SPECIFICATIONS PRIOR TO RELEASE. THERE WAS NO INDICATION OF A DEVICE MALFUNCTION FROM THE AVAILABLE INFORMATION. THE NXSTAGE SYSTEM ONE USER GUIDE OUTLINES RISKS ASSOCIATED WITH PERFORMING HEMODIALYSIS THERAPY AND WARNS ALL TREATMENTS MUST BE ADMINISTERED UNDER A PHYSICIAN'S PRESCRIPTION AND MUST BE PERFORMED BY A TRAINED AND QUALIFIED PERSON, CONSIDERED TO BE COMPETENT IN THE USE OF THIS DEVICE BY THE PRESCRIBING PHYSICIAN. THE NXSTAGE SYSTEM ONE USER GUIDE SUPPLEMENT FOR SOLO HOME HEMODIALYSIS OUTLINES RISKS, WARNINGS, AND REQUIREMENTS FOR THE SOLO HOME HEMODIALYSIS PATIENT. UDI: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21645151&pc=FNJ | 1423537-2025-00130 |  | Death | CARDINAL HEALTH | 2025-03-19 04:00 | FNJ | UNKNOWN PRODUCT-PR | Adverse Event Without Identified Device or Use Problem | Unspecified Infection |  |  | 1 | Event Description: PLAINTIFF (B)(6) ALLEGES THAT HER HUSBAND, (B)(6), WAS HOSPITALIZED WITH COVID-19 AT (B)(6) ON (B)(6) 2021. WHILE THERE, MR. (B)(6) ALLEGEDLY WAS EXPOSED TO AND INFECTED WITH BURKHOLDERIA CEPACIA BACTERIA FROM CONTAMINATED PRODUCTS, SOME OF WHICH ALLEGEDLY WERE LATER RECALLED DUE TO CONTAMINATION AND RESULTING INFECTIONS AND DEATHS. NO PRODUCTS ARE IDENTIFIED IN THE COMPLAINT. PLAINTIFF ASSERTS A VARIETY OF PRODUCTS LIABILITY CLAIMS. Manufacturer Narrative: PRODUCT CODE FNJ WAS USED ON THIS REPORT DUE TO A SYSTEM REQUIREMENT FOR USE OF A PRODUCT CODE. HOWEVER, THE TRUE PRODUCT CODE COULD NOT BE DETERMINED AS THERE WAS NO PRODUCT INFORMATION RECEIVED WITH THE COMPLAINT. A NON-CONFORMANCE REVIEW COULD NOT BE PERFORMED BECAUSE NO LOT INFORMATION WAS PROVIDED WITHIN THE COMPLAINT. PRIOR TO A LOT¿S RELEASE, THE LOT MUST BE DEEMED ACCEPTABLE BY-PASSING INSPECTIONS THAT ARE BASED ON A VALID SAMPLING PLAN. DURING PRODUCTION, INSPECTORS ROUTINELY EXAMINE A STATISTICAL SAMPLE BOTH PHYSICALLY AND VISUALLY. THERE WERE NO DEVICES RETURNED FOR EVALUATION; THEREFORE, THE AFFECTED PRODUCT(S) COULD NOT BE EVALUATED. AS SUCH, A ROOT CAUSE COULD NOT BE DETERMINED. WE WILL CONTINUE TO MONITOR RELATED REPORTS TO DETERMINE IF ADDITIONAL ACTIONS ARE NECESSARY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21645153&pc=MHX | 1218950-2025-000173 | 2025-02-14 05:00 | Death | PHILIPS MEDICAL SYSTEMS | 2025-03-19 04:00 | MHX | PATIENT INFORMATION CENTER IX | Defective Alarm | Insufficient Information | K183387 |  | 1 | Event Description: IT WAS REPORTED THE CUSTOMER REQUESTED ASSISTANCE WITH THE CLINICAL AUDIT LOGS FOLLOWING A PATIENT DEATH. THE NURSE MANAGER WANTED TO CONFIRM WHETHER THE ALARMS WERE GENERATED AT THE TIME OF THE EVENT. REMOTE SERVICE WAS ABLE TO REVIEW THE AUDIT LOGS WITH THE CUSTOMER, REVEALING THE SPO2 ALARMS WERE TURNED OFF FOR THIS PATIENT BY THE USER AND NOT TURNED BACK ON PRIOR TO THE EVENT. Manufacturer Narrative: A REMOTE SERVICE ENGINEER (RSE) SPOKE TO THE CUSTOMER AND DETERMINED THAT THE SPO2 ALARMS WERE TURNED OFF AT THE TIME OF THE EVENT. BASED ON THE INFORMATION AVAILABLE AND THE TESTING CONDUCTED, THE CAUSE OF THE REPORTED PROBLEM WAS THE USER. THE REPORTED PROBLEM WAS NOT CONFIRMED. THE DEVICE WAS CONFIRMED TO BE OPERATING PER SPECIFICATIONS AND NO FAILURE WAS IDENTIFIED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21645205&pc=OZP | 2032227-2025-138822 | 2025-02-17 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-03-19 04:00 | OZP | PUMP MMT-1880L MM770G BLE MG | Insufficient Flow or Under Infusion; Adverse Event Without Identified Device or Use Problem | Myocardial Infarction; Dementia | P160017 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2025. THE PRIMARY CAUSE OF DEATH WAS HEART ATTACK. THE CUSTOMER WAS HOSPITALIZED ON (B)(6) 2025 WITH A BLOOD GLUCOSE VALUE OF ABOVE 500 MG/DL. THE PUMP WAS NOT WORN AT THE TIME OF PASSING AS IT WAS REMOVED AT THE TIME OF HOSPITALIZATION ON (B)(6) 2025 AND ALSO NOTICED THAT REASON FOR NOT WORN PUMP THAT IT WAS NOT WORKING. THE CUSTOMER ALSO HAD A DEMENTIA WHICH WAS ALSO A CONTRIBUTION FACTORS THAT MAY LEAD TO THIS EVENT. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: MMT-1884, MMT-242A, MMT-332A, MMT-7841ZN, MMT-7040A. TROUBLESHOOTING WAS PERFORMED AND IT WAS NOTICED THAT THE CUSTOMER WAS USING THE SENSOR BUT IT WAS NOT WORN DURING THE EVENT. NO PRODUCT RETURN IS REQUIRED FOR MMT-1884. NO PRODUCT RETURN IS REQUIRED FOR MMT-242A. NO PRODUCT RETURN IS REQUIRED FOR MMT-332A. NO PRODUCT RETURN IS REQUIRED FOR MMT-7841ZN. NO PRODUCT RETURN IS REQUIRED FOR MMT-7040A. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21645282&pc=CBK | 2518422-2025-103169 | 2025-03-13 04:00 | Death | RESPIRONICS, INC. | 2025-03-19 04:00 | CBK | TRILOGY | Adverse Event Without Identified Device or Use Problem; Insufficient Information | Insufficient Information | K093905 |  | 1 | Event Description: BASED ON THE INFORMATION CURRENTLY PROVIDED AND AVAILABLE, IT WAS ALLEGED THAT DURING NON-INVASIVE CLINICAL AND THERAPEUTIC USE OF THE DEVICE, THE PATIENT WAS FOUND UNRESPONSIVE AND EXPIRED. IT WAS REPORTED THAT THE DEVICE'S APNEA ALARM HAD NOT TRIGGERED, WITH FURTHER NOTATION STATING THAT THE DEVICE BACKUP RATE WAS SET TO 12 BPM AND IT WAS OBSERVED THE PATIENT WAS RECEIVING 30 BPM. IT WAS FURTHER NOTED THAT THE PATIENT HAD BEEN ON THIS SPECIFIC WARD FOR 21 DAYS. THE REPORTER HAS REQUESTED FURTHER GUIDANCE ON DETERMINING IF THE APNEA ALARM WAS ACTIVATED/MUTED/DISABLED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM DETAILS OF THE DEVICE'S OPERATION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. CAUSE AND/OR CONTRIBUTION OF THE DEVICE TO THE PATIENT OUTCOME CANNOT BE CONFIRMED NOR REFUTED, AS FURTHER GOOD FAITH EFFORTS ARE REQUIRED. Manufacturer Narrative: THE BECOME AWARE DATE HAS BEEN CORRECTED FROM 03/14/2025 TO 03/13/2025. Manufacturer Narrative: THE TRILOGY 202 VENTILATOR WAS RETURNED TO A 3RD-PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS SENT IN FOR INVESTIGATION FOLLOWING A PATIENT INCIDENT. THE CUSTOMER STATED THEY BELIEVED THE APNEA ALARM DID NOT TRIGGER AND WANTED THE DEVICE INSPECTED. THE UNIT WAS INSPECTED VISUALLY FOR ANY DAMAGE AND THE ERRORS LOGS CHECKED. NO COSMETIC DAMAGE WAS FOUND. THE ERROR LOGS SHOWED THAT ON (B)(6) 2024, THE APNEA ALARM WAS TRIGGERED FOLLOWED BY THE LOW CIRCUIT LEAK ALARM. THE ALARM PAUSE KEY WAS PRESSED. THIS ACTION PAUSES THE AUDIO FOR ANY FURTHER ALARMS UNTIL THE RESET KEY IS PRESSED, OR UNTIL THE DEVICE IS TURNED OFF AND BACK ON AGAIN. THE PRODUCT VALIDATION TEST WAS PERFORMED AT THE REQUEST OF PHILIPS TO CHECK THE FUNCTIONALITY OF THE DEVICE. THE UNIT PASSED THIS TEST WITH NO FAULTS FOUND. THE UNIT HAS ALSO BEEN RUN WITH THE APNEA ALARM TURNED ON AND A SIMULATED SCENARIO WAS CARRIED OUT TO CHECK THAT THE ALARM FUNCTIONED AS EXPECTED. THE ALARM TRIGGERED AFTER 10 SECONDS AS SET IN THE MENU AND WHEN RESET AND TRIED AGAIN, THE ALARM TRIGGERED AGAIN AS EXPECTED. THE UNIT COMPLETED AND PASSED ALL TESTING AND IS DETERMINED TO BE FUNCTIONING AS EXPECTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21645579&pc=MCX | 2124215-2025-16729 | 2025-03-04 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-03-19 04:00 | MCX | ROTAPRO? | Adverse Event Without Identified Device or Use Problem | Vascular Dissection | P900056 |  | 1 | Event Description: IT WAS REPORTED THAT A DISSECTION AND DEATH OCCURRED. THE TARGET LESION WAS LOCATED IN THE PROXIMAL LEFT ANTERIOR DESCENDING (LAD) ARTERY. A 1.50MM ROTAPRO WAS SELECTED FOR USE IN A PERCUTANEOUS CORONARY INTERVENTION (PCI) PROCEDURE. THE PATIENT WAS NOT DOING WELL AFTER THE PROCEDURE. A TYPE E OR F DISSECTION OCCURRED IN THE MID LEFT ANTERIOR DESCENDING ARTERY, WHICH WAS FLOW LIMITING, WITH NO FLOW RIGHT AFTER. THE PROCEDURE WAS ABORTED DUE TO COMPLICATIONS, AND THE PATIENT DIED POST-PROCEDURE.Event Description: IT WAS REPORTED THAT A DISSECTION AND DEATH OCCURRED. THE TARGET LESION WAS LOCATED IN THE PROXIMAL LEFT ANTERIOR DESCENDING (LAD) ARTERY. A 1.50MM ROTAPRO WAS SELECTED FOR USE IN A PERCUTANEOUS CORONARY INTERVENTION (PCI) PROCEDURE. THE PATIENT WAS NOT DOING WELL AFTER THE PROCEDURE. A TYPE E OR F DISSECTION OCCURRED IN THE MID LEFT ANTERIOR DESCENDING ARTERY, WHICH WAS FLOW LIMITING, WITH NO FLOW RIGHT AFTER. THE PROCEDURE WAS ABORTED DUE TO COMPLICATIONS, AND THE PATIENT DIED POST-PROCEDURE. IT WAS FURTHER REPORTED THAT THE 98% STENOSED TARGET LOCATION WAS LOCATED IN THE NOT TORTUOUS, SEVERELY CALCIFIED SMALL VESSEL. THE AUTOPSY STATUS IS UNKNOWN, AND THE OFFICIAL CAUSE OF DEATH HAS NOT BEEN DETERMINED. THERE IS NO INFORMATION ON THE PATIENT'S MEDICAL HISTORY. DURING THE PROCEDURE, SHOCKWAVE THERAPY AND MEDTRONIC STENTS WERE USED, BUT NO IVUS WAS PERFORMED. DETAILS ABOUT OTHER DEVICES USED ARE UNAVAILABLE. THE STAFF STATED THAT WHILE THEY DIDN'T THINK THE ROTAPRO DEVICE CAUSED THE COMPLICATION, THEY WERE UNSURE. THE PHYSICIAN ATTEMPTED TO STENT THE ENTIRE ARTERY, BUT IT SPIRALLY DISSECTED, REQUIRING TWO OR THREE MEDTRONIC STENTS. THE SIZES AND LENGTHS OF THE STENTS ARE UNKNOWN. THE PHYSICIAN CONSIDERED USING A BALLOON PUMP BUT DECIDED AGAINST IT AND DECLINED TO USE A NON-BOSTON SCIENTIFIC DEVICE. NO FURTHER INFORMATION HAS BEEN SHARED BY THE PHYSICIAN. Manufacturer Narrative: E1 - INITIAL REPORTER ADDRESS 1: (B)(6). Manufacturer Narrative: E1 - INITIAL REPORTER ADDRESS 1: (B)(6).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21645922&pc=DSQ | 2916596-2025-01933 | 2025-02-26 05:00 | Death | THORATEC CORPORATION | 2025-03-19 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Fall; Head Injury; Intracranial Hemorrhage | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD A CLOSED HEAD INJURY POST FALL THAT OCCURRED ON (B)(6) 2025 AT HOME AT NIGHT. THE PATIENT REPORTED TO THE EMERGENCY DEPARTMENT ON (B)(6) 2025. THE PATIENT PASSED AWAY DUE TO INTRACEREBRAL HEMORRHAGE. THE EVENT AND OUTCOME WERE NOT CONSIDERED DEVICE OR THERAPY RELATED; THE DEVICE OPERATED AS INTENDED.Event Description: IT WAS REPORTED THAT AN AUTOPSY WAS NOT PERFORMED. THE CAUSE OF THE FALL WAS ALSO SAID TO BE UNKNOWN SINCE IT WAS UNWITNESSED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENT COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. IT WAS REPORTED THAT THE DEVICE WILL NOT BE EXPLANTED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU), DOCUMENT (B)(4), REV. C IS CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ OF THIS DOCUMENT LISTS BLEEDING AND DEATH AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. ADDITIONALLY, SECTION 6 ¿PATIENT CARE AND MANAGEMENT¿ PROVIDES INFORMATION REGARDING THE RECOMMENDED ANTICOAGULATION THERAPY AND INR (INTERNATIONAL NORMALIZED RATIO) RANGE, AS WELL AS CONSIDERATIONS FOR WHEN THERE IS A RISK OF BLEEDING. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21646559&pc=DSQ | 2916596-2025-01531 | 2024-12-20 05:00 | Death | THORATEC CORPORATION | 2025-03-19 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. THE OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. NO AUTOPSY WAS PERFORMED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: SECTION A2, A4: PATIENT AGE AND WEIGHT NOT PROVIDED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. DUE TO HOSPITAL PRIVACY POLICY, NO ADDITIONAL INFORMATION WILL BE PROVIDED. THE HEARTMATE 3 LVAS IFU, REV. G IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: E1, ADDRESS - LINE 1: (B)(6). NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21646590&pc=NIO | 2024168-2025-02870 | 2014-01-01 05:00 | Death | ABBOTT VASCULAR | 2025-03-19 04:00 | NIO | ABSOLUTE PRO LL SELF-EXPANDING STENT SYSTEM | Adverse Event Without Identified Device or Use Problem | Unspecified Vascular Problem | P110028 |  | 1 | Event Description: IT WAS REPORTED IN A RESEARCH SUMMARY ARTICLE THAT THE AIM OF THE STUDY WAS TO EVALUATE THE SHORT- AND MID-TERM RESULTS OF THE EFFECTIVENESS OF LASER CUT NITINOL STENT (CNS) AND INTERWOVEN NITINOL STENT (INS) IN THE REVASCULARIZATION OF LONG FEMOROPOPLITEAL OCCLUSIONS. THE DEVICES USED WERE SUPERA SELF EXPANDING STENTS (INTERWOVEN NITINOL STENT) AND ABSOLUTE PRO LL SELF EXPANDING STENTS (LASER CUT NITINOL STENT). BETWEEN JANUARY 2012 AND NOVEMBER 2020, 126 PATIENTS WITH SUPERA INS AND 125 PATIENTS WITH ABSOLUTE PRO CNS WERE REVIEWED. THE TECHNICAL SUCCESS WAS 100% IN BOTH GROUPS AND NO DEATHS WERE NOTED IN THE PERIOPERATIVE PERIOD IN BOTH GROUPS; HOWEVER 2 DISSECTIONS OCCURRED INTEROPERATIVELY IN THE ABSOLUTE PRO CNS GROUP. IN THE 30 DAY POSTOPERATIVE PERIOD, THERE WERE 8 CASES OF IN STENT THROMBOSIS IN THE SUPERA INS GROUP AND 4 CASES OF IN STENT THROMBOSIS IN THE ABSOLUTE PRO CNS GROUP. THERE WAS ONE POST PUNCTURE HEMATOMA REQUIRING OPEN SURGERY IN THE SUPERA INS GROUP. IN THE ABSOLUTE PRO CNS GROUP THERE WERE 4 POST PUNCTURE HEMATOMAS REQUIRING SURGICAL INTERVENTION. HOSPITALIZATION IN THE SUPERA INS GROUP WAS 5 DAYS WHILE HOSPITALIZATION IN THE ABSOLUTE PRO CNS GROUP WAS 7 DAYS. MAJOR AND MINOR AMPUTATIONS AND RESTENOSIS OCCURRED IN BOTH GROUPS. THERE WERE NO STENT FRACTURES IN THE SUPERA INS GROUP, WHILE THERE WERE 37 STENT FRACTURES IN THE ABSOLUTE PRO CNS GROUP IN THE TWO YEAR OUTCOMES. REINTERVENTION WAS REQUIRED IN BOTH GROUPS AT 2 YEAR OUTCOMES. BOTH GROUPS HAD DEATH IN THE TWO YEAR OUTCOMES. AFTER PROPENSITY SCORE MATCHING, INS SHOWED COMPARABLE RESULTS WITH CNS FOR THE WHOLE COHORT. HOWEVER, INS SEEMS TO ACHIEVE BETTER OUTCOMES FOR FEMOROPOPLITEAL LESION EXTENDED TO THE BELOW-THE-KNEE LEVEL. ADDITIONAL INFORMATION CAN BE FOUND IN THE ARTICLE "INTERWOVEN NITINOL STENT VERSUS LASER-CUT-TYPE NITINOL STENT FOR LONG FEMOROPOPLITEAL OCCLUSIVE DISEASE: A PROPENSITY MATCHED ANALYSIS". Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTIVE ACTIONS (CAPA) REVIEW WAS PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A PRODUCTION RECORD REVIEW AND A REVIEW OF THE COMPLAINT HANDLING DATABASE WAS NOT PERFORMED BECAUSE THE PART AND LOT NUMBER WAS NOT REPORTED. A CONCLUSIVE CAUSE FOR THE REPORTED DEATH AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED. BASED ON THE REPORTED INFORMATION AND RESULTS OF THE COMPLAINT INVESTIGATION THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO THE DESIGN, MANUFACTURE, OR LABELING OF THE DEVICE. B3: DATE OF EVENT - ESTIMATED AS 01/01/2014. D6A: DATE OF IMPLANT- ESTIMATED AS (B)(6)2012. D4: THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. THE ADDITIONAL DEVICE REFERENCED IN B5 IS FILED UNDER A SEPARATE MEDWATCH REPORT NUMBER. THE ADDITIONAL PATIENT EFFECTS AND MALFUNCTIONS REPORTED IN THE ARTICLE(S) ARE CAPTURED UNDER SEPARATE MEDWATCH REPORTS. LITERATURE ATTACHMENT: ARTICLE TITLE: INTERWOVEN NITINOL STENT VERSUS LASER-CUT-TYPE NITINOL STENT FOR LONG FEMOROPOPLITEAL OCCLUSIVE DISEASE: A PROPENSITY MATCHED ANALYSIS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21646767&pc=DSQ | 2916596-2025-01532 | 2025-02-25 05:00 | Death | THORATEC CORPORATION | 2025-03-19 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Renal Failure; Multiple Organ Failure; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY DUE TO MULTI-SYSTEM ORGAN FAILURE. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. THE PATIENT WAS ON EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) PRIOR IMPLANT, POST IMPLANT KIDNEY FAILURE AND SEVERE RIGHT HEART FAILURE. THE PATIENT WAS UNABLE TO WEAN FROM ECMO. PUMP PERFORMED AS INTENDED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED MULTISYSTEM ORGAN FAILURE (MSOF) AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RESPIRATORY FAILURE, RIGHT HEART FAILURE, RENAL FAILURE, HEPATIC DYSFUNCTION) AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 OF THE IFU (UNDER "RIGHT HEART FAILURE") DISCUSSES THE POTENTIAL DEVELOPMENT OF RIGHT HEART FAILURE FOLLOWING IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS, INCLUDING INOTROPES AND RIGHT VENTRICULAR ASSIST DEVICE (RVAD) PLACEMENT. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21646930&pc=NIP | 2024168-2025-02874 | 2014-01-01 05:00 | Death | ABBOTT VASCULAR | 2025-03-19 04:00 | NIP | SUPERA | Adverse Event Without Identified Device or Use Problem | Unspecified Vascular Problem | P120020 |  | 1 | Event Description: IT WAS REPORTED IN A RESEARCH SUMMARY ARTICLE THAT THE AIM OF THE STUDY WAS TO EVALUATE THE SHORT- AND MID-TERM RESULTS OF THE EFFECTIVENESS OF LASER CUT NITINOL STENT (CNS) AND INTERWOVEN NITINOL STENT (INS) IN THE REVASCULARIZATION OF LONG FEMOROPOPLITEAL OCCLUSIONS. THE DEVICES USED WERE SUPERA SELF EXPANDING STENTS (INTERWOVEN NITINOL STENT) AND ABSOLUTE PRO LL SELF EXPANDING STENTS (LASER CUT NITINOL STENT). BETWEEN JANUARY 2012 AND NOVEMBER 2020, 126 PATIENTS WITH SUPERA INS AND 125 PATIENTS WITH ABSOLUTE PRO CNS WERE REVIEWED. THE TECHNICAL SUCCESS WAS 100% IN BOTH GROUPS AND NO DEATHS WERE NOTED IN THE PERIOPERATIVE PERIOD IN BOTH GROUPS; HOWEVER 2 DISSECTIONS OCCURRED INTEROPERATIVELY IN THE ABSOLUTE PRO CNS GROUP. IN THE 30 DAY POSTOPERATIVE PERIOD, THERE WERE 8 CASES OF IN STENT THROMBOSIS IN THE SUPERA INS GROUP AND 4 CASES OF IN STENT THROMBOSIS IN THE ABSOLUTE PRO CNS GROUP. THERE WAS ONE POST PUNCTURE HEMATOMA REQUIRING OPEN SURGERY IN THE SUPERA INS GROUP. IN THE ABSOLUTE PRO CNS GROUP THERE WERE 4 POST PUNCTURE HEMATOMAS REQUIRING SURGICAL INTERVENTION. HOSPITALIZATION IN THE SUPERA INS GROUP WAS 5 DAYS WHILE HOSPITALIZATION IN THE ABSOLUTE PRO CNS GROUP WAS 7 DAYS. MAJOR AND MINOR AMPUTATIONS AND RESTENOSIS OCCURRED IN BOTH GROUPS. THERE WERE NO STENT FRACTURES IN THE SUPERA INS GROUP, WHILE THERE WERE 37 STENT FRACTURES IN THE ABSOLUTE PRO CNS GROUP IN THE TWO YEAR OUTCOMES. REINTERVENTION WAS REQUIRED IN BOTH GROUPS AT 2 YEAR OUTCOMES. BOTH GROUPS HAD DEATH IN THE TWO YEAR OUTCOMES. AFTER PROPENSITY SCORE MATCHING, INS SHOWED COMPARABLE RESULTS WITH CNS FOR THE WHOLE COHORT. HOWEVER, INS SEEMS TO ACHIEVE BETTER OUTCOMES FOR FEMOROPOPLITEAL LESION EXTENDED TO THE BELOW-THE-KNEE LEVEL. ADDITIONAL INFORMATION CAN BE FOUND IN THE ARTICLE "INTERWOVEN NITINOL STENT VERSUS LASER-CUT-TYPE NITINOL STENT FOR LONG FEMOROPOPLITEAL OCCLUSIVE DISEASE: A PROPENSITY MATCHED ANALYSIS". Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTIVE ACTIONS (CAPA) REVIEW WAS PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A PRODUCTION RECORD REVIEW AND A REVIEW OF THE COMPLAINT HANDLING DATABASE WAS NOT PERFORMED BECAUSE THE PART AND LOT NUMBER WAS NOT REPORTED. A CONCLUSIVE CAUSE FOR THE REPORTED DEATH AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED. THE REPORTED PATIENT EFFECT OF DEATH IS LISTED IN THE SUPERA PERIPHERAL STENT SYSTEMS INSTRUCTIONS FOR USE (IFU) AS A POTENTIAL ADVERSE EFFECT OF PERIPHERAL PERCUTANEOUS INTERVENTION. BASED ON THE REPORTED INFORMATION AND RESULTS OF THE COMPLAINT INVESTIGATION THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO THE DESIGN, MANUFACTURE, OR LABELING OF THE DEVICE. B3 - DATE OF EVENT - ESTIMATED AS 01/01/2014. D6A: DATE OF IMPLANT- ESTIMATED AS 01/01/2012. D4 - PRIMARY UDI NUMBER - THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: ARTICLE TITLE: INTERWOVEN NITINOL STENT VERSUS LASER-CUT-TYPE NITINOL STENT FOR LONG FEMOROPOPLITEAL OCCLUSIVE DISEASE: A PROPENSITY MATCHED ANALYSIS. THE ADDITIONAL DEVICE REFERENCED IN B5 IS FILED UNDER A SEPARATE MEDWATCH REPORT NUMBER. THE ADDITIONAL PATIENT EFFECTS REPORTED IN THE ARTICLE ARE CAPTURED UNDER A SEPARATE MEDWATCH REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21647049&pc=NPT | 2025587-2025-01990 | 2025-03-12 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-19 04:00 | NPT | ENVEO PRO DELIVERY SYSTEM | Device Dislodged or Dislocated; Insufficient Information; Patient Device Interaction Problem | Cardiac Arrest; Vascular Dissection; Unspecified Nervous System Problem; No Clinical Signs, Symptoms or Conditions | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT PRIOR TO IMPLANT OF A TRANSCATHETER AORTIC VALVE, A PRE-IMPLANT BALLOON DILATION WAS PERFORMED WITH A 22 MM NON-COMPLIANT BALLOON DUE TO SEVERE CALCIFICATION OF THE VALVE AND A HIGH INITIAL GRADIENT. FOLLOWING DILATION, HYPOTENSION AND BRADYCARDIA WERE OBSERVED. IMMEDIATE ANGIOGRAPHIC CONTROL SHOWED NO ANNULAR RUPTURE OR DISSECTION, AND AN ECHOCARDIOGRAPHIC CHECK CONFIRMED THE ABSENCE OF ANY EFFUSION. DUE TO PERSISTENT HYPOTENSION AND PROGRESSIVE BRADYCARDIA, RESUSCITATION MANEUVERS AND PHARMACOLOGICAL SUPPORT WERE IMPLEMENTED. ONCE HEMODYNAMICS IMPROVED, THE VALVE IMPLANT WAS ATTEMPTED. ON THE FIRST DEPLOYMENT ATTEMPT, THE VALVE DISLODGED INTO THE AORTA, NECESSITATING A SECOND DEPLOYMENT ATTEMPT. ANGIOGRAPHIC CONTROL BEFORE THE SECOND RELEASE REVEALED EVIDENCE OF A DISSECTION AT THE SINOTUBULAR JUNCTION, PROMPTING AN EMERGENCY RELEASE OF THE VALVE. CARDIAC ARREST OCCURRED AND RESUSCITATION MANEUVERS WERE PERFORMED. EXTRACORPOREAL CIRCULATION WAS INITIATED, AND THE PROCEDURE WAS CONVERTED TO SURGICAL INTERVENTION FOR AORTIC VALVE AND ASCENDING AORTA REPLACEMENT.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT DURING THE VALVE IMPLANT, FOLLOWING THE BALLOON DILATION, PULSELESS ELECTRICAL ACTIVITY (PEA) WAS ALSO OBSERVED AND PRIOR TO RESUSCITATION MANEUVERS BEING PERFORMED. ECHOCARDIOGRAM REVEALED SEVERE LEFT VENTRICULAR HYPERTROPHY AND REDUCED PRELOAD FOR SMALL INTRAVENTRICULAR CAVITY. DURING THE FIRST DEPLOYMENT OF THE VALVE, CONTROL ANGIOGRAPHY SHOWED A MYOINTIMAL FLAP LOCALIZED IN AT THE SINOTUBULAR JUNCTION (STJ). ANGIOGRAPHY WAS PERFORMED WITH EVIDENCE OF ASCENDING AORTIC DISSECTION INVOLVING ANONYMOUS TRUNK FROM THE FLAP PREVIOUSLY MENTIONED LIKELY DUE TO THE DISLODGED OF THE PROSTHETIC VALVE DURING THE FIRST RELEASE. FOLLOWING ADMISSION THE INTENSIVE CARE UNIT (ICU), THE PATIENT DIED 19 DAYS LATER. THE CAUSE OF DEATH WAS NOT REPORTED.Event Description: IT WAS REPORTED THAT PRIOR TO IMPLANT OF A TRANSCATHETER AORTIC VALVE, A PRE-IMPLANT BALLOON DILATION WAS PERFORMED WITH A 22 MM NON-COMPLIANT BALLOON DUE TO SEVERE CALCIFICATION OF THE VALVE AND A HIGH INITIAL GRADIENT. FOLLOWING DILATION, HYPOTENSION AND BRADYCARDIA WERE OBSERVED. IMMEDIATE ANGIOGRAPHIC CONTROL SHOWED NO ANNULAR RUPTURE OR DISSECTION, AND AN ECHOCARDIOGRAPHIC CHECK CONFIRMED THE ABSENCE OF ANY EFFUSION. DUE TO PERSISTENT HYPOTENSION AND PROGRESSIVE BRADYCARDIA, RESUSCITATION MANEUVERS AND PHARMACOLOGICAL SUPPORT WERE IMPLEMENTED. ONCE HEMODYNAMICS IMPROVED, THE VALVE IMPLANT WAS ATTEMPTED. ON THE FIRST DEPLOYMENT ATTEMPT, THE VALVE DISLODGED INTO THE AORTA, NECESSITATING A SECOND DEPLOYMENT ATTEMPT. ANGIOGRAPHIC CONTROL BEFORE THE SECOND RELEASE REVEALED EVIDENCE OF A DISSECTION AT THE SINOTUBULAR JUNCTION, PROMPTING AN EMERGENCY RELEASE OF THE VALVE. CARDIAC ARREST OCCURRED AND RESUSCITATION MANEUVERS WERE PERFORMED. EXTRACORPOREAL CIRCULATION WAS INITIATED, AND THE PROCEDURE WAS CONVERTED TO SURGICAL INTERVENTION FOR AORTIC VALVE AND ASCENDING AORTA REPLACEMENT. ADDITIONAL INFORMATION WAS RECEIVED THAT PER THE PHYSICIAN, THE CAUSE OF THE DISSECTION WAS ATTRIBUTED TO THE VALVE DISLODGING DURING THE POSITIONING ATTEMPT AND RELEASE, AND NOT THE DELIVERY CATHETER SYSTEM (DCS).Event Description: IT WAS REPORTED THAT PRIOR TO IMPLANT OF A TRANSCATHETER AORTIC VALVE, A PRE-IMPLANT BALLOON DILATION WAS PERFORMED WITH A 22 MM NON-COMPLIANT BALLOON DUE TO SEVERE CALCIFICATION OF THE VALVE AND A HIGH INITIAL GRADIENT. FOLLOWING DILATION, HYPOTENSION AND BRADYCARDIA WERE OBSERVED. IMMEDIATE ANGIOGRAPHIC CONTROL SHOWED NO ANNULAR RUPTURE OR DISSECTION, AND AN ECHOCARDIOGRAPHIC CHECK CONFIRMED THE ABSENCE OF ANY EFFUSION. DUE TO PERSISTENT HYPOTENSION AND PROGRESSIVE BRADYCARDIA, RESUSCITATION MANEUVERS AND PHARMACOLOGICAL SUPPORT WERE IMPLEMENTED. ONCE HEMODYNAMICS IMPROVED, THE VALVE IMPLANT WAS ATTEMPTED. ON THE FIRST DEPLOYMENT ATTEMPT, THE VALVE DISLODGED INTO THE AORTA, NECESSITATING A SECOND DEPLOYMENT ATTEMPT. ANGIOGRAPHIC CONTROL BEFORE THE SECOND RELEASE REVEALED EVIDENCE OF A DISSECTION AT THE SINOTUBULAR JUNCTION, PROMPTING AN EMERGENCY RELEASE OF THE VALVE. CARDIAC ARREST OCCURRED AND RESUSCITATION MANEUVERS WERE PERFORMED. EXTRACORPOREAL CIRCULATION WAS INITIATED, AND THE PROCEDURE WAS CONVERTED TO SURGICAL INTERVENTION FOR AORTIC VALVE AND ASCENDING AORTA REPLACEMENT. ADDITIONAL INFORMATION WAS RECEIVED THAT PER THE PHYSICIAN, THE CAUSE OF THE DISSECTION WAS ATTRIBUTED TO THE VALVE DISLODGING DURING THE POSITIONING ATTEMPT AND RELEASE, AND NOT THE DELIVERY CATHETER SYSTEM (DCS). ADDITIONAL INFORMATION WAS RECEIVED THAT DURING THE VALVE IMPLANT, FOLLOWING THE BALLOON DILATION, PULSELESS ELECTRICAL ACTIVITY (PEA) WAS ALSO OBSERVED AND PRIOR TO RESUSCITATION MANEUVERS BEING PERFORMED. ECHOCARDIOGRAM REVEALED SEVERE LEFT VENTRICULAR HYPERTROPHY AND REDUCED PRELOAD FOR SMALL INTRAVENTRICULAR CAVITY. DURING THE FIRST DEPLOYMENT OF THE VALVE, CONTROL ANGIOGRAPHY SHOWED A MYOINTIMAL FLAP LOCALIZED IN AT THE SINOTUBULAR JUNCTION (STJ). ANGIOGRAPHY WAS PERFORMED WITH EVIDENCE OF ASCENDING AORTIC DISSECTION INVOLVING ANONYMOUS TRUNK FROM THE FLAP PREVIOUSLY MENTIONED LIKELY DUE TO THE DISLODGED OF THE PROSTHETIC VALVE DURING THE FIRST RELEASE. FOLLOWING ADMISSION THE THE INTENSIVE CARE UNIT (ICU), THE PATIENT DIED 19 DAYS LATER. THE CAUSE OF DEATH WAS NOT REPORTED. ADDITIONAL INFORMATION WAS RECEIVED FROM THE PHYSICIAN TO NOTE A DEEP ANALYSIS WAS PERFORMED OF THE PRE-CASE PLANNING COMPUTED TOMOGRAPHY. THE PHYSICIAN REPORTED THE PATIENT'S UNDERLYING MEDICAL HISTORY WAS A CONTRIBUTING FACTOR TO THE PATIENT'S DEATH BY NOTING IN "A FIRST STEP OF [TRANSCATHETER AORTIC VALVE IMPLANTATION] TAVI PROCEDURE A 'SUICIDE VENTRICLE SYNDROME' OCCURRED; THIS COMPLICATION WAS MANAGED WITH VASOPRESSOR MEDICATION. THE CAUSE OF THE CARDIAC ARREST WAS THE AORTIC DISSECTION. THE PHYSICIAN NOTED THE PATIENT WAS HOSPITALIZED FOR NEUROLOGICAL IMPAIRMENT DUE TO BLOOD HYPOPERFUSION DURING THE TAVI PROCEDURE AND WAS A CAUSE OF THE PATIENT'S DEATH. Manufacturer Narrative: UPDATED H.6 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. A FOURTH PARAGRAPH WAS ADDED. H6. ADDITIONAL CODES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B2 B5 H1 H2 H6 ADDITIONAL CODES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED: B5, H6 - ADDED CODES APPLICABLE TO THE DELIVERY CATHETER SYSTEM (DCS). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME EVOLUT R TRANSCATHETER AORTIC VALVE; PRODUCT ID EVOLUTR-29 (SERIAL: (B)(6)); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE (B)(6) 2025; EXPLANT DATE (B)(6) 2025 SECTION D REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: PRODUCT ID L-ENVPRO-14 (LOT: 0012099633); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE N/A; EXPLANT DATE N/A. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21647135&pc=HAW | 1723170-2025-01384 | 2024-11-18 05:00 | Death | MEDTRONIC NAVIGATION, INC | 2025-03-19 04:00 | HAW | MEDTRONIC NAVIGATION | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Cardiac Arrest; Hemorrhage/Bleeding; Intracranial Hemorrhage; Sepsis; Respiratory Failure; Hydrocephalus; Insufficient Information |  |  | 1 | Event Description: CITATION: TEKLE W, BENITES G, MILLER S, ET AL, J NEUROLNTERVENT SURG EPUB DOI:10.1136LJNIS-2024-022237 SUMMARY: BACKGROUND THERE IS GROWING INTEREST AND EVIDENCE IN SPONTANEOUS INTRACEREBRAL HEMORRHAGE (ICH) EVACUATION WITH MINIMALLY INVASIVE S URGERY (MIS). IF EARLY ICH EVACUATION BECOMES THE STANDARD OF CARE, TRAINING NEUROINTERVENTIONALISTS TO PERFORM MIS WOULD EXPAND GLOBAL ACCESS TOTREATMENT. WE PRESENT A RETROSPECTIVE ANALYSIS OF PATIENTS WHO UNDERWENT MIS-ICH EVACUATION PERFORMED BY INTERVENTIONAL NEUROLOGISTS IN COLLABORATION WITH NEUROSURGEONS. METHOD PATIENTS MEETING PRESPECIFIED CRITERIA UNDERWENT MIS-ICH EVACUATION USING THE ARTEMIS (PENUMBRA) BY AN INTERVENTIONAL NEUROLO GISTNEUROSURGEON TEAM. BASELINE DEMOGRAPHIC, CLINICAL, AND RADIOGRAPHIC CHARACTERISTICS WERECOLLECTED. PROCEDURE LOCATION WAS RECORDED. THE PRIMARY OUTCOME OF INTEREST WAS THE RATE OF SYMPTOMATIC REBLEEDING. SECONDARY OUTCOMES INCLUDED HEMATOMA REDUCTION, SERIOUS ADVERSE EVENTS, LENGTH OF STAY, DISPOSITION, AND DISCHARGE AND 6 MONTH FUNCTIONAL STATUS. RESULTS 19 PATIENTS WERE INCLUDED IN THIS ANALYSIS. ONE THIRD OF CASES WERE PERFORMED IN THE NEUROANGIOGRAPHY SUITE USING INTRAPROCEDURAL FLAT PANEL CT AND THE REST WERE PERFORMED IN THE OPERATING ROOM. ALL WERE PERFORMED UNDER NEURONAVIGATION USING AXIEM (MEDTRONIC-STEALTH-STATION). THERE WAS A MEDIAN 80% HEMATOMA REDUCTION FROM A MEDIAN PREOPERATIVE ICH VOLUME OF 31.1 ML (IQR 26.2-56.4). A POST-PROCEDURAL HEMATOMA VOLUME OF 15 ML WAS ACHIEVED IN 67% OF CASES, COMPARABLE WITH THAT SEEN IN THE ENRICH (EARLY MINIMALLY INVASIVE REMOVAL OF LNTRACEREBRAL HEMORRHAGE) TRIAL (72.7%). NO PATIENTS DEVELOPED SYMPTOMATIC POSTPROCEDURAL HEMATOMA EXPANSION. CONCLUSION THIS STUDY SUGGESTS THAT MIS-ICH EVACUATION CAN BE PERFORMED SAFELY AND EFFECTIVELY BY TRAINED NEUROINTERVENTIONALISTS. OUR EXPERIENCE ALSO SUPPORTS THE ABILITY TO PERFORM MIS-ICH EVACUATION IN THE NEUROANGIOGRAPHY SUITE. WE ADVOCATE FOR THE DEVELOPMENT OF A STANDARDIZED NEUROINTERVENTIONAL TRAINING PROTOCOL AND CERTIFICATION PATHWAY FOR THE PERFORMANCE OF MIS-ICH EVACUATION WITH THE GOAL OF IMPROVING GLOBAL ACCESS TO CARE. REPORTED EVENT(S): THIS ARTICLE DOCUMENTED THE CASES OF 19 PATIENTS. THE MEDIAN AGE OF THE PATIENTS WAS 62 YEARS, AND 47% WERE WOMEN. THREE PATIENTS DIED WITHIN 7 DAYS. ONE PATIENT UNDERWENT CRANIECTOMY >72 HOURS POST-MIS FOR SYMPTOMATIC CEREBRAL EDEMA IMMEDIATELY AFTER WHICH HE DEVELOPED MASSIVE ICH WITH INTRAVENTRICULAR EXTENSION, HYDROCEPHALUS, AND MIDLINE SHIFT, AND PROGRESSED TO BRAIN DEATH. THE SECOND PATIENT DIED FROM MULTIORGAN FAILURE IN THE SETTING OF SEPSIS. THE THIRD PATIENT DIED LATER DURING HOSPITALIZATION FROM CARDIAC ARREST IN THE SETTING OF HYPOXIC RESPIRATORY FAILURE DUE TO COVID-19 PNEUMONIA. NO OTHER SERIOUS ADVERSE EVENTS WERE RECORDED. Manufacturer Narrative: E0117: HYDROCEPHALUS E0118: INTRACRANIAL HEMORRHAGE E0306: SEPSIS E0742: HYPOXIC RESPIRATORY FAILURE E1901: PNEUMONIA E2401: MULTIORGAN FAILURE, DEATH E0506: INTRAVENTRICULAR EXTENSION E0602: CARDIAC ARREST H3, H6: NO PRODUCTS WERE RETURNED TO MEDTRONIC FOR ANALYSIS. CODES B17, C20, AND D15 ARE APPLICABLE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21647147&pc=BTR | 3000219639-2025-00035 | 2025-02-17 05:00 | Death | PARKER MEDICAL | 2025-03-19 04:00 | BTR | UNKNOWN, PARKER ENDOTRACHEAL TUBE | Break | Brain Injury | UNKNOWN |  | 1 | Event Description: THE PATIENT'S FATHER REPORTED, AN UNSPECIFIED AIR LIFE 'PRODUCT WITH A TUBE FAILED/BROKE; IT HAD A RARE BREAK [INTERNAL BREAK] IN IT AND IT CAUSED MY DAUGHTER TO END UP IN A VEGETATIVE STATE [BRAIN DAMAGED] WHERE I JUST LET HER GO. THE EVENT REPORTEDLY OCCURRED AT (B)(6) HOSPITAL; NO ADDITONAL INFORMATION WAS PROVIDED CONCERNING ANY OTHER PERFORMED MEDICAL INTERVENTION(S). Manufacturer Narrative: THE ACTUAL COMPLAINT PRODUCT WAS NOT RETURNED FOR EVALUATION. A REVIEW OF THE DEVICE HISTORY RECORD IS NOT POSSIBLE AS NO LOT NUMBER WAS PROVIDED. WITHOUT A SAMPLE TO PERFORM FUNCTIONAL EVALUATION OR PHOTOGRAPHIC /VIDEO EVIDENCE THE ROOT CAUSE COULD NOT BE DETERMINED. ALL INFORMATION REASONABLY KNOWN AS OF 19 MAR 2025 HAS BEEN INCLUDED IN THIS HEALTH AUTHORITY REPORT. SHOULD ADDITIONAL INFORMATION BE OBTAINED, A FOLLOW-UP HEALTH AUTHORITY REPORT WILL BE PROVIDED. THE INFORMATION PROVIDED BY SUN MED HOLDINGS LLC. REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO SUN MED HOLDINGS LLC. SUN MED HOLDINGS LLC. HAS NO INDEPENDENT KNOWLEDGE OF THE EVENT REPORTED BUT IS RELAYING THE INFORMATION THAT WAS PROVIDED BY THE USER FACILITY WHERE THE INCIDENT OCCURRED. THIS PRODUCT INCIDENT IS DOCUMENTED IN THE SUN MED HOLDINGS COMPLAINT DATABASE AND IDENTIFIED AS COMPLAINT- (B)(4). THIS INFORMATION IS SUBMITTED PURSUANT TO 21CFR803, IN COMPLIANCE WITH THE MEDICAL DEVICE REPORTING REQUIREMENT AND SHOULD NOT BE CONSIDERED TO BE AN ADMISSION THAT AN SUN MED HOLDINGS LLC. PRODUCT IS DEFECTIVE OR CAUSED SERIOUS INJURY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21647254&pc=NPT | 2015691-2025-02069 | 2025-02-21 05:00 | Death | EDWARDS LIFESCIENCES | 2025-03-19 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Obstruction/Occlusion | P140031 |  | 1 | Event Description: AS LEARNED THROUGH EDWARDS IMPLANT PATIENT REGISTRY, A PATIENT WITH A 23MM SAPIEN 3 ULTRA AORTIC VALVE EXPIRED ON THE DAY OF IMPLANT. ADDITIONAL INFORMATION PROVIDED IN THE PATIENT PORTAL REGISTRATION INDICATED "CAUSE OF DEATH: LEFT MAIN OCCLUSION, HEMODYNAMIC INSTABILITY, ASYSTOLE, APNEIC. LEFT MAIN CORONARY STENTING REQUIRED, PERICARDIAL EFFUSION, IMPELLA PLACEMENT." PER RECEIVED RECORDS, THE PATIENT UNDERWENT TRANSCATHETER AORTIC VALVE REPLACEMENT WITH EDWARDS SAPIEN 23 MM VALVE VIA RIGHT COMMON FEMORAL ARTERY ACCESS. AFTER TAVR DEPLOYMENT, THE PATIENT BECAME HEMODYNAMICALLY UNSTABLE AND HYPOTENSIVE WITH SYSTOLIC PRESSURES IN THE 40S. REPEAT ANGIOGRAPHY SHOWED OCCLUSION OF THE LEFT MAIN CORONARY ARTERY. AT THIS TIME A BALLOON WAS DELIVERED AND ANGIOPLASTY PERFORMED WHILE CHEST COMPRESSIONS BEGAN SIMULTANEOUSLY. A STENT WAS THEN DELIVERED INTO THE LEFT MAIN CORONARY ARTERY WITH APPROPRIATE POSITIONING INTO THE AORTIC SINUS. REPEAT ANGIOGRAPHY SHOWED GOOD FLOW IN BOTH THE CIRCUMFLEX AND ANTERIOR DESCENDING DISTRIBUTION. WITH ONGOING CHEST COMPRESSIONS AND RESUSCITATION, THE EDWARDS SHEATH WAS EXCHANGED OVER A WIRE FOR AN IMPELLA CP SHEATH. AN IMPELLA WAS THEN POSITIONED ACROSS THE AORTIC VALVE UNDER FLUOROSCOPY. THE PATIENT MAINTAINED A NATIVE SINUS RHYTHM. AT THIS TIME TTE WAS PERFORMED WHICH SHOWED MODERATE PERICARDIAL EFFUSION. A PERCUTANEOUS PERICARDIOCENTESIS WAS PERFORMED AND INITIAL 1.7 L OF HEMOPERICARDIUM WAS OBTAINED. THE PATIENT'S HEMODYNAMICS IMPROVED WITH THIS. A TEE PROBE WAS THEN PLACED WHICH SHOWED RESOLUTION OF HER PERICARDIAL EFFUSION THOUGH THERE WAS SOME CLOT NOTED IN THE ANTERIOR AND LATERAL ASPECTS OF THE HEART. SOME IMPINGEMENT OF THE RIGHT ATRIUM WAS NOTED. GIVEN HER ONGOING HEMODYNAMIC INSTABILITY, IT WAS ELECTED TO PERFORM A SUBXIPHOID PERICARDIAL WINDOW. SIMULTANEOUSLY ANESTHESIA HAD OBTAINED ENDOTRACHEAL INTUBATION FOR AIRWAY. A SWAN-GANZ CATHETER WAS THEN ADVANCED. THIS CAUSED A COMPLETE HEART BLOCK IN THIS PATIENT WITH A PRE-EXISTING LEFT BUNDLE BRANCH BLOCK. THE TRANSVENOUS PACER WAS THEN REPLACED THROUGH THE RIGHT COMMON FEMORAL VENOUS ACCESS AND THE PATIENT WAS PACED AT 80 BPM WITH APPROPRIATE CAPTURE. AFTER RESUSCITATION INCLUDING 4 PRBC IN ADDITION TO ALBUMIN AND CRYSTALLOID THE PATIENT'S HEMODYNAMICS IMPROVED. SHE REMAINED HEMODYNAMICALLY STABLE AND WAS THEN TRANSFERRED TO THE CCU IN CRITICAL CONDITION. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), CORONARY FLOW OBSTRUCTION IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT PROCEDURE. THE IFU CAUTIONS THAT SAFETY AND EFFECTIVENESS HAVE NOT BEEN ESTABLISHED FOR PATIENTS WITH BULKY CALCIFIED AORTIC VALVE LEAFLETS NEAR CORONARY OSTIA. IN ADDITION, IT WARNS THAT CAUTION SHOULD BE EXERCISED IN IMPLANTING A BIOPROSTHESIS IN PATIENTS WITH CLINICALLY SIGNIFICANT CORONARY ARTERY DISEASE. CORONARY OBSTRUCTION CAN RESULT IN MYOCARDIAL ISCHEMIA OR INFARCTION DUE TO OBSTRUCTION OF THE CORONARY BLOOD FLOW AND MAY REQUIRE INTERVENTION (E.G., PERCUTANEOUS CORONARY INTERVENTION (PCI)). MULTIPLE PATIENT FACTORS COULD PUT THE PATIENT AT RISK FOR CORONARY FLOW OBSTRUCTION DURING THE THV PROCEDURE, INCLUDING SIGNIFICANT UNDERLYING CORONARY ARTERY DISEASE AND BULKY CALCIFICATION OF THE NATIVE LEAFLETS AND ROOT. DISPLACEMENT OF CALCIUM DEPOSITS WITH EMBOLIZATION OF DEBRIS INTO ONE OF THE ARTERIES, OR AORTIC DISSECTION WITH CONTINUITY OF THE RUPTURE INTO THE INTIMA OF ONE OF THE CORONARY OSTIA, CAN RESULT IN THIS COMPLICATION. ADDITIONALLY, MINIMAL DISTANCE BETWEEN THE NATIVE ANNULUS AND THE CORONARY OSTIA, BULKY CALCIFICATION, LONG NATIVE LEAFLETS, AND OBLITERATED CORONARY SINUSES. PROCEDURAL FACTORS SUCH AS TORN NATIVE LEAFLET DURING BALLOON VALVULOPLASTY, PLAQUE SHIFT, DEPLOYMENT OF THE BIOPROSTHETIC HEART VALVE TOO AORTIC, AND SIGNIFICANT VALVE OVER SIZING COULD ALSO CONTRIBUTE TO THIS COMPLICATION. THE EDWARDS THV TRAINING MANUALS ADVISE THE OPERATOR ON PRE-PROCEDURE ASSESSMENT OF THE AORTIC VALVE, ROOT, AND CORONARY ANATOMY. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE TRANSCATHETER HEART VALVE (THV). TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES. THE PHYSICIAN IS INSTRUCTED TO EVALUATE THIS RISK EARLY IN THE PATIENT SCREENING PROCESS IN ALL PATIENTS THE FOLLOWING FACTORS SHOULD BE CONSIDERED: DEGREE OF CALCIFICATION ON LEAFLETS, ANNULUS TO CORONARY OSTIA DISTANCE, LENGTH OF THE VALVE LEAFLET, WIDTH OF THE VALSALVA SINUSES, MOVEMENT OF THE LEAFLETS DURING BAV, PATENCY OF CORONARIES DURING BAV, AND EXPANDED HEIGHT OF THE INTENDED THV. THE TRAINING MANUALS ALSO PROVIDE THE FOLLOWING TIPS FOR DETECTING RISK FOR LEFT MAIN OCCLUSION: (1) AORTOGRAM OR TEE BEFORE THV IMPLANTATION TO REVEAL BULKY CALCIFIED LEAFLETS; (2) DURING PRE-DILATATION, NOTE BULKY CALCIFICATION ON VALVE MOVING TOWARDS OSTIUM ON LEFT MAIN; AND (3) CONSIDER AORTOGRAM DURING VALVULOPLASTY TO ASSESS CORONARY FLOW. THE EDWARDS THV MANUALS ALSO ADVISE THE OPERATOR ON PRE-PROCEDURE ASSESSMENT OF THE AORTIC VALVE, ROOT, AND CORONARY ANATOMY. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE TRANSCATHETER HEART VALVE (THV). TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. BASED ON THE AVAILABLE INFORMATION, THE ROOT CAUSE OF THE CORONARY OBSTRUCTION REMAINS INDETERMINABLE BUT MAY BE RELATED TO THE FACTORS DESCRIBED ABOVE. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. NO CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21648847&pc=GCX | 3011053282-2025-00001 | 2024-10-17 04:00 | Death | LIFEVAC LLC | 2025-03-19 04:00 | GCX | LIFEVAC | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: IT WAS REPORTED THAT A CHILD ALLEGEDLY CHOKED DURING LUNCH AT THE (B)(6) ELEMENTARY SCHOOL. AFTER THAT, WE HAVE CONTACTED (B)(6) AS THEY HAVE PREVIOUSLY PURCHASED LIFEVAC DEVICES FOR DISTRIBUTION TO LOCAL SCHOOLS INCLUDING THE (B)(6) ELEMENTARY SCHOOL IN (B)(6). AS A PRE-CONDITION TO DEVICE PURCHASING, (B)(6) HEALTHCARE CONDUCTED AN INDEPENDENT TESTING OF LIFEVAC TO CONFIRM THE DEVICES WERE ACCEPTABLE. UNFORTUNATELY, (B)(6) WAS NOT ABLE TO PROVIDE ANY FACTUAL INFORMATION ABOUT THE ABOVE REFERENCED EVENT. THEY DID HOWEVER HAVE KNOWLEDGE ON THE CHILD'S OVERALL MEDICAL HISTORY WHICH INCLUDED SPECIAL NEEDS MEDICAL CONDITIONS. OTHER THAN THAT, WE WERE NEVER CONTACTED BY THE SCHOOL, A DEVICE USER, EMS STAFF OR THE CHILD'S FAMILY. NEITHER THE SCHOOL NOR THE HOSPITAL WOULD RELEASE ANY INCIDENT INFORMATION DIRECTLY TO THE DEVICE MANUFACTURER DUE TO THE PRIVACY LAWS. SO THE ONLY SOURCE OF INFORMATION WE HAVE IS WHAT WAS REPORTED IN THE NEWS. WE FILE THIS MDR OUT OF AN ABUNDANCE OF CAUTION. THROUGH THE INFORMATION REPORTED IN THE NEWS, WE HAVE DETERMINED THERE WAS NO FIRST-HAND ACCOUNT OF THE INCIDENT. THE YOUTH RELATIONS DEPUTY WHO WAS INTERVIEWED CONVEYED THE SECOND-HAND REPORT STATING"THE NURSE'S AIDE HAD THE LIFEVAC IN HER HAND AND FRANTICALLY STATED"HE'S CHOKING! "IT'S NOT WORKING." HOWEVER, THERE WAS NO ACTUAL ACCOUNT OF ANYBODY SEEING THE CHILD CHOKE OR THE CHILD ASKING FOR HELP WITH CHOKING. IT WAS ALSO REPORTED THE CHILD WAS LYING MOTIONLESS ON THE BENCH OF THE TABLE WITH HIS TONGUE HANGING OUT OF HIS MOUTH, PALE IN COLOR, UNCONSCIOUS, AND NOT BREATHING. IT APPEARS THE CHILD NEVER GOT UP FROM HIS SEAT. IT IS OUR EXPERIENCE WITH NUMEROUS SCHOOLS INCIDENT REPORTS THAT A CHOKING STUDENT WOULD GET UP AND SEEK HELP. WE HAVE NEVER RECEIVED A REPORT WHERE THE CHOKING STUDENT SITS MOTIONLESSLY UNTIL THEY GO UNCONSCIOUS. WHILE HISTORICALLY, A CHOKING VICTIM WILL BECOME BLUE AS THE HEART WORKS HARDER TO CIRCULATE THE REMAINING OXYGENATED BLOOD, BEING "PALE" IN COLOR IS MORE INDICATIVE OF CARDIAC ARREST, WHERE THE BLOOD BEGINS TO SETTLE. ANOTHER NEWS REPORT ALSO STATED THAT THE HEIMLICH WAS ATTEMPTED BUT FAILED, CHEST COMPRESSIONS WERE TRIED AS WELL BUT FAILED, AND THEY SENT FOR THE AED. THERE WAS NO RECORD OF THE AIRWAY BEING OBSTRUCTED AND THEN CLEARED. AT THAT POINT, AND AT NO OTHER TIME SINCE, WAS THERE ANY INDICATION OR HINT OF THE AIRWAY BEING CLEARED. ADDITIONALLY, NO REPORT MENTIONED THE USE OF THE MIGUEL FORCEPS OR INTUBATION AT THE HOSPITAL. IT IS IMPORTANT TO NOTE THE CHILD SURVIVED FOR TWO DAYS AFTER THE INCIDENT. THAT WOULD NOT BE POSSIBLE WITH AN OBSTRUCTED AIRWAY. IN THE VIDEO INTERVIEW WITH DR. (B)(6) WHO IS THE CHIEF MEDICAL DIRECTOR OF NAPLES FIRE & RESCUE STATED "WE DO NOT KNOW WHAT REALLY CAUSED THE EVENT AND DO NOT NECESSARILY KNOW WHERE THINGS MAY OR MAY NOT HAVE EVEN FAILED. Manufacturer Narrative: LIFEVAC HAS NOT BEEN DIRECTLY CONTACTED BY THE SCHOOL, EMS, OR THE FAMILY. WE WERE NOT ABLE TO RECEIVE ANY INFORMATION DIRECTLY FROM THE SCHOOL OR THE TREATING HOSPITAL. GIVEN THE CIRCUMSTANCES AND THE ABSENCE OF A FIRSTHAND ACCOUNTS AS WELL AS THE SPECIFIC STATEMENT FROM THE CHIEF MEDICAL DIRECTOR, WE DETERMINED THERE WAS NO CLEAR INDICATION THE CHILD CHOKED OR THAT THE LIFEVAC USE WAS APPROPRIATE OR THAT IT WAS INEFFECTIVE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21649157&pc=LWR | 2025587-2025-01992 | 2025-02-13 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-19 04:00 | LWR | FREESTYLE AORTIC ROOT BIOPROSTHESIS | Patient Device Interaction Problem | Rupture | P970031 |  | 1 | Event Description: MEDTRONIC RECEIVED INFORMATION THAT 1 WEEK POST IMPLANT OF THIS 25MM AORTIC BIOPROSTHETIC VALVE, THE PATIENT DIED. THE CAUSE OF DEATH WAS REPORTED AS A FALSE LUMEN RUPTURE FROM AORTIC DISSECTION. IT WAS FURTHER REPORTED THAT THE IMPLANT WAS A COMPLICATED AORTIC CASE. HOWEVER, RELATEDNESS OF THE DEATH TO THE VALVE COULD NOT BE EXCLUDED. IT WAS STATED THAT THE HEALTH CARE PROFESSIONAL (HCP) DOES NOT BELIEVE THAT THE DEATH WAS RELATED TO THE VALVE. IT IS UNKNOWN WHETHER AN AUTOPSY WAS PERFORMED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21649426&pc=OUT | 2029214-2025-00718 | 2025-02-24 05:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-03-19 04:00 | OUT | PIPELINE | Adverse Event Without Identified Device or Use Problem | Ruptured Aneurysm |  |  | 1 | Event Description: LV, M., DONG, L., WANG, C., WU, X., XU, H., WU, X., ZHANG, Y.; QUANT IMAGING MED SURG; 2025; 15(3):1977-1989; MIDTERM OUTCOMES AND COSTS OF PIPELINE EMBOLIZATION DEVICE ALONE VERSUS ATLAS STENT-ASSISTED COILING FOR UNRUPTURED ANTERIOR CIRCULATION ANEURYSMS: A PROPENSITY SCORE MATCHED COMPARATIVE ANALYSIS; DOI.ORG/10.21037/QIMS-24-1581 LITERATURE WAS REVIEWED REGARDING: MIDTERM OUTCOMES AND COSTS OF PIPELINE EMBOLIZATION DEVICE ALONE VERSUS ATLAS STENT-ASSISTED COILING FOR UNRUPTURED ANTERIOR CIRCULATION ANEURYSMS: A PROPENSITY SCORE MATCHED COMPARATIVE ANALYSIS. THE TIME FRAME OF THIS STUDY WAS: JANUARY 2018 TO JUNE 2023. THE FOLLOWING MEDTRONIC DEVICES WERE USED: PIPELINE EMBOLIZATION DEVICE PED DELIVERED AND DEPLOYED THROUGH A MARKSMAN MICROCATHETER. 299 PATIENTS WERE TREATED IN THE GROUP SEPARATED FOR PED USE. DEATHS OCCURRED IN THE STUDY POPULATION: ONE DEATH IN THE PED GROUP DUE TO RUPTURE OF THE TARGET ANEURYSM. AMONG PATIENT ADVERSE EVENTS INCLUDED: IN-STENT STENOSIS (ISS), HIGHER IN THE PED GROUP (4.2%). ONE PATIENT REQUIRED RETREATMENT. IN PED GROUP: 15 (5.0%) PATIENTS WITH COM PLICATIONS: 7 ISCHEMIC (4 TRANSIENT ISCHEMIC ATTACK (TIA), 3 CEREBRAL INFARCTION), 6 HEMORRHAGIC (4 DELAYED ANEURYSM RUPTURE, 2 DISTAL INTRAPARENCHYMAL HEMORRHAGE), 2 EMERGING COMPRESSION SYMPTOMS REFERING TO CRANIAL NEUROPATHY OR BRAINSTEM SYMPTOMS ASSOCIATED WITH ANEURYSM COMPRESSION. EIGHT (2.6%) PATIENTS WERE TREATED WITH OVERLAPPED PED DEVICES. 2 PATIENTS IN THE PED GROUP DISCHARGED WITH POOR NEUROLOGICAL STATUS. FOUR (1.3%) CASES REQUIRED STENT ADJUSTMENT. PATIENTS OF PED GROUP WERE RETREATED IF (I) THE PED MAY FAIL TO FULLY COVER THE ANEURYSM NECK, OR THE DISTAL END OF THE PED MAY FALL INTO THE ANEURYSMAL SAC. Manufacturer Narrative: IMF CODE CORRECTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SEE ATTACHMENTS FOR LITERATURE ARTICLE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21649452&pc=KRA | 2029214-2025-00719 | 2025-02-24 05:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-03-19 04:00 | KRA | MARKSMAN | Adverse Event Without Identified Device or Use Problem | Ruptured Aneurysm |  |  | 1 | Event Description: LV, M., DONG, L., WANG, C., WU, X., XU, H., WU, X., & ZHANG, Y.; QUANT IMAGING MED SURG; 2025; 15(3):1977-1989; MIDTERM OUTCOMES AND COSTS OF PIPELINE EMBOLIZATION DEVICE ALONE VERSUS ATLAS STENT-ASSISTED COILING FOR UNRUPTURED ANTERIOR CIRCULATION ANEURYSMS: A PROPENSITY SCORE MATCHED COMPARATIVE ANALYSIS; DOI.ORG/10.21037/QIMS-24-1581 LITERATURE WAS REVIEWED REGARDING: MIDTERM OUTCOMES AND COSTS OF PIPELINE EMBOLIZATION DEVICE ALONE VERSUS ATLAS STENT-ASSISTED COILING FOR UNRUPTURED ANTERIOR CIRCULATION ANEURYSMS: A PROPENSITY SCORE MATCHED COMPARATIVE ANALYSIS. THE TIME FRAME OF THIS STUDY WAS: JANUARY 2018 TO JUNE 2023. THE FOLLOWING MEDTRONIC DEVICES WERE USED: PIPELINE EMBOLIZATION DEVICE PED DELIVERED AND DEPLOYED THROUGH A MARKSMAN MICROCATHETER. 299 PATIENT WERE TREATED IN THE GROUP SEPARATED FOR PED USE. DEATHS OCCURRED IN THE STUDY POPULATION: ONE DEATH IN THE PED GROUP DUE TO RUPTURE OF THE TARGET ANEURYSM. AMONG PATIENT ADVERSE EVENTS INCLUDED: IN-STENT STENOSIS (ISS), HIGHER IN THE PED GROUP (4.2%). ONE PATIENT REQUIRED RETREATMENT. IN PED GROUP: 15 (5.0%) PATIENTS WITH COM PLICATIONS: 7 ISCHEMIC (4 TRANSIENT ISCHEMIC ATTACK (TIA), 3 CEREBRAL INFARCTION), 6 HEMORRHAGIC (4 DELAYED ANEURYSM RUPTURE, 2 DISTAL INTRAPARENCHYMAL HEMORRHAGE), 2 EMERGING COMPRESSION SYMPTOMS REFERING TO CRANIAL NEUROPATHY OR BRAINSTEM SYMPTOMS ASSOCIATED WITH ANEURYSM COMPRESSION. EIGHT (2.6%) PATIENTS WERE TREATED WITH OVERLAPPED PED DEVICES. 2 PATIENTS IN THE PED GROUP DISCHARGED WITH POOR NEUROLOGICAL STATUS. FOUR (1.3%) CASES REQUIRED STENT ADJUSTMENT. PATIENTS OF PED GROUP WERE RETREATED IF (I) THE PED MAY FAIL TO FULLY COVER THE ANEURYSM NECK, OR THE DISTAL END OF THE PED MAY FALL INTO THE ANEURYSMAL SAC. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21649665&pc=NKM | 2135147-2025-01424 | 2019-03-14 04:00 | Death | ABBOTT MEDICAL | 2025-03-19 04:00 | NKM | MITRACLIP® SYSTEM CLIP DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P100009 |  | 1 | Event Description: THE ARTICLE "THE COAPT SECONDARY MR POST-APPROVAL STUDY ANNUAL PROGRESS REPORT 2025" WAS REVIEWED. THE ARTICLE PRESENTED A PROSPECTIVE MULTI-CENTER STUDY, TO EVALUATE REAL-WORLD SURVEILLANCE TO CONFIRM THE LONG-TERM SAFETY AND EFFECTIVENESS OF THE MITRACLIP SYSTEM IN HEART FAILURE PATIENTS WITH SECONDARY MITRAL REGURGITATION. DEVICES MENTIONED INCLUDE MITRACLIP. THE ARTICLE CONCLUDED THAT THE OUTCOMES FROM THE STUDY SO FAR ARE AS EXPECTED FROM THIS ELDERLY POPULATION WITH HIGH PROPORTION OF COMORBIDITIES AND ARE CONSISTENT WITH OTHER MITRACLIP STUDIES. NO NEW SAFETY RISKS HAVE BEEN OBSERVED DURING THIS STUDY. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS JANANI AIYER, CLINICAL SCIENTIST, CLINICAL AFFAIRS, ABBOTT STRUCTURAL HEART, WITH CORRESPONDING EMAIL JANANI.AIYER@ABBOTT.COM.] THE TIME FRAME OF THE STUDY WAS 14 MARCH 2019 TO 09 MAY 2023. ABBOTT USED CENTERS FOR MEDICARE AND MEDICAID SERVICES (CMS) FEE-FOR-SERVICE (FFS) ANNUAL DATA FOR 2019-2022 AND RECEIVED ACCESS TO CMS FFS QUARTERLY DATA THROUGH 31 DECEMBER 2023 ON 11 SEPTEMBER 2024. A TOTAL OF 5000 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH ALL RECEIVED AN ABBOTT DEVICE. LONG TERM DATA (OVER 12 MONTHS) WAS AVAILABLE IN 2396 OF THE 5000 PATIENTS. THE AVERAGE AGE OF THE PATIENTS ENROLLED WAS 76. THE MAJORITY GENDER WAS MALE. AVERAGE WEIGHT WAS 78 KG. COMORBIDITIES INCLUDED: HEART FAILURE, PRIOR MYOCARDIAL INFARCTION, ANGINA, ISCHEMIC CARDIOMYOPATHY, CARDIOGENIC SHOCK, CARDIAC ARREST, PORCELAIN AORTA, ATRIAL FIBRILLATION, INFECTIVE ENDOCARDITIS, HEART FAILURE HOSPITALIZATION, PERMANENT PACEMAKER, PRIOR PERCUTANEOUS INTERVENTION, PRIOR CORONARY ARTERY BYPASS GRAFT, PRIOR AORTIC VALVE REPAIR/REPLACEMENT, PRIOR MITRAL REPAIR/REPLACEMENT, PRIOR TRICUSPID REPAIR/REPLACEMENT, PRIOR STROKE, TRANSIENT ISCHEMIC ATTACK, CAROTID STENOSIS, PERIPHERAL ARTERIAL DISEASE, CURRENT SMOKER, HYPERTENSION, DIABETES MELLITUS, DIALYSIS, CHRONIC LUNG DISEASE, SUPPLEMENTAL OXYGEN, HOSTILE CHEST, IMMUNOCOMPROMISE, AORTIC REGURGITATION, FUNCTIONAL/DEGENERATIVE/MIXED MITRAL REGURGITATION, MITRAL STENOSIS, MITRAL LEAFLET CALCIFICATION, LEAFLET TETHERING, MITRAL ANNULAR CALCIFICATION. PERI AND POST-PROCEDURAL COMPLICATIONS INCLUDED DEATH, HOSPITALIZATION, STROKE, MITRAL VALVE RE-INTERVENTION, ENDOCARDITIS, MYOCARDIAL INFARCTION, MITRAL VALVE STENOSIS, LEFT VENTRICULAR ASSIST DEVICE (LVAD), RENAL FAILURE, SURGICAL REINTERVENTION. Manufacturer Narrative: BASED ON THE INFORMATION REVIEWED AND DUE TO THE LIMITED INFORMATION AVAILABLE FROM THE ARTICLE AND THE ARTICLE COVERING MULTIPLE PATIENTS, A CAUSE FOR THE REPORTED DEATHS CANNOT BE DETERMINED. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. A REVIEW OF THE LOT HISTORY RECORD COULD NOT BE PERFORMED AS THE PART/LOT INFORMATION WAS NOT PROVIDED. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. THEREFORE, NO REMEDIAL ACTION IS REQUIRED. LITERATURE ATTACHMENT: ARTICLE TITLE: THE COAPT SECONDARY MR POST-APPROVAL STUDY ANNUAL PROGRESS REPORT 2025.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21628374&pc=DSP | 2248146-2025-0000146 | 2025-02-23 05:00 | Death | DATASCOPE CORP. | 2025-03-18 04:00 | DSP | SENSATION PLUS 7.5FR. 40CC IAB & ACCESSORIES | Optical Problem | Cardiogenic Shock | K122628 |  | 1 | Event Description: N/AEvent Description: A FIBER OPTIC FAILURE WAS REPORTED IN THE INTRA-AORTIC BALLOON. HOWEVER, THE IAB WAS SUCCESSFULLY USED WITH A STANDARD ARTERIAL LINE.Event Description: A FIBER OPTIC FAILURE WAS REPORTED IN THE INTRA-AORTIC BALLOON. HOWEVER, THE IAB WAS SUCCESSFULLY USED WITH A STANDARD ARTERIAL LINE. THE PATIENT DIED IN THE ICU. Manufacturer Narrative: UPDATED FIELDS: B4, D9, G3, G6, H2, H3, H6 (TYPE OF INVESTIGATION, INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS), H11. CORRECTED FIELD: B5. NO DECON OR VISUAL INSPECTION PERFORMED SINCE PRODUCT WAS NOT RETURNED AND COULD NOT BE EVALUATED. THEREFORE, WE WERE UNABLE TO CONFIRM OR DETERMINE A ROOT CAUSE FOR THE REPORTED FAILURE. COMPLAINT ID: (B)(4). Manufacturer Narrative: A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF OUR INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21629064&pc=OZO | 2032227-2025-137506 | 2025-01-10 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-03-18 04:00 | OZO | 630G INSULIN PUMP MMT-1715KL 630G | Communication or Transmission Problem; Adverse Event Without Identified Device or Use Problem | Viral Infection; Diabetic Ketoacidosis | P150001 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER'S FATHER PASSED AWAY AT HOME ON (B)(6) 2025. THE PATIENT WAS NOT ADMITTED TO A HOSPITAL PRIOR TO THE REPORTED INCIDENT. THE CAUSE OF DEATH WAS COVID AND POSSIBLE DIABETIC KETOACIDOSIS. PATIENT BLOOD GLUCOSE WAS UNKNOWN AT THE TIME OF INCIDENT. THE LAST KNOWN PRODUCT(S) FOR THIS PATIENT ARE AS FOLLOWS: MMT-1715KL, UNOMEDICAL, MMT-332A. TROUBLESHOOTING WAS PERFORMED FOR DECEASED REPORTING, PATIENT WORN INSULIN PUMP AT THE TIME OF PASSING. PATIENT WAS USING THE INSULIN PUMP SYSTEM WITHIN 48 HOURS OF REPORTED EVENT AND PATIENT WAS NOT USING SENSORS. THERE IS NO MENTION OF THE AUTO MODE FEATURE AT THE TIME OF THE EVENT. CUSTOMER AGREED TO RETURN THE INSULIN PUMP. MMT-1715KL WAS REQUESTED AND CUSTOMER RESPONSE WAS THE DEVICE WILL BE RETURNED. NO PRODUCT RETURN IS REQUIRED FOR UNOMEDICAL. NO PRODUCT RETURN IS REQUIRED FOR MMT-332A. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. THE DEVICE WILL BE RETURNED FOR ANALYSIS AND FURTHER INFORMATION WILL FOLLOW ONCE THE ANALYSIS HAS BEEN COMPLETED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21629726&pc=DTZ | 8010762-2025-0000119 | 2025-02-16 05:00 | Death | MAQUET CARDIOPULMONARY GMBH | 2025-03-18 04:00 | DTZ | TUBING SET | Failure to Prime; Incomplete or Inadequate Priming; Air/Gas in Device | Cardiac Arrest; No Clinical Signs, Symptoms or Conditions | K112360 |  | 1 | Event Description: COMPLAINT ID# (B)(4).Event Description: THE EVENT OCCURRED IN THE USA DURING PRIMING. FOLLOWING WAS REPORTED IN THE IMPORTER REPORT # (B)(4). CVL (CARDIOVASCULAR LAB) STEMI (ST-SEGMENT ELEVATION MYOCARDIAL INFARCTION) TEAM CALLED IN TO HAVE ECMO (EXTRACORPOREAL MEMBRANE OXYGENATION) CANNULATION AND INITIATION STARTED ON A CCU (CORONARY CARE UNIT) PATIENT. THE PATIENT WAS BROUGHT DOWN TO CATHETER LAB AND ECLS (EXTRACORPOREAL LIFE SUPPORT) TEAM STARTED SETTING UP ECMO DEVICE WHILE THE PROVIDER AND CVL STAFF BEGAN GETTING ACCESSES FOR SHEATH CANNULATION. SHEATH CANNULATION WAS COMPLETED. THE ECMO DEVICE WAS FILLING AND CIRCULATING FOAM AND WAS UNABLE TO BE CONNECTED TO PATIENT. ECLS TEAM NEEDED TO GO SET UP A NEW KIT AND PUMP UPSTAIRS. THE PATIENT WAS CONNECTED TO A NEW ECMO SET 1 HOUR AFTER CATHETER LAB TEAM ACHIEVED CANNULATION OF BILATERAL FEMORAL AND RETROGRADE ACCESS. ADDITIONAL 15 MINUTES FROM THE TIME THE PATIENT WAS CONNECTED TO THE PUMP TO THE PATIENT RECEIVING OXYGENATED BLOOD WERE NEEDED. NO HARM TO ANY PERSON HAS BEEN REPORTED. DUE TO THE DELAY IN TREATMENT OF THE PATIENT, A REPORT IS REQUIRED. COMPLAINT ID# (B)(4).Event Description: COMPLAINT ID# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED ON 2025-03-18/19 THAT THE PATIENT PASSED AWAY 24 HOURS AFTER ON THE SECOND HLS SET CONNECTED DUE TO THE FAMILY DECIDED TO GAVE UP TREATMENT. THE PATIENT WAS A STEMI THAT CODED MULTIPLE TIMES. THE PHYSICIAN INITIATED THE ECMO THERAPY AFTER THE PATIENT WAS REVIVED FROM MULTIPLE CODES AND WAS UNSTABLE. IT WAS STATED THAT THE FIRST USED SET WAS NEVER CONNECTED TO THE PATIENT AND THE CUSTOMER STATED THAT THE MOST PROBABLE ROOT CAUSE WITH THE FIRST HLS SET WAS AN IN-EXPERIENCED NEW ECMO- SPECIALIST WHO PRIMED THE CIRCUIT AND DID NOT DE AIR THE CIRCUIT CORRECTLY CAUSING THE FOAM. THE SECOND HLS SET WHICH WAS THAN USED ON THE PATIENT HAD NO FAILURE. THE PATIENT PASSED AWAY ON (B)(6) 2025. BOTH SETS ARE NOT AVAILABLE FOR INVESTIGATION, AS THEY WERE DISCARDED BY THE CUSTOMER. FURTHER THE CUSTOMER CONFIRMED THAT THE PATIENT PASSED AWAY FROM HIS PRE-EXISTING CONDITIONS AND NOT DUE TO THE EVENT WITH THE FIRST HLS SET, WHICH CAUSED A DELAY IN TREATMENT. A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. Manufacturer Narrative: THE EVENT OCCURRED IN THE USA DURING PRIMING. FOLLOWING WAS REPORTED IN THE IMPORTER REPORT 1600820000-2025-8020. CVL (CARDIOVASCULAR LAB) STEMI (ST-SEGMENT ELEVATION MYOCARDIAL INFARCTION) TEAM CALLED IN TO HAVE ECMO (EXTRACORPOREAL MEMBRANE OXYGENATION) CANNULATION AND INITIATION STARTED ON A CCU (CORONARY CARE UNIT) PATIENT. THE PATIENT WAS BROUGHT DOWN TO CATHETER LAB AND ECLS (EXTRACORPOREAL LIFE SUPPORT) TEAM STARTED SETTING UP ECMO DEVICE WHILE THE PROVIDER AND CVL STAFF BEGAN GETTING ACCESSES FOR SHEATH CANNULATION. SHEATH CANNULATION WAS COMPLETED. THE ECMO DEVICE WAS FILLING AND CIRCULATING FOAM AND WAS UNABLE TO BE CONNECTED TO PATIENT. ECLS TEAM NEEDED TO GO SET UP A NEW KIT AND PUMP UPSTAIRS. THE PATIENT WAS CONNECTED TO A NEW ECMO SET 1 HOUR AFTER CATHETER LAB TEAM ACHIEVED CANNULATION OF BILATERAL FEMORAL AND RETROGRADE ACCESS. ADDITIONAL 15 MINUTES FROM THE TIME THE PATIENT WAS CONNECTED TO THE PUMP TO THE PATIENT RECEIVING OXYGENATED BLOOD WERE NEEDED. ADDITIONAL INFORMATION WAS RECEIVED ON 2025-03-18/19 THAT THE PATIENT PASSED AWAY 24 HOURS AFTER HE WAS ON THE SECOND HLS SET CONNECTED DUE TO THE FAMILY DECIDED TO GIVE UP TREATMENT. THE PATIENT WAS A STEMI THAT CODED MULTIPLE TIMES. THE PHYSICIAN INITIATED THE ECMO THERAPY AFTER THE PATIENT WAS REVIVED FROM MULTIPLE CODES AND WAS UNSTABLE. IT WAS STATED THAT THE FIRST USED SET WAS NEVER CONNECTED TO THE PATIENT AND THE CUSTOMER STATED THAT THE MOST PROBABLE ROOT CAUSE WITH THE FIRST HLS SET WAS AN IN-EXPERIENCED NEW ECMO- SPECIALIST WHO PRIMED THE CIRCUIT AND DID NOT DE AIR THE CIRCUIT CORRECTLY CAUSING THE FOAM. THE SECOND HLS SET WHICH WAS THAN USED ON THE PATIENT HAD NO FAILURE. THE PATIENT PASSED AWAY ON (B)(6) 2025. BOTH SETS ARE NOT AVAILABLE FOR INVESTIGATION, AS THEY WERE DISCARDED BY THE CUSTOMER. FURTHER THE CUSTOMER CONFIRMED THAT THE PATIENT PASSED AWAY FROM HIS PRE-EXISTING CONDITIONS AND NOT FROM THE FAILURE OF THE FIRST HLS SET, WHICH CAUSED A DELAY IN TREATMENT. A MEDICAL REVIEW WAS PERFORMED BY GETINGE MEDICAL AFFAIRS ON 2025-03-31 WITH FOLLOWING CONCLUSION: "ACCORDING TO THE COMPLAINT NARRATIVE, A STEMI (ST-SEGMENT ELEVATION MYOCARDIAL INFARCTION)TEAM WAS CALLED TO INITIATE ECMO (EXTRACORPOREAL MEMBRANE OXYGENATION) CANNULATION FOR A PATIENT IN THE CORONARY CARE UNIT (CCU). THE PATIENT WAS TRANSFERRED TO THE CATH LAB, WHERE THE EXTRACORPOREAL LIFE SUPPORT (ECLS) TEAM SET UP THE ECMO DEVICE WHILE THE PROVIDER AND CVL STAFF PERFORMED VASCULAR ACCESS FOR SHEATH CANNULATION. CANNULATION OF BILATERAL FEMORAL AND RETROGRADE ACCESS WAS COMPLETED SUCCESSFULLY. DURING PRIMING, THE ECMO CIRCUIT EXHIBITED EXCESSIVE FOAM, PREVENTING PATIENT CONNECTION AND DELAYING EXTRACORPOREAL PERFUSION. WITH NO BACKUP SET ON-SITE, THE ECLS TEAM RETRIEVED A REPLACEMENT CIRCUIT AND PUMP FROM ANOTHER LOCATION, DELAYING TREATMENT. THE PATIENT WAS CONNECTED TO THE NEW ECMO SET ONE HOUR AFTER CANNULATION, WITH AN ADDITIONAL 15-MINUTE DELAY BEFORE OXYGENATED BLOOD CIRCULATION BEGAN. THE INTENDED PROCEDURE WAS TO PRIME THE ECMO CIRCUIT WITHOUT VISIBLE BUBBLES FOR IMMEDIATE PATIENT CONNECTION AND SUPPORT. HOWEVER, FOAMING COMPROMISED THE CIRCUIT, DELAYING PATIENT CONNECTION BY OVER AN HOUR. THE LACK OF A BACKUP SET FURTHER EXTENDED THIS DELAY, POTENTIALLY IMPACTING PATIENT STABILITY. A STUDY PUBLISHED IN ACS OMEGA EXAMINED THE STABILITY OF LIQUID FILMS FORMED BY BUBBLES IN BOVINE SERUM ALBUMIN SOLUTIONS. THE RESEARCH HIGHLIGHTED THAT THE AMPHIPATHIC CHARACTERISTICS OF ALBUMIN CONTRIBUTE TO THE FORMATION AND STABILIZATION OF FOAM FILMS, EMPHASIZING THE ROLE OF ALBUMIN AT THE AIR-LIQUID INTERFACE. THIS SUGGESTS THAT THE EXCESSIVE FOAM OBSERVED DURING PRIMING MAY HAVE BEEN DUE TO THE DIRECT INCLUSION OF ALBUMIN INTO THE PRIMING SOLUTION, LEADING TO INCREASED BUBBLE FORMATION. THE PATIENT, A STEMI CASE, HAD SUFFERED MULTIPLE CARDIAC ARRESTS BEFORE ECMO INITIATION. AFTER RESUSCITATION AND STABILIZATION, THE PHYSICIAN INITIATED CARDIOHELP THERAPY DUE TO CONTINUED INSTABILITY. THE PATIENT¿S FAMILY OPTED TO DISCONTINUE SUPPORT AFTER 24 HOURS. THE ATTENDING PHYSICIAN CONFIRMED THAT THE PATIENT¿S DEATH RESULTED FROM PRE-EXISTING CARDIAC CONDITIONS RATHER THAN EQUIPMENT FAILURE. ADDITIONALLY, FEEDBACK FROM THE ECMO MANAGER SUGGESTED THAT THE INITIAL CIRCUIT PRIMING ISSUE MAY HAVE RESULTED FROM IMPROPER DE-AIRING. FURTHERMORE, IT WAS CONFIRMED THAT THE PATIENT EXPIRED DUE TO PRE-EXISTING CONDITIONS RATHER THAN ANY FAILURE OF THE ECMO SYSTEM. IN CONCLUSION, THERE WAS NO EVIDENCE OF DEVICE MALFUNCTION IN THIS INCIDENT. HOWEVER, THE DELAY IN INITIATING EXTRACORPOREAL PERFUSION DUE TO INITIAL PRIMING ISSUES MAY HAVE CONTRIBUTED TO THE OVERALL PATIENT OUTCOME. THE EXCESSIVE FOAM OBSERVED DURING PRIMING, WHICH PREVENTED IMMEDIATE PATIENT CONNECTION, MAY HAVE BEEN INFLUENCED BY THE INCLUSION OF ALBUMIN IN THE PRIMING SOLUTION. RESEARCH HAS SHOWN THAT ALBUMIN'S AMPHIPATHIC PROPERTIES CONTRIBUTE TO FOAM STABILIZATION, POTENTIALLY LEADING TO INCREASED BUBBLE FORMATION. AS BOTH ECMO SETS INVOLVED IN THIS CASE WERE DISCARDED BY THE CUSTOMER, NO FURTHER INVESTIGATION CAN BE CONDUCTED, AND A DEFINITIVE ROOT CAUSE CANNOT BE DETERMINED." ACCORDING TO THE INSTRUCTION OF USE OF THE HLS SET (CHAPTER "SAFETY INSTRUCTIONS FOR THE OXYGENATOR") IT IS STATED THAT THE DE-AIRING MEMBRANE MUST BE OPEN THROUGHOUT PRIMING TO ENSURE SAFE DE-AIRING OF THE OXYGENATOR AND TO ENSURE THAT THE DE-AIRING MEMBRANE OF THE OXYGENATOR IS CLOSED WITH THE YELLOW PROTECTIVE CAP DURING OPERATION. THE FLOW/BUBBLE SENSOR MUST ALWAYS BE AFFIXED ON THE ARTERIAL SIDE OF THE SET IF YOU ARE USING THE CARDIOHELP-I. WITH AN ACTIVATED INTERVENTION, THE DETECTION OF BUBBLES BY THE FLOW/BUBBLE SENSOR TRIGGERS A PUMP STOP. IN CHAPTER "PRIMING THE SYSTEM" IT IS MENTIONED THAT PRE-PRIMING CAN CAUSE AIR BUBBLES TO FORM IN THE SYSTEM. REFERRING TO THE CHAPTER "START PERFUSION" THE CUSTOMER SHOULD ENSURE THAT THE TUBE SYSTEM HAS BEEN COMPLETELY DE-AIRED BEFORE STARTING THE PERFUSION. IN ORDER TO AVOID REOCCURRENCE OF THE REPORTED FAILURE, THE CUSTOMER WILL BE INFORMED BY THE GETINGE SALES AND SERVICE UNIT (SSU) TO FOLLOW THE CHAPTER PRIMING THE SYSTEM AND TO RE-TRAIN THE NEW MEDICAL STUFF. THE PRODUCTION RECORDS OF THE AFFECTED PRODUCT WERE REVIEWED ON 2025-03-31. ACCORDING TO THE FINAL TEST RESULTS, THE PRODUCT PASSED THE TESTS AS PER SPECIFICATIONS. PRODUCTION RELATED INFLUENCES ARE UNLIKELY. THE REVIEW OF SCRAP, REWORK, ENHANCEMENTS AND DESIGN CHANGES WERE REVIEWED AN NO ABNORMALITIES IN REGARDS TO THE REPORTED FAILURE WERE FOUND. IN ADDITION, A REVIEW FOR POTENTIAL FIELD ACTIONS AND CAPAS RELATED TO THE FAILURE AND PRODUCT IN THIS COMPLAINT WAS PERFORMED. THIS COMPLAINT IS NOT IN SCOPE OF ANY ONGOING FIELD ACTIONS AND/OR CAPAS. THE REVIEW OF THE NON-CONFORMITIES HAS BEEN PERFORMED FOR THE REVIEWED TIME PERIOD. IT DOES NOT SHOW ANY NON-CONFORMITY IN REGARDS TO THE REPORTED FAILURE. BASED ON THE RESULTS THE REPORTED FAILURE "AIR IN PRODUCT" COULD BE CONFIRMED, BUT WAS NOT A PRODUCT RELATED MALFUNCTION. THE CUSTOMER WILL BE INFORMED ABOUT THE RESULTS BY THE GETINGE SALES AND SERVICE UNIT. THE OCCURRENCE RATE WAS CALCULATED FOR THE REPORTED ISSUE AND IT WAS DETERMINED THAT THIS IS NOT A SYSTEMIC ISSUE. THEREFORE, NO REMEDIAL ACTION IS REQUIRED. THE OCCURRENCE RATE RELATED TO THE REPORTED ISSUE IS CURRENTLY BEING MONITORED AS PART OF MAQUET CARDIOPULMONARY¿ S TRENDING PROGRAM AND ADDITIONAL INVESTIGATIONS OR CORRECTIONS WILL BE IMPLEMENTED IN CASE OF ADVERSE TRENDING. Manufacturer Narrative: A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. NOTE: NO UDI NUMBER HAS BEEN INCLUDED IN THE SECTION D4 UNIQUE DEVICE IDENTIFIER (UDI) SINCE THE LOT NUMBER OF THE AFFECTED SET WAS NOT PROVIDED. THEREFORE, IT IS NOT POSSIBLE TO IDENTIFY THE SET LOT NUMBER OF THE DEVICE IN QUESTION AND THEREFORE ITS UDI NUMBER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21629850&pc=MIH | 9612164-2025-01369 | 2025-02-21 05:00 | Death | MEDTRONIC IRELAND | 2025-03-18 04:00 | MIH | ENDURANT II BIFURCATED STENT GRAFT | Leak/Splash; Adverse Event Without Identified Device or Use Problem | Rupture | P100021 |  | 1 | Event Description: AN ENDURANT IIS STENT GRAFT SYSTEM WAS IMPLANTED DURING THE ENDOVASCULAR TREATMENT OF A RUPTURED AAA. IT WAS REPORTED THAT THE PATIENT PRESENTED EMERGENTLY 11 YEARS POST-INDEX PROCEDURE WITH A RUPTURED ANEURYSM. THE PHYSICIAN CONVERTED TO OPEN REPAIR. THE STENT GRAFTS WERE EXPLANTED AND AN AORTA BIFEMORAL BYPASS GRAFT WAS IMPLANTED. THE PHYSICIAN COULD NOT IDENTIFY ANY POSSIBLE LEAK SOURCE DUE TO EXCESSIVE BLEEDING BUT IT WAS CONFIRMED AN ENDOLEAK WAS SUSPECTED TO BE THE CAUSE OF THE RUPTURE. IT WAS UNCLEAR THE EXACT TYPE BUT IT WAS REPORTED THAT A TYPE II ENDOLEAK WAS MOST LIKELY. THE PATIENT PASSED AWAY TWO DAYS LATER. Manufacturer Narrative: THE EXACT CAUSE OF THE PATIENT'S DEATH IS NOT AVAILABLE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: SECTION D INFORMATION REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER RELEVANT DEVICE(S) ARE: PRODUCT ID: ETEW2020C82E, SERIAL/LOT #: (B)(6), UBD: 14-OCT-2015, UDI#: (B)(4); PRODUCT ID: ETLW1616C124E, SERIAL/LOT #: (B)(6), UBD: 16-DEC-2015, UDI#: (B)(4); PRODUCT ID: ETLW1616C82E, SERIAL/LOT #: (B)(6), UBD: 28-AUG-2015, UDI#: (B)(4); PRODUCT ID: ETLW1620C124E, SERIAL/LOT #: (B)(6), UBD: 16-SEP-2015, UDI#: (B)(4); PRODUCT ID: ETLW1616C156E, SERIAL/LOT #: (B)(6), UBD: 22-AUG-2024, UDI#: (B)(4). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21629968&pc=PNJ | 9612164-2025-01370 | 2025-02-25 05:00 | Death | MEDTRONIC IRELAND | 2025-03-18 04:00 | PNJ | MICRA¿ AV2 | Adverse Event Without Identified Device or Use Problem | Cardiac Tamponade; Cardiac Perforation | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT INTRAOPERATIVE THE PATIENT EXPERIENCED CARDIAC PERFORATION AND CARDIAC TAMPONADE DURING AN ATTEMPT TO IMPLANT A LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM. CONTRAST WAS INJECTED TO CONFIRM THE IPG POSITION IN THE LOW SEPTUM. TIP PRESSURE WAS APPLIED TO THE DELIVERY SYSTEM AND "GOOSE NECK" WAS OBSERVED. AT THIS CRITICAL MOMENT, THE PATIENT WAS SUSPECTED TO HAVE TAKEN A DEEP BREATH, WHICH CAUSED THE IPG SYSTEM TO ADVANCE UNEXPECTEDLY TOWARDS THE APEX. THE DEPLOYMENT PROCESS WAS ABORTED. CARDIAC TAMPONADE WAS OBSERVED ON ECHOCARDIOGRAPHY. THE IPG SYSTEM WAS ATTEMPTED/NOT USED AND WAS REPLACED. PERICARDIOCENTESIS WAS PERFORMED TO DRAIN EXCESS FLUID. RESUSCITATION WAS ATTEMPTED, AND MEDICATION WAS ADMINISTERED. IT WAS REPORTED THAT THE PATIENT DIED AFTER A FEW HOURS. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE DELIVERY SYSTEM WAS RETURNED WITH THE DEVICE INTACT IN THE DEVICE CUP, ANALYZED. ANALYSIS INDICATED THE LUMEN OF THE DELIVERY SYSTEM WAS TORN. THE ARTICULATION OF THE DELIVERY SYSTEM WAS OUT OF PLANE. THE DELIVERY SYSTEM OUTER SHAFT WAS BENT. THE DELIVERY SYSTEM INNER SHAFT WAS MECHANICALLY KINKED/BENT. D9: YES, RETURN DATE: 28-MAR-2025 H3: YES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE DEVICE WAS RETURNED AND ANALYZED. RETURNED PRODUCT ANALYSIS WAS PERFORMED AND NO ANOMALIES WERE FOUND. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTED: H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SECTION D REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [MICRA] PRODUCT ID [MC2AVR1-DELSYS] (SERIAL: (B)(6). D9: NO. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21630060&pc=DSQ | 2916596-2025-01676 | 2024-10-17 04:00 | Death | THORATEC CORPORATION | 2025-03-18 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Renal Failure; Respiratory Failure; Gastrointestinal Hemorrhage; Pancreatitis | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS ADMITTED ON (B)(6) 2024 FOR RECURRENT TYPE 2 RESPIRATORY FAILURE REQUIRING TRACHEOSTOMY. THE PATIENT ALSO HAD HYPOACTIVE DELIRIUM WITH ONGOING LUNG INFECTION AND BACTEREMIA, POSSIBLY EARLIER EPISODE OF PANCREATITIS WITH FORMATION OF A PSEUDOCYST. SUBSEQUENTLY THE PATIENT DEVELOPED ACUTE KIDNEY DISEASE REQUIRING RENAL DIALYSIS. THE PATIENT WAS NOTED TO HAVE GRADUAL DETERIORATION AND ON (B)(6) 2025 THE PATIENT STARTED HAVING COFFEE GROUND ASPIRATES FROM THE NASOGASTRIC TUBE (NGT). DECISION WAS MADE BY THE MEDICAL TEAM TO HOLD OFF BLOOD TRANSFUSION IN VIEW OF MEDICAL FUTILITY AS WELL AS ANURIC PHASE WITH RISK OF OVERLOAD. IN VIEW OF MEDICAL FUTILITY, THE FAMILY DECIDED ON MEDICAL MANAGEMENT AND COMFORT CARE. THE PATIENT PASSED AWAY ON (B)(6) 2025. SITE NOTED NO FURTHER QUESTIONS WOULD BE ANSWERED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE ACCOUNT INDICATED THAT NO FURTHER INFORMATION WILL BE PROVIDED. HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), WAS NOT RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATION. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) AND THE HEARTMATE 3 PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ""INTRODUCTION"", LISTS RESPIRATORY FAILURE, INFECTION, RENAL FAILURE, AND BLEEDING AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, LISTS INFECTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. THIS SECTION ALSO PROVIDES INFORMATION REGARDING THE RECOMMENDED ANTICOAGULATION THERAPY AND INR (INTERNATIONAL NORMALIZED RATIO) RANGE, AS WELL AS CONSIDERATIONS FOR WHEN THERE IS A RISK OF BLEEDING. SEVERAL SECTIONS OF THE IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21630414&pc=OZD | 1220648-2025-26786 | 2025-02-19 05:00 | Death | ABIOMED, INC. | 2025-03-18 04:00 | OZD | IMPELLA CP | Restricted Flow rate | Hemolysis; Hemorrhage/Bleeding; Tachycardia | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT IN CARDIOGENIC SHOCK FROM AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS). THE PATIENT EXPERIENCED LOW PUMP FLOW, ACCESS SITE BLEEDING AND HEMOLYSIS AND SUBSEQUENTLY EXPIRED. FIVE DAYS PRIOR TO THE IMPELLA IMPLANT, THE PATIENT UNDERWENT THREE-VESSEL CORONARY ARTERY DISEASE COMPLICATED BY CARDIAC ARREST REQUIRING VENOARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION (VA ECMO). THE IMPELLA CP WAS EMERGENTLY PLACED SUCCESSFULLY. THEREAFTER, THE PATIENT WAS TRANSPORTED TO THE UNIT ON ECMO AND IMPELLA SUPPORT. TWO DAYS LATER IT WAS NOTED THE PATIENT UNDERWENT PERCUTANEOUS CORONARY ARTERY INTERVENTION OF DRUG-ELUTING STENT PLACEMENT TO THE LEFT ANTERIOR DESCENDING (LAD) ARTERY. THE NEXT DAY, LOW PUMP FLOW WAS NOTED. THE IMPELLA CP PUMP WAS UNABLE TO GO ABOVE PERFORMANCE LEVEL P-1 WITHOUT SUCTION. THE PHYSICIANS NOTED BLOOD IN CHEST CREATING HEART CONSTRICTION. CHEST TUBE WAS PLACED AT BEDSIDE. ECHOCARDIOGRAM SHOWED LITTLE FILLING IN LEFT VENTRICLE AND RIGHT VENTRICLE. RIGHT LUNG WHITED OUT ON X-RAY. ECMO WAS TURNED DOWN TO 3.0LPM DUE TO CHUGGING THE NIGHT PRIOR. THE PATIENT WOULD RECEIVE FOUR UNITS OF PACKED RED BLOOD CELLS AND TWO UNITS OF FRESH FROZEN PLASMA (UNKNOWN IF FOR THE CHEST BLEED OR LOW PUMP VOLUME). URINE OUTPUT WAS AROUND 30ML/HR. WITH PINK TINGE. FIVE DAYS LATER BLEEDING AT THE ACCESS SITE WAS NOTED REQUIRING A DRESSING CHANGE. THERE WAS UNCERTAINTY IF ECMO CANNULA OR IMPELLA WAS THE SITE OF BLEEDING. THE PHYSICIAN PLACED A SUTURE AT THE IMPELLA ACCESS SITE AND ECMO CANNULAS. THE NEXT DAY PATIENT WAS EXPERIENCED HEMOLYSIS. THE PATIENT BEING WEANED OFF OF VA ECMO BUT WAS UNABLE TO. THE PATIENT WAS ACCEPTED TO AN OUTSIDE FACILITY FOR VENTRICULAR ASSIST DEVICE EVALUATION. HOWEVER, THE PATIENT'S FAMILY MEMBER DECIDED AGAINST THIS TRANSITION, OPTED FOR COMFORT CARE, AND THE PATIENT WOULD EXPIRE LATER THIS SAME DAY.Event Description: ALSO, THE PATIENT EXPERIENCED TORSADES, THE PATIENT WAS SHOCKED 64 TIMES. THE PATIENT CONTINUED TO HAVE TACHYCARDIA DURING SUPPORT. Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING, HEMOLYSIS, LOW PUMP FLOW, AND VF/VT/AF/AT HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. DATA LOG REVIEWED: MANY SUCTIONS DURING SUPPORT OCCURRED. NO POSITIONING ISSUE SEEN. MANY IMPELLA POSITION UNKNOWN OBSERVED. ON (B)(6) 2019 A DIP IN PLACEMENT SIGNAL (PS) AND CONCURRENT DROP IN PUMP FLOW WAS OBSERVED. ECHOCARDIOGRAM SHOWS LITTLE FILLING IN LEFT VENTRICLE AND RIGHT VENTRICLE ON (B)(6) 2019. SOON AFTER THE DIPS, A MOTOR CURRENT SPIKE IS OBSERVED. CLINICAL STATED: ON (B)(6) 2025 BLEEDING FROM EVERY TUBE INCLUDING EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) CANNULAS, IMPELLA, FEEDING TUBE ETC. AND BLOOD GIVEN WITH ADDING SUTURE OF SUPPORT. THE CAUSE OF THE ACCESS SITE BLEEDING - MAJOR MOST LIKELY WAS PATIENT CONDITION RELATED SINCE BLEEDING OBSERVED FROM MULTISITE OF SUPPORT. THE CAUSE OF THE HEMOLYSIS ISSUE CANNOT BE DETERMINED SINCE COMPLAINT DEVICE NOT RETURNED FOR ANALYSIS AND INSUFFICIENT CLINICAL DATA PROVIDED. THE CAUSE OF THE LOW PUMP FLOW WAS MOST LIKELY PATIENT CONDITION RELATED SINCE ECHOCARDIOGRAM SHOWS LITTLE FILLING IN LEFT VENTRICLE AND RIGHT VENTRICLE AT THE TIME OF THE ISSUE. AS THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE VT/VF WAS NOT DETERMINED. B5 ADDITIONAL INFORMATION WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26786. H6 HEALTH EFFECT - CLINICAL CODE 2095 AND HEALTH EFFECT - IMPACT CODE 4647 WERE INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT: USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS: ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY¿ (INCLUDING VENTRICULAR PERFORATION). SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER: ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE: ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS: ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION: ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21630850&pc=OZD | 1220648-2025-26790 | 2024-11-09 05:00 | Death | ABIOMED, INC. | 2025-03-18 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Atrial Fibrillation; Cardiac Arrest | P140003 |  | 1 | Event Description: A NOTIFICATION WAS RECEIVED VIA ABIOMED'S LONG-TERM OUTCOME AND QUALITY INDICATED IMPELLA REGISTRY AND NOTED A PATIENT EXPERIENCED ATRIAL FIBRILLATION WITH RAPID VENTRICULAR RESPONSE. THE RELATIONSHIP TO IMPELLA CP DEVICE IS LISTED AS "POSSIBLY RELATED." THE PATIENT WOULD ULTIMATELY EXPIRE. THE PATIENT WAS ALTERNATED BETWEEN NORMAL SINUS RHYTHM AND ATRIAL FIBRILLATION DURING INDEX STAY PRIOR TO IMPELLA CP DEVICE PROCEDURE. IN THE AFTERNOON, POST IMPELLA CP DEVICE PLACEMENT, THE PATIENT WENT INTO ATRIAL FIBRILLATION WITH VENTRICULAR RATES OF 110-130 BEATS PER MINUTE. THE PATIENT WAS CARDIOVERTED TWICE, AMIODARONE INTRAVENOUS BOLUS WAS GIVEN AND AMIODARONE DRIP WAS INITIATED. THE PATIENT REMAINED IN ATRIAL FIBRILLATION UNTIL DEATH. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21631690&pc=OZD | 1220648-2025-26793 | 2025-02-19 05:00 | Death | ABIOMED, INC. | 2025-03-18 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Atrial Fibrillation; Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT AND EXPERIENCED BLEEDING AND ARRHYTHMIAS. THE PATIENT WOULD SUBSEQUENTLY EXPIRED. THE IMPELLA WAS PLACED, AND PERCUTANEOUS CORONARY INTERVENTION WAS COMPLETED TO THE LEFT MAIN ARTERY. THE PATIENT REMAINED HEMODYNAMICALLY STABLE THROUGHOUT CASE AND THEREAFTER WAS SENT TO THE UNIT ON SUPPORT. OVERNIGHT, THERE WAS SOME ECTOPY NOTED, AND AMIODARONE WAS STARTED. ADDITIONALLY, IT WAS NOTED THERE WAS BLEEDING FROM THE CONTRALATERAL SHEATH SITE. THE SHEATH WAS REMOVED, HEPARIN WAS WITHHELD, AND THE PHYSICIAN MONITORED THE SITUATION. THE FOLLOWING DAY THE PATIENT WAS FREE OF ARRHYTHMIAS THROUGH THE NIGHT AND AMIODARONE DRIP WAS DISCONTINUED. THE PATIENT WAS FOLLOWING SOME COMMANDS AND DISTAL PULSES WERE PRESENT ON DOPPLER. FOLLOWING DAY, THE PATIENT PACED AT 80BPM, AND NURSE REPORTS ASYSTOLE UNDERLING RHYTHM. THE PATIENT REMAINED INTUBATED AND ON VASOACTIVE MEDICATIONS. HEPARIN, RESTARTED, WAS WITHHELD THE PRIOR NIGHT DUE TO BLOOD-TINGED SPUTUM. DISTAL PULSES PRESENT ON DOPPLER. NOW DAY 7 ON SUPPORT THE PATIENT WAS IN ATRIAL FIBRILLATION WITH RAPID VENTRICULAR RESPONSE. THE FOLLOWING DAY THE PATIENT EXPIRED; CARE WAS WITHDRAWN. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: IMPELLA CP WITH SMARTASSIST SYSTEM. SECTION: GENERAL PATIENT CARE CONSIDERATIONS: ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT, USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS: ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY¿ (INCLUDING VENTRICULAR PERFORATION).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21632138&pc=DQY | 2029214-2025-00691 | 2023-07-18 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-03-18 04:00 | DQY | REFLEX CATHETER | Adverse Event Without Identified Device or Use Problem | Bronchopneumonia |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING: THE STUDY TITLED "A SINGLE-CENTER RETROSPECTIVE STUDY OF THE COCO TECHNIQUE IN THE TREATMENT OF CHRONIC INTERNAL CAROTID ARTERY OCCLUSION." THE OBJECTIVE IS TO INTRODUCE A NOVEL ENDOVASCULAR RECANALIZATION METHOD (THE COCO TECHNIQUE) AND INVESTIGATE ITS SUCCESS RATE, PERIPROCEDURAL COMPLICATIONS, AND EARLY OUTCOMES IN PATIENTS WITH CHRONIC INTERNAL CAROTID ARTERY OCCLUSION (CICAO). THE TIME FRAME OF THIS STUDY WAS: THE STUDY REVIEWED DATA FROM PATIENTS WHO UNDERWENT ENDOVASCULAR RECANALIZATION BETWEEN JANUARY 2019 AND DECEMBER 2022. MULTIPLE MANUFACTURER¿S DEVICES WERE USED IN THE STUDY POPULATION: 6F/90 CM SHUTTLE CATHETER (OPTMED), 5F/125 CM VERTEBRAL ANGIOGRAPHIC CATHETER (CORDIS), 6F/115 CM INTERMEDIATE CATHETER (MEDTRONIC), MICROCATHETER (SL-10, STRYKER), 0.035 INCH STIFF SLIPPERY GUIDEWIRE (TERUMO), 2.5 MM BALLOON (SACSPEED, ACHEVA), ENTERPRISE 2 STENT (CODMAN), CEREBRAL PROTECTION DEVICE (SPIDERFX, MEDTRONIC) THE FOLLOWING MEDTRONIC DEVICES WERE USED: 6F/115 CM INTERMEDIATE CATHETER, CEREBRAL PROTECTION DEVICE (SPIDERFX) DEATHS OCCURRED IN THE STUDY POPULATION: DURING THE FOLLOW-UP, ONE PATIENT DIED OF SEVERE PNEUMONIA. AMONG PATIENT ADVERSE EVENTS INCLUDED: INTRAOPERATIVE COMPLICATIONS: SLIGHT SUBARACHNOID HEMORRHAGE (2 PATIENTS), ASYMPTOMATIC DISSECTION HEMORRHAGIC COMPLICATIONS (3 PATIENTS) AND POSTOPERATIVE ISCHEMIC COMPLICATIONS (3 PATIENTS), ALL IMPROVED AFTER CONSERVATIVE TREATMENT AND REHABILITATION. MALFUNCTIONS, COMPLICATIONS, SIDE EFFECTS: INTRAOPERATIVE COMPLICATIONS: DISSECTION, DISTAL EMBOLIZATION, CAROTID-CAVERNOUS ARTERIOVENOUS FISTULA, AND VESSEL PERFORATION. PERIOPERATIVE NEUROLOGICAL COMPLICATIONS: HEMORRHAGIC OR ISCHEMIC. RESTENOSIS WAS OBSERVED IN SIX (15.8%) PATIENTS. POST-OPERATIVE ISSUES: DEVELOPMENT OF ISCHEMIC EVENTS IN TWO PATIENTS DURING THE FOLLOW-UP PERIOD. CLINICAL OUTCOMES OF INTEREST: THE TECHNICAL SUCCESS RATE OF THE COCO TECHNIQUE WAS 89.8% (44/49). THE MODIFIED RANKIN SCALE SCORES WERE LOWER AT THE 3-MONTH FOLLOW-UP THAN AT BASELINE (1 (0¿2) VS 2 (1¿2), P=0.04). ONLY TWO PATIENTS EXPERIENCED RECURRENT ISCHEMIC EVENTS, INDICATING A LOWER RATE OF RECURRENT ISCHEMIA ASSOCIATED WITH SUCCESSFUL RECANALIZATION. NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID UNK-NV-RFX (LOT: UNKNOWN); PRODUCT TYPE: ; IMPLANT DATE N/A; EXPLANT DATE N/A PRODUCT ID UNK-NV-RFX (UNKNOWN); PRODUCT TYPE: ; IMPLANT DATE N/A; EXPLANT DATE N/A G2: CITATION: AUTHORS: CHEN ,L., XU, Y., LI, L., JI, T., WANG, Y., ZHU, W., WANG, F., ZHANG, Q.. A SINGLE-CENTER RETROSPECTIVE STUDY OF THE COCO TECHNIQUE IN THE TREATMENT OF CHRONIC INTERNAL CAROTID ARTERY OCCLUSION. J NEUROINTERVENT SURG 16:756¿763 2024. DOI: 10.1136/JNIS-2023-020451 EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21632203&pc=NPT | 2135147-2025-01371 | 2025-02-21 05:00 | Death | ABBOTT MEDICAL | 2025-03-18 04:00 | NPT | NAVITOR TRANSCATHETER AORTIC VALVE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | P190023 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 29MM NAVITOR VALVE WAS SELECTED FOR IMPLANT USING A LARGE FLEXNAV DELIVERY SYSTEM (DS). THE PATIENT HAD HIGH COMPLEXITY IN ACCESS, SO SUBCLAVIAN ACCESS WAS USED. DURING THE PROCEDURE, UPON INITIATING THE PRE-DILATION, THE PATIENT WENT INTO CARDIORESPIRATORY ARREST. RESUSCITATION AND INITIATION OF EXTERNAL PACEMAKER ACTIVITY WAS USED AND THE PATIENT WAS ABLE TO RETURN TO SINUS RHYTHM. UPON STARTING THE IMPLANT, THE VALVE APPEARED TO BE LOW, REQUIRING RECAPTURE AND REPOSITIONING. WHEN THIS WAS DONE, THE PATIENT WENT INTO CARDIORESPIRATORY ARREST AGAIN. THE PHYSICIAN IMPLANTED THE NAVITOR AT A DEPTH OF 3MM IN AN ATTEMPT TO RETURN THE PATIENT TO SINUS RHYTHM BUT WAS WITHOUT SUCCESS. THE PATIENT ENDED UP PASSING AWAY DUE TO CARDIORESPIRATORY ARREST.Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 29MM NAVITOR VALVE WAS SELECTED FOR IMPLANT USING A LARGE FLEXNAV DELIVERY SYSTEM (DS). THE PATIENT HAD HIGH COMPLEXITY IN ACCESS, SO SUBCLAVIAN ACCESS WAS USED. DURING THE PROCEDURE, UPON INITIATING THE PRE-DILATION, THE PATIENT WENT INTO CARDIORESPIRATORY ARREST. RESUSCITATION AND INITIATION OF EXTERNAL PACEMAKER ACTIVITY WAS USED AND THE PATIENT WAS ABLE TO RETURN TO SINUS RHYTHM. UPON STARTING THE IMPLANT, THE VALVE APPEARED TO BE LOW, REQUIRING RECAPTURE AND REPOSITIONING. WHEN THIS WAS DONE, THE PATIENT WENT INTO CARDIORESPIRATORY ARREST AGAIN. THE PHYSICIAN IMPLANTED THE NAVITOR AT A DEPTH OF 3MM IN AN ATTEMPT TO RETURN THE PATIENT TO SINUS RHYTHM, BUT WAS WITHOUT SUCCESS. THE PATIENT ENDED UP PASSING AWAY DUE TO CARDIORESPIRATORY ARREST. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. AN EVENT OF CARDIAC ARREST AND PATIENT DEATH WAS REPORTED. INFORMATION FROM THE FIELD INDICATED THAT RESUSCITATION AND INITIATION OF EXTERNAL PACEMAKER ACTIVITY WAS USED AND THE PATIENT WAS ABLE TO RETURN TO SINUS RHYTHM. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED, AND THE PRODUCT MET ALL SPECIFICATIONS. BASED ON ALL AVAILABLE INFORMATION, CAUSE FOR THE REPORTED CARDIAC ARREST AND PATIENT DEATH COULD NOT BE CONCLUSIVELY DETERMINED. IT IS POSSIBLE THAT THEY ARE RELATED TO THE IMPLANT PROCEDURE. HOWEVER, THIS CANNOT BE CONFIRMED. THE REPORTED DEATH APPEARS TO BE RELATED TO THE CARDIORESPIRATORY ARREST. THE REPORTED UNEXPECTED MEDICAL INTERVENTION WAS THE RESULT OF CASE-SPECIFIC CIRCUMSTANCES, AS TEMPORARY PACEMAKER IMPLANT WAS PERFORMED. THE EVENT AND IMAGING WAS REVIEWED BY AN ABBOTT SENIOR DIRECTOR OF MEDICAL AFFAIRS. THE REVIEWER STATED, "THE CAUSE OF DEATH WOULD APPEAR TO BE PROCEDURE RELATED. THERE DOES NOT APPEAR TO BE ANY DEVICE MALFUNCTION. NO FURTHER ANALYSIS IS POSSIBLE GIVEN THE LIMITED INFORMATION PROVIDED. IF ADDITIONAL INFORMATION OR IMAGING IS OBTAINED, AN ADDENDUM MEDICAL REVIEW WILL BE COMPLETED. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21632352&pc=NKM | 2135147-2025-01373 | 2016-03-01 05:00 | Death | ABBOTT MEDICAL | 2025-03-18 04:00 | NKM | MITRACLIP® SYSTEM CLIP DELIVERY SYSTEM | Off-Label Use | No Clinical Signs, Symptoms or Conditions | P100009 |  | 1 | Event Description: THE ARTICLE "MALNUTRITION AND OUTCOMES IN PATIENTS WITH TRICUSPID REGURGITATION UNDERGOING TRANSCATHETER TRICUSPID VALVE REPAIR" WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE MULTI CENTER STUDY, TO EVALUATE THE IMPACT OF MALNUTRITION IN PATIENTS WITH TRICUSPID REGURGITATION (TR) UNDERGOING TRICUSPID TRANSCATHETER EDGE-TO-EDGE REPAIR (T-TEER) IS NOT WELL ESTABLISHED. WE EVALUATED THE IMPACT OF MALNUTRITION AMONG PATIENTS WITH SYMPTOMATIC TR UNDERGOING T-TEER. DEVICES MENTIONED INCLUDE TRICLIP, MITRACLIP, AND A NON-ABBOTT DEVICE. THE ARTICLE CONCLUDED THAT IN THE LARGE, REAL-WORLD, MULTICENTRE EUROTR REGISTRY, MALNUTRITION WAS PRESENT IN ONE OUT OF FIVE PATIENTS WITH SYMPTOMATIC TR UNDERGOING T-TEER AND WAS INDEPENDENTLY ASSOCIATED WITH INCREASED MORTALITY. THE PROGNOSTIC BENEFIT OF SUCCESSFUL T-TEER IN REDUCING MORTALITY WAS CONSISTENTLY OBSERVED IN PATIENTS WITH AND WITHOUT MALNUTRITION. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS MARIANNA ADAMO AT UNIVERSITY OF BRESCIA INSTITUTION WITH CORRESPONDING EMAIL MARIANNAADAMO@HOTMAIL.COM]. THE TIME FRAME OF THE STUDY WAS MARCH 2016 TO FEBRUARY 2024. A TOTAL OF 1034 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH 645 RECEIVED AN ABBOTT DEVICE. AS THIS EVENT IS FROM A LITERATURE REVIEW, THERE IS NO RELEVANT PATIENT INFORMATION (DATE OF BIRTH, AGE, GENDER, WEIGHT, AND MEDICAL HISTORY) TO REPORT. COMORBIDITIES INCLUDED HYPERTENSION, DIABETES, ATRIAL FIBRILLATION, PRIOR CARDIAC SURGERY. (B)(6), UNK MITRACLIP: PERI- AND POST-PROCEDURAL COMPLICATIONS INCLUDED HEART FAILURE, PROLONGED HOSPITALIZATION, DEATH. (B)(6), UNK TRICLIP: PERI- AND POST-PROCEDURAL COMPLICATIONS INCLUDED HEART FAILURE, PROLONGED HOSPITALIZATION, DEATH. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF MITRACLIP DEVICE WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING HYPERTENSION, DIABETES, ATRIAL FIBRILLATION, PRIOR CARDIAC SURGERY. COMPLICATIONS REPORTED INCLUDED HEART FAILURE, PROLONGED HOSPITALIZATION, DEATH; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. THE REPORTED OFF-LABEL USE OF THE DEVICE WAS ASSOCIATED WITH THE MITRACLIP DEVICE BEING USED ON THE TRICUSPID VALVE TO TREAT TRICUSPID REGURGITATION (TR). A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. Manufacturer Narrative: IT IS UNKNOWN IF THE DEVICE IS RETURNING FOR ANALYSIS. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21632358&pc=NGV | 2135147-2025-01374 | 2025-02-21 05:00 | Death | ABBOTT MEDICAL | 2025-03-18 04:00 | NGV | AMPLATZER AMULET | Migration or Expulsion of Device; Patient-Device Incompatibility | Hemorrhage/Bleeding; Cardiac Tamponade; Pericardial Effusion | P200049 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 28MM AMPLATZER AMULET LEFT ATRIAL APPENDAGE OCCLUDER WAS CHOSEN FOR IMPLANT USING A 14F AMULET DELIVERY SHEATH. THE SIZING OF THE DETERMINED VIA TRANSESOPHAGEAL ECHOCARDIOGRAPHY (TEE) IN FOUR DIFFERENT VIEWS (0,45, 90, 135). DURING PROCEDURE, THE LEFT ATRIAL PRESSURE AT TIME OF PROCEDURE WAS 13MMHG AND FLUID WAS GIVEN. THE ACTIVATED CLOTTING LEVELS (ACT) WERE ABOVE 250S AND 5000 UNITS. THE AMULET WAS IMPLANTED AFTER ALL THE CLOSE CRITERIA SATISFIED AND TENSION TEST WAS NOT PERFORMED. THERE WAS NO MALFUNCTION NOTED ON THE DELIVERY SHEATH. AFTER THE APPENDAGE WAS CLOSED, THE DOCTORS NOTICED BLOOD LEAKAGE INTO THE PERICARDIUM, INDICATING A CIRCUMFERENTIAL PERICARDIAL EFFUSION AND A CARDIAC SURGERY WAS CALLED. THE PHYSICIAN MENTIONED THE CAUSE OF PERICARDIAL EFFUSION WAS AMULET DEVICE. THE PATIENT WAS THEN TAKEN TO CARDIAC SURGERY WITH TAMPONADE. WHEN THE PHYSICIAN TRIED TO PERFORM A PERICARDIOCENTESIS, THE AMULET DEVICE WAS EMBOLIZED INTO MITRAL VALVE. THEY DECIDED TO PERFORM A SURGICAL RETRIEVAL BECAUSE THE PATIENT WAS BLEEDING A LOT AND DIDN¿T RESIST TO THE PROCEDURE. THEY DON'T HAVE TIME TO RUN FOR SURGICAL INTERVENTION. THE EMBOLIZATION OCCURRED DUE SURGICAL MANIPULATION. THE DEVICE WAS SURGICALLY REMOVED. AFTER THE PROCEDURE, THE PATIENT STARTED TO BLEED AND LATER DIED. THE PHYSICIAN MENTIONED THE CAUSE OF DEATH WAS PROCEDURE.Event Description: N/A. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. Manufacturer Narrative: AN EVENT OF MIGRATION OR EXPULSION OF DEVICE (DEVICE EMBOLIZATION) AND PATIENT-DEVICE INCOMPATIBILITY (IMPLANT-RELATED TISSUE DAMAGE) AND PERICARDIAL EFFUSION WAS REPORTED. A RETURNED DEVICE ASSESSMENT COULD NOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED FOR ANALYSIS. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED, AND THE PRODUCT MET ALL SPECIFICATIONS. THE PROVIDED PROCEDURAL IMAGING AND INFORMATION WERE REVIEWED BY MEDICAL AFFAIRS. BASED ON THE INFORMATION RECEIVED, THE CAUSE FOR THE REPORTED IMPLANT-RELATED TISSUE DAMAGE COULD NOT BE DETERMINED. IT IS POSSIBLE DUE TO PROCEDURAL CONDITIONS (E.G. CLOSE CRITERIA HAD NOT BEEN ACHIEVED PRIOR TO RELEASE, "NOT COMPRESSED ENOUGH" AT TIME OF RELEASE);. IT WAS INDICATED THAT DEVICE IMPLANTED AFTER ALL THE CLOSE CRITERIA SATISFIED AND TENSION TEST WAS NOT PERFORMED. THE REPORTED DEVICE EMBOLIZATION APPEARS TO BE RELATED TO PROCEDURAL CONDITIONS (EMBOLIZATION OCCURRED DUE SURGICAL MANIPULATION). IT WAS ALSO INDICATED THAT CAUSE OF PATIENT DEATH WAS PROCEDURE. THE REPORTED UNEXPECTED MEDICAL INTERVENTIONS AND SURGICAL INTERVENTION WERE A RESULT OF CASE-SPECIFIC CIRCUMSTANCES AS THE DEVICE WAS SNARED AND PERICARDIOCENTESIS WAS PREFORMED. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH REGARDS TO MANUFACTURE, DESIGN, OR LABELING. CODES REMOVED: HEALTH EFFECT-CLINICAL CODE: COMPONENT CODE: 4755 PART/COMPONENT/SUB-ASSEMBLY TERM NOT APPLICABLE. TYPE OF INVESTIGATION: 4118 TYPES OF INVESTIGATION NOT YET DETERMINED. INVESTIGATION FINDINGS: INVESTIGATION CONCLUSIONS: 11 CONCLUSIONS: NOT YET AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21632575&pc=MNT | 2518422-2025-103124 | 2025-03-10 04:00 | Death | RESPIRONICS, INC. | 2025-03-18 04:00 | MNT | BIPAP A40 SYSTEM SILVER, JP | Electrical /Electronic Property Problem | Respiratory Arrest; Insufficient Information | K121623 |  | 1 | Event Description: ON 13 MARCH 2025, THE MANUFACTURER RECEIVED INFORMATION INDICATING A PATIENT USING A BIPAP A40 VERSION 3 DEVICE HAD DIED. THE HEALTH FACILITY CONTACTED PHILIPS' CALL CENTER REQUESTING TO RETURN OF THE SUBJECT DEVICE. THE PHILIPS SALES REPRESENTATIVE LEARNED OF THE PATIENT'S DEATH WHEN THEY WENT TO THE FACILITY TO RETRIEVE THE DEVICE AS REQUESTED. IT WAS REPORTED A VENTILATOR INOPERATIVE ALARM HAD OCCURRED ON THE DEVICE AND THE VENTILATOR HAD STOPPED OPERATING. THE MEDICAL STAFF RESPONDED TO THE PATIENT EMERGENCY WITH MANUAL RESUSCITATION OF THE PATIENT USING AN AMBU BAG. HOWEVER, THERE WAS NO CHANGE IN THE PATIENT'S CONDITION. THE PATIENT WAS TRANSFERRED TO A HOSPITAL BY AMBULANCE WHERE THEY DIED THAT NIGHT. THE CAUSE OF DEATH REMAINS UNKNOWN TO THE MANUFACTURER AT THIS TIME. THIS ISSUE HAS BEEN IDENTIFIED UNDER THE FSN 2023-CC-SRC-039. Manufacturer Narrative: THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER AND EVALUATION COMPLETED WITH THE FOLLOWING RESULTS: A VENTILATOR INOPERATIVE ERROR CODE WAS CONFIRMED TO BE PRESENT IN THE DOWNLOAD ERROR LOG INDICATING THE DEVICE WAS UNABLE TO WRITE TO THE VENT LOG. EVALUATION CONFIRMED THE ALLEGED VENTILATOR INOPERATIVE CONDITION. THE MAIN PRINTED CIRCUIT ASSEMBLY WAS REPLACED TO ADDRESS THIS ISSUE. THE DEVICE PASSED FINAL TESTING AND WAS CONFIRMED TO BE OPERATING PROPERLY. Manufacturer Narrative: THE BIPAP A40 PRO (UPDATE MATERIAL # FROM RA) IS SUBSTANTIALLY SIMILAR TO THE BIPAP A40 (1111169) AND WILL BE REPORTED IN THE UNITED STATES UNDER BIPAP A40, 501K NUMBER: K121623.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21632883&pc=NGV | 2124215-2025-16482 | 2025-02-05 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-03-18 04:00 | NGV | WATCHMAN FLX PRO | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P130013 |  | 1 | Event Description: IT WAS REPORTED THAT DEATH OCCURRED. A CONCOMITANT ABLATION AND LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS PERFORMED. THE ABLATION WAS PERFORMED WITH A NON-BOSTON SCIENTIFIC (BSC) DEVICE AND WAS COMPLETED IN 2.5 HOURS, THEN THE LAA CLOSURE PROCEDURE WAS INITIATED. A 24MM WATCHMAN FLX PRO CLOSURE DEVICE WAS IMPLANTED WITH COMPLETE DEVICE SEAL NOTED. THE PATIENT WAS STABLE POST PROCEDURE WITH NO PERICARDIAL EFFUSION NOTED, AND NO ISSUES OCCURRING DURING THE PROCEDURE. AT AN UNKNOWN DATE APPROXIMATELY ONE (1) WEEK POST INDEX PROCEDURE THE PATIENT DIED. NO FURTHER DETAILS WERE PROVIDED. Manufacturer Narrative: B2 AND B3: EVENT DATE AND DEATH DATE ESTIMATED TO BE ONE (1) WEEK POST INDEX PROCEDURE, AS THE DATE OF DEATH AND DATE OF EVENT IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21633161&pc=NPT | 2025587-2025-01963 | 2024-11-22 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-18 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Insufficient Information | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 1 YEAR AND 2 MONTHS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT WAS DISCHARGED FROM THE HOSPITAL DUE TO A CANCER OF UNKNOWN ORIGIN. FOLLOWING DISCHARGE, THE PATIENT DIED. THE CAUSE OF DEATH WAS UNKNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21633491&pc=NPT | 2015691-2025-01765 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-18 04:00 | NPT | EDWARDS SAPIEN 3 TRANSCATHETER HEART VALVE | Insufficient Information | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P140031 | E2016006 | 7 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 VALVE. THIS REPORT SUMMARIZES 7 AORTIC CARDIAC SURGERY OR INTERVENTION - OTHER UNPLANNED DEATH EVENT(S) FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 68 - 86 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 3 FEMALE(S) AND 4 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 7 CARDIAC SURGERY OR INTERVENTION - OTHER UNPLANNED DEATH EVENT(S) FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 1.57 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 TRANSCATHETER HEART VALVE ARE: (B)(4). THE INSTRUCTIONS FOR USE (IFU) LIST REOPERATION, INCLUDING EMERGENCY CARDIAC SURGERY, AS ADDITIONAL POTENTIAL RISKS ASSOCIATED WITH THE USE OF THE THV, DELIVERY SYSTEM, AND/OR ACCESSORIES. THIS EVENT TYPE (CARDIAC SURGERY OR INTERVENTION - OTHER UNPLANNED) IS DEFINED AS: THE PATIENT SUBSEQUENTLY UNDERWENT OTHER TYPES OF CARDIAC SURGERY OR CATH LAB INTERVENTIONS (NOT OTHERWISE SPECIFIED) THAT WAS UNPLANNED. THIS COULD INCLUDE REOPERATION FOR DAMAGE TO ADJACENT STRUCTURES, SUCH AS THE MITRAL APPARATUS, OR POTENTIALLY VALVE EXPLANT RELATED TO EMBOLIZATION, OR OTHER EVENT TYPES. CARDIOVASCULAR INJURY, INCLUDING PERFORATION OR DISSECTION OF VALVULAR STRUCTURES (E.G., INJURY TO THE MITRAL VALVE/ APPARATUS), AND VALVE EMBOLIZATION REQUIRING INTERVENTION ARE KNOWN POTENTIAL COMPLICATIONS LISTED IN THE PRODUCT INSTRUCTIONS FOR USE (IFU). CHORDAE TENDINEA RUPTURE IN TRANSCATHETER VALVE REPLACEMENT CAN OCCUR DURING ADVANCEMENT OF THE GUIDEWIRE, BAV CATHETER, OR DELIVERY SYSTEM AND IS MOST LIKELY TO OCCUR WITH ANTEGRADE APPROACHES, HOWEVER, CAN OCCUR WITH A RETROGRADE APPROACH. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE TRANSCATHETER HEART VALVES (THV). TRAINING INCLUDES PROPER GUIDEWIRE POSITIONING, AND CAREFUL MANIPULATION OF DEVICES. PER THE THV TA TRAINING MANUALS, AFTER GAINING LV ACCESS WITH THE 0.035" SOFT GUIDEWIRE, THE WIRE SHOULD BE JIGGLED UNDER TEE IMAGING OF THE MITRAL VALVE. AN INCREASE IN MR SUGGESTS WIRE ENTANGLEMENT IN THE MITRAL SUB-VALVULAR APPARATUS. IF ENTANGLEMENT IS SUSPECTED, THE GUIDEWIRE SHOULD BE COMPLETELY REMOVED FROM THE VENTRICLE; THE OPERATOR SHOULD CHANGE DIRECTION OF THE NEEDLE AND REINSERT THE GUIDEWIRE INTO THE VENTRICLE, CHECKING AGAIN FOR WIRE ENTANGLEMENT. THE TF TRAINING MANUAL ALSO PROVIDES GUIDANCE FOR VALVE CROSSING AND WIRE EXCHANGE: EXCHANGE 0.035" AMPLATZ EXTRA-STIFF WIRE WITH PRE-SHAPED DISTAL END IN LEFT VENTRICLE. ENSURE GUIDE CATHETER IS ADVANCED UPON WIRE EXCHANGE TO ENSURE EXCHANGE WIRE DOES NOT GET CAUGHT IN THE MITRAL CHORDAE. THERE ARE MULTIPLE PATIENT AND PROCEDURAL FACTORS THAT ALONE OR IN COMBINATION CAN CAUSE OR CONTRIBUTE TO VENTRICULAR EMBOLIZATION, INCLUDING IMPROPER POSITIONING PRIOR TO DEPLOYMENT, POOR IMAGE INTENSIFIER ANGLE, POOR COAXIAL ALIGNMENT OF THE VALVE/DELIVERY SYSTEM, A NARROW, CALCIFIED SINOTUBULAR JUNCTION, MINIMALLY OR BULKY/SEVERELY CALCIFIED AORTIC LEAFLETS, RAPID DEPLOYMENT, RELEASE OF STORED TENSION DURING DEPLOYMENT, AND MOVEMENT OF THE DELIVERY SYSTEM BY THE OPERATOR. THE THV TRAINING MANUALS INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN THV. TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS AND PROCTORED PROCEDURES. THE CORRECT ALIGNMENT AND POSITIONING OF THE DEVICE AT THE POINT OF DEPLOYMENT IS EMPHASIZED AS A KEY FACTOR TO THE PLACEMENT AND FIXATION OF THE DEVICE. OPERATORS ARE ALSO INSTRUCTED TO USE FLUOROSCOPY AS THE PRIMARY METHOD OF VISUALIZATION FOR POSITIONING AND DEPLOYMENT. IN PATIENTS WITH HIGH-RISK ANATOMICAL FEATURES FOR VENTRICULAR EMBOLIZATION (I.E., SMALL, CALCIFIED STJ, MINIMAL LEAFLET CALCIFICATION), BAV MAY PROVIDE INDICATION OF POTENTIAL BALLOON MOVEMENT DURING VALVE DEPLOYMENT. I IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21633493&pc=NPT | 2015691-2025-01771 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-18 04:00 | NPT | EDWARDS SAPIEN 3 TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P140031 | E2016006 | 1 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 VALVE. THIS REPORT SUMMARIZES 1 AORTIC DEVICE MIGRATION DEATH EVENT(S) FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 81 - 81 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 0 FEMALE(S) AND 1 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER: E2016006 AND SUMMARIZES 1 DEVICE MIGRATION DEATH EVENT(S) FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 0 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 TRANSCATHETER HEART VALVE ARE: (B)(4). PER THE INSTRUCTIONS FOR USE, VALVE MIGRATION REQUIRING INTERVENTION IS A POTENTIAL ADVERSE EVENT ASSOCIATED WITH TRANSCATHETER AORTIC VALVE REPLACEMENT. ACCORDING TO LITERATURE REVIEW, VALVE MIGRATION RESULTS WHEN FORCES ACTING ON THE TRANSCATHETER HEART VALVE (THV) OVERCOME THE STRENGTH OF ATTACHMENT OF THE VALVE TO THE ANNULUS. STENT VALVES ARE SUBJECTED TO ANTEGRADE EJECTION FORCES DURING SYSTOLE. LESS-THAN-SEVERE AND NON-UNIFORMLY DISTRIBUTED CALCIFICATION OF THE NATIVE LEAFLETS, INCORRECT BIOPROSTHETIC VALVE SIZING, AND INCOMPLETE FRAME EXPANSION CAN CONTRIBUTE TO VALVE MIGRATION. ADDITIONALLY, RESIDUAL OVERHANGING LEAFLETS CAN EXERT DOWNWARDS FORCE DURING DIASTOLE, CAUSING MIGRATION OF THE THV TOWARDS THE LEFT VENTRICLE. THE EDWARDS THV PATIENT SCREENING MANUAL ADVISES THE OPERATOR ON PRE-PROCEDURE ASSESSMENT OF THE AORTIC VALVE AND ROOT, TAKING INTO CONSIDERATION THE DEGREE AND DISTRIBUTION OF NATIVE LEAFLET CALCIFICATION. THE PROCEDURAL DIDACTIC INSTRUCTS THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE TRANSCATHETER HEART VALVE (THV). TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS AND PROCTORED PROCEDURES. THE CORRECT SIZING, ALIGNMENT AND POSITIONING OF THE DEVICE ARE EMPHASIZED AS KEY FACTORS TO THE PLACEMENT AND FIXATION OF THE DEVICE. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21633500&pc=NPT | 2015691-2025-01768 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-18 04:00 | NPT | EDWARDS SAPIEN 3 TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P140031 | E2016006 | 2 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 VALVE. THIS REPORT SUMMARIZES 2 AORTIC DEVICE EMBOLIZATION DEATH EVENT(S) FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 73 - 81 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 0 FEMALE(S) AND 2 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 2 DEVICE EMBOLIZATION DEATH EVENT(S) FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 111 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 TRANSCATHETER HEART VALVE ARE: (B)(4). PER THE INSTRUCTIONS FOR USE, VALVE EMBOLIZATION IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT PROCEDURE. THERE ARE MULTIPLE PATIENT AND PROCEDURAL FACTORS THAT ALONE OR IN COMBINATION CAN CAUSE OR CONTRIBUTE TO EMBOLIZATION, INCLUDING IMPROPER POSITIONING PRIOR TO DEPLOYMENT, POOR IMAGE INTENSIFIER ANGLE, POOR COAXIAL ALIGNMENT OF THE VALVE AND DELIVERY SYSTEM, MINIMALLY OR BULKY/SEVERELY CALCIFIED AORTIC LEAFLETS, PRESERVED EJECTION FRACTION, LOSS OF PACING CAPTURE, RAPID DEPLOYMENT, RELEASE OF STORED TENSION DURING DEPLOYMENT, INACCURATE MEASUREMENT OF THE LANDING ZONE ,AND MOVEMENT OF THE DELIVERY SYSTEM BY THE OPERATOR. THE THV TRAINING MANUALS INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE THV. TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS AND PROCTORED PROCEDURES. THE CORRECT ALIGNMENT AND POSITIONING OF THE DEVICE AT THE POINT OF DEPLOYMENT IS EMPHASIZED AS A KEY FACTOR TO THE PLACEMENT AND FIXATION OF THE DEVICE. OPERATORS ARE ALSO INSTRUCTED TO USE FLUOROSCOPY AS THE PRIMARY METHOD OF VISUALIZATION FOR POSITIONING AND DEPLOYMENT. IN PATIENTS WITH HIGH-RISK ANATOMICAL FEATURES FOR EMBOLIZATION (I.E. MINIMAL LEAFLET CALCIFICATION), BAV MAY PROVIDE INDICATION OF POTENTIAL BALLOON MOVEMENT DURING VALVE DEPLOYMENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21633504&pc=NPT | 2015691-2025-01769 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-18 04:00 | NPT | EDWARDS SAPIEN 3 TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P140031 | E2016006 | 2 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 VALVE. THIS REPORT SUMMARIZES 2 AORTIC PERCUTANEOUS CORONARY INTERVENTION DEATH EVENT(S) FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 84 - 88 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 2 FEMALE(S) AND 0 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER: E2016006 AND SUMMARIZES 2 PERCUTANEOUS CORONARY INTERVENTION DEATH EVENT(S) FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 0 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 TRANSCATHETER HEART VALVE ARE: (B)(4). PER THE INSTRUCTIONS FOR USE (IFU), CORONARY FLOW OBSTRUCTION IS A POTENTIAL ADVERSE EVENT ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT PROCEDURE. THE IFU CAUTIONS THAT THE SAFETY AND EFFECTIVENESS HAVE NOT BEEN ESTABLISHED IN PATIENTS WITH BULKY CALCIFIED AORTIC VALVE LEAFLETS IN CLOSE PROXIMITY TO THE CORONARY OSTIA. IN ADDITION, IT WARNS THAT CAUTION SHOULD BE EXERCISED IN IMPLANTING A BIOPROSTHESIS IN PATIENTS WITH CLINICALLY SIGNIFICANT CORONARY ARTERY DISEASE. CORONARY OBSTRUCTION CAN RESULT IN MYOCARDIAL ISCHEMIA OR INFARCTION DUE TO OBSTRUCTION OF THE CORONARY BLOOD FLOW AND MAY REQUIRE INTERVENTION (E.G., PCI). THERE ARE MULTIPLE PATIENT FACTORS THAT COULD PUT THE PATIENT AT RISK FOR CORONARY FLOW OBSTRUCTION DURING THE TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING SIGNIFICANT UNDERLYING CORONARY ARTERY DISEASE AND BULKY CALCIFICATION OF THE NATIVE LEAFLETS AND ROOT. DISPLACEMENT OF CALCIUM DEPOSITS WITH EMBOLIZATION OF DEBRIS INTO ONE OF THE ARTERIES, OR AORTIC DISSECTION WITH CONTINUITY OF THE RUPTURE INTO THE INTIMA OF ONE OF THE CORONARY OSTIA, CAN RESULT IN THIS COMPLICATION. ADDITIONALLY, MINIMAL DISTANCE BETWEEN THE NATIVE ANNULUS AND THE CORONARY OSTIA, BULKY CALCIFICATION, LONG NATIVE LEAFLETS, AND OBLITERATED CORONARY SINUSES. PROCEDURAL FACTORS SUCH AS TORN NATIVE LEAFLET DURING BALLOON VALVULOPLASTY, PLAQUE SHIFT, DEPLOYMENT OF THE BIOPROSTHETIC HEART VALVE TOO AORTIC, AND SIGNIFICANT VALVE OVER SIZING COULD ALSO CONTRIBUTE TO THIS COMPLICATION. THE EDWARDS THV MANUALS ADVISE THE OPERATOR ON PRE-PROCEDURE ASSESSMENT OF THE AORTIC VALVE, ROOT, AND CORONARY ANATOMY. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE TRANSCATHETER HEART VALVE (THV). TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS AND PROCTORED PROCEDURES. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21633506&pc=NPT | 2015691-2025-01773 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-18 04:00 | NPT | EDWARDS SAPIEN 3 TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Myocardial Infarction | P140031 | E2016006 | 1 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 VALVE. THIS REPORT SUMMARIZES 1 AORTIC MYOCARDIAL INFARCTION DEATH EVENT(S) FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 84 - 84 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 FEMALE(S) AND 0 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER: E2016006 AND SUMMARIZES 1 MYOCARDIAL INFARCTION DEATH EVENT(S) FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 0 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 TRANSCATHETER HEART VALVE ARE: (B)(4). PER THE INSTRUCTIONS FOR USE (IFU), THROMBUS FORMATION, PLAQUE DISLODGEMENT, AND EMBOLIZATION THAT MAY RESULT IN MYOCARDIAL INFARCTION (MI) ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE. THE IFU CAUTIONS THAT THE SAFETY AND EFFECTIVENESS HAVE NOT BEEN ESTABLISHED FOR PATIENTS WITH BULKY CALCIFIED AORTIC VALVE LEAFLETS IN CLOSE PROXIMITY TO THE CORONARY OSTIA. IN ADDITION, CAUTION SHOULD BE EXERCISED WHEN IMPLANTING A BIOPROSTHESIS IN PATIENTS WITH CLINICALLY SIGNIFICANT CORONARY ARTERY DISEASE. MI RELATED TO THE TRANSCATHETER VALVE REPLACEMENT PROCEDURE AND THE VALVE IMPLANT WILL MANIFEST INTRA-PROCEDURALLY OR IN THE IMMEDIATE POST-OPERATIVE PERIOD AND IS TYPICALLY DUE TO A COMBINATION OF PATIENT AND/OR PROCEDURAL FACTORS. AS DEFINED IN THE VALVE ACADEMIC RESEARCH CONSORTIUM (VARC) PUBLICATION ON TRANSCATHETER VALVE REPLACEMENT COMPLICATIONS, THE PERI-PROCEDURAL INTERVAL IS INCLUSIVE OF ALL EVENTS THAT BEGIN WITHIN 72 HRS. OF THE INDEX PROCEDURE. MI'S THAT OCCUR AFTER THE PERI-PROCEDURAL PERIOD (>72 HRS.) ARE TYPICALLY RELATED TO THE PATIENT'S UNDERLYING CORONARY DISEASE. THERE ARE MULTIPLE PATIENT FACTORS THAT COULD PUT THE PATIENT AT RISK FOR MI DURING THE TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING SIGNIFICANT UNDERLYING CORONARY ARTERY DISEASE, AND BULKY CALCIFICATION OF THE NATIVE LEAFLETS AND ROOT. DISPLACEMENT OF CALCIUM DEPOSITS WITH EMBOLIZATION OF DEBRIS INTO ONE OF THE ARTERIES, OR AORTIC DISSECTION WITH CONTINUITY OF THE RUPTURE INTO THE INTIMA OF ONE OF THE CORONARY OSTIA, CAN RESULT IN THIS COMPLICATION. THE EDWARDS THV MANUALS ADVISE THE OPERATOR ON PRE-PROCEDURE ASSESSMENT OF THE AORTIC VALVE, ROOT, AND CORONARY ANATOMY. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE TRANSCATHETER HEART VALVE (THV). TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS AND PROCTORED PROCEDURES. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21633508&pc=NPT | 2015691-2025-01772 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-18 04:00 | NPT | EDWARDS SAPIEN 3 TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Thrombosis/Thrombus | P140031 | E2016006 | 1 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 VALVE. THIS REPORT SUMMARIZES 1 AORTIC DEVICE THROMBOSIS DEATH EVENT(S) FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS (B)(6) YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 0 FEMALE(S) AND 1 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER: E2016006 AND SUMMARIZES 1 DEVICE THROMBOSIS DEATH EVENT(S) FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 9 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 TRANSCATHETER HEART VALVE ARE: (B)(4). PER THE INSTRUCTIONS FOR USE (IFU), VALVE THROMBOSIS IS A POTENTIAL RISK ASSOCIATED WITH THE USE OF THE TRANSCATHETER HEART VALVES (THV). THROMBOSIS IS A RARE AND WELL-RECOGNIZED COMPLICATION ASSOCIATED WITH THE USE OF BIOPROSTHETIC HEART VALVES. VALVE THROMBOSIS IS THE FORMATION OF SIGNIFICANT BLOOD CLOTS FORMING ON THE VALVE. POTENTIAL PATIENT RISK FACTORS SUCH AS ATRIAL FIBRILLATION, SYSTEMIC DISEASE (E.G., SYSTEMIC LUPUS ERYTHEMATOSUS, INFLAMMATION, AND DAMAGE TO VARIOUS BODY TISSUES, INCLUDING JOINTS, SKIN, KIDNEYS, HEART, LUNGS, BLOOD VESSELS, AND BRAIN), THE VALVULAR DISEASE ITSELF, AND REDUCED CARDIAC EJECTION FRACTION CAN CONTRIBUTE TO INCREASED RISK OF THROMBUS/THROMBOSIS. SUB-OPTIMAL ANTICOAGULATION DURING THE PROCEDURE AND UNDERLYING PATIENT CONDITIONS CAN ALSO RESULT IN INCREASED THROMBOGENICITY. THESE PATIENTS ARE ANTICOAGULATED FOR THE PROCEDURE AND INTERVENTIONAL BEST PRACTICES MANDATE METICULOUS PREPARATION, ASPIRATION, AND FLUSHING OF THE DEVICES TO PREVENT AND/OR REMOVE CLOT(S). SHORT-TERM ANTICOAGULATION THERAPY MAY ALSO BE NECESSARY AFTER VALVE REPAIR, ANTICOAGULATION IS PRESCRIBED PER INSTITUTIONAL GUIDELINES. INTRA-PROCEDURAL INTRA-CARDIAC THROMBUS AND POST PROCEDURE OR LATE VALVE THROMBOSIS ARE COMPLEX PROCESSES TRIGGERED BY THE INTERACTION BETWEEN THE HOST AND THE DEVICE WHICH ARE HIGHLY VARIABLE AMONG PATIENTS. IT IS THE NATURAL TENDENCY OF THE BODY TO FORM A CLOT ON FOREIGN OBJECTS IN THE VASCULAR SPACE AND A DEFINITIVE ROOT CAUSE CANNOT ALWAYS BE CONFIRMED. THE DEVICE TRAINING MANUALS INSTRUCT THE OPERATOR TO CONSIDER ALL PROCEDURAL AND ANATOMICAL FACTORS. EDWARDS EXTENSIVELY TRAINS PHYSICIANS BEFORE THEY ARE QUALIFIED TO USE THE DEVICE SYSTEM. TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES. OPERATORS ARE ALSO INSTRUCTED TO USE FLUOROSCOPY IN CONJUNCTION WITH ECHOCARDIOGRAPHY FOR OPTIMAL VISUALIZATION DURING POSITIONING AND DEPLOYMENT AND CAUTIONS INCLUDE THE MAINTENANCE OF THE PATIENT'S ANTICOAGULATION STATUS (ACT AT >250 SECONDS) DURING THE PROCEDURE. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21633511&pc=NPT | 2015691-2025-01767 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-18 04:00 | NPT | EDWARDS SAPIEN 3 TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Heart Block | P140031 | E2016006 | 5 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 VALVE. THIS REPORT SUMMARIZES 5 AORTIC PERMANENT PACEMAKER DEATH EVENT(S) FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 57 - 86 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 2 FEMALE(S) AND 3 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER: E2016006 AND SUMMARIZES 5 PERMANENT PACEMAKER DEATH EVENT(S) FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 9 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 TRANSCATHETER HEART VALVE ARE: (B)(4). PER THE INSTRUCTIONS FOR USE, CONDUCTION SYSTEM DEFECTS (HEART BLOCK) ARRHYTHMIAS AND CONDUCTION SYSTEM DEFECTS THAT MAY REQUIRE A PERMANENT PACEMAKER ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH BALLOON AORTIC VALVULOPLASTY, DEPLOYMENT OF THE PROSTHETIC VALVE, AND THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE. ACCORDING TO THE VALVE ACADEMIC RESEARCH CONSORTIUM (VARC) GUIDELINES, THE CLOSE ANATOMICAL RELATIONSHIP BETWEEN THE AORTIC VALVE COMPLEX AND THE BRANCHING ATRIOVENTRICULAR BUNDLE MAY PROVIDE AN EXPLANATION FOR THESE COMPLICATIONS OF THE TRANSCATHETER VALVE REPLACEMENT PROCEDURE. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY -CONDUCTION DISTURBANCES/ HEART BLOCK WRITTEN BY EDWARDS LIFESCIENCES, ATRIOVENTRICULAR CONDUCTION DISTURBANCES AFTER THV PLACEMENT ARE ASSOCIATED WITH MANY PATIENT RELATED AND PROCEDURAL RELATED FACTORS, INCLUDING PRE-OPERATIVE CO-MORBID STATUS, THE DEGREE AND BULKINESS OF AORTIC VALVE AND ANNULAR CALCIFICATION, INTER-VENTRICULAR SEPTAL THICKNESS, PRE-EXISTING ELECTROCARDIOGRAM ABNORMALITIES, THE DEPTH OF PROSTHESIS IMPLANTATION, AND THE PROFILE OF THE IMPLANTED PROSTHESIS. UNLIKE CONVENTIONAL AVR, WHERE THERE MAY BE LOCALIZED TRAUMA DUE TO DECALCIFICATION OF THE ANNULUS AND/OR SUTURE PLACEMENT IN THE PROXIMITY OF THE AV NODE OR THE BUNDLES, TRANSCATHETER VALVE REPLACEMENT MAY CAUSE CONDUCTION ABNORMALITIES THROUGH MECHANICAL IMPINGEMENT OF THE CONDUCTION SYSTEM BY THE PROSTHESIS. THE MECHANISMS OF THE DEVELOPMENT OF HEART BLOCK AFTER THV PLACEMENT ARE WELL DOCUMENTED AND DESCRIBED IN THE LITERATURE. IT IS ALSO DOCUMENTED THAT PRE-EXISTING HEART BLOCK IS COMMON IN PATIENTS UNDERGOING THV PLACEMENT OR SURGICAL AVR AND ANOTHER 4-6 % WILL DEVELOP POSTOPERATIVE HEART BLOCK, POTENTIALLY REQUIRING A PERMANENT PACEMAKER. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE, HOWEVER THE CONDUCTION DISTURBANCE IS LIKELY RELATED TO THE MECHANISM DESCRIBED ABOVE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21633516&pc=NPT | 2015691-2025-01770 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-18 04:00 | NPT | EDWARDS SAPIEN 3 TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Ischemia Stroke | P140031 | E2016006 | 2 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 VALVE. THIS REPORT SUMMARIZES 2 AORTIC STROKE - ISCHEMIC DEATH EVENT(S) FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 73 - 80 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 FEMALE(S) AND 1 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER: E2016006 AND SUMMARIZES 2 STROKE - ISCHEMIC DEATH EVENT(S) FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 7.5 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 TRANSCATHETER HEART VALVE ARE: (B)(4). PER THE INSTRUCTIONS FOR USE, PERMANENT OR TRANSIENT NEUROLOGICAL EVENTS INCLUDING STROKE ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH INVASIVE CARDIAC PROCEDURE INCLUDING TRANSCATHETER HEART VALVE PROCEDURES AND THE USE OF THE EDWARDS THV DEVICES. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, STROKE IS RECOGNIZED IN THE LITERATURE AS A WELL-KNOWN COMPLICATION IN A SMALL NUMBER OF PATIENTS UNDERGOING TRANSCATHETER VALVE REPLACEMENT PROCEDURE. RISK FACTORS CORRELATING WITH A NUMBER OF PATIENT CO-MORBIDITIES HAVE BEEN IDENTIFIED. ALTHOUGH IN MANY CASES THE ROOT CAUSE OF THE EVENT IS UNABLE TO BE DETERMINED, STROKES DURING TRANSCATHETER VALVE REPLACEMENT PROCEDURE ARE UNDOUBTEDLY MULTIFACTORIAL, THE DOMINANT ETIOLOGY LIKELY BEING INTRA-PROCEDURE EMBOLIC EVENTS. A TRANSCRANIAL DOPPLER STUDY DURING TRANSCATHETER VALVE REPLACEMENT PROCEDURE DEMONSTRATED THAT THE MAJORITY OF PROCEDURAL EMBOLIC EVENTS OCCURRED DURING BALLOON VALVULOPLASTY, MANIPULATION OF CATHETERS ACROSS THE AORTIC VALVE, AND VALVE IMPLANTATION. AN ANALYSIS IN PATIENTS UNDERGOING VALVE SURGERY REVEALED FOUR BASELINE CHARACTERISTICS AND TWO PROCEDURAL EVENTS THAT WERE ASSOCIATED WITH EARLY POST-PROCEDURE STROKE: FEMALE SEX, EF < 30%, DIABETES, AGE OLDER THAN 70 YEARS, BYPASS PROCEDURE TIME> 120 MIN, AND CALCIFICATION OF THE ASCENDING AORTA. PREDICTORS OF LATE STROKE HAVE INCLUDED FEMALE SEX, AGE OLDER THAN 75 YEARS, ATRIAL FIBRILLATION, AND A HISTORY OF OR CURRENT SMOKING. THERE WERE NO IMPORTANT DIFFERENCES IN THE FREQUENCY OF LATE STROKES BETWEEN TRANSCATHETER VALVE REPLACEMENT PROCEDURE AND AVR PATIENTS. AFTER TRANSCATHETER VALVE REPLACEMENT PROCEDURE, THERE APPEARS TO BE A MORE SIGNIFICANT PROPORTION OF EARLY STROKES OCCURRING < 24 H POST-PROCEDURE, BUT TRANSCATHETER VALVE REPLACEMENT PROCEDURE PATIENTS WITH MULTIPLE CO MORBIDITIES ARE PROBABLY AT HIGHER RISK OF BOTH EARLY AND LATE STROKES. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE HOWEVER, THE EVENT MAY BE RELATED TO THE MECHANISM DESCRIBED ABOVE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21633517&pc=NPT | 2015691-2025-01766 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-18 04:00 | NPT | EDWARDS SAPIEN 3 TRANSCATHETER HEART VALVE | Insufficient Information | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P140031 | E2016006 | 6 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 VALVE. THIS REPORT SUMMARIZES 6 AORTIC READMISSION (VALVE RELATED) DEATH EVENT(S) FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 72 - 92 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 0 FEMALE(S) AND 6 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER: E2016006 AND SUMMARIZES 6 READMISSION (VALVE RELATED) DEATH EVENT(S) FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 176.67 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 TRANSCATHETER HEART VALVE ARE: (B)(4). VALVE RELATED RE-ADMISSION IN THE FOLLOW-UP PERIOD IS LIKELY DUE TO REOCCURRENCE OF SYMPTOMS. THIS MAY RESULT FROM REGURGITATION (CENTRAL OR PVL) OR PROGRESSION OF THE PRE-EXISTING DISEASE PROCESS AND MAY RESULT IN HEART FAILURE. CAUSES OF HEART FAILURE CAN BE MULTI-FACTORIAL. PER THE INSTRUCTION FOR USE (IFU), HEART FAILURE, VALVE REGURGITATION, AND VALVE DEGENERATION INCLUDING STENOSIS ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH AORTIC VALVE REPLACEMENT AND BIOPROSTHETIC HEART VALVES. CONGESTIVE HEART FAILURE CAN HAVE MULTIPLE ETIOLOGIES AND IS OFTEN DUE TO THE PROGRESSION OF UNDERLYING DISEASE PROCESSES, INCLUDING CORONARY ARTERY DISEASE, UNCONTROLLED HYPERTENSION, MYOCARDIAL INFARCTION, CARDIOMYOPATHY, FLUID OVERLOAD, OR VALVULAR DYSFUNCTION. RISK FACTORS THAT INCREASE A PATIENT'S RISK OF SUFFERING FROM CHF INCLUDE OBESITY, ADVANCED AGE, AND A HISTORY OF SMOKING. REGURGITATION WHICH DEVELOPS PROGRESSIVELY OVER TIME CAN BE DUE TO A NUMBER OF ISSUES INCLUDING PATIENT RELATED FACTORS OR STRUCTURAL VALVE DETERIORATION, INCLUDING CALCIFICATION, NON-CALCIFIC DEGENERATION, LEAFLET THICKENING OR FIBROSIS, OR A COMBINATION OF THESE. REGURGITATION MAY ALSO DEVELOP PROGRESSIVELY IF HOST FIBROTIC TISSUE, OR PANNUS, GROWS ONTO THE BIOPROSTHETIC VALVE. PANNUS, A CAUSE OF NONSTRUCTURAL DYSFUNCTION, MAY INTERFERE WITH FUNCTIONALITY OF THE DEVICE BY RESTRICTING THE LEAFLET MOTION LEADING TO ABNORMAL COAPTATION. THE MECHANISM BEHIND WORSENING OR LATE PVL IS NOT WELL UNDERSTOOD BUT MAY BE RELATED TO CARDIAC REMODELING. VALVE STENOSIS MAY RESULT IN SYMPTOMS SUCH AS SOB AND DECREASED EXERCISE TOLERANCE, WHICH MAY BE ACCOMPANIED BY AN INCREASED GRADIENT ACROSS THE VALVE. THIS COULD BE DUE TO EARLY CALCIFICATION OF THE LEAFLETS, HOST TISSUE OVERGROWTH, OR IN RARE CASES A NON-FUNCTIONING LEAFLET. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21633648&pc=OZD | 1220648-2025-26804 | 2025-02-20 05:00 | Death | ABIOMED, INC. | 2025-03-18 04:00 | OZD | IMPELLA CP | Failure to Pump | Ischemia | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT ON IMPELLA CP FOR ACUTE MYOCARDIAL INFARCTION. IT WAS NOTED THAT THE PATIENT HAD DIMINISHED BILATERAL PEDAL PULSES. THE IMPELLA SIDE WAS THE WORSE OF THE TWO. VASCULAR SURGERY WAS FOLLOWING, HOWEVER, THERE WAS NO PLAN TO REMOVE THE IMPELLA OR SURGICAL INTERVENTION PLANNED. ADDITIONALLY, THE UNIT NURSE STATED THAT THE PUMP STOPPED SPONTANEOUS. ATTEMPTS WERE MADE TO RESTART. HOWEVER, THEY WERE UNSUCCESSFUL. THE AUTOMATED IMPELLA CONTROLLER WAS SWAP OUT AND WAS ALSO UNSUCCESSFUL. THEY NOTICED A MOTOR CURRENT SPIKE ON CONNECT. ULTIMATELY, THE PATIENT CODED AND EVENTUALLY WAS MADE DO NOT RESUSCITATE AND EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH RISK PERCUTANEOUS CORONARY INTERVENTION. SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA CATHETER ¿ACHIEVING AN ACT = 250 SECONDS PRIOR TO REMOVING THE DILATOR WILL HELP PREVENT A THROMBUS FROM ENTERING THE CATHETER AND CAUSING A SUDDEN STOP ON STARTUP.¿ SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA RP WITH SMARTASSIST SYSTEM CATHETER. ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA RP WITH SMARTASSIST SYSTEM CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP. MONITOR MOTOR CURRENT AND CONSIDER REPLACING IMPELLA RP WITH SMARTASSIST SYSTEM CATHETER WHENEVER A RISE IN MOTOR CURRENT IS SEEN.¿ IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK: SECTION: TROUBLESHOOTING THE PURGE SYSTEM ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP.¿ SECTION: IMPELLA STOPPED: ¿IF THE IMPELLA CATHETER HAS STOPPED SUDDENLY: 1. TRY TO RESTART THE CATHETER AT PREVIOUS P-LEVEL. 2. IF THE IMPELLA DOES NOT RESTART, TRY TO RESTART AT P-2. 3. IF THE IMPELLA DOES NOT RESTART OR STOPS AGAIN, WAIT 1 MINUTE AND TRY TO RESTART AGAIN. 4. IF THE IMPELLA RESTARTS, WEAN DOWN TO P-2 AS THE PATIENT CAN TOLERATE. UNDER THESE CIRCUMSTANCES, CATHETER FUNCTION IS NOT RELIABLE AND THE IMPELLA MAY STOP AGAIN. 5. IF THE IMPELLA DOES NOT RESTART, REMOVE THE IMPELLA FROM THE VENTRICLE AS SOON AS POSSIBLE TO AVOID AORTIC INSUFFICIENCY.¿ IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH RISK PERCUTANEOUS CORONARY INTERVENTION. SECTION: WARNINGS, CAUTIONS, AND PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21633702&pc=DQX | 2025587-2025-01967 | 2025-02-20 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-18 04:00 | DQX | CONFIDA BRECKER CURVE GUIDEWIRE | Patient Device Interaction Problem | Bradycardia; Low Blood Pressure/ Hypotension; Perforation | K132623 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE PROCEDURE INVOLVING THE EVFXPLUS-29, THERE WAS A DROP IN BLOOD PRESSURE AND HEART RATE AS THE DELIVERY CATHETER SYSTEM WAS ADVANCED THROUGH THE DESCENDING AORTA AND AROUND THE AORTIC ARCH. THE VALVE WAS POSITIONED ACROSS THE NATIVE VALVE, AND THE BLOOD PRESSURE AND HEART RATE CONTINUED TO DECREASE, LEADING TO THE VALVE BEING PULLED BACK INTO THE ASCENDING AORTA. CARDIOPULMONARY RESUSCITATION WAS INITIATED, AND THE PATIENT WAS PREPPED FOR OPEN HEART SURGERY AND PLACED ON EXTRACORPOREAL MEMBRANE OXYGENATION. IT WAS NOTED THAT THE GUIDEWIRE PERFORATED THE LEFT VENTRICLE (LV), RESULTING IN A LV PERFORATION. THE PROCEDURE WAS ULTIMATELY ABORTED. THE PATIENT PASSED AWAY THE SAME DAY. THE CAUSE OF DEATH WAS LV PERFORATION.Event Description: ADDITIONAL INFORMATION WAS RECEIVED FROM THE PHYSICIAN TO NOTE THE DCS CANNOT BE EXCLUDED AS A CONTRIBUTING FACTOR TO THE PATIENT'S DEATH. Manufacturer Narrative: CONTINUATION OF D10: SECTION D INFORMATION REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER RELEVANT DEVICE(S) ARE: PRODUCT ID: D -EVOLUTFX-2329. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: A4. B5. E1. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21634317&pc=NPS | 2135147-2025-01393 | 2016-03-01 05:00 | Death | ABBOTT MEDICAL | 2025-03-18 04:00 | NPS | TRICLIP DELIVERY SYSTEM (TCDS) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P230007 |  | 1 | Event Description: THE ARTICLE "MALNUTRITION AND OUTCOMES IN PATIENTS WITH TRICUSPID REGURGITATION UNDERGOING TRANSCATHETER TRICUSPID VALVE REPAIR" WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE MULTI CENTER STUDY, TO EVALUATE THE IMPACT OF MALNUTRITION IN PATIENTS WITH TRICUSPID REGURGITATION (TR) UNDERGOING TRICUSPID TRANSCATHETER EDGE-TO-EDGE REPAIR (T-TEER) IS NOT WELL ESTABLISHED. WE EVALUATED THE IMPACT OF MALNUTRITION AMONG PATIENTS WITH SYMPTOMATIC TR UNDERGOING T-TEER. DEVICES MENTIONED INCLUDE TRICLIP, MITRACLIP, AND A NON-ABBOTT DEVICE. THE ARTICLE CONCLUDED THAT IN THE LARGE, REAL-WORLD, MULTICENTRE EUROTR REGISTRY, MALNUTRITION WAS PRESENT IN ONE OUT OF FIVE PATIENTS WITH SYMPTOMATIC TR UNDERGOING T-TEER AND WAS INDEPENDENTLY ASSOCIATED WITH INCREASED MORTALITY. THE PROGNOSTIC BENEFIT OF SUCCESSFUL T-TEER IN REDUCING MORTALITY WAS CONSISTENTLY OBSERVED IN PATIENTS WITH AND WITHOUT MALNUTRITION. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS MARIANNA ADAMO AT UNIVERSITY OF BRESCIA INSTITUTION WITH CORRESPONDING EMAIL MARIANNAADAMO@HOTMAIL.COM]. THE TIME FRAME OF THE STUDY WAS MARCH 2016 TO FEBRUARY 2024. A TOTAL OF 1034 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH 645 RECEIVED AN ABBOTT DEVICE. AS THIS EVENT IS FROM A LITERATURE REVIEW, THERE IS NO RELEVANT PATIENT INFORMATION (DATE OF BIRTH, AGE, GENDER, WEIGHT, AND MEDICAL HISTORY) TO REPORT. COMORBIDITIES INCLUDED HYPERTENSION, DIABETES, ATRIAL FIBRILLATION, PRIOR CARDIAC SURGERY. (B)(6), UNK MITRACLIP; PERI- AND POST-PROCEDURAL COMPLICATIONS INCLUDED HEART FAILURE, PROLONGED HOSPITALIZATION, DEATH. (B)(6), UNK TRICLIP; PERI- AND POST-PROCEDURAL COMPLICATIONS INCLUDED HEART FAILURE, PROLONGED HOSPITALIZATION, DEATH. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF TRICLIP DEVICE WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING HYPERTENSION, DIABETES, ATRIAL FIBRILLATION, PRIOR CARDIAC SURGERY. COMPLICATIONS REPORTED INCLUDED HEART FAILURE, PROLONGED HOSPITALIZATION, DEATH; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. Manufacturer Narrative: IT IS UNKNOWN IF THE DEVICE IS RETURNING FOR ANALYSIS. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21634437&pc=DSQ | 2916596-2025-01599 | 2025-03-01 05:00 | Death | THORATEC CORPORATION | 2025-03-18 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO CARDIAC ARREST. THE OUTCOME WAS CONSIDERED DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. AN AUTOPSY WAS NOT PERFORMED, AND THE DEVICE WAS NOT EXPLANTED OR RETURNED FOR EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN DETAILS LEADING UP TO THE CARDIAC ARREST AND THE PATIENT¿S PASSING; HOWEVER, NO RESPONSE WAS RECEIVED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, REV. C IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21634450&pc=MKJ | 3030677-2025-000882 | 2023-08-04 04:00 | Death | PHILIPS NORTH AMERICA LLC | 2025-03-18 04:00 | MKJ | HEARTSTART FRX | Failure to Deliver Shock/Stimulation | Cardiac Arrest | P180028 |  | 1 | Event Description: IT HAS BEEN REPORTED THAT THE DEVICE DID NOT SHOCK. THE PATIENT DID NOT SURVIVE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21634452&pc=NPT | 2025587-2025-01972 | 2025-03-10 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-18 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Perforation; Pericardial Effusion | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT UPON DEPLOYMENT OF A TRANSCATHETER AORTIC VALVE, ANGIOGRAPHY AND ECHOCARDIOGRAPHY WERE PERFORMED THAT REVEALED PERICARDIAL EFFUSION. SUBSEQUENTLY, SURGICAL REPAIR OF THE LEFT VENTRICLE WAS PERFORMED. THE FOLLOWING EVENING, THE PATIENT DIED IN THE INTENSIVE CARE UNIT (ICU). PER THE PHYSICIAN, THE PROCEDURE CONTRIBUTED TO THE DEATH.Event Description: IT WAS REPORTED THAT UPON DEPLOYMENT OF A TRANSCATHETER AORTIC VALVE, ANGIOGRAPHY AND ECHOCARDIOGRAPHY WERE PERFORMED THAT REVEALED PERICARDIAL EFFUSION. SUBSEQUENTLY, SURGICAL REPAIR OF THE LEFT VENTRICLE WAS PERFORMED. THE FOLLOWING EVENING, THE PATIENT DIED IN THE INTENSIVE CARE UNIT (ICU). PER THE PHYSICIAN, THE PROCEDURE CONTRIBUTED TO THE DEATH.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE CAUSE OF DEATH WAS DUE TO A LEFT VENTRICLE PERFORATION. PER THE PHYSICIAN, IT WAS UNKNOWN IF THE DELIVERY CATHETER SYSTEM (DCS) CAUSED OR CONTRIBUTED TO THE PERICARDIAL EFFUSION AND SUBSEQUENT DEATH, BUT THE VALVE DID NOT CAUSE OR CONTRIBUTE TO THE PERICARDIAL EFFUSION AND SUBSEQUENT DEATH. PER THE PHYSICIAN, THE PERICARDIAL EFFUSION WAS POSSIBLY CAUSED BY THE GUIDEWIRE.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE PERFORATION MAY HAVE BEEN CAUSED BY THE CROSSING GUIDEWIRE, WHICH WAS NOT A MEDTRONIC DEVICE. Manufacturer Narrative: UPDATED DATA: B5. D4. FULL UDI. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. H6. D10. CONTINUATION OF D10: PRODUCT ID {GUIDEWIRE}; PRODUCT LOT/SERIAL NUMBER {UNKNOWN}; PRODUCT TYPE: {GUIDEWIRE}; IMPLANT DATE {N/A}; EXPLANT DATE {N/A} MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: SECTION D INFORMATION REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER RELEVANT DEVICE(S) ARE: PRODUCT ID: D- EVOLUTFX-2329, SERIAL/LOT #: UNKNOWN, UBD: UNKNOWN, UDI#: UNKNOWN. THE CODES PRESENT IN SECTION H6 CORRESPOND TO MULTIPLE COMPONENTS/PRODUCTS THAT COMPRISE THIS REPORTED EVENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: PRODUCT ID: D-EVOLUTFX-2329, LOT #: 0012460560 USE BEFORE DATE: 2026-09-25 FULL UDI#:(B)(4), MANUFACTURED DATE: 2024-09-25, NON-IMPLANTABLE PRODUCT. H6. EVAL CODE METHOD FOR THE SYSTEM REPORTED DEVICE WAS CHANGED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21634903&pc=DSQ | 2916596-2025-01876 | 2025-02-15 05:00 | Death | THORATEC CORPORATION | 2025-03-18 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE HEARTMATE 3 LVAS IFU, REV. G IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21634904&pc=DSQ | 2916596-2025-01491 | 2025-02-22 05:00 | Death | THORATEC CORPORATION | 2025-03-18 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage; Loss of consciousness | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT WAS FOUND UNRESPONSIVE AT HOME AND WAS TAKEN TO THE HOSPITAL. COMPUTED TOMOGRAPHY WAS PERFORMED AND SHOWED A MASSIVE INTRACRANIAL BLEED; HER INTERNATIONAL NORMALIZED RATIO (INR) WAS 3.1 SECONDS ON ADMISSION. THE PATIENT WAS NOT ON ASPIRIN. THE PATIENT PASSED AWAY DUE TO THE LARGE INTRACRANIAL BLEED OF AN UNKNOWN CAUSE. THE PATIENT WAS ALONE DURING THE EVENT, AND IT WAS UNKNOWN OF ANY SYMPTOMS WERE EXPERIENCED. THERE WAS NO KNOW CHANGES TO THE PATIENT'S ANTICOAGULATION STATUS PRIOR TO DEATH. PER LOG FILE FINDINGS THE PUMP APPEARED TO HAVE BEEN OPERATING AS INTENDED. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: THE PUMP REMAINS IN USE SUPPORTING THE PATIENT. A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF STROKE COULD NOT BE CONCLUSIVELY DETERMINED. STROKE IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. STROKE HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21635122&pc=NKM | 2135147-2025-01401 | 2025-02-21 05:00 | Death | ABBOTT MEDICAL | 2025-03-18 04:00 | NKM | MITRACLIP G4 CLIP DELIVERY SYSTEM (MDR) | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P100009 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A MITRACLIP PROCEDURE WAS PERFORMED TO TREAT DEGENERATIVE MITRAL REGURGITATION (MR) GRADE 4. A XTW MITRACLIP WAS IMPLANTED SUCCESSFULLY, REDUCING MR TO TRACE. ON (B)(6)2025, THE PATIENT RETURNED WITH HEMOPTYSIS. THE CAUSE WAS UNKNOWN. THE MITRACLIP REMAINED STABLE ON BOTH LEAFLETS. PATIENT DIED ON (B)(6) 2025.Event Description: SUBSEQUENT TO THE PREVIOUSLY FILED REPORT, ADDITIONAL INFORMATION WAS RECEIVED: THE CAUSE OF DEATH WAS REPORTED TO BE UNKNOWN. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A REVIEW OF THE LOT HISTORY RECORD REVEALED NO MANUFACTURING NONCONFORMITIES ISSUED TO THE REPORTED LOT THAT WOULD HAVE CONTRIBUTED TO THIS EVENT. THE INVESTIGATION WAS UNABLE TO DETERMINE A CAUSE FOR REPORTED DEATH AND HEMORRHAGE. DEATH AND HEMORRHAGE ARE LISTED IN THE INSTRUCTIONS FOR USE AS KNOWN POSSIBLE COMPLICATIONS ASSOCIATED WITH THE MITRACLIP PROCEDURES. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21635187&pc=NKM | 2135147-2025-01405 | 2025-02-21 05:00 | Death | ABBOTT MEDICAL | 2025-03-18 04:00 | NKM | MITRACLIP G4 CLIP DELIVERY SYSTEM (MDR) | Adverse Event Without Identified Device or Use Problem | Pulmonary Embolism; Dyspnea | P100009 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A MITRACLIP PROCEDURE WAS PERFORMED TO TREAT DEGENERATIVE MITRAL REGURGITATION (MR) GRADE 4. TWO XTW MITRACLIPS WERE IMPLANTED SUCCESSFULLY REDUCING MR TO TRACE. POST PROCEDURE ON CT, PULMONARY EMBOLISM WAS OBSERVED WHICH WAS TREATED SUCCESSFULLY WITH ANTICOAGULANTS.Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A MITRACLIP PROCEDURE WAS PERFORMED TO TREAT DEGENERATIVE MITRAL REGURGITATION (MR) GRADE 4. TWO XTW MITRACLIPS WERE IMPLANTED SUCCESSFULLY REDUCING MR TO TRACE. ON (B)(6) 2025, CT DIAGNOSED A PULMONARY EMBOLISM WHICH WAS TREATED SUCCESSFULLY WITH ANTICOAGULANTS. PATIENT WAS SENT TO HOSPICE AND DIED ON (B)(6) 2025 FROM MULTIPLE COMORBIDITIES. Manufacturer Narrative: H1 TYPE OF REPORTABLE EVENT: CHANGED FROM SERIOUS INJURY TO DEATH. H6 HEALTH EFFECT - IMPACT CODE: 1802 DEATH ADDED. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A REVIEW OF THE LOT HISTORY RECORD REVEALED NO MANUFACTURING NONCONFORMITIES ISSUED TO THE REPORTED LOT THAT WOULD HAVE CONTRIBUTED TO THIS EVENT. THE INVESTIGATION WAS UNABLE TO DETERMINE A CAUSE FOR THE REPORTED PULMONARY EMBOLISM AND DEATH. THE REPORTED DYSPNEA APPEARS TO BE A CASCADING EFFECT OF THE EMBOLISM. PULMONARY EMBOLISM, DYSPNEA, AND DEATH ARE LISTED IN THE INSTRUCTIONS FOR USE ARE KNOWN POSSIBLE COMPLICATIONS ASSOCIATED WITH MITRACLIP PROCEDURES. THE REPORTED MEDICATION REQUIRED WAS A RESULT OF CASE-SPECIFIC CIRCUMSTANCES AS ANTICOAGULANTS WAS PROVIDED TO TREAT THE PULMONARY EMBOLISM. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURE, DESIGN, OR LABELING. Manufacturer Narrative: THE DEVICE IS NOT RETURNING FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21635432&pc=OZD | 1220648-2025-26813 | 2025-02-21 05:00 | Death | ABIOMED, INC. | 2025-03-18 04:00 | OZD | IMPELLA CP | Malposition of Device; Detachment of Device or Device Component | Cardiac Arrest | P140003 |  | 1 | Event Description: THE COMPLAINANT HAD PLACED A IMPELLA CP PUMP TO SUPPORT A PATIENT ADMITTED IN WITH MULTIPLE CARDIAC COMORBIDITIES AND SUFFERING FROM ACUTE MYOCARDIAL INFARCTION AND CARDIOGENIC SHOCK. THE PATIENT HAD BEEN SUPPORTED BY AN INTRA-AORTIC BALLOON PUMP (IABP) ON THE RIGHT SIDE PRIOR TO THE IMPELLA CP. THE IMPELLA CP WAS INSERTED AND POSITIONED IN LEFT VENTRICLE SUCCESSFULLY. THERE WAS DIFFICULTY REMOVED IABP ULTIMATELY ENDING WITH APPLYING FEMSTOP. IT WAS NOTED THAT IABP EXPLANT TOOK APPROX 20 MINUTES. SHORTLY AFTER FEMSTOP WAS APPLIED, IT WAS NOTED THAT IMPELLA WAS ENCOUNTERING MORE FREQUENT SUCTION. IMPELLA POSITION UNDER FLURO WAS NOTED TO BE UNCHANGED HOWEVER SUPPORT DECREASED TO THE POINT WHERE IMPELLA WAS ABLE TO ACHIEVE FLOWS HIGHER THAN P3. ECHOCARDIOGRAM (ECHO)WAS REQUESTED BY LOCAL SUPPORT TO ASSESS RIGHT VENTRICLE (RV) FUNCTION. PRESSURE WAS NOTED TO BE DROPPING. THE DOCTOR INITIALLY WANTED TO TRANSFER PATIENT BACK TO INTENSIVE CARE UNIT HOWEVER PATIENT CODED BEFORE TRANSFER COULD BE DONE. ECHO ARRIVED AND LIMITED STUDY WAS PERFORMED. RV SHOWED LITTLE TO NO FUNCTION AT TIME OF STUDY. RETURN OF SPONTANEOUS CIRCULATION WAS ACHIEVED AND THE DOCTOR OPTED TO TRANSFER PATIENT TO CARDIOVASCULAR INTENSIVE CARE UNIT (ICU). UPON ARRIVAL TO ICU, AUTOMATED IMPELLA CONTROLLER WAVEFORMS INDICATED DEVICE IN AORTA. IMPELLA ACCESS WAS INSPECTED BY LOCAL SUPPORT SHOWING IMPELLA STERILE SLEEVE HAD BECOME DISCONNECTED AND DEVICE APPEARED TO HAVE BEEN PULLED BACK. PATIENT¿S PRESSURE DROPPED AND PATIENT ULTIMATELY CODED AGAIN. ECHO ARRIVED AND LIMITED STUDY AGAIN PERFORMED SHOWING DEVICE NO LONGER ACROSS ATRIOVENTRICULAR (AV). THE DOCTOR ATTEMPTED TO RECROSS AV BUT WAS NOT SUCCESSFUL. PATIENT CONTINUED TO BE CODED HOWEVER, DESPITE AGGRESSIVE EFFORTS THE PATIENT EXPIRED. THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21635498&pc=DSQ | 2916596-2025-01653 | 2025-02-28 05:00 | Death | THORATEC CORPORATION | 2025-03-18 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. NO AUTOPSY WOULD BE PERFORMED, AND THE PUMP WOULD NOT BE EXPLANTED. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. IT WAS COMMUNICATED THAT DUE TO PATIENT PRIVACY LAWS, THE MANAGING HOSPITAL WILL NOT COMMUNICATE ANY FURTHER PATIENT OUTCOME INFORMATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REV. C IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21635638&pc=DSQ | 2916596-2025-01643 | 2025-01-28 05:00 | Death | THORATEC CORPORATION | 2025-03-18 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT'S FAMILY DECIDED TO INITIATE COMFORT CARE, AND THE PATIENT PASSED AWAY ON (B)(6) 2025. THE PUMP WOULD NOT RETURN. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICLE ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21635977&pc=LWS | 2649622-2025-07594 | 2025-02-20 05:00 | Death | MPRI | 2025-03-18 04:00 | LWS | SPRINT QUATTRO SECURE S MRI SURESCAN | Under-Sensing | Cardiac Arrest; Septic Shock | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE RIGHT ATRIAL (RA) LEAD EXHIBITED UNDERSENSING TRIGGERING DEVICE CLASSIFIED FALSE TERMINATION OF ATRIAL TACH YCARDIA/ATRIAL FIBRILLATION (AT/AF) EVENTS(S) AT TIMES. THE RIGHT VENTRICULAR (RV) LEAD UNDERSENSED A SINGLE BEAT ON A STORED VENTRICULAR TACHYCARDIA (VT) EPISODE. THREE DAYS LATER, THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) PATIENT EXPIRED. THE CAUSE OF DEATH WAS DUE TO CARDIOPULMONARY ARREST AND SEPTIC SHOCK. THE SOURCE OF THE INFECTION WAS NOT PROVIDED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21635986&pc=DTB | 2649622-2025-07595 | 2025-02-20 05:00 | Death | MPRI | 2025-03-18 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Under-Sensing | Cardiac Arrest; Septic Shock | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE RIGHT ATRIAL (RA) LEAD EXHIBITED UNDERSENSING TRIGGERING DEVICE CLASSIFIED FALSE TERMINATION OF ATRIAL TACH YCARDIA/ATRIAL FIBRILLATION (AT/AF) EVENTS(S) AT TIMES. THE RIGHT VENTRICULAR (RV) LEAD UNDERSENSED A SINGLE BEAT ON A STORED VENTRICULAR TACHYCARDIA (VT) EPISODE. THREE DAYS LATER, THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) PATIENT EXPIRED. THE CAUSE OF DEATH WAS DUE TO CARDIOPULMONARY ARREST AND SEPTIC SHOCK. THE SOURCE OF THE INFECTION WAS NOT PROVIDED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21636004&pc=LWS | 9614453-2025-01087 | 2025-02-20 05:00 | Death | MEDTRONIC EUROPE SARL | 2025-03-18 04:00 | LWS | EVERA MRI XT DR SURESCAN | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Septic Shock | P980016 |  | 1 | Event Description: IT WAS REPORTED THAT THE RIGHT ATRIAL (RA) LEAD EXHIBITED UNDERSENSING TRIGGERING DEVICE CLASSIFIED FALSE TERMINATION OF ATRIAL TACH YCARDIA/ATRIAL FIBRILLATION (AT/AF) EVENTS(S) AT TIMES. THE RIGHT VENTRICULAR (RV) LEAD UNDERSENSED A SINGLE BEAT ON A STORED VENTRICULAR TACHYCARDIA (VT) EPISODE. THREE DAYS LATER, THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) PATIENT EXPIRED. THE CAUSE OF DEATH WAS DUE TO CARDIOPULMONARY ARREST AND SEPTIC SHOCK. THE SOURCE OF THE INFECTION WAS NOT PROVIDED. Manufacturer Narrative: CORRECTION: H6 EVAL CODE CONCLUSION (FDC/ANNEX D). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21636783&pc=KRA | 2124215-2025-16351 | 2024-12-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-03-18 04:00 | KRA | EKOSONIC KIT 106CM 12CM TZ | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K213422 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. TWO EKOSONIC CATHETERS WERE SELECTED FOR USE DURING A BILATERAL PULMONARY EMBOLISM EKOS PROCEDURE. THE PROTOCOL USED WAS 1MG/HR/CATHETER OF TISSUE PLASMINOGEN ACTIVATOR (TPA) FOR 6 HOURS. THE PATIENT WAS ALSO BEING TREATED WITH ANTICOAGULATION. THE PATIENT DID NOT IMPROVE FOLLOWING THE PROTOCOL, AND ADDITIONAL INTERVENTION WAS PERFORMED; HOWEVER, WAS UNSUCCESSFUL. THE PATIENT LATER PASSED AWAY DURING THE HOSPITAL STAY. Manufacturer Narrative: B2: DATE OF DEATH AND B3: DATE OF EVENT: APPROXIMATE, AS THE REPORTED INFORMATION NOTED THIS OCCURRED IN 'LATE 2024'. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION, BUT IT WAS UNABLE TO BE OBTAINED. DETAILED PHYSICIAN ADDRESS AND FACILITY INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION, BUT IT WAS UNABLE TO BE OBTAINED. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION, BUT IT WAS UNABLE TO BE OBTAINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21636785&pc=KRA | 2124215-2025-16359 | 2025-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-03-18 04:00 | KRA | EKOSONIC KIT 106CM 12CM TZ | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K213422 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. TWO EKOSONIC CATHETERS WERE SELECTED FOR USE DURING A BILATERAL PULMONARY EMBOLISM EKOS PROCEDURE. THE PROTOCOL USED WAS 1MG/HR/CATHETER OF TISSUE PLASMINOGEN ACTIVATOR (TPA) FOR 6 HOURS. THE PATIENT WAS ALSO BEING TREATED WITH ANTICOAGULATION. THE PATIENT DID NOT IMPROVE FOLLOWING THE PROTOCOL, AND ADDITIONAL INTERVENTION WAS PERFORMED; HOWEVER, WAS UNSUCCESSFUL. THE PATIENT LATER PASSED AWAY DURING THE HOSPITAL STAY. Manufacturer Narrative: B2: DATE OF DEATH AND B3: DATE OF EVENT: APPROXIMATE, AS THE REPORTED INFORMATION NOTED THIS OCCURRED IN 'EARLY (B)(6) 2025. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION, BUT IT WAS UNABLE TO BE OBTAINED. DETAILED PHYSICIAN ADDRESS AND FACILITY INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION, BUT IT WAS UNABLE TO BE OBTAINED. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION, BUT IT WAS UNABLE TO BE OBTAINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21636790&pc=KRA | 2124215-2025-16353 | 2025-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-03-18 04:00 | KRA | EKOSONIC KIT 106CM 12CM TZ | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K213422 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. TWO EKOSONIC CATHETERS WERE SELECTED FOR USE DURING A BILATERAL PULMONARY EMBOLISM EKOS PROCEDURE. THE PROTOCOL USED WAS 1MG/HR/CATHETER OF TISSUE PLASMINOGEN ACTIVATOR (TPA) FOR 6 HOURS. THE PATIENT WAS ALSO BEING TREATED WITH ANTICOAGULATION. THE PATIENT DID NOT IMPROVE FOLLOWING THE PROTOCOL, AND ADDITIONAL INTERVENTION WAS PERFORMED; HOWEVER, WAS UNSUCCESSFUL. THE PATIENT LATER PASSED AWAY DURING THE HOSPITAL STAY. Manufacturer Narrative: B2: DATE OF DEATH AND B3: DATE OF EVENT: APPROXIMATE, AS THE REPORTED INFORMATION NOTED THIS OCCURRED IN 'EARLY (B)(6) 2025'. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION, BUT IT WAS UNABLE TO BE OBTAINED. DETAILED PHYSICIAN ADDRESS AND FACILITY INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION, BUT IT WAS UNABLE TO BE OBTAINED. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION, BUT IT WAS UNABLE TO BE OBTAINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21636793&pc=KRA | 2124215-2025-16344 | 2024-12-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-03-18 04:00 | KRA | EKOSONIC KIT 106CM 12CM TZ | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K213422 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. TWO EKOSONIC CATHETERS WERE SELECTED FOR USE DURING A BILATERAL PULMONARY EMBOLISM EKOS PROCEDURE. THE PROTOCOL USED WAS 1MG/HR/CATHETER OF TISSUE PLASMINOGEN ACTIVATOR (TPA) FOR 6 HOURS. THE PATIENT WAS ALSO BEING TREATED WITH ANTICOAGULATION. THE PATIENT DID NOT IMPROVE FOLLOWING THE PROTOCOL, AND ADDITIONAL INTERVENTION WAS PERFORMED; HOWEVER, WAS UNSUCCESSFUL. THE PATIENT LATER PASSED AWAY DURING THE HOSPITAL STAY. Manufacturer Narrative: B2: DATE OF DEATH AND B3: DATE OF EVENT: APPROXIMATE, AS THE REPORTED INFORMATION NOTED THIS OCCURRED IN 'LATE 2024'. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION, BUT IT WAS UNABLE TO BE OBTAINED. DETAILED PHYSICIAN ADDRESS AND FACILITY INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION, BUT IT WAS UNABLE TO BE OBTAINED. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION, BUT IT WAS UNABLE TO BE OBTAINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21636862&pc=DSQ | 2916596-2025-01625 | 2025-02-28 05:00 | Death | THORATEC CORPORATION | 2025-03-18 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTI-SYSTEM ORGAN FAILURE. THE PATIENT SUFFERED AND EXPIRED FROM A PROGRESSIVE MULTISYSTEM ORGAN FAILURE (MOF). MOF WAS NOT RELATED TO THE LVAD THERAPY. THE DEVICE WAS RUNNING AS EXPECTED. NO ISSUES WERE MENTIONED. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED . Manufacturer Narrative: A4 PATIENT WEIGHT NOT PROVIDED BY CUSTOMER. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEART MATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. THE IFU LISTS ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21637031&pc=NPU | 2015691-2025-01785 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-18 04:00 | NPU | COMMANDER DELIVERY SYSTEM UNKNOWN | Adverse Event Without Identified Device or Use Problem | Cardiac Perforation | P140031 | E2016006 | 2 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR MITRAL DEATH EVENTS FOR THE SAPIEN 3 VALVE. THIS REPORT SUMMARIZES 2 MITRAL CARDIAC PERFORATION DEATH EVENT(S) FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 78 - 80 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 FEMALE(S) AND 1 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 2 CARDIAC PERFORATION DEATH EVENT(S) FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE IN THE MITRAL POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 0 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS COMMANDER DELIVERY SYSTEM ARE: (B)(4), (B)(4), (B)(4), (B)(4). PER THE INSTRUCTIONS FOR USE, CARDIOVASCULAR INJURIES SUCH AS PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, MYOCARDIAL OR VALVULAR STRUCTURES ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE AND MAY REQUIRE INTERVENTION. THERE ARE SEVERAL POTENTIAL ETIOLOGIES FOR CARDIAC PERFORATION DURING A TRANSCATHETER VALVE REPLACEMENT PROCEDURE INCLUDING PERFORATION BY THE GUIDEWIRE, THE DELIVERY SYSTEM, OR THE TRANSVENOUS PACER (TVP) LEAD. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE TRANSCATHETER HEART VALVES. TRAINING INCLUDES PROPER GUIDEWIRE POSITIONING, FIXATION OF THE TVP TO PREVENT VENTRICLE PERFORATION, AND CAREFUL MANIPULATION OF DEVICES. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21637032&pc=DYB | 2015691-2025-01784 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-18 04:00 | DYB | EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | Adverse Event Without Identified Device or Use Problem | Hematoma | K200258 | E2016006 | 1 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR MITRAL DEATH EVENTS FOR THE SAPIEN 3 VALVE. THIS REPORT SUMMARIZES 1 MITRAL BLEEDING - HEMATOMA AT ACCESS SITE DEATH EVENT(S) FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 82 - 82 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 0 FEMALE(S) AND 1 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 1 BLEEDING - HEMATOMA AT ACCESS SITE DEATH EVENT(S) FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE IN THE MITRAL POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 9 DAY(S). DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). THE INSTRUCTIONS FOR USE (IFU) LIST HEMORRHAGE REQUIRING TRANSFUSION OR INTERVENTION AS A POTENTIAL RISK ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING POTENTIAL ACCESS COMPLICATIONS ASSOCIATED WITH STANDARD CARDIAC CATHETERIZATION. THE MAJORITY OF TRANSFEMORAL ACCESS RELATED BLEEDING COMPLICATIONS ARE RELATED TO DISEASED ILEO-FEMORAL VESSELS AND/OR PROCEDURAL TECHNIQUE DURING THE INSERTION OR REMOVAL OF THE SHEATH AND/OR DILATORS. THERE ARE CASES WHERE THE BLEEDING IS SIGNIFICANT AND MORE COMPLEX INTERVENTION OR TRANSFUSION IS REQUIRED TO TREAT/PREVENT A PERMANENT INJURY FROM OCCURRING. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE VALVE CAN BE DELIVERED TRANSFEMORALLY. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. HOWEVER, A TRANSFUSION OF BLOOD WAS REPORTED IN THE REGISTRY. BY VARC DEFINITION, THE MEET CRITERIA FOR MAJOR VASCULAR COMPLICATION THEREFORE, THE EVENT IS BEING REPORTED. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21637033&pc=NPT | 2015691-2025-01779 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-18 04:00 | NPT | COMMANDER DELIVERY SYSTEM UNKNOWN | Adverse Event Without Identified Device or Use Problem |  | P140031 | E2016006 | 1 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 VALVE. THIS REPORT SUMMARIZES 1 AORTIC DISSECTION DEATH EVENT(S) FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 68 - 68 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 0 FEMALE(S) AND 1 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER: E2016006 AND SUMMARIZES 1 AORTIC DISSECTION DEATH EVENT(S) FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 0 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS COMMANDER DELIVERY SYSTEM ARE: (B)(4). PER THE INSTRUCTIONS FOR USE (IFU) CARDIOVASCULAR INJURY, SUCH AS PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, MYOCARDIUM, OR VALVULAR STRUCTURES, IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT PROCEDURE. ASCENDING AORTIC DISSECTION MAY OCCUR WHEN MULTIPLE ATTEMPTS ARE MADE TO CROSS THE STENOTIC NATIVE VALVE, AND/OR WHEN EXCESSIVE FORCE IS USED. AORTIC DISSECTION IN A DISTAL SECTION OF THE AORTA CAN OCCUR WITH DEVICE MANIPULATION AND NAVIGATION THROUGH ANATOMIC TORTUOSITY. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE TRANSCATHETER HEART VALVE (THV). THE THV TRAINING MANUALS PROVIDE GUIDANCE TO FACILITATE SAFE CROSSING OF THE NATIVE VALVE, INCLUDING CAMERA PROJECTIONS, HANDLING DURING ADVANCEMENT, AND TROUBLESHOOTING TECHNIQUES IF DIFFICULTY IS ENCOUNTERED. AS STATED, EXCESSIVE FORCE SHOULD NOT BE USED WHEN THE DEVICE HAS DIFFICULTY CROSSING THE STENOTIC VALVE. ADDING TENSION TO THE WIRE, PULLING BACK THE SYSTEM TO RE-ORIENT THE VALVE, AS NEEDED, AND TORQUEING OF THE FLEX CATHETER MAY BE HELPFUL IN SOLVING THE PROBLEM. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21637034&pc=DYB | 2015691-2025-01774 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-18 04:00 | DYB | EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | Adverse Event Without Identified Device or Use Problem | Unspecified Vascular Problem | K200258 | E2016006 | 3 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 VALVE. THIS REPORT SUMMARIZES 3 AORTIC VASCULAR COMPLICATION - MAJOR DEATH EVENT(S) FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 78 - 82 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 0 FEMALE(S) AND 3 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER: E2016006 AND SUMMARIZES 3 VASCULAR COMPLICATION - MAJOR DEATH EVENT(S) FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 31.33 DAY(S). DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). ACCORDING TO THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR COMPLICATIONS, INCLUDING PERFORATION OR DISSECTION OF VESSELS WHICH MAY REQUIRE INTERVENTION, ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE TRANSFEMORAL TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, VASCULAR COMPLICATIONS ARE A WELL-RECOGNIZED COMPLICATION OF THE TRANSFEMORAL TRANSCATHETER VALVE REPLACEMENT PROCEDURE IN THIS ELDERLY POPULATION WITH MULTIPLE CO-MORBIDITIES. EDWARDS HAS REVIEWED MANY REPORTS, INCLUDING SCREENING DATA RECORDS AND SOURCE DOCUMENTATION OF VASCULAR COMPLICATIONS, AND HAS FOUND THAT THE ROOT CAUSE IS TYPICALLY RELATED TO A COMBINATION OF VESSEL SIZE, TORTUOSITY, AND CALCIFICATIONS. ALTHOUGH THE INCIDENCE IS DECREASING WITH SMALLER SHEATH/DELIVERY SYSTEM SIZES AND PHYSICIAN EXPERIENCE, THERE WILL CONTINUE TO BE CASES IN WHICH VASCULAR COMPLICATIONS WILL OCCUR. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE TRAINING MANUAL INSTRUCTS THE OPERATOR ON PROPER SHEATH INSERTION AND WITHDRAWAL TECHNIQUES, INCLUDING PRE-DILATING THE VESSEL WITH THE EDWARDS DILATORS, AS NEEDED. IT ALSO NOTES THAT CALCIFICATION MAY REDUCE LUMEN DIAMETER AND LIMIT OR PREVENT TRANSFEMORAL PASSAGE OF THE DEVICES. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE THV CAN BE DELIVERED TRANSFEMORALLY. ASSESSMENT OF LOCATION AND AMOUNT OF CIRCUMFERENTIAL CALCIUM WILL AID IN DETERMINING AREAS OF REDUCED VESSEL DIAMETERS. THE OPERATORS ARE TRAINED TO MEASURE MINIMUM VESSEL DIAMETER TAKING CALCIUM INTO ACCOUNT. THE PHYSICIAN TRAINING MANUAL ALSO LISTS THE MINIMUM RECOMMENDED VESSEL SIZE FOR EACH SIZE DEVICE. DESPITE THE BEST SCREENING TOOLS, A SMALL PERCENTAGE OF PATIENTS WILL HAVE FEMORAL/ILIAC VESSELS THAT ARE NOT AMENABLE TO THE TRANS-FEMORAL APPROACH OR WHERE INCREASED RESISTANCE IS ENCOUNTERED DURING INSERTION OF DEVICES. IN MANY CASES, THE VESSEL MINIMUM LUMINAL DIAMETER (MLD) MAY BE BORDERLINE OR BELOW THE INDICATED SIZE. IN ADDITION, SIGNIFICANT CALCIFICATION AND/OR TORTUOSITY, NOT ALWAYS APPRECIABLE ON IMAGING, COULD BE CONTRIBUTING FACTORS TO THE EVENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21637036&pc=DYB | 2015691-2025-01777 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-18 04:00 | DYB | EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | Adverse Event Without Identified Device or Use Problem | Unspecified Vascular Problem | K200258 | E2016006 | 1 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 VALVE. THIS REPORT SUMMARIZES 1 AORTIC VASCULAR SURGERY OR INTERVENTION - UNPLANNED DEATH EVENT(S) FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 81 - 81 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 FEMALE(S) AND 0 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER: E2016006 AND SUMMARIZES 1 VASCULAR SURGERY OR INTERVENTION - UNPLANNED DEATH EVENT(S) FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 0 DAY(S). DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). ACCORDING TO THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR COMPLICATIONS, INCLUDING PERFORATION OR DISSECTION OF VESSELS WHICH MAY REQUIRE INTERVENTION, ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE TRANSFEMORAL TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, VASCULAR COMPLICATIONS ARE A WELL-RECOGNIZED COMPLICATION OF THE TRANSFEMORAL TRANSCATHETER VALVE REPLACEMENT PROCEDURE IN THIS ELDERLY POPULATION WITH MULTIPLE CO-MORBIDITIES. EDWARDS HAS REVIEWED MANY REPORTS, INCLUDING SCREENING DATA RECORDS AND SOURCE DOCUMENTATION OF VASCULAR COMPLICATIONS, AND HAS FOUND THAT THE ROOT CAUSE IS TYPICALLY RELATED TO A COMBINATION OF VESSEL SIZE, TORTUOSITY, AND CALCIFICATIONS. ALTHOUGH THE INCIDENCE IS DECREASING WITH SMALLER SHEATH/DELIVERY SYSTEM SIZES AND PHYSICIAN EXPERIENCE, THERE WILL CONTINUE TO BE CASES IN WHICH VASCULAR COMPLICATIONS WILL OCCUR. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE TRAINING MANUAL INSTRUCTS THE OPERATOR ON PROPER SHEATH INSERTION AND WITHDRAWAL TECHNIQUES, INCLUDING PRE-DILATING THE VESSEL WITH THE EDWARDS DILATORS, AS NEEDED. IT ALSO NOTES THAT CALCIFICATION MAY REDUCE LUMEN DIAMETER AND LIMIT OR PREVENT TRANSFEMORAL PASSAGE OF THE DEVICES. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE SAPIEN VALVE CAN BE DELIVERED TRANSFEMORALLY. ASSESSMENT OF LOCATION AND AMOUNT OF CIRCUMFERENTIAL CALCIUM WILL AID IN DETERMINING AREAS OF REDUCED VESSEL DIAMETERS. THE OPERATORS ARE TRAINED TO MEASURE MINIMUM VESSEL DIAMETER TAKING CALCIUM INTO ACCOUNT. THE PHYSICIAN TRAINING MANUAL ALSO LISTS THE MINIMUM RECOMMENDED VESSEL SIZE FOR EACH SIZE DEVICE. DESPITE THE BEST SCREENING TOOLS, A SMALL PERCENTAGE OF PATIENTS WILL HAVE FEMORAL/ILIAC VESSELS THAT ARE NOT AMENABLE TO THE TRANS-FEMORAL APPROACH OR WHERE INCREASED RESISTANCE IS ENCOUNTERED DURING INSERTION OF DEVICES. IN MANY CASES, THE VESSEL MINIMUM LUMINAL DIAMETER (MLD) MAY BE BORDERLINE OR BELOW THE INDICATED SIZE. IN ADDITION, SIGNIFICANT CALCIFICATION AND/OR TORTUOSITY, NOT ALWAYS APPRECIABLE ON IMAGING, COULD BE CONTRIBUTING FACTORS TO THE EVENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21637037&pc=DYB | 2015691-2025-01775 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-18 04:00 | DYB | EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | K200258 | E2016006 | 1 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 VALVE. THIS REPORT SUMMARIZES 1 AORTIC BLEEDING - ACCESS SITE DEATH EVENT(S) FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 81 - 81 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 FEMALE(S) AND 0 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 1 BLEEDING - ACCESS SITE DEATH EVENT(S) FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 0 DAY(S). DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). THE INSTRUCTIONS FOR USE (IFU) LIST HEMORRHAGE REQUIRING TRANSFUSION OR INTERVENTION AS A POTENTIAL RISK ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING POTENTIAL ACCESS COMPLICATIONS ASSOCIATED WITH STANDARD CARDIAC CATHETERIZATION. THE MAJORITY OF TRANSFEMORAL ACCESS RELATED BLEEDING COMPLICATIONS ARE RELATED TO DISEASED ILEO-FEMORAL VESSELS AND/OR PROCEDURAL TECHNIQUE DURING THE INSERTION OR REMOVAL OF THE SHEATH AND/OR DILATORS. THERE ARE CASES WHERE THE BLEEDING IS SIGNIFICANT AND MORE COMPLEX INTERVENTION OR TRANSFUSION IS REQUIRED TO TREAT/PREVENT A PERMANENT INJURY FROM OCCURRING. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE VALVE CAN BE DELIVERED TRANSFEMORALLY. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE; HOWEVER, A TRANSFUSION OF BLOOD WAS REPORTED IN THE REGISTRY. BY VARC DEFINITION, THE MEET CRITERIA FOR MAJOR VASCULAR COMPLICATION, THEREFORE THE EVENT IS BEING REPORTED. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21637038&pc=PZP | 3012977056-2025-00083 | 2025-02-20 05:00 | Death | PROCEPT BIOROBOTICS CORPORATION | 2025-03-18 04:00 | PZP | AQUABEAM ROBOTIC SYSTEM | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | DEN170024 |  | 1 | Event Description: A MALE PATIENT UNDERWENT AQUABLATION THERAPY TO TREAT SYMPTOMATIC BENIGN PROSTATIC HYPERPLASIA (BPH). PROCEPT BIOROBOTICS CORPORATION (PROCEPT) BECAME AWARE THAT AFTER AQUABLATION THERAPY AND DURING FOCAL BLADDER NECK CAUTERY, THE PATIENT CODED. THE PATIENT WAS RESUSCITATED AND TRANSPORTED TO INTERVENTIONAL RADIOLOGY WHERE UNSPECIFIED MEDICAL INTERVENTION WAS PERFORMED. IT WAS REPORTED THAT THE PATIENT PASSED AWAY THE WEEKEND AFTER HIS AQUABLATION THERAPY. THE TREATING SURGEON REPORTED THAT THE PATIENT'S CARDIAC ARREST MAY HAVE BEEN RELATED TO UNSPECIFIED PRE-EXISTING MEDICAL CONDITIONS. PROCEPT HAS MADE MULTIPLE FOLLOW-UP ATTEMPTS TO OBTAIN ADDITIONAL DETAILS REGARDING THE PATIENT'S DEATH; HOWEVER, TO DATE, NO ADDITIONAL DETAILS HAVE BEEN RECEIVED. NO MALFUNCTION OF THE AQUABEAM ROBOTIC SYSTEM WAS REPORTED. Manufacturer Narrative: THE AQUABEAM ROBOTIC SYSTEM IS A REUSABLE DEVICE; THEREFORE, IT IS STILL CURRENTLY IN POSSESSION OF THE USER FACILITY. THE INVESTIGATION OF THIS EVENT CONSISTED OF A REVIEW OF THE DEVICE HISTORY RECORD (DHR AND INSTRUCTIONS FOR USE (IFU). A REVIEW OF THE DEVICE HISTORY RECORD (DHR) FOR AB2000-B/SERIAL NUMBER (B)(6) WAS PERFORMED, WHICH CONFIRMED THAT THERE WERE NO NON-CONFORMANCES, FAILURES, DISCREPANCIES, OR MISSED STEPS DURING THE MANUFACTURING PROCESS THAT COULD BE RELATED TO THE REPORTED EVENT. THE REVIEW INDICATED THAT THE DEVICE MET ALL DESIGN AND MANUFACTURING SPECIFICATIONS WHEN RELEASED FOR DISTRIBUTION. THE AQUABEAM ROBOTIC SYSTEM INSTRUCTIONS FOR USE (IFU), IFU0101-00, REV. E, WAS REVIEWED. 3. CONTRAINDICATIONS: DO NOT USE THE AQUABEAM ROBOTIC SYSTEM IN PATIENTS WHO DO NOT MEET THE INDICATION FOR THE SYSTEM¿S INTENDED USE. THE ROOT CAUSE OF THE EVENT WAS UNABLE TO BE DETERMINED. IT WAS REPORTED THAT THE PATIENT PASSED AWAY A FEW DAYS POST AQUABLATION THERAPY. THE TREATING SURGEON CONFIRMED THAT THE PATIENT HAD UNDISCLOSED PRE-EXISTING CONDITIONS THAT POSSIBLY ATTRIBUTED TO THE PATIENT'S DEATH. THE CAUSE OF DEATH REMAINS UNKNOWN AS THE TREATING SURGEON DID NOT PROVIDE ANY FURTHER DETAILS DUE TO THE SENSITIVE NATURE OF THE EVENT DESPITE MULTIPLE ATTEMPTS TO OBTAIN INFORMATION SUCH AS CAUSE OF DEATH AND RELATIONSHIP TO THE DEVICE OR PROCEDURE. SHALL NEW INFORMATION BECOMES AVAILABLE IN THE FUTURE, THEN FURTHER INVESTIGATION WILL BE CONDUCTED. SUBMISSION OF THIS REPORT DOES NOT CONSTITUTE AN ADMISSION THAT THE MANUFACTURER'S PRODUCT CAUSED OR CONTRIBUTED TO THE EVENT. Manufacturer Narrative: ROOT CAUSE OF THE REPORTED EVENT HAS NOT YET BEEN ESTABLISHED. INVESTIGATION BY MANUFACTURER IS CURRENTLY IN PROCESS. SUBMISSION OF THIS REPORT DOES NOT CONSTITUTE AN ADMISSION THAT THE MANUFACTURER'S PRODUCT CAUSED OR CONTRIBUTED TO THE EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21637040&pc=NPT | 2015691-2025-01778 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-18 04:00 | NPT | COMMANDER DELIVERY SYSTEM UNKNOWN | Adverse Event Without Identified Device or Use Problem | Cardiac Perforation | P140031 | E2016006 | 2 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 VALVE. THIS REPORT SUMMARIZES 2 AORTIC CARDIAC PERFORATION DEATH EVENT(S) FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 68 - 80 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 FEMALE(S) AND 1 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER: E2016006 AND SUMMARIZES 2 CARDIAC PERFORATION DEATH EVENT(S) FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 0 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS COMMANDER DELIVERY SYSTEM ARE: (B)(4). PER THE INSTRUCTIONS FOR USE, CARDIOVASCULAR INJURIES SUCH AS PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, MYOCARDIAL OR VALVULAR STRUCTURES ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE AND MAY REQUIRE INTERVENTION. THERE ARE SEVERAL POTENTIAL ETIOLOGIES FOR PERFORATION DURING A TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING PERFORATION BY THE GUIDEWIRE, THE DELIVERY SYSTEM, OR THE TRANSVENOUS PACER (TVP) LEAD. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE TRANSCATHETER HEART VALVE (THV). TRAINING INCLUDES PROPER GUIDEWIRE POSITIONING, FIXATION OF THE TVP TO PREVENT VENTRICLE PERFORATION, AND CAREFUL MANIPULATION OF DEVICES. PER THE PROCEDURE DIDACTIC, PATIENTS WITH SMALL VENTRICLES ARE AT PARTICULARLY HIGH RISK FOR VENTRICULAR PERFORATION. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21637041&pc=DYB | 2015691-2025-01776 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-18 04:00 | DYB | EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | Adverse Event Without Identified Device or Use Problem | Hematoma | K200258 | E2016006 | 1 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 VALVE. THIS REPORT SUMMARIZES 1 AORTIC BLEEDING - HEMATOMA AT ACCESS SITE DEATH EVENT(S) FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 78 - 78 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 0 FEMALE(S) AND 1 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER: E2016006 AND SUMMARIZES 1 BLEEDING - HEMATOMA AT ACCESS SITE DEATH EVENT(S) FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 0 DAY(S). DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). THE INSTRUCTIONS FOR USE (IFU) LIST HEMORRHAGE REQUIRING TRANSFUSION OR INTERVENTION AS A POTENTIAL RISK ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING POTENTIAL ACCESS COMPLICATIONS ASSOCIATED WITH STANDARD CARDIAC CATHETERIZATION. THE MAJORITY OF TRANSFEMORAL ACCESS RELATED BLEEDING COMPLICATIONS ARE RELATED TO DISEASED ILEO-FEMORAL VESSELS AND/OR PROCEDURAL TECHNIQUE DURING THE INSERTION OR REMOVAL OF THE SHEATH AND/OR DILATORS. THERE ARE CASES WHERE THE BLEEDING IS SIGNIFICANT AND MORE COMPLEX INTERVENTION OR TRANSFUSION IS REQUIRED TO TREAT/PREVENT A PERMANENT INJURY FROM OCCURRING. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE VALVE CAN BE DELIVERED TRANSFEMORALLY. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. HOWEVER, A TRANSFUSION OF BLOOD WAS REPORTED IN THE REGISTRY. BY VARC DEFINITION, THE MEET CRITERIA FOR MAJOR VASCULAR COMPLICATION THEREFORE, THE EVENT IS BEING REPORTED. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21637042&pc=NPU | 2015691-2025-01781 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-18 04:00 | NPU | EDWARDS SAPIEN 3 TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P140031 | E2016006 | 1 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR MITRAL DEATH EVENTS FOR THE SAPIEN 3 VALVE. THIS REPORT SUMMARIZES 1 MITRAL DEVICE MIGRATION DEATH EVENT(S) FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 76 - 76 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 0 FEMALE(S) AND 1 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER: E2016006 AND SUMMARIZES 1 DEVICE MIGRATION DEATH EVENT(S) FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE IN THE MITRAL POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 0 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 TRANSCATHETER HEART VALVE ARE: (B)(4). PER THE INSTRUCTIONS FOR USE, VALVE MIGRATION REQUIRING INTERVENTION IS A POTENTIAL ADVERSE EVENT ASSOCIATED WITH THE USE OF TRANSCATHETER HEART VALVES. VALVE MIGRATION RESULTS WHEN FORCES ACTING ON THE THV OVERCOME THE STRENGTH OF ATTACHMENT OF THE VALVE TO THE SITE OF DEPLOYMENT. STENT VALVES ARE SUBJECTED TO ANTEGRADE EJECTION FORCES DURING SYSTOLE. INCORRECT BIOPROSTHETIC VALVE SIZING AND INCOMPLETE FRAME EXPANSION CAN CONTRIBUTE TO VALVE MIGRATION. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE THV. TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS AND PROCTORED PROCEDURES. THE CORRECT SIZING, ALIGNMENT AND POSITIONING OF THE DEVICE ARE EMPHASIZED AS KEY FACTORS TO THE PLACEMENT AND FIXATION OF THE DEVICE. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21637045&pc=DYB | 2015691-2025-01783 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-18 04:00 | DYB | EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | Adverse Event Without Identified Device or Use Problem | Unspecified Vascular Problem | K200258 | E2016006 | 2 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR MITRAL DEATH EVENTS FOR THE SAPIEN 3 VALVE. THIS REPORT SUMMARIZES 2 MITRAL VASCULAR COMPLICATION - MAJOR DEATH EVENT(S) FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 78 - 80 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 FEMALE(S) AND 1 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 2 VASCULAR COMPLICATION - MAJOR DEATH EVENT(S) FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE IN THE MITRAL POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 0 DAY(S). DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). ACCORDING TO THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR COMPLICATIONS, INCLUDING PERFORATION OR DISSECTION OF VESSELS WHICH MAY REQUIRE INTERVENTION, ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE TRANSFEMORAL TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, VASCULAR COMPLICATIONS ARE A WELL-RECOGNIZED COMPLICATION OF THE TRANSFEMORAL TRANSCATHETER VALVE REPLACEMENT PROCEDURE IN THIS ELDERLY POPULATION WITH MULTIPLE CO-MORBIDITIES. EDWARDS HAS REVIEWED MANY REPORTS, INCLUDING SCREENING DATA RECORDS AND SOURCE DOCUMENTATION OF VASCULAR COMPLICATIONS, AND HAS FOUND THAT THE ROOT CAUSE IS TYPICALLY RELATED TO A COMBINATION OF VESSEL SIZE, TORTUOSITY, AND CALCIFICATIONS. ALTHOUGH THE INCIDENCE IS DECREASING WITH SMALLER SHEATH/DELIVERY SYSTEM SIZES AND PHYSICIAN EXPERIENCE, THERE WILL CONTINUE TO BE CASES IN WHICH VASCULAR COMPLICATIONS WILL OCCUR. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE TRAINING MANUAL INSTRUCTS THE OPERATOR ON PROPER SHEATH INSERTION AND WITHDRAWAL TECHNIQUES, INCLUDING PRE-DILATING THE VESSEL WITH THE EDWARDS DILATORS, AS NEEDED. IT ALSO NOTES THAT CALCIFICATION MAY REDUCE LUMEN DIAMETER AND LIMIT OR PREVENT TRANSFEMORAL PASSAGE OF THE DEVICES. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE THV CAN BE DELIVERED TRANSFEMORALLY. ASSESSMENT OF LOCATION AND AMOUNT OF CIRCUMFERENTIAL CALCIUM WILL AID IN DETERMINING AREAS OF REDUCED VESSEL DIAMETERS. THE OPERATORS ARE TRAINED TO MEASURE MINIMUM VESSEL DIAMETER TAKING CALCIUM INTO ACCOUNT. THE PHYSICIAN TRAINING MANUAL ALSO LISTS THE MINIMUM RECOMMENDED VESSEL SIZE FOR EACH SIZE DEVICE. DESPITE THE BEST SCREENING TOOLS, A SMALL PERCENTAGE OF PATIENTS WILL HAVE FEMORAL/ILIAC VESSELS THAT ARE NOT AMENABLE TO THE TRANS-FEMORAL APPROACH OR WHERE INCREASED RESISTANCE IS ENCOUNTERED DURING INSERTION OF DEVICES. IN MANY CASES, THE VESSEL MINIMUM LUMINAL DIAMETER (MLD) MAY BE BORDERLINE OR BELOW THE INDICATED SIZE. IN ADDITION, SIGNIFICANT CALCIFICATION AND/OR TORTUOSITY, NOT ALWAYS APPRECIABLE ON IMAGING, COULD BE CONTRIBUTING FACTORS TO THE EVENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21637048&pc=NPU | 2015691-2025-01782 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-18 04:00 | NPU | EDWARDS SAPIEN 3 TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Thrombosis/Thrombus | P140031 | E2016006 | 1 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR MITRAL DEATH EVENTS FOR THE SAPIEN 3 VALVE. THIS REPORT SUMMARIZES 1 MITRAL DEVICE THROMBOSIS DEATH EVENT(S) FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 71 - 71 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 FEMALE(S) AND 0 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 1 DEVICE THROMBOSIS DEATH EVENT(S) FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE IN THE MITRAL POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 51 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 TRANSCATHETER HEART VALVE ARE: (B)(4). PER THE INSTRUCTIONS FOR USE (IFU), VALVE THROMBOSIS IS A POTENTIAL RISK ASSOCIATED WITH THE USE OF THE TRANSCATHETER HEART VALVES (THV). THROMBOSIS IS A RARE AND WELL-RECOGNIZED COMPLICATION ASSOCIATED WITH THE USE OF BIOPROSTHETIC HEART VALVES. VALVE THROMBOSIS IS THE FORMATION OF SIGNIFICANT BLOOD CLOTS FORMING ON THE VALVE. POTENTIAL PATIENT RISK FACTORS SUCH AS ATRIAL FIBRILLATION, SYSTEMIC DISEASE (E.G., SYSTEMIC LUPUS ERYTHEMATOSUS, INFLAMMATION, AND DAMAGE TO VARIOUS BODY TISSUES, INCLUDING JOINTS, SKIN, KIDNEYS, HEART, LUNGS, BLOOD VESSELS, AND BRAIN), THE VALVULAR DISEASE ITSELF, AND REDUCED CARDIAC EJECTION FRACTION CAN CONTRIBUTE TO INCREASED RISK OF THROMBUS/THROMBOSIS. SUB-OPTIMAL ANTICOAGULATION DURING THE PROCEDURE AND UNDERLYING PATIENT CONDITIONS CAN ALSO RESULT IN INCREASED THROMBOGENICITY. THESE PATIENTS ARE ANTICOAGULATED FOR THE PROCEDURE AND INTERVENTIONAL BEST PRACTICES MANDATE METICULOUS PREPARATION, ASPIRATION, AND FLUSHING OF THE DEVICES TO PREVENT AND/OR REMOVE CLOT(S). SHORT-TERM ANTICOAGULATION THERAPY MAY ALSO BE NECESSARY AFTER VALVE REPAIR, ANTICOAGULATION IS PRESCRIBED PER INSTITUTIONAL GUIDELINES. INTRA-PROCEDURAL INTRA-CARDIAC THROMBUS AND POST PROCEDURE OR LATE VALVE THROMBOSIS ARE COMPLEX PROCESSES TRIGGERED BY THE INTERACTION BETWEEN THE HOST AND THE DEVICE WHICH ARE HIGHLY VARIABLE AMONG PATIENTS. IT IS THE NATURAL TENDENCY OF THE BODY TO FORM A CLOT ON FOREIGN OBJECTS IN THE VASCULAR SPACE AND A DEFINITIVE ROOT CAUSE CANNOT ALWAYS BE CONFIRMED. THE DEVICE TRAINING MANUALS INSTRUCT THE OPERATOR TO CONSIDER ALL PROCEDURAL AND ANATOMICAL FACTORS. EDWARDS EXTENSIVELY TRAINS PHYSICIANS BEFORE THEY ARE QUALIFIED TO USE THE DEVICE SYSTEM. TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES. OPERATORS ARE ALSO INSTRUCTED TO USE FLUOROSCOPY IN CONJUNCTION WITH ECHOCARDIOGRAPHY FOR OPTIMAL VISUALIZATION DURING POSITIONING AND DEPLOYMENT AND CAUTIONS INCLUDE THE MAINTENANCE OF THE PATIENT'S ANTICOAGULATION STATUS (ACT AT >250 SECONDS) DURING THE PROCEDURE. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21637053&pc=NPU | 2015691-2025-01780 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-18 04:00 | NPU | EDWARDS SAPIEN 3 TRANSCATHETER HEART VALVE | Insufficient Information | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P140031 | E2016006 | 1 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR MITRAL DEATH EVENTS FOR THE SAPIEN 3 VALVE. THIS REPORT SUMMARIZES 1 MITRAL CARDIAC SURGERY OR INTERVENTION - OTHER UNPLANNED DEATH EVENT(S) FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 80 - 80 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 FEMALE(S) AND 0 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER: E2016006 AND SUMMARIZES 1 CARDIAC SURGERY OR INTERVENTION - OTHER UNPLANNED DEATH EVENT(S) FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE IN THE MITRAL POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 0 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 TRANSCATHETER HEART VALVE ARE: (B)(4). THE INSTRUCTIONS FOR USE (IFU) LIST REOPERATION, INCLUDING EMERGENCY CARDIAC SURGERY, AS ADDITIONAL POTENTIAL RISKS ASSOCIATED WITH THE USE OF THE THV, DELIVERY SYSTEM, AND/OR ACCESSORIES. THIS EVENT TYPE (CARDIAC SURGERY OR INTERVENTION - OTHER UNPLANNED) IS DEFINED AS: THE PATIENT SUBSEQUENTLY UNDERWENT OTHER TYPES OF CARDIAC SURGERY OR CATH LAB INTERVENTIONS (NOT OTHERWISE SPECIFIED) THAT WAS UNPLANNED. THIS COULD INCLUDE REOPERATION FOR DAMAGE TO ADJACENT STRUCTURES, OR POTENTIALLY VALVE EXPLANT RELATED TO EMBOLIZATION, OR OTHER EVENT TYPES. CARDIOVASCULAR INJURY, INCLUDING PERFORATION OR DISSECTION OF VALVULAR STRUCTURES AND VALVE EMBOLIZATION REQUIRING INTERVENTION ARE KNOWN POTENTIAL COMPLICATIONS LISTED IN THE PRODUCT INSTRUCTIONS FOR USE (IFU). PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE TRANSCATHETER HEART VALVES (THV). TRAINING INCLUDES PROPER GUIDEWIRE POSITIONING, AND CAREFUL MANIPULATION OF DEVICES. PER THE THV TA TRAINING MANUALS, AFTER GAINING LV ACCESS WITH THE 0.035'' SOFT GUIDEWIRE, THE WIRE SHOULD BE JIGGLED UNDER TEE IMAGING OF THE MITRAL VALVE. AN INCREASE IN MR SUGGESTS WIRE ENTANGLEMENT IN THE MITRAL SUB-VALVULAR APPARATUS. IF ENTANGLEMENT IS SUSPECTED, THE GUIDEWIRE SHOULD BE COMPLETELY REMOVED FROM THE VENTRICLE; THE OPERATOR SHOULD CHANGE DIRECTION OF THE NEEDLE AND REINSERT THE GUIDEWIRE INTO THE VENTRICLE, CHECKING AGAIN FOR WIRE ENTANGLEMENT. THE TF TRAINING MANUAL ALSO PROVIDES GUIDANCE FOR VALVE CROSSING AND WIRE EXCHANGE: EXCHANGE 0.035'' AMPLATZ EXTRA-STIFF WIRE WITH PRE-SHAPED DISTAL END IN LEFT VENTRICLE. ENSURE THE GUIDE CATHETER IS ADVANCED UPON WIRE EXCHANGE TO ENSURE EXCHANGE WIRE DOES NOT GET CAUGHT IN THE MITRAL CHORDAE. THERE ARE MULTIPLE PATIENT AND PROCEDURAL FACTORS THAT ALONE OR IN COMBINATION CAN CAUSE OR CONTRIBUTE TO VENTRICULAR EMBOLIZATION, INCLUDING IMPROPER POSITIONING PRIOR TO DEPLOYMENT, POOR IMAGE INTENSIFIER ANGLE, POOR COAXIAL ALIGNMENT OF THE VALVE/DELIVERY SYSTEM, RAPID DEPLOYMENT, RELEASE OF STORED TENSION DURING DEPLOYMENT, AND MOVEMENT OF THE DELIVERY SYSTEM BY THE OPERATOR. THE THV TRAINING MANUALS INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN THV. TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS AND PROCTORED PROCEDURES. THE CORRECT ALIGNMENT AND POSITIONING OF THE DEVICE AT THE POINT OF DEPLOYMENT IS EMPHASIZED AS A KEY FACTOR TO THE PLACEMENT AND FIXATION OF THE DEVICE. OPERATORS ARE ALSO INSTRUCTED TO USE FLUOROSCOPY AS THE PRIMARY METHOD OF VISUALIZATION FOR POSITIONING AND DEPLOYMENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21637759&pc=MHX | 3003832357-2025-000203 | 2025-02-14 05:00 | Death | REMOTE DIAGNOSTIC TECHNOLOGIES LTD. | 2025-03-18 04:00 | MHX | TEMPUS PRO | Incorrect Measurement | Cardiac Arrest | K201746 |  | 1 | Event Description: THIS REPORT IS BASED ON INFORMATION PROVIDED BY REMOTE SERVICE ENGINEER RSE AND IS CURRENTLY UNDER INVESTIGATION BY THE PHILIPS COMPLAINT HANDLING TEAM. PHILIPS RECEIVED A COMPLAINT ON THE TEMPUS PRO INDICATING THE DEVICE WAS DISPLAYING ASYSTOLE, YET THE CREW ON SCENE COULD FEEL A PALPABLE PULSE. THE PATIENT WENT INTO CARDIAC ARREST AND THE DEVICE DID PROVIDE SHOCKS TO THE PATIENT. Manufacturer Narrative: CORRECTION OF MFR REPORT NUMBER 3003832357-2025-000203. UPDATE TO EVENT AND DEATH DATE. ALSO, HEALTH IMPACT GRID.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21638780&pc=NPT | 2015691-2025-01793 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-18 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P140031 | E2016006 | 1 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA VALVE. THIS REPORT SUMMARIZES 1 AORTIC DEVICE MIGRATION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 79 - 79 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 0 FEMALE(S) AND 1 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 1 DEVICE MIGRATION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 0 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE ARE: (B)(4). PER THE INSTRUCTIONS FOR USE, VALVE MIGRATION REQUIRING INTERVENTION IS A POTENTIAL ADVERSE EVENT ASSOCIATED WITH TRANSCATHETER AORTIC VALVE REPLACEMENT. ACCORDING TO LITERATURE REVIEW, VALVE MIGRATION RESULTS WHEN FORCES ACTING ON THE TRANSCATHETER HEART VALVE (THV) OVERCOME THE STRENGTH OF ATTACHMENT OF THE VALVE TO THE ANNULUS. STENT VALVES ARE SUBJECTED TO ANTEGRADE EJECTION FORCES DURING SYSTOLE. LESS-THAN-SEVERE AND NON-UNIFORMLY DISTRIBUTED CALCIFICATION OF THE NATIVE LEAFLETS, INCORRECT BIOPROSTHETIC VALVE SIZING, AND INCOMPLETE FRAME EXPANSION CAN CONTRIBUTE TO VALVE MIGRATION. ADDITIONALLY, RESIDUAL OVERHANGING LEAFLETS CAN EXERT DOWNWARDS FORCE DURING DIASTOLE, CAUSING MIGRATION OF THE THV TOWARDS THE LEFT VENTRICLE. THE EDWARDS THV PATIENT SCREENING MANUAL ADVISES THE OPERATOR ON PRE-PROCEDURE ASSESSMENT OF THE AORTIC VALVE AND ROOT, TAKING INTO CONSIDERATION THE DEGREE AND DISTRIBUTION OF NATIVE LEAFLET CALCIFICATION. THE PROCEDURAL DIDACTIC INSTRUCTS THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE TRANSCATHETER HEART VALVE (THV). TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS AND PROCTORED PROCEDURES. THE CORRECT SIZING, ALIGNMENT AND POSITIONING OF THE DEVICE ARE EMPHASIZED AS KEY FACTORS TO THE PLACEMENT AND FIXATION OF THE DEVICE. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21638781&pc=NPT | 2015691-2025-01792 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-18 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem |  | P140031 | E2016006 | 3 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA VALVE. THIS REPORT SUMMARIZES 3 AORTIC CORONARY ARTERY COMPRESSION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 70 - 86 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 2 FEMALE(S) AND 1 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 3 CORONARY ARTERY COMPRESSION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 0 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE ARE: (B)(4). PER THE INSTRUCTIONS FOR USE (IFU), CORONARY FLOW OBSTRUCTION IS A POTENTIAL ADVERSE EVENT ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT PROCEDURE. THE IFU CAUTIONS THAT THE SAFETY AND EFFECTIVENESS HAVE NOT BEEN ESTABLISHED FOR PATIENTS WITH BULKY CALCIFIED AORTIC VALVE LEAFLETS IN CLOSE PROXIMITY TO CORONARY OSTIA. IN ADDITION, IT WARNS THAT CAUTION SHOULD BE EXERCISED IN IMPLANTING A BIOPROSTHESIS IN PATIENTS WITH CLINICALLY SIGNIFICANT CORONARY ARTERY DISEASE. CORONARY OCCLUSION CAN RESULT IN MYOCARDIAL ISCHEMIA OR INFARCTION DUE TO OBSTRUCTION OF THE CORONARY BLOOD FLOW AND MAY REQUIRE INTERVENTION (E.G., PCI). MULTIPLE PATIENT FACTORS COULD PUT THE PATIENT AT RISK FOR CORONARY FLOW OBSTRUCTION DURING THE THV PROCEDURE, INCLUDING SIGNIFICANT UNDERLYING CORONARY ARTERY DISEASE AND BULKY CALCIFICATION OF THE NATIVE LEAFLETS AND ROOT. DISPLACEMENT OF CALCIUM DEPOSITS WITH EMBOLIZATION OF DEBRIS INTO ONE OF THE ARTERIES, OR AORTIC DISSECTION WITH CONTINUITY OF THE RUPTURE INTO THE INTIMA OF ONE OF THE CORONARY OSTIA, CAN RESULT IN THIS COMPLICATION. ADDITIONALLY, MINIMAL DISTANCE BETWEEN THE NATIVE ANNULUS AND THE CORONARY OSTIA, BULKY CALCIFICATION, LONG NATIVE LEAFLETS, AND OBLITERATED CORONARY SINUSES. PROCEDURAL FACTORS SUCH AS TORN NATIVE LEAFLET DURING BALLOON VALVULOPLASTY, PLAQUE SHIFT, DEPLOYMENT OF THE BIOPROSTHETIC HEART VALVE TOO AORTIC, AND SIGNIFICANT VALVE OVER SIZING COULD ALSO CONTRIBUTE TO THIS COMPLICATION. THE EDWARDS THV TRAINING MANUALS ADVISE THE OPERATOR ON PRE-PROCEDURE ASSESSMENT OF THE AORTIC VALVE, ROOT, AND CORONARY ANATOMY. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE TRANSCATHETER HEART VALVE (THV). TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS AND PROCTORED PROCEDURES. THE PHYSICIAN IS INSTRUCTED TO EVALUATE THIS RISK EARLY IN THE PATIENT SCREENING PROCESS IN ALL PATIENTS. THE FOLLOWING FACTORS SHOULD BE CONSIDERED: DEGREE OF CALCIFICATION ON LEAFLETS, ANNULUS TO CORONARY OSTIA DISTANCE, LENGTH OF THE VALVE LEAFLET, WIDTH OF THE VALSALVA SINUSES, MOVEMENT OF THE LEAFLETS DURING BAV, PATENCY OF CORONARIES DURING BAV, AND EXPANDED HEIGHT OF THE INTENDED THV. THE TRAINING MANUALS ALSO PROVIDE THE FOLLOWING TIPS FOR DETECTING RISK FOR LEFT MAIN OCCLUSION: (1) AORTOGRAM OR TEE PRIOR TO THV IMPLANTATION TO REVEAL BULKY CALCIFIED LEAFLETS; (2) DURING PRE-DILATATION, NOTE BULKY CALCIFICATION ON VALVE MOVING TOWARDS OSTIUM ON LEFT MAIN; AND (3) CONSIDER AORTOGRAM DURING VALVULOPLASTY TO ASSESS CORONARY FLOW. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21638782&pc=NPT | 2015691-2025-01789 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-18 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Ischemia Stroke | P140031 | E2016006 | 6 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA VALVE. THIS REPORT SUMMARIZES 6 AORTIC STROKE - ISCHEMIC DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 80 - 91 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 3 FEMALE(S) AND 3 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 6 STROKE - ISCHEMIC DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 3.33 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE ARE: (B)(4). PER THE INSTRUCTIONS FOR USE, PERMANENT OR TRANSIENT NEUROLOGICAL EVENTS INCLUDING STROKE ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH INVASIVE CARDIAC PROCEDURE INCLUDING TRANSCATHETER HEART VALVE PROCEDURES AND THE USE OF THE EDWARDS THV DEVICES. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, STROKE IS RECOGNIZED IN THE LITERATURE AS A WELL-KNOWN COMPLICATION IN A SMALL NUMBER OF PATIENTS UNDERGOING TRANSCATHETER VALVE REPLACEMENT PROCEDURE. RISK FACTORS CORRELATING WITH A NUMBER OF PATIENT CO-MORBIDITIES HAVE BEEN IDENTIFIED. ALTHOUGH IN MANY CASES THE ROOT CAUSE OF THE EVENT IS UNABLE TO BE DETERMINED, STROKES DURING TRANSCATHETER VALVE REPLACEMENT PROCEDURE ARE UNDOUBTEDLY MULTIFACTORIAL, THE DOMINANT ETIOLOGY LIKELY BEING INTRA-PROCEDURE EMBOLIC EVENTS. A TRANSCRANIAL DOPPLER STUDY DURING TRANSCATHETER VALVE REPLACEMENT PROCEDURE DEMONSTRATED THAT THE MAJORITY OF PROCEDURAL EMBOLIC EVENTS OCCURRED DURING BALLOON VALVULOPLASTY, MANIPULATION OF CATHETERS ACROSS THE AORTIC VALVE, AND VALVE IMPLANTATION. AN ANALYSIS IN PATIENTS UNDERGOING VALVE SURGERY REVEALED FOUR BASELINE CHARACTERISTICS AND TWO PROCEDURAL EVENTS THAT WERE ASSOCIATED WITH EARLY POST-PROCEDURE STROKE: FEMALE SEX, EF < 30%, DIABETES, AGE OLDER THAN 70 YEARS, BYPASS PROCEDURE TIME> 120 MIN, AND CALCIFICATION OF THE ASCENDING AORTA. PREDICTORS OF LATE STROKE HAVE INCLUDED FEMALE SEX, AGE OLDER THAN 75 YEARS, ATRIAL FIBRILLATION, AND A HISTORY OF OR CURRENT SMOKING. THERE WERE NO IMPORTANT DIFFERENCES IN THE FREQUENCY OF LATE STROKES BETWEEN TRANSCATHETER VALVE REPLACEMENT PROCEDURE AND AVR PATIENTS. AFTER TRANSCATHETER VALVE REPLACEMENT PROCEDURE, THERE APPEARS TO BE A MORE SIGNIFICANT PROPORTION OF EARLY STROKES OCCURRING < 24 H POST-PROCEDURE, BUT TRANSCATHETER VALVE REPLACEMENT PROCEDURE PATIENTS WITH MULTIPLE CO MORBIDITIES ARE PROBABLY AT HIGHER RISK OF BOTH EARLY AND LATE STROKES. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE HOWEVER, THE EVENT MAY BE RELATED TO THE MECHANISM DESCRIBED ABOVE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21638783&pc=NPT | 2015691-2025-01794 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-18 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Myocardial Infarction | P140031 | E2016006 | 1 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA VALVE. THIS REPORT SUMMARIZES 1 AORTIC MYOCARDIAL INFARCTION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 78 - 78 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 FEMALE(S) AND 0 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 1 MYOCARDIAL INFARCTION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 0 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE ARE: (B)(4). PER THE INSTRUCTIONS FOR USE (IFU), THROMBUS FORMATION, PLAQUE DISLODGEMENT, AND EMBOLIZATION THAT MAY RESULT IN MYOCARDIAL INFARCTION (MI) ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE. THE IFU CAUTIONS THAT THE SAFETY AND EFFECTIVENESS HAVE NOT BEEN ESTABLISHED FOR PATIENTS WITH BULKY CALCIFIED AORTIC VALVE LEAFLETS IN CLOSE PROXIMITY TO THE CORONARY OSTIA. IN ADDITION, CAUTION SHOULD BE EXERCISED WHEN IMPLANTING A BIOPROSTHESIS IN PATIENTS WITH CLINICALLY SIGNIFICANT CORONARY ARTERY DISEASE. MI RELATED TO THE TRANSCATHETER VALVE REPLACEMENT PROCEDURE AND THE VALVE IMPLANT WILL MANIFEST INTRA-PROCEDURALLY OR IN THE IMMEDIATE POST-OPERATIVE PERIOD AND IS TYPICALLY DUE TO A COMBINATION OF PATIENT AND/OR PROCEDURAL FACTORS. AS DEFINED IN THE VALVE ACADEMIC RESEARCH CONSORTIUM (VARC) PUBLICATION ON TRANSCATHETER VALVE REPLACEMENT COMPLICATIONS, THE PERI-PROCEDURAL INTERVAL IS INCLUSIVE OF ALL EVENTS THAT BEGIN WITHIN 72 HRS. OF THE INDEX PROCEDURE. MI'S THAT OCCUR AFTER THE PERI-PROCEDURAL PERIOD (>72 HRS.) ARE TYPICALLY RELATED TO THE PATIENT'S UNDERLYING CORONARY DISEASE. THERE ARE MULTIPLE PATIENT FACTORS THAT COULD PUT THE PATIENT AT RISK FOR MI DURING THE TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING SIGNIFICANT UNDERLYING CORONARY ARTERY DISEASE, AND BULKY CALCIFICATION OF THE NATIVE LEAFLETS AND ROOT. DISPLACEMENT OF CALCIUM DEPOSITS WITH EMBOLIZATION OF DEBRIS INTO ONE OF THE ARTERIES, OR AORTIC DISSECTION WITH CONTINUITY OF THE RUPTURE INTO THE INTIMA OF ONE OF THE CORONARY OSTIA, CAN RESULT IN THIS COMPLICATION. THE EDWARDS THV MANUALS ADVISE THE OPERATOR ON PRE-PROCEDURE ASSESSMENT OF THE AORTIC VALVE, ROOT, AND CORONARY ANATOMY. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE TRANSCATHETER HEART VALVE (THV). TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS AND PROCTORED PROCEDURES. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21638784&pc=NPT | 2015691-2025-01787 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-18 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Rupture | P140031 | E2016006 | 9 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA VALVE. THIS REPORT SUMMARIZES 9 AORTIC ANNULAR RUPTURE DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 75 - (B)(6) YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 7 FEMALE(S) AND 2 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 9 ANNULAR RUPTURE DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 0 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE ARE: (B)(4). PER THE INSTRUCTIONS FOR USE (IFU) CARDIOVASCULAR INJURIES, SUCH AS PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, MYOCARDIAL OR VALVULAR STRUCTURES, AND HEMATOMA ARE A KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT PROCEDURE. ACCORDING TO THE THV TRAINING MANUALS, RISK FACTORS FOR AORTIC DISSECTION, HEMATOMA, OR ANNULAR RUPTURE DURING THE TRANSCATHETER VALVE REPLACEMENT PROCEDURE INCLUDE SIGNIFICANT THV OVERSIZING, SEVERELY OBLITERATED SINUSES OF VALSALVA, PORCELAIN AORTA AND/OR PRESENCE OF BULKY CALCIFICATION AND NARROW CALCIFIED STJ. IN ADDITION, ADVANCED AGE, FEMALE GENDER, SMALL BODY WEIGHT, AND STEROID DEPENDENCY CAN ALSO BE CONTRIBUTING FACTORS. THE SAPIEN VALVE RELIES ON NATIVE VALVE CALCIUM TO SECURELY ANCHOR TO THE ANNULUS. DESPITE THIS BENEFICIAL ASPECT OF CALCIUM, BULKY CALCIUM CAN INCREASE THE RISK OF CALCIFIC NODULE DISPLACEMENT INTO THE VASCULATURE, WHICH CAN LEAD TO VASCULAR INJURY. AT TIMES THE EXTENT AND DISTRIBUTION OF CALCIUM CAN IMPAIR EASE OF DELIVERY OF THE VALVE, CORRECT POSITIONING OF THE VALVE, DEPLOYMENT OF THE VALVE AND PROCEDURAL SUCCESS. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21638786&pc=NPT | 2015691-2025-01791 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-18 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Insufficient Information | Aortic Valve Insufficiency/ Regurgitation | P140031 | E2016006 | 5 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA VALVE. THIS REPORT SUMMARIZES 5 AORTIC REINTERVENTION - AORTIC VALVE DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 55 - 72 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 FEMALE(S) AND 4 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 5 REINTERVENTION - AORTIC VALVE DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 99.8 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE ARE: (B)(4). INTRA-PROCEDURAL AORTIC VALVE RE-INTERVENTION WILL TYPICALLY RESULT FROM VALVE MALPOSITION OR REGURGITATION (PVL OR CENTRAL, INCLUDING LEAFLET RESTRICTION). THESE CONDITIONS ARE KNOWN POTENTIAL RISKS ASSOCIATED WITH THE USE OF THE THV, DELIVERY SYSTEM, AND/OR ACCESSORIES. THERE ARE MULTIPLE PATIENT AND PROCEDURAL FACTORS THAT ALONE OR IN COMBINATION CAN CAUSE OR CONTRIBUTE TO AORTIC MALPOSITION, INCLUDING IMPROPER POSITIONING PRIOR TO DEPLOYMENT, POOR IMAGE INTENSIFIER ANGLE, POOR COAXIAL ALIGNMENT OF THE VALVE AND DELIVERY SYSTEM, SEVERE SEPTAL HYPERTROPHY, MINIMALLY OR BULKY/SEVERELY CALCIFIED AORTIC LEAFLETS, PRESERVED EJECTION FRACTION, SIGNIFICANT MITRAL ANNULAR CALCIFICATION (MAC), LOSS OF PACING CAPTURE, RAPID DEPLOYMENT, RELEASE OF STORED TENSION DURING DEPLOYMENT, AND MOVEMENT OF THE DELIVERY SYSTEM BY THE OPERATOR. DEPLOYMENT OF THE SAPIEN VALVE TOO AORTIC HAS THE POTENTIAL TO CONTRIBUTE TO SUBOPTIMAL COAPTATION OF THE SAPIEN VALVE LEAFLETS AND CAUSE CENTRAL AORTIC INSUFFICIENCY; IT CAN OBSTRUCT THE CORONARY OSTIA; AND LEAD TO EMBOLIZATION OF THE PROSTHESIS INTO THE ASCENDING AORTA. THE THV TRAINING MANUALS INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE THV. TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS AND PROCTORED PROCEDURES. THE CORRECT ALIGNMENT AND POSITIONING OF THE DEVICE AT THE POINT OF DEPLOYMENT IS EMPHASIZED AS A KEY FACTOR TO THE PLACEMENT AND FIXATION OF THE DEVICE. OPERATORS ARE ALSO INSTRUCTED TO USE FLUOROSCOPY AS THE PRIMARY METHOD OF VISUALIZATION FOR POSITIONING AND DEPLOYMENT. IN PATIENTS WITH HIGH-RISK ANATOMICAL FEATURES FOR AORTIC MALPOSITION (I.E., MINIMAL LEAFLET CALCIFICATION, SEVERE SEPTAL HYPERTROPHY), BAV MAY PROVIDE INDICATION OF POTENTIAL BALLOON MOVEMENT DURING VALVE DEPLOYMENT. AORTIC VALVE RE-INTERVENTION IN THE FOLLOW-UP PERIOD (IN THE ABSENCE OF PROSTHETIC CARDIAC VALVE THROMBOSIS) WILL TYPICALLY RESULT FROM ON-GOING OR WORSENING REGURGITATION, VALVE DEGENERATION RELATED TO THE FORMATION OF CALCIFICATION OR PANNUS, OR VALVE MIGRATION. THESE EVENTS ARE IDENTIFIED IN THE PRODUCT INSTRUCTIONS FOR USE (IFU) AS POTENTIAL RISKS ASSOCIATED WITH THE USE OF THE THV. REGURGITATION WHICH DEVELOPS PROGRESSIVELY OVER TIME CAN BE DUE TO A NUMBER OF ISSUES INCLUDING PATIENT RELATED FACTORS OR STRUCTURAL VALVE DETERIORATION, INCLUDING CALCIFICATION, NON-CALCIFIC DEGENERATION, LEAFLET THICKENING OR FIBROSIS, OR A COMBINATION OF THESE. REGURGITATION MAY ALSO DEVELOP PROGRESSIVELY IF HOST FIBROTIC TISSUE, OR PANNUS, GROWS ONTO THE BIOPROSTHETIC VALVE. PANNUS, A CAUSE OF NONSTRUCTURAL DYSFUNCTION, MAY INTERFERE WITH FUNCTIONALITY OF THE DEVICE BY RESTRICTING THE LEAFLET MOTION LEADING TO ABNORMAL COAPTATION. CALCIFICATION IS A WELL-RECOGNIZED FAILURE MODE OF BIOPROSTHETIC VALVES. THE MECHANISM OF CALCIFICATION OF BIOMATERIALS IS NOT COMPLETELY UNDERSTOOD BUT IS PROBABLY RELATED TO AN INABILITY OF THE NON-VIABLE CELLS TO MAINTAIN THEIR NORMALLY LOW INTRACELLULAR CONCENTRATION OF CALCIUM. THE MECHANISM BEHIND WORSENING OR LATE PVL IS NOT WELL UNDERSTOOD BUT MAY BE RELATED TO CARDIAC REMODELING. ACCORDING TO LITERATURE REVIEW, VALVE MIGRATION RESULTS WHEN FORCES ACTING ON THE TRANSCATHETER HEART VALVE (THV) OVERCOME THE STRENGTH OF ATTACHMENT OF THE VALVE TO THE AORTIC WALL. STENT VALVES ARE SUBJECTED TO ANTEGRADE EJECTION FORCES DURING SYSTOLE. LESS-THAN-SEVERE AND NON-UNIFORMLY DISTRIBUTED CALCIFICATION OF THE NATIVE LEAFLETS, INCORRECT BIOPROSTHETIC VALVE SIZING, AND INCOMPLETE FRAME EXPANSION CAN CONTRIBUTE TO VALVE MIGRATION. ADDITIONALLY, RESIDUAL OVERHANGING LEAFLETS CAN EXERT DOWNWARDS FORCE DURING DIASTOLE, CAUSING MIGRATION OF THE THV TOWARDS THE LEFT VENTRICLE. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21638793&pc=NPT | 2015691-2025-01790 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-18 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Heart Block | P140031 | E2016006 | 5 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA VALVE. THIS REPORT SUMMARIZES 5 AORTIC PERMANENT PACEMAKER DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 77 - (B)(6) YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 3 FEMALE(S) AND 2 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 5 PERMANENT PACEMAKER DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 14.4 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE ARE: (B)(4). PER THE INSTRUCTIONS FOR USE, CONDUCTION SYSTEM DEFECTS (HEART BLOCK) ARRHYTHMIAS AND CONDUCTION SYSTEM DEFECTS THAT MAY REQUIRE A PERMANENT PACEMAKER ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH BALLOON AORTIC VALVULOPLASTY, DEPLOYMENT OF THE PROSTHETIC VALVE, AND THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE. ACCORDING TO THE VALVE ACADEMIC RESEARCH CONSORTIUM (VARC) GUIDELINES, THE CLOSE ANATOMICAL RELATIONSHIP BETWEEN THE AORTIC VALVE COMPLEX AND THE BRANCHING ATRIOVENTRICULAR BUNDLE MAY PROVIDE AN EXPLANATION FOR THESE COMPLICATIONS OF THE TRANSCATHETER VALVE REPLACEMENT PROCEDURE. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY -CONDUCTION DISTURBANCES/ HEART BLOCK WRITTEN BY EDWARDS LIFESCIENCES, ATRIOVENTRICULAR CONDUCTION DISTURBANCES AFTER THV PLACEMENT ARE ASSOCIATED WITH MANY PATIENT RELATED AND PROCEDURAL RELATED FACTORS, INCLUDING PRE-OPERATIVE CO-MORBID STATUS, THE DEGREE AND BULKINESS OF AORTIC VALVE AND ANNULAR CALCIFICATION, INTER-VENTRICULAR SEPTAL THICKNESS, PRE-EXISTING ELECTROCARDIOGRAM ABNORMALITIES, THE DEPTH OF PROSTHESIS IMPLANTATION, AND THE PROFILE OF THE IMPLANTED PROSTHESIS. UNLIKE CONVENTIONAL AVR, WHERE THERE MAY BE LOCALIZED TRAUMA DUE TO DECALCIFICATION OF THE ANNULUS AND/OR SUTURE PLACEMENT IN THE PROXIMITY OF THE AV NODE OR THE BUNDLES, TRANSCATHETER VALVE REPLACEMENT MAY CAUSE CONDUCTION ABNORMALITIES THROUGH MECHANICAL IMPINGEMENT OF THE CONDUCTION SYSTEM BY THE PROSTHESIS. THE MECHANISMS OF THE DEVELOPMENT OF HEART BLOCK AFTER THV PLACEMENT ARE WELL DOCUMENTED AND DESCRIBED IN THE LITERATURE. IT IS ALSO DOCUMENTED THAT PRE-EXISTING HEART BLOCK IS COMMON IN PATIENTS UNDERGOING THV PLACEMENT OR SURGICAL AVR AND ANOTHER 4-6 % WILL DEVELOP POSTOPERATIVE HEART BLOCK, POTENTIALLY REQUIRING A PERMANENT PACEMAKER. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE, HOWEVER THE CONDUCTION DISTURBANCE IS LIKELY RELATED TO THE MECHANISM DESCRIBED ABOVE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21638795&pc=NPT | 2015691-2025-01795 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-18 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P140031 | E2016006 | 1 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA VALVE. THIS REPORT SUMMARIZES 1 AORTIC STROKE - UNDETERMINED DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 81 - 81 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 FEMALE(S) AND 0 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 1 STROKE - UNDETERMINED DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 1 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE ARE: (B)(4). PER THE INSTRUCTIONS FOR USE, PERMANENT OR TRANSIENT NEUROLOGICAL EVENTS INCLUDING STROKE ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH INVASIVE CARDIAC PROCEDURE INCLUDING TRANSCATHETER HEART VALVE PROCEDURES AND THE USE OF THE EDWARDS THV DEVICES. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, STROKE IS RECOGNIZED IN THE LITERATURE AS A WELL-KNOWN COMPLICATION IN A SMALL NUMBER OF PATIENTS UNDERGOING TRANSCATHETER VALVE REPLACEMENT PROCEDURE. RISK FACTORS CORRELATING WITH SEVERAL PATIENT CO-MORBIDITIES HAVE BEEN IDENTIFIED. ALTHOUGH IN MANY CASES THE ROOT CAUSE OF THE EVENT IS UNABLE TO BE DETERMINED, STROKES DURING TRANSCATHETER VALVE REPLACEMENT PROCEDURE ARE UNDOUBTEDLY MULTIFACTORIAL, THE DOMINANT ETIOLOGY LIKELY BEING INTRA-PROCEDURE EMBOLIC EVENTS. A TRANSCRANIAL DOPPLER STUDY DURING TRANSCATHETER VALVE REPLACEMENT DEMONSTRATED THAT THE MAJORITY OF PROCEDURAL EMBOLIC EVENTS OCCURRED DURING BALLOON VALVULOPLASTY, MANIPULATION OF CATHETERS ACROSS THE AORTIC VALVE, AND VALVE IMPLANTATION. AN ANALYSIS IN PATIENTS UNDERGOING VALVE SURGERY REVEALED FOUR BASELINE CHARACTERISTICS AND TWO PROCEDURAL EVENTS THAT WERE ASSOCIATED WITH EARLY POST-PROCEDURE STROKE: FEMALE SEX, EF < 30%, DIABETES, AGE OLDER THAN 70 YEARS, BYPASS PROCEDURE TIME> 120 MIN, AND CALCIFICATION OF THE ASCENDING AORTA. PREDICTORS OF LATE STROKE HAVE INCLUDED FEMALE SEX, AGE OLDER THAN 75 YEARS, ATRIAL FIBRILLATION, AND A HISTORY OF OR CURRENT SMOKING. THERE WERE NO IMPORTANT DIFFERENCES IN THE FREQUENCY OF LATE STROKES BETWEEN TRANSCATHETER VALVE REPLACEMENT AND AVR PATIENTS. AFTER TAVR, THERE APPEARS TO BE A MORE SIGNIFICANT PROPORTION OF EARLY STROKES OCCURRING < 24 H POST-PROCEDURE, BUT TRANSCATHETER VALVE REPLACEMENT PROCEDURE PATIENTS WITH MULTIPLE CO MORBIDITIES ARE PROBABLY AT HIGHER RISK OF BOTH EARLY AND LATE STROKES. IN THIS CASE, THE EVENT IS LISTED AS UNDETERMINED STROKE. IT IS NOT KNOWN IF THE MECHANISM OF THE STROKE IS IN QUESTION, OR IF THE STROKE ITSELF IS IN QUESTION. THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. IT IS POSSIBLE THAT THE MECHANISM DESCRIBED ABOVE MAY HAVE CONTRIBUTED TO THE EVENT. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21638796&pc=NPT | 2015691-2025-01786 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-18 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Insufficient Information | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P140031 | E2016006 | 14 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA VALVE. THIS REPORT SUMMARIZES 14 AORTIC CARDIAC SURGERY OR INTERVENTION - OTHER UNPLANNED DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 75 - (B)(6) YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 12 FEMALE(S) AND 2 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 14 CARDIAC SURGERY OR INTERVENTION - OTHER UNPLANNED DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 1.5 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE ARE: (B)(4). THE INSTRUCTIONS FOR USE (IFU) LIST REOPERATION, INCLUDING EMERGENCY CARDIAC SURGERY, AS ADDITIONAL POTENTIAL RISKS ASSOCIATED WITH THE USE OF THE THV, DELIVERY SYSTEM, AND/OR ACCESSORIES. THIS EVENT TYPE (CARDIAC SURGERY OR INTERVENTION - OTHER UNPLANNED) IS DEFINED AS: THE PATIENT SUBSEQUENTLY UNDERWENT OTHER TYPES OF CARDIAC SURGERY OR CATH LAB INTERVENTIONS (NOT OTHERWISE SPECIFIED) THAT WAS UNPLANNED. THIS COULD INCLUDE REOPERATION FOR DAMAGE TO ADJACENT STRUCTURES, SUCH AS THE MITRAL APPARATUS, OR POTENTIALLY VALVE EXPLANT RELATED TO EMBOLIZATION, OR OTHER EVENT TYPES. CARDIOVASCULAR INJURY, INCLUDING PERFORATION OR DISSECTION OF VALVULAR STRUCTURES (E.G., INJURY TO THE MITRAL VALVE/ APPARATUS), AND VALVE EMBOLIZATION REQUIRING INTERVENTION ARE KNOWN POTENTIAL COMPLICATIONS LISTED IN THE PRODUCT INSTRUCTIONS FOR USE (IFU). CHORDAE TENDINEA RUPTURE IN TRANSCATHETER VALVE REPLACEMENT CAN OCCUR DURING ADVANCEMENT OF THE GUIDEWIRE, BAV CATHETER, OR DELIVERY SYSTEM AND IS MOST LIKELY TO OCCUR WITH ANTEGRADE APPROACHES, HOWEVER, CAN OCCUR WITH A RETROGRADE APPROACH. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE TRANSCATHETER HEART VALVES (THV). TRAINING INCLUDES PROPER GUIDEWIRE POSITIONING, AND CAREFUL MANIPULATION OF DEVICES. PER THE THV TA TRAINING MANUALS, AFTER GAINING LV ACCESS WITH THE 0.035'' SOFT GUIDEWIRE, THE WIRE SHOULD BE JIGGLED UNDER TEE IMAGING OF THE MITRAL VALVE. AN INCREASE IN MR SUGGESTS WIRE ENTANGLEMENT IN THE MITRAL SUB-VALVULAR APPARATUS. IF ENTANGLEMENT IS SUSPECTED, THE GUIDEWIRE SHOULD BE COMPLETELY REMOVED FROM THE VENTRICLE; THE OPERATOR SHOULD CHANGE DIRECTION OF THE NEEDLE AND REINSERT THE GUIDEWIRE INTO THE VENTRICLE, CHECKING AGAIN FOR WIRE ENTANGLEMENT. THE TF TRAINING MANUAL ALSO PROVIDES GUIDANCE FOR VALVE CROSSING AND WIRE EXCHANGE: EXCHANGE 0.035'' AMPLATZ EXTRA-STIFF WIRE WITH PRE-SHAPED DISTAL END IN LEFT VENTRICLE. ENSURE GUIDE CATHETER IS ADVANCED UPON WIRE EXCHANGE TO ENSURE EXCHANGE WIRE DOES NOT GET CAUGHT IN THE MITRAL CHORDAE. THERE ARE MULTIPLE PATIENT AND PROCEDURAL FACTORS THAT ALONE OR IN COMBINATION CAN CAUSE OR CONTRIBUTE TO VENTRICULAR EMBOLIZATION, INCLUDING IMPROPER POSITIONING PRIOR TO DEPLOYMENT, POOR IMAGE INTENSIFIER ANGLE, POOR COAXIAL ALIGNMENT OF THE VALVE/DELIVERY SYSTEM, A NARROW, CALCIFIED SINOTUBULAR JUNCTION, MINIMALLY OR BULKY/SEVERELY CALCIFIED AORTIC LEAFLETS, RAPID DEPLOYMENT, RELEASE OF STORED TENSION DURING DEPLOYMENT, AND MOVEMENT OF THE DELIVERY SYSTEM BY THE OPERATOR. THE THV TRAINING MANUALS INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN THV. TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS AND PROCTORED PROCEDURES. THE CORRECT ALIGNMENT AND POSITIONING OF THE DEVICE AT THE POINT OF DEPLOYMENT IS EMPHASIZED AS A KEY FACTOR TO THE PLACEMENT AND FIXATION OF THE DEVICE. OPERATORS ARE ALSO INSTRUCTED TO USE FLUOROSCOPY AS THE PRIMARY METHOD OF VISUALIZATION FOR POSITIONING AND DEPLOYMENT. IN PATIENTS WITH HIGH-RISK ANATOMICAL FEATURES FOR VENTRICULAR EMBOLIZATION (I.E., SMALL, CALCIFIED STJ, MINIMAL LEAFLET CALCIFICATION), BAV MAY PROVIDE INDICATION OF POTENTIAL BALLOON MOVEMENT DURING VALVE DEPLOYMENT. I IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21638851&pc=NPT | 2015691-2025-01788 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-18 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Insufficient Information | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P140031 | E2016006 | 9 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA VALVE. THIS REPORT SUMMARIZES 9 AORTIC READMISSION (VALVE RELATED) DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 55 - 83 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 4 FEMALE(S) AND 5 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER: E2016006 AND SUMMARIZES 9 READMISSION (VALVE RELATED) DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 142.78 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE ARE: (B)(4). VALVE RELATED RE-ADMISSION IN THE FOLLOW-UP PERIOD IS LIKELY DUE TO REOCCURRENCE OF SYMPTOMS. THIS MAY RESULT FROM REGURGITATION (CENTRAL OR PVL) OR PROGRESSION OF THE PRE-EXISTING DISEASE PROCESS AND MAY RESULT IN HEART FAILURE. CAUSES OF HEART FAILURE CAN BE MULTI-FACTORIAL. PER THE INSTRUCTION FOR USE (IFU), HEART FAILURE, VALVE REGURGITATION, AND VALVE DEGENERATION INCLUDING STENOSIS ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH AORTIC VALVE REPLACEMENT AND BIOPROSTHETIC HEART VALVES. CONGESTIVE HEART FAILURE CAN HAVE MULTIPLE ETIOLOGIES AND IS OFTEN DUE TO THE PROGRESSION OF UNDERLYING DISEASE PROCESSES, INCLUDING CORONARY ARTERY DISEASE, UNCONTROLLED HYPERTENSION, MYOCARDIAL INFARCTION, CARDIOMYOPATHY, FLUID OVERLOAD, OR VALVULAR DYSFUNCTION. RISK FACTORS THAT INCREASE A PATIENT'S RISK OF SUFFERING FROM CHF INCLUDE OBESITY, ADVANCED AGE, AND A HISTORY OF SMOKING. REGURGITATION WHICH DEVELOPS PROGRESSIVELY OVER TIME CAN BE DUE TO A NUMBER OF ISSUES INCLUDING PATIENT RELATED FACTORS OR STRUCTURAL VALVE DETERIORATION, INCLUDING CALCIFICATION, NON-CALCIFIC DEGENERATION, LEAFLET THICKENING OR FIBROSIS, OR A COMBINATION OF THESE. REGURGITATION MAY ALSO DEVELOP PROGRESSIVELY IF HOST FIBROTIC TISSUE, OR PANNUS, GROWS ONTO THE BIOPROSTHETIC VALVE. PANNUS, A CAUSE OF NONSTRUCTURAL DYSFUNCTION, MAY INTERFERE WITH FUNCTIONALITY OF THE DEVICE BY RESTRICTING THE LEAFLET MOTION LEADING TO ABNORMAL COAPTATION. THE MECHANISM BEHIND WORSENING OR LATE PVL IS NOT WELL UNDERSTOOD BUT MAY BE RELATED TO CARDIAC REMODELING. VALVE STENOSIS MAY RESULT IN SYMPTOMS SUCH AS SOB AND DECREASED EXERCISE TOLERANCE, WHICH MAY BE ACCOMPANIED BY AN INCREASED GRADIENT ACROSS THE VALVE. THIS COULD BE DUE TO EARLY CALCIFICATION OF THE LEAFLETS, HOST TISSUE OVERGROWTH, OR IN RARE CASES A NON-FUNCTIONING LEAFLET. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21638902&pc=DYB | 2015691-2025-01799 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-18 04:00 | DYB | EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | K200258 | E2016006 | 3 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA VALVE. THIS REPORT SUMMARIZES 3 AORTIC BLEEDING - ACCESS SITE DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 78 - 89 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 FEMALE(S) AND 2 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 3 BLEEDING - ACCESS SITE DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 0 DAY(S). DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). THE INSTRUCTIONS FOR USE (IFU) LIST HEMORRHAGE REQUIRING TRANSFUSION OR INTERVENTION AS A POTENTIAL RISK ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING POTENTIAL ACCESS COMPLICATIONS ASSOCIATED WITH STANDARD CARDIAC CATHETERIZATION. THE MAJORITY OF TRANSFEMORAL ACCESS RELATED BLEEDING COMPLICATIONS ARE RELATED TO DISEASED ILEO-FEMORAL VESSELS AND/OR PROCEDURAL TECHNIQUE DURING THE INSERTION OR REMOVAL OF THE SHEATH AND/OR DILATORS. THERE ARE CASES WHERE THE BLEEDING IS SIGNIFICANT AND MORE COMPLEX INTERVENTION OR TRANSFUSION IS REQUIRED TO TREAT/PREVENT A PERMANENT INJURY FROM OCCURRING. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE VALVE CAN BE DELIVERED TRANSFEMORALLY. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE; HOWEVER, A TRANSFUSION OF BLOOD WAS REPORTED IN THE REGISTRY. BY VARC DEFINITION, THE MEET CRITERIA FOR MAJOR VASCULAR COMPLICATION, THEREFORE THE EVENT IS BEING REPORTED. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21638903&pc=DYB | 2015691-2025-01800 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-18 04:00 | DYB | EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | Adverse Event Without Identified Device or Use Problem | Hematoma | K200258 | E2016006 | 2 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA VALVE. THIS REPORT SUMMARIZES 2 AORTIC BLEEDING - HEMATOMA AT ACCESS SITE DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 78 - 85 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 FEMALE(S) AND 1 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER: E2016006 AND SUMMARIZES 2 BLEEDING - HEMATOMA AT ACCESS SITE DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 0 DAY(S). DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). THE INSTRUCTIONS FOR USE (IFU) LIST HEMORRHAGE REQUIRING TRANSFUSION OR INTERVENTION AS A POTENTIAL RISK ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING POTENTIAL ACCESS COMPLICATIONS ASSOCIATED WITH STANDARD CARDIAC CATHETERIZATION. THE MAJORITY OF TRANSFEMORAL ACCESS RELATED BLEEDING COMPLICATIONS ARE RELATED TO DISEASED ILEO-FEMORAL VESSELS AND/OR PROCEDURAL TECHNIQUE DURING THE INSERTION OR REMOVAL OF THE SHEATH AND/OR DILATORS. THERE ARE CASES WHERE THE BLEEDING IS SIGNIFICANT AND MORE COMPLEX INTERVENTION OR TRANSFUSION IS REQUIRED TO TREAT/PREVENT A PERMANENT INJURY FROM OCCURRING. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE VALVE CAN BE DELIVERED TRANSFEMORALLY. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. HOWEVER, A TRANSFUSION OF BLOOD WAS REPORTED IN THE REGISTRY. BY VARC DEFINITION, THE MEET CRITERIA FOR MAJOR VASCULAR COMPLICATION THEREFORE, THE EVENT IS BEING REPORTED. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21638904&pc=NPU | 2015691-2025-01803 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-18 04:00 | NPU | EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Insufficient Information | Mitral Valve Insufficiency/ Regurgitation | P140031 | E2016006 | 1 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR MITRAL DEATH EVENTS FOR THE SAPIEN 3 ULTRA VALVE. THIS REPORT SUMMARIZES 1 MITRAL REINTERVENTION - MITRAL VALVE DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 80 - 80 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 0 FEMALE(S) AND 1 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER: E2016006 AND SUMMARIZES 1 REINTERVENTION - MITRAL VALVE DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE IN THE MITRAL POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 83 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE ARE: (B)(4). INTRA-PROCEDURAL AND IN-HOSPITAL MITRAL VALVE RE-INTERVENTION WILL TYPICALLY RESULT FROM VALVE MALPOSITION OR REGURGITATION (PVL OR CENTRAL, INCLUDING LEAFLET RESTRICTION). THESE CONDITIONS ARE KNOWN POTENTIAL RISKS ASSOCIATED WITH THE USE OF THE THV, DELIVERY SYSTEM, AND/OR ACCESSORIES. THERE ARE MULTIPLE PATIENT AND PROCEDURAL FACTORS THAT ALONE OR IN COMBINATION CAN CAUSE OR CONTRIBUTE TO VALVE MALPOSITION, INCLUDING IMPROPER POSITIONING PRIOR TO DEPLOYMENT, POOR IMAGE INTENSIFIER ANGLE, POOR COAXIAL ALIGNMENT OF THE VALVE AND DELIVERY SYSTEM, CHARACTERISTICS OF THE PRE-EXISTING VALVE OR RING, LOSS OF PACING CAPTURE, RAPID DEPLOYMENT, RELEASE OF STORED TENSION DURING DEPLOYMENT, AND MOVEMENT OF THE DELIVERY SYSTEM BY THE OPERATOR. VALVE MALPOSITION HAS THE POTENTIAL TO CONTRIBUTE TO SUBOPTIMAL COAPTATION OF THE THV VALVE LEAFLETS AND CAUSE CENTRAL INSUFFICIENCY; IT CAN LEAD TO MIGRATION OR EMBOLIZATION OF THE PROSTHESIS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE THV. TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS AND PROCTORED PROCEDURES. THE CORRECT ALIGNMENT AND POSITIONING OF THE DEVICE AT THE POINT OF DEPLOYMENT IS EMPHASIZED AS A KEY FACTOR TO THE PLACEMENT AND FIXATION OF THE DEVICE. OPERATORS ARE ALSO INSTRUCTED TO USE FLUOROSCOPY AS THE PRIMARY METHOD OF VISUALIZATION FOR POSITIONING AND DEPLOYMENT. MITRAL VALVE RE-INTERVENTION IN THE FOLLOW-UP PERIOD (IN THE ABSENCE OF PROSTHETIC CARDIAC VALVE THROMBOSIS) WILL TYPICALLY RESULT FROM ON-GOING OR WORSENING REGURGITATION, VALVE DEGENERATION RELATED TO THE FORMATION OF CALCIFICATION OR PANNUS, OR VALVE MIGRATION. THESE EVENTS ARE IDENTIFIED IN THE PRODUCT INSTRUCTIONS FOR USE (IFU) AS POTENTIAL RISKS ASSOCIATED WITH THE USE OF THE THV. REGURGITATION WHICH DEVELOPS PROGRESSIVELY OVER TIME CAN BE DUE TO SEVERAL ISSUES INCLUDING PATIENT RELATED FACTORS OR STRUCTURAL VALVE DETERIORATION, INCLUDING CALCIFICATION, NON-CALCIFIC DEGENERATION, LEAFLET THICKENING OR FIBROSIS, OR A COMBINATION OF THESE. REGURGITATION MAY ALSO DEVELOP PROGRESSIVELY IF HOST FIBROTIC TISSUE, OR PANNUS, GROWS ONTO THE BIOPROSTHETIC VALVE. PANNUS, A CAUSE OF NONSTRUCTURAL DYSFUNCTION, MAY INTERFERE WITH FUNCTIONALITY OF THE DEVICE BY RESTRICTING THE LEAFLET MOTION LEADING TO ABNORMAL COAPTATION. CALCIFICATION IS A WELL-RECOGNIZED FAILURE MODE OF BIOPROSTHETIC VALVES. THE MECHANISM OF CALCIFICATION OF BIOMATERIALS IS NOT COMPLETELY UNDERSTOOD BUT IS PROBABLY RELATED TO AN INABILITY OF THE NON-VIABLE CELLS TO MAINTAIN THEIR NORMALLY LOW INTRACELLULAR CONCENTRATION OF CALCIUM. ANOTHER POTENTIAL CONDITION THAT MAY LEAD TO RE-INTERVENTION IS PATIENT PROSTHESIS MISMATCH AND/OR UNDER-EXPANDED VALVES DUE TO THV RESTRICTION FROM THE EXISTING VALVE FRAME. INCORRECT BIOPROSTHETIC VALVE SIZING, AND INCOMPLETE FRAME EXPANSION, CAN CONTRIBUTE TO VALVE MIGRATION. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21638905&pc=DYB | 2015691-2025-01796 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-18 04:00 | DYB | EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | Adverse Event Without Identified Device or Use Problem | Unspecified Vascular Problem | K200258 | E2016006 | 27 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA VALVE. THIS REPORT SUMMARIZES 27 AORTIC VASCULAR COMPLICATION - MAJOR DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 56 - (B)(6) YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 15 FEMALE(S) AND 12 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 27 VASCULAR COMPLICATION - MAJOR DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 29.67 DAY(S). DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). ACCORDING TO THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR COMPLICATIONS, INCLUDING PERFORATION OR DISSECTION OF VESSELS WHICH MAY REQUIRE INTERVENTION, ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE TRANSFEMORAL TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, VASCULAR COMPLICATIONS ARE A WELL-RECOGNIZED COMPLICATION OF THE TRANSFEMORAL TRANSCATHETER VALVE REPLACEMENT PROCEDURE IN THIS ELDERLY POPULATION WITH MULTIPLE CO-MORBIDITIES. EDWARDS HAS REVIEWED MANY REPORTS, INCLUDING SCREENING DATA RECORDS AND SOURCE DOCUMENTATION OF VASCULAR COMPLICATIONS, AND HAS FOUND THAT THE ROOT CAUSE IS TYPICALLY RELATED TO A COMBINATION OF VESSEL SIZE, TORTUOSITY, AND CALCIFICATIONS. ALTHOUGH THE INCIDENCE IS DECREASING WITH SMALLER SHEATH/DELIVERY SYSTEM SIZES AND PHYSICIAN EXPERIENCE, THERE WILL CONTINUE TO BE CASES IN WHICH VASCULAR COMPLICATIONS WILL OCCUR. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE TRAINING MANUAL INSTRUCTS THE OPERATOR ON PROPER SHEATH INSERTION AND WITHDRAWAL TECHNIQUES, INCLUDING PRE-DILATING THE VESSEL WITH THE EDWARDS DILATORS, AS NEEDED. IT ALSO NOTES THAT CALCIFICATION MAY REDUCE LUMEN DIAMETER AND LIMIT OR PREVENT TRANSFEMORAL PASSAGE OF THE DEVICES. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE THV CAN BE DELIVERED TRANSFEMORALLY. ASSESSMENT OF LOCATION AND AMOUNT OF CIRCUMFERENTIAL CALCIUM WILL AID IN DETERMINING AREAS OF REDUCED VESSEL DIAMETERS. THE OPERATORS ARE TRAINED TO MEASURE MINIMUM VESSEL DIAMETER TAKING CALCIUM INTO ACCOUNT. THE PHYSICIAN TRAINING MANUAL ALSO LISTS THE MINIMUM RECOMMENDED VESSEL SIZE FOR EACH SIZE DEVICE. DESPITE THE BEST SCREENING TOOLS, A SMALL PERCENTAGE OF PATIENTS WILL HAVE FEMORAL/ILIAC VESSELS THAT ARE NOT AMENABLE TO THE TRANS-FEMORAL APPROACH OR WHERE INCREASED RESISTANCE IS ENCOUNTERED DURING INSERTION OF DEVICES. IN MANY CASES, THE VESSEL MINIMUM LUMINAL DIAMETER (MLD) MAY BE BORDERLINE OR BELOW THE INDICATED SIZE. IN ADDITION, SIGNIFICANT CALCIFICATION AND/OR TORTUOSITY, NOT ALWAYS APPRECIABLE ON IMAGING, COULD BE CONTRIBUTING FACTORS TO THE EVENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21638906&pc=DYB | 2015691-2025-01797 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-18 04:00 | DYB | EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | Insufficient Information | Unspecified Vascular Problem | K200258 | E2016006 | 8 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA VALVE. THIS REPORT SUMMARIZES 8 AORTIC VASCULAR SURGERY OR INTERVENTION - UNPLANNED DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 77 - 89 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 6 FEMALE(S) AND 2 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 8 VASCULAR SURGERY OR INTERVENTION - UNPLANNED DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 3.25 DAY(S). DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). ACCORDING TO THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR COMPLICATIONS, INCLUDING PERFORATION OR DISSECTION OF VESSELS WHICH MAY REQUIRE INTERVENTION, ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE TRANSFEMORAL TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, VASCULAR COMPLICATIONS ARE A WELL-RECOGNIZED COMPLICATION OF THE TRANSFEMORAL TRANSCATHETER VALVE REPLACEMENT PROCEDURE IN THIS ELDERLY POPULATION WITH MULTIPLE CO-MORBIDITIES. EDWARDS HAS REVIEWED MANY REPORTS, INCLUDING SCREENING DATA RECORDS AND SOURCE DOCUMENTATION OF VASCULAR COMPLICATIONS, AND HAS FOUND THAT THE ROOT CAUSE IS TYPICALLY RELATED TO A COMBINATION OF VESSEL SIZE, TORTUOSITY, AND CALCIFICATIONS. ALTHOUGH THE INCIDENCE IS DECREASING WITH SMALLER SHEATH/DELIVERY SYSTEM SIZES AND PHYSICIAN EXPERIENCE, THERE WILL CONTINUE TO BE CASES IN WHICH VASCULAR COMPLICATIONS WILL OCCUR. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE TRAINING MANUAL INSTRUCTS THE OPERATOR ON PROPER SHEATH INSERTION AND WITHDRAWAL TECHNIQUES, INCLUDING PRE-DILATING THE VESSEL WITH THE EDWARDS DILATORS, AS NEEDED. IT ALSO NOTES THAT CALCIFICATION MAY REDUCE LUMEN DIAMETER AND LIMIT OR PREVENT TRANSFEMORAL PASSAGE OF THE DEVICES. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE SAPIEN VALVE CAN BE DELIVERED TRANSFEMORALLY. ASSESSMENT OF LOCATION AND AMOUNT OF CIRCUMFERENTIAL CALCIUM WILL AID IN DETERMINING AREAS OF REDUCED VESSEL DIAMETERS. THE OPERATORS ARE TRAINED TO MEASURE MINIMUM VESSEL DIAMETER TAKING CALCIUM INTO ACCOUNT. THE PHYSICIAN TRAINING MANUAL ALSO LISTS THE MINIMUM RECOMMENDED VESSEL SIZE FOR EACH SIZE DEVICE. DESPITE THE BEST SCREENING TOOLS, A SMALL PERCENTAGE OF PATIENTS WILL HAVE FEMORAL/ILIAC VESSELS THAT ARE NOT AMENABLE TO THE TRANS-FEMORAL APPROACH OR WHERE INCREASED RESISTANCE IS ENCOUNTERED DURING INSERTION OF DEVICES. IN MANY CASES, THE VESSEL MINIMUM LUMINAL DIAMETER (MLD) MAY BE BORDERLINE OR BELOW THE INDICATED SIZE. IN ADDITION, SIGNIFICANT CALCIFICATION AND/OR TORTUOSITY, NOT ALWAYS APPRECIABLE ON IMAGING, COULD BE CONTRIBUTING FACTORS TO THE EVENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21638907&pc=NPT | 2015691-2025-01801 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-18 04:00 | NPT | COMMANDER DELIVERY SYSTEM UNKNOWN | Adverse Event Without Identified Device or Use Problem | Cardiac Perforation | P140031 | E2016006 | 8 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA VALVE. THIS REPORT SUMMARIZES 8 AORTIC CARDIAC PERFORATION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 75 - 92 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 5 FEMALE(S) AND 3 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 8 CARDIAC PERFORATION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 0 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS COMMANDER DELIVERY SYSTEM ARE: (B)(4). PER THE INSTRUCTIONS FOR USE, CARDIOVASCULAR INJURIES SUCH AS PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, MYOCARDIAL OR VALVULAR STRUCTURES ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE AND MAY REQUIRE INTERVENTION. THERE ARE SEVERAL POTENTIAL ETIOLOGIES FOR PERFORATION DURING A TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING PERFORATION BY THE GUIDEWIRE, THE DELIVERY SYSTEM, OR THE TRANSVENOUS PACER (TVP) LEAD. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE TRANSCATHETER HEART VALVE (THV). TRAINING INCLUDES PROPER GUIDEWIRE POSITIONING, FIXATION OF THE TVP TO PREVENT VENTRICLE PERFORATION, AND CAREFUL MANIPULATION OF DEVICES. PER THE PROCEDURE DIDACTIC, PATIENTS WITH SMALL VENTRICLES ARE AT PARTICULARLY HIGH RISK FOR VENTRICULAR PERFORATION. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21638908&pc=DYB | 2015691-2025-01798 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-18 04:00 | DYB | EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | Adverse Event Without Identified Device or Use Problem | Retroperitoneal Hemorrhage | K200258 | E2016006 | 6 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA VALVE. THIS REPORT SUMMARIZES 6 AORTIC BLEEDING - RETROPERITONEAL DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 79 - 88 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 5 FEMALE(S) AND 1 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 6 BLEEDING - RETROPERITONEAL DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 0.67 DAY(S). DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). THE INSTRUCTIONS FOR USE (IFU) LIST HEMORRHAGE REQUIRING TRANSFUSION OR INTERVENTION AS A POTENTIAL RISK ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING POTENTIAL ACCESS COMPLICATIONS ASSOCIATED WITH STANDARD CARDIAC CATHETERIZATION. THE MAJORITY OF TRANSFEMORAL ACCESS RELATED BLEEDING COMPLICATIONS ARE RELATED TO DISEASED ILEO-FEMORAL VESSELS AND/OR PROCEDURAL TECHNIQUE DURING THE INSERTION OR REMOVAL OF THE SHEATH AND/OR DILATORS. THERE ARE CASES WHERE THE BLEEDING IS SIGNIFICANT AND MORE COMPLEX INTERVENTION OR TRANSFUSION IS REQUIRED TO TREAT/PREVENT A PERMANENT INJURY FROM OCCURRING. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE VALVE CAN BE DELIVERED TRANSFEMORALLY. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE; ; HOWEVER, A TRANSFUSION OF BLOOD WAS REPORTED IN THE REGISTRY. BY VARC DEFINITION, THE MEET CRITERIA FOR MAJOR VASCULAR COMPLICATION THEREFORE, THE EVENT IS BEING REPORTED. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21638910&pc=NPT | 2015691-2025-01802 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-18 04:00 | NPT | COMMANDER DELIVERY SYSTEM UNKNOWN | Adverse Event Without Identified Device or Use Problem |  | P140031 | E2016006 | 2 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA VALVE. THIS REPORT SUMMARIZES 2 AORTIC AORTIC DISSECTION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 84 - 87 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 0 FEMALE(S) AND 2 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 2 AORTIC DISSECTION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 0.5 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS COMMANDER DELIVERY SYSTEM ARE: (B)(4). PER THE INSTRUCTIONS FOR USE (IFU) CARDIOVASCULAR INJURY, SUCH AS PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, MYOCARDIUM, OR VALVULAR STRUCTURES, IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT PROCEDURE. ASCENDING AORTIC DISSECTION MAY OCCUR WHEN MULTIPLE ATTEMPTS ARE MADE TO CROSS THE STENOTIC NATIVE VALVE, AND/OR WHEN EXCESSIVE FORCE IS USED. AORTIC DISSECTION IN A DISTAL SECTION OF THE AORTA CAN OCCUR WITH DEVICE MANIPULATION AND NAVIGATION THROUGH ANATOMIC TORTUOSITY. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE TRANSCATHETER HEART VALVE (THV). THE THV TRAINING MANUALS PROVIDE GUIDANCE TO FACILITATE SAFE CROSSING OF THE NATIVE VALVE, INCLUDING CAMERA PROJECTIONS, HANDLING DURING ADVANCEMENT, AND TROUBLESHOOTING TECHNIQUES IF DIFFICULTY IS ENCOUNTERED. AS STATED, EXCESSIVE FORCE SHOULD NOT BE USED WHEN THE DEVICE HAS DIFFICULTY CROSSING THE STENOTIC VALVE. ADDING TENSION TO THE WIRE, PULLING BACK THE SYSTEM TO RE-ORIENT THE VALVE, AS NEEDED, AND TORQUEING OF THE FLEX CATHETER MAY BE HELPFUL IN SOLVING THE PROBLEM. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21638938&pc=DSQ | 2916596-2025-01492 | 2025-02-25 05:00 | Death | THORATEC CORPORATION | 2025-03-18 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. DUE TO PATIENT PRIVACY LAWS THE MANAGING HOSPITAL WOULD NOT COMMUNICATE PATIENT OUTCOME INFORMATION. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED.. Manufacturer Narrative: E1: REPORTER ADDRESS: (B)(6) NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. G IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21638955&pc=DSQ | 2916596-2025-01635 | 2025-02-20 05:00 | Death | THORATEC CORPORATION | 2025-03-18 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Unspecified Infection | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO INFECTION. Manufacturer Narrative: CORRECTED DATA: SECTION H3: DEVICE NOT EVALUATED BY MANUFACTURER MANUFACTURER¿S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF INFECTION COULD NOT BE CONCLUSIVELY DETERMINED. INFECTION IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. INFECTION HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21639259&pc=NPT | 2015691-2025-01808 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-18 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Rupture | P140031 | E2016006 | 10 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 10 AORTIC ANNULAR RUPTURE DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 62 - 91 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 4 FEMALE(S) AND 6 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 10 ANNULAR RUPTURE DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 0.1 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE ARE: (B)(4). PER THE INSTRUCTIONS FOR USE (IFU) CARDIOVASCULAR INJURIES, SUCH AS PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, MYOCARDIAL OR VALVULAR STRUCTURES, AND HEMATOMA ARE A KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT PROCEDURE. ACCORDING TO THE THV TRAINING MANUALS, RISK FACTORS FOR AORTIC DISSECTION, HEMATOMA, OR ANNULAR RUPTURE DURING THE TRANSCATHETER VALVE REPLACEMENT PROCEDURE INCLUDE SIGNIFICANT THV OVERSIZING, SEVERELY OBLITERATED SINUSES OF VALSALVA, PORCELAIN AORTA AND/OR PRESENCE OF BULKY CALCIFICATION AND NARROW CALCIFIED STJ. IN ADDITION, ADVANCED AGE, FEMALE GENDER, SMALL BODY WEIGHT, AND STEROID DEPENDENCY CAN ALSO BE CONTRIBUTING FACTORS. THE SAPIEN VALVE RELIES ON NATIVE VALVE CALCIUM TO SECURELY ANCHOR TO THE ANNULUS. DESPITE THIS BENEFICIAL ASPECT OF CALCIUM, BULKY CALCIUM CAN INCREASE THE RISK OF CALCIFIC NODULE DISPLACEMENT INTO THE VASCULATURE, WHICH CAN LEAD TO VASCULAR INJURY. AT TIMES THE EXTENT AND DISTRIBUTION OF CALCIUM CAN IMPAIR EASE OF DELIVERY OF THE VALVE, CORRECT POSITIONING OF THE VALVE, DEPLOYMENT OF THE VALVE AND PROCEDURAL SUCCESS. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21639260&pc=NPT | 2015691-2025-01809 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-18 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P140031 | E2016006 | 4 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 4 AORTIC PERCUTANEOUS CORONARY INTERVENTION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 70 - 81 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 2 FEMALE(S) AND 2 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER: E2016006 AND SUMMARIZES 4 PERCUTANEOUS CORONARY INTERVENTION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 0 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE ARE: (B)(4). PER THE INSTRUCTIONS FOR USE (IFU), CORONARY FLOW OBSTRUCTION IS A POTENTIAL ADVERSE EVENT ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT PROCEDURE. THE IFU CAUTIONS THAT THE SAFETY AND EFFECTIVENESS HAVE NOT BEEN ESTABLISHED IN PATIENTS WITH BULKY CALCIFIED AORTIC VALVE LEAFLETS IN CLOSE PROXIMITY TO THE CORONARY OSTIA. IN ADDITION, IT WARNS THAT CAUTION SHOULD BE EXERCISED IN IMPLANTING A BIOPROSTHESIS IN PATIENTS WITH CLINICALLY SIGNIFICANT CORONARY ARTERY DISEASE. CORONARY OBSTRUCTION CAN RESULT IN MYOCARDIAL ISCHEMIA OR INFARCTION DUE TO OBSTRUCTION OF THE CORONARY BLOOD FLOW AND MAY REQUIRE INTERVENTION (E.G., PCI). THERE ARE MULTIPLE PATIENT FACTORS THAT COULD PUT THE PATIENT AT RISK FOR CORONARY FLOW OBSTRUCTION DURING THE TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING SIGNIFICANT UNDERLYING CORONARY ARTERY DISEASE AND BULKY CALCIFICATION OF THE NATIVE LEAFLETS AND ROOT. DISPLACEMENT OF CALCIUM DEPOSITS WITH EMBOLIZATION OF DEBRIS INTO ONE OF THE ARTERIES, OR AORTIC DISSECTION WITH CONTINUITY OF THE RUPTURE INTO THE INTIMA OF ONE OF THE CORONARY OSTIA, CAN RESULT IN THIS COMPLICATION. ADDITIONALLY, MINIMAL DISTANCE BETWEEN THE NATIVE ANNULUS AND THE CORONARY OSTIA, BULKY CALCIFICATION, LONG NATIVE LEAFLETS, AND OBLITERATED CORONARY SINUSES. PROCEDURAL FACTORS SUCH AS TORN NATIVE LEAFLET DURING BALLOON VALVULOPLASTY, PLAQUE SHIFT, DEPLOYMENT OF THE BIOPROSTHETIC HEART VALVE TOO AORTIC, AND SIGNIFICANT VALVE OVER SIZING COULD ALSO CONTRIBUTE TO THIS COMPLICATION. THE EDWARDS THV MANUALS ADVISE THE OPERATOR ON PRE-PROCEDURE ASSESSMENT OF THE AORTIC VALVE, ROOT, AND CORONARY ANATOMY. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE TRANSCATHETER HEART VALVE (THV). TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS AND PROCTORED PROCEDURES. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21639261&pc=NPT | 2015691-2025-01821 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-18 04:00 | NPT | COMMANDER DELIVERY SYSTEM UNKNOWN | Adverse Event Without Identified Device or Use Problem | Cardiac Perforation | P140031 | E2016006 | 14 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 14 AORTIC CARDIAC PERFORATION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 65 - (B)(6) YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 8 FEMALE(S) AND 6 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 14 CARDIAC PERFORATION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 0.07 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS COMMANDER DELIVERY SYSTEM ARE: (B)(4). PER THE INSTRUCTIONS FOR USE, CARDIOVASCULAR INJURIES SUCH AS PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, MYOCARDIAL OR VALVULAR STRUCTURES ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE AND MAY REQUIRE INTERVENTION. THERE ARE SEVERAL POTENTIAL ETIOLOGIES FOR PERFORATION DURING A TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING PERFORATION BY THE GUIDEWIRE, THE DELIVERY SYSTEM, OR THE TRANSVENOUS PACER (TVP) LEAD. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE TRANSCATHETER HEART VALVE (THV). TRAINING INCLUDES PROPER GUIDEWIRE POSITIONING, FIXATION OF THE TVP TO PREVENT VENTRICLE PERFORATION, AND CAREFUL MANIPULATION OF DEVICES. PER THE PROCEDURE DIDACTIC, PATIENTS WITH SMALL VENTRICLES ARE AT PARTICULARLY HIGH RISK FOR VENTRICULAR PERFORATION. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21639262&pc=NPT | 2015691-2025-01804 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-18 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Insufficient Information | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P140031 | E2016006 | 23 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 23 AORTIC CARDIAC SURGERY OR INTERVENTION - OTHER UNPLANNED DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 62 - 90 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 13 FEMALE(S) AND 10 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER: E2016006 AND SUMMARIZES 23 CARDIAC SURGERY OR INTERVENTION - OTHER UNPLANNED DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 0.91 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE ARE: (B)(4). THE INSTRUCTIONS FOR USE (IFU) LIST REOPERATION, INCLUDING EMERGENCY CARDIAC SURGERY, AS ADDITIONAL POTENTIAL RISKS ASSOCIATED WITH THE USE OF THE THV, DELIVERY SYSTEM, AND/OR ACCESSORIES. THIS EVENT TYPE (CARDIAC SURGERY OR INTERVENTION - OTHER UNPLANNED) IS DEFINED AS: THE PATIENT SUBSEQUENTLY UNDERWENT OTHER TYPES OF CARDIAC SURGERY OR CATH LAB INTERVENTIONS (NOT OTHERWISE SPECIFIED) THAT WAS UNPLANNED. THIS COULD INCLUDE REOPERATION FOR DAMAGE TO ADJACENT STRUCTURES, SUCH AS THE MITRAL APPARATUS, OR POTENTIALLY VALVE EXPLANT RELATED TO EMBOLIZATION, OR OTHER EVENT TYPES. CARDIOVASCULAR INJURY, INCLUDING PERFORATION OR DISSECTION OF VALVULAR STRUCTURES (E.G., INJURY TO THE MITRAL VALVE/ APPARATUS), AND VALVE EMBOLIZATION REQUIRING INTERVENTION ARE KNOWN POTENTIAL COMPLICATIONS LISTED IN THE PRODUCT INSTRUCTIONS FOR USE (IFU). CHORDAE TENDINEA RUPTURE IN TRANSCATHETER VALVE REPLACEMENT CAN OCCUR DURING ADVANCEMENT OF THE GUIDEWIRE, BAV CATHETER, OR DELIVERY SYSTEM AND IS MOST LIKELY TO OCCUR WITH ANTEGRADE APPROACHES, HOWEVER, CAN OCCUR WITH A RETROGRADE APPROACH. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE TRANSCATHETER HEART VALVES (THV). TRAINING INCLUDES PROPER GUIDEWIRE POSITIONING, AND CAREFUL MANIPULATION OF DEVICES. PER THE THV TA TRAINING MANUALS, AFTER GAINING LV ACCESS WITH THE 0.035'' SOFT GUIDEWIRE, THE WIRE SHOULD BE JIGGLED UNDER TEE IMAGING OF THE MITRAL VALVE. AN INCREASE IN MR SUGGESTS WIRE ENTANGLEMENT IN THE MITRAL SUB-VALVULAR APPARATUS. IF ENTANGLEMENT IS SUSPECTED, THE GUIDEWIRE SHOULD BE COMPLETELY REMOVED FROM THE VENTRICLE; THE OPERATOR SHOULD CHANGE DIRECTION OF THE NEEDLE AND REINSERT THE GUIDEWIRE INTO THE VENTRICLE, CHECKING AGAIN FOR WIRE ENTANGLEMENT. THE TF TRAINING MANUAL ALSO PROVIDES GUIDANCE FOR VALVE CROSSING AND WIRE EXCHANGE: EXCHANGE 0.035'' AMPLATZ EXTRA-STIFF WIRE WITH PRE-SHAPED DISTAL END IN LEFT VENTRICLE. ENSURE GUIDE CATHETER IS ADVANCED UPON WIRE EXCHANGE TO ENSURE EXCHANGE WIRE DOES NOT GET CAUGHT IN THE MITRAL CHORDAE. THERE ARE MULTIPLE PATIENT AND PROCEDURAL FACTORS THAT ALONE OR IN COMBINATION CAN CAUSE OR CONTRIBUTE TO VENTRICULAR EMBOLIZATION, INCLUDING IMPROPER POSITIONING PRIOR TO DEPLOYMENT, POOR IMAGE INTENSIFIER ANGLE, POOR COAXIAL ALIGNMENT OF THE VALVE/DELIVERY SYSTEM, A NARROW, CALCIFIED SINOTUBULAR JUNCTION, MINIMALLY OR BULKY/SEVERELY CALCIFIED AORTIC LEAFLETS, RAPID DEPLOYMENT, RELEASE OF STORED TENSION DURING DEPLOYMENT, AND MOVEMENT OF THE DELIVERY SYSTEM BY THE OPERATOR. THE THV TRAINING MANUALS INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN THV. TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS AND PROCTORED PROCEDURES. THE CORRECT ALIGNMENT AND POSITIONING OF THE DEVICE AT THE POINT OF DEPLOYMENT IS EMPHASIZED AS A KEY FACTOR TO THE PLACEMENT AND FIXATION OF THE DEVICE. OPERATORS ARE ALSO INSTRUCTED TO USE FLUOROSCOPY AS THE PRIMARY METHOD OF VISUALIZATION FOR POSITIONING AND DEPLOYMENT. IN PATIENTS WITH HIGH-RISK ANATOMICAL FEATURES FOR VENTRICULAR EMBOLIZATION (I.E., SMALL, CALCIFIED STJ, MINIMAL LEAFLET CALCIFICATION), BAV MAY PROVIDE INDICATION OF POTENTIAL BALLOON MOVEMENT DURING VALVE DEPLOYMENT. I IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21639263&pc=NPT | 2015691-2025-01806 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-18 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Ischemia Stroke | P140031 | E2016006 | 12 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 12 AORTIC STROKE - ISCHEMIC DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 69 - 95 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 5 FEMALE(S) AND 7 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 12 STROKE - ISCHEMIC DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 5.67 DAY(S).THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE ARE: (B)(4). PER THE INSTRUCTIONS FOR USE, PERMANENT OR TRANSIENT NEUROLOGICAL EVENTS INCLUDING STROKE ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH INVASIVE CARDIAC PROCEDURE INCLUDING TRANSCATHETER HEART VALVE PROCEDURES AND THE USE OF THE EDWARDS THV DEVICES. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, STROKE IS RECOGNIZED IN THE LITERATURE AS A WELL-KNOWN COMPLICATION IN A SMALL NUMBER OF PATIENTS UNDERGOING TRANSCATHETER VALVE REPLACEMENT PROCEDURE. RISK FACTORS CORRELATING WITH A NUMBER OF PATIENT CO-MORBIDITIES HAVE BEEN IDENTIFIED. ALTHOUGH IN MANY CASES THE ROOT CAUSE OF THE EVENT IS UNABLE TO BE DETERMINED, STROKES DURING TRANSCATHETER VALVE REPLACEMENT PROCEDURE ARE UNDOUBTEDLY MULTIFACTORIAL, THE DOMINANT ETIOLOGY LIKELY BEING INTRA-PROCEDURE EMBOLIC EVENTS. A TRANSCRANIAL DOPPLER STUDY DURING TRANSCATHETER VALVE REPLACEMENT PROCEDURE DEMONSTRATED THAT THE MAJORITY OF PROCEDURAL EMBOLIC EVENTS OCCURRED DURING BALLOON VALVULOPLASTY, MANIPULATION OF CATHETERS ACROSS THE AORTIC VALVE, AND VALVE IMPLANTATION. AN ANALYSIS IN PATIENTS UNDERGOING VALVE SURGERY REVEALED FOUR BASELINE CHARACTERISTICS AND TWO PROCEDURAL EVENTS THAT WERE ASSOCIATED WITH EARLY POST-PROCEDURE STROKE: FEMALE SEX, EF < 30%, DIABETES, AGE OLDER THAN 70 YEARS, BYPASS PROCEDURE TIME> 120 MIN, AND CALCIFICATION OF THE ASCENDING AORTA. PREDICTORS OF LATE STROKE HAVE INCLUDED FEMALE SEX, AGE OLDER THAN 75 YEARS, ATRIAL FIBRILLATION, AND A HISTORY OF OR CURRENT SMOKING. THERE WERE NO IMPORTANT DIFFERENCES IN THE FREQUENCY OF LATE STROKES BETWEEN TRANSCATHETER VALVE REPLACEMENT PROCEDURE AND AVR PATIENTS. AFTER TRANSCATHETER VALVE REPLACEMENT PROCEDURE, THERE APPEARS TO BE A MORE SIGNIFICANT PROPORTION OF EARLY STROKES OCCURRING < 24 H POST-PROCEDURE, BUT TRANSCATHETER VALVE REPLACEMENT PROCEDURE PATIENTS WITH MULTIPLE CO MORBIDITIES ARE PROBABLY AT HIGHER RISK OF BOTH EARLY AND LATE STROKES. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE HOWEVER, THE EVENT MAY BE RELATED TO THE MECHANISM DESCRIBED ABOVE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21639264&pc=DYB | 2015691-2025-01820 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-18 04:00 | DYB | EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | Adverse Event Without Identified Device or Use Problem | Retroperitoneal Hemorrhage | K200258 | E2016006 | 5 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 5 AORTIC BLEEDING - RETROPERITONEAL DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 76 - 86 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 2 FEMALE(S) AND 3 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 5 BLEEDING - RETROPERITONEAL DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 2 DAY(S). DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). THE INSTRUCTIONS FOR USE (IFU) LIST HEMORRHAGE REQUIRING TRANSFUSION OR INTERVENTION AS A POTENTIAL RISK ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING POTENTIAL ACCESS COMPLICATIONS ASSOCIATED WITH STANDARD CARDIAC CATHETERIZATION. THE MAJORITY OF TRANSFEMORAL ACCESS RELATED BLEEDING COMPLICATIONS ARE RELATED TO DISEASED ILEO-FEMORAL VESSELS AND/OR PROCEDURAL TECHNIQUE DURING THE INSERTION OR REMOVAL OF THE SHEATH AND/OR DILATORS. THERE ARE CASES WHERE THE BLEEDING IS SIGNIFICANT AND MORE COMPLEX INTERVENTION OR TRANSFUSION IS REQUIRED TO TREAT/PREVENT A PERMANENT INJURY FROM OCCURRING. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE VALVE CAN BE DELIVERED TRANSFEMORALLY. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE; HOWEVER, A TRANSFUSION OF BLOOD WAS REPORTED IN THE REGISTRY. BY VARC DEFINITION, THE MEET CRITERIA FOR MAJOR VASCULAR COMPLICATION THEREFORE, THE EVENT IS BEING REPORTED. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21639265&pc=NPT | 2015691-2025-01813 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-18 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P140031 | E2016006 | 2 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 2 AORTIC DEVICE EMBOLIZATION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 84 - 84 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 0 FEMALE(S) AND 2 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 2 DEVICE EMBOLIZATION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 0 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE ARE: (B)(4). PER THE INSTRUCTIONS FOR USE, VALVE EMBOLIZATION IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT PROCEDURE. THERE ARE MULTIPLE PATIENT AND PROCEDURAL FACTORS THAT ALONE OR IN COMBINATION CAN CAUSE OR CONTRIBUTE TO EMBOLIZATION, INCLUDING IMPROPER POSITIONING PRIOR TO DEPLOYMENT, POOR IMAGE INTENSIFIER ANGLE, POOR COAXIAL ALIGNMENT OF THE VALVE AND DELIVERY SYSTEM, , MINIMALLY OR BULKY/SEVERELY CALCIFIED AORTIC LEAFLETS, PRESERVED EJECTION FRACTION, LOSS OF PACING CAPTURE, RAPID DEPLOYMENT, RELEASE OF STORED TENSION DURING DEPLOYMENT, INACCURATE MEASUREMENT OF THE LANDING ZONE ,AND MOVEMENT OF THE DELIVERY SYSTEM BY THE OPERATOR. THE THV TRAINING MANUALS INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE THV. TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS AND PROCTORED PROCEDURES. THE CORRECT ALIGNMENT AND POSITIONING OF THE DEVICE AT THE POINT OF DEPLOYMENT IS EMPHASIZED AS A KEY FACTOR TO THE PLACEMENT AND FIXATION OF THE DEVICE. OPERATORS ARE ALSO INSTRUCTED TO USE FLUOROSCOPY AS THE PRIMARY METHOD OF VISUALIZATION FOR POSITIONING AND DEPLOYMENT. IN PATIENTS WITH HIGH-RISK ANATOMICAL FEATURES FOR EMBOLIZATION (I.E. MINIMAL LEAFLET CALCIFICATION), BAV MAY PROVIDE INDICATION OF POTENTIAL BALLOON MOVEMENT DURING VALVE DEPLOYMENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21639266&pc=NPT | 2015691-2025-01807 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-18 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Heart Block | P140031 | E2016006 | 11 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 11 AORTIC PERMANENT PACEMAKER DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 70 - (B)(6) YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 4 FEMALE(S) AND 7 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 11 PERMANENT PACEMAKER DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 8.09 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE ARE: (B)(4). PER THE INSTRUCTIONS FOR USE, CONDUCTION SYSTEM DEFECTS (HEART BLOCK) ARRHYTHMIAS AND CONDUCTION SYSTEM DEFECTS THAT MAY REQUIRE A PERMANENT PACEMAKER ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH BALLOON AORTIC VALVULOPLASTY, DEPLOYMENT OF THE PROSTHETIC VALVE, AND THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE. ACCORDING TO THE VALVE ACADEMIC RESEARCH CONSORTIUM (VARC) GUIDELINES, THE CLOSE ANATOMICAL RELATIONSHIP BETWEEN THE AORTIC VALVE COMPLEX AND THE BRANCHING ATRIOVENTRICULAR BUNDLE MAY PROVIDE AN EXPLANATION FOR THESE COMPLICATIONS OF THE TRANSCATHETER VALVE REPLACEMENT PROCEDURE. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY -CONDUCTION DISTURBANCES/ HEART BLOCK WRITTEN BY EDWARDS LIFESCIENCES, ATRIOVENTRICULAR CONDUCTION DISTURBANCES AFTER THV PLACEMENT ARE ASSOCIATED WITH MANY PATIENT RELATED AND PROCEDURAL RELATED FACTORS, INCLUDING PRE-OPERATIVE CO-MORBID STATUS, THE DEGREE AND BULKINESS OF AORTIC VALVE AND ANNULAR CALCIFICATION, INTER-VENTRICULAR SEPTAL THICKNESS, PRE-EXISTING ELECTROCARDIOGRAM ABNORMALITIES, THE DEPTH OF PROSTHESIS IMPLANTATION, AND THE PROFILE OF THE IMPLANTED PROSTHESIS. UNLIKE CONVENTIONAL AVR, WHERE THERE MAY BE LOCALIZED TRAUMA DUE TO DECALCIFICATION OF THE ANNULUS AND/OR SUTURE PLACEMENT IN THE PROXIMITY OF THE AV NODE OR THE BUNDLES, TRANSCATHETER VALVE REPLACEMENT MAY CAUSE CONDUCTION ABNORMALITIES THROUGH MECHANICAL IMPINGEMENT OF THE CONDUCTION SYSTEM BY THE PROSTHESIS. THE MECHANISMS OF THE DEVELOPMENT OF HEART BLOCK AFTER THV PLACEMENT ARE WELL DOCUMENTED AND DESCRIBED IN THE LITERATURE. IT IS ALSO DOCUMENTED THAT PRE-EXISTING HEART BLOCK IS COMMON IN PATIENTS UNDERGOING THV PLACEMENT OR SURGICAL AVR AND ANOTHER 4-6 % WILL DEVELOP POSTOPERATIVE HEART BLOCK, POTENTIALLY REQUIRING A PERMANENT PACEMAKER. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE, HOWEVER THE CONDUCTION DISTURBANCE IS LIKELY RELATED TO THE MECHANISM DESCRIBED ABOVE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21639267&pc=NPT | 2015691-2025-01814 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-18 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Myocardial Infarction | P140031 | E2016006 | 2 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 2 AORTIC MYOCARDIAL INFARCTION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 70 - 77 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 0 FEMALE(S) AND 2 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER: E2016006 AND SUMMARIZES 2 MYOCARDIAL INFARCTION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 1 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE ARE: (B)(4). PER THE INSTRUCTIONS FOR USE (IFU), THROMBUS FORMATION, PLAQUE DISLODGEMENT, AND EMBOLIZATION THAT MAY RESULT IN MYOCARDIAL INFARCTION (MI) ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE. THE IFU CAUTIONS THAT THE SAFETY AND EFFECTIVENESS HAVE NOT BEEN ESTABLISHED FOR PATIENTS WITH BULKY CALCIFIED AORTIC VALVE LEAFLETS IN CLOSE PROXIMITY TO THE CORONARY OSTIA. IN ADDITION, CAUTION SHOULD BE EXERCISED WHEN IMPLANTING A BIOPROSTHESIS IN PATIENTS WITH CLINICALLY SIGNIFICANT CORONARY ARTERY DISEASE. MI RELATED TO THE TRANSCATHETER VALVE REPLACEMENT PROCEDURE AND THE VALVE IMPLANT WILL MANIFEST INTRA-PROCEDURALLY OR IN THE IMMEDIATE POST-OPERATIVE PERIOD AND IS TYPICALLY DUE TO A COMBINATION OF PATIENT AND/OR PROCEDURAL FACTORS. AS DEFINED IN THE VALVE ACADEMIC RESEARCH CONSORTIUM (VARC) PUBLICATION ON TRANSCATHETER VALVE REPLACEMENT COMPLICATIONS, THE PERI-PROCEDURAL INTERVAL IS INCLUSIVE OF ALL EVENTS THAT BEGIN WITHIN 72 HRS. OF THE INDEX PROCEDURE. MI'S THAT OCCUR AFTER THE PERI-PROCEDURAL PERIOD (>72 HRS.) ARE TYPICALLY RELATED TO THE PATIENT'S UNDERLYING CORONARY DISEASE. THERE ARE MULTIPLE PATIENT FACTORS THAT COULD PUT THE PATIENT AT RISK FOR MI DURING THE TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING SIGNIFICANT UNDERLYING CORONARY ARTERY DISEASE, AND BULKY CALCIFICATION OF THE NATIVE LEAFLETS AND ROOT. DISPLACEMENT OF CALCIUM DEPOSITS WITH EMBOLIZATION OF DEBRIS INTO ONE OF THE ARTERIES, OR AORTIC DISSECTION WITH CONTINUITY OF THE RUPTURE INTO THE INTIMA OF ONE OF THE CORONARY OSTIA, CAN RESULT IN THIS COMPLICATION. THE EDWARDS THV MANUALS ADVISE THE OPERATOR ON PRE-PROCEDURE ASSESSMENT OF THE AORTIC VALVE, ROOT, AND CORONARY ANATOMY. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE TRANSCATHETER HEART VALVE (THV). TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS AND PROCTORED PROCEDURES. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21639268&pc=NPT | 2015691-2025-01811 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-18 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P140031 | E2016006 | 3 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 3 AORTIC DEVICE MIGRATION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 51 - 84 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 0 FEMALE(S) AND 3 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 3 DEVICE MIGRATION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 0 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE ARE: (B)(4). PER THE INSTRUCTIONS FOR USE, VALVE MIGRATION REQUIRING INTERVENTION IS A POTENTIAL ADVERSE EVENT ASSOCIATED WITH TRANSCATHETER AORTIC VALVE REPLACEMENT. ACCORDING TO LITERATURE REVIEW, VALVE MIGRATION RESULTS WHEN FORCES ACTING ON THE TRANSCATHETER HEART VALVE (THV) OVERCOME THE STRENGTH OF ATTACHMENT OF THE VALVE TO THE ANNULUS. STENT VALVES ARE SUBJECTED TO ANTEGRADE EJECTION FORCES DURING SYSTOLE. LESS-THAN-SEVERE AND NON-UNIFORMLY DISTRIBUTED CALCIFICATION OF THE NATIVE LEAFLETS, INCORRECT BIOPROSTHETIC VALVE SIZING, AND INCOMPLETE FRAME EXPANSION CAN CONTRIBUTE TO VALVE MIGRATION. ADDITIONALLY, RESIDUAL OVERHANGING LEAFLETS CAN EXERT DOWNWARDS FORCE DURING DIASTOLE, CAUSING MIGRATION OF THE THV TOWARDS THE LEFT VENTRICLE. THE EDWARDS THV PATIENT SCREENING MANUAL ADVISES THE OPERATOR ON PRE-PROCEDURE ASSESSMENT OF THE AORTIC VALVE AND ROOT, TAKING INTO CONSIDERATION THE DEGREE AND DISTRIBUTION OF NATIVE LEAFLET CALCIFICATION. THE PROCEDURAL DIDACTIC INSTRUCTS THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE TRANSCATHETER HEART VALVE (THV). TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS AND PROCTORED PROCEDURES. THE CORRECT SIZING, ALIGNMENT AND POSITIONING OF THE DEVICE ARE EMPHASIZED AS KEY FACTORS TO THE PLACEMENT AND FIXATION OF THE DEVICE. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21639269&pc=NPT | 2015691-2025-01812 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-18 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem |  | P140031 | E2016006 | 2 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 2 AORTIC CORONARY ARTERY COMPRESSION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 70 - 72 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 2 FEMALE(S) AND 0 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 2 CORONARY ARTERY COMPRESSION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 0 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE ARE: (B)(4). PER THE INSTRUCTIONS FOR USE (IFU), CORONARY FLOW OBSTRUCTION IS A POTENTIAL ADVERSE EVENT ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT PROCEDURE. THE IFU CAUTIONS THAT THE SAFETY AND EFFECTIVENESS HAVE NOT BEEN ESTABLISHED FOR PATIENTS WITH BULKY CALCIFIED AORTIC VALVE LEAFLETS IN CLOSE PROXIMITY TO CORONARY OSTIA. IN ADDITION, IT WARNS THAT CAUTION SHOULD BE EXERCISED IN IMPLANTING A BIOPROSTHESIS IN PATIENTS WITH CLINICALLY SIGNIFICANT CORONARY ARTERY DISEASE. CORONARY OCCLUSION CAN RESULT IN MYOCARDIAL ISCHEMIA OR INFARCTION DUE TO OBSTRUCTION OF THE CORONARY BLOOD FLOW AND MAY REQUIRE INTERVENTION (E.G., PCI). MULTIPLE PATIENT FACTORS COULD PUT THE PATIENT AT RISK FOR CORONARY FLOW OBSTRUCTION DURING THE THV PROCEDURE, INCLUDING SIGNIFICANT UNDERLYING CORONARY ARTERY DISEASE AND BULKY CALCIFICATION OF THE NATIVE LEAFLETS AND ROOT. DISPLACEMENT OF CALCIUM DEPOSITS WITH EMBOLIZATION OF DEBRIS INTO ONE OF THE ARTERIES, OR AORTIC DISSECTION WITH CONTINUITY OF THE RUPTURE INTO THE INTIMA OF ONE OF THE CORONARY OSTIA, CAN RESULT IN THIS COMPLICATION. ADDITIONALLY, MINIMAL DISTANCE BETWEEN THE NATIVE ANNULUS AND THE CORONARY OSTIA, BULKY CALCIFICATION, LONG NATIVE LEAFLETS, AND OBLITERATED CORONARY SINUSES. PROCEDURAL FACTORS SUCH AS TORN NATIVE LEAFLET DURING BALLOON VALVULOPLASTY, PLAQUE SHIFT, DEPLOYMENT OF THE BIOPROSTHETIC HEART VALVE TOO AORTIC, AND SIGNIFICANT VALVE OVER SIZING COULD ALSO CONTRIBUTE TO THIS COMPLICATION. THE EDWARDS THV TRAINING MANUALS ADVISE THE OPERATOR ON PRE-PROCEDURE ASSESSMENT OF THE AORTIC VALVE, ROOT, AND CORONARY ANATOMY. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE TRANSCATHETER HEART VALVE (THV). TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS AND PROCTORED PROCEDURES. THE PHYSICIAN IS INSTRUCTED TO EVALUATE THIS RISK EARLY IN THE PATIENT SCREENING PROCESS IN ALL PATIENTS. THE FOLLOWING FACTORS SHOULD BE CONSIDERED: DEGREE OF CALCIFICATION ON LEAFLETS, ANNULUS TO CORONARY OSTIA DISTANCE, LENGTH OF THE VALVE LEAFLET, WIDTH OF THE VALSALVA SINUSES, MOVEMENT OF THE LEAFLETS DURING BAV, PATENCY OF CORONARIES DURING BAV, AND EXPANDED HEIGHT OF THE INTENDED THV. THE TRAINING MANUALS ALSO PROVIDE THE FOLLOWING TIPS FOR DETECTING RISK FOR LEFT MAIN OCCLUSION: (1) AORTOGRAM OR TEE PRIOR TO THV IMPLANTATION TO REVEAL BULKY CALCIFIED LEAFLETS; (2) DURING PRE-DILATATION, NOTE BULKY CALCIFICATION ON VALVE MOVING TOWARDS OSTIUM ON LEFT MAIN; AND (3) CONSIDER AORTOGRAM DURING VALVULOPLASTY TO ASSESS CORONARY FLOW. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21639270&pc=NPT | 2015691-2025-01805 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-18 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Insufficient Information | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P140031 | E2016006 | 17 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 17 AORTIC READMISSION (VALVE RELATED) DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 21 - 92 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 8 FEMALE(S) AND 9 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER: E2016006 AND SUMMARIZES 17 READMISSION (VALVE RELATED) DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 143.82 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE ARE: (B)(4). VALVE RELATED RE-ADMISSION IN THE FOLLOW-UP PERIOD IS LIKELY DUE TO REOCCURRENCE OF SYMPTOMS. THIS MAY RESULT FROM REGURGITATION (CENTRAL OR PVL) OR PROGRESSION OF THE PRE-EXISTING DISEASE PROCESS AND MAY RESULT IN HEART FAILURE. CAUSES OF HEART FAILURE CAN BE MULTI-FACTORIAL. PER THE INSTRUCTION FOR USE (IFU), HEART FAILURE, VALVE REGURGITATION, AND VALVE DEGENERATION INCLUDING STENOSIS ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH AORTIC VALVE REPLACEMENT AND BIOPROSTHETIC HEART VALVES. CONGESTIVE HEART FAILURE CAN HAVE MULTIPLE ETIOLOGIES AND IS OFTEN DUE TO THE PROGRESSION OF UNDERLYING DISEASE PROCESSES, INCLUDING CORONARY ARTERY DISEASE, UNCONTROLLED HYPERTENSION, MYOCARDIAL INFARCTION, CARDIOMYOPATHY, FLUID OVERLOAD, OR VALVULAR DYSFUNCTION. RISK FACTORS THAT INCREASE A PATIENT'S RISK OF SUFFERING FROM CHF INCLUDE OBESITY, ADVANCED AGE, AND A HISTORY OF SMOKING. REGURGITATION WHICH DEVELOPS PROGRESSIVELY OVER TIME CAN BE DUE TO A NUMBER OF ISSUES INCLUDING PATIENT RELATED FACTORS OR STRUCTURAL VALVE DETERIORATION, INCLUDING CALCIFICATION, NON-CALCIFIC DEGENERATION, LEAFLET THICKENING OR FIBROSIS, OR A COMBINATION OF THESE. REGURGITATION MAY ALSO DEVELOP PROGRESSIVELY IF HOST FIBROTIC TISSUE, OR PANNUS, GROWS ONTO THE BIOPROSTHETIC VALVE. PANNUS, A CAUSE OF NONSTRUCTURAL DYSFUNCTION, MAY INTERFERE WITH FUNCTIONALITY OF THE DEVICE BY RESTRICTING THE LEAFLET MOTION LEADING TO ABNORMAL COAPTATION. THE MECHANISM BEHIND WORSENING OR LATE PVL IS NOT WELL UNDERSTOOD BUT MAY BE RELATED TO CARDIAC REMODELING. VALVE STENOSIS MAY RESULT IN SYMPTOMS SUCH AS SOB AND DECREASED EXERCISE TOLERANCE, WHICH MAY BE ACCOMPANIED BY AN INCREASED GRADIENT ACROSS THE VALVE. THIS COULD BE DUE TO EARLY CALCIFICATION OF THE LEAFLETS, HOST TISSUE OVERGROWTH, OR IN RARE CASES A NON-FUNCTIONING LEAFLET. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21639271&pc=NPT | 2015691-2025-01810 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-18 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Insufficient Information | Aortic Valve Insufficiency/ Regurgitation | P140031 | E2016006 | 4 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 4 AORTIC REINTERVENTION - AORTIC VALVE DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 51 - (B)(6) YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 2 FEMALE(S) AND 2 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER: E2016006 AND SUMMARIZES 4 REINTERVENTION - AORTIC VALVE DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 6.25 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE ARE: (B)(4). INTRA-PROCEDURAL AORTIC VALVE RE-INTERVENTION WILL TYPICALLY RESULT FROM VALVE MALPOSITION OR REGURGITATION (PVL OR CENTRAL, INCLUDING LEAFLET RESTRICTION). THESE CONDITIONS ARE KNOWN POTENTIAL RISKS ASSOCIATED WITH THE USE OF THE THV, DELIVERY SYSTEM, AND/OR ACCESSORIES. THERE ARE MULTIPLE PATIENT AND PROCEDURAL FACTORS THAT ALONE OR IN COMBINATION CAN CAUSE OR CONTRIBUTE TO AORTIC MALPOSITION, INCLUDING IMPROPER POSITIONING PRIOR TO DEPLOYMENT, POOR IMAGE INTENSIFIER ANGLE, POOR COAXIAL ALIGNMENT OF THE VALVE AND DELIVERY SYSTEM, SEVERE SEPTAL HYPERTROPHY, MINIMALLY OR BULKY/SEVERELY CALCIFIED AORTIC LEAFLETS, PRESERVED EJECTION FRACTION, SIGNIFICANT MITRAL ANNULAR CALCIFICATION (MAC), LOSS OF PACING CAPTURE, RAPID DEPLOYMENT, RELEASE OF STORED TENSION DURING DEPLOYMENT, AND MOVEMENT OF THE DELIVERY SYSTEM BY THE OPERATOR. DEPLOYMENT OF THE SAPIEN VALVE TOO AORTIC HAS THE POTENTIAL TO CONTRIBUTE TO SUBOPTIMAL COAPTATION OF THE SAPIEN VALVE LEAFLETS AND CAUSE CENTRAL AORTIC INSUFFICIENCY; IT CAN OBSTRUCT THE CORONARY OSTIA; AND LEAD TO EMBOLIZATION OF THE PROSTHESIS INTO THE ASCENDING AORTA. THE THV TRAINING MANUALS INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE THV. TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS AND PROCTORED PROCEDURES. THE CORRECT ALIGNMENT AND POSITIONING OF THE DEVICE AT THE POINT OF DEPLOYMENT IS EMPHASIZED AS A KEY FACTOR TO THE PLACEMENT AND FIXATION OF THE DEVICE. OPERATORS ARE ALSO INSTRUCTED TO USE FLUOROSCOPY AS THE PRIMARY METHOD OF VISUALIZATION FOR POSITIONING AND DEPLOYMENT. IN PATIENTS WITH HIGH-RISK ANATOMICAL FEATURES FOR AORTIC MALPOSITION (I.E., MINIMAL LEAFLET CALCIFICATION, SEVERE SEPTAL HYPERTROPHY), BAV MAY PROVIDE INDICATION OF POTENTIAL BALLOON MOVEMENT DURING VALVE DEPLOYMENT. AORTIC VALVE RE-INTERVENTION IN THE FOLLOW-UP PERIOD (IN THE ABSENCE OF PROSTHETIC CARDIAC VALVE THROMBOSIS) WILL TYPICALLY RESULT FROM ON-GOING OR WORSENING REGURGITATION, VALVE DEGENERATION RELATED TO THE FORMATION OF CALCIFICATION OR PANNUS, OR VALVE MIGRATION. THESE EVENTS ARE IDENTIFIED IN THE PRODUCT INSTRUCTIONS FOR USE (IFU) AS POTENTIAL RISKS ASSOCIATED WITH THE USE OF THE THV. REGURGITATION WHICH DEVELOPS PROGRESSIVELY OVER TIME CAN BE DUE TO A NUMBER OF ISSUES INCLUDING PATIENT RELATED FACTORS OR STRUCTURAL VALVE DETERIORATION, INCLUDING CALCIFICATION, NON-CALCIFIC DEGENERATION, LEAFLET THICKENING OR FIBROSIS, OR A COMBINATION OF THESE. REGURGITATION MAY ALSO DEVELOP PROGRESSIVELY IF HOST FIBROTIC TISSUE, OR PANNUS, GROWS ONTO THE BIOPROSTHETIC VALVE. PANNUS, A CAUSE OF NONSTRUCTURAL DYSFUNCTION, MAY INTERFERE WITH FUNCTIONALITY OF THE DEVICE BY RESTRICTING THE LEAFLET MOTION LEADING TO ABNORMAL COAPTATION. CALCIFICATION IS A WELL-RECOGNIZED FAILURE MODE OF BIOPROSTHETIC VALVES. THE MECHANISM OF CALCIFICATION OF BIOMATERIALS IS NOT COMPLETELY UNDERSTOOD BUT IS PROBABLY RELATED TO AN INABILITY OF THE NON-VIABLE CELLS TO MAINTAIN THEIR NORMALLY LOW INTRACELLULAR CONCENTRATION OF CALCIUM. THE MECHANISM BEHIND WORSENING OR LATE PVL IS NOT WELL UNDERSTOOD BUT MAY BE RELATED TO CARDIAC REMODELING. ACCORDING TO LITERATURE REVIEW, VALVE MIGRATION RESULTS WHEN FORCES ACTING ON THE TRANSCATHETER HEART VALVE (THV) OVERCOME THE STRENGTH OF ATTACHMENT OF THE VALVE TO THE AORTIC WALL. STENT VALVES ARE SUBJECTED TO ANTEGRADE EJECTION FORCES DURING SYSTOLE. LESS-THAN-SEVERE AND NON-UNIFORMLY DISTRIBUTED CALCIFICATION OF THE NATIVE LEAFLETS, INCORRECT BIOPROSTHETIC VALVE SIZING, AND INCOMPLETE FRAME EXPANSION CAN CONTRIBUTE TO VALVE MIGRATION. ADDITIONALLY, RESIDUAL OVERHANGING LEAFLETS CAN EXERT DOWNWARDS FORCE DURING DIASTOLE, CAUSING MIGRATION OF THE THV TOWARDS THE LEFT VENTRICLE. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21639272&pc=DYB | 2015691-2025-01818 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-18 04:00 | DYB | EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | K200258 | E2016006 | 5 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 5 AORTIC BLEEDING - ACCESS SITE DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 68 - (B)(6) YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 3 FEMALE(S) AND 2 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 5 BLEEDING - ACCESS SITE DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 1.4 DAY(S). DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). THE INSTRUCTIONS FOR USE (IFU) LIST HEMORRHAGE REQUIRING TRANSFUSION OR INTERVENTION AS A POTENTIAL RISK ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING POTENTIAL ACCESS COMPLICATIONS ASSOCIATED WITH STANDARD CARDIAC CATHETERIZATION. THE MAJORITY OF TRANSFEMORAL ACCESS RELATED BLEEDING COMPLICATIONS ARE RELATED TO DISEASED ILEO-FEMORAL VESSELS AND/OR PROCEDURAL TECHNIQUE DURING THE INSERTION OR REMOVAL OF THE SHEATH AND/OR DILATORS. THERE ARE CASES WHERE THE BLEEDING IS SIGNIFICANT AND MORE COMPLEX INTERVENTION OR TRANSFUSION IS REQUIRED TO TREAT/PREVENT A PERMANENT INJURY FROM OCCURRING. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE VALVE CAN BE DELIVERED TRANSFEMORALLY. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE; HOWEVER, A TRANSFUSION OF BLOOD WAS REPORTED IN THE REGISTRY. BY VARC DEFINITION, THE MEET CRITERIA FOR MAJOR VASCULAR COMPLICATION, THEREFORE THE EVENT IS BEING REPORTED. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21639273&pc=NPT | 2015691-2025-01822 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-18 04:00 | NPT | COMMANDER DELIVERY SYSTEM UNKNOWN | Adverse Event Without Identified Device or Use Problem |  | P140031 | E2016006 | 4 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 4 AORTIC AORTIC DISSECTION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 66 - 100 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 3 FEMALE(S) AND 1 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 4 AORTIC DISSECTION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 0 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS COMMANDER DELIVERY SYSTEM ARE: (B)(4). PER THE INSTRUCTIONS FOR USE (IFU) CARDIOVASCULAR INJURY, SUCH AS PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, MYOCARDIUM, OR VALVULAR STRUCTURES, IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT PROCEDURE. ASCENDING AORTIC DISSECTION MAY OCCUR WHEN MULTIPLE ATTEMPTS ARE MADE TO CROSS THE STENOTIC NATIVE VALVE, AND/OR WHEN EXCESSIVE FORCE IS USED. AORTIC DISSECTION IN A DISTAL SECTION OF THE AORTA CAN OCCUR WITH DEVICE MANIPULATION AND NAVIGATION THROUGH ANATOMIC TORTUOSITY. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE TRANSCATHETER HEART VALVE (THV). THE THV TRAINING MANUALS PROVIDE GUIDANCE TO FACILITATE SAFE CROSSING OF THE NATIVE VALVE, INCLUDING CAMERA PROJECTIONS, HANDLING DURING ADVANCEMENT, AND TROUBLESHOOTING TECHNIQUES IF DIFFICULTY IS ENCOUNTERED. AS STATED, EXCESSIVE FORCE SHOULD NOT BE USED WHEN THE DEVICE HAS DIFFICULTY CROSSING THE STENOTIC VALVE. ADDING TENSION TO THE WIRE, PULLING BACK THE SYSTEM TO RE-ORIENT THE VALVE, AS NEEDED, AND TORQUEING OF THE FLEX CATHETER MAY BE HELPFUL IN SOLVING THE PROBLEM. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21639274&pc=DYB | 2015691-2025-01819 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-18 04:00 | DYB | EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | Adverse Event Without Identified Device or Use Problem | Hematoma | K200258 | E2016006 | 5 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 5 AORTIC BLEEDING - HEMATOMA AT ACCESS SITE DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 59 - 81 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 3 FEMALE(S) AND 2 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 5 BLEEDING - HEMATOMA AT ACCESS SITE DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 1.2 DAY(S). DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). THE INSTRUCTIONS FOR USE (IFU) LIST HEMORRHAGE REQUIRING TRANSFUSION OR INTERVENTION AS A POTENTIAL RISK ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING POTENTIAL ACCESS COMPLICATIONS ASSOCIATED WITH STANDARD CARDIAC CATHETERIZATION. THE MAJORITY OF TRANSFEMORAL ACCESS RELATED BLEEDING COMPLICATIONS ARE RELATED TO DISEASED ILEO-FEMORAL VESSELS AND/OR PROCEDURAL TECHNIQUE DURING THE INSERTION OR REMOVAL OF THE SHEATH AND/OR DILATORS. THERE ARE CASES WHERE THE BLEEDING IS SIGNIFICANT AND MORE COMPLEX INTERVENTION OR TRANSFUSION IS REQUIRED TO TREAT/PREVENT A PERMANENT INJURY FROM OCCURRING. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE VALVE CAN BE DELIVERED TRANSFEMORALLY. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. HOWEVER, A TRANSFUSION OF BLOOD WAS REPORTED IN THE REGISTRY. BY VARC DEFINITION, THE MEET CRITERIA FOR MAJOR VASCULAR COMPLICATION THEREFORE, THE EVENT IS BEING REPORTED. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21639275&pc=NPT | 2015691-2025-01815 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-18 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Thrombosis/Thrombus | P140031 | E2016006 | 1 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 1 AORTIC DEVICE THROMBOSIS DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 81 - 81 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 0 FEMALE(S) AND 1 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER: E2016006 AND SUMMARIZES 1 DEVICE THROMBOSIS DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 0 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE ARE: (B)(4). PER THE INSTRUCTIONS FOR USE (IFU), VALVE THROMBOSIS IS A POTENTIAL RISK ASSOCIATED WITH THE USE OF THE TRANSCATHETER HEART VALVES (THV). THROMBOSIS IS A RARE AND WELL-RECOGNIZED COMPLICATION ASSOCIATED WITH THE USE OF BIOPROSTHETIC HEART VALVES. VALVE THROMBOSIS IS THE FORMATION OF SIGNIFICANT BLOOD CLOTS FORMING ON THE VALVE. POTENTIAL PATIENT RISK FACTORS SUCH AS ATRIAL FIBRILLATION, SYSTEMIC DISEASE (E.G., SYSTEMIC LUPUS ERYTHEMATOSUS, INFLAMMATION, AND DAMAGE TO VARIOUS BODY TISSUES, INCLUDING JOINTS, SKIN, KIDNEYS, HEART, LUNGS, BLOOD VESSELS, AND BRAIN), THE VALVULAR DISEASE ITSELF, AND REDUCED CARDIAC EJECTION FRACTION CAN CONTRIBUTE TO INCREASED RISK OF THROMBUS/THROMBOSIS. SUB-OPTIMAL ANTICOAGULATION DURING THE PROCEDURE AND UNDERLYING PATIENT CONDITIONS CAN ALSO RESULT IN INCREASED THROMBOGENICITY. THESE PATIENTS ARE ANTICOAGULATED FOR THE PROCEDURE AND INTERVENTIONAL BEST PRACTICES MANDATE METICULOUS PREPARATION, ASPIRATION, AND FLUSHING OF THE DEVICES TO PREVENT AND/OR REMOVE CLOT(S). SHORT-TERM ANTICOAGULATION THERAPY MAY ALSO BE NECESSARY AFTER VALVE REPAIR, ANTICOAGULATION IS PRESCRIBED PER INSTITUTIONAL GUIDELINES. INTRA-PROCEDURAL INTRA-CARDIAC THROMBUS AND POST PROCEDURE OR LATE VALVE THROMBOSIS ARE COMPLEX PROCESSES TRIGGERED BY THE INTERACTION BETWEEN THE HOST AND THE DEVICE WHICH ARE HIGHLY VARIABLE AMONG PATIENTS. IT IS THE NATURAL TENDENCY OF THE BODY TO FORM A CLOT ON FOREIGN OBJECTS IN THE VASCULAR SPACE AND A DEFINITIVE ROOT CAUSE CANNOT ALWAYS BE CONFIRMED. THE DEVICE TRAINING MANUALS INSTRUCT THE OPERATOR TO CONSIDER ALL PROCEDURAL AND ANATOMICAL FACTORS. EDWARDS EXTENSIVELY TRAINS PHYSICIANS BEFORE THEY ARE QUALIFIED TO USE THE DEVICE SYSTEM. TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES. OPERATORS ARE ALSO INSTRUCTED TO USE FLUOROSCOPY IN CONJUNCTION WITH ECHOCARDIOGRAPHY FOR OPTIMAL VISUALIZATION DURING POSITIONING AND DEPLOYMENT AND CAUTIONS INCLUDE THE MAINTENANCE OF THE PATIENT'S ANTICOAGULATION STATUS (ACT AT >250 SECONDS) DURING THE PROCEDURE. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21639276&pc=DYB | 2015691-2025-01816 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-18 04:00 | DYB | EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | Adverse Event Without Identified Device or Use Problem | Unspecified Vascular Problem | K200258 | E2016006 | 31 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 31 AORTIC VASCULAR COMPLICATION - MAJOR DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 21 - 94 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 17 FEMALE(S) AND 14 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 31 VASCULAR COMPLICATION - MAJOR DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 20.97 DAY(S). DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4).THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). ACCORDING TO THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR COMPLICATIONS, INCLUDING PERFORATION OR DISSECTION OF VESSELS WHICH MAY REQUIRE INTERVENTION, ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE TRANSFEMORAL TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, VASCULAR COMPLICATIONS ARE A WELL-RECOGNIZED COMPLICATION OF THE TRANSFEMORAL TRANSCATHETER VALVE REPLACEMENT PROCEDURE IN THIS ELDERLY POPULATION WITH MULTIPLE CO-MORBIDITIES. EDWARDS HAS REVIEWED MANY REPORTS, INCLUDING SCREENING DATA RECORDS AND SOURCE DOCUMENTATION OF VASCULAR COMPLICATIONS, AND HAS FOUND THAT THE ROOT CAUSE IS TYPICALLY RELATED TO A COMBINATION OF VESSEL SIZE, TORTUOSITY, AND CALCIFICATIONS. ALTHOUGH THE INCIDENCE IS DECREASING WITH SMALLER SHEATH/DELIVERY SYSTEM SIZES AND PHYSICIAN EXPERIENCE, THERE WILL CONTINUE TO BE CASES IN WHICH VASCULAR COMPLICATIONS WILL OCCUR. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE TRAINING MANUAL INSTRUCTS THE OPERATOR ON PROPER SHEATH INSERTION AND WITHDRAWAL TECHNIQUES, INCLUDING PRE-DILATING THE VESSEL WITH THE EDWARDS DILATORS, AS NEEDED. IT ALSO NOTES THAT CALCIFICATION MAY REDUCE LUMEN DIAMETER AND LIMIT OR PREVENT TRANSFEMORAL PASSAGE OF THE DEVICES. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE THV CAN BE DELIVERED TRANSFEMORALLY. ASSESSMENT OF LOCATION AND AMOUNT OF CIRCUMFERENTIAL CALCIUM WILL AID IN DETERMINING AREAS OF REDUCED VESSEL DIAMETERS. THE OPERATORS ARE TRAINED TO MEASURE MINIMUM VESSEL DIAMETER TAKING CALCIUM INTO ACCOUNT. THE PHYSICIAN TRAINING MANUAL ALSO LISTS THE MINIMUM RECOMMENDED VESSEL SIZE FOR EACH SIZE DEVICE. DESPITE THE BEST SCREENING TOOLS, A SMALL PERCENTAGE OF PATIENTS WILL HAVE FEMORAL/ILIAC VESSELS THAT ARE NOT AMENABLE TO THE TRANS-FEMORAL APPROACH OR WHERE INCREASED RESISTANCE IS ENCOUNTERED DURING INSERTION OF DEVICES. IN MANY CASES, THE VESSEL MINIMUM LUMINAL DIAMETER (MLD) MAY BE BORDERLINE OR BELOW THE INDICATED SIZE. IN ADDITION, SIGNIFICANT CALCIFICATION AND/OR TORTUOSITY, NOT ALWAYS APPRECIABLE ON IMAGING, COULD BE CONTRIBUTING FACTORS TO THE EVENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21639277&pc=DYB | 2015691-2025-01817 | 2025-01-01 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-03-18 04:00 | DYB | EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | Insufficient Information | Unspecified Vascular Problem | K200258 | E2016006 | 15 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 1 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 15 AORTIC VASCULAR SURGERY OR INTERVENTION - UNPLANNED DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 59 - 94 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 6 FEMALE(S) AND 9 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 15 VASCULAR SURGERY OR INTERVENTION - UNPLANNED DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 2.6 DAY(S). DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). ACCORDING TO THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR COMPLICATIONS, INCLUDING PERFORATION OR DISSECTION OF VESSELS WHICH MAY REQUIRE INTERVENTION, ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE TRANSFEMORAL TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, VASCULAR COMPLICATIONS ARE A WELL-RECOGNIZED COMPLICATION OF THE TRANSFEMORAL TRANSCATHETER VALVE REPLACEMENT PROCEDURE IN THIS ELDERLY POPULATION WITH MULTIPLE CO-MORBIDITIES. EDWARDS HAS REVIEWED MANY REPORTS, INCLUDING SCREENING DATA RECORDS AND SOURCE DOCUMENTATION OF VASCULAR COMPLICATIONS, AND HAS FOUND THAT THE ROOT CAUSE IS TYPICALLY RELATED TO A COMBINATION OF VESSEL SIZE, TORTUOSITY, AND CALCIFICATIONS. ALTHOUGH THE INCIDENCE IS DECREASING WITH SMALLER SHEATH/DELIVERY SYSTEM SIZES AND PHYSICIAN EXPERIENCE, THERE WILL CONTINUE TO BE CASES IN WHICH VASCULAR COMPLICATIONS WILL OCCUR. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE TRAINING MANUAL INSTRUCTS THE OPERATOR ON PROPER SHEATH INSERTION AND WITHDRAWAL TECHNIQUES, INCLUDING PRE-DILATING THE VESSEL WITH THE EDWARDS DILATORS, AS NEEDED. IT ALSO NOTES THAT CALCIFICATION MAY REDUCE LUMEN DIAMETER AND LIMIT OR PREVENT TRANSFEMORAL PASSAGE OF THE DEVICES. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE SAPIEN VALVE CAN BE DELIVERED TRANSFEMORALLY. ASSESSMENT OF LOCATION AND AMOUNT OF CIRCUMFERENTIAL CALCIUM WILL AID IN DETERMINING AREAS OF REDUCED VESSEL DIAMETERS. THE OPERATORS ARE TRAINED TO MEASURE MINIMUM VESSEL DIAMETER TAKING CALCIUM INTO ACCOUNT. THE PHYSICIAN TRAINING MANUAL ALSO LISTS THE MINIMUM RECOMMENDED VESSEL SIZE FOR EACH SIZE DEVICE. DESPITE THE BEST SCREENING TOOLS, A SMALL PERCENTAGE OF PATIENTS WILL HAVE FEMORAL/ILIAC VESSELS THAT ARE NOT AMENABLE TO THE TRANS-FEMORAL APPROACH OR WHERE INCREASED RESISTANCE IS ENCOUNTERED DURING INSERTION OF DEVICES. IN MANY CASES, THE VESSEL MINIMUM LUMINAL DIAMETER (MLD) MAY BE BORDERLINE OR BELOW THE INDICATED SIZE. IN ADDITION, SIGNIFICANT CALCIFICATION AND/OR TORTUOSITY, NOT ALWAYS APPRECIABLE ON IMAGING, COULD BE CONTRIBUTING FACTORS TO THE EVENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21615832&pc=MCM | 6000034-2025-00938 |  | Death | COCHLEAR LIMITED | 2025-03-17 04:00 | MCM | NUCLEUS CI622 COCHLEAR IMPLANT WITH SLIM STRAIGHT ELECTRODE | Appropriate Term/Code Not Available | Abscess; Pain |  |  | 1 | Event Description: PER THE CLINIC, THE PATIENT EXPERIENCED POST-AURAL PAIN AND ABSCESS. THE PATIENT WAS TREATED WITH IV ANTIBIOTICS FOR A COUPLE OF WEEKS; HOWEVER, THE ISSUE DID NOT RESOLVE. SUBSEQUENTLY, THE DEVICE WAS EXPLANTED ON (B)(6) 2025.Event Description: CORRECTION: THE DEATH OUTCOME MARKED IN B2 OF THE PREVIOUS OR INITIAL MDR SUBMITTED ON MARCH 17, 2025, WAS FILED INADVERTENTLY. THE PATIENT ONLY UNDERWENT AN EXPLANTATION AND IS DOING WELL. DURING EXPLANT SURGERY ON (B)(6) 2025, IT WAS FOUND THAT CHOLESTEATOMA HAD RECURRED AND HAD ENGULFED THE ELECTRODE CABLES IN THE MASTOID CAVITY, THUS, LEADING TO THE DECISION OF EXPLANTATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21615943&pc=DYE | 2015691-2025-01962 | 2025-02-10 05:00 | Death | EDWARDS LIFESCIENCES | 2025-03-17 04:00 | DYE | CARPENTIER-EDWARDS PERIMOUNT MAGNA MITRAL EASE PERICARDIAL BIOPROSTHESIS | Entrapment of Device; Mechanical Jam | Insufficient Information | P860057 |  | 1 | Event Description: EDWARDS RECEIVED NOTIFICATION FROM COLOMBIA THAT A VALVE MODEL 7300TFX27 IMPLANTED IN THE MITRAL POSITION WAS EXPLANTED AT IMPLANT DUE TO LEAFLETS NOT OPENING WELL AND ADEQUATELY. REPORTEDLY, THE VALVE WAS PLACED CORRECTLY USING FULL STERNOTOMY APPROACH. THERE WERE NO DIFFICULTIES USING THE TRICENTRIX HOLDER, WHICH WAS USED PER IFU. HOWEVER, IN THE INTRAOPERATIVE ECHO IT WAS OBSERVED THAT TWO (2) OF THE LEAFLETS WERE NOT OPENING CORRECTLY, AS IF THEY WERE DAMAGED. THUS, DECISION WAS MADE TO EXCHANGE THE DEVICE AND ANOTHER BIOPROSTHETIC VALVE WAS IMPLANTED IN REPLACEMENT. REPORTEDLY, THE PATIENT DIED TWO DAYS AFTER PROCEDURE DUE TO VENTRICULAR DISRUPTION. Manufacturer Narrative: ADDITIONAL INFORMATION: H6: TYPE OF INVESTIGATION. H3: PRODUCT EVALUATION: CUSTOMER REPORT THAT "LEAFLETS NOT OPENING WELL AND ADEQUATELY" WAS UNABLE TO BE CONFIRMED. AS RECEIVED INDENTATIONS WERE OBSERVED ON THE FREE MARGINS OF LEAFLETS 2 AND 3 NEAR COMMISSURE 3. INDENTATIONS WERE WIDER THAN THE TYPICAL SUTURE LOOP. CLOTH IMPRINT MARKS WERE OBSERVED ON LEAFLET 3 NEAR COMMISSURE 3. SEWING RING CLOTH WAS CUT IN MULTIPLE AREAS AROUND THE VALVE AND EXPOSED THE SEWING RING SILICONE. X-RAY DEMONSTRATED WIREFORM INTACT. MULTIPLE SUTURE HOLES WERE VISIBLE AROUND THE SEWING RING. THE INVESTIGATION IS STILL IN PROGRESS; THEREFORE, A CONCLUSION HAS YET TO BE ESTABLISHED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY UPON INVESTIGATION COMPLETION. Manufacturer Narrative: CORRECTED, UPDATED OR ADDED INFORMATION: H6: DEVICE CODE, TYPE OF INVESTIGATION, INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS. H11: ADDITIONAL MANUFACTURER NARRATIVE: PER R&D FUNCTIONAL EVALUATION. CLOSE UP INSPECTION SHOWS LEAFLET INDENTATION ON LEAFLETS 2 & 3 NEAR COMMISSURE 3. THESE FEATURES ARE CONSISTENT WITH A FORCE PRESSING ON THE LEAFLET FREE EDGES IN THE RETROGRADE DIRECTION, USUALLY DUE TO SUTURE ENTRAPMENT BY A SUTURE WITH OR WITHOUT PLEDGETS, INTERFERENCE WITH THE PRESERVED VALVULAR SUBVALVULAR APPARATUS SUCH AS CHORDAE, OR SURGICAL TOOLS. THE RETURNED VALVE COMPETENCY WAS EVALUATED USING EDWARDS STANDARDIZED IN VITRO TESTS FOR PULSATILE FLOW AND STEADY BACKPRESSURE LEAKAGE UNDER SIMULATED NORMAL PHYSIOLOGICAL CONDITIONS. ALL THREE LEAFLETS PARTIALLY OPENED, REPRODUCING PARTIALLY THE REPORTED LEAFLET NOT OPENING WELL . ALL THREE LEAFLETS CLOSED COMPLETELY. PARTIAL OPENING OF THE LEAFLETS IS MOST LIKELY DUE TO DEHYDRATION, BLOOD EXPOSURE, AND SUBSEQUENT STORAGE IN FORMALIN CAUSING THE LEAFLETS TO STIFFEN. DAMAGE TO LEAFLETS 2 AND 3 CAN ALSO CONTRIBUTE TO THE PARTIAL OPENING OF LEAFLETS 2 AND 3. HOWEVER, VALVE STILL MET ALL REQUIREMENTS IN TERMS OF EFECTIVE ORIFICE AREA (EOA) AND REGURGITANT FRACTION. DURING IMPLANTATION OF THE VALVE, ANY IMPINGEMENT IN THE LEAFLET CAN CAUSE IMPROPER LEAFLET MOTION AND COAPTATION, RESULTING IN VALVE FAILURE. THE RESULTS FROM IN VITRO TESTING MAY BE DIFFERENT FROM THOSE OBTAINED IN VIVO DUE TO HEMODYNAMIC AND ENVIRONMENTAL DIFFERENCES, WHICH WERE NOT PRESENT WHEN THE VALVE WAS RETURNED TO EDWARDS. A DEVICE HISTORY RECORD (DHR) WAS PERFORMED, AND NO RELEVANT NON-CONFORMANCES WERE IDENTIFIED. DURING THE COMPLAINT HISTORY REVIEW, ALL APPLICABLE COMPLAINTS WERE CAUSED BY SIMILAR PROCEDURAL FACTORS. PER EDWARD'S MANUFACTURING PROCESS, LEAFLET SORTING, VALVE FLOW AND COAPTATION TESTING, AND VISUAL VALVE INSPECTIONS WERE COMPLETED TO ENSURE LEAFLET TISSUE QUALITY, PRESSURE GRADIENT VALUES, FLOW/COAPTATION, OUTFLOW/INFLOW GENERAL APPEARANCE AND WIREFORM/STITCHING QUALITY ALL MET SPECIFICATIONS. THEREFORE, IT IS NOT LIKELY THAT THE INDENTATIONS IN THE LEAFLETS WERE CAUSED BY THE MANUFACTURING PROCESS. BASED ON THE INFORMATION AVAILABLE, THE COMPLAINT FOR LEAFLET IMMOBILITY RESTRICTED MOTION IS ABLE TO BE CONFIRMED AND MOST LIKELY RESULTED FROM PROCEDURAL FACTORS, INCLUDING EXTERNAL FORCE PRESSING ON THE LEAFLET FREE EDGES DUE TO SUTURE ENTRAPMENT BY A SUTURE WITH OR WITHOUT PLEDGETS, INTERFERENCE WITH THE PRESERVED VALVULAR/SUBVALVULAR APPARATUS SUCH AS CHORDAE, OR SURGICAL TOOLS. AN EDWARDS MANUFACTURING DEFECT HAS NOT BEEN CONFIRMED. Manufacturer Narrative: H11: ADDITIONAL MANUFACTURER NARRATIVE: DEVICE EVALUATION ANTICIPATED, BUT NOT YET BEGUN. THE INVESTIGATION IS STILL IN PROGRESS; THEREFORE, A CONCLUSION HAS YET TO BE ESTABLISHED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY UPON INVESTIGATION COMPLETION. EDWARDS WILL CONTINUE TO REVIEW AND MONITOR ALL REPORTED EVENTS. TRENDS ARE MONITORED ON A MONTHLY BASIS AND IF ACTION IS REQUIRED, APPROPRIATE INVESTIGATION WILL BE PERFORMED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21617274&pc=MIH | 2017233-2025-05942 | 2025-02-14 05:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-03-17 04:00 | MIH | GORE® TAG® CONFORMABLE THORACIC ENDOPROSTHESIS | Contamination | Chest Pain; Unspecified Infection; Inflammation | P040043 |  | 1 | Event Description: IT WAS REPORTED TO GORE, AT AN UNKNOWN DATE THIS PATIENT UNDERWENT THE THORACIC ENDOVASCULAR ANEURYSM REPAIR (TEVAR) PROCEDURE WITH A GORE® TAG® THORACIC BRANCH ENDOPROSTHESIS AT NORTHERN GENERAL HOSPITAL IN SHEFFIELD. DUE TO AN INFECTION, AN EXPLANT TOOK PLACE MID-END FEBRUARY (EXACT DATE UNKNOWN) AT ST BARTHOLOMEW'S HOSPITAL. DURING THIS PROCEDURE, THE SURGEON PROCEEDED WITH A FROZEN ELEPHANT TRUNK GRAFT. THE PATIENT UNFORTUNATELY DID NOT SURVIVE THE EXPLANT AND ADJUNCTIVE PROCEDURE. THE SURGEON REPORTED THAT THE CAUSE OF DEATH MIGHT BE DUE TO THROMBOTIC SHOWER.Event Description: IT WAS REPORTED TO GORE, ON (B)(6) 2024, A PATIENT PRESENTED WITH A THORACOABDOMINAL ANEURYSM UNDERWENT A THORACIC ENDOVASCULAR ANEURYSM REPAIR (TEVAR) PROCEDURE WITH A GORE® TAG® THORACIC BRANCH ENDOPROSTHESIS (TBE, SERIAL TAC: (B)(6)) AT (B)(6) HOSPITAL IN (B)(6) AND WAS IMPLANTED IN ZONE 2. A GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM (TGMR313120E, SERIAL: (B)(6)) WAS IMPLANTED DISTALLY. DURING THE PROCEDURE, A PERCUTANEOUS ACCESS WAS NOT POSSIBLE AND A RIGHT FEMORAL CUTDOWN AND SURGICAL REPAIR WAS REQUIRED, THE INVOLVED DEVICE WAS A NON-GORE DEVICE FROM ABBOTT (PERCLOSE PROGLIDE¿). THE PATIENT REQUIRED A TWO STAGE REPAIR; TEVAR AND IN SECOND STAGE AN INNER BRANCH ENDOVASCULAR AORTIC REPAIR (IBEVAR) WITH ARTRIVION (NON-GORE) DEVICE, WHICH WAS PLANNED IN 6-8 WEEKS¿ TIME. POST OPERATION, THE PATIENT HAD CHEST INFECTION TREATED WITH ANTIBIOTICS AND RECOVERED WELL AND WAS DISCHARGED. HOWEVER, 4 WEEKS LATER, THE PATIENT PRESENTED WITH SEVERE CHEST PAIN. CTA SHOWED SLIGHT INCREASE IN THE DISTAL THORACIC ANEURYSM WITH SIGNS OF SYMPTOMATIC THORACIC ANEURYSM. BLOOD RESULTS WERE NORMAL. AN IBEVAR WAS PERFORMED ON (B)(6) 2024 AS EMERGENCY. THE BRAND OF THE IMPLANTED DEVICE WAS ARTRIVION. TWO DAYS LATER, THE PATIENT DEVELOPED MEDIASTINITIS TREATED WITH ANTIBIOTICS BUT THE INFECTION CONTINUED PROGRESSING AND WAS NOT RESPONDING WITH ANTIBIOTICS. THE SURGEON STATED THAT THIS MEDIASTINITIS IS THE CAUSE OF THE GRAFT INFECTION. THE CAUSE OF THE MEDIASTINITIS IS UNKNOWN. DUE TO THIS INFECTION, AN EXPLANT TOOK PLACE MID-END (B)(6) 2025 (EXACT DATE UNKNOWN) AT (B)(6) HOSPITAL. DURING THIS PROCEDURE, THE SURGEON PROCEEDED WITH A FROZEN ELEPHANT TRUNK (FET) GRAFT. THE PATIENT UNFORTUNATELY DID NOT SURVIVE THE EXPLANT AND ADJUNCTIVE PROCEDURE. THE EXPLANT SURGEON REPORTED THAT THE CAUSE OF DEATH MIGHT BE DUE TO THROMBOTIC SHOWER. IT WAS ADDITIONALLY REPORTED, THE PATIENT HAD UNDERLYING CONDITION(S) THAT MAY HAVE TRIGGERED THE INFECTION AT LAST PROCEDURE FROM 2ND STAGE, SUCH HISTORY OF PERIODIC FEVER SYNDROME AND EMPHYSEMA. Manufacturer Narrative: UPDATED G3. B5: MORE INFORMATION IS DESCRIBED AND EVENT DESCRIPTION UPDATED. B7: UPDATED WITH NEW INFORMATION. B3: THE EVENT DATE REMAINS UNKNOWN WITH THE INFECTION ONSET DATE. NO CONCRETE EVENT DATE, WE SELECTED THE DATE WHEN THE FIRST TIME GORE BECAME AWARE. LATER, WE COULD DETERMINE MORE INFORMATION TO THE INFECTION SO THAT THE EXACT ONSET DATE FOR THE GRAFT WAS EARLIER. HOWEVER, IT WAS KEPT THE (B)(6) 2025 AS NO EXACT DATE OF INCIDENT. THERE ARE POSSIBLE EARLIER EVENTS THAT HAVE BEEN OCCURRED AND POSSIBLY PRIOR THE IBEVAR IMPLANT DATE (BASED ON THE PATIENT HISTORY). THUS, IT CANNOT BE DETERMINED WHEN WAS A FIRST INFECTION OBSERVED, THEN WHEN MEDIASTINITIS HAS BEEN DIAGNOSED AND THEN STARTED THE REINTERVENTION BY PHYSICIAN. B2: DATE OF PATIENT DEATH, IT REMAINS UNKNOWN TO GORE WHEN THE PATIENT DIED. THE CAUSE OF DEATH WAS ONLY REPORTED AND THE DATE OF DEATH SEEMED TO BE IN (B)(6) 2025. D6A: DATE OF IMPLANTATION, DATE ADDED. H6, CODING: - INVESTIGATION FINDINGS, CODE C19 ADDED. C21 NO LONGER APPLICABLE. - INVESTIGATION CONCLUSIONS: CODE D15, D12 ADDED. D16 NO LONGER APPLICABLE. - HEATH EFFECT-CLINICAL CODE: CODE E233001 AND E2326 ADDED. - HEALTH EFFECT-IMPACT CODE: CODE F2301 ADDED. - TYPE OF INVESTIGATION: CODE B14 ADDED, CODE B11 ADDED. THE INVESTIGATION WAS COMPLETED WITH THAT RESULT: A REVIEW OF THE MANUFACTURING & STERILIZATION RECORDS INDICATED THE LOTS MET ALL PRE-RELEASE SPECIFICATIONS. NEITHER IMAGES ENABLING DIRECT ASSESSMENT OF PRODUCT PERFORMANCE NOR THE PRODUCT ITSELF WERE RETURNED FOR EVALUATION. BASED ON THE INCIDENT DESCRIPTION AND THE SUBSEQUENT INVESTIGATION, NO FURTHER INFORMATION WAS PROVIDED TO GORE, WE ARE UNABLE TO DETERMINE THE CAUSE OF THIS INCIDENT AND ASSIGN A ROOT CAUSE. PER THE GORE® TAG® THORACIC BRANCH ENDOPROSTHESIS INSTRUCTIONS FOR USE (IFU), POSSIBLE ADVERSE EVENTS AND COMPLICATIONS THAT MAY OCCUR WITH THE USE OF GORE® TAG® THORACIC BRANCH ENDOPROSTHESIS INCLUDE, BUT ARE NOT LIMITED TO: DEATH, SURGICAL CONVERSION, INFECTION. Manufacturer Narrative: A1: NO PATIENT SPECIFIC DETAILS HAVE BEEN PROVIDED. THEREFORE, THE PATIENT IDENTIFIER REFLECT THE W. L. GORE INTERNAL CASE NUMBER. B3: DATE OF EVENT WAS CHOSEN AS 14-FEB-2025 AS THE EXACT DATE FOR EXPLANT WAS NOT PROVIDED YET. H3: THE STATUS OF THE EXPLANTED DEVICE IS UNKNOWN. THE SERIAL NUMBER IS ALSO UNKNOWN. H6, FURTHER INTERVIEWS ARE NECESSARY AND TRY TO GET FURTHER DETAILS TO THE EVENT. D10: CONCOMITANT MEDICAL PRODUCTS: (RELEVANT ASSOCIATED DEVICES USED WITH THE DEVICE BEING REPORTED ON): GORE® TAG® THORACIC BRANCH ENDOPROSTHESIS (AORTIC COMPONENT). W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21617499&pc=OZD | 1220648-2025-26742 | 2025-02-09 05:00 | Death | ABIOMED, INC. | 2025-03-17 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: A NOTIFICATION WAS RECEIVED VIA ABIOMED¿S LONG-TERM OUTCOME AND QUALITY INDICATOR (LOQI) IMPELLA REGISTRY REGARDING A PATIENT WHO EXPERIENCED BLEEDING DURING IMPELLA CP SUPPORT. ACTIONS TAKEN ARE UNKNOWN. THE PATIENT EXPIRED WHILE ON IMPELLA SUPPORT. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21617653&pc=MIH | 3002808486-2025-00048 | 2021-03-21 04:00 | Death | WILLIAM COOK EUROPE | 2025-03-17 04:00 | MIH | ZENITH TX2 DISSECTION ENDOVASCULAR GRAFT TAPERED COMPONENT | Adverse Event Without Identified Device or Use Problem | Shock; Ruptured Aneurysm |  |  | 1 | Event Description: DESCRIPTION OF EVENT ACCORDING TO STUDY: (B)(6): ZENITH ALPHA THORACIC POST MARKET RETROSPECTIVE STUDY ON THE DAY OF THE PROCEDURE, THE PATIENT DIED AFTER EXPERIENCING HEMORRHAGIC SHOCK. ON (B)(6) 2021, THIS 68-YEAR-OLD FEMALE PATIENT WAS URGENTLY TREATED FOR FUSIFORM THORACIC AORTIC ANEURYSM WITH PLACEMENT OF A ZTA-P-36-161 AND A ZDEG-PT-38-30-199-PF. THE PATIENT WAS HEMODYNAMICALLY STABLE PRIOR TO THE PROCEDURE. NO PRE-PROCEDURE OR PROCEDURE IMAGING DATA HAVE BEEN ENTERED. ON THE DAY OF THE PROCEDURE, THE PATIENT EXPERIENCED HEMORRHAGIC SHOCK, SUSPECTED AORTIC RUPTURE DUE TO THORACOABDOMINAL AORTIC ANEURYSM (TAAA). RESUSCITATION MEASURES WERE UNSUCCESSFUL AND THE PATIENT DIED. PATIENT OUTCOME: ON (B)(6) 2021, THE PATIENT DIED. THE CAUSE OF DEATH IS NOTED AS ¿HEMORRHAGIC SHOCK SOME HOURS AFTER PROCEDURE. SUSPECTED AORTIC RUPTURE DUE TO TAAA.¿ THE CLINICAL REVIEW NURSE CANNOT RULE OUT RELATIONSHIP OF DEATH TO DEVICE AND/OR PROCEDURE. Manufacturer Narrative: MANUFACTURERS REF#: (B)(4). G4) SIMILAR TO DEVICE MARKETED UNDER PMA/510(K): P180001. INVESTIGATION IS STILL IN PROGRESS. THIS REPORT IS REQUIRED BY THE FDA UNDER 21CFR PART 803. THIS REPORT IS BASED ON UNCONFIRMED INFORMATION SUBMITTED BY OTHERS. NEITHER THE SUBMISSION OF THIS REPORT NOR ANY STATEMENT MADE IN IT IS INTENDED TO BE AN ADMISSION THAT ANY COOK DEVICE IS DEFECTIVE OR MALFUNCTIONED; THAT A DEATH OR SERIOUS INJURY OCCURRED; OR THAT ANY COOK DEVICE CAUSED OR CONTRIBUTED TO; OR IS LIKELY TO CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY IF A MALFUNCTION OCCURRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21618030&pc=DQA | 2936999-2025-00179 | 2024-12-15 05:00 | Death | MEDIANA CO. LTD. | 2025-03-17 04:00 | DQA | NELLCOR | Device Alarm System | Distress | K142865 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WAS CO-SLEEPING WITH PARENTS AND USING A SPO2 MONITORING DEVICE WHEN THE PATIENT PASSED AWAY WHILE SLEEPING OVERNIGHT. THE FAMILY STATED THAT THE DEVICE DID NOT ALARM AT THE TIME OF THE PATIENT'S DEATH. THE PATIENT, A 2-MONTH-OLD FEMALE WEIGHING 4 KG, WAS EXPERIENCING RESPIRATORY DISTRESS SYNDROME OF NEWBORNS. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21618178&pc=OZD | 1220648-2025-26745 | 2025-02-14 05:00 | Death | ABIOMED, INC. | 2025-03-17 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Renal Failure; Liver Failure | P140003 |  | 1 | Event Description: A NOTIFICATION WAS RECEIVED VIA ABIOMED¿S LONG-TERM OUTCOME AND QUALITY INDICATOR (LOQI) IMPELLA REGISTRY REGARDING A PATIENT WHO EXPERIENCED RENAL AND LIVER FAILURE DURING IMPELLA CP SUPPORT. ACTIONS TAKEN ARE UNKNOWN. THE ISSUE WAS NOT RESOLVED. ADDITIONAL INFORMATION NOTED THAT CARE WAS WITHDRAWN. THE PATIENT WAS PLACED ON HOSPICE AND SUBSEQUENTLY EXPIRED. Manufacturer Narrative: CORRECTED INFORMATION PROVIDED IN B2, H1 AND H6. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA IFU: IMPELLA CP WITH SMARTASSIST: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21618278&pc=OZD | 1220648-2025-26746 | 2025-01-17 05:00 | Death | ABIOMED, INC. | 2025-03-17 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Cardiac Perforation | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. THE PATIENT HAD SHUNTING CAUSED BY VENTRICULAR SEPTAL PERFORATION (VSP). ALTHOUGH IMPELLA MAY HAVE INCREASED THE SHUNT, THE VSP ITSELF HAD EXPANDED OVER TIME, MAKING IT DIFFICULT TO SUPPRESS THE SHUNT. THE PATIENT EXPIRED. THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA IFU: IMPELLA CP WITH SMARTASSIST SYSTEM: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿ SECTION: WARNINGS & CAUTIONS: WARNINGS, SECTION: PRE-SUPPORT EVALUATION, SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER, SECTION: DIRECT AORTIC INSERTION, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION. [ADDITION FOR 5.0, 5.5 ONLY] IN INSTANCES WHERE THE IMPELLA PUMP HAS BEEN PLACED PRIOR TO PERFORMING CARDIAC SURGERY WITH AORTIC CROSS CLAMPING AND CARDIOPLEGIC ARREST, CARE SHOULD BE TAKEN WHEN MANIPULATING THE HEART WHEN THE PUMP POSITION IS FIXED WITH APPLICATION OF THE AORTIC CROSS CLAMP ACROSS THE PUMP CATHETER.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE. SECTION: WARNINGS & CAUTIONS: WARNINGS, SECTION: PRE-SUPPORT EVALUATION, SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER, SECTION: DIRECT AORTIC INSERTION, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: ¿TO REDUCE THE RISK OF CARDIAC INJURY (INCLUDING VENTRICULAR PERFORATION), PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY IN THE SETTINGS OF ACUTE INFARCTION WITH TISSUE NECROSIS.¿ ¿TO REDUCE THE RISK OF VASCULAR INJURY, PHYSICIANS SHOULD EXERCISE CAUTION WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX PERIPHERAL VASCULAR ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: UNREPAIRED ABDOMINAL AORTIC ANEURYSM, SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM, DISSECTION OF THE ASCENDING/ TRANSVERSE/DESCENDING AORTA, CHRONIC ANATOMICAL CHANGES IN THE RELATIONSHIP OF THE AORTA/AORTIC VALVE/VENTRICULAR ALIGNMENT, SIGNIFICANT MOBILE ATHEROMATOUS DISEASE IN THE THORACIC OR ABDOMINAL AORTA OR PERIPHERAL VESSELS.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING A RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY IN THE LEFT VENTRICLE AFTER CPR WITH ECHOCARDIOGRAPHY GUIDANCE. SECTION: WARNINGS & CAUTIONS: CAUTIONS: ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: CAUTIONS: ¿HANDLE WITH CARE. THE IMPELLA CATHETER CAN BE DAMAGED DURING REMOVAL FROM PACKAGING, PREPARATION, INSERTION, AND REMOVAL. DO NOT BEND, PULL, OR PLACE EXCESS PRESSURE ON THE CATHETER OR MECHANICAL COMPONENTS AT ANY TIME.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿AVOID MANUAL COMPRESSION OF THE INLET AND OUTLET AREAS OF THE CANNULA ASSEMBLY.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS. SECTION: PRE-SUPPORT EVALUATION. SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER. SECTION: DIRECT AORTIC INSERTION. SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21618942&pc=NGV | 2124215-2025-15698 | 2024-11-09 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-03-17 04:00 | NGV | WATCHMAN LAA CLOSURE DEVICE & DELIVERY SYSTEM | Difficult to Open or Close | Stroke/CVA; Hemorrhage/Bleeding; Myocardial Infarction; Transient Ischemic Attack; Pericardial Effusion; Embolism/Embolus |  |  | 1 | Event Description: REPORTED VIA JOURNAL ARTICLE: IT WAS REPORTED THAT PATIENT DEATH(S) OCCURRED. METHODS: THIS RETROSPECTIVE STUDY INCLUDED CONSECUTIVE PATIENTS WITH NON-VALVULAR ATRIAL FIBRILLATION (NVAF) WHO UNDERWENT LEFT ATRIAL APPENDAGE CLOSURE (LAAC) USING THE WATCHMAN LAA CLOSURE DEVICE & DELIVERY SYSTEM (WDS) BETWEEN (B)(6) 2020 AND (B)(6) 2023. THE WATCHMAN 2.5 DEVICE WAS USED UNTIL (B)(6) 2021, AFTER WHICH THE WATCHMAN FLX WAS IMPLEMENTED. DURING THE SAME OBSERVATION PERIOD, A SUBSET OF PATIENTS UNDERWENT CONCOMITANT TRANSCATHETER EDGE-TO-EDGE REPAIR (TEER) FOR MODERATE-TO-SEVERE OR GREATER MITRAL REGURGITATION (MR). THE PRIMARY OBJECTIVE WAS TO COMPARE THE EFFICACY AND FEASIBILITY OF THE COMBINED LAAC AND TEER APPROACH VERSUS LAAC ALONE. RESULTS: A TOTAL OF 192 PATIENTS WERE INCLUDED IN THIS STUDY. ELEVEN PATIENTS UNDERWENT LAAC AND TEER, WHILE THE REMAINING 181 UNDERWENT LAAC ALONE. THERE WERE NO CASES OF MR GREATER THAN 3PLUS IN THE LAAC ALONE GROUP. OF THE 11 PATIENTS IN THE LAAC AND TEER GROUP, 7 PATIENTS RECEIVED LAAC AFTER TEER, AND THE OTHER 4 PATIENTS UNDERWENT LAAC CONCOMITANTLY WITH TEER. NO PATIENTS EXPERIENCED DEVICE MIGRATION OR DEVICE EMBOLIZATION DURING THE PROCEDURE. ONE PATIENT REQUIRED A TRANSFUSION DUE TO ANEMIA, WHILE ONE PATIENT HAD A PERICARDIAL EFFUSION THAT DID NOT NECESSITATE ANY INTERVENTION. TWO PATIENTS EXPERIENCED NON-PROCEDURAL EVENTS: ONE FATAL MYOCARDIAL INFARCTION AND ONE CASE OF GASTROINTESTINAL BLEEDING REQUIRING TRANSFUSION, BOTH OCCURRING 5 DAYS POST-IMPLANTATION. 13 PATIENTS DIED: 3 FROM MALIGNANCY, 2 FROM CEREBROVASCULAR EVENTS, 2 FROM PULMONARY PNEUMONIA, 2 FROM HEMODIALYSIS-RELATED EVENTS, 1 FROM HEART FAILURE, 1 FROM MYOCARDIAL INFARCTION, AND 2 FROM UNKNOWN CAUSES. 13 PATIENTS HAD BLEEDING EVENTS, INCLUDING 7 CASES OF GASTROINTESTINAL BLEEDING REQUIRED TRANSFUSION. 2 PATIENTS EXPERIENCED A CEREBRAL VASCULAR ACCIDENT, 2 PATIENTS EXPERIENCED A TRANSIENT ISCHEMIC ATTACK, 1 PATIENT EXPERIENCED A SYSTEMIC EMBOLISM, 2 PATIENTS EXPERIENCED MINOR BLEEDING EVENTS, AND 11 PATIENTS EXPERIENCED 11 MAJOR BLEEDING EVENTS. TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) IMAGING AT THE 45 DAY FOLLOW UP AND THE ONE YEAR FOLLOW UP SHOWED THAT 1 PATIENT HAD A PERIDEVICE LEAK OF GREATER THAN 5MM. Manufacturer Narrative: B2 DATE OF DEATH - ARTICLE PUBLISH DATE USED AS DEATH DATE IS UNKNOWN. B3 DATE OF EVENT - ARTICLE PUBLISH DATE USED AS EVENT DATE IS UNKNOWN. D3 MODEL NUMBER, D3 LOT NUMBER, D4 EXPIRATION DATE, D4 UNIQUE IDENTIFIER (UDI) # ARE UNKNOWN AS THE INFORMATION WAS NOT PROVIDED IN THE LITERATURE ARTICLE. D6A IMPLANT DATE WAS REPORTED AS (B)(6) 2020 THROUGH (B)(6) 2023. LITERATURE CITATION: FUKUDA, N., IMAMURA, T., TANAKA, S., KATAOKA, N., USHIJIMA, R., UENO, H., & KINUGAWA, K. (2024). FEASIBILITY OF COMBINED THERAPY: PERCUTANEOUS LEFT ATRIAL APPENDAGE CLOSURE AND TRANSCATHETER EDGE-TO-EDGE REPAIR. CARDIOVASCULAR INTERVENTION AND THERAPEUTICS, 1-14.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21618946&pc=NGV | 2124215-2025-15704 | 2024-11-09 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-03-17 04:00 | NGV | WATCHMAN FLX LEFT ATRIAL APPENDAGE CLOSURE DEVICE WITH DELIVERY SYSTEM | Difficult to Open or Close | Stroke/CVA; Hemorrhage/Bleeding; Myocardial Infarction; Transient Ischemic Attack; Pericardial Effusion; Embolism/Embolus |  |  | 1 | Event Description: REPORTED VIA JOURNAL ARTICLE: IT WAS REPORTED THAT PATIENT DEATH(S) OCCURRED. METHODS: THIS RETROSPECTIVE STUDY INCLUDED CONSECUTIVE PATIENTS WITH NON-VALVULAR ATRIAL FIBRILLATION (NVAF) WHO UNDERWENT LEFT ATRIAL APPENDAGE CLOSURE (LAAC) USING THE WATCHMAN LAA CLOSURE DEVICE & DELIVERY SYSTEM (WDS) BETWEEN (B)(6) 2020 AND (B)(6) 2023. THE WATCHMAN 2.5 DEVICE WAS USED UNTIL (B)(6) 2021, AFTER WHICH THE WATCHMAN FLX WAS IMPLEMENTED. DURING THE SAME OBSERVATION PERIOD, A SUBSET OF PATIENTS UNDERWENT CONCOMITANT TRANSCATHETER EDGE-TO-EDGE REPAIR (TEER) FOR MODERATE-TO-SEVERE OR GREATER MITRAL REGURGITATION (MR). THE PRIMARY OBJECTIVE WAS TO COMPARE THE EFFICACY AND FEASIBILITY OF THE COMBINED LAAC AND TEER APPROACH VERSUS LAAC ALONE. RESULTS: A TOTAL OF 192 PATIENTS WERE INCLUDED IN THIS STUDY. ELEVEN PATIENTS UNDERWENT LAAC AND TEER, WHILE THE REMAINING 181 UNDERWENT LAAC ALONE. THERE WERE NO CASES OF MR GREATER THAN 3PLUS IN THE LAAC ALONE GROUP. OF THE 11 PATIENTS IN THE LAAC AND TEER GROUP, 7 PATIENTS RECEIVED LAAC AFTER TEER, AND THE OTHER 4 PATIENTS UNDERWENT LAAC CONCOMITANTLY WITH TEER. NO PATIENTS EXPERIENCED DEVICE MIGRATION OR DEVICE EMBOLIZATION DURING THE PROCEDURE. ONE PATIENT REQUIRED A TRANSFUSION DUE TO ANEMIA, WHILE ONE PATIENT HAD A PERICARDIAL EFFUSION THAT DID NOT NECESSITATE ANY INTERVENTION. TWO PATIENTS EXPERIENCED NON-PROCEDURAL EVENTS: ONE FATAL MYOCARDIAL INFARCTION AND ONE CASE OF GASTROINTESTINAL BLEEDING REQUIRING TRANSFUSION, BOTH OCCURRING 5 DAYS POST-IMPLANTATION. 13 PATIENTS DIED: 3 FROM MALIGNANCY, 2 FROM CEREBROVASCULAR EVENTS, 2 FROM PULMONARY PNEUMONIA, 2 FROM HEMODIALYSIS-RELATED EVENTS, 1 FROM HEART FAILURE, 1 FROM MYOCARDIAL INFARCTION, AND 2 FROM UNKNOWN CAUSES. 13 PATIENTS HAD BLEEDING EVENTS, INCLUDING 7 CASES OF GASTROINTESTINAL BLEEDING REQUIRED TRANSFUSION. 2 PATIENTS EXPERIENCED A CEREBRAL VASCULAR ACCIDENT, 2 PATIENTS EXPERIENCED A TRANSIENT ISCHEMIC ATTACK, 1 PATIENT EXPERIENCED A SYSTEMIC EMBOLISM, 2 PATIENTS EXPERIENCED MINOR BLEEDING EVENTS, AND 11 PATIENTS EXPERIENCED 11 MAJOR BLEEDING EVENTS. TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) IMAGING AT THE 45 DAY FOLLOW UP AND THE ONE YEAR FOLLOW UP SHOWED THAT 1 PATIENT HAD A PERIDEVICE LEAK OF GREATER THAN 5MM. Manufacturer Narrative: B2 DATE OF DEATH - ARTICLE PUBLISH DATE USED AS DEATH DATE IS UNKNOWN. B3 DATE OF EVENT - ARTICLE PUBLISH DATE USED AS EVENT DATE IS UNKNOWN. D3 MODEL NUMBER, D3 LOT NUMBER, D4 EXPIRATION DATE, D4 UNIQUE IDENTIFIER (UDI) # ARE UNKNOWN AS THE INFORMATION WAS NOT PROVIDED IN THE LITERATURE ARTICLE. D6A IMPLANT DATE WAS REPORTED AS (B)(6) 2020 THROUGH (B)(6) 2023. LITERATURE CITATION: FUKUDA, N., IMAMURA, T., TANAKA, S., KATAOKA, N., USHIJIMA, R., UENO, H., & KINUGAWA, K. (2024). FEASIBILITY OF COMBINED THERAPY: PERCUTANEOUS LEFT ATRIAL APPENDAGE CLOSURE AND TRANSCATHETER EDGE-TO-EDGE REPAIR. CARDIOVASCULAR INTERVENTION AND THERAPEUTICS, 1-14.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21619246&pc=OZD | 1220648-2025-26754 | 2025-02-21 05:00 | Death | ABIOMED, INC. | 2025-03-17 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemolysis | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT A PATIENT DEVELOPED HEMOLYSIS DURING IMPELLA CP SUPPORT. THE PUMP WAS REPOSITIONED, BUT THE CONDITION PERSISTED. SUPPORT WAS CONTINUED FOR TWO MORE DAYS UNTIL THE DECISION WAS MADE TO WITHDRAW CARE DUE TO THE PATIENT'S SIGNIFICANT COMORBIDITIES AND POOR PROGNOSIS. THE PATIENT SUBSEQUENTLY PASSED AWAY. IT IS UNKNOWN IF THE CONDITION CONTRIBUTED TO THE PATIENT'S PASSING. Manufacturer Narrative: THE INVESTIGATION OF HEMOLYSIS HAS BEEN COMPLETED. NO PRODUCT WAS RETURNED FOR ANALYSIS. THE CLINICAL DETAILS MENTIONED THE PUMP WAS PULLED 10 CENTIMETERS BACK. CLINICAL INFORMATION MENTIONED THE HEMOLYSIS DID NOT RESOLVE WITH REPOSITIONING AND WITH PUMP REMOVAL. CLINICAL DETAILS ALSO MENTIONED THE PATIENT HAD POOR PROGNOSIS AND SIGNIFICANT COMORBIDITIES. THE DATA LOGS RETURNED DID NOT SPAN THE ENTIRE CASE DURATION. LOGS ONLY SHOW 20-FEB-2025 AND NO MOTOR CURRENT SPIKES. THE ROOT CAUSE OF THE HEMOLYSIS ISSUE COULD NOT BE DEFINITIVELY DETERMINED AS NO PRODUCT WAS RETURNED FOR ANALYSIS AND INSUFFICIENT DATA LOGS WERE PROVIDED. G.1 REPORTING CONTACT US PHONE NUMBER SHOULD OF BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT 1220648-2025-26754. ADDED MANUFACTURING SITE FAX NUMBER AS IT WAS INADVERTENTLY OMITTED FROM INITIAL MANUFACTURER DEVICE REPORT 1220648-2025-26754. H.1 TYPE OF REPORTABLE EVENT WAS PREVIOUSLY ENTERED INCORRECTLY ON THE INITIAL MANUFACTURER DEVICE REPORT 1220648-2025-26754. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE: ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS: ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿. ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION: ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21619326&pc=NPT | 2025587-2025-01930 | 2025-02-19 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-17 04:00 | NPT | EVOLUT FX DCS | Patient Device Interaction Problem | Perforation | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE INVOLVING THE 29MM EVOLUT FX+, INITIAL ARTERIAL ACCESS WAS PERFORMED ON THE RIGHT FEMORAL ARTERY, BUT DUE TO ISSUES CROSSING THE RIGHT COMMON ILIAC ARTERY, THE PHYSICIAN SWITCHED TO THE LEFT FEMORAL ARTERY. THERE WERE DIFFICULTIES PASSING DEVICES THROUGH THE SUPERIOR PORTION OF THE LEFT COMMON ILIAC ARTERY, AND DAMAGE WAS NOTED TO THE TIP OF THE DELIVERY CATHETER SYSTEM (DCS), RAISING CONCERNS ABOUT THE SAFETY OF THE DCS AND THE VALVE PADDLES. THE ORIGINAL DCS WAS DISCARDED AND A NEW DCS WAS OPENED TO COMPLETE THE PROCEDURE. SUBSEQUENTLY, A 16FR NON-MEDTRONIC SHEATH WAS PASSED THROUGH THE LEFT ILIAC ARTERY. THE DCS WAS PASSED THROUGH THE SHEATH, ADVANCED TO THE ANNULUS, AND THE VALVE WAS SUCCESSFULLY DEPLOYED. VASCULAR SURGERY WAS PERFORMED TO CLOSE BOTH ARTERIAL ACCESS POINTS. THE PATIENT REMAINED STABLE IN THE CATHETER LAB FOR ABOUT 30-45 MINUTES. APPROXIMATELY THREE HOURS AFTER THE PROCEDURE, THE PATIENT DIED DUE TO BLEEDING FROM A PERFORATION IN TH E LEFT COMMON ILIAC ARTERY.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT DURING THE VALVE IMPLANT, THE FIRST DELIVERY CATHETER SYSTEM (DCS) THAT HAD DAMAGE AT THE TIP OF THE DELIVERY CATHETER SYSTEM (DCS) AND WAS OPENED AND NOT USED. PER THE PHYSICIAN, THE SECOND DELIVERY CATHETER SYSTEM (DCS) DID NOT CAUSE OR CONTRIBUTE TO THE VASCULAR INJURY. THE CAUSE WAS NOT KNOWN, BUT THE PHYSICIAN SPECULATED EITHER THE GUIDEWIRE DISSECTED AND/OR PERFORATED OR THE TRANSITION POINT ON THE DILATOR/SHEATH DISSECTED THE ARTERIAL WALL. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID: EVFXPLUS-29 (SERIAL: (B)(6)); PRODUCT TYPE: 0195-HEART VALVES. PRODUCT ID: D-EVOLUTFX-2329 (0012426256); PRODUCT TYPE: 0195-HEART VALVES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5 D4 H4 H6. CORRECTED DATA: H8 - USAGE OF DEVICE MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21619371&pc=NPT | 2025587-2025-01929 | 2025-02-19 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-17 04:00 | NPT | EVOLUT FX DCS | No Apparent Adverse Event; Appropriate Term/Code Not Available; Patient Device Interaction Problem | Perforation; No Clinical Signs, Symptoms or Conditions | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE INVOLVING THE 29 MM EVOLUT FX+, INITIAL ARTERIAL ACCESS WAS PERFORMED ON THE RIGHT FEMORAL ARTERY, BUT DUE TO ISSUES CROSSING THE RIGHT COMMON ILIAC ARTERY, THE PHYSICIAN SWITCHED TO THE LEFT FEMORAL ARTERY. THERE WERE DIFFICULTIES PASSING DEVICES THROUGH THE SUPERIOR PORTION OF THE LEFT COMMON ILIAC ARTERY, AND DAMAGE WAS NOTED TO THE TIP OF THE DELIVERY CATHETER SYSTEM (DCS), RAISING CONCERNS ABOUT THE SAFETY OF THE DCS AND THE VALVE PADDLES. THE ORIGINAL DCS WAS DISCARDED AND A NEW DCS WAS OPENED TO COMPLETE THE PROCEDURE. SUBSEQUENTLY, A 16FR NON-MEDTRONIC SHEATH WAS PASSED THROUGH THE LEFT ILIAC ARTERY. THE DCS WAS PASSED THROUGH THE SHEATH, ADVANCED TO THE ANNULUS, AND THE VALVE WAS SUCCESSFULLY DEPLOYED. VASCULAR SURGERY WAS PERFORMED TO CLOSE BOTH ARTERIAL ACCESS POINTS. THE PATIENT REMAINED REMAINED STABLE IN THE CATHETER LAB FOR ABOUT 30-45 MINUTES. APPROXIMATELY THREE HOURS AFTER THE PROCEDURE, THE PATIENT DIED DUE TO BLEEDING FROM A PERFORATION IN TH E LEFT COMMON ILIAC ARTERY.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT DURING THE VALVE IMPLANT, THE FIRST DELIVERY CATHETER SYSTEM (DCS) THAT HAD DAMAGE AT THE TIP OF THE DELIVERY CATHETER SYSTEM (DCS) AND WAS OPENED AND NOT USED. PER THE PHYSICIAN, THE SECOND DELIVERY CATHETER SYSTEM (DCS) DID NOT CAUSE OR CONTRIBUTE TO THE VASCULAR INJURY. THE CAUSE WAS NOT KNOWN, BUT THE PHYSICIAN SPECULATED EITHER THE GUIDEWIRE DISSECTED AND/OR PERFORATED OR THE TRANSITION POINT ON THE DILATOR/SHEATH DISSECTED THE ARTERIAL WALL. Manufacturer Narrative: UPDATED DATA: B5 H6 ADDITIONAL CODES. MEDTRONIC RECEIVED ADDITIONAL INFORMATION THAT CHANGED THE EVENT DESCRIPTION AND DEVICE RELATEDNESS OF THIS PREVIOUSLY REPORTED EVENT. THERE IS NO EVIDENCE TO SUGGEST THE DEVICE CAUSED OR CONTRIBUTED TO A DEATH OR SERIOUS INJURY. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID EVFXPLUS-29 (SERIAL: (B)(6)); PRODUCT TYPE: 0195-HEART VALVES; PRODUCT ID D-EVOLUTFX-2329 (UNKNOWN); PRODUCT TYPE: 0195-HEART VALVES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21619580&pc=BZD | 2518422-2025-103080 | 2024-11-15 05:00 | Death | RESPIRONICS, INC. | 2025-03-17 04:00 | BZD | DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER HAS BEEN INFORMED OF THE PASSING OF A PATIENT WHO WAS USING A DREAMSTATION AUTO CPAP DEVICE. THE CALLER EXPLAINED THAT SHE HAD RECEIVED THE DEVICE FROM A DECEASED INDIVIDUAL AND INQUIRED ABOUT THE RETURN PROCESS. SHE WAS PROVIDED WITH A LINK TO THE RECALL WEBSITE TO REGISTER THE DEVICE, AND WAS ASSURED THAT A REPRESENTATIVE WOULD REACH OUT TO HER FOR FURTHER DETAILS. NO ADDITIONAL INFORMATION WAS SHARED, AND THERE ARE NO CLAIMS INDICATING THAT THE DEVICE PLAYED ANY ROLE IN THE PATIENT'S DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21619600&pc=OZD | 1220648-2025-26761 | 2025-02-22 05:00 | Death | ABIOMED, INC | 2025-03-17 04:00 | OZD | AUTOMATED IMPELLA CONTROLLER | Application Program Problem; Device Sensing Problem | Cardiovascular Insufficiency | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WAS ON IMPELLA 5.5 PUMP SUPPORT FOR 14 DAYS WHILE THE PATIENT WAS AWAITING TRANSPLANT/LEFT VENTRICULAR ASSIST DEVICE (LVAD) DECISIONS. THE AUTOMATED IMPELLA CONTROLLER (AIC) EXHIBITIED A CONTROLLER ALARM YELLOW FOR A FAILURE. THE AIC WAS REMOVED/REPLACED AND SUPPORT WAS CONTINUED WITHOUT HARM OR INJURY FOR JUST UNDER 21 DAYS UNTIL THE DECISION WAS MADE TO WITHDRAW CARE AND THE PATIENT SUBSEQUENTLY EXPIRED. THE PATIENT'S OUTCOME WAS REVIEWED BY ABIOMED'S MEDICAL SAFETY OFFICER (MSO) AND IT WAS DETERMINED THAT THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA 5.5 AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. ALTHOUGH THE CONSOLE ERROR OCCURRED, THE PATIENT'S LVAD, CRITICAL CARDIAC NATURE, AND MULTIPLE OTHER ISSUES CONTRIBUTED TO THE PATIENT'S DEMISE. Manufacturer Narrative: THE INVESTIGATION OF AUTOMATED IMPELLA CONTROLLER (AIC) - CONTROLLER ERROR (YELLOW ALARM) HAS BEEN COMPLETED. CONSOLE LOGS FROM THE REPORTED DAY OF THE EVENT ARE CONSISTENT WITH THE COMPLAINT. THE CONSOLE WAS SENT BACK TO FIELD SERVICE FOR FURTHER INVESTIGATION AND WAS TESTED AFTER ARRIVING IN FS. VOLTAGE FROM THE POWER BATTERY MANAGER (PBM) AT CONNECTOR (J3) WAS 4.9V. THE VOLTAGE ON THE OPTICAL BENCH LAMP CONNECTOR (P4) WAS 4.8V. THE LAMP POWER WAS 2.65UW, WHICH IS BELOW THE SPECIFIED VALUE STATED IN THE PM PROCEDURE, BUT DID NOT CAUSE THE REPORTED ISSUES. THE ANALOG OUTPUT FROM THE OPTICAL BENCH (ENVELOPE/FRINGE) WAS OBSERVED WITHOUT ANY DISTORTIONS. THE CONTROLLER ERROR WAS IDENTIFIED AS A RESULT OF AN ANOMALY IN THE OPTICAL BENCH FIRMWARE COMMUNICATION PROTOCOL, LEADING TO MISALIGNMENT OF THE DATA COMMUNICATION PACKETS RECEIVED BY THE EPC. MEANWHILE, THE AIC SOFTWARE FUNCTIONED AS INTENDED. FURTHER EXAMINATION OF THE AIC REVEALED THAT THE 250 HZ PLACEMENT SIGNAL WAS EXPERIENCING SAMPLE LOSS, RESULTING IN NEGATIVE SPIKES. ROOT CAUSE: THE ROOT CAUSE OF THE ¿CONTROLLER ERROR (OPM COMM CRC INVALID) [OPTICAL PUMP CONNECTED] - ALARM ISSUED (I.E. AIC - LOSS OF OPM DATA)" AND THE "CONTROLLER ERROR (OPTICAL BENCH AMBIENT PRESSURE OUT OF RANGE) [OPTICAL PUMP CONNECTED] - ALARM ISSUED" WAS DUE TO A FIRMWARE ISSUE (OPTICAL BENCH FW ANOMALY). THE ROOT CAUSE FOR THE DISTORTION ON THE PLACEMENT SIGNAL WAS MOST LIKELY DUE TO THE OPTICAL BENCH MALFUNCTION. Manufacturer Narrative: THE IMPELLA DEVICE WAS RETURNED, AND AN EVALUATION IS UNDERWAY. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21619880&pc=QZI | MW5167712 | 2025-02-12 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-03-17 04:00 | QZI | FARAPULSE PULSED FIELD ABLATION |  | Atrial Fibrillation; Bradycardia; Dyspnea; Low Blood Pressure/ Hypotension; Ventricular Fibrillation |  |  | 1 | Event Description: IT WAS REPORTED THAT AFTER THE ATRIAL FIBRILLATION CASE HAD COMPLETED, THE PATIENT HAD PASSED AWAY. THE PATIENT HAD TO BE EXTUBATED, AND THEY WERE PREPARING TO TRANSPORT THE PATIENT TO RECOVERY, WHEN "EXTREME BRADYCARDIA AND LOW BLOOD PRESSURE" WERE NOTICED IN THE PATIENT. THE PATIENT ALSO DEMONSTRATED "DIFFICULTY BREATHING." THE PHYSICIAN WAS ATTEMPTING TO GAIN ACCESS IN ORDER TO IMPLANT A PACER WIRE INTO THE PATIENT. AT SOME POINT, THE PATIENT WENT INTO VENTRICULAR FIBRILLATION (VFIB) AND PASSED AWAY. MEDICAL INTERVENTION WAS PROVIDED TO PATIENT; HOWEVER, THE DETAILS COULD NOT BE CONFIRMED. THE PATIENT PASSED AWAY ON (B)(6) 2025. FARAPULSE SYSTEM WAS IN USE. THE PHYSICIAN'S OPINION ON THE CAUSE OF THE ADVERSE EVENT IS PATIENT CONDITION, HOWEVER THE CAUSE IS UNKNOWN. "THIS REPORT REFLECTS INFORMATION RECEIVED BY FDA IN THE FORM OF A NOTIFICATION PER 803.22 (B)(2)".
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21620125&pc=NKM | 2135147-2025-01337 | 2000-01-01 05:00 | Death | ABBOTT MEDICAL | 2025-03-17 04:00 | NKM | MITRACLIP® SYSTEM CLIP DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P100009 |  | 1 | Event Description: THE ARTICLE "ONE-YEAR RECURRENT TRICUSPID REGURGITATION AFTER SUCCESSFUL TRANSCATHETER EDGE TO EDGE REPAIR: THE TRI-SPA REGISTRY" WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE MULTI CENTER STUDY, TO EVALUATE THE PREDICTIVE FACTORS AND MID-TERM OUTCOMES OF RECURRENT TRICUSPID REGURGITATION (TR) FOLLOWING SUCCESSFUL TRANSCATHETER EDGE-TO-EDGE REPAIR (TEER). DEVICES MENTIONED INCLUDE MITRACLIP, TRICLIP, PASCAL P10, AND PASCAL ACE. THE ARTICLE CONCLUDED THAT RECURRENT TR WAS ASSOCIATED WITH WORSE OUTCOMES. RIGHT VENTRICULAR FRACTIONAL AREA CHANGE, RESIDUAL TR AND PRIMARY TR WERE INDEPENDENT PREDICTORS FOR RECURRENT TR. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS RODRIGO ESTÉVEZ-LOUREIRO AT UNIVERSITY HOSPITAL ALVARO CUNQUEIRO, VIGO, SPAIN WITH CORRESPONDING EMAIL: ROIESTEVEZ@HOTMAIL.COM]. THE TIME FRAME OF THE STUDY WAS 2000 TO 2023. A TOTAL OF 280 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH 142 RECEIVED AN ABBOTT DEVICE. MEDIAN AGE WAS 76 AND MAJORITY GENDER WAS FEMALE. COMORBIDITIES INCLUDE PREVIOUS VALVE SURGERY, DIABETES MELLITUS, ARTERIAL HYPERTENSION, DYSLIPIDEMIA, SMOKER, ATRIAL FIBRILLATION, PACING, CHRONIC KIDNEY DISEASE, CHRONIC LUNG DISEASE, CABG, EDEMA, ASCITES, HEART FAILURE, TRICUSPID REGURGITATION (PRIMARY, SECONDARY). PERI AND POST-PROCEDURAL COMPLICATIONS INCLUDED RECURRENT TR, REINTERVENTION, DEATH, HEART FAILURE HOSPITALIZATION, SINGLE LEAFLET DEVICE ATTACHMENT, OFF LABEL USE.Event Description: THE ARTICLE "ONE-YEAR RECURRENT TRICUSPID REGURGITATION AFTER SUCCESSFUL TRANSCATHETER EDGE TO EDGE REPAIR: THE TRI-SPA REGISTRY" WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE MULTI CENTER STUDY, TO EVALUATE THE PREDICTIVE FACTORS AND MID-TERM OUTCOMES OF RECURRENT TRICUSPID REGURGITATION (TR) FOLLOWING SUCCESSFUL TRANSCATHETER EDGE-TO-EDGE REPAIR (TEER) . DEVICES MENTIONED INCLUDE MITRACLIP, TRICLIP, PASCAL P10, AND PASCAL ACE. THE ARTICLE CONCLUDED THAT RECURRENT TR WAS ASSOCIATED WITH WORSE OUTCOMES. RIGHT VENTRICULAR FRACTIONAL AREA CHANGE, RESIDUAL TR AND PRIMARY TR WERE INDEPENDENT PREDICTORS FOR RECURRENT TR. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS RODRIGO ESTÉVEZ-LOUREIRO AT UNIVERSITY HOSPITAL ALVARO CUNQUEIRO, VIGO, SPAIN WITH CORRESPONDING EMAIL ROIESTEVEZ@HOTMAIL.COM]. THE TIME FRAME OF THE STUDY WAS 2000 TO 2023. A TOTAL OF 280 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH 142 RECEIVED AN ABBOTT DEVICE. MEDIAN AGE WAS 76 AND MAJORITY GENDER WAS FEMALE. COMORBIDITIES INCLUDE PREVIOUS VALVE SURGERY, DIABETES MELLITUS, ARTERIAL HYPERTENSION, DYSLIPIDEMIA, SMOKER, ATRIAL FIBRILLATION, PACING, CHRONIC KIDNEY DISEASE, CHRONIC LUNG DISEASE, CABG, EDEMA, ASCITES, HEART FAILURE, TRICUSPID REGURGITATION (PRIMARY, SECONDARY). PERI AND POST-PROCEDURAL COMPLICATIONS INCLUDED RECURRENT TR, REINTERVENTION, DEATH, HEART FAILURE HOSPITALIZATION, SINGLE LEAFLET DEVICE ATTACHMENT, OFF LABEL USE Manufacturer Narrative: IT IS UNKNOWN IF THE DEVICE IS RETURNING FOR ANALYSIS. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION B2 DEATH DATE IS ESTIMATED. B3 EVENT DATE IS ESTIMATED. LITERATURE ATTACHMENT: ONE-YEAR RECURRENT TRICUSPID REGURGITATION AFTER SUCCESSFUL TRANSCATHETER EDGE TO EDGE REPAIR: THE TRI-SPA REGISTRY. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF MITRACLIP WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING PREVIOUS VALVE SURGERY, DIABETES MELLITUS, ARTERIAL HYPERTENSION, DYSLIPIDEMIA, SMOKER, ATRIAL FIBRILLATION, PACING, CHRONIC KIDNEY DISEASE, CHRONIC LUNG DISEASE, CABG, EDEMA, ASCITES, HEART FAILURE, TRICUSPID REGURGITATION (PRIMARY, SECONDARY). COMPLICATIONS REPORTED INCLUDED RECURRENT TR, REINTERVENTION, DEATH, HEART FAILURE HOSPITALIZATION, SINGLE LEAFLET DEVICE ATTACHMENT, OFF LABEL USE; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. THE REPORTED OFF-LABEL USE WAS ASSOCIATED WITH USING A MITRACLIP ON A TRICUSPID VALVE. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. B2 DEATH DATE AND B3 EVENT DATE ARE ESTIMATED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21620224&pc=NPS | 2135147-2025-01341 | 2000-01-01 05:00 | Death | ABBOTT MEDICAL | 2025-03-17 04:00 | NPS | TRICLIP DELIVERY SYSTEM (TCDS) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P230007 |  | 1 | Event Description: THE ARTICLE "ONE-YEAR RECURRENT TRICUSPID REGURGITATION AFTER SUCCESSFUL TRANSCATHETER EDGE TO EDGE REPAIR: THE TRI-SPA REGISTRY" WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE MULTI CENTER STUDY, TO EVALUATE THE PREDICTIVE FACTORS AND MID-TERM OUTCOMES OF RECURRENT TRICUSPID REGURGITATION (TR) FOLLOWING SUCCESSFUL TRANSCATHETER EDGE-TO-EDGE REPAIR (TEER). DEVICES MENTIONED INCLUDE MITRACLIP, TRICLIP, PASCAL P10, AND PASCAL ACE. THE ARTICLE CONCLUDED THAT RECURRENT TR WAS ASSOCIATED WITH WORSE OUTCOMES. RIGHT VENTRICULAR FRACTIONAL AREA CHANGE, RESIDUAL TR AND PRIMARY TR WERE INDEPENDENT PREDICTORS FOR RECURRENT TR. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS RODRIGO ESTÉVEZ-LOUREIRO AT UNIVERSITY HOSPITAL ALVARO CUNQUEIRO, VIGO, SPAIN WITH CORRESPONDING EMAIL ROIESTEVEZ@HOTMAIL.COM]. THE TIME FRAME OF THE STUDY WAS 2000 TO 2023. A TOTAL OF 280 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH 142 RECEIVED AN ABBOTT DEVICE. MEDIAN AGE WAS 76 AND MAJORITY GENDER WAS FEMALE. COMORBIDITIES INCLUDE PREVIOUS VALVE SURGERY, DIABETES MELLITUS, ARTERIAL HYPERTENSION, DYSLIPIDEMIA, SMOKER, ATRIAL FIBRILLATION, PACING, CHRONIC KIDNEY DISEASE, CHRONIC LUNG DISEASE, CABG, EDEMA, ASCITES, HEART FAILURE, TRICUSPID REGURGITATION (PRIMARY, SECONDARY). PERI AND POST-PROCEDURAL COMPLICATIONS INCLUDED RECURRENT TR, REINTERVENTION, DEATH, HEART FAILURE HOSPITALIZATION, SINGLE LEAFLET DEVICE ATTACHMENT, OFF LABEL USE.Event Description: THE ARTICLE "ONE-YEAR RECURRENT TRICUSPID REGURGITATION AFTER SUCCESSFUL TRANSCATHETER EDGE TO EDGE REPAIR: THE TRI-SPA REGISTRY" WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE MULTI CENTER STUDY, TO EVALUATE THE PREDICTIVE FACTORS AND MID-TERM OUTCOMES OF RECURRENT TRICUSPID REGURGITATION (TR) FOLLOWING SUCCESSFUL TRANSCATHETER EDGE-TO-EDGE REPAIR (TEER) . DEVICES MENTIONED INCLUDE MITRACLIP, TRICLIP, PASCAL P10, AND PASCAL ACE. THE ARTICLE CONCLUDED THAT RECURRENT TR WAS ASSOCIATED WITH WORSE OUTCOMES. RIGHT VENTRICULAR FRACTIONAL AREA CHANGE, RESIDUAL TR AND PRIMARY TR WERE INDEPENDENT PREDICTORS FOR RECURRENT TR. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS RODRIGO ESTÉVEZ-LOUREIRO AT UNIVERSITY HOSPITAL ALVARO CUNQUEIRO, VIGO, SPAIN WITH CORRESPONDING EMAIL ROIESTEVEZ@HOTMAIL.COM]. THE TIME FRAME OF THE STUDY WAS 2000 TO 2023. A TOTAL OF 280 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH 142 RECEIVED AN ABBOTT DEVICE. MEDIAN AGE WAS 76 AND MAJORITY GENDER WAS FEMALE. COMORBIDITIES INCLUDE PREVIOUS VALVE SURGERY, DIABETES MELLITUS, ARTERIAL HYPERTENSION, DYSLIPIDEMIA, SMOKER, ATRIAL FIBRILLATION, PACING, CHRONIC KIDNEY DISEASE, CHRONIC LUNG DISEASE, CABG, EDEMA, ASCITES, HEART FAILURE, TRICUSPID REGURGITATION (PRIMARY, SECONDARY). PERI AND POST-PROCEDURAL COMPLICATIONS INCLUDED RECURRENT TR, REINTERVENTION, DEATH, HEART FAILURE HOSPITALIZATION, SINGLE LEAFLET DEVICE ATTACHMENT, OFF LABEL USE Manufacturer Narrative: IT IS UNKNOWN IF THE DEVICE IS RETURNING FOR ANALYSIS. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. B2 DEATH DATE IS ESTIMATED. B3 EVENT DATE IS ESTIMATED. LITERATURE ATTACHMENT: ONE-YEAR RECURRENT TRICUSPID REGURGITATION AFTER SUCCESSFUL TRANSCATHETER EDGE TO EDGE REPAIR: THE TRI-SPA REGISTRY. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF TRICLIP WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING PREVIOUS VALVE SURGERY, DIABETES MELLITUS, ARTERIAL HYPERTENSION, DYSLIPIDEMIA, SMOKER, ATRIAL FIBRILLATION, PACING, CHRONIC KIDNEY DISEASE, CHRONIC LUNG DISEASE, CABG, EDEMA, ASCITES, HEART FAILURE, TRICUSPID REGURGITATION (PRIMARY, SECONDARY). COMPLICATIONS REPORTED INCLUDED RECURRENT TR, REINTERVENTION, DEATH, HEART FAILURE HOSPITALIZATION, SINGLE LEAFLET DEVICE ATTACHMENT; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. B2 DEATH DATE AND B3 EVENT DATE ARE ESTIMATED. LITERATURE: ONE-YEAR RECURRENT TRICUSPID REGURGITATION AFTER SUCCESSFUL TRANSCATHETER EDGE TO EDGE REPAIR: THE TRI-SPA REGISTRY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21620246&pc=NIK | 2017865-2025-27619 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-03-17 04:00 | NIK | GALLANT DR | Adverse Event Without Identified Device or Use Problem | Unspecified Infection | P030054 |  | 1 | Event Description: RELATED MANUFACTURER REPORT NUMBER: 2017865-2025-27620. RELATED MANUFACTURER REPORT NUMBER: 2017865-2025-27621. IT WAS REPORTED THAT A PATIENT PRESENTED WITH E. COLI ENDOCARDITIS WITH VEGETATION ON THE ATRIAL LEAD. THE PATIENT'S IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) AND LEADS WERE EXPLANTED ON (B)(6) 2024. THE PATIENT PASSED AWAY (B)(6) 2024. IT IS UNKNOWN IF THE PATIENT¿S DEATH WAS CAUSED OR CONTRIBUTED BY ANY OF THE PATIENT¿S ABBOTT PRODUCTS OR ANY PROCEDURE INVOLVING THE ABBOTT PRODUCTS. Manufacturer Narrative: DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES. REVIEW OF THE STERILIZATION RECORDS CONFIRMED NORMAL STERILIZATION CYCLES FOR THE PRODUCTS. THE CAUSE OF INFECTION COULD NOT BE DETERMINED. INTERROGATION OF THE DEVICE REVEALED IT WAS ABOVE ELECTIVE REPLACEMENT INDICATOR (ERI) WHEN RECEIVED. THE DEVICE¿S SENSING, PACING, LEAD IMPEDANCE, HIGH VOLTAGE CHARGING, AND HIGH VOLTAGE SHOCK IMPEDANCE WERE TESTED, AND THE DEVICE¿S FUNCTION WAS NORMAL. Manufacturer Narrative: DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES. REVIEW OF THE STERILIZATION RECORDS CONFIRMED NORMAL STERILIZATION CYCLES FOR THE PRODUCTS. THE CAUSE OF INFECTION COULD NOT BE DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21620253&pc=NVN | 2017865-2025-27620 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-03-17 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | Unspecified Infection | P960013 |  | 1 | Event Description: RELATED MANUFACTURER REPORT NUMBER: 2017865-2025-27619. RELATED MANUFACTURER REPORT NUMBER: 2017865-2025-27621. IT WAS REPORTED THAT A PATIENT PRESENTED WITH E. COLI ENDOCARDITIS WITH VEGETATION ON THE ATRIAL LEAD. THE PATIENT'S IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) AND LEADS WERE EXPLANTED ON (B)(6) 2024. THE PATIENT PASSED AWAY (B)(6) 2024. IT IS UNKNOWN IF THE PATIENT¿S DEATH WAS CAUSED OR CONTRIBUTED BY ANY OF THE PATIENT¿S ABBOTT PRODUCTS OR ANY PROCEDURE INVOLVING THE ABBOTT PRODUCTS. Manufacturer Narrative: DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES. REVIEW OF THE STERILIZATION RECORDS CONFIRMED NORMAL STERILIZATION CYCLES FOR THE PRODUCTS. THE CAUSE OF INFECTION COULD NOT BE DETERMINED. VISUAL EXAMINATION FOUND NO MALFUNCTIONS. FULL ANALYSIS NOT NEEDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21620262&pc=NVY | 2017865-2025-27621 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-03-17 04:00 | NVY | DURATA STS OPTIM ACTIVE FIXATION, DF-4 CONNECTOR | Adverse Event Without Identified Device or Use Problem | Unspecified Infection | P950022 |  | 1 | Event Description: RELATED MANUFACTURER REPORT NUMBER: 2017865-2025-27619. RELATED MANUFACTURER REPORT NUMBER: 2017865-2025-27620. IT WAS REPORTED THAT A PATIENT PRESENTED WITH E. COLI ENDOCARDITIS WITH VEGETATION ON THE ATRIAL LEAD. THE PATIENT'S IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) AND LEADS WERE EXPLANTED ON (B)(4) 2024. THE PATIENT PASSED AWAY (B)(4) 2024. IT IS UNKNOWN IF THE PATIENT¿S DEATH WAS CAUSED OR CONTRIBUTED BY ANY OF THE PATIENT¿S ABBOTT PRODUCTS OR ANY PROCEDURE INVOLVING THE ABBOTT PRODUCTS. Manufacturer Narrative: DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES. REVIEW OF THE STERILIZATION RECORDS CONFIRMED NORMAL STERILIZATION CYCLES FOR THE PRODUCTS. THE CAUSE OF INFECTION COULD NOT BE DETERMINED. VISUAL EXAMINATION FOUND NO MALFUNCTIONS. FULL ANALYSIS NOT NEEDED. Manufacturer Narrative: DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES. REVIEW OF THE STERILIZATION RECORDS CONFIRMED NORMAL STERILIZATION CYCLES FOR THE PRODUCTS. THE CAUSE OF INFECTION COULD NOT BE DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21620293&pc=LWR | 2135147-2025-01334 | 2023-03-17 04:00 | Death | ABBOTT MEDICAL | 2025-03-17 04:00 | LWR | TRIFECTA¿ GT VALVE | Insufficient Information | Insufficient Information | P100029 |  | 1 | Event Description: MEDWATCH REPORT: MW5166215. IT WAS REPORTED THAT ON AN UNKNOWN DATE, A 19MM TRIFECTA VALVE WAS IMPLANTED. ON AN UNKNOWN DATE, THE PATIENT PRESENTED FOR A DIVERTICULITIS SCAN AND IT WAS NOTED THAT VALVE WAS MALFUNCTIONING. ON (B)(6) 2023, A VALVE REPLACEMENT PROCEDURE WAS PERFORMED AND THE PATIENT PASSED AWAY. Manufacturer Narrative: AN EVENT OF STRUCTURAL VALVE DETERIORATION AND INSUFFICIENT INFORMATION FOR EXPLANT DEVICE WAS REPORTED. A RETURNED DEVICE ASSESSMENT COULD NOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED FOR ANALYSIS. WITHOUT LOT NUMBER, DEVICE HISTORY RECORD AND COMPLAINT REVIEW DETAILS COULD NOT BE PERFORMED. THE PROVIDED PROCEDURAL IMAGING AND INFORMATION WERE REVIEWED BY MEDICAL AFFAIRS. BASED ON THE INFORMATION RECEIVED, THE CAUSE FOR THE REPORTED STRUCTURAL VALVE DETERIORATION COULD NOT BE DETERMINED. THE REPORTED INFORMATION FOR THE EXPLANT DEVICE WAS INSUFFICIENT. THE PATIENT EFFECTS OF DEATH WAS DUE TO SECOND PROCEDURE, PER CASE DETAIL. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH REGARDS TO MANUFACTURE, DESIGN, OR LABELING. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21620730&pc=NIQ | 2124215-2025-15112 | 2025-02-20 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-03-17 04:00 | NIQ | SYNERGY MEGATRON | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Obstruction/Occlusion |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. THE 85% STENOSED LESION WAS LOCATED IN THE MODERATELY TORTUOUS AND SEVERELY CALCIFIED PROXIMAL CIRCUMFLEX TO MID LEFT MAIN. A 3.50 X 16MM SYNERGY MEGATRON WAS DEPLOYED AND PLAQUE DISPLACEMENT OCCURRED. THE STENT WAS THEN REMOVED FROM THE VESSEL AND THE PHYSICIAN RECANALIZED THE LEFT ANTERIOR DESCENDING ARTERY TO BALLOON AND PLACE ANOTHER STENT. PRIOR TO PLACEMENT OF THE ADDITIONAL STENT, THE PATIENT WENT INTO CARDIAC ARREST AND WAS RESUSCITATED. MEDICATION WAS PROVIDED; HOWEVER, THE PATIENT WENT INTO CARDIAC ARREST AGAIN AND DIED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21620978&pc=OZD | 1220648-2025-26768 | 2025-02-24 05:00 | Death | ABIOMED, INC. | 2025-03-17 04:00 | OZD | IMPELLA 5.5 | Increase in Pressure | Cardiovascular Insufficiency | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. THE PUMP EXPERIENCED HIGH PURGE PRESSURE DURING SUPPORT. LYSIS THERAPY WAS INITIATED AND THE PURGE LEVELS STABILIZED. THE PATIENT PASSED AWAY ON SUPPORT 15 DAYS LATER. THERE IS NOT ENOUGH INFORMATION TO DISASSOCIATE THE DEATH WITH THE USE OF THE IMPELLA. Manufacturer Narrative: THE INVESTIGATION OF HIGH PURGE PRESSURE HAS BEEN COMPLETED. CLINICAL STATES THAT PURGE PRESSURE OF 1070MMHG AND PURGE FLOW OF 1.6ML/HR WAS OBSERVED. THE PUMP AND DATA LOGS WERE NOT RETURNED. THEREFORE, THE ROOT CAUSE OF THE HIGH PURGE PRESSURE ISSUE WAS NOT DETERMINED SINCE PRODUCT AND DATA LOGS WERE NOT RETURNED. (ONLY THE CASSETTE WAS RETURNED) B1 REVISED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26768 TO REFLECT BOTH ADVERSE EVENT AND PRODUCT PROBLEM. D9 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26768, AS THE PUMP WAS NOT RECEIVED FOR INVESTIGATION, ONLY THE CASSETTE. G1 REPORTING CONTACT ADDRESS AND COUNTRY WERE INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26768. G1 MANUFACTURING SITE NAME, ADDRESS, AND FAX NUMBER WERE ERRONEOUSLY ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26768. H3 SHOULD HAVE REFLECTED THE PUMP WAS NOT RETURNED FOR INVESTIGATION ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26768. Manufacturer Narrative: THE IMPELLA DEVICE WAS RETURNED, AND AN EVALUATION IS UNDERWAY. UPON COMPLETION OF OUR ANALYSIS, A SUPPLEMENTAL REPORT WILL BE FILED. AS NOTED IN THE IMPELLA IFU: IMPELLA 5.5 WITH SMARTASSIST SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA CATHETER ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP. CLINICIANS SHOULD MONITOR MOTOR CURRENT AND CONSIDER REPLACING THE IMPELLA CATHETER WHENEVER A RISE IN MOTOR CURRENT IS SEEN.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21621174&pc=DSQ | 2916596-2025-01537 | 2025-02-18 05:00 | Death | THORATEC CORPORATION | 2025-03-17 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTI-SYSTEM ORGAN FAILURE. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. THE PUMP WAS WORKING AS EXPECTED. NO PUMP RELATED ISSUES WERE REPORTED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT'S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. IT WAS REPORTED NO AUTOPSY WAS PERFORMED, AND HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), WAS NOT EXPLANTED FOR EVALUATION. NO FURTHER INFORMATION WAS PROVIDED BY THE CUSTOMER. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REVISION D IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (INCLUDING RIGHT HEART, RESPIRATORY, RENAL, AND HEPATIC FAILURE) AND DEATH AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: H6- 3261 - MULTIPLE ORGAN DYSFUNCTION SYNDROME. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21621228&pc=DSQ | 2916596-2025-01494 | 2025-02-25 05:00 | Death | THORATEC CORPORATION | 2025-03-17 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT OUTCOME THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. IT WAS UNKNOWN IF AN AUTOPSY WOULD BE PERFORMED. THE PUMP WOULD NOT BE EXPLANTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE CUSTOMER COMMUNICATED THAT NO ADDITIONAL INFORMATION WILL BE PROVIDED DUE TO PATIENT PRIVACY LAWS. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED THROUGH THE MANUFACTURING EXECUTION SYSTEM (MES) AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. G, IS CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21621339&pc=DSQ | 2916596-2025-01493 | 2025-02-14 05:00 | Death | THORATEC CORPORATION | 2025-03-17 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED A LOW FLOW ALARM. IN THE EVENING, THE PATIENT'S FAMILY CONFIRMED THAT THE PATIENT WAS ALIVE. WHEN THE FAMILY WENT TO CHECK ON THE PATIENT AFTER HEARING AN ALARM AT NIGHT, THEY FOUND THAT THE PATIENT'S BODY WAS COLD, AND THE AMBULANCE ARRIVED AND CONFIRMED THAT THE PATIENT WAS DECEASED.Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY FROM HEMODYNAMIC INSUFFICIENCY.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT NO LOG FILES WERE AVAILABLE FOR EVALUATION; HOWEVER, THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: SECTION D4: MODEL AND CATALOG NUMBERS CORRECTED. SECTION E1: CUSTOMER SITE WAS (B)(6) HOSPITAL. MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. FURTHERMORE, THE REPORTED LOW FLOW ALARM WAS UNABLE TO BE CONFIRMED AS NO LOG FILES WERE SUBMITTED FOR REVIEW AND THE PRODUCT WAS NOT RETURNED FOR EVALUATION. WHETHER THE PATIENT'S OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED. THE PATIENT WAS CREMATED, AND THE PUMP WAS NOT EXPLANTED FOR EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. THIS SECTION ALSO PROVIDES AN EXPLANATION OF PUMP PARAMETERS, INCLUDING FLOW. THIS SECTION EXPLAINS THAT PUMP FLOW IS A CALCULATED VALUE THAT IS ESTIMATED BASED ON PUMP POWER. SECTION 4 OF THE IFU, ¿SYSTEM MONITOR¿, PROVIDES INFORMATION ABOUT THE PUMP FLOW DISPLAY AND THE LOW FLOW HAZARD ALARM CONDITION. THIS SECTION STATES THAT THE LOW FLOW HAZARD ALARM WILL BE TRIGGERED WHEN THE ESTIMATED PUMP FLOW IS LESS THAN 2.5 LPM AND EXPLAINS THAT CHANGES IN PATIENT CONDITIONS CAN RESULT IN LOW FLOW. SECTION 7 OF THE IFU, ¿ALARMS AND TROUBLESHOOTING¿, AND SECTION 5 OF THE PATIENT HANDBOOK, "ALARMS AND TROUBLESHOOTING" OUTLINE SYSTEM CONTROLLER ALARMS, AS WELL AS HOW TO RESPOND TO EACH ALARM CONDITION. THESE DOCUMENTS INSTRUCT THE USER THAT IN THE EVENT OF A LOW FLOW HAZARD ALARM, CALL YOUR HOSPITAL CONTACT IMMEDIATELY FOR DIAGNOSIS AND INSTRUCTIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21621389&pc=DSQ | 2916596-2025-01538 | 2025-02-27 05:00 | Death | THORATEC CORPORATION | 2025-03-17 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Cardiac Arrest; Tachycardia; Ventricular Fibrillation | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO CARDIAC ARREST. THE PATIENT DIED AFTER RECURRENCE OF VENTRICULAR TACHYCARDIA THAT DEGENERATED INTO VENTRICULAR FIBRILLATION. THE PUMP WAS WORKING AS INTENDED DURING TIME OF SUPPORT, THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. G IS CURRENTLY AVAILABLE. SECTION 1, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING DEATH AND ARRHYTHMIA. SECTION 6, "PATIENT CARE AND MANAGEMENT," LISTS ARRHYTHMIA AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21621518&pc=DSQ | 21621518 | 2025-02-25 05:00 | Death | THORATEC CORPORATION | 2025-03-17 04:00 | DSQ | THORATEC® HEARTMATE 3¿ |  | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: PATIENT PRESENTED TO ESTABLISH HOSPICE AND WAS ADMITTED TO INPATIENT HOSPICE. PATIENT WAS STARTED ON COMFORT CARE MEDICATIONS AND PASSED AT 0303 WITH FAMILY AT SIDE. PROVIDERS HAVE EXPRESSED CONCERNS OVER THE LAST FEW MONTHS REGARDING PATIENT'S DRIVELINE INFECTION AND NEED FOR ADMISSION FOR IV ANTIBIOTICS. PATIENT REFUSED ADMISSION ON MULTIPLE OCCASIONS OVER THE LAST COUPLE OF MONTHS, STATING THAT HE UNDERSTOOD THE RISKS BUT WOULD PREFER TO CONTINUE CARE ON AN OUTPATIENT BASIS TO SPEND AS MUCH MEANINGFUL TIME WITH HIS FAMILY AS POSSIBLE. UPON ARRIVAL TO EMERGENCY DEPARTMENT, FAMILY WAS COUNSELED ABOUT WHAT TO EXPECT AT END OF LIFE, ROLE OF MEDICATIONS AND RISK FOR LIMITED PROGNOSIS. SPIRITUAL NEEDS ADDRESSED. PATIENT ULTIMATELY PASSED ON INPATIENT HOSPICE SURROUNDED BY FAMILY AT BEDSIDE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21621537&pc=DSQ | 2916596-2025-01588 | 2025-02-13 05:00 | Death | THORATEC CORPORATION | 2025-03-17 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Tachycardia; Ventricular Fibrillation; Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2025 DUE TO MULTIORGAN FAILURE.Event Description: ADDITIONAL INFORMATION RECEIVED STATED THAT THE PATIENT HAD AN IMPLANTABLE CARDIOVERTER-DEFIBRILLATOR (ICD) SHOCK AND PRESENTED TO OUTSIDE HOSPITAL. THEY HAD ADDITIONAL EPISODES OF VENTRICULAR TACHYCARDIA (VT) AND VENTRICULAR FIBRILLATION (VF). THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6) AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE CONTROLLER LOG FILES WERE RETRIEVED FROM THE RETURNED CONTROLLER AND REVIEWED. THE SYSTEM APPEARED TO BE OPERATING AS INTENDED AT THE SET SPEED. PUMP SUPPORT WAS DISCONTINUED IN THE EARLY MORNING OF (B)(6) 2022 PER THE LOG FILE TIMESTAMPS. THERE WERE NO NOTABLE FINDINGS OR ABNORMALITIES. IT WAS REPORTED THAT THE PUMP WILL NOT BE RETURNED FOR EVALUATION. NO FURTHER INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO RELEVANT DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ OF THIS DOCUMENT LISTS DEATH AND CARDIAC ARRHYTHMIA AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. THE IFU ALSO LISTS MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RIGHT HEART FAILURE, RESPIRATORY FAILURE, RENAL DYSFUNCTION, AND HEPATIC DYSFUNCTION) AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6 ¿PATIENT CARE AND MANAGEMENT¿ LISTS ARRHYTHMIA AS A POTENTIAL LATE POSTIMPLANT COMPLICATION THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21621569&pc=NPT | 2015691-2025-01980 | 2025-02-25 05:00 | Death | EDWARDS LIFESCIENCES | 2025-03-17 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Rupture | P140031 |  | 1 | Event Description: AS REPORTED BY OUR POLISH AFFILIATES, DURING IMPLANT OF A 26MM SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE IMPLANT IN AORTIC POSITION, AFTER THE VALVULOPLASTY WITH A 23MM BALLOON, IMPLANTATION OF THE VALVE WENT SMOOTHLY. THE FIRST ANGIOGRAPHIC CHECK REVEALED A MASSIVE TRANSVALVULAR RETROGRADE WAVE AND JUST SHORTLY THEREAFTER, THE PATIENT'S BLOOD PRESSURE DROPPED. AN ECHOCARDIOGRAPHIC CHECK CONFIRMED PROPER VALVE FUNCTION BUT ALSO DETECTED THE PRESENCE OF FLUID IN PERICARDIAL SAC. FOLLOWING FLUID REMOVAL, A SECOND ANGIOGRAPHIC CHECK WAS PERFORMED, WHICH SHOWED NO TRACE OF PARAVALVULAR LEAK AND NO SIGNS OF ANNULUS OR AORTIC RAPTURE. DUE TO THE PATIENT'S INSTABILITY AND RISK OF FURTHER PERICARDIAL BLEEDING, A DECISION WAS MADE TO CONVERT TO OPEN SURGERY. THE CAUSE OF BLEEDING WAS IDENTIFIED AS AN ANNULUS RUPTURE AND MINOR BLEEDING IN THE SINOTUBULAR JUNCTION (STJ) AREA. THIS RUPTURE AND BLEEDING WERE SEALED WITH THE SAPIEN 3 ULTRA VALVE IN THE CORRECT POSITION AND THE PATIENT, IN STABLE CONDITION, WAS TRANSFERRED TO THE INTENSIVE CARE UNIT. THE PATIENT REMAINED STABLE BUT FEW DAYS LATER PASSED AWAY. Manufacturer Narrative: THE EVENT REPORTED IS AN ANTICIPATED IN THE RISK MANAGEMENT DOCUMENTATION FOR TRANSCATHETER HEART VALVE PROCEDURES. A PREVIOUS INVESTIGATION INTO THIS TYPE OF EVENT IS CAPTURED IN AN EDWARDS LIFESCIENCES TECHNICAL SUMMARY AND APPLIES TO THIS COMPLAINT. ADDITIONAL ASSESSMENT OF THE FAILURE MODE IS NOT REQUIRED AT THIS TIME. DUE TO THE UNAVAILABILITY OF THE COMPLAINT DEVICE, ENGINEERING WAS UNABLE TO PERFORM ANY VISUAL INSPECTION, FUNCTIONAL TESTING, OR DIMENSIONAL ANALYSIS. PER THE TECHNICAL SUMMARY: THE IFU, CURRENT RISK MITIGATION'S INCLUDING DESIGN AND MANUFACTURING CONTROLS, AND TRAINING MANUALS HAVE BEEN REVIEWED, AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. ACCORDING TO THE DEVICE TRAINING MANUALS, RISK FACTORS FOR AORTIC DISSECTION, CARDIAC/AORTIC HEMATOMA, OR ANNULAR RUPTURE DURING THE TVR PROCEDURE INCLUDE SIGNIFICANT VALVE OVER SIZING, SEVERELY OBLITERATED SINUSES OF VALSALVA, PORCELAIN AORTA AND/OR PRESENCE OF BULKY CALCIFICATION, AND NARROW CALCIFIED SINOTUBULAR JUNCTION. IN ADDITION, ADVANCED AGE, FEMALE GENDER, SMALL BODY WEIGHT, AND STEROID DEPENDENCY CAN ALSO BE CONTRIBUTING FACTORS. THE SAPIEN TRANSCATHETER HEART VALVE (ALL MODELS) RELIES CALCIUM TO SECURELY ANCHOR IN THE LANDING ZONE. DESPITE THIS BENEFICIAL ASPECT OF CALCIUM, BULKY CALCIUM CAN INCREASE THE RISK OF CALCIFIC NODULE DISPLACEMENT INTO THE VASCULATURE, WHICH CAN LEAD TO VASCULAR INJURY. AT TIMES, THE EXTENT AND DISTRIBUTION OF CALCIUM CAN IMPAIR EASE OF DELIVERY OF THE VALVE, CORRECT POSITIONING OF THE VALVE, DEPLOYMENT OF THE VALVE, AND PROCEDURAL SUCCESS. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST/INDICATE A COMPLICATED PATIENT ANATOMY WITH A HORIZONTAL AORTA AND BULKY CALCIFICATION OF NATIVE LEAFLETS COULD HAVE CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21621914&pc=EOQ | 2955842-2025-08656 | 2025-02-12 05:00 | Death | INTUITIVE SURGICAL, INC | 2025-03-17 04:00 | EOQ | ION | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Dyspnea; Insufficient Information | K182188 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN ION ENDOLUMINAL LUNG BIOPSY PROCEDURE AND SUBSEQUENTLY EXPIRED AT HOME. THE PHYSICIAN STATED THERE WERE NO MAJOR ISSUES OR COMPLICATIONS DURING OR AFTER THE PROCEDURE; THE PHYSICIAN STATED THAT RADIAL ENDOBRONCHIAL ULTRASOUND (EBUS) WAS UTILIZED MORE FREQUENTLY DURING THE PROCEDURE. THE PATIENT WAS DISCHARGED ON THE SAME DAY AS THE PROCEDURE, WHICH WAS COMPLETED AS PLANNED WITHOUT ANY DELAYS. THERE WAS NO DEVICE MALFUNCTION ASSOCIATED WITH THE EVENT. THE PHYSICIAN STATED THE PATIENT PRESENTED TO A COMMUNITY HOSPITAL THREE DAYS POST-PROCEDURE WITH SHORTNESS OF BREATH. A CHEST CT SCAN AT THE TIME SHOWED NO BLEEDING OR ACUTE CHANGES, AND OTHER UNSPECIFIED WORKUPS WERE UNREMARKABLE; THE PATIENT WAS DISCHARGED AND SUBSEQUENTLY EXPIRED AT HOME ON AN UNKNOWN DATE. THE PHYSICIAN REPORTED THAT THE PATIENT OUTCOME WAS NOT RELATED TO THE ION PRODUCTS OR THE PROCEDURE AND THE CAUSE OF DEATH WAS THOUGHT TO HAVE BEEN DUE TO A CARDIAC EVENT. Manufacturer Narrative: SYSTEM LOG REVIEW FOR THE PROCEDURE DATE FOUND THERE WERE NO RELATED SYSTEM ERRORS TO HAVE OCCURRED DURING THE PROCEDURE THAT WERE RELEVANT TO THE REPORTED POST-PROCEDURE EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT A PATIENT OF UNKNOWN AGE UNDERWENT AN ION LUNG BIOPSY. THE PROCEDURE WAS COMPLETED, AND THE PATIENT WAS DISCHARGED HOME THE SAME DAY. THE PATIENT WAS SEEN IN THE ER AT ANOTHER HOSPITAL 3 DAYS AFTER THE PROCEDURE WITH SHORTNESS OF BREATH. A CT SCAN WAS REPORTEDLY WITHOUT SIGNIFICANT FINDINGS, AND THE PATIENT WAS DISCHARGED HOME. IT IS UNKNOWN WHAT OTHER EVALUATION WAS PERFORMED. THE PATIENT THEN DIED AT HOME ON AN UNKNOWN DATE FROM AN UNKNOWN CAUSE. BASED ON THE AVAILABLE DATA, THE REPORTED EVENT WAS NEITHER PROCEDURE NOR DEVICE RELATED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21622328&pc=MHX | 1218950-2025-000163 | 2025-02-20 05:00 | Death | PHILIPS MEDICAL SYSTEMS | 2025-03-17 04:00 | MHX | PATIENT INFORMATION CENTER IX | Defective Alarm | Hypoxia; Heart Failure/Congestive Heart Failure | K183387 |  | 1 | Event Description: IT WAS REPORTED THERE WAS AN ALARM ISSUE AND THE PATIENT PASSED AWAY. THE CUSTOMER INDICATED AN ALLEGED ALARM MALFUNCTION BUT COULD NOT PROVIDE WHAT ALARMS FAILED. IT WAS INDICATED THE VASCULAR AMPUTEE PATIENT WAS BEING MONITORED DUE TO CARDIAC DECOMPENSATION RELATED TO AORTIC STENOSIS. THE HOSPITAL REPORT FORM WAS PROVIDED, INDICATING AROUND 0645 THE NIGHT NURSES LEFT THE PATIENT'S ROOM AFTER PERFORMING A TREATMENT AND THE PATIENT WAS WELL WITH OXYGEN AND TELEMETRY IN PLACE. AT 0715 THE DAY SHIFT STARTED HER SHIFT AND HEARD THE ALARM SOUNDING FROM THE PATIENT ROOM. UPON ENTERING THE ROOM, SHE FOUND THE PATIENT IN CARDIAC ARREST WITH HIS OXYGEN AND ECG ELECTRODES REMOVED. CPR WAS INITIATED BUT THE RESUSCITATION WAS UNSUCCESSFUL AND THE PATIENT PASSED AWAY. THE CUSTOMER CONFIRMED THE REPORTED DEATH DATE AND TIME. IN ADDITION, THE CAUSE OF DEATH WAS THOUGHT TO BE RELATED TO HYPOXIA AS A PART OF OVERALL CARDIAC DECOMPENSATION AND A VENTRICULAR RHYTHM DISORDER IN ADVANCED HEART DISEASE AS WELL AS POSSIBLE TECHNICAL ISSUES. Manufacturer Narrative: E1: (B)(6). PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE. Manufacturer Narrative: A PHILIPS PRODUCT SPECIALIST ENGINEER (PSE) AND A CLINICAL APPLICATION SPECIALIST (CAS) REVIEWED THE AUDIT LOGS PROVIDED BY THE CUSTOMER. BASED ON THE LIMITED INFORMATION PROVIDED THE LOGS SHOW THAT ALARMS DID SOUND BETWEEN 0645 AND 0745 ON FEBRUARY 20TH IN ROOM 223 BUT THE LOGS SHOW THERE WERE NO ALARMS ON BED 211 OR ANY ACTIVITY ON FEBRUARY 20TH DURING THAT TIME. THE CUSTOMER WAS PROVIDED INFORMATION ABOUT THE LOGS. IT HAS BEEN CONCLUDED THAT NO FURTHER ACTION IS REQUIRED AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED THE COMPLAINT FILE WILL BE REOPENED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21622992&pc=NVN | 2017865-2025-27648 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-03-17 04:00 | NVN | TENDRIL SDX LEAD | Failure to Capture | Insufficient Information | P960013 |  | 1 | Event Description: VOLUNTARY MEDWATCH FORM RECEIVED NOTES THAT A PATIENT'S RIGHT VENTRICULAR (RV) LEAD WAS NOT CAPTURING. THE PATIENT PASSED AWAY; THE CAUSE OF DEATH IS NOT KNOWN. Manufacturer Narrative: VOLUNTARY MEDWATCH MDR REPORT #: MW5167560.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21624638&pc=FKX | 0008030665-2025-00612 | 2025-03-07 05:00 | Death | ERIKA DE REYNOSA, S.A. DE C.V. | 2025-03-17 04:00 | FKX | LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K222318 |  | 1 | Event Description: A PERITONEAL DIALYSIS REGISTERED NURSE (PDRN) REPORTED TO FRESENIUS CUSTOMER SUPPORT THAT A PATIENT EXPIRED. THE PATIENT WAS ON PD TREATMENT WHEN HE PASSED PEACEFULLY WHILE SLEEPING. FOLLOW-UP WITH THE PATIENT¿S PD REGISTERED NURSE (PDRN) REVEALED THE PATIENT BEGAN HIS TREATMENT ON THE EVENING OF (B)(6) 2025 WITHOUT ANY KNOWN ISSUES. THE FOLLOWING MORNING THE PATIENT¿S SPOUSE HEARD THE LIBERTY SELECT CYCLER ALARM INDICATING THE TREATMENT HAD BEEN COMPLETED, AFTER WHICH SHE TURNED OFF THE DEVICE AND RETURNED TO BED. THE PATIENT¿S SPOUSE AWOKE AT APPROXIMATELY 9-9:30 AM TO DISCOVER THE PATIENT HAD EXPIRED PEACEFULLY IN HIS SLEEP ON THE MORNING OF (B)(6) 2025. PER THE PDRN, THE CAUSE OF DEATH WAS "NATURAL CAUSES," AND WAS UNRELATED TO ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S).Event Description: A PERITONEAL DIALYSIS REGISTERED NURSE (PDRN) REPORTED TO FRESENIUS CUSTOMER SUPPORT THAT A PATIENT EXPIRED. THE PATIENT WAS ON PD TREATMENT WHEN HE PASSED PEACEFULLY WHILE SLEEPING. FOLLOW-UP WITH THE PATIENT¿S PD REGISTERED NURSE (PDRN) REVEALED THE PATIENT BEGAN HIS TREATMENT ON THE EVENING OF (B)(6) 2025 WITHOUT ANY KNOWN ISSUES. THE FOLLOWING MORNING THE PATIENT¿S SPOUSE HEARD THE LIBERTY SELECT CYCLER ALARM INDICATING THE TREATMENT HAD BEEN COMPLETED, AFTER WHICH SHE TURNED OFF THE DEVICE AND RETURNED TO BED. THE PATIENT¿S SPOUSE AWOKE AT APPROXIMATELY 9-9:30 AM TO DISCOVER THE PATIENT HAD EXPIRED PEACEFULLY IN HIS SLEEP ON THE MORNING OF (B)(6) 2025. PER THE PDRN, THE CAUSE OF DEATH WAS "NATURAL CAUSES," AND WAS UNRELATED TO ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S). Manufacturer Narrative: ADDITIONAL INFORMATION: D9, G4, H3. PLANT INVESTIGATION: THE ACTUAL DEVICE WAS RETURNED TO THE MANUFACTURER FOR PHYSICAL EVALUATION. AN EXTERNAL VISUAL INSPECTION WAS PERFORMED ON THE CYCLER WITH NO PHYSICAL DAMAGE NOTED. ALTHOUGH THERE WAS EVIDENCE OF DRIED FLUID PRESENT WITHIN THE CASSETTE COMPARTMENT, THERE WERE NO BURRS OR SHARP EDGES IN THE CASSETTE AREA THAT COULD HAVE PUNCTURED THE CASSETTE MEMBRANE. THERE WERE NO VISUAL INDICATIONS OF PARTICULATES WITHIN THE CASSETTE AREA. A SIMULATED TREATMENT USING (AS-RECEIVED) TREATMENT SETTINGS WITH REDUCED DWELL TIMES WAS PERFORMED AND COMPLETED WITHOUT FAILURES. NO FLUID LEAKS IN THE TEST CASSETTE DURING THE TREATMENT TEST. THE CYCLER UNDERWENT AND PASSED A SYSTEM AIR LEAK TEST AND A VALVE ACTUATION TEST. AN INTERNAL VISUAL INSPECTION OF THE RETURNED CYCLER WAS PERFORMED. THERE WERE NO VISUAL INDICATIONS OF DRIED FLUID UNDER THE PUMP ASSEMBLY ON THE BOTTOM COVER. A REVIEW OF THE DEVICE MANUFACTURING RECORDS WAS CONDUCTED BY THE MANUFACTURER. THERE WERE NO DEVIATIONS OR NON-CONFORMANCES DURING THE MANUFACTURING PROCESS. IN ADDITION, A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND VERIFIED THAT THE RESULTS OF THE IN-PROGRESS AND FINAL QUALITY CONTROL (QC) TESTING MET ALL REQUIREMENTS. UPON COMPLETION OF THE EVALUATION, THERE WERE NO MALFUNCTIONS THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE CYCLER PERFORMED AS DESIGNED AND AN ASSOCIATED CAUSE COULD NOT BE DETERMINED. Manufacturer Narrative: CLINICAL INVESTIGATION: A TEMPORAL RELATIONSHIP EXISTS BETWEEN CCPD THERAPY UTILIZING A LIBERTY SELECT CYCLER, AND THE SERIOUS ADVERSE EVENT OF DEATH, AS THE PATIENT WAS ACTIVELY UNDERGOING TREATMENT WHEN THE SERIOUS ADVERSE EVENT OCCURRED. THE PDRN REPORTED THE PATIENT¿S PRIMARY CAUSE OF DEATH WAS DUE TO NATURAL CAUSES AND WAS UNRELATED TO THE PATIENT¿S UTILIZATION OF ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S). THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. ADULTS WITH ESRD HAVE MORTALITY RATES UP TO 30-FOLD HIGHER THAN THE GENERAL POPULATION. BASED ON THE INFORMATION AVAILABLE, THE LIBERTY SELECT CYCLER CAN BE DISASSOCIATED FROM THE SERIOUS ADVERSE EVENT. THERE IS NO ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) CAUSED OR CONTRIBUTED TO THE SERIOUS ADVERSE EVENT. FURTHERMORE, THERE WAS NO REPORT A FRESENIUS PRODUCT(S) AND/OR DEVICE(S) FAILED TO MEET THE USERS¿ EXPECTATIONS AND/OR THE MANUFACTURERS¿ SPECIFICATIONS. THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21625618&pc=DSQ | 2916596-2025-01496 | 2025-02-27 05:00 | Death | THORATEC CORPORATION | 2025-03-17 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY IN THE INTENSIVE CARE UNIT AFTER TRANSITIONING TO COMFORT CARE. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: CORRECTED DATA: SECTION G2: CORRECTED REPORT SOURCE TO INCLUDE "OTHER". MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE CAUSE OF DEATH WAS NOT PROVIDED. WHETHER THE PATIENT¿S OUTCOME WAS CONSIDERED TO BE DEVICE OR THERAPY RELATED WAS NOT PROVIDED. THE DEVICE WAS NOT RETURNED FOR EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21625649&pc=DSQ | 2916596-2025-01648 | 2025-02-27 05:00 | Death | THORATEC CORPORATION | 2025-03-17 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTIPLE SYSTEM ORGAN FAILURE. THE DEVICE OPERATED AS INTENDED AND THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. THE DEVICE WOULD NOT RETURN. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RESPIRATORY FAILURE, RIGHT HEART FAILURE, RENAL FAILURE, HEPATIC DYSFUNCTION) AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: A1-A4: PATIENT INFORMATION CANNOT BE PROVIDED DUE TO PERSONAL DATA PRIVACY LEGISLATION/POLICY. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21625676&pc=NPT | 2015691-2025-01989 | 2025-02-20 05:00 | Death | EDWARDS LIFESCIENCES | 2025-03-17 04:00 | NPT | COMMANDER DELIVERY SYSTEM | Burst Container or Vessel; Difficult or Delayed Positioning; Inflation Problem | Foreign Body In Patient; Aortic Valve Insufficiency/ Regurgitation | P140031 |  | 1 | Event Description: AS REPORTED BY THE FIELD CLINICAL SPECIALIST (FCS), DURING THE TRANSFEMORAL TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) PROCEDURE WITH A 26 MM SAPIEN 3 ULTRA RESILIA VALVE, THERE WAS DIFFICULTY CROSSING THE NATIVE VALVE WITH THE 26 MM COMMANDER DELIVERY SYSTEM AND VALVE DUE TO TORTUOSITY AND CALCIUM BURDEN. THE PATIENT WAS NOTED TO HAVE TORTUOSITY IN THE ILIAC ARTERIES AND DESCENDING THORACIC AORTA. THE 14FR ESHEATH+ HAD BEEN SUCCESSFULLY PLACED IN THE PATIENT. AFTER SEVERAL TECHNIQUES WERE USED, AN ADDITIONAL ONE CC WAS ADDED TO THE COMMANDER DELIVERY SYSTEM BALLOON AND THE COMMANDER DELIVERY SYSTEM AND VALVE WAS ABLE TO CROSS THE NATIVE VALVE. THE PUSHER WAS WITHDRAWN AND VALVE DEPLOYMENT BEGAN. DURING VALVE DEPLOYMENT, IT APPEARED THAT THE VALVE WAS OFFSET FROM THE DELIVERY SYSTEM BALLOON 3-4 MM VENTRICULARLY. AS THE DELIVERY SYSTEM BALLOON WAS BEING INFLATED, THE VALVE APPEARED TO BE MORE AORTIC THAN INTENDED AND AS A RESULT, THE COMMANDER DELIVERY SYSTEM WAS PUSHED FORWARD. AS IT WAS BEING PUSHED, THE BALLOON DISENGAGED FROM THE SEMI-DEPLOYED VALVE TO THE POINT WHERE ONLY THE SHOULDERS OF THE DELIVERY SYSTEM BALLOON WERE IN CONTACT WITH THE VENTRICULAR PORTION OF THE VALVE. AT THIS POINT, THE DELIVERY SYSTEM BALLOON RUPTURED. THE DELIVERY SYSTEM BALLOON WAS ABLE TO BE WITHDRAWN FROM THE 14FR ESHEATH+ WITHOUT ANY ISSUES. THE PARTIALLY DEPLOYED VALVE WAS BROUGHT BACK TO THE LEVEL OF THE ABDOMINAL AORTA BELOW THE RENAL ARTERIES AND IT WAS SUCCESSFULLY DEPLOYED IN THE ABDOMINAL AORTA. THE PATIENT APPEARED TO BE IN STABLE CONDITION POST TAVR PROCEDURE. IT WAS BELIEVED THAT THE EVENT WAS DUE TO PATIENT TORTUOSITY IT WAS INDICATED THAT FOR ANY FUTURE TAVR PROCEDURES, IT WOULD BE PERFORMED VIA A CAROTID APPROACH. SUBSEQUENTLY, POST TAVR PROCEDURE, THE PATIENT WAS OBSERVED WITH INCREASING NATIVE VALVE AORTIC REGURGITATION. AS A RESULT, THE PATIENT WAS TAKEN TO THE OPERATING ROOM (OR) FOR TREATMENT. THE PATIENT THEN CODED AND PASSED AWAY. ADDITIONAL INFORMATION WAS RECEIVED VIA MEDICAL RECORDS REVEALING THE PATIENT HAD BEEN ADMITTED TO THE HOSPITAL AFTER A SERIES OF FALLS. DURING THE HOSPITALIZATION, THE PATIENT WAS FOUND TO HAVE SEVERE AORTIC STENOSIS AND WAS EVALUATED FOR A TAVR IMPLANT. SUBSEQUENTLY, THE PATIENT UNDERWENT THE TAVR PROCEDURE WITH A 26 MM SAPIEN 3 ULTRA RESILIA VALVE WHICH WAS UNSUCCESSFUL DUE TO THE RUPTURE OF THE DELIVERY SYSTEM BALLOON ON DEPLOYMENT. POST PROCEDURE, THE PATIENT EXPERIENCED MODERATE AORTIC INSUFFICIENCY (AI) AND WAS TAKEN TO THE OPERATING ROOM (OR) FOR EMERGENT SURGICAL VALVE REPLACEMENT VIA STERNOTOMY. UPON ARRIVAL TO THE UNIT POST SURGICAL VALVE REPLACEMENT, THE PATIENT WAS EXTREMELY HYPOTENSIVE AND FOUND TO BE WITHOUT A PULSE. CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED, THE PACEMAKER WAS TURNED OFF AND THE PATIENT WAS FOUND TO BE IN VENTRICULAR FIBRILLATION. A DECISION WAS MADE TO OPEN THE PATIENT'S CHEST TO LOOK FOR BLEEDING, BUT NO SOURCE OF BLEEDING WAS IDENTIFIED. RETURN OF SPONTANEOUS CIRCULATION WAS NOT OBSERVED AND RESUSCITATIVE EFFORTS WERE STOPPED. THE EXACT CAUSE OF DEATH WAS NOT PROVIDED. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR CORRECTION AND INCLUDES ADDITIONAL INFORMATION BASED ON THE INVESTIGATION. THE FOLLOWING SECTIONS OF THIS REPORT HAVE BEEN CORRECTED/UPDATED: H6 (TYPE OF INVESTIGATION, INVESTIGATION FINDINGS, INVESTIGATION CONCLUSIONS) AND H11 (PROVIDE NARRATIVE/DATA). THE EVENTS REPORTED ARE ANTICIPATED IN THE RISK MANAGEMENT DOCUMENTATION FOR TRANSCATHETER HEART VALVE PROCEDURES. ADDITIONAL ASSESSMENT OF THE FAILURE MODES IS NOT REQUIRED AT THIS TIME. THE DEVICE WAS NOT RETURNED FOR EVALUATION AS IT WAS DISCARDED. DUE TO THE UNAVAILABILITY OF THE DEVICE, NO VISUAL INSPECTION, FUNCTIONAL TESTING, OR DIMENSIONAL ANALYSIS COULD BE PERFORMED. THERE WAS IMAGERY PROVIDED FOR EVALUATION. A REVIEW OF THE IMAGERY REVEALED THERE WAS TORTUOSITY IN THE DESCENDING THORACIC AORTA. THERE WAS ALSO CALCIFICATION IN THE LANDING ZONE. THE INSTRUCTIONS FOR USE (IFU), AND TRAINING MANUALS HAVE BEEN REVIEWED. NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. THROUGH EXTENSIVE COMPLAINT INVESTIGATIONS, DIFFICULTY OR INABILITY IN TRACKING THE DELIVERY SYSTEM TO THE LANDING ZONE, INCLUDING TRAVERSING THE AORTIC ARCH AND CROSSING THE NATIVE ANNULUS/BIOPROSTHESIS HAS NOT BEEN ASSOCIATED WITH DEVICE MALFUNCTIONS OR MANUFACTURING NONCONFORMANCES. RATHER, THE ROOT CAUSE FOR THESE EVENTS HAVE HISTORICALLY BEEN DUE TO PATIENT FACTORS, POOR GUIDEWIRE MANAGEMENT, DAMAGED GUIDEWIRE, AND/OR EXCESSIVE MANIPULATION OF THE FLEX WHEEL. IN THIS CASE, THE DIFFICULTY ENCOUNTERED CROSSING THE NATIVE ANNULUS WAS UNABLE TO BE CONFIRMED. THE AVAILABLE INFORMATION SUGGESTS THAT PATIENT FACTORS (TORTUOSITY AND CALCIFICATION) LIKELY CONTRIBUTED TO THE EVENT AS IT WAS REPORTED THAT "THERE WAS DIFFICULTY CROSSING THE NATIVE VALVE WITH THE 26 MM COMMANDER DELIVERY SYSTEM AND VALVE DUE TO TORTUOSITY AND CALCIUM BURDEN. THE PATIENT WAS NOTED TO HAVE TORTUOSITY IN THE ILIAC ARTERIES AND DESCENDING THORACIC AORTA. IT WAS BELIEVED THAT THE EVENT WAS DUE TO THE TORTUOSITY, AND IT WAS INDICATED THAT FOR ANY FUTURE TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) PROCEDURES, IT WOULD BE PERFORMED VIA A CAROTID APPROACH." SUBSEQUENTLY, "ADDITIONAL INFORMATION WAS PROVIDED REVEALING THE PATIENT HAD SEVERE TORTUOSITY AND MODERATE CALCIFICATION." ADDITIONALLY, PER REVIEW OF THE PROVIDED IMAGERY, TORTUOSITY WAS PRESENT IN THE DESCENDING THORACIC AORTA. PATIENT FACTORS (I.E. CALCIFICATION AND TORTUOSITY) MAY CAUSE CONSTRAINED SECTIONS OF THE ANATOMY THAT INTERFERES WITH PASSAGE OF THE DELIVERY SYSTEM AND CAUSES DIFFICULTY DURING TRACKING/CROSSING. IN REGARD TO THE VALVE NOT BEING BETWEEN THE ALIGNMENT MARKERS AND DEPLOYED THAT RESULTED IN THE VALVE BEING ONLY SEMI-DEPLOYED AND HAVING TO THEN BE DEPLOYED IN THE ABDOMINAL AORTA DUE TO THE COMMANDER DELIVERY SYSTEM BALLOON BURST, THIS EVENT WAS ALSO UNABLE TO BE CONFIRMED. A REVIEW OF THE DEVICE HISTORY RECORDS (DHR) DID NOT PROVIDE ANY INDICATION THAT A MANUFACTURING NON-CONFORMANCE WOULD HAVE CONTRIBUTED TO THE EVENT. A REVIEW OF THE IFU/TRAINING MATERIALS REVEALED NO DEFICIENCIES. FURTHERMORE, NO ABNORMALITIES WERE REPORTED DURING DEVICE UNPACKING OR PREPARATION. AS REPORTED, "DURING VALVE DEPLOYMENT, IT APPEARED THAT THE VALVE WAS OFFSET FROM THE DELIVERY SYSTEM BALLOON 3-4 MM VENTRICULARLY. AS THE DELIVERY SYSTEM BALLOON WAS BEING INFLATED, THE VALVE APPEARED TO BE MORE AORTIC THAN INTENDED AND AS A RESULT, THE COMMANDER DELIVERY SYSTEM WAS PUSHED FORWARD. AS IT WAS BEING PUSHED, THE BALLOON DISENGAGED FROM THE SEMI-DEPLOYED VALVE TO THE POINT WHERE ONLY THE SHOULDERS OF THE DELIVERY SYSTEM BALLOON WERE IN CONTACT WITH THE VENTRICULAR PORTION OF THE VALVE." ADDITIONALLY, IT WAS REPORTED THAT "THE VALVE WAS SLIGHTLY MORE AORTIC THAN THE HIGHEST ALIGNMENT MARKER." PER THE IFU AND TRAINING MANUAL, THE INSTRUCTIONS STATE TO CHECK TO ENSURE THE TRANSCATHETER HEART VALVE (THV) IS EXACTLY BETWEEN THE VALVE ALIGNMENT MARKERS PRIOR TO DEPLOYMENT. IN THIS CASE, THE AVAILABLE INFORMATION SUGGESTS THAT NOT HAVING THE THV CENTERED BETWEEN THE ALIGNMENT MARKERS PRIOR TO VALVE DEPLOYMENT MAY HAVE CONTRIBUTED TO THE EVENT. THROUGH EXTENSIVE COMPLAINT INVESTIGATIONS, BALLOON BURST EVENTS DURING VALVE DEPLOYMENT HAVE NOT BEEN ASSOCIATED WITH DEVICE MALFUNCTIONS OR MANUFACTURING NONCONFORMANCES. RATHER, THE ROOT CAUSE FOR THESE EVENTS HAVE HISTORICALLY BEEN DUE TO CALCIFICATION IN THE LANDING ZONE OR OVER-INFLATION OF THE BALLOON. IN THIS CASE, THE COMMANDER DELIVERY SYSTEM BALLOON BURST THAT CONTRIBUTED TO THE SEMI-DEPLOYED VALVE HAVING TO BE DEPLOYED IN THE ABDOMINAL AORTA WAS ALSO UNABLE TO BE CONFIRMED. THE AVAILABLE INFORMATION SUGGESTS THAT PATIENT FACTORS (CALCIFICATION) LIKELY CONTRIBUTED TO THE EVENT AS CALCIFICATION WAS OBSERVED IN THE LANDING ZONE PER REVIEW OF THE PROVIDED IMAGERY. THE PRESENCE OF CALCIFICATION CAN CREATE A CHALLENGING ANATOMY FOR BALLOON INFLATION. WHILE THE BALLOONS ARE SUFFICIENTLY DESIGNED AND TESTED TO ENSURE THE BURST PRESSURE IS AT OR ABOVE THE RATED BURST PRESSURE, CALCIFIED NODULES CAN COMPROMISE THE STRUCTURE OF THE BALLOON WALL VIA THE FOLLOWING MECHANISMS SUCH AS PUNCTURE, LOCAL OVERSTRETCHING, OPEN CELL IMPINGEMENT, OR STRESS CONCENTRATION. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. SINCE NO EDWARDS PRODUCT DEFECTS OR LABELING DEFICIENCIES WERE IDENTIFIED, NO CORRECTIVE OR PREVENTATIVE ACTION IS REQUIRED. Manufacturer Narrative: THE INVESTIGATION IS ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21625891&pc=FKX | 0008030665-2025-00613 | 2025-02-28 05:00 | Death | ERIKA DE REYNOSA, S.A. DE C.V. | 2025-03-17 04:00 | FKX | LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | K181108 |  | 1 | Event Description: A NURSE REPORTED THAT A PERITONEAL DIALYSIS (PD) PATIENT WENT INTO CARDIAC ARREST WHILE CONNECTED TO THE LIBERTY SELECT CYCLER DURING PD TREATMENT. ADDITIONAL INFORMATION WAS OBTAINED THROUGH FOLLOW-UP WITH THE PATIENT¿S PERITONEAL DIALYSIS REGISTERED NURSE (PDRN). THE PATIENT WAS IN TREATMENT IN THE EARLY MORNING HOURS ON (B)(6) 2025 WHEN A ROOMMATE HEARD HIM COLLAPSE. THE ROOMMATE FOUND THE PATIENT UNRESPONSIVE. 911 WAS CALLED AND THE ROOMMATE INITIATED CARDIOPULMONARY RESUSCITATION (CPR). EMERGENCY MEDICAL SERVICES (EMS) TRANSPORTED THE PATIENT TO THE HOSPITAL WHERE THE PATIENT WAS INTUBATED. THE DECISION WAS MADE TO PLACE THE PATIENT ON COMFORT MEASURES ONLY. THE PATIENT PASSED AWAY IN THE EVENING ON (B)(6) 2025. THE CAUSE OF DEATH IS CARDIAC ARREST. THIS PATIENT HAS SIGNIFICANT CARDIAC COMORBIDITIES INCLUDING A CARDIAC PACEMAKER, AND RECENT OPEN-HEART SURGERY (DATES UNKNOWN). A REVIEW OF THE PATIENT¿S TREATMENT RECORDS DID NOT SHOW ANY ISSUES DURING THE TREATMENT.Event Description: A NURSE REPORTED THAT A PERITONEAL DIALYSIS (PD) PATIENT WENT INTO CARDIAC ARREST WHILE CONNECTED TO THE LIBERTY SELECT CYCLER DURING PD TREATMENT. ADDITIONAL INFORMATION WAS OBTAINED THROUGH FOLLOW-UP WITH THE PATIENT¿S PERITONEAL DIALYSIS REGISTERED NURSE (PDRN). THE PATIENT WAS IN TREATMENT IN THE EARLY MORNING HOURS ON (B)(6) 2025 WHEN A ROOMMATE HEARD HIM COLLAPSE. THE ROOMMATE FOUND THE PATIENT UNRESPONSIVE. 911 WAS CALLED AND THE ROOMMATE INITIATED CARDIOPULMONARY RESUSCITATION (CPR). EMERGENCY MEDICAL SERVICES (EMS) TRANSPORTED THE PATIENT TO THE HOSPITAL WHERE THE PATIENT WAS INTUBATED. THE DECISION WAS MADE TO PLACE THE PATIENT ON COMFORT MEASURES ONLY. THE PATIENT PASSED AWAY IN THE EVENING ON (B)(6) 2025. THE CAUSE OF DEATH IS CARDIAC ARREST. THIS PATIENT HAS SIGNIFICANT CARDIAC COMORBIDITIES INCLUDING A CARDIAC PACEMAKER, AND RECENT OPEN-HEART SURGERY (DATES UNKNOWN). A REVIEW OF THE PATIENT¿S TREATMENT RECORDS DID NOT SHOW ANY ISSUES DURING THE TREATMENT. Manufacturer Narrative: ADDITIONAL INFORMATION: D.9.,H.3. PLANT INVESTIGATION: THE ACTUAL DEVICE WAS RETURNED. A VISUAL INSPECTION OF THE RETURNED CYCLER EXTERIOR SHOWED NO SIGNS OF PHYSICAL DAMAGE. THERE WERE NO VISUAL INDICATIONS OF DRIED FLUID WITHIN THE CASSETTE COMPARTMENT ON THE PUMP. THERE WERE NO VISUAL INDICATIONS OF PARTICULATES WITHIN THE CASSETTE AREA. THERE WERE NO BURRS OR SHARP EDGES IN CASSETTE AREA THAT MAY HAVE PUNCTURED A CASSETTE MEMBRANE. A (AS-RECEIVED WITH REDUCED DWELL) SIMULATED TREATMENT WAS PERFORMED AND COMPLETED. VALVE ACTUATION TEST ¿ PASSED. SYSTEM AIR LEAK TEST ¿ PASSED. THE DEVICE HISTORY RECORD DID NOT REVEAL ANY ISSUES OR PROBLEMS RELATED TO THE REPORTED SYMPTOM CODE(S)..THE DEVICE HISTORY RECORD DID NOT REVEAL ANY ISSUES OR PROBLEMS RELATED TO THE REPORTED SYMPTOM CODE(S).UPON COMPLETION OF THE EVALUATION, THERE WERE NO MALFUNCTIONS THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. AN ASSOCIATED CAUSE COULD NOT BE DETERMINED. Manufacturer Narrative: CLINICAL REVIEW: THERE IS A TEMPORAL RELATIONSHIP BETWEEN PERITONEAL DIALYSIS UTILIZING THE LIBERTY SELECT CYCLER AND THE PATIENT EVENT OF CARDIAC ARREST WITH HOSPITALIZATION AND DEATH. HOWEVER, THERE IS NO DOCUMENTATION IN THE COMPLAINT FILE TO SHOW A CAUSAL RELATIONSHIP BETWEEN THE ADVERSE EVENT AND THE USE OF THE LIBERTY SELECT CYCLER. ADDITIONALLY, THERE IS NO ALLEGATION OF A MALFUNCTION OR DEFICIENCY REPORTED FOR THIS EVENT. THE CAUSE OF THE PATIENT¿S DEATH WAS CARDIAC ARREST. THE PATIENT HAS SIGNIFICANT CARDIAC HISTORY INCLUDING A CARDIAC PACEMAKER AND RECENT OPEN-HEART SURGERY. PATIENTS WITH CHRONIC KIDNEY DISEASE (CKD) EXHIBIT AN ELEVATED CARDIOVASCULAR RISK MANIFESTING AS CORONARY ARTERY DISEASE, HEART FAILURE, ARRHYTHMIAS, AND SUDDEN CARDIAC DEATH. CARDIOVASCULAR RATHER THAN END-STAGE KIDNEY DISEASE (CKD STAGE 5) IS THE LEADING CAUSE OF DEATH IN THIS HIGH-RISK POPULATION. BASED ON THE AVAILABLE INFORMATION AND NO ALLEGATION OR EVIDENCE OF A MALFUNCTION OR DEFICIENCY, THE LIBERTY SELECT CYCLER CAN BE EXCLUDED AS THE CAUSE OF THE PATIENT¿S CARDIAC ARREST AND DEATH. THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21627253&pc=OZD | 1220648-2025-26779 | 2025-02-21 05:00 | Death | ABIOMED, INC. | 2025-03-17 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED 76 YEAR OLD FEMALE PATIENT WAS ON IMPELLA CP SUPPORT AND DURING SUPPORT, THE PATIENT HAD LIMB ISCHEMIA. THE PATIENT HAD NO PULSES PRESENT IN RIGHT LEG DISTAL TO IMPELLA ACCESS. THE IMPELLA WAS REMOVED. ADDITIONALLY, FURTHER DETAILS STATE THAT CARE WAS WITHDRAWN AND THE PATIENT SUBSEQUENTLY EXPIRED. THERE WERE NO DETAILS PROVIDED AS TO THE CAUSE OF THE DEATH. Manufacturer Narrative: THE INVESTIGATION OF LIMB ISCHEMIA HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. AS ALL THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE LIMB ISCHEMIA WAS NOT DETERMINED. B.7 REVISED AND ADDED INFORMATION SINCE THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26779 WAS SUBMITTED. E.1 ADDED US PHONE NUMBER AS IT WAS INADVERTENTLY OMITTED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26779 WAS SUBMITTED. E.4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26779 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. G.1 REVISED REPORTING CONTACT FACILITY NAME AND FAX NUMBER AS THEY WERE INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26779. H.6 COMPONENT CODE 925 WAS REPORTED INCORRECTLY ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26779. A NEW COMPONENT CODE WAS ADDED. CODE 4641 WAS REPORTED INCORRECTLY ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26779. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMART ASSIST SYSTEM STATES THE FOLLOWING: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21627276&pc=OZD | 1220648-2025-26781 | 2025-02-19 05:00 | Death | ABIOMED, INC. | 2025-03-17 04:00 | OZD | IMPELLA CP | Pumping Stopped | Cardiac Arrest | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WITH CARDIOMYOPATHY WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS). THE PATIENT WAS A TRANSFER FROM AN OUTSIDE HOSPITAL ON AN INTRA-AORTA BALLOON PUMP FOR A LEFT VENTRICULAR ASSIST DEVICE WORKUP. APPROXIMATELY NINE DAYS AFTER START OF IMPELLA MCS, THE PUMP FLOW SATURATION WAS NOTED AND THE PUMP STOPPED. PUMP STOPPED. THE AUTOMATED IMPELLA CONTROLLER (AIC) WAS EXCHANGED; HOWEVER, THE PUMP DID NOT RESTART. FOLLOWING, THE DECISION WAS MADE TO PULL IMPELLA ACROSS THE AORTIC VALVE AND START COMFORT CARE. THE PATIENT WOULD SUBSEQUENTLY EXPIRE. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER AT THE TIME OF THIS MDR WRITING; THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA CATHETER ¿ACHIEVING AN ACT = 250 SECONDS PRIOR TO REMOVING THE DILATOR WILL HELP PREVENT A THROMBUS FROM ENTERING THE CATHETER AND CAUSING A SUDDEN STOP ON STARTUP.¿ SECTION: TROUBLESHOOTING THE PURGE SYSTEM ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP.¿ SECTION: IMPELLA STOPPED ¿IF THE IMPELLA CATHETER HAS STOPPED SUDDENLY: 1. TRY TO RESTART THE CATHETER AT PREVIOUS P-LEVEL. 2. IF THE IMPELLA DOES NOT RESTART, TRY TO RESTART AT P-2. 3. IF THE IMPELLA DOES NOT RESTART OR STOPS AGAIN, WAIT 1 MINUTE AND TRY TO RESTART AGAIN. 4. IF THE IMPELLA RESTARTS, WEAN DOWN TO P-2 AS THE PATIENT CAN TOLERATE. UNDER THESE CIRCUMSTANCES, CATHETER FUNCTION IS NOT RELIABLE AND THE IMPELLA MAY STOP AGAIN. 5. IF THE IMPELLA DOES NOT RESTART, REMOVE THE IMPELLA FROM THE VENTRICLE AS SOON AS POSSIBLE TO AVOID AORTIC INSUFFICIENCY.¿ SECTION: WARNINGS, CAUTIONS, AND PRECAUTIONS ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21627546&pc=LWP | 2017865-2025-27689 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-03-17 04:00 | LWP | ASSURITY MRI | Adverse Event Without Identified Device or Use Problem | Arrhythmia; No Clinical Signs, Symptoms or Conditions | P140033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD DECEASED DUE TO AN UNKNOWN CAUSE. NO ADDITIONAL INFORMATION WAS REPORTED.Event Description: IT WAS REPORTED THAT THE PATIENT HAS DECEASED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THE DEATH WAS DEVICE RELATED. THE CAUSE OF DEATH WAS CARDIAC ARREST. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: CORRECTION: TO H6 CODING SHOULD HAVE BEEN ARRHYTHMIA CODING INSTEAD OF JUST DEATH. Manufacturer Narrative: CORRECTION: B5 FOR CARDIAC ARREST AND PHYSICIAN BELIEVING THE DEATH WAS RELATED TO THE PACEMAKER MISSING IN THE STATEMENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21627830&pc=CBK | 2518422-2025-103122 | 2025-02-14 05:00 | Death | RESPIRONICS, INC. | 2025-03-17 04:00 | CBK | TRILOGY EVO | Date/Time-Related Software Problem | Insufficient Information | K181166 |  | 1 | Event Description: THE CUSTOMER REPORTED THAT, WHILE INVESTIGATING THE DEATH OF A PATIENT, THE PATIENT MAY HAVE BEEN CONNECTED TO THE DEVICE AT THE TIME OF DEATH. THE CUSTOMER REPORTED THAT THE TIMES RECORDED IN THE DEVICE'S EVENT LOG SHOW A DISCREPANCY COMPARED TO WHAT IS DISPLAYED IN CARE ORCHESTRATOR ESSENCE. WHILE THE DEVICE CANNOT BE RETURNED AT THIS TIME DUE TO BEING HELD IN POLICE CUSTODY, THE MANUFACTURER HAS ACCESSED PARTIAL THERAPY DATA IN CARE ORCHESTRATOR ESSENCE. THE EVENT IN QUESTION OCCURRED ON (B)(6) 2025. THE CUSTOMER INITIALLY STATES THAT THE VENTILATOR DISPLAYED THE FOLLOWING DATA: BEFORE 14:00 THERE WERE SEVERAL HIGH-PRESSURE ALARMS, AT 14:18 (SCREEN ON THE TRILOGY) A LOW-PRESSURE ALARM, AT 14:35 A DISCONNECTION ALARM, AT 15:09 A HIGH INSPIRATORY PRESSURE ALARM, AND AT 16:19 A DISCONNECTION ALARM (LAST ALARM). THE CUSTOMER THEN PROVIDED ADDITIONAL CONTEXT AFTER SPEAKING WITH THE POLICE AND BEING GRANTED ACCESS TO THE DEVICE TO DOWNLOAD SUPPLEMENTAL DATA. THE SCREEN LOG SHOWS A VT PRESCRIPTION CHANGE AT 13:05 (WHICH IS ACCURATE PER THE CUSTOMER) WHILE CARE ORCHESTRATOR HAS IT REGISTERED AT 11:43. IT IS NOTED THAT THE MAIN CONCERN IS THE DIFFERENCE BETWEEN THOSE TWO TIMES. PRODUCT SUPPORT REVIEWED THE INITIAL THERAPY DATA IN CARE ORCHESTRATOR ESSENCE AND CONFIRMED THE TIME DIFFERENCE OF WHEN THE SETTING CHANGES AND ALARM TIMES WERE DIFFERENT FROM THE TIME OF THE DEVICE VERSUS THE TIME SHOWN OF THE EVENTS IN CARE ORCHESTRATOR ESSENCE. THE CUSTOMER ALSO STATED THAT THE VENTILATOR IS A HOSPITAL UTILIZED DEVICE AND MAY DISPLAY DATA FROM PREVIOUS PATIENTS. AT THIS TIME WITH THE PRELIMINARY INFORMATION GIVEN, PRODUCT SUPPORT IS UNABLE TO CONFIRM WHY THERE IS A DIFFERENCE IN TIME. SINCE A DEATH IS CONNECTED TO THIS CASE, THE DEVICE WILL NEED TO BE RETURNED TO THE MANUFACTURER FOR FURTHER INVESTIGATION, AND TO DOWNLOAD THE INTERNAL LOGS, IN ORDER TO PROVIDE A FINAL REPORT, PER THE FDA REQUIREMENTS. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER PREVIOUSLY REPORTED THAT WHILE INVESTIGATING THE DEATH OF A PATIENT, THE PATIENT MAY HAVE BEEN CONNECTED TO THE DEVICE AT THE TIME OF DEATH. THE CUSTOMER REPORTED THAT THE TIMES RECORDED IN THE DEVICE'S EVENT LOG SHOW A DISCREPANCY COMPARED TO WHAT IS DISPLAYED IN CARE ORCHESTRATOR ESSENCE. WHILE THE DEVICE CANNOT BE RETURNED AT THIS TIME DUE TO BEING HELD IN POLICE CUSTODY, THE MANUFACTURER HAS ACCESSED PARTIAL THERAPY DATA IN CARE ORCHESTRATOR ESSENCE. THE EVENT IN QUESTION OCCURRED ON 02/14/2025. THE CUSTOMER INITIALLY STATES THAT THE VENTILATOR DISPLAYED THE FOLLOWING DATA: BEFORE 14:00 THERE WERE SEVERAL HIGH-PRESSURE ALARMS, AT 14:18 (SCREEN ON THE TRILOGY) A LOW-PRESSURE ALARM, AT 14:35 A DISCONNECTION ALARM, AT 15:09 A HIGH INSPIRATORY PRESSURE ALARM, AND AT 16:19 A DISCONNECTION ALARM (LAST ALARM). THE CUSTOMER THEN PROVIDED ADDITIONAL CONTEXT AFTER SPEAKING WITH THE POLICE AND BEING GRANTED ACCESS TO THE DEVICE TO DOWNLOAD SUPPLEMENTAL DATA. THE SCREEN LOG SHOWS A VT PRESCRIPTION CHANGE AT 13:05 (WHICH IS ACCURATE PER THE CUSTOMER) WHILE CARE ORCHESTRATOR HAS IT REGISTERED AT 11:43. IT IS NOTED THAT THE MAIN CONCERN IS THE DIFFERENCE BETWEEN THOSE TWO TIMES. PRODUCT SUPPORT REVIEWED THE INITIAL THERAPY DATA IN CARE ORCHESTRATOR ESSENCE AND CONFIRMED THE TIME DIFFERENCE OF WHEN THE SETTING CHANGES AND ALARM TIMES WERE DIFFERENT FROM THE TIME OF THE DEVICE VERSUS THE TIME SHOWN OF THE EVENTS IN CARE ORCHESTRATOR ESSENCE. THE CUSTOMER ALSO STATED THAT THE VENTILATOR IS A HOSPITAL UTILIZED DEVICE AND MAY DISPLAY DATA FROM PREVIOUS PATIENTS. AT THIS TIME WITH THE PRELIMINARY INFORMATION GIVEN, PRODUCT SUPPORT IS UNABLE TO CONFIRM WHY THERE IS A DIFFERENCE IN TIME. SINCE A DEATH IS CONNECTED TO THIS CASE, THE DEVICE WILL NEED TO BE RETURNED TO THE MANUFACTURER FOR FURTHER INVESTIGATION, AND TO DOWNLOAD THE INTERNAL LOGS, IN ORDER TO PROVIDE A FINAL REPORT, PER THE FDA REQUIREMENTS. THE DEVICE'S EVENT LOGS HAD A PRELIMINARY EVALUATION DONE. THE FULL DEVICE EVALUATION HAS NOT BEEN COMPLETED. BASED ON THE EVALUATION OF THE SUBJECT DEVICE¿S DOWNLOADED EVENT LOG, NO EVIDENCE WAS FOUND OF A MALFUNCTION OR FAILURE TO PERFORM WITHIN SPECIFICATIONS; THE SUBJECT DEVICE OPERATED AND ALARMED AS DESIGNED AND DID NOT CONTAIN ANY EVENTS THAT WOULD CONSTITUTE A DEVICE CONTRIBUTION TO THE REPORTED INCIDENT. HOWEVER, A FULL INVESTIGATION MUST BE PERFORMED UPON RETURNING THE SUBJECT DEVICE TO A MANUFACTURER FACILITY IN ORDER TO COMPLETELY RULE OUT ANY MALFUNCTIONS AS A CAUSE OF THE ALARMS. ALTHOUGH A PRELIMINARY EVENT LOG EVALUATION WAS DONE, THE FULL DEVICE EVALUATION HAS YET TO BE COMPLETED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE SERVICE CENTER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21628009&pc=LPB | 2029046-2025-00819 | 2025-02-20 05:00 | Death | BIOSENSE WEBSTER INC | 2025-03-17 04:00 | LPB | THERMOCOOL SMARTTOUCH SF | Adverse Event Without Identified Device or Use Problem; Patient Device Interaction Problem | Cardiac Tamponade | P030031 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A PAROXYSMAL PULMONARY VEIN ISOLATION ABLATION PROCEDURE WITH A SMART TOUCH BIDIRECTIONAL SF AND THE PATIENT EXPERIENCED PERICARDIAL EFFUSION THAT REQUIRED PERICARDIOCENTESIS, SURGICAL INTERVENTION AND PATIENT DIED. DURING THE ABLATION WITH THE SMART TOUCH BIDIRECTIONAL SF, A STEAM POP OCCURRED ALONG WITH A PERICARDIAL EFFUSION; THE FLUID WAS DRAINED BUT CONTINUED TO ACCUMULATE. CARDIAC SURGERY WAS CONSULTED AND ATTEMPTED TO SAVE THE PATIENT, BUT UNFORTUNATELY, THEY WERE UNSUCCESSFUL DUE TO THE SIZE OF THE DEFECT. THE PATIENT HAD A PRE-EXISTING POOR PUMP FUNCTION. THEY HAD ABLATED AT 50 (FIFTY) WATTS, DURING A 10 (TEN) SECONDS ABLATION, AND WITH AN ABLATION INDEX AT 400 (FOUR HUNDRED). AN ANTERIOR LINE WAS CREATED AND ABLATED WITH 50 W. THE INCIDENT HAS BEEN REPORTED AS FATAL FOR THE PATIENT. Manufacturer Narrative: ON 27-MAR-2025, ADDITIONAL INFORMATION WAS RECEIVED INDICATING THE PRODUCT CODE OF THE CATHETER USED IS D134805. AS SUCH, THE FOLLOWING FIELDS HAVE BEEN UPDATED: FIELD D1. BRAND NAME HAS BEEN UPDATED FROM "SMART TOUCH BIDIRECTIONAL SF" TO "THERMOCOOL SMARTTOUCH SF". FIELD D4. CATALOG HAS BEEN UPDATED FROM "UNK_SMART TOUCH BIDIRECTIONAL SF" TO "D134805". FIELD D4: UDI: AS THE LOT NUMBER FOR THE DEVICE INVOLVED IN THE EVENT WAS NOT PROVIDED, THE FULL UDI IS CURRENTLY NOT AVAILABLE. IF ADDITIONAL INFORMATION IS RECEIVED REGARDING THIS EVENT, A SUPPLEMENTAL 3500A REPORT WILL BE SUBMITTED TO THE FDA. MANUFACTURER'S REF. # (B)(4). Manufacturer Narrative: DEVICE INVESTIGATION DETAILS: AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. AN EVALUATION OF THE MANUFACTURING RECORD COULD NOT BE PERFORMED AS THE REQUIRED PRODUCT IDENTIFICATION NUMBER WAS NOT PROVIDED TO COMPLETE THE EVALUATION. HOWEVER, IF THE PRODUCT OR PRODUCT ID NUMBERS ARE RECEIVED AT A LATER DATE, THE INVESTIGATION WILL BE UPDATED AS APPLICABLE. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. NOTE: FIELD D4. CATALOG SHOULD BE ¿UNK_SMART TOUCH BIDIRECTIONAL SF ¿. NOTE: FIELD D4: UDI: AS THE CATALOG/MODEL NUMBER WAS NOT PROVIDED, THE (01)GTIN IS NOT AVAILABLE. NOTE: FIELD E1. INITIAL REPORTER PHONE: (B)(6). IF ADDITIONAL INFORMATION IS RECEIVED REGARDING THIS EVENT, A SUPPLEMENTAL 3500A REPORT WILL BE SUBMITTED TO THE FDA. MANUFACTURER'S REF. # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21729246&pc=PNJ | 9612164-2025-01598 | 2025-03-10 04:00 | Death | MEDTRONIC IRELAND | 2025-03-31 04:00 | PNJ | MICRA¿ AV2 | Failure to Capture; Adverse Event Without Identified Device or Use Problem; Positioning Problem | Cardiac Tamponade; Cardiac Perforation | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE IMPLANT PROCEDURE THE PATIENTS HEART WAS SMALL AND THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) DELIVERY CATHETER WAS HARD TO MOVE TO THE RIGHT VENTRICULAR (RV) LEAD SEPTUM. WHEN PLACED THE LEADLESS IPG HAD NO CAPTURE AT HIGH OUTPUT. THE DEVICE WAS RECAPTURED AND REPOSITIONED TO A NEW LOCATION. CONTRAST WAS GIVEN FOR SEPTAL CONFIRMATION AND THE DEVICE WAS DEPLOYED. THE PATIENT BECAME SYMPTOMATIC EXPERIENCING A CARDIAC PERFORATION LEADING TO CARDIAC TAMPONADE. A CODE WAS CALLED AND AN ECHOCARDIOGRAM, PERICARDIOCENTESIS AND RESUSCITATIVE MEASURES WERE PERFORMED. RESCUE WAS INEFFECTIVE AND THE PATIENT PASSED AWAY. Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [MICRA] PRODUCT ID [MC2AVR1-DELSYS] (SERIAL: (B)(6). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTED: H6, ADDITIONAL CODES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTED: G3 (FOLLOW-UP #001 AWARE DATE = 15-MAY-2025), H6 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTED: H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21729939&pc=MUE | MW5168382 | 2025-03-13 04:00 | Death | HOLOGIC, INC. | 2025-03-31 04:00 | MUE | MAMMOGRAPHY-DIGITAL | Use of Device Problem | Jaundice; Cancer; Liver Failure |  |  | 1 | Event Description: PATIENT'S LAST SCREENING MAMMOGRAM WAS (B)(6) 2025. PATIENT WAS ADMITTED TO HOSPITAL ON (B)(6) 2025 DUE TO OBSTRUCTIVE JAUNDICE. PATIENT UNDERWENT EUS WITH EGD ON (B)(6) 2025 WITH BIOPSY AND WHIPPLE PROCEDURE (B)(6) 2025. PATHOLOGY FINDINGS OF PERIAMPULLARY ADENOCARCINOMA. POST-SURGERY, PATIENT DEVELOPED ACUTE LIVER FAILURE AND INCREASING HEMODYNAMIC INSTABILITY. DIED ON (B)(6) 2025 WHILE IN HOSPITAL. THIS WAS WITHIN 30 DAYS OF ANY STUDY RELATED ACTIVITY, AND THUS EXPEDITED REPORTING IS REQUIRED PER PROTOCOL. Event Description: ADDITIONAL INFORMATION RECEIVED FOR REPORT MW5168382 ON 4/04/2025.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21730009&pc=KNT | MW5168383 | 2024-12-10 05:00 | Death | AVANOS MEDICAL, INC. (FORMERLY HALYARD HEALTH INC.) | 2025-03-31 04:00 | KNT | CORTRACK |  | Cardiac Arrest; Fistula; Pleural Effusion |  |  | 1 | Event Description: PT ADMITTED FOR A FALL AT HOME RESULTING IN A FRACTURE OF THE DISTAL CLAVICLE W MINIMAL DISPLACEMENT. NGT (NASOGASTRIC TUBE)/ CORTRACK WAS ORDERED THAT RESULTED IN 3-MINUTE CARDIAC ARREST. INJURY TO THE RIGHT MAINSTEM AND THROUGH THE LUNG. ABDOMINAL X-RAY SHOWED FEEDING TUBE WAS LOOPED OVER RIGHT HEMITHORAX CONCERNING FOR TRAUMATIC BRONCHOPLEURAL FISTULA AND POSSIBLE EXTENDING INTO THE PLEURAL SPACE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21730041&pc=MKJ | MW5168384 | 2025-03-24 04:00 | Death | ZOLL MEDICAL CORPORATION | 2025-03-31 04:00 | MKJ | ZOLL DEFIBRILLATOR |  | Cardiac Arrest |  |  | 1 | Event Description: ON (B)(6) 2025 05:58AM, THE PATIENT WAS FOUND PULSELESS, AND CODE BLUE WAS ACTIVATED. DURING CODE BLUE, R ZOLL DEFIBRILLATOR/MONITOR WAS UTILIZED. RESUSCITATION EFFORTS WERE UNSUCCESSFUL, AND THE PATIENT WAS PRONOUNCED DEAD AT 06:42AM. AFTER THE CODE, THE STRIPS FROM R ZOLL DEFIBRILLATOR WERE SHOWING YEAR 2005 INSTEAD OF YEAR 2025. FACILITY BIO-MED CONTACTED THE MANUFACTURER AND WAS TOLD THAT THE COIN BATTERY MAY NEED REPLACEMENT TO SHOW THE CORRECT YEAR. WE CURRENTLY DO NOT HAVE THE RESULTS OF MACHINE INTERROGATION BY THE MANUFACTURER, AND BASED ON THE CURRENT INFORMATION WE HAVE, WE DO NOT HAVE EVIDENCE THAT THE MACHINE CONTRIBUTED TO THE UNFORTUNATE OUTCOME. HOWEVER, THIS REPORT IS BEING MADE OUT OF ABUNDANCE OF CAUTION AND TO COMPLY WITH REPORTABLE MANDATES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21730067&pc=QDA | MW5168385 | 2025-02-17 05:00 | Death | UNK | 2025-03-31 04:00 | QDA | HBAND |  | Insufficient Information |  |  | 1 | Event Description: BECAUSE OF A BLOOD GLUCOSE WATCH MY ELDERLY UNCLE STOPPED USING HIS GLUCOSE METER AND TAKING HIS MEDS AND IN TURN DIED FROM USING AND RELYING ON THESE DEVICES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21730144&pc=OZD | 1220648-2025-27085 | 2025-03-03 05:00 | Death | ABIOMED, INC. | 2025-03-31 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Atrial Fibrillation; Atrial Flutter; Hematoma; Hemorrhage/Bleeding; Thrombocytopenia; Gastrointestinal Hemorrhage | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT THERE WAS SOME OOZING AND A MEDIUM SIZED HEMATOMA AT THE IMPELLA CP ACCESS SITE. THE MEDICAL DOCTOR ADJUSTED REPOSITIONING SHEATH AND MANIPULATED THE PERCLOSE AND OOZING/HEMATOMA RESOLVED. ADDITIONALLY, THE PATIENT DEVELOPED THROMBOCYTOPENIA AND BLOOD PRODUCTS WERE ADMINISTERED. THE PATIENT ALSO DEVELOPED A GASTROINTESTINAL BLEED. THE FAMILY DECIDED TO WITHDREW CARE AND THE PATIENT EXPIRED.Event Description: IT WAS FURTHER REPORTED THE PATIENT EXPERIENCED RUNS OF ATRIAL FIBRILLATION AND ATRIAL FLUTTER. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE,¿AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM: SECTION: GENERAL PATIENT CARE CONSIDERATIONS: ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT: USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING, THROMBOCYTOPENIA, VF/VT, AND VASCULAR DAMAGE - PERIPHERAL VESSEL HAS BEEN COMPLETED. ACCESS SITE BLEEDING: PRODUCT WAS NOT RETURNED FOR INVESTIGATION. AS PER THE CLINICAL INFORMATION PROVIDED, THE REPOSITION SHEATH AND THE PERCLOSE HAD TO BE ADJUSTED TO STOP THE OOZING. THEREFORE, THE ROOT CAUSE OF THE ACCESS SITE BLEEDING ISSUE WAS MOST LIKELY USE RELATED TO IMPROPER TECHNIQUE. THROMBOCYTOPENIA: THE ROOT CAUSE OF THE THROMBOCYTOPENIA ISSUE WAS NOT DETERMINED DUE TO INSUFFICIENT CLINICAL DETAILS. VASCULAR DAMAGE: PUMP WAS NOT RETURNED. CLINICAL INFORMATION PROVIDED WAS VERY LIMITED AND NO INFORMATION PROVIDED ABOUT THE REASON OF THE GI BLEED. THEREFORE, THE ROOT CAUSE OF THE VASCULAR DAMAGE ISSUE WAS NOT DETERMINED DUE TO INSUFFICIENT CLINICAL INFORMATION AND UNKNOWN RELATION TO IMPELLA WITH GI BLEED. VF/VT: AS THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE VT/VF WAS NOT DETERMINED. B5 REVISED TO INCLUDE ADDITIONAL INFORMATION INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER: 1220648-2025-27085. H6 HEALTH EFFECT - CLINICAL CODE 1729 AND 1730 WERE INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER: 1220648-2025-27085.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21730146&pc=OZD | 1220648-2025-27086 | 2024-12-19 05:00 | Death | ABIOMED, INC. | 2025-03-31 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia | P140003 |  | 1 | Event Description: A NOTIFICATION WAS RECEIVED VIA ABIOMED¿S LONG-TERM OUTCOME AND QUALITY INDICATOR (LOQI) IMPELLA REGISTRY REGARDING A PATIENT THAT ENDURED RIGHT LOWER EXTREMITY (RLE) ISCHEMIA WITH "DEFINITELY RELATED" TO THE IMPELLA. LOQI STATES: PATIENT SHOWED MOTTLING IN THE RLE, AND NO SIGNAL FOUND IN RLE BELOW GROIN. REQUIRED EMERGENT DISTAL PERFUSION CANNULA PLACEMENT AT BEDSIDE. ADDITIONALLY, CARE WAS WITHDRAWN AND PATIENT EXPIRED. THERE IS NOT ENOUGH INFORMATION TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMART ASSIST SYSTEM STATES THE FOLLOWING: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21730223&pc=OZD | 1220648-2025-27087 | 2025-03-04 05:00 | Death | ABIOMED, INC. | 2025-03-31 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Thrombocytopenia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS INITIALLY IMPLANTED WITH AN IMPELLA CP DEVICE WITH VENO-ARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION (VA ECMO) SUPPORT FOR APPROXIMATELY 10 DAYS UNTIL ESCALATION IN THERAPY TO AN IMPELLA 5.5 DEVICE FOR CONTINUED MECHANICAL CIRCULATORY SUPPORT (MCS). DURING IMPELLA 5.5 SUPPORT, THE PATIENT EXPERIENCED BLEEDING AT THE GRAFT AND THROMBOCYTOPENIA REQUIRING BLOOD PRODUCTS TRANSFUSION. THE IMPELLA CP DEVICE WAS SUCCESSFULLY WEANED AND EXPLANTED BY UTILIZING THE DOUBLE BARREL METHOD INSERTING THE IMPELLA 5.5. THE IMPELLA 5.5 PUMP WAS STARTED AT PERFORMANCE LEVEL P-2. THERE WAS AN ATTEMPT TO WEAN OFF OF ECMO, BUT THE PATIENT'S HEMODYNAMICS WERE NOT TOLERATING THIS WEAN. SIMULTANEOUSLY, THE WALLS OF THE GRAFT WERE SWEATING NOTICEABLE AMOUNTS OF BLOOD. CONSEQUENTLY, PERFORMANCE LEVEL WAS LEFT AT P-2 AND MAINTAINED SUPPORT ON ECMO, WITH PLAN TO RETURN AT A LATER DATE FOR REMOVAL OF ECMO SUPPORT. FOR THE GRAFT, THIS WAS KNOWN TO BE COAGULAOPATHIC. PRESSURE WAS HELD, HEMOSTATIC AGENT WAS APPLIED, AND BLOOD PRODUCT WERE GIVEN WITH EVENTUAL HEMOSTASIS ON THE GRAFT. THE SITE WAS THEN CLOSED AND DRESSED APPROPRIATELY. THREE-POINT FIXATION WAS APPLIED, AND THE PATIENT IN STABLE MANNER WAS TRANSFERRED TO UNIT ON MCS. THE FOLLOWING DAY, ONE UNIT OF PLATELETS WAS GIVEN FOR THROMBOCYTOPENIA (UNCLEAR RELATIONSHIP TO THE IMPELLA 5.5). THE PATIENT CONTINUED ON SUPPORT AND WAS LAST NOTED TO BE IN STABLE CONDITION.Event Description: IT WAS FURTHER REPORTED THE FAMILY DECIDED TO WITHDRAW CARE AND THE PATIENT EXPIRED. IT IS UNKNOWN IF THE USE OF THE IMPELLA WAS RELATED TO PATIENT DEATH, THUS THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED AND THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 FOR USE DURING CARDIOGENIC SHOCK SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING AND THROMBOCYTOPENIA HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. THE ROOT CAUSE OF ACCESS SITE BLEEDING WAS MOST LIKELY USE ISSUE SINCE NON-ABIOMED PRODUCT WAS USED WHICH IS NOT RECOMMENDED PER INSTRUCTIONS FOR USE. THE ROOT CAUSE OF THROMBOCYTOPENIA CANNOT BE DETERMINED SINCE THE PRODUCT WAS NOT RETURNED AND INSUFFICIENT CLINICAL DETAILS WERE PROVIDED. Manufacturer Narrative: B2 UPDATED TO INCLUDE DEATH AND DATE OF DEATH B5 UPDATED WITH ADDITIONAL INFORMATION RECEIVED FROM THE CLINICAL SITE D6B EXPLANT DATE ADDED H1 TYPE OF REPORT REVISED BASED ON THE ADDITIONAL INFORMATION RECEIVED FROM THE CLINICAL SITE H6 REVISED TO REFLECT DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21730397&pc=NPT | 2025587-2025-02220 | 2025-03-27 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-31 04:00 | NPT | EVOLUT PRO PLUS VALVE | Device Dislodged or Dislocated; Insufficient Information; Patient Device Interaction Problem | Cardiac Arrest; Foreign Body In Patient | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER HEART VALVE PROCEDURE INVOLVING THE EVOLUT PRO + 26 VALVE, A PRE-IMPLANT BALLOON DILATION WAS PERFORMED USING A 23MM X 40MM BALLOON. THE VALVE WAS IMPLANTED AT A DEPTH OF 3MM; HOWEVER, DISLODGED AND REMAINED ABOVE THE VALVE ANNULUS. A SECOND EVOLUT PRO + 29 VALVE WAS SUBSEQUENTLY IMPLANTED. POST-IMPLANT BALLOON DILATION WAS PERFORMED THREE TIMES WITH A 25MM X 40MM SEMI-COMPLIANT BALLOON TO EXPAND THE VALVE. THE PATIENT EXPERIENCED CARDIORESPIRATORY ARREST, AND RESUSCITATION MANEUVERS WERE PERFORMED FOR 30 MINUTES WITHOUT SUCCESS, RESULTING IN THE PATIENT'S DEATH. PER THE PHYSICIAN, THE CAUSE OF DEATH WAS UNKNOWN.Event Description: ADDITIONAL INFORMATION WAS RECEIVED INDICATING THAT, PER THE PHYSICIAN, THE VALVES AND DELIVERY CATHETER SYSTEM (DCS) DID NOT CONTRIBUTE TO THE PATIENT¿S DEATH. Manufacturer Narrative: IMAGE REVIEW: THERE WERE 49 CINE IMAGES OF THE EVENT PROVIDED FOR REVIEW. IMAGING SHOWED THE PATIENT¿S NATIVE ANATOMY AND CORONARY ARTERY LOCATION. IMAGING SHOWED THAT A PRE-IMPLANT DILATATION WAS PERFORMED USING A 23MM X 40MM BALLOON. EVIDENCE SHOWS THE VALVE WAS RECAPTURED AFTER ONE DEPLOYMENT AT 0MM. EVIDENCE SHOWS THAT THE VALVE WAS DEPLOYED AT APPROXIMATELY 3MM. THE VALVE WAS THEN VISUALIZED IN A LEFT ANTERIOR OBLIQUE (LAO) PROJECTION AND RELEASED. SOME TIME BETWEEN FINAL RELEASE AND WITHDRAWAL OF THE DELIVERY CATHETER SYSTEM (DCS), THE VALVE DISLODGED TO THE LEVEL OF THE SINOTUBULAR JUNCTION (STJ). THERE IS NO EVIDENCE THAT SNARING THE DISLODGED VALVE WAS ATTEMPTED, AND A SECOND VALVE WAS IMPLANTED. THERE IS EVIDENCE OF CORONARY PERFUSION IN ALL IMAGES UP UNTIL THIS POINT. THE SECOND VALVE WAS RELEASED AND A POST-IMPLANT BALLOON DILATION WAS PERFORMED IN THREE DIFFERENT LOCATIONS. THERE IS EVIDENCE THAT CARDIOPULMONARY RESUSCITATION WAS INITIATED AFTER THE POST-IMPLANT BALLOON DILATION. THE FINAL ANGIOGRAM WAS PERFORMED WITH A LARGER AMOUNT OF CONTRAST, SHOWING NO PERFUSION TO THE RIGHT CORONARY ARTERY. UPDATED: B5, H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: H6 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID: D-EVPROP23-29 (LOT: 0012426257); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: NON-IMPLANTABLE DEVICE PRODUCT ID L-EVPROP23-29 (LOT: 0012063636); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: NON-IMPLANTABLE DEVICE PRODUCT ID GWBC30 (LOT: 92106989); PRODUCT TYPE: 0193-GUIDEWIRE; IMPLANT DATE: NON-IMPLANTABLE DEVICE PRODUCT ID EVPROPLUS-29 (J292681); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: ON (B)(6) 2025; EXPLANT DATE SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21730422&pc=NPT | 2025587-2025-02222 | 2025-03-27 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-31 04:00 | NPT | EVOLUT PRO PLUS VALVE | Insufficient Information; Activation Failure; Patient Device Interaction Problem | Cardiac Arrest | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER HEART VALVE PROCEDURE INVOLVING THE EVOLUT PRO + 26 VALVE, A PRE-IMPLANT BALLOON DILATION WAS PERFORMED USING A 23MM X 40MM BALLOON. THE VALVE WAS IMPLANTED AT A DEPTH OF 3MM; HOWEVER, DISLODGED AND REMAINED ABOVE THE VALVE ANNULUS. A SECOND EVOLUT PRO + 29 VALVE WAS SUBSEQUENTLY IMPLANTED. POST-IMPLANT BALLOON DILATION WAS PERFORMED THREE TIMES WITH A 25MM X 40MM SEMI-COMPLIANT BALLOON TO EXPAND THE VALVE. THE PATIENT EXPERIENCED CARDIORESPIRATORY ARREST, AND RESUSCITATION MANEUVERS WERE PERFORMED FOR 30 MINUTES WITHOUT SUCCESS, RESULTING IN THE PATIENT'S DEATH. PER THE PHYSICIAN, THE CAUSE OF DEATH WAS UNKNOWN.Event Description: ADDITIONAL INFORMATION WAS RECEIVED INDICATING THAT, PER THE PHYSICIAN, THE VALVES AND DELIVERY CATHETER SYSTEM (DCS) DID NOT CONTRIBUTE TO THE PATIENT¿S DEATH. Manufacturer Narrative: IMAGE REVIEW: THERE WERE 49 CINE IMAGES OF THE EVENT PROVIDED FOR REVIEW. IMAGING SHOWED THE PATIENT¿S NATIVE ANATOMY AND CORONARY ARTERY LOCATION. IMAGING SHOWED THAT A PRE-IMPLANT DILATATION WAS PERFORMED USING A 23MM X 40MM BALLOON. EVIDENCE SHOWS THE VALVE WAS RECAPTURED AFTER ONE DEPLOYMENT AT 0MM. EVIDENCE SHOWS THAT THE VALVE WAS DEPLOYED AT APPROXIMATELY 3MM. THE VALVE WAS THEN VISUALIZED IN A LEFT ANTERIOR OBLIQUE (LAO) PROJECTION AND RELEASED. SOME TIME BETWEEN FINAL RELEASE AND WITHDRAWAL OF THE DELIVERY CATHETER SYSTEM (DCS), THE VALVE DISLODGED TO THE LEVEL OF THE SINOTUBULAR JUNCTION (STJ). THERE IS NO EVIDENCE THAT SNARING THE DISLODGED VALVE WAS ATTEMPTED, AND A SECOND VALVE WAS IMPLANTED. THERE IS EVIDENCE OF CORONARY PERFUSION IN ALL IMAGES UP UNTIL THIS POINT. THE SECOND VALVE WAS RELEASED AND A POST-IMPLANT BALLOON DILATION WAS PERFORMED IN THREE DIFFERENT LOCATIONS. THERE IS EVIDENCE THAT CARDIOPULMONARY RESUSCITATION WAS INITIATED AFTER THE POST-IMPLANT BALLOON DILATION. THE FINAL ANGIOGRAM WAS PERFORMED WITH A LARGER AMOUNT OF CONTRAST, SHOWING NO PERFUSION TO THE RIGHT CORONARY ARTERY. UPDATED: B5, H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED H.6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID D-EVPROP23-29 (LOT: 0012426257); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: NON-IMPLANTABLE DEVICE PRODUCT ID L-EVPROP23-29 (LOT: 0012063636); PRODUCT TYPE: 0195-HEART VALVES; NON-IMPLANTABLE DEVICE PRODUCT ID GWBC30 (LOT: 92106989); PRODUCT TYPE: 0193-GUIDEWIRE; IMPLANT DATE: NON-IMPLANTABLE DEVICE PRODUCT ID EVPROPLUS-26 (B)(6); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE (B)(6) 2025; EXPLANT DATE SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21730543&pc=LYJ | 1644487-2025-00349 | 2024-06-11 04:00 | Death | CYBERONICS - HOUSTON | 2025-03-31 04:00 | LYJ | PULSE GEN MODEL 1000 | Adverse Event Without Identified Device or Use Problem | Convulsion, Clonic | P970003 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD PASSED AWAY DUE TO STATUS EPILEPTICUS. THE PHYSICIAN EXPRESSED CONCERNS ABOUT GENERATOR PEOS FIELD SAFETY NOTIFICATION POSSIBLY BEING ASSOCIATED WITH THIS DEATH DUE TO STATUS EPILEPTICUS. LATER THE PHYSICIAN REPORTED THAT THEY DID NOT BELIEVE THAT VNS CAUSED THIS DEATH HOWEVER HE COULDN'T BE CONCLUSIVE WITH THIS DETERMINATION GIVEN THE FACT THAT THE PATIENT WAS INCLUDED ON THE PREMATURE END OF SERVICE ISSUE AFFLICTED SERIAL NUMBER LIST. HE STATED THAT THE PATIENT WAS INTERROGATED ON (B)(6) 2024 AND THERE WERE NO ISSUES OR LOW BATTERY WARNINGS ON THE DEVICE. THE PATIENT WAS THEN ADMITTED DUE TO AN INCREASE IN SEIZURES AND HAD PASSED AWAY IN THE HOSPITAL. THE PHYSICIAN HAS REPORTED THAT THE PATIENT ALSO WAS EXPERIENCING AN UNITARY INFECTION DUE TO A DISLODGING OF SUPER PUBIC CATHETER AND THIS COULD HAVE ALSO BEEN THE CAUSE FOR THE STATUS EPILEPTICUS WHICH RESULTED IN DEATH. A VNS DEATH FOLLOW-UP FORM WAS COMPLETED BY THE FOLLOWING NEUROLOGIST. THE BELIEVED CAUSE OF DEATH IS LISTED AS" "INCREASED SEIZURES FREQUENCY IN THE SETTING OF UNRAINY TRACT INFECTION. CANNOT SPEAK TO THE FUNCTIONAL STATUS OF VNS DEVICE DURING ADMISSION" THE RELATIONSHIP OF THE CAUSE OF DEATH IN ASSOCIATION WITH VNS SYSTEM IS UNKNOWN. THE PROVIDER HAS NO ADDITIONAL DETAILS ABOUT AUTOPSY. NO OTHER RELEVANT INFORMATION HAS BEEN RECEIVED TO DATE. THE SUSPECT DEVICE HAS NOT BEEN RECEIVE BY THE MANUFACTURER TO DATE.Event Description: INTERNAL GENERATOR DATA WAS RECEIVED FROM PRIOR TO THE PATIENT'S PASSING. NO ANOMALIES WERE FOUND. DESPITE FOLLOW-UP WITH THE PHYSICIAN THE CAUSE OF DEATH REMAINS UNKNOWN. Manufacturer Narrative: LIVANOVA USA, INC. SUBMITS THIS REPORT TO COMPLY WITH 21 CFR PART 803, THE MEDICAL DEVICE REPORTING REGULATION, BASED ON INFORMATION THAT LIVANOVA HAS OBTAINED, BUT MAY NOT HAVE BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION, OR A CONCLUSION BY FDA OR ANYONE ELSE, THAT THE DEVICE, LIVANOVA, OR LIVANOVA'S EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS" OR "MALFUNCTIONS". THESE WORDS ARE INCORPORATED INTO THE FDA 3500A MEDWATCH FORM BY THE FDA, AND LIVANOVA OBJECTS TO THEIR USE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21730586&pc=OTD | 9612164-2025-01612 | 2024-07-04 04:00 | Death | MEDTRONIC IRELAND | 2025-03-31 04:00 | OTD | ENDO ANCHOR SYSTEM - HELI-FX AAA | Adverse Event Without Identified Device or Use Problem | Pain; Confusion/ Disorientation; Constipation; Insufficient Information | K140036 |  | 1 | Event Description: AN ENDURANT IIS BIFURCATE STENT GRAFTS SYSTEM AND ENDOANCHORS WERE IMPLANTED IN THE ENDOVASCULAR TREATMENT OF A AAA AS PART OF THE HERCULES POST MARKET CLINICAL STUDY. APPROXIMATELY 5 WEEKS POST INDEX PROCEDURE, THE PATIENT WAS ADMITTED WITH LOWER ABDOMINAL PAIN AND CONSTIPATION. THE PATIENT WAS ALSO NOTED TO BE EXPERIENCING DELIRIUM. THE PATIENT HAD A CT SCAN THE FOLLOWING DAY WHICH SHOWED A NON-ENHANCEMENT OF THE LOWER AND MID POLES OF THE RIGHT KIDNEY IN KEEPING WITH AN ACUTE INFARCTION. THERE WAS A PERINEPHRIC INFLAMMATORY STRANDING WHICH WAS SAID TO BE LIKELY THE CAUSE OF THE PATIENTS SYMPTOMS. ABRUPT FILLING DEFECT IN THE DISTAL HILAR ASPECT OF THE MORE LATERAL BRANCH OF THE RIGHT RENAL ARTERY IS CONSISTENT WITH ARTERIAL INFARCT. MINOR INFARCT SEEN ON THE LEFT UPPER POLE OF THE LEFT KIDNEY. IT IS REPORTED THAT APPROXIMATELY 5 MONTHS LATER THE PATIENT PASSED AWAY. THE CAUSE OF DEATH IS UNKNOWN. THE DEATH WAS ASSESSED BY THE MEDICAL MONITOR TO BE POSSIBLY RELATED TO THE STENT GRAFT , NOT RELATED TO THE ENDOANCHORS, NOT RELATED TO THE INDEX PROCEDURE , NOT RELATED TO THE STENT GRAFT DEPLOYMENT AND POSSIBLY RELATED TO THE ANEURYSM. THE DEATH WAS ASSESSED BY THE SITE AS NOT RELATED TO THE STENT GRAFT, ENDOANCHORS, INDEX PROCEDURE, STENT GRAFT DEPLOYMENT OR ANEURYSM. Manufacturer Narrative: B5; ADDITIONAL INFORMATION RECEIVED: IT WAS REPORTED THE PATIENT'S WIFE HAD CALLED TO CANCEL APPOINTMENTS AND THEN REPORTED THE PATIENT'S DEATH. THERE ARE NO DETAILS AVAILABLE ABOUT THE DEATH AND IT IS UNKNOWN WHETHER IT TOOK PLACE IN A HOSPITAL OR AT HOME. THE INFORMATION REGARDING THE LOWER ABDOMINAL PAIN, CONSTIPATION, DELIRIUM AND KIDNEY INFRACTION WERE REPORTED IN ERROR AND DID NOT OCCUR IN THIS PATIENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21730659&pc=MIH | 9612164-2025-01613 | 2024-07-04 04:00 | Death | MEDTRONIC IRELAND | 2025-03-31 04:00 | MIH | ENDURANT IIS BIFURCATED STENT GRAFT | Adverse Event Without Identified Device or Use Problem | Pain; Confusion/ Disorientation; Constipation; Insufficient Information | P100021 |  | 1 | Event Description: AN ENDURANT IIS BIFURCATE STENT GRAFTS SYSTEM AND ENDOANCHORS WERE IMPLANTED IN THE ENDOVASCULAR TREATMENT OF A AAA AS PART OF THE HERCULES POST MARKET CLINICAL STUDY. APPROXIMATELY 5 WEEKS POST INDEX PROCEDURE, THE PATIENT WAS ADMITTED WITH LOWER ABDOMINAL PAIN AND CONSTIPATION. THE PATIENT WAS ALSO NOTED TO BE EXPERIENCING DELIRIUM. THE PATIENT HAD A CT SCAN THE FOLLOWING DAY WHICH SHOWED A NON-ENHANCEMENT OF THE LOWER AND MID POLES OF THE RIGHT KIDNEY IN KEEPING WITH AN ACUTE INFARCTION. THERE WAS A PERINEPHRIC INFLAMMATORY STRANDING WHICH WAS SAID TO BE LIKELY THE CAUSE OF THE PATIENTS SYMPTOMS. ABRUPT FILLING DEFECT IN THE DISTAL HILAR ASPECT OF THE MORE LATERAL BRANCH OF THE RIGHT RENAL ARTERY IS CONSISTENT WITH ARTERIAL INFARCT. MINOR INFARCT SEEN ON THE LEFT UPPER POLE OF THE LEFT KIDNEY. IT IS REPORTED THAT APPROXIMATELY 5 MONTHS LATER THE PATIENT PASSED AWAY. THE CAUSE OF DEATH IS UNKNOWN. THE DEATH WAS ASSESSED BY THE MEDICAL MONITOR TO BE POSSIBLY RELATED TO THE STENT GRAFT , NOT RELATED TO THE ENDOANCHORS, NOT RELATED TO THE INDEX PROCEDURE , NOT RELATED TO THE STENT GRAFT DEPLOYMENT AND POSSIBLY RELATED TO THE ANEURYSM. THE DEATH WAS ASSESSED BY THE SITE AS NOT RELATED TO THE STENT GRAFT, ENDOANCHORS, INDEX PROCEDURE, STENT GRAFT DEPLOYMENT OR ANEURYSM. Manufacturer Narrative: CONTINUATION OF D10: SECTION D INFORMATION REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER RELEVANT DEVICE(S) ARE: PRODUCT ID: E TLW1620C124E, SERIAL/LOT#: (B)(6), UBD: 20-FEB-2026, UDI#: (B)(4), PRODUCT ID: ETLW1616C124E, SERIAL/LOT#: (B)(6), UBD: 14-JUN-2025, UDI#: (B)(4). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: B5; ADDITIONAL INFORMATION RECEIVED : IT WAS REPORTED THE PATIENTS WIFE HAD CALLED TO CANCEL APPOINTMENTS AND THEN REPORTED THE PATIENTS DEATH. THERE ARE NO DETAILS AVAILABLE ABOUT THE DEATH AND IT IS UNKNOWN WHETHER IT TOOK PLACE IN A HOSPITAL OR AT HOME. THE INFORMATION REGARDING THE LOWER ABDOMINAL PAIN, CONSTIPATION , DELIRIUM AND KIDNEY INFRACTION WERE REPORTED IN ERROR AND DID NOT OCCUR IN THIS PATIENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21730670&pc=LWP | 2017865-2025-38908 | 2025-03-11 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-03-31 04:00 | LWP | ASSURITY MRI | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P140033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. IT WAS REPORTED THAT THE CAUSE OF DEATH IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21730677&pc=NVN | 2017865-2025-38909 | 2025-03-11 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-03-31 04:00 | NVN | ISOFLEX S LEAD | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. IT WAS REPORTED THAT THE CAUSE OF DEATH IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21730680&pc=NVN | 2017865-2025-38910 | 2025-03-11 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-03-31 04:00 | NVN | ISOFLEX S LEAD | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. IT WAS REPORTED THAT THE CAUSE OF DEATH IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21730820&pc=DYB | 2015691-2025-02424 | 2025-03-04 05:00 | Death | EDWARDS LIFESCIENCES | 2025-03-31 04:00 | DYB | ESHEATH PLUS INTRODUCER SHEATH | Material Puncture/Hole; Difficult to Remove; Failure to Advance | Hemorrhage/Bleeding; Vascular Dissection | K200258 |  | 1 | Event Description: PER REPORT RECEIVED FROM UNITED KINGDOM, IT WAS A CASE OF AN IMPLANT OF A 29MM SAPIEN 3 ULTRA RESILIA VALVE IN AORTIC POSITION BY SUBCLAVIAN APPROACH. THE ESHEATH WAS INSERTED WITHOUT PROBLEMS. HOWEVER, DURING THE INSERTION OF THE COMMANDER DELIVERY SYSTEM WITH THE VALVE INTO THE ESHEATH+, DIFFICULTIES DUE TO THE INCREASING OF THE PUSH FORCE WERE ENCOUNTERED. IT WAS OBSERVED BY FLUORO THAT THE ESHEATH+ WAS BENT SO IT WAS DECIDED TO REMOVE THE SYSTEM AS A UNIT. DURING THE REMOVAL ATTEMPT, THE SYSTEM WAS UNABLE TO BE RETRACTED AS IT WAS STUCK. THE PROCEDURE WAS CONVERTED INTO VASCULAR SURGERY TO REMOVE THE SYSTEM. A CUTDOWN WAS PERFORMED TO FORCE THE REMOVAL OF THE SYSTEM, AND DURING THIS MANEUVER, WHEN PULLING THE COMMANDER DELIVERY SYSTEM OUT, THE VALVE WAS COMPLETELY DISLODGED FROM THE COMMANDER INTO THE ARTERY. THE SYSTEM WAS EVENTUALLY REMOVED WITH GREAT DIFFICULTY BUT A SUBCLAVIAN VESSEL DISSECTION OCCURRED. THE BLEEDING OF THE VESSEL REQUIRED BLOOD TRANSFUSION AND MEDICATION TO STABILIZE THE BLOOD PRESSURE. VASCULAR SURGEON REPAIRED THE SUBCLAVIAN AND THE PATIENT WAS TRANSFERRED TO THE ICU WHERE PASSED AWAY DURING THE EVENING DUE TO THE EXCESSIVE BLOOD LOSS. AS PER THE IMAGES RECEIVED, THE VALVE HAD A BENT STRUT WHICH WAS CONFIRMED TO HAVE OCCURRED DURING THE PROCEDURE WHILE ATTEMPTING RETRIEVAL OF THE DEVICES.Event Description: AS REPORTED BY A FIELD CLINICAL SPECIALIST, A PATIENT IN THE ALTERRA PRE-STENT (ALP) CLINICAL STUDY UNDERWENT A TRANSFEMORAL TPVR PROCEDURE WITH A 29MM SAPIEN 3 VALVE WITH ALTERRA PRESTENT IN THE PULMONIC POSITION. IT WAS CONFIRMED THAT THE ALTERRA PRESTENT WAS SUCCESSFULLY IMPLANTED IN THE PATIENT AS INTENDED. NEXT, DURING WITHDRAWAL OF THE ALTERRA DS DEVICE FROM THE PRE-STENT, THE TEAM ENCOUNTERED DIFFICULTY/RESISTANCE AT THE LEVEL OF THE TRICUSPID ANATOMY. A NON-EDWARDS SHEATH WAS IN PLACE DURING THE INSERTION AND WITHDRAWAL OF THE ALTERRA PRESENT SYSTEM. AT THIS TIME, THERE WAS NO TR OBSERVED. NEXT, THE 29MM S3 WAS IMPLANTED SUCCESSFULLY WITHIN THE PRE-STENT.. DURING THE WITHDRAWAL OF THE PDS, THE TEAM AGAIN ENCOUNTERED WITHDRAWAL DIFFICULTY AT THE SAME ANATOMICAL PLACE. THIS IS WHERE THEY BELIEVE THE TRICUSPID VALVE INJURY LIKELY TOOK PLACE. FOLLOWING THE REMOVAL OF ALL DEVICES, AN INJURED TRICUSPID CHORDAE WAS OBSERVED ON ICE CAUSING MODERATE TR. THE PATIENT WAS HEMODYNAMICALLY STABLE AND THERE IS NO PLAN FOR INTERVENTION TO TREAT THE TR. Manufacturer Narrative: CORRECTED: B1, B2, B5, H1, H6. UPDATED B4, H2, G3, H10, H11. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED DUE TO ENGINEERING EVALUATION FINDINGS. SECTIONS B4, D9, G3, G6, H2, H6 TYPE OF INVESTIGATION, INVESTIGATION FINDINGS, AND INVESTIGATION CONCLUSIONS AND H11 HAS BEEN UPDATED. THE EVENTS REPORTED ARE ANTICIPATED IN THE RISK MANAGEMENT DOCUMENTATION FOR TRANSCATHETER HEART VALVE PROCEDURES. A PREVIOUS INVESTIGATION INTO THIS TYPE OF EVENT IS CAPTURED IN AN EDWARDS LIFESCIENCES TECHNICAL SUMMARY AND APPLIES TO THIS COMPLAINT. ADDITIONAL ASSESSMENT OF THE FAILURE MODE(S) IS NOT REQUIRED AT THIS TIME. THE DEVICE WAS NOT RETURNED TO EDWARDS LIFESCIENCES FOR EVALUATION. DUE TO THE UNAVAILABILITY OF THE COMPLAINT DEVICE, ENGINEERING WAS UNABLE TO PERFORM ANY VISUAL INSPECTION, FUNCTIONAL TESTING, OR DIMENSIONAL ANALYSIS. PER THE TECHNICAL SUMMARY, THE INSTRUCTIONS FOR USE (IFU), TRAINING MANUALS AND CURRENT RISK MITIGATIONS HAVE BEEN REVIEWED. NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. THROUGH EXTENSIVE COMPLAINT INVESTIGATIONS, EVENTS REPORTING RESISTANCE DURING DELIVERY SYSTEM INSERTION/ADVANCEMENT THROUGH THE SHEATH AND POTENTIAL VALVE FRAME DAMAGE USING THE S3U/S3UR VALVE CONFIGURATION HAVE NOT BEEN ASSOCIATED WITH DEVICE MALFUNCTIONS OR MANUFACTURING NONCONFORMANCES. RATHER, THE ROOT CAUSE FOR THESE EVENTS HAVE HISTORICALLY BEEN DUE TO VESSEL TORTUOSITY, CALCIFICATION, UNDERSIZED VESSELS, AND/OR STEEP INSERTION ANGLE. IN ADDITION, VALVE FRAME AND/OR SHEATH DAMAGE CAN BE A RESULT OF INCREASED PUSH FORCE AND ANY EXCESSIVE DEVICE MANIPULATION OR SHEATH-VALVE INTERACTION. THE COMPLAINT FOR INABILITY TO ADVANCE THROUGH SHEATH WAS CONFIRMED BASED ON THE EVALUATION OF PROVIDED IMAGERY. AVAILABLE INFORMATION SUGGESTS THAT PATIENT FACTORS (CALCIFICATION) LIKELY CONTRIBUTED TO THE EVENT AS PER INFORMATION PROVIDED THERE WAS MODERATE DEGREE OF CALCIFICATION IN THE ACCESS VESSEL. CALCIFICATION CAN REDUCE THE VESSEL LUMEN DIAMETER AND MAY INCREASE RESTRICTION LEADING TO RESISTANCE. CALCIFICATION CAN ALSO RESULT IN THE CREATION OF SUB-OPTIMAL ANGLES DURING DELIVERY SYSTEM INSERTION THAT MAY LEAD TO RESISTANCE. THE COMPLAINTS FOR SHEATH LINER PUNCTURED AND KINKED/BENT WERE CONFIRMED BASED ON THE EVALUATION OF PROVIDED IMAGERY. AVAILABLE INFORMATION SUGGESTS THAT PATIENT FACTORS (CALCIFICATION) AND PROCEDURAL FACTORS (EXCESSIVE DEVICE MANIPULATION) LIKELY CONTRIBUTED TO THE EVENT AS PER INFORMATION PROVIDED PATIENT THERE WAS MODERATE DEGREE OF CALCIFICATION IN THE ACCESS VESSEL. IT IS POSSIBLE THAT ADDITIONAL DEVICE MANIPULATION TO OVERCOME THE RESISTANCE FELT DURING DELIVERY SYSTEM ADVANCEMENT MAY BE APPLIED RESULTING IN DAMAGE THE SHEATH KINK AND LINER TORN. DURING DELIVERY SYSTEM ADVANCEMENT THROUGH A CHALLENGING PATHWAY, IT IS POSSIBLE THAT THE DEVICES MAY NOT BE COAXIALLY ALIGNED. THIS CAN LEAD TO THE CRIMPED VALVE TO CATCH ONTO THE SHEATH HDPE AND LINER, LEADING TO THE SHEATH KINK AND LINER TORN. EXCESSIVE DEVICE MANIPULATION APPLIED TO OVERCOME THE RESISTANCE CAN FURTHER DAMAGE THE SHEATH THROUGH CONTINUED INTERACTION BETWEEN THE CRIMPED VALVE AND SHEATH. VESSEL CALCIFICATIONS PRESENT AT THE ACCESS VESSEL CAN EXASPERATE THE INTERACTION BETWEEN THE CRIMPED VALVE AND SHEATH. SHARP CALCIFIED NODULES CAN DIRECTLY WEAKEN THE SHEATH SHAFT MAKING IT MORE SUSCEPTIBLE TO DAMAGE AS THE DELIVERY SYSTEM WITH CRIMPED VALVE IS ADVANCED THROUGH. EXCESSIVE MANIPULATION CAN LEAD TO SHEATH SHAFT DAMAGE (SHEATH SHAFT KINK AND LINER TORN) IF COMPOUNDED WITH VESSEL CALCIFICATION. THE COMPLAINT FOR INABILITY TO REMOVE SHEATH WAS CONFIRMED BY REVIEW OF THE PROVIDED IMAGERY. NO MANUFACTURING NON-CONFORMANCES THAT WOULD HAVE CONTRIBUTED TO THE COMPLAINT EVENT WERE IDENTIFIED. REVIEWS OF THE DHR, LOT HISTORY, AND COMPLAINT HISTORY DID NOT PROVIDE ANY INDICATION THAT A MANUFACTURING NON-CONFORMANCE WOULD HAVE CONTRIBUTED TO THE COMPLAINT EVENT. A REVIEW OF IFU/TRAINING MATERIALS REVEALED NO DEFICIENCIES. FURTHERMORE, THERE WAS NO REPORT OF ANY ISSUES WITH THE SHEATH DURING DEVICE UNPACKING OR PREPARATION. AS REPORTED "DURING THE REMOVAL ATTEMPT, THE SYSTEM WAS UNABLE TO BE RETRACTED AS IT WAS STUCK. THE PROCEDURE WAS CONVERTED INTO VASCULAR SURGERY TO REMOVE THE SYSTEM. A CUTDOWN WAS PERFORMED TO FORCE THE REMOVAL OF THE SYSTEM, AND DURING THIS MANEUVER, WHEN PULLING THE COMMANDER DELIVERY SYSTEM OUT, THE VALVE WAS COMPLETELY DISLODGED FROM THE COMMANDER INTO THE ARTERY. THE SYSTEM WAS EVENTUALLY REMOVED WITH GREAT DIFFICULTY, BUT A SUBCLAVIAN VESSEL DISSECTION OCCURRED." PER IFU/TRAINING MANUAL "FOR SUBCLAVIAN-AXILLARY APPROACH, KEEP DELIVERY SYSTEM INSIDE SHEATH UNTIL READY TO REMOVE ALL DEVICES AS ONE UNIT." AS PER PROVIDED IMAGERY, THE DELIVERY SYSTEM WAS OBSERVED PROTRUDING THROUGH THE LINER TEAR. IT IS POSSIBLE THAT THE CRIMPED VALVE EXITED THROUGH SHEATH LINER TEAR BEING MORE SUSCEPTIBLE TO CATCH ON THE LINER AND/OR THE VASCULATURE DURING WITHDRAWAL LEADING TO THE REPORTED WITHDRAWAL DIFFICULTY. ADDITIONALLY THE SHEATH SHAFT WAS OBSERVED TO HAVE A KINK, BOTH CONDITIONS (LINER TEAR AND SHEATH SHAFT KINK) MAY HAVE CREATED A CHALLENGING CONDITION TO REMOVE THE DEVICES. IN ADDITION, PER PATIENT INFORMATION THERE WAS PRESENCE OF MODERATE CALCIFICATION AT THE ACCESS VESSEL. CALCIFICATION PRESENT IN THE PATIENT ANATOMY COULD CREATE A CONSTRAINED CONDITION OR SUBJECT THE DEVICES TO SUBOPTIMAL ANGLES, WHICH MAY FURTHER CONTRIBUTE TO THE REPORTED WITHDRAWAL DIFFICULTY. AS SUCH, AVAILABLE INFORMATION SUGGESTS THAT PROCEDURAL FACTORS (SHEATH LINER TORN) AND PATIENT FACTORS (CALCIFICATION) MAY HAVE CONTRIBUTED TO THE REPORTED EVENT. HOWEVER, A DEFINITIVE ROOT CAUSE IS UNABLE TO BE DETERMINED AT THIS TIME. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. SINCE NO EDWARDS PRODUCT DEFECTS OR LABELING DEFICIENCIES WERE IDENTIFIED, NO CORRECTIVE OR PREVENTATIVE ACTION IS REQUIRED. Manufacturer Narrative: INVESTIGATION IS ONGOING. THIS IS ONE OF THREE REPORTS BEING SUBMITTED FOR THIS CASE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21730821&pc=DSQ | 2916596-2025-01722 | 2025-03-04 05:00 | Death | THORATEC CORPORATION | 2025-03-31 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PRE-OPERATIVELY, THE PATIENT WAS CONSIDERED FOR POTENTIAL RIGHT VENTRICULAR ASSIST DEVICE (RVAD) IMPLANTATION. AT THE TIME OF OPERATION, THE PATIENT WAS IMPLANTED WITH A THIRD-PARTY (PROTEKDUO) CANNULA AND THEN PUT ON THE CENTRIMAG CONSOLE. THE VENTRICULAR ASSIST DEVICE (VAD) TEAM WAS UNABLE TO WEAN RVAD SUPPORT POST-OPERATIVELY, AND THE PATIENT FAMILY DECIDED TO WITHDRAW VAD SUPPORT IN COMPLIANCE WITH THE PATIENT'S WISHES AND DUE TO NO HOPE FOR A MEANINGFUL QUALITY OF LIFE. THE PATIENT UNDERWENT MULTI-SYSTEM ORGAN FAILURE (MSOF) AND PASSED AWAY. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED; THE VAD HAD FUNCTIONED AS INTENDED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (INCLUDING RIGHT HEART FAILURE), THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 OF THE IFU, UNDER ¿RIGHT HEART FAILURE¿, ALSO STATES THAT PATIENTS MAY DEVELOP RIGHT VENTRICULAR FAILURE DURING OR SHORTLY AFTER IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21730860&pc=OZD | 1220648-2025-27093 | 2025-03-03 05:00 | Death | ABIOMED, INC. | 2025-03-31 04:00 | OZD | IMPELLA CP-JP | Device Sensing Problem | Great Vessel Perforation | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT ON THE IMPELLA CP FOR ACUTE MYOCARDIAL INFARCTION. IT WAS REPORTED THAT THE PUMP HAD SUCTION, POSITION UNKNOWN, AND POSITION IN AORTA ALARMS WHILE ON PATIENT SUPPORT. THE POSITIONING OF THE IMPELLA WAS CHECKED AND NO PROBLEMS WERE NOTED. IT WAS FURTHER REPORTED THAT TRANSTHORACIC ECHOCARDIOGRAM REVEALED PERICARDIAL EFFUSION FROM VENTRICULAR SEPTAL PERFORATION (VSP) AND OOZING RUPTURE. THE RELATIONSHIP BETWEEN OOZING RUPTURE/PERICARDIAL EFFUSION AND IMPELLA IS UNKNOWN. THE PATIENT EVENTUALLY EXPIRED ON SUPPORT. Manufacturer Narrative: THE INVESTIGATION OF OPTICAL SIGNAL ISSUE AND VENTRICLE PERFORATION HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. E4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27093 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. G1 REPORTING CONTACT FACILITY NAME AND FAX NUMBER WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27093. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMART ASSIST SYSTEM STATES THE FOLLOWING: SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21730871&pc=OZD | 1220648-2025-27094 | 2025-03-04 05:00 | Death | ABIOMED, INC. | 2025-03-31 04:00 | OZD | IMPELLA CP | Failure to Advance; Malposition of Device | Cardiac Arrest | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THE INSERTION OF IMPELLA CP DEVICE TO A PATIENT (UNKNOWN RACE AND ETHNICITY) IN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS UNSUCCESSFUL AND THE PATIENT WOULD EXPIRE ON THE PROCEDURE TABLE IN THE CARDIOVASCULAR LAB. THE PATIENT PRESENTED TO AN OUTSIDE FACILITY AND WAS FOUND TO BE HAVING A ST-SEGEMENT EVALUATION MYOCARDIAL INFARCTION (STEMI). THE PATIENT WAS GIVEN TISSUE PLASMINOGEN ACTIVATOR (TPA) AND EMERGENTLY TRANSFERRED IN. ON ARRIVAL, A REPEAT ELECTROCARDIOGRAM CONFIRMED STEMI; THE PATIENT WAS EMERGENTLY BROUGHT TO THE CARDIOVASCULAR LAB. A LEFT HEART CATHETERIZATION FOUND 100% OCCLUDED LEFT ANTERIOR DESCENDING ARTERY (LAD). THE PHYSICIAN UTILIZED INTRACORONARY THROMBOLYSIS TO DISSOLVE A THROMBUS IN THE LAD WHICH RESTORED THROMBOLYSIS IN MYOCARDIAL INFARCTION 3 FLOW. UPON REPERFUSION, THE PATIENT WENT PULSELESS AND CARDIOPULMONARY RESUSCITATION (CPR) WAS INITIATED. MID-CPR, THE IMPELLA WAS EMERGENTLY PREPPED AND PRIMED. THE PEEL-AWAY SHEATH WAS INSERTED. DUE TO CPR, THE PLACEMENT GUIDEWIRE SLIPPED OUT OF THE LEFT VENTRICLE WHILE ADVANCING THE IMPELLA AND THE PHYSICIAN WAS UNABLE TO SUCCESSFULLY PASS THE PIGTAIL ACROSS THE AORTIC VALVE. CPR CONTINUED WHILE THE PHYSICIAN ATTEMPTED PLACEMENT. A PULSE CHECK WAS DONE, AND THE PHYSICIAN MADE THE DECISION TO ABORT BOTH CPR AND IMPELLA PLACEMENT AFTER INABILITY TO ACHIEVE RETURN OF SPONTANEOUS CIRCULATION. THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA CP AS AN ASSOCIATED FACTOR TO THE EVENT. HOWEVER, IT SHOULD BE NOTED THE PATIENT WAS IN A STATE OF DECLINE WITHOUT PULSE AT THE ONSET OF IMPELLA CP PLACEMENT ATTEMPT. Manufacturer Narrative: THE INVESTIGATION OF DELIVERY ISSUES HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. BASED ON THE CLINICAL DETAILS THE ROOT CAUSE OF DELIVERY ISSUE IS DETERMINED TO BE USE ISSUE BECAUSE THE ISSUE HAPPENED DURING PERFORMING THE CPR. INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH RISK PERCUTANEOUS CORONARY INTERVENTION SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿. SECTION: WARNINGS & CAUTIONS: WARNINGS SECTION: PRE-SUPPORT EVALUATION SECTION: DIRECT AORTIC INSERTION SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES ¿TO REDUCE THE RISK OF CARDIAC INJURY (INCLUDING VENTRICULAR PERFORATION), PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY IN THE SETTINGS OF ACUTE INFARCTION WITH TISSUE NECROSIS.¿ ¿TO REDUCE THE RISK OF VASCULAR INJURY, PHYSICIANS SHOULD EXERCISE CAUTION WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX PERIPHERAL VASCULAR ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: UNREPAIRED ABDOMINAL AORTIC ANEURYSM, SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM, DISSECTION OF THE ASCENDING/ TRANSVERSE/DESCENDING AORTA, CHRONIC ANATOMICAL CHANGES IN THE RELATIONSHIP OF THE AORTA/AORTIC VALVE/VENTRICULAR ALIGNMENT, SIGNIFICANT MOBILE ATHEROMATOUS DISEASE IN THE THORACIC OR ABDOMINAL AORTA OR PERIPHERAL VESSELS.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING A RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY IN THE LEFT VENTRICLE AFTER CPR WITH ECHOCARDIOGRAPHY GUIDANCE. SECTION: WARNINGS & CAUTIONS: CAUTIONS: ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS, SECTION: PRE-SUPPORT EVALUATION, SECTION: DIRECT AORTIC INSERTION, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿ D6A AND D6B WERE INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-2709. E4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27094 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. H10 INSTRUCTIONS FOR USE WERE INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-2709. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21730934&pc=NIQ | 2024168-2025-03244 | 2024-11-30 05:00 | Death | ABBOTT VASCULAR INC. | 2025-03-31 04:00 | NIQ | XIENCE SKYPOINT¿ | Adverse Event Without Identified Device or Use Problem | Unspecified Heart Problem | P110019 |  | 1 | Event Description: IT WAS REPORTED THAT THE PROCEDURE ON (B)(6) 2024 WAS TO TREAT THE RIGHT CORONARY ARTERY (RCA) WITH NO CALCIFICATION AND MILD TORTUOSITY. THE LESION WAS CONFIRMED WITH INTRAVASCULAR ULTRASOUND (IVUS), PRE-DILATED WITH A NON-COMPLIANT BALLOON AND THE 3.0 X 38 MM XIENCE SKYPOINT STENT WAS IMPLANTED. THE STENT WAS POST-DILATED AND THE PROCEDURE WAS COMPLETED WITH NO COMPLICATIONS. ON (B)(6) 2024 THE PATIENT RETURNED FOR A PLANNED PROCEDURE TO TREAT THE LEFT ANTERIOR DESCENDING (LAD) CORONARY ARTERY WITH MODERATE CALCIFICATION AND MILD TORTUOSITY. THE LESION WAS CONFIRMED WITH FRACTIONAL FLOW RESERVE (FFR) AND OPTICAL COHERENCE TOMOGRAPHY (OCT). INTRAVASCULAR LITHOTRIPSY (IVL) WAS PERFORMED DUE TO THE MODERATE CALCIFICATION AND THE 2.5 X 33 MM XIENCE SKYPOINT STENT WAS IMPLANTED AND POST-DILATED. THE PROCEDURE WAS SUCCESSFUL WITH NO COMPLICATIONS. THE PATIENT DID NOT VISIT THE OUTPATIENT CLINIC FOR THE 1 YEAR FOLLOW UP AND IT WAS FOUND THAT THE PATIENT EXPIRED ON (B)(6) 2024 DUE TO AN UNKNOWN CAUSE. THE MEDICAL INSTITUTION'S OPINION IS THAT THE EVENT IS UNRELATED TO THE STUDY; HOWEVER, THE DEATH IS POTENTIALLY RELATED TO THE STENTS, PER THE PHYSICIAN. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. PRODUCTION RECORD AND CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A QUERY OF THE COMPLAINT HANDLING DATABASE FOR THE REPORTED LOT REVEALED THERE IS NO INDICATION OF A LOT SPECIFIC ISSUE. A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECTS, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED. THE REPORTED PATIENT EFFECT OF DEATH IS LISTED IN THE XIENCE SKYPOINT EVEROLIMUS ELUTING CORONARY STENT SYSTEMS INSTRUCTIONS FOR USE AS A KNOWN PATIENT EFFECT OF CORONARY STENTING PROCEDURES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN, OR LABELING; THEREFORE, NO PRODUCT-RELATED CORRECTIVE ACTION WILL BE IMPLEMENTED IN THIS CASE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21730955&pc=NPT | 2015691-2025-02425 | 2025-03-04 05:00 | Death | EDWARDS LIFESCIENCES | 2025-03-31 04:00 | NPT | SAPIEN 3 ULTRA RESILIA VALVE | Material Deformation | Vascular Dissection | P140031 |  | 1 | Event Description: PER REPORT RECEIVED FROM UNITED KINGDOM, IT WAS A CASE OF AN IMPLANT OF A 29MM SAPIEN 3 ULTRA RESILIA VALVE IN AORTIC POSITION BY SUBCLAVIAN APPROACH. THE ESHEATH WAS INSERTED WITHOUT PROBLEMS. HOWEVER, DURING THE INSERTION OF THE COMMANDER DELIVERY SYSTEM WITH THE VALVE INTO THE ESHEATH+, DIFFICULTIES DUE TO THE INCREASING OF THE PUSH FORCE WERE ENCOUNTERED. IT WAS OBSERVED BY FLUORO THAT THE ESHEATH+ WAS BENT SO IT WAS DECIDED TO REMOVE THE SYSTEM AS A UNIT. DURING THE REMOVAL ATTEMPT, THE SYSTEM WAS UNABLE TO BE RETRACTED AS IT WAS STUCK. THE PROCEDURE WAS CONVERTED INTO VASCULAR SURGERY TO REMOVE THE SYSTEM. A CUTDOWN WAS PERFORMED TO FORCE THE REMOVAL OF THE SYSTEM, AND DURING THIS MANEUVER, WHEN PULLING THE COMMANDER DELIVERY SYSTEM OUT, THE VALVE WAS COMPLETELY DISLODGED FROM THE COMMANDER INTO THE ARTERY. THE SYSTEM WAS EVENTUALLY REMOVED WITH GREAT DIFFICULTY BUT A SUBCLAVIAN VESSEL DISSECTION OCCURRED. THE BLEEDING OF THE VESSEL REQUIRED BLOOD TRANSFUSION AND MEDICATION TO STABILIZE THE BLOOD PRESSURE. VASCULAR SURGEON REPAIRED THE SUBCLAVIAN AND THE PATIENT WAS TRANSFERRED TO THE ICU WHERE PASSED AWAY DURING THE EVENING DUE TO THE EXCESSIVE BLOOD LOSS. AS PER THE IMAGES RECEIVED, THE VALVE HAD A BENT STRUT WHICH WAS CONFIRMED TO HAVE OCCURRED DURING THE PROCEDURE WHILE ATTEMPTING RETRIEVAL OF THE DEVICES. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED DUE TO ENGINEERING EVALUATION FINDINGS. SECTIONS B4, G3, G6, H2, H6 COMPONENT CODE, TYPE OF INVESTIGATION, INVESTIGATION FINDINGS, INVESTIGATION CONCLUSIONS AND H11 HAS BEEN UPDATED. THE COMPLAINT FOR FRAME DAMAGE WAS CONFIRMED BASED ON IMAGERY PROVIDED FOR EVALUATION. AS THE DEVICE WAS NOT RETURNED, ENGINEERING WAS UNABLE TO PERFORM ANY VISUAL EXAMINATION, FUNCTIONAL TESTING, OR DIMENSIONAL ANALYSIS. THEREFORE, THE PRESENCE OF A MANUFACTURING NON-CONFORMANCE WAS UNABLE TO BE DETERMINED. HOWEVER, REVIEW OF THE DHR AND LOT HISTORY DID NOT PROVIDE ANY INDICATION THAT A MANUFACTURING NON-CONFORMANCE WOULD HAVE CONTRIBUTED TO THE COMPLAINT EVENT. A REVIEW OF IFU/TRAINING MATERIALS REVEALED NO DEFICIENCIES. FURTHERMORE, NO ABNORMALITIES WERE REPORTED DURING DEVICE UNPACKING OR PREPARATION. AS REPORTED, "DURING THE REMOVAL ATTEMPT, THE SYSTEM WAS UNABLE TO BE RETRACTED AS IT WAS STUCK. THE PROCEDURE WAS CONVERTED INTO VASCULAR SURGERY TO REMOVE THE SYSTEM. A CUTDOWN WAS PERFORMED TO FORCE THE REMOVAL OF THE SYSTEM, AND DURING THIS MANEUVER, WHEN PULLING THE COMMANDER DELIVERY SYSTEM OUT, THE VALVE WAS COMPLETELY DISLODGED FROM THE COMMANDER INTO THE ARTERY. THE SYSTEM WAS EVENTUALLY REMOVED WITH GREAT DIFFICULTY BUT A SUBCLAVIAN VESSEL DISSECTION OCCURRED." DIFFICULTY WAS ENCOUNTERED DURING WITHDRAWAL OF THE DELIVERY SYSTEM WITH CRIMPED VALVE, SO ADDITIONAL MANIPULATION WAS APPLIED. IF EXCESSIVE FORCE WAS APPLIED TO MANIPULATE THE DEVICE, IT CAN LEAD TO THE CRIMPED VALVE INTERACTING WITH THE SHEATH AND THE DELIVERY SYSTEM FLEX TIP, RESULTING IN THE VALVE GETTING CAUGHT ON THE SHEATH/FLEX TIP AND STRUT DAMAGE AT OUTFLOW SIDE. PER TRAINING MANUAL, "WITHDRAW THE DELIVERY SYSTEM AND SHEATH AS A SINGLE UNIT COMPLETELY FROM THE ARTERIOTOMY WHILE MAINTAIN GUIDEWIRE POSITION." ADDITIONALLY, PER PATIENT INFORMATION THERE WAS PRESENCE OF MODERATE CALCIFICATION AT THE ACCESS VESSEL. CALCIFICATION CAN REDUCE THE VESSEL DIAMETER AND MAY INCREASE RESTRICTION LEADING TO RESISTANCE. CALCIFICATION CAN RESULT IN THE CREATION OF SUB-OPTIMAL ANGLES DURING DELIVERY SYSTEM WITHDRAWAL THAT MAY LEAD TO RESISTANCE. AS SUCH, AVAILABLE INFORMATION SUGGESTS THAT PROCEDURAL FACTORS (EXCESSIVE MANIPULATION, WITHDRAWAL OF CRIMPED VALVE) AND PATIENT FACTORS (CALCIFICATION) MAY HAVE CONTRIBUTED TO THE COMPLAINT EVENT. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME. Manufacturer Narrative: INVESTIGATION IS ONGOING. THIS IS ONE OF TWO REPORTS BEING SUBMITTED FOR THIS CASE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21731043&pc=KZL | 3012307300-2025-03499 | 2024-12-12 05:00 | Death | SMITHS HEALTHCARE MANUFACTURING S.A. DE C.V. | 2025-03-31 04:00 | KZL | LEVEL 1 FAST FLOW TRAUMA FLUID WARMER | Collapse; Fluid/Blood Leak; Insufficient Flow or Under Infusion | Insufficient Information | UNKNOWN |  | 1 | Event Description: IT WAS REPORTED THAT SPIKING THE BLOOD TUBING WAS DIFFICULT AND THE DEVICE HAD A LOW FLOW RATE. THERE WAS LEAKING FROM THE BAG AROUND THE SPIKE AND DRIP CHAMBERS, WHICH HAD COLLAPSED DURING INFUSION. THE PATIENT WAS PART OF A MASSIVE TRANSFUSION PROTOCOL AND HAD RECEIVED MULTIPLE BLOOD PRODUCTS AS WELL AS MEDICATIONS. THOUGH THE PATIENT DID RECEIVE BLOOD AND TREATMENT, THE PATIENT ULTIMATELY PASSED AWAY. THE PRODUCT THAT WAS INITIALLY USED IS STILL AVAILABLE AND DOES HAVE COAGULATED BLOOD THROUGHOUT THE TUBING WHICH WAS UNABLE TO BE FLUSHED. A REVIEW BY THE ICU MEDICAL TEAM WAS COMPLETED AND THEY FOUND THAT THE LACK OF DETAILS REGARDING THIS CLINICAL EVENT PRECLUDES A COMPREHENSIVE MEDICAL ASSESSMENT. BASED ON A REVIEW OF THE EVENTS IN THIS CASE, THE CUMULATIVE EVIDENCE DOES NOT ALLOW FOR ANY CONCLUSIONS TO BE DRAWN FOR THIS INCIDENT. IF ADDITIONAL INFORMATION BECOMES AVAILABLE, THIS MEDICAL ASSESSMENT WILL BE REVISED AS NEEDED. Manufacturer Narrative: THIS FILE WAS ORIGINALLY INCORRECTLY FILED UNDER REGISTRATION NUMBER: 9617604-2025-00084. THE DATE OF THAT SUBMISSION WAS 03-FEB-2025. D4. PRIMARY UDI NUMBER: THE PRIMARY DI# HAS BEEN USED AS THE LOT/SERIAL INFORMATION IS UNKNOWN. DEVICE EVALUATION: AS RECEIVED, THE FLUID WARMER SET HAD NO DAMAGES OBSERVED. BLOOD RESIDUALS WERE OBSERVED INSIDE OF THE LOWER FILTER TUBE. THE TWO USED IV BAGS CONTAINED COAGULATED BLOOD THROUGHOUT THE TUBING. THE SAMPLE SET WAS INSTALLED ONTO A LEVEL 1 INFUSER IN AN ATTEMPT TO REPLICATE CLINICAL USE. THE INFUSER DID NOT SET OFF ANY ALARMS. NO LEAKS OR OTHER ANOMALIES WERE OBSERVED. THE SPIKES WERE REMOVED FROM THE IV BAGS RECEIVED AND REINSERTED. THERE WERE NO DIFFICULTIES IN REINSERTION. DURING REFILLING OF THE IV BAGS, NO LEAK WAS OBSERVED DURING AND AFTER THE INFUSER OPERATION. THE SAMPLE SET AND THE LEVEL 1 INFUSER WERE RUN THREE TIMES, WITH 500ML OF FLUID. ONCE EACH WITH RO WATER, 0.9% SALINE SOLUTION, AND WITH SIMULATED BLOOD. PROPER PRIMING PROCEDURES WERE PERFORMED BEFORE OPERATION. NO LEAKS NOR ALARMS WERE OBSERVED WITH EACH RUN. THE TWO SPIKES WERE MEASURED WITH CALIPERS AND A GAGE SET TO DETERMINE IF THEY ARE WITHIN SPECIFICATIONS. THE SPIKE TIPS WERE OBSERVED TO BE WITHIN MEASUREMENT SPECS. THE COMPLAINT OF SPIKING DIFFICULTY, DRIP CHAMBER COLLAPSE, IV BAG LEAKAGE, AND LOW FLOW RATE COULD NOT BE CONFIRMED OR REPLICATED. A LOT HISTORY REVIEW COULD NOT BE CONDUCTED BECAUSE NO LOT NUMBER(S) WAS/WERE IDENTIFIED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21731120&pc=DSQ | 2916596-2025-02014 | 2025-01-05 05:00 | Death | THORATEC CORPORATION | 2025-03-31 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTIPLE SYSTEM ORGAN FAILURE. THE DEVICE WAS NOT THE CAUSE OF DEATH, AND THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: A4: PATIENT WEIGHT NOT PROVIDED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT'S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE DEVICE WAS NOT EXPLANTED AND WOULD NOT BE RETURNED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION AND DEATH AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21731179&pc=LLZ | 2249582-2025-00001 | 2025-03-12 04:00 | Death | AGFA HEALTHCARE N.V. | 2025-03-31 04:00 | LLZ | IMPAX | Installation-Related Problem; Insufficient Information | Cardiac Arrest; Pneumonia |  |  | 1 | Event Description: ON MARCH 12, 2025, A CUSTOMER REPORTED THE FOLLOWING INCIDENT TO AGFA HEALTHCARE THROUGH COMPLAINT (B)(4): A 40-YEAR-OLD FEMALE PATIENT ARRIVED IN THE EMERGENCY IN CARDIAC AND RESPIRATORY ARREST. WHEN THE RADIOLOGIST, READING FROM A REMOTE HOME STATION, ATTEMPTED TO VIEW THE IMAGES, AN ERROR MESSAGE 'PROBLEM DOWNLOADING IMAGE, WILL ATTEMPT TO RETRY" DISPLAYED AND COULD NOT VIEW THE CT AND CTA IMAGES. THE RADIOLOGIST USER RESTARTED THE WORKSTATION WHICH TOOK APPROXIMATELY 4-5 MINUTES AFTER WHICH THE IMAGES COULD BE SUCCESSFULLY VIEWED. THE CUSTOMER REPORTED THAT THE INABILITY TO VIEW THE IMAGES DURING THE EMERGENCY SITUATION CAUSED THEM TO LOSE VALUABLE MINUTES IN MAKING THE DIAGNOSIS. THE VIEWING ISSUE HAPPENED AROUND 8 AM AND THE PATIENT PASSED AWAY AT 10:25 PM LATER THAT DAY. SUMMARY: TAKING ALL FACTORS INTO CONSIDERATION, INCLUDING PATIENT HISTORY, CLINICAL PRESENTATION UPON ED ARRIVAL WITH FIXED PUPILS, DIAGNOSTIC IMAGING FINDINGS, MEDICAL MANAGEMENT PRE AND POST DIAGNOSTIC IMAGING ACQUISITION, AND CAUSE OF DEATH; BASED ON CLINICAL EXPERIENCE, AGFA HEALTHCARE BELIEVES THAT THE 5 MINUTE DELAY IN IMAGE VIEWING BY THE READING RADIOLOGIST DID NOT CAUSE OR CONTRIBUTE TO THE PATIENT'S DEATH. ALSO, AGFA'S RISK ANALYSIS INDICATES THAT THE POTENTIAL FOR HARM FALLS BELOW THE REPORTING THRESHOLD (THE DEVICE WOULD BE NOT LIKELY TO CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY IF THE ISSUE WERE TO RECUR). AS THE CUSTOMER SUBMITTED AN INCIDENT INVOLVING THE DEATH OF A PATIENT, AGFA HEALTHCARE DECIDED TO REPORT THIS INCIDENT REGARDLESS.Event Description: ON (B)(6) 2025, A CUSTOMER REPORTED THE FOLLOWING INCIDENT TO AGFA HEALTHCARE THROUGH COMPLAINT (B)(4): A 40-YEAR-OLD FEMALE PATIENT ARRIVED IN THE EMERGENCY IN CARDIAC AND RESPIRATORY ARREST. WHEN THE RADIOLOGIST, READING FROM A REMOTE HOME STATION, ATTEMPTED TO VIEW THE IMAGES, AN ERROR MESSAGE 'PROBLEM DOWNLOADING IMAGE, WILL ATTEMPT TO RETRY" DISPLAYED AND COULD NOT VIEW THE CT AND CTA IMAGES. THE RADIOLOGIST USER RESTARTED THE WORKSTATION WHICH TOOK APPROXIMATELY 4-5 MINUTES AFTER WHICH THE IMAGES COULD BE SUCCESSFULLY VIEWED. THE CUSTOMER REPORTED THAT THE INABILITY TO VIEW THE IMAGES DURING THE EMERGENCY SITUATION CAUSED THEM TO LOSE VALUABLE MINUTES IN MAKING THE DIAGNOSIS. THE VIEWING ISSUE HAPPENED AROUND 8 AM AND THE PATIENT PASSED AWAY AT 10:25 PM LATER THAT DAY. SUMMARY: TAKING ALL FACTORS INTO CONSIDERATION, INCLUDING PATIENT HISTORY, CLINICAL PRESENTATION UPON ED ARRIVAL WITH FIXED PUPILS, DIAGNOSTIC IMAGING FINDINGS, MEDICAL MANAGEMENT PRE AND POST DIAGNOSTIC IMAGING ACQUISITION, AND CAUSE OF DEATH; BASED ON CLINICAL EXPERIENCE, AGFA HEALTHCARE BELIEVES THAT THE 5 MINUTE DELAY IN IMAGE VIEWING BY THE READING RADIOLOGIST DID NOT CAUSE OR CONTRIBUTE TO THE PATIENT'S DEATH. ALSO, AGFA'S RISK ANALYSIS INDICATES THAT THE POTENTIAL FOR HARM FALLS BELOW THE REPORTING THRESHOLD (THE DEVICE WOULD BE NOT LIKELY TO CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY IF THE ISSUE WERE TO RECUR). AS THE CUSTOMER SUBMITTED AN INCIDENT INVOLVING THE DEATH OF A PATIENT, AGFA HEALTHCARE DECIDED TO REPORT THIS INCIDENT REGARDLESS. ROOT CAUSE OF 5 MINUTE DELAY: THE AGFA PERFORMED TECHNICAL INVESTIGATION DISCOVERED WHEN REVIEWING THE WORKFLOW MANAGER (WFM) SERVER LOGS AT THE SITE THAT TICKET EXPIRY ERRORS APPEARED. IT WAS CONFIRMED THAT THE WFM SERVER WAS BRIEFLY OUT OF SYNC, AS THE TIME WAS AHEAD BY APPROXIMATELY 80 SECONDS. THE TECHNICAL INVESTIGATION REVEALED THAT THE BRIEF DELAY IN IMAGE VIEWING OF 4 -5 MINUTES WAS CAUSED BY THE RADIOLOGIST'S DECISION TO RESTART THE HOME DIAGNOSTIC WORKSTATION WHEN THE IMAGES WERE NOT IMMEDIATELY AVAILABLE, DUE TO THE CONFIGURATION OF THE NTP SERVER AT TIME OF GO-LIVE ON (B)(6) 2025. NTP SERVER PROVIDES ACCURATE TIME INFORMATION TO OTHER DEVICES ON A NETWORK, ENSURING THAT ALL SYSTEMS MAINTAIN TIME SYNCHRONIZATION. THIS ISSUE IS NOT THE RESULT OF AN AGFA PRODUCT DEFECT. THIS TYPE OF SYNCHRONIZATION ISSUE IN THE WORKFLOW MANAGER IS RELATED TO A MISSED SERVER CONFIGURATION AT THIS PARTICULAR SITE AND NOT RELATED TO SOFTWARE/PRODUCT. CORRECTIVE ACTION: A CONFIGURATION CHANGE WAS COMPLETED ON (B)(6) 2025, TO ADDRESS THE TIME SYNC ISSUE. THE NTP SERVER WAS ADDED INTO THE ESX SERVER CONFIGURATION AT THE SITE AND THE NTPD SERVICE. Manufacturer Narrative: THE CUSTOMER STATED THAT THE 40-YEAR-OLD FEMALE PATIENT WAS FOUND IN CARDIOPULMONARY ARREST AT 5:00AM AT HOME AND WAS TRANSPORTED TO THE HOSPITAL VIA AMBULANCE, ARRIVING AT THE EMERGENCY DEPARTMENT (ED) AT 5:50AM, ADMITTED IN CRITICAL, PRIORITY 1 STATUS. A CARDIAC ULTRASOUND AND ECG WERE PERFORMED IN THE ED UPON ARRIVAL. THE PATIENT'S HISTORY PROVIDED BY THE CUSTOMER IS: POORLY CONTROLLED ASTHMA/COPD, BRONCHOSPASM, PNEUMONIA, SMOKING AND OPIOID DEPENDENCE. CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED, AND THE PATIENT'S CARDIAC RHYTHM RETURNED TO SINUS RHYTHM AT 6:09AM. THE PATIENT WAS INTUBATED AT 6:30AM. IN ADDITION TO THE ENDOTRACHEAL INTUBATION (ETI) AND CPR, THE CUSTOMER HAS CONFIRMED THAT THE FOLLOWING MEDICATIONS WERE ADMINISTERED PRIOR TO TRANSPORT TO THE RADIOLOGY DEPARTMENT: INTRAVENOUS (IV) EPINEPHRINE X2, IV FENTANYL, IV VERSED, IV KETAMINE, IV VENTOLIN, IV SOLUMEDROL, IV VANCOMYCIN, AND VENTOLIN (ROUTE OF ADMINISTRATION WAS NOT SHARED HOWEVER FOR MECHANICALLY VENTILATED PATIENTS, BASED ON CLINICAL EXPERIENCE, THE USUAL ROUTE OF ADMINISTRATION IS VIA NEBULIZER OR METERED DOSE INHALER (MDI)). THE PATIENT ARRIVED AT THE RADIOLOGY DEPARTMENT AT 7:40AM. THE CUSTOMER CONFIRMED THE PATIENT HAD BEEN STABILIZED IN THE ED BETWEEN 5:50AM AND 7:40AM, PRIOR TO TRANSPORT TO RADIOLOGY. A NON-CONTRAST BRAIN CT STUDY WAS ACQUIRED FOLLOWED BY CONTRAST-ENHANCED THORACO-ABDOMINOPELVIC CT ANGIOGRAPHY (CTA) WERE ACQUIRED AT 7:58AM WHILE IN THE RADIOLOGY DEPARTMENT. THE CLINICAL INFORMATION/REASON FOR STUDY WAS DOCUMENTED AS "CARDIO-RESPIRATORY ARREST. CAUSE? FIXED PUPILS. DISTENDED ABDOMEN. NORMAL HEART ULTRASOUND." THE CUSTOMER CONFIRMED THE PATIENT WAS ALIVE FOR BOTH STUDIES AND RETURNED TO THE ED AND VITAL SIGNS WERE RECORDED. PER THE LOGS COLLECTED BY AGFA SUPPORT, THE RADIOLOGIST, DR (B)(6), READING FROM A REMOTE HOME STATION, INITIALLY ATTEMPTED TO VIEW THE IMAGES AT 8:03AM, INSTEAD A "PROBLEM DOWNLOADING IMAGE, WILL ATTEMPT TO RETRY" ERROR MESSAGE DISPLAYED TO THE USER. THE IMAGES WERE CONFIRMED TO BE ACCESSIBLE ON THE MODALITY, HOWEVER, THE RADIOLOGIST DID NOT PROCEED WITH ALTERNATIVE DIAGNOSTIC STEPS AS HE WAS FOCUSED ON DISPLAYING THE IMAGES IN THE IMPAX CLIENT ON HIS HOME DIAGNOSTIC CONSOLE. THE RADIOLOGIST USER DECIDED TO RESTART THE HOME STATION WHICH TOOK APPROXIMATELY 4-5 MINUTES. THE LOGS CONFIRMED THAT THE IMAGES WERE SUCCESSFULLY VIEWED BY THE RADIOLOGIST AT 8:08AM, AND THE DICTATION WORKFLOW COMPLETED. THE PATIENT WAS TRANSFERRED TO THE ICU AT 8:15AM, AND A CHEST XRAY ACQUIRED AT 11:15AM. FOLLOWING A DISCUSSION WITH THE FAMILY, AS THE PATIENT WAS CONFIRMED TO BE BRAIN DEAD, THE PATIENT WAS PLACED IN COMFORT CARE AND EXTUBATED AT 9:15PM. THE PATIENT WAS PRONOUNCED DECEASED IN THE ICU WITH A TIME OF DEATH OF 10:28PM ON (B)(6) 2025. IT WAS REPORTED BY THE CUSTOMER THAT THE INABILITY TO VIEW THE IMAGES DURING THE EMERGENCY SITUATION CAUSED THEM "TO LOSE VALUABLE MINUTES IN MAKING THE DIAGNOSIS. THE CAUSE OF DEATH WAS CONFIRMED TO BE REFRACTORY BRONCHOSPASM DUE TO BILATERAL PNEUMONIA. SUMMARY: TAKING ALL FACTORS INTO CONSIDERATION, INCLUDING PATIENT HISTORY, CLINICAL PRESENTATION UPON ED ARRIVAL WITH FIXED PUPILS, DIAGNOSTIC IMAGING FINDINGS, MEDICAL MANAGEMENT PRE AND POST DIAGNOSTIC IMAGING ACQUISITION, AND CAUSE OF DEATH; BASED ON CLINICAL EXPERIENCE, AGFA HEALTHCARE BELIEVES THAT THE 5 MINUTE DELAY IN IMAGE VIEWING BY THE READING RADIOLOGIST DID NOT CAUSE OR CONTRIBUTE TO THE PATIENT'S DEATH. ALSO, AGFA'S RISK ANALYSIS INDICATES THAT THE POTENTIAL FOR HARM FALLS BELOW THE REPORTING THRESHOLD (THE DEVICE WOULD BE NOT LIKELY TO CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY IF THE ISSUE WERE TO RECUR). AS THE CUSTOMER SUBMITTED AN INCIDENT INVOLVING THE DEATH OF A PATIENT, AGFA HEALTHCARE DECIDED TO REPORT THIS INCIDENT REGARDLESS. SOFTWARE VERSION: 6.7.1 SU02; UDI-DI: (B)(4); BASIC UDI: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21731184&pc=DSQ | 3007042319-2025-00705 | 2025-02-11 05:00 | Death | HEARTWARE, INC. | 2025-03-31 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT DEATH OCCURRED, WHICH WAS NOT CORRELATED TO THE USE OF THE VENTRICULAR ASSIST DEVICE. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: THE VENTRICULAR ASSIST DEVICE (VAD) (B)(6) WAS NOT RETURNED FOR EVALUATION AS IT REMAINS IN USE. THE COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. REVIEW OF THE MANUFACTURING DOCUMENTATION CONFIRMED THAT THE ASSOCIATED DEVICE MET ALL REQUIREMENTS FOR RELEASE. LOG FILE ANALYSIS WAS NOT PERFORMED SINCE LOG FILES WERE NOT AVAILABLE FOR ANALYSIS. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH IS UNKNOWN. BASED ON THE LIMITED INFORMATION AVAILABLE, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, DEATH IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. BASED ON REVIEW OF PAST ADVERSE EVENTS FOR THIS PATIENT, IT WAS NOTED THAT THE PATIENT HAD A HISTORY OF THROMBUS. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21731205&pc=LLZ | 9616389-2025-00001 | 2025-03-12 04:00 | Death | AGFA HEALTHCARE N.V. | 2025-03-31 04:00 | LLZ | IMPAX | Installation-Related Problem; Insufficient Information | Cardiac Arrest; Pneumonia | K022292 |  | 1 | Event Description: ON (B)(6) 2025, A CUSTOMER REPORTED THE FOLLOWING INCIDENT TO AGFA HEALTHCARE THROUGH COMPLAINT (B)(4): A 40-YEAR-OLD FEMALE PATIENT ARRIVED IN THE EMERGENCY IN CARDIAC AND RESPIRATORY ARREST. WHEN THE RADIOLOGIST, READING FROM A REMOTE HOME STATION, ATTEMPTED TO VIEW THE IMAGES, AN ERROR MESSAGE 'PROBLEM DOWNLOADING IMAGE, WILL ATTEMPT TO RETRY" DISPLAYED AND COULD NOT VIEW THE CT AND CTA IMAGES. THE RADIOLOGIST USER RESTARTED THE WORKSTATION WHICH TOOK APPROXIMATELY 4-5 MINUTES AFTER WHICH THE IMAGES COULD BE SUCCESSFULLY VIEWED. THE CUSTOMER REPORTED THAT THE INABILITY TO VIEW THE IMAGES DURING THE EMERGENCY SITUATION CAUSED THEM TO LOSE VALUABLE MINUTES IN MAKING THE DIAGNOSIS. THE VIEWING ISSUE HAPPENED AROUND 8 AM AND THE PATIENT PASSED AWAY AT 10:25 PM LATER THAT DAY. SUMMARY: TAKING ALL FACTORS INTO CONSIDERATION, INCLUDING PATIENT HISTORY, CLINICAL PRESENTATION UPON ED ARRIVAL WITH FIXED PUPILS, DIAGNOSTIC IMAGING FINDINGS, MEDICAL MANAGEMENT PRE AND POST DIAGNOSTIC IMAGING ACQUISITION, AND CAUSE OF DEATH; BASED ON CLINICAL EXPERIENCE, AGFA HEALTHCARE BELIEVES THAT THE 5 MINUTE DELAY IN IMAGE VIEWING BY THE READING RADIOLOGIST DID NOT CAUSE OR CONTRIBUTE TO THE PATIENT'S DEATH. ALSO, AGFA'S RISK ANALYSIS INDICATES THAT THE POTENTIAL FOR HARM FALLS BELOW THE REPORTING THRESHOLD (THE DEVICE WOULD BE NOT LIKELY TO CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY IF THE ISSUE WERE TO RECUR). AS THE CUSTOMER SUBMITTED AN INCIDENT INVOLVING THE DEATH OF A PATIENT, AGFA HEALTHCARE DECIDED TO REPORT THIS INCIDENT REGARDLESS.Event Description: ON MARCH 12, 2025, A CUSTOMER REPORTED THE FOLLOWING INCIDENT TO (B)(6) HEALTHCARE THROUGH COMPLAINT (B)(4): A 40-YEAR-OLD FEMALE PATIENT ARRIVED IN THE EMERGENCY IN CARDIAC AND RESPIRATORY ARREST. WHEN THE RADIOLOGIST, READING FROM A REMOTE HOME STATION, ATTEMPTED TO VIEW THE IMAGES, AN ERROR MESSAGE 'PROBLEM DOWNLOADING IMAGE, WILL ATTEMPT TO RETRY" DISPLAYED AND COULD NOT VIEW THE CT AND CTA IMAGES. THE RADIOLOGIST USER RESTARTED THE WORKSTATION WHICH TOOK APPROXIMATELY 4-5 MINUTES AFTER WHICH THE IMAGES COULD BE SUCCESSFULLY VIEWED. THE CUSTOMER REPORTED THAT THE INABILITY TO VIEW THE IMAGES DURING THE EMERGENCY SITUATION CAUSED THEM TO LOSE VALUABLE MINUTES IN MAKING THE DIAGNOSIS. THE VIEWING ISSUE HAPPENED AROUND 8 AM AND THE PATIENT PASSED AWAY AT 10:25 PM LATER THAT DAY. SUMMARY: TAKING ALL FACTORS INTO CONSIDERATION, INCLUDING PATIENT HISTORY, CLINICAL PRESENTATION UPON ED ARRIVAL WITH FIXED PUPILS, DIAGNOSTIC IMAGING FINDINGS, MEDICAL MANAGEMENT PRE AND POST DIAGNOSTIC IMAGING ACQUISITION, AND CAUSE OF DEATH; BASED ON CLINICAL EXPERIENCE, (B)(6) HEALTHCARE BELIEVES THAT THE 5 MINUTE DELAY IN IMAGE VIEWING BY THE READING RADIOLOGIST DID NOT CAUSE OR CONTRIBUTE TO THE PATIENT'S DEATH. ALSO, (B)(6)'S RISK ANALYSIS INDICATES THAT THE POTENTIAL FOR HARM FALLS BELOW THE REPORTING THRESHOLD (THE DEVICE WOULD BE NOT LIKELY TO CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY IF THE ISSUE WERE TO RECUR). AS THE CUSTOMER SUBMITTED AN INCIDENT INVOLVING THE DEATH OF A PATIENT, (B)(6) HEALTHCARE DECIDED TO REPORT THIS INCIDENT REGARDLESS. ROOT CAUSE OF 5 MINUTE DELAY: THE (B)(6) PERFORMED TECHNICAL INVESTIGATION DISCOVERED WHEN REVIEWING THE WORKFLOW MANAGER (WFM) SERVER LOGS AT THE SITE THAT TICKET EXPIRY ERRORS APPEARED. IT WAS CONFIRMED THAT THE WFM SERVER WAS BRIEFLY OUT OF SYNC, AS THE TIME WAS AHEAD BY APPROXIMATELY 80 SECONDS. THE TECHNICAL INVESTIGATION REVEALED THAT THE BRIEF DELAY IN IMAGE VIEWING OF 4 -5 MINUTES WAS CAUSED BY THE RADIOLOGIST'S DECISION TO RESTART THE HOME DIAGNOSTIC WORKSTATION WHEN THE IMAGES WERE NOT IMMEDIATELY AVAILABLE, DUE TO THE CONFIGURATION OF THE NTP SERVER AT TIME OF GO-LIVE ON JANUARY 15, 2025. NTP SERVER PROVIDES ACCURATE TIME INFORMATION TO OTHER DEVICES ON A NETWORK, ENSURING THAT ALL SYSTEMS MAINTAIN TIME SYNCHRONIZATION. THIS ISSUE IS NOT THE RESULT OF AN (B)(6) PRODUCT DEFECT. THIS TYPE OF SYNCHRONIZATION ISSUE IN THE WORKFLOW MANAGER IS RELATED TO A MISSED SERVER CONFIGURATION AT THIS PARTICULAR SITE AND NOT RELATED TO SOFTWARE/PRODUCT. CORRECTIVE ACTION: A CONFIGURATION CHANGE WAS COMPLETED ON MARCH 18, 2025, TO ADDRESS THE TIME SYNC ISSUE. THE NTP SERVER WAS ADDED INTO THE ESX SERVER CONFIGURATION AT THE SITE AND THE NTPD SERVICE. Manufacturer Narrative: THE CUSTOMER STATED THAT THE 40-YEAR-OLD FEMALE PATIENT WAS FOUND IN CARDIOPULMONARY ARREST AT 5:00AM AT HOME AND WAS TRANSPORTED TO THE HOSPITAL VIA AMBULANCE, ARRIVING AT THE EMERGENCY DEPARTMENT (ED) AT 5:50AM, ADMITTED IN CRITICAL, PRIORITY 1 STATUS. A CARDIAC ULTRASOUND AND ECG WERE PERFORMED IN THE ED UPON ARRIVAL. THE PATIENT'S HISTORY PROVIDED BY THE CUSTOMER IS: POORLY CONTROLLED ASTHMA/COPD, BRONCHOSPASM, PNEUMONIA, SMOKING AND OPIOID DEPENDENCE. CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED, AND THE PATIENT'S CARDIAC RHYTHM RETURNED TO SINUS RHYTHM AT 6:09AM. THE PATIENT WAS INTUBATED AT 6:30AM. IN ADDITION TO THE ENDOTRACHEAL INTUBATION (ETI) AND CPR, THE CUSTOMER HAS CONFIRMED THAT THE FOLLOWING MEDICATIONS WERE ADMINISTERED PRIOR TO TRANSPORT TO THE RADIOLOGY DEPARTMENT: INTRAVENOUS (IV) EPINEPHRINE X2, IV FENTANYL, IV VERSED, IV KETAMINE, IV VENTOLIN, IV SOLUMEDROL, IV VANCOMYCIN, AND VENTOLIN (ROUTE OF ADMINISTRATION WAS NOT SHARED HOWEVER FOR MECHANICALLY VENTILATED PATIENTS, BASED ON CLINICAL EXPERIENCE, THE USUAL ROUTE OF ADMINISTRATION IS VIA NEBULIZER OR METERED DOSE INHALER (MDI)). THE PATIENT ARRIVED AT THE RADIOLOGY DEPARTMENT AT 7:40AM. THE CUSTOMER CONFIRMED THE PATIENT HAD BEEN STABILIZED IN THE ED BETWEEN 5:50AM AND 7:40AM, PRIOR TO TRANSPORT TO RADIOLOGY. A NON-CONTRAST BRAIN CT STUDY WAS ACQUIRED FOLLOWED BY CONTRAST-ENHANCED THORACO-ABDOMINOPELVIC CT ANGIOGRAPHY (CTA) WERE ACQUIRED AT 7:58AM WHILE IN THE RADIOLOGY DEPARTMENT. THE CLINICAL INFORMATION/REASON FOR STUDY WAS DOCUMENTED AS "CARDIO-RESPIRATORY ARREST. CAUSE? FIXED PUPILS. DISTENDED ABDOMEN. NORMAL HEART ULTRASOUND." THE CUSTOMER CONFIRMED THE PATIENT WAS ALIVE FOR BOTH STUDIES AND RETURNED TO THE ED AND VITAL SIGNS WERE RECORDED. PER THE LOGS COLLECTED BY AGFA SUPPORT, THE RADIOLOGIST, DR DE LA FONTAINE, READING FROM A REMOTE HOME STATION, INITIALLY ATTEMPTED TO VIEW THE IMAGES AT 8:03AM, INSTEAD A "PROBLEM DOWNLOADING IMAGE, WILL ATTEMPT TO RETRY" ERROR MESSAGE DISPLAYED TO THE USER. THE IMAGES WERE CONFIRMED TO BE ACCESSIBLE ON THE MODALITY, HOWEVER, THE RADIOLOGIST DID NOT PROCEED WITH ALTERNATIVE DIAGNOSTIC STEPS AS HE WAS FOCUSED ON DISPLAYING THE IMAGES IN THE IMPAX CLIENT ON HIS HOME DIAGNOSTIC CONSOLE. THE RADIOLOGIST USER DECIDED TO RESTART THE HOME STATION WHICH TOOK APPROXIMATELY 4-5 MINUTES. THE LOGS CONFIRMED THAT THE IMAGES WERE SUCCESSFULLY VIEWED BY THE RADIOLOGIST AT 8:08AM, AND THE DICTATION WORKFLOW COMPLETED. THE PATIENT WAS TRANSFERRED TO THE ICU AT 8:15AM, AND A CHEST XRAY ACQUIRED AT 11:15AM. FOLLOWING A DISCUSSION WITH THE FAMILY, AS THE PATIENT WAS CONFIRMED TO BE BRAIN DEAD, THE PATIENT WAS PLACED IN COMFORT CARE AND EXTUBATED AT 9:15PM. THE PATIENT WAS PRONOUNCED DECEASED IN THE ICU WITH A TIME OF DEATH OF 10:28PM ON (B)(6) 2025. IT WAS REPORTED BY THE CUSTOMER THAT THE INABILITY TO VIEW THE IMAGES DURING THE EMERGENCY SITUATION CAUSED THEM "TO LOSE VALUABLE MINUTES IN MAKING THE DIAGNOSIS. THE CAUSE OF DEATH WAS CONFIRMED TO BE REFRACTORY BRONCHOSPASM DUE TO BILATERAL PNEUMONIA. SUMMARY: TAKING ALL FACTORS INTO CONSIDERATION, INCLUDING PATIENT HISTORY, CLINICAL PRESENTATION UPON ED ARRIVAL WITH FIXED PUPILS, DIAGNOSTIC IMAGING FINDINGS, MEDICAL MANAGEMENT PRE AND POST DIAGNOSTIC IMAGING ACQUISITION, AND CAUSE OF DEATH; BASED ON CLINICAL EXPERIENCE, AGFA HEALTHCARE BELIEVES THAT THE 5 MINUTE DELAY IN IMAGE VIEWING BY THE READING RADIOLOGIST DID NOT CAUSE OR CONTRIBUTE TO THE PATIENT'S DEATH. ALSO, AGFA'S RISK ANALYSIS INDICATES THAT THE POTENTIAL FOR HARM FALLS BELOW THE REPORTING THRESHOLD (THE DEVICE WOULD BE NOT LIKELY TO CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY IF THE ISSUE WERE TO RECUR). AS THE CUSTOMER SUBMITTED AN INCIDENT INVOLVING THE DEATH OF A PATIENT, AGFA HEALTHCARE DECIDED TO REPORT THIS INCIDENT REGARDLESS. SOFTWARE VERSION: 6.7.1 SU02. UDI-DI: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21731666&pc=NPT | 2025587-2025-02233 | 2025-03-03 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-31 04:00 | NPT | EVOLUT FX DCS | Patient Device Interaction Problem | Vascular Dissection | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE IN A PATIENT WITH A WITH A 70-DEGREE HORIZONTAL AORTA AND AN ENLARGED ASCENDING AORTA, THE VALVE PASSED THROUGH THE GREATER CURVATURE WHILE ADVANCING THE DELIVERY CATHETER SYSTEM (DCS). WHEN PLACING THE VALVE, IT WAS SLIGHTLY TILTED TOWARDS THE GREATER CURVATURE. INTRAOPERATIVELY AND IMMEDIATELY POSTOPERATIVELY, THE VITAL SIGNS WERE STABLE, AND THE PATIENT HAD NO SYMPTOMS. AN ECHOCARDIOGRAM IN THE NEXT EVENING REVEALED A POSSIBLE DISSECTION IN THE ASCENDING AORTA, PROMPTING A CONTRAST COMPUTED TOMOGRAPHY (CT) SCAN. IT WAS DIAGNOSED AS AN TYPE A ASCENDING AORTA DISSECTION WITH AN ENTRY POINT JUST ABOVE THE VALVE OUTFLOW. PER THE PHYSICIAN, THE ANGLE BETWEEN THE DCS AND THE ASCENDING AORTA WAS POOR. IT IS BELIEVED THAT THE CAUSE WAS A VASCULAR WALL INJURY WAS THE DEVICE EDGE. OF NOTE, THE DISSECTION CAVITY DID NOT EXTEND TO THE BRACHIOCEPHALIC ARTERY, AND CEREBRAL BLOOD FLOW WAS UNAFFECTED. ON THE SECOND POSTOPERATIVE DAY, A CONTRAST CT SHOWED NO ENLARGEMENT OF THE DISSECTION CAVITY. CONTINUED BLOOD PRESSURE CONTROL AND REST WERE MAINTAINED AND THE PATIENT WAS MONITORED. ON THE 3RD DAY POSTOPERATIVE, THE DISSECTION PROGRESSED. THE PATIENT DIED ON THE 4TH DAY POSTOPERATIVE. EMERGENCY SURGERY WAS NOT PERFORMED PER THE PATIENT'S REQUEST. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID: L-EVOLUTFX-2329 (LOT: 0012449276); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: (B)(6) 2025. SECTION D REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: BRAND NAME: EVOLUT FX VALVE; PRODUCT ID: EVOLUTFX-26 (SERIAL: (B)(6)); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: (B)(6) 2025. SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21731772&pc=DSQ | 2916596-2025-02026 | 2025-03-12 04:00 | Death | THORATEC CORPORATION | 2025-03-31 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2025 DUE TO MULTISYSTEM ORGAN FAILURE (MSOF). THE DEVICE WAS NOTED TO HAVE WORKED AS INTENDED AND WAS NOT EXPLANTED OR RETURNED. THE DEATH WAS DETERMINED NOT TO BE DEVICE OR THERAPY RELATED. THE CUSTOMER STATED THAT NO ADDITIONAL FOLLOW-UP WOULD BE PROVIDED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6) , AND THE REPORTED EVENT COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE ACCOUNT INDICATED THAT NO ADDITIONAL INFORMATION WILL BE PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. CHAPTER 1, ¿INTRODUCTION¿, LISTS MULTIPLE ORGAN FAILURES/DYSFUNCTIONS AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21732125&pc=LWP | 2124215-2025-19611 | 2025-03-13 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-03-31 04:00 | LWP | ACCOLADE MRI DR | High Capture Threshold | Insufficient Information | P150012 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WITH THIS NEWLY IMPLANTED PACEMAKER DIED ON (B)(6) 2025, THE DAY OF THEIR IMPLANT PROCEDURE. THE PATIENT WAS ORIGINALLY BEING TREATED FOR BRADYCARDIA/JUNCTIONAL RHYTHM POST A TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE WHICH HAD OCCURRED BEFOREHAND. ACCORDING TO THE PHYSICIAN, THE PACEMAKER IMPLANT PROCEDURE WAS SUCCESSFUL, HOWEVER ON MULTIPLE OCCASIONS DURING SURGERY, THE PATIENT WAS REPORTED TO NOT HAVE HAD NO PULSE, EVEN WITH THE PACEMAKER PACING THE HEART. AFTER THE PROCEDURE WAS COMPLETED, THE PATIENT EXPERIENCED MULTIPLE ROUNDS OF VENTRICULAR TACHYCARDIA/VENTRICULAR FIBRILLATION WHICH REQUIRED EXTERNAL SHOCKS TO CONVERT. A RISE IN RV THRESHOLDS WAS ALSO NOTED DURING A TWO-HOUR PERIOD, HOWEVER NO ALLEGATION WAS MADE AGAINST THE IMPLANTED LEADS. THE PATIENT WAS GIVEN MEDICATIONS, INTUBATED, AN ECHOCARDIOGRAM WAS PERFORMED, AN ARTERIAL LINE MADE, AND A BLOOD TRANSFUSION PERFORMED, BUT THEY ENDED UP SUCCUMBING AND PASSING AWAY FROM THEIR COMPLICATIONS. IT WAS UNKNOWN IF THE PACEMAKER OR IMPLANT PROCEDURE CONTRIBUTED TO THE PATIENT'S DEATH AT THIS TIME. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21732188&pc=OZD | 1220648-2025-27105 | 2025-03-04 05:00 | Death | ABIOMED, INC. | 2025-03-31 04:00 | OZD | IMPELLA CP | Restricted Flow rate; Malposition of Device | Cardiac Arrest | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. THE PATIENT UNDERWENT PERCUTANEOUS CORONARY INTERVENTION: DRUG-ELUTING STENT TO THE LEFT ANTERIOR DESCENDING ARTERY (LAD) AND RIGHT CORONARY ARTERY (RCA). THE PATIENT BEGAN FEELING "BAD" THROUGHOUT THE NIGHT AND AN ELECTROCARDIOGRAM SHOWED ANTERIOR ST-SEGMENT ELEVATION THE FOLLOWING MORNING. THE PATIENT WAS TAKEN TO THE CATHETERIZATION LAB AND IN-STENT THROMBOSIS OF THE LAD WAS NOTED. UPON WIRING THE LAD THE PATIENT WENT INTO VENTRICULAR FIBRILLATION ARREST AND WAS SHOCKED FIVE TIMES. THE DECISION WAS MADE TO IMPLANT AN IMPELLA CP DEVICE, WHICH WAS PLACED VIA THE RIGHT FEMORAL ARTERY WITHOUT INCIDENT. THE PATIENT WAS INTUBATED. PENUMBRA AND PERCUTANEOUS TRANSLUMINAL ANGIOPLASTY RESTORED FLOW TO THE LAD. IT WAS REPORTED, HOWEVER, AFTER START OF THE IMPELLA MCS, THE PUMP BEGAN TO HAVE SUCTION. A FLUID BOLUS WAS GIVEN. THE PATIENT BECAME VERY ACIDOTIC AND MULTIPLE PRESSORS WERE ADDED. MULTIPLE BICARB PUSHES AND DRIP WERE STARTED. AN ECHOCARDIOGRAM WAS COMPLETED AND SHOWED THE IMPELLA CP PUMP POINTED TOWARD MITRAL INSTEAD OF TOWARD APEX. THE PHYSICIAN WAS UNABLE TO GET BETTER POSITIONING WITH THE DEVICE. THE PATIENT CONTINUED TO DECLINE AND ULTIMATELY EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21732344&pc=OZD | 1220648-2025-27106 | 2022-12-10 05:00 | Death | ABIOMED, INC. | 2025-03-31 04:00 | OZD | IMPELLA 5.5 | Increase in Pressure | Insufficient Information | 140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT IN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. THE PATIENT WAS THEN ESCALATED FROM THE IMPELLA CP TO AN IMPELLA 5.5 FOR INCREASED CIRCULATORY SUPPORT AFTER 4 DAYS OF SUPPORT ON THE IMPELLA CP. WHEN THE PATIENT WAS BROUGHT TO THE OPERATING ROOM (OR) TO REPLACE THE IMPELLA CP WITH IMPELLA 5.5 THE DECISION WAS MADE FOR THE PATIENT TO HAVE CORONARY ARTERY BYPASS (CABG). THE PATIENT WAS PLACED ON BYPASS, AND THE IMPELLA WAS PLACE DIRECTLY THROUGH THE AORTA AFTER AN AORTIC VALVE REPLACEMENT WAS PERFORMED. CABG WAS PERFORMED SUCCESSFULLY AND WITHOUT COMPLICATIONS. THE IMPELLA 5.5 PLACEMENT WAS CONFIRMED WITH TRANSESOPHAGEAL ECHOCARDIOGRAPHY (TEE), SUPPORT WAS INITIATED AND THE PATIENT WAS WEANED OFF BYPASS. TEE OF THE RIGHT VENTRICLE REVEALED WORSENING RIGHT HEART FUNCTION. THE PATIENT WAS GIVEN AN INCREASED INOTROPE SUPPORT AND TRANSPORTED BACK TO THE ICU. THE FOLLOWING DAY, ON (B)(6) 2022, THE PATIENT WAS TAKEN BACK TO THE OR FOR VENOARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION CANNULATION (VA ECMO) DUE TO THE FAILING RIGHT VENTRICLE AND ASSOCIATED DECOMPENSATION. DURING THIS TIME, IT WAS REPORTED THAT THE IMPELLA 5.5 HAD ALARMS FOR HIGH PURGE PRESSURE. THE MEDICAL TEAM DECIDED TO PLACE TISSUE PLASMINOGEN ACTIVATOR IN THE PURGE SOLUTION. THE HIGH PURGE PRESSURE ALARM WAS RESOLVED; HOWEVER, IT WAS NOTED THAT THE PRESSURE REMAINED HIGH, WITH THE MEAN MOTOR CURRENT WITHIN EXPECTED RANGES. THERE WERE NO FURTHER IMPELLA COMPLICATIONS AND THE PATIENT WAS TRANSPORTED BACK TO THE ICU. ON (B)(6) 2022, THE PATIENT'S RIGHT LUNG WAS DISCOVERED TO BE NECROTIC. THE SURGICAL TEAM MADE PLANS FOR RESECTION AT A LATER TIME, HOWEVER THE PATIENT CONTINUED TO DECLINE AND REQUIRED INCREASED SUPPORT DESPITE THE PATIENT BEING SUPPORTED BY IMPELLA 5.5, VA ECMO, AND HEMODIALYSIS. THE FAMILY MADE THE DECISION TO WITHDRAW CARE, AND THE PATIENT EXPIRED. ALTHOUGH THE PATIENT'S PRESENTING CONDITION MAY BE MORE LIKELY HAVE IMPACTED THE EVENT THE IMPELLA CANNOT BE EXCLUDED AS A POSSIBLE CONTRIBUTING FACTOR DUE TO THE HIGH PURGE PRESSURE WHICH IS A PRODUCT MALFUNCTION AND CANNOT EXCLUDE THIS FROM THE OUTCOME OF DEMISED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH-RISK PCI INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA CATHETER ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP. CLINICIANS SHOULD MONITOR MOTOR CURRENT AND CONSIDER REPLACING THE IMPELLA CATHETER WHENEVER A RISE IN MOTOR CURRENT IS SEEN.¿ THIS REPORT IS BEING FILED AS PART OF A RETROSPECTIVE REVIEW OF HISTORICAL RECORDS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21732384&pc=OZD | 1220648-2025-27107 | 2021-12-05 05:00 | Death | ABIOMED, INC. | 2025-03-31 04:00 | OZD | IMPELLA 5.5 | Restricted Flow rate; Malposition of Device | Bradycardia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT IN NON-ST ELEVATION MYOCARDIAL INFARCTION WITH HYPOTENSION, ACUTE SEVERE MITRAL REGURGITATION, AND LOW EJECTION FRACTION WAS IMPLANTED WITH AN IMPELLA 5.5 DEVICE FOR MECHANICAL CIRCULATORY SUPPORT TO STABILIZE AND UNLOAD THE PATIENT'S LEFT HEART. DURING DELIVERY OF THE IMPELLA INTO THE LEFT VENTRICLE, THE IMPELLA 5.5 WAS PLACED DIRECTLY INFO THE PAPILLARY WALL AND WAS UNABLE TO BE MOVED. THE MEDICAL TEAM WAS ADVISED TO PULL THE GUIDEWIRE, AND THEN TURN THE IMPELLA AS THERE WAS SUSPICION THE GUIDEWIRE WAS IN THE CHORDAE. THE PATIENT THEN WENT INTO BRADYCARDIA AND LOST BLOOD PRESSURES. THE IMPELLA 5.5 WAS SUCCESSFULLY TURNED TOWARDS THE APEX AFTER REMOVAL OF THE GUIDEWIRE AND SUPPORT WAS THEN INITIATED. IT WAS NOTED BY THE PHYSICIAN THEY BELIEVE THAT THE PATIENT'S PRESSURES DROPPED DUE TO THE RIGHT CORONARY ARTERY BEING COMPLETELY BLOCKED AND HAVING NO PERFUSION. DURING SUPPORT, THERE WERE SUCTION ALARMS. THE POWER LEVEL OF THE IMPELLA 5.5 WAS DROPPED FROM P9 TO P8 AND THE PATIENT WAS GIVEN FLUID VOLUME TO RESOLVE THE ALARMS. THE PATIENT REMAINED ON SUPPORT FOR 5 DAYS, AND IT WAS DETERMINED THE PATIENT WAS UNLIKELY TO SURVIVE ANY FURTHER REVASCULARIZATION AND THE DECISION WAS MADE TO WITHDRAW CARE. CARE WAS WITHDRAWN, AND THE PATIENT EXPIRED. ALTHOUGH THE PATIENT'S PRESENTING CONDITION MAY BE MORE LIKELY HAVE IMPACTED THE EVENT THE IMPELLA CANNOT BE EXCLUDED AS A POSSIBLE CONTRIBUTING FACTOR IN THE PATIENT'S DEMISE. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH-RISK PCI INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): "ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)." THIS REPORT IS BEING FILED AS PART OF A RETROSPECTIVE REVIEW OF HISTORICAL RECORDS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21732571&pc=DSQ | 2916596-2025-01921 | 2025-03-01 05:00 | Death | THORATEC CORPORATION | 2025-03-31 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY FROM A SPONTANEOUS INTRACRANIAL BLEED (ICB) WHICH WAS RELATED TO THE OUTCOME AT LEAST. NO ANTICOAGULATION, NO DEVICE RELATED, OR OTHER THERAPEUTIC ISSUES WERE MENTIONED PRIOR OUTCOME. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF STROKE COULD NOT BE CONCLUSIVELY DETERMINED. STROKE IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. STROKE HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21732710&pc=OZD | 1220648-2025-27109 | 2025-03-04 05:00 | Death | ABIOMED, INC. | 2025-03-31 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT WITH CARDIOMYOPATHY WAS IMPLANTED WITH AN IMPELLA 5.5 DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. IT WAS REPORTED THAT THE PATIENT DEVELOPED ACCESS SITE BLEEDING AND ANTICOAGULATION WAS HELD WHILE THE SUPPORT CONTINUED. THE DECISION WAS MADE TO WITHDRAW CARE AND THE PATIENT EXPIRED. IT IS UNKNOWN IF THE USE OF THE IMPELLA WAS RELATED TO PATIENT DEATH, THUS THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA PUMP WAS NOT RETURNED FOR INVESTIGATION. SHOULD NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21732831&pc=DSQ | 2916596-2025-01868 | 2025-02-16 05:00 | Death | THORATEC CORPORATION | 2025-03-31 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTIPLE SYSTEM ORGAN FAILURE (MSOF). THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT'S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE DEVICE WAS NOT RETURNED FOR ANALYSIS. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION AND DEATH AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21732857&pc=LGW | 3006630150-2025-02004 | 2025-02-22 05:00 | Death | BOSTON SCIENTIFIC NEUROMODULATION CORPORATION | 2025-03-31 04:00 | LGW | WAVEWRITER ALPHA | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030017 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY, THE CAUSE OF DEATH IS UNKNOWN, AND NO ADDITIONAL INFORMATION WAS PROVIDED DESPITE MULTIPLE GOOD FAITH EFFORTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21733410&pc=NIK | 3004209178-2025-05559 | 2025-03-03 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-03-31 04:00 | NIK | VIVA¿ QUAD XT CRT-D | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P010031 |  | 1 | Event Description: IT WAS REPORTED THAT THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) BECAME INFECTED. THE CRT-D SYSTEM WAS EXPLANTED AND REPLACED WITH A LEADLESS PACEMAKER. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS REPORTED THAT THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) BECAME INFECTED. THE CRT-D SYSTEM WAS EXPLANTED AND REPLACED WITH A LEADLESS PACEMAKER. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT IT WAS FURTHER REPORTED THAT THE PATIENT WAS DECEASED. THE DEATH OCCURRED APPROXIMATELY TWO WEEKS POST EXPLANT OF THE CRT-D SYSTEM. THE DEATH WAS NOTED TO BE A RESULT OF COMPLEX MEDICAL CONDITIONS AND THE EFFECTS OF A PROLONGED HOSPITALIZATION. Manufacturer Narrative: CONTINUATION OF D10: 439888 LEAD IMPLANTED: (B)(6) 2017 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21733412&pc=OZD | 1220648-2025-27120 | 2025-02-28 05:00 | Death | ABIOMED, INC. | 2025-03-31 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Perforation | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT THE PATIENT DEVELOPED CARDIAC TAMPONADE DUE TO ACUTE MYOCARDIAL INFARCTION, AND IMPELLA CP WAS INSERTED TO STABILIZE HEMODYNAMICS. AFTER PERCUTANEOUS CORONARY INTERVENTION, IT WAS DIFFICULT TO CONTROL THE BLEEDING DUE TO CARDIAC TAMPONADE. THE PATIENT WAS MOVED TO THE OPERATING ROOM AND THORACOTOMY WAS PERFORMED TO STOP THE BLEEDING. WHEN THE THORACOTOMY WAS PERFORMED, IT WAS VISUALLY CONFIRMED THAT THE IMPELLER HAD PERFORATED THE VENTRICLE. BLEEDING WAS NOT ABLE TO BE CONTROLLED, AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM AND IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿. SECTION: WARNINGS & CAUTIONS: WARNINGS. SECTION: PRE-SUPPORT EVALUATION. SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER. SECTION: DIRECT AORTIC INSERTION. SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES. ¿TO REDUCE THE RISK OF CARDIAC INJURY (INCLUDING VENTRICULAR PERFORATION), PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY IN THE SETTINGS OF ACUTE INFARCTION WITH TISSUE NECROSIS.¿ ¿TO REDUCE THE RISK OF VASCULAR INJURY, PHYSICIANS SHOULD EXERCISE CAUTION WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX PERIPHERAL VASCULAR ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: UNREPAIRED ABDOMINAL AORTIC ANEURYSM, SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM, DISSECTION OF THE ASCENDING/ TRANSVERSE/DESCENDING AORTA, CHRONIC ANATOMICAL CHANGES IN THE RELATIONSHIP OF THE AORTA/AORTIC VALVE/VENTRICULAR ALIGNMENT, SIGNIFICANT MOBILE ATHEROMATOUS DISEASE IN THE THORACIC OR ABDOMINAL AORTA OR PERIPHERAL VESSELS.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING A RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY IN THE LEFT VENTRICLE AFTER CPR WITH ECHOCARDIOGRAPHY GUIDANCE. SECTION: WARNINGS & CAUTIONS: CAUTIONS. ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ Manufacturer Narrative: B3 DATE OF EVENT WAS ERRONEOUSLY ENTERED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27120. Manufacturer Narrative: THE INVESTIGATION OF VENTRICLE PERFORATION HAS BEEN COMPLETED. THE PUMP WAS NOT RETURNED FOR INVESTIGATION. A DATA LOG REVIEW WAS NOT REQUIRED. ACCORDING TO THE CLINICAL REPORT, THE PATIENT WAS REPORTED TO HAVE A FRAGILE AREA OF THE MYOCARDIUM. THE DOCTOR CONFIRMED, BASED ON VISUAL INSPECTION, THAT THE IMPELLA HAD PERFORATED THE VENTRICLE. THE CAUSE OF THE VENTRICLE PERFORATION ISSUE WAS MOST LIKELY PATIENT CONDITION RELATED TO POOR HEART CONDITION, BASED ON PROVIDED CLINICAL DETAILS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21733515&pc=LWS | 2649622-2025-08699 | 2025-03-03 05:00 | Death | MPRI | 2025-03-31 04:00 | LWS | SPRINT QUATTRO SECURE MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) BECAME INFECTED. THE CRT-D SYSTEM WAS EXPLANTED AND REPLACED WITH A LEADLESS PACEMAKER. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS REPORTED THAT THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) BECAME INFECTED. THE CRT-D SYSTEM WAS EXPLANTED AND REPLACED WITH A LEADLESS PACEMAKER. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. IT WAS FURTHER REPORTED THAT THE PATIENT WAS DECEASED. THE DEATH OCCURRED APPROXIMATELY TWO WEEKS POST EXPLANT OF THE CRT-D SYSTEM. THE DEATH WAS NOTED TO BE A RESULT OF COMPLEX MEDICAL CONDITIONS AND THE EFFECTS OF A PROLONGED HOSPITALIZATION. Manufacturer Narrative: CONTINUATION OF D10: 407652 LEAD IMPLANTED: (B)(6) 2009. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21733531&pc=DTB | 2649622-2025-08700 | 2025-03-03 05:00 | Death | MPRI | 2025-03-31 04:00 | DTB | CAPSUREFIX NOVUS MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) BECAME INFECTED. THE CRT-D SYSTEM WAS EXPLANTED AND REPLACED WITH A LEADLESS PACEMAKER. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS REPORTED THAT THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) BECAME INFECTED. THE CRT-D SYSTEM WAS EXPLANTED AND REPLACED WITH A LEADLESS PACEMAKER. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. IT WAS FURTHER REPORTED THAT THE PATIENT WAS DECEASED. THE DEATH OCCURRED APPROXIMATELY TWO WEEKS POST EXPLANT OF THE CRT-D SYSTEM. THE DEATH WAS NOTED TO BE A RESULT OF COMPLEX MEDICAL CONDITIONS AND THE EFFECTS OF A PROLONGED HOSPITALIZATION. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: 694765 LEAD IMPLANTED: (B)(6) 2009. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21733570&pc=OJX | 2649622-2025-08702 | 2025-03-03 05:00 | Death | MPRI | 2025-03-31 04:00 | OJX | ATTAIN PERFORMA STRAIGHT MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P080006 |  | 1 | Event Description: IT WAS REPORTED THAT THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) BECAME INFECTED. THE CRT-D SYSTEM WAS EXPLANTED AND REPLACED WITH A LEADLESS PACEMAKER. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS REPORTED THAT THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) BECAME INFECTED. THE CRT-D SYSTEM WAS EXPLANTED AND REPLACED WITH A LEADLESS PACEMAKER. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. IT WAS FURTHER REPORTED THAT THE PATIENT WAS DECEASED. THE DEATH OCCURRED APPROXIMATELY TWO WEEKS POST EXPLANT OF THE CRT-D SYSTEM. THE DEATH WAS NOTED TO BE A RESULT OF COMPLEX MEDICAL CONDITIONS AND THE EFFECTS OF A PROLONGED HOSPITALIZATION. Manufacturer Narrative: CONTINUATION OF D10: DTBA1Q1 CRT-D IMPLANTED: (B)(6) 2017. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21733601&pc=MIH | 3013164176-2025-02458 | 2025-03-09 05:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-03-31 04:00 | MIH | GORE® EXCLUDER®AAA ENDOPROSTHESIS | Insufficient Information | Insufficient Information | P020004 |  | 1 | Event Description: THE FIELD SALES ASSOCIATE (FSA) REPORTED THE FOLLOWING TO GORE: ON (B)(6) 2025, THE PATIENT UNDERWENT AN ENDOVASCULAR PROCEDURE TO TREAT AN ANEURYSM RUPTURE UTILIZING THE GORE® EXCLUDER® CONFORMABLE AAA IN ZONE 9. IT WAS REPORTED THAT THE PATIENT PRESENTED SYMPTOMATIC OF A RUPTURED ANEURYSM. HIS LEFT RENAL WAS INTENTIONALLY COVERED BECAUSE THE ANEURYSM EXTENDED TO THE LEFT RENAL. ALSO, TWO CONTRALATERAL LEGS WERE PLACED IN THE PATIENT. A FINAL ANGIOGRAM WAS PERFORMED THAT SHOWED NO EVIDENCE OF ENDOLEAK AND THE CONTAINMENT OF THE ANEURYSM SAC. THE RIGHT RENAL WAS FUNCTIONAL WITH NO ISSUES. THE PATIENT TOLERATED THE PROCEDURE. ON (B)(6) 2025, THE PATIENT'S RENAL FUNCTION DECLINED. THE PHYSICIAN PERFORMED AN OPEN RENAL REVASCULARIZATION WITH ENDOSCOPIC INTERVENTION. THE PHYSICIAN ALSO PERFORMED A LAPAROTOMY. AT SOME POINT DURING THESE PROCEDURES, AN AORTIC EXTENDER AND NON-GORE BARE METAL STENT WAS PLACED IN THE PATIENT. THE PATIENT TOLERATED THE PROCEDURE. ON (B)(6), THE FSA WAS IN COMMUNICATION WITH THE PHYSICIAN. THE PHYSICIAN STATED THE PATIENT EXPIRED. THE EXACT EXPIRATION DATE IS UNKNOWN. NO ADDITIONAL INFORMATION IS AVAILABLE. THE CAUSE OF THE REPORTED EVENT IS UNKNOWN. Manufacturer Narrative: ADDED G3/G4, H1/H2. CORRECTION: IT SHOULD BE NOTED THAT, PER THE GORE® EXCLUDER® AAA ENDOPROSTHESIS INSTRUCTIONS FOR USE (IFU), ADVERSE EVENTS THAT MAY OCCUR AND / OR REQUIRE INTERVENTION INCLUDE BUT ARE NOT LIMITED TO DEATH. Manufacturer Narrative: ADDED: G3/G4, H1/H2. CORRECTED: B3. Manufacturer Narrative: AS IT IS UNKNOWN WHICH DEVICE(S) IS DIRECTLY IMPLICATED IN THIS DEATH EVENT, THE FOLLOWING DEVICE(S) WILL ALSO BE INCLUDED IN THIS REPORT: CATALOG #PLC181000/ SERIAL #(B)(6)/ UDI #(B)(4), CATALOG #PLC141200/ SERIAL #(B)(6)/ UDI #(B)(4). H6: CODE A26- USED TO CAPTURE INSUFFICIENT INFORMATION AVAILABLE TO DETERMINE IF GORE DEVICE CONTRIBUTED TO DEATH. H6: CODE E2401- USED TO CAPTURE INSUFFICIENT INFORMATION AVAILABLE TO DETERMINE IF ANY HEALTH EFFECT APPEARS TO HAVE OCCURRED. IT SHOULD BE NOTED THAT, PER THE GORE® EXCLUDER® CONFORMABLE AAA ENDOPROSTHESIS INSTRUCTIONS FOR USE (IFU), ADVERSE EVENTS THAT MAY OCCUR AND / OR REQUIRE INTERVENTION INCLUDE BUT ARE NOT LIMITED TO DEATH. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY HAVE NOT BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21733710&pc=DSQ | 2916596-2025-01964 | 2024-11-19 05:00 | Death | THORATEC CORPORATION | 2025-03-31 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Skin Infection | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD A LONG HISTORY OF SERRATIA DRIVELINE INFECTIONS AND WAS PREVIOUSLY ADMITTED ON (B)(6) 2024 WITH A COMPLEX VIRIDANS STREPTOCOCCI, SERRATIA, AND DIPHTHEROID DRIVELINE INFECTION, WHICH WAS TREATED WITH ANTIBIOTICS AND A WASHOUT, THEN WAS DISCHARGED A FEW DAYS LATER (MFR# 2916596-2024-07607). ON (B)(6) 2024, THE DRIVELINE WAS RELOCATED WITH AN OMENTUM FLAP. THE PATIENT WAS RE-HOSPITALIZED FROM (B)(6) 2024 WITH A THORACOTOMY INFECTION, WHERE THE PUMP WAS VISIBLE THROUGH THE WOUND DUE TO SERRATIA. THE PATIENT WAS TREATED WITH ONGOING KEFLEX (CEFALEXIN). THE INFECTION HEALED WELL, AND THE PATIENT WAS ONGOING WITH INTRAVENOUS FLUIDS MEROPENEM (MERREM) THROUGH (B)(6) 2025. THE PATIENT APPEARED TO HAVE BEEN HEALING WELL WITH A WOUND VAC (VACUUM-ASSISTED CLOSURE) BUT BECAME RE-INFECTED. THE PATIENT WAS TRANSITIONED TO IN-HOME PALLIATIVE CARE AND THEN PASSED AWAY AT HOME ON (B)(6) 2025. Manufacturer Narrative: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF INFECTION COULD NOT BE CONCLUSIVELY DETERMINED. INFECTION ARE LISTED IN THE INSTRUCTIONS FOR USE AS POTENTIAL ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. INFECTION HAVE BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21733984&pc=DSQ | 2916596-2025-01850 | 2025-03-09 05:00 | Death | THORATEC CORPORATION | 2025-03-31 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Multiple Organ Failure; No Clinical Signs, Symptoms or Conditions | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTI-SYSTEM ORGAN FAILURE. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED.Event Description: IT WAS REPORTED THAT THE SOURCE OF THE MULTIORGAN FAILURE WAS RIGHT VENTRICULAR FAILURE WHICH WAS PRESENT BEFORE THE IMPLANT AND THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) THERAPY DID NOT WORSEN THE MULTISYSTEM ORGAN FAILURE (MSOF). Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: NO DEVICE RELATED ISSUES WERE REPORTED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. ALTHOUGH NO DEVICE RELATED ISSUES WERE REPORTED BY THE ACCOUNT, THE IFU LISTS ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: E1: REPORTER ESTABLISHMENT NAME: (B)(6). NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21734123&pc=DSQ | 2916596-2025-01875 | 2025-03-08 05:00 | Death | THORATEC CORPORATION | 2025-03-31 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. D IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21734186&pc=OZD | 1220648-2025-27125 | 2022-03-15 04:00 | Death | ABIOMED, INC. | 2025-03-31 04:00 | OZD | IMPELLA 5.5 | Failure to Advance | Insufficient Information | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT, WHO WAS PREVIOUSLY IMPLANTED WITH AN IMPELLA CP WAS BEING WORKED UP FOR A DURABLE LEFT VENTRICULAR ASSIST DEVICE WHEN THEY STARTED DECOMPENSATING. THE PATIENT WENT INTO ATRIAL FIBRILLATION. AND QUICKLY INCREASED DECOMPENSATING AND WAS ADMINISTERED AMIODARONE WITH NITROPRUSSIDE AND MILRINONE. THE PATIENT DID NOT RESPOND TO THE TREATMENT AND THE DECISION WAS MADE TO PLACE AN IMPELLA 5.5 FOR MECHANICAL CIRCULATORY SUPPORT. DURING IMPELLA 5.5 DELIVERY, THE PHYSICIAN WAS UNABLE TO ADVANCE THE IMPELLA FROM THE AXILLARY ARTERY INTO THE AORTA. THE IMPELLA WAS REMOVED AND WASHED IN A STERILE CANISTER. AN AMPLASE WIRE PLACED AND THE IMPELLA IMPLANTED OVER THE IMPELLA WIRE. WITH NOTED DIFFICULTY GETTING THE IMPELLA AROUND THE BEND IN THE ARTERY, THE IMPELLA WAS EVENTUALLY SUCCESSFULLY DELIVERED TO THE LEFT VENTRICLE. SUPPORT WAS INITIATED AND THE IMPELLA WAS SET TO POWER LEVEL P9. THE PATIENT REMAINED ON IMPELLA SUPPORT FOR 6 DAYS. AFTER FURTHER DECLINING IN CONDITION, THE FAMILY DECIDED TO WITHDRAW CARE. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. ALTHOUGH THE PATIENT'S PRESENTING CONDITION MAY BE MORE LIKELY HAVE IMPACTED THE EVENT THE IMPELLA CANNOT BE EXCLUDED AS A POSSIBLE CONTRIBUTING FACTOR IN THE PATIENT'S DEMISE. Manufacturer Narrative: THE INVESTIGATION FOR THE REPORTED DELIVERY ISSUE HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. HOWEVER, CLINICAL DETAILS PROVIDED WERE SUFFICIENT TO DETERMINE A ROOT CAUSE. THE ROOT CAUSE OF THE DELIVERY ISSUE WAS MOST LIKELY DUE TO PATIENT CONDITION. IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH-RISK PCI INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ THIS REPORT IS BEING FILED AS PART OF A RETROSPECTIVE REVIEW OF HISTORICAL RECORDS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21734241&pc=MHY | 3006630150-2025-02019 | 2025-03-06 05:00 | Death | BOSTON SCIENTIFIC NEUROMODULATION CORPORATION | 2025-03-31 04:00 | MHY | VERCISE GEVIA | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P150031 |  | 1 | Event Description: IT WAS REPORTED THAT THE DEEP BRAIN STIMULATION (DBS) PATIENT PASSED AWAY. NO ADDITIONAL INFORMATION WAS AVAILABLE.Event Description: IT WAS REPORTED THAT THE DEEP BRAIN STIMULATION (DBS) PATIENT PASSED AWAY. NO ADDITIONAL INFORMATION WAS AVAILABLE. ADDITIONAL INFORMATION WAS RECEIVED THAT THE DBS SYSTEM WAS NOT INVOLVED IN THE PATIENTS PASSING. Manufacturer Narrative: BLOCK B3: APPROXIMATED BASED ON THE DATE THE MANUFACTURER BECAME AWARE OF THE EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21734289&pc=HGM | 9610816-2025-000260 | 2025-01-31 05:00 | Death | PHILIPS MEDIZIN SYSTEME BÖBLINGEN GMBH | 2025-03-31 04:00 | HGM | AVALON FETAL MONITOR FM50 | Incorrect, Inadequate or Imprecise Result or Readings | Respiratory Acidosis | K140535 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT WAS BORN WITH SEVERE ACIDOSIS AND CARED FOR IMMEDIATELY BUT PASSED AWAY AFTER 55 MINUTES. THE USER THOUGHT THE PROBLEM MAY HAVE BEEN THAT THE DEVICE SWITCHED THE FETAL AND MATERNAL PULSE ON THE PRINTED TRACE. THE DEVICE REMAINS IN USE. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTION WAS REPORTED; THEREFORE, GOOD FAITH EFFORTS ARE BEING PERFORMED.Event Description: PHILIPS RECEIVED A COMPLAINT ON THE AVALON FM50 FETAL MONITOR INDICATING THAT A PATIENT WAS BORN WITH SEVERE ACIDOSIS AND CARED FOR IMMEDIATELY BUT PASSED AWAY AFTER 55 MINUTES. THE USER THOUGHT THE PROBLEM MAY HAVE BEEN THAT THE DEVICE SWITCHED THE FETAL AND MATERNAL PULSE ON THE PRINTED TRACE. THE PATIENT PRESENTED TO THE EMERGENCY DEPARTMENT AT 10:59 DUE TO LOSS OF AMNIOTIC FLUID AND REGULAR CONTRACTIONS, WITH FETAL MOVEMENT AND FETAL HEART RATE (FHR) PRESENT. FHR WAS MONITORED THROUGH 14:30 WHEN THE PATIENT TOOK A BATH UNTIL 16:35 WITHOUT FETAL MONITORING. AT 16:45, THE FHR WAS AGAIN MONITORED WITH A FHR AT 120 BPM. AT 17:45, THE FHR WAS IN THE 140S WITH LATE DECELERATIONS AT 100BPM, AND AT 18:45, THE FHR WAS 'NOT WELL CAPTURED' DUE TO PATIENT MOVEMENT, WITH FHR AT 140 BPM AND VARIABLE DECELERATIONS TO 100BPM. AT 18:55, THE FHR WAS NOT PICKED UP AND A SENSOR WIRE WAS IMPLEMENTED AT 19:00, WITH FHR AT 140BPM. AT 19:55, IT WAS INDICATED THE FHR WAS ¿NOT WELL CAPTURED¿ AND THE BABY WAS DELIVERED AT 20:05 WITH APGARS AT 0/0/0. RESUSCITATION WAS ATTEMPTED THROUGH 21:45, BUT DEATH WAS DECLARED DUE TO THE TOTAL ABSENCE OF SPONTANEOUS CARDIAC ACTIVITY. Manufacturer Narrative: THE COMPLAINT WAS ESCALATED FOR TECHNICAL INVESTIGATION, AND THE CHRONOLOGY OF THE EVENT WAS REVIEWED BY A CLINICAL SCIENTIST. IT WAS INDICATED THE MOTHER AT 37+4 WEEKS GESTATION WITH A NORMAL PREGNANCY PRESENTED DUE TO BROKEN WATER WITH REGULAR CONTRACTIONS. THE APGAR OF THE BABY AT BIRTH WAS 0/0/0. DURING THE COURSE OF LABOR, VISUAL AND AUDIBLE ALARMS WERE PRESENT FOR TACHYCARDIA, BRADYCARDIA, AND COINCIDENCE OR CROSS-CHANNEL VERIFICATION (CCV) PARAMETERS. BASED ON THE CHRONOLOGY PROVIDED, THE DOCUMENTATION DOES NOT DESCRIBE ACKNOWLEDGEMENT/HANDLING OF ALARMS. IT WAS DOCUMENTED THAT FHR WAS DIFFICULT TO FIND AT TIMES. AT 1900, THE WIRELESS TRANSDUCER WAS EXCHANGED FOR A WIRED TRANSDUCER; HOWEVER, THE TOCO + MATERNAL PULSE (MP) TRANSDUCER WAS NOT APPLIED FROM 1910-1946; THEREFORE, THE CCV FUNCTIONALITY WOULD NOT ALARM DURING THIS TIME PERIOD DUE TO NO MATERNAL INPUT AND OTHER MEANS OF MEASURING MATERNAL PULSE WERE NOT DESCRIBED IN THE DOCUMENTATION. IT REMAINED UNCLEAR IF THE STAFF WAS WORKING WITH THE FM50 MONITOR OR THE NON-PHILIPS CENTRAL MONITOR DURING LABOR. AT 1946, THE TOCO +MP WAS APPLIED WITH TRACES OF 80-95 BPM REGISTERED THREE TIMES DURING A PERIOD OF ONE MINUTE EACH BY THE TRANSDUCER. IT COULD NOT BE CONCLUDED THIS HEART RATE WAS FHR SINCE THE MATERNAL PULSE HAD BEEN >100BPM SINCE FETAL MONITORING STARTED. CCV ALARMS WERE ACTIVATED AT ONE POINT. ADDITIONAL CHRONOLOGY WAS REVIEWED BY THE CLINICAL SCIENTIST THROUGH THE TIME OF BIRTH WAS NOTED. THE CONCLUSION IS AS FOLLOWS: "THE CTG-RECORDINGS OF THE AVALON FM 50 RAN FOR APPROXIMATELY 4 HOURS. COINCIDENCE ALARMING AND BRADYCARDIA ALARMING WERE ACTIVE AND REGULARLY TRIGGERED. THE AVAILABLE DOCUMENTATION AND INFORMATION SHOW THAT THERE WAS SOME DIFFICULTY IN DERIVING THE FHR, ESPECIALLY WHEN THE WOMAN IN LABOR WAS MOBILE AND IN AN ¿ALL FOURS¿ POSITION. HOWEVER, NO EVIDENCE IS AVAILABLE FROM THE PROVIDED SOURCES, THAT A COUNTERCHECK OF THE MATERNAL PULSE HAS BEEN PERFORMED. BY THE RECORDINGS OF THE FETAL MONITOR, IT WAS CONFIRMED THAT NO SPO2 FINGER SENSOR NOR A MATERNAL ELECTROCARDIOGRAM (MECG) HAS BEEN USED. ACCORDING TO THE INFORMATION PROVIDED BY THE HOSPITAL, DURING THE PERIOD OF CCV ALARMING, AN OBSTETRIC SONOGRAPHY OF THE FETAL CONDITION HAS NOT BEEN PERFORMED. THERE IS NO EVIDENCE THAT A COUNTERCHECK OF THE MATERNAL PULSE BY MANUAL DETERMINATION OF THE MATERNAL PULSE WAS COMPARED WITH THE FETAL HEART RATE SOUND SIGNALS FROM THE US TRANSDUCER. THE REGISTERED MATERNAL PULSE OF THE WOMAN IN LABOR WAS MAINLY ABOVE 110 BPM AND THEREFORE CAUSING CCVS ON A REGULAR BASIS IF THE TOCO+MP TRANSDUCER WAS APPLIED, WHICH HAS NOT BEEN THE CASE DURING ABOUT 50 MINUTES IN TOTAL. BASED ON THE INFORMATION AVAILABLE AND THE TESTING CONDUCTED, THE DEVICE WAS PROVIDING COINCIDENCE AND OTHER ALARMS THROUGHOUT THE COURSE OF LABOR. THERE WAS NO ISSUE FOUND WITH THE DEVICE. THE REPORTED PROBLEM WAS NOT CONFIRMED. THE DEVICE WAS CONFIRMED TO BE OPERATING PER SPECIFICATIONS AND NO FAILURE WAS IDENTIFIED. Manufacturer Narrative: A PHILIPS FIELD SERVICE ENGINEER (FSE) WENT ONSITE, AND THE DEVICE PASSED OPERATING AND PERFORMANCE TESTS. THE FSE RECOVERED THE INTERNAL LOGS OF THE DEVICE FOR ANALYSIS. THE PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE. SECTION E REPORTER PHONE #: (B)(6).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21734765&pc=MIH | 3007284313-2025-03866 | 2025-03-09 05:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-03-31 04:00 | MIH | GORE® EXCLUDER® CONFORMABLE AAA ENDOPROSTHESIS | Insufficient Information | Insufficient Information | P200030 |  | 1 | Event Description: THE FIELD SALES ASSOCIATE (FSA) REPORTED THE FOLLOWING TO GORE: ON (B)(6) 2025, THE PATIENT UNDERWENT AN ENDOVASCULAR PROCEDURE TO TREAT AN ANEURYSM RUPTURE UTILIZING THE GORE® EXCLUDER® CONFORMABLE AAA IN ZONE 9. IT WAS REPORTED THAT THE PATIENT PRESENTED SYMPTOMATIC OF A RUPTURED ANEURYSM. HIS LEFT RENAL WAS INTENTIONALLY COVERED BECAUSE THE ANEURYSM EXTENDED TO THE LEFT RENAL. ALSO, TWO PLCS WERE PLACED IN THE PATIENT. A FINAL ANGIOGRAM WAS PERFORMED THAT SHOWED NO EVIDENCE OF ENDOLEAK AND THE CONTAINMENT OF THE ANEURYSM SAC. THE RIGHT RENAL WAS FUNCTIONAL WITH NO ISSUES. THE PATIENT TOLERATED THE PROCEDURE. ON (B)(6), 2025, THE PATIENT'S RENAL FUNCTION DECLINED. THE PHYSICIAN PERFORMED AN OPEN RENAL REVASCULARIZATION WITH ENDOSCOPIC INTERVENTION. THE PHYSICIAN ALSO PERFORMED A LAPAROTOMY. AT SOME POINT DURING THESE PROCEDURES, AN AORTIC EXTENDER AND NON GORE BARE METAL STENT WAS PLACED IN THE PATIENT. THE PATIENT TOLERATED THE PROCEDURE. ON MARCH 10, THE FSA WAS IN COMMUNICATION WITH THE PHYSICIAN. THE PHYSICIAN STATED THE PATIENT EXPIRED. THE EXACT EXPIRATION DATE IS UNKNOWN. NO ADDITIONAL INFORMATION IS AVAILABLE. THE CAUSE OF THE REPORTED EVENT IS UNKNOWN. Manufacturer Narrative: H6: CODE A26- USED TO CAPTURE INSUFFICIENT INFORMATION AVAILABLE TO DETERMINE IF GORE DEVICE CONTRIBUTED TO DEATH. H6: CODE E2401- USED TO CAPTURE INSUFFICIENT INFORMATION AVAILABLE TO DETERMINE IF ANY HEALTH EFFECT APPEARS TO HAVE OCCURRED.. IT SHOULD BE NOTED THAT, PER THE GORE® EXCLUDER® AAA ENDOPROSTHESIS INSTRUCTIONS FOR USE (IFU), ADVERSE EVENTS THAT MAY OCCUR AND / OR REQUIRE INTERVENTION INCLUDE, BUT ARE NOT LIMITED TO DEATH. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. Manufacturer Narrative: ADDED G3/G4, H1/H2. CORRECTION: IT SHOULD BE NOTED THAT, PER THE GORE® EXCLUDER® CONFORMABLE AAA ENDOPROSTHESIS INSTRUCTIONS FOR USE (IFU), ADVERSE EVENTS THAT MAY OCCUR AND / OR REQUIRE INTERVENTION INCLUDE BUT ARE NOT LIMITED TO DEATH. Manufacturer Narrative: ADDED G3/G4, H1/H2. CORRECTED B3.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21734767&pc=NAW | 2124215-2025-19639 | 2022-01-20 05:00 | Death | BIOCOMPATIBLES UK LIMITED | 2025-03-31 04:00 | NAW | THERASPHERE Y-90 | Adverse Event Without Identified Device or Use Problem | Dyspnea; Fatigue; Liver Damage/Dysfunction; Ascites; Swelling/ Edema |  |  | 1 | Event Description: IT WAS REPORTED THAT THIS PATIENT PASSED AWAY RELATED TO OEDEMA DECOMPENSATION. ON (B)(6) 2021, THE SUBJECT WAS ENROLLED INTO THE STUDY AND THE TREATMENT WITH THERASPHERE WAS PERFORMED ON THE SAME DAY TO THE SELECTIVE LIVER. THE CATHETER POSITIONING DURING THE ADMINISTRATION PROCEDURE WAS NOT DOCUMENTED. TWO VIALS WERE ADMINISTERED TO THE PATIENT; 4.72 GBQ OF THERASPHERE WAS ADMINISTERED TO THE SELECTIVE LIVER THROUGH VIAL 1 AND 4.56 GBQ OF THERASPHERE WAS ADMINISTERED TO THE SELECTIVE LIVER THROUGH VIAL 2. TOTAL THERASPHERE ACTIVITY ADMINISTERED TO THE SELECTIVE LIVER WAS 9.28 GBQ. ON (B)(6) 2022, 223 DAYS POST-INDEX PROCEDURE, THE SUBJECT WAS DIAGNOSED WITH OEDEMA DECOMPENSATION. ON THE SAME DAY, THE SUBJECT WAS ADMITTED TO EMERGENCY ROOM DUE TO ALTERATION OF THEIR GENERAL CONDITION. DURING THE HOSPITALIZATION ASCITES WITH SIGNIFICANT LOWER LIMB EDEMA WAS DIAGNOSED, THUS DIURETICS WERE ADMINISTERED, AND REGULAR ASCITES PUNCTURE WAS PERFORMED TO TREAT THE EVENT. ADDITIONALLY, MAINTENANCE OF DIURETICS WAS ALSO PRESCRIBED. ON (B)(6) 2022, THE OUTCOME OF THE EVENT WAS CONSIDERED RESOLVED AND THE SUBJECT WAS DISCHARGED ON THE SAME DAY. POST-DISCHARGE, THE SUBJECT HAS HAD PERSISTENT ASCITES OF GREAT ABUNDANCE (ECOG, 2), WHICH WAS TREATED WITH REPEATED PUNCTURE. ON (B)(6) 2022, THE SUBJECT PASSED AWAY DUE TO PROGRESSION OF LIVER DISEASE AND NOT DUE TO THE PROGRESSION OF THEIR CANCER, WHERE THEIR HEPATOPATHY WAS DECOMPENSATED AS A RESULT OF RADIO-EMBOLIZATION.Event Description: IT WAS REPORTED THAT THIS PATIENT PASSED AWAY RELATED TO OEDEMA DECOMPENSATION. ON (B)(6) 2021, THE SUBJECT WAS ENROLLED INTO THE STUDY AND THE TREATMENT WITH THERASPHERE WAS PERFORMED ON THE SAME DAY TO THE SELECTIVE LIVER. THE CATHETER POSITIONING DURING THE ADMINISTRATION PROCEDURE WAS NOT DOCUMENTED. TWO VIALS WERE ADMINISTERED TO THE PATIENT; 4.72 GBQ OF THERASPHERE WAS ADMINISTERED TO THE SELECTIVE LIVER THROUGH VIAL 1 AND 4.56 GBQ OF THERASPHERE WAS ADMINISTERED TO THE SELECTIVE LIVER THROUGH VIAL 2. TOTAL THERASPHERE ACTIVITY ADMINISTERED TO THE SELECTIVE LIVER WAS 9.28 GBQ. ON (B)(6) 2022, 223 DAYS POST-INDEX PROCEDURE, THE SUBJECT WAS DIAGNOSED WITH OEDEMA DECOMPENSATION. ON THE SAME DAY, THE SUBJECT WAS ADMITTED TO EMERGENCY ROOM DUE TO ALTERATION OF THEIR GENERAL CONDITION. DURING THE HOSPITALIZATION ASCITES WITH SIGNIFICANT LOWER LIMB EDEMA WAS DIAGNOSED, THUS DIURETICS WERE ADMINISTERED, AND REGULAR ASCITES PUNCTURE WAS PERFORMED TO TREAT THE EVENT. ADDITIONALLY, MAINTENANCE OF DIURETICS WAS ALSO PRESCRIBED. ON (B)(6) 2022, THE OUTCOME OF THE EVENT WAS CONSIDERED RESOLVED AND THE SUBJECT WAS DISCHARGED ON THE SAME DAY. POST-DISCHARGE, THE SUBJECT HAS HAD PERSISTENT ASCITES OF GREAT ABUNDANCE (ECOG, 2), WHICH WAS TREATED WITH REPEATED PUNCTURE. ON (B)(6) 2022, THE SUBJECT PASSED AWAY DUE TO PROGRESSION OF LIVER DISEASE AND NOT DUE TO THE PROGRESSION OF THEIR CANCER, WHERE THEIR HEPATOPATHY WAS DECOMPENSATED AS A RESULT OF RADIO-EMBOLIZATION. FOLLOW-UP INFORMATION WAS PROVIDED WITH ADDITIONAL DETAILS THAT OCCURRED BETWEEN (B)(6) 2022 AND WHEN THE PATIENT PASSED AWAY. THE PATIENT WAS SEEN AT A FOLLOW-UP VISIT ON (B)(6) 2021 DUE TO FATIGUE. FROM (B)(6) 2022, THE PATIENT WAS OBSERVED FOR DYSPNEA AND FATIGUE, WHERE THE EXPLORATION CONCLUDED LIVER DECOMPENSATION WITH ASCITES AND OEDEMA, HYPOALBUMINEMIA 25G/L (BUT NORMAL BILIRUBIN). THE CT SCAN DID NOT IDENTIFY A TUMOR BUT SIRT SEQUELAE, ASCITES, LARGE PELVIC AND ABDOMINAL VEIN, AND LARGE UMBILICAL VEIN. PATIENT WAS DISCHARGED WITH "DIURETICS" ON (B)(6) 2020 (AS PREVIOUSLY REPORTED). ON (B)(6) 2022 A FOLLOW-UP PROVIDED THE PATIENT HAD DETERIORATED, ECOG2, NO CLINICAL ASCITES, LOWER LIMB OEDEMA. ON (B)(6) 2022 A CT SCAN DID NOT IDENTIFY A TUMOR BUT SIRT SEQUELAE AND ASCITES. THE PATIENT UNDERSTOOD THAT WITH THEIR LIVER FUNCTION DETERIORATION NO MORE TREATMENT WAS PROPOSED AND WAS PLACED ON IN-HOME SUPPORTIVE CARE. Manufacturer Narrative: D3 & G1: MANUFACTURER ZIP/POSTAL CODE: GU9 8QL. E1 INITIAL REPORTER ADDRESS 1: (B)(6). THE UPN AND LOT NUMBERS WERE NOT PROVIDED BY THE STUDY; THUS, THE UDI AND OTHER PRODUCT SPECIFIC INFORMATION IS NOT AVAILABLE. Manufacturer Narrative: D3 & G1: MANUFACTURER ZIP/POSTAL CODE: (B)(6). E1 INITIAL REPORTER ADDRESS 1: (B)(6). THE UPN AND LOT NUMBERS WERE NOT PROVIDED BY THE STUDY; THUS, THE UDI AND OTHER PRODUCT SPECIFIC INFORMATION IS NOT AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21735298&pc=LWS | 2124215-2025-19838 | 2025-03-07 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-03-31 04:00 | LWS | EMBLEM? MRI S-ICD | Adverse Event Without Identified Device or Use Problem | Bradycardia; Asystole | P110042 |  | 1 | Event Description: IT WAS REPORTED THAT THIS PATIENT EXPIRED DURING THE SUBCUTANEOUS IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (S-ICD) IMPLANT PROCEDURE. AFTER THE SYSTEM WAS PLACED, THE PHYSICIAN TRIED TO INDUCE THE PATIENT FOR DEFIBRILLATION THRESHOLD (DFT) TESTING BUT THE PATIENT SELF-CONVERTED OUT OF THE VENTRICULAR FIBRILLATION (VF). A COMMANDED 10 JOULE SHOCK WAS DELIVERED TO EVALUATE THE SHOCK IMPEDANCE, THEN THE PATIENT DECOMPENSATED AND WAS UNABLE TO BE REVIVED. THE RHYTHM HAD DECOMPENSATED, THEN BRADYCARDIA OCCURRED, AND EVENTUALLY THE PATIENT FLATLINED. THE PHYSICIAN WAS UNABLE TO GAIN ACCESS TO PLACE A TEMPORARY PACER. IT WAS NOTED THE S-ICD SYSTEM REMAINED IMPLANTED IN THE PATIENT AND WAS NOT GOING TO BE EXPLANTED FOR RETURN. Manufacturer Narrative: IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21735679&pc=BZD | 2518422-2025-040474 | 2022-08-17 04:00 | Death | RESPIRONICS, INC. | 2025-03-31 04:00 | BZD | DS2ADV AUTO CPAP | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: IN THIS REPORT, SECTION BOX B: ADVERSE EVENT OR PRODUCT PROBLEM (ADVERSE EVENT/PRODUCT PROBLEM), BOX D: SUSPECT MEDICAL DEVICE (OPERATOR OF DEVICE), BOX G: ALL MANUFACTURERS (REPORT SOURCE), BOX H: DEVICE MANUFACTURERS (TYPE OF REPORTED COMPLAINT, EVALUATION METHOD CODE GRID, EVALUATION RESULTS CODE GRID, CONCLUSION CODE GRID) HAVE BEEN UPDATED/CORRECTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21728549&pc=MIH | 2017233-2025-06008 | 2025-02-14 05:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-03-30 04:00 | MIH | GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM | Contamination; Adverse Event Without Identified Device or Use Problem | Chest Pain; Unspecified Infection; Inflammation | P040043 |  | 1 | Event Description: IT WAS REPORTED TO GORE, ON (B)(6) 2024, A PATIENT PRESENTED WITH A THORACOABDOMINAL ANEURYSM UNDERWENT A THORACIC ENDOVASCULAR ANEURYSM REPAIR (TEVAR) PROCEDURE WITH A GORE® TAG® THORACIC BRANCH ENDOPROSTHESIS (TBE, SERIAL TAC: (B)(6)) AT (B)(6). THIS DEVICE (SERIAL NUMBER UNKNOWN) WAS IMPLANTED IN ZONE 2. A GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM (TGMR313120E, SERIAL: (B)(6)) WAS IMPLANTED DISTALLY. DURING THE PROCEDURE, A PERCUTANEOUS ACCESS WAS NOT POSSIBLE AND A RIGHT FEMORAL CUTDOWN AND SURGICAL REPAIR WAS REQUIRED, THE INVOLVED DEVICE WAS A NON-GORE DEVICE FROM ABBOTT (PERCLOSE PROGLIDE¿). THE PATIENT REQUIRED A TWO STAGE REPAIR; TEVAR AND IN SECOND STAGE AN INNER BRANCH ENDOVASCULAR AORTIC REPAIR (IBEVAR) WITH ARTRIVION (NON-GORE) DEVICE, WHICH WAS PLANNED IN 6-8 WEEKS¿ TIME. POST OPERATION, THE PATIENT HAD CHEST INFECTION TREATED WITH ANTIBIOTICS AND RECOVERED WELL AND WAS DISCHARGED. HOWEVER, 4 WEEKS LATER, THE PATIENT PRESENTED WITH SEVERE CHEST PAIN. CTA SHOWED SLIGHT INCREASE IN THE DISTAL THORACIC ANEURYSM WITH SIGNS OF SYMPTOMATIC THORACIC ANEURYSM. BLOOD RESULTS WERE NORMAL. AN IBEVAR WAS PERFORMED ON (B)(6) 2024 AS EMERGENCY. THE BRAND OF THE IMPLANTED DEVICE WAS ARTRIVION. TWO DAYS LATER, THE PATIENT DEVELOPED MEDIASTINITIS TREATED WITH ANTIBIOTICS BUT THE INFECTION CONTINUED PROGRESSING AND WAS NOT RESPONDING WITH ANTIBIOTICS. THE SURGEON STATED THAT THIS MEDIASTINITIS IS THE CAUSE OF THE GRAFT INFECTION. THE CAUSE OF THE MEDIASTINITIS IS UNKNOWN. DUE TO THIS INFECTION, AN EXPLANT TOOK PLACE MID-END (B)(6) 2025 (EXACT DATE UNKNOWN) AT (B)(6) HOSPITAL. DURING THIS PROCEDURE, THE SURGEON PROCEEDED WITH A FROZEN ELEPHANT TRUNK GRAFT. THE PATIENT UNFORTUNATELY DID NOT SURVIVE THE EXPLANT AND ADJUNCTIVE PROCEDURE. THE EXPLANT SURGEON REPORTED THAT THE CAUSE OF DEATH MIGHT BE DUE TO THROMBOTIC SHOWER. IT WAS ADDITIONALLY REPORTED, THE PATIENT HAD UNDERLYING CONDITION(S) THAT MAY HAVE TRIGGERED THE INFECTION AT LAST PROCEDURE FROM 2ND STAGE, SUCH HISTORY OF PERIODIC FEVER SYNDROME AND EMPHYSEMA.Event Description: IT WAS REPORTED TO GORE, ON (B)(6) 2024, A PATIENT PRESENTED WITH A THORACOABDOMINAL ANEURYSM UNDERWENT A THORACIC ENDOVASCULAR ANEURYSM REPAIR (TEVAR) PROCEDURE WITH A GORE® TAG® THORACIC BRANCH ENDOPROSTHESIS (TBE, SERIAL TAC: (B)(6)) AT (B)(6) HOSPITAL IN (B)(6) AND WAS IMPLANTED IN ZONE 2. A GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM (TGMR313120E, SERIAL: (B)(6)) WAS IMPLANTED DISTALLY. DURING THE PROCEDURE, A PERCUTANEOUS ACCESS WAS NOT POSSIBLE AND A RIGHT FEMORAL CUTDOWN AND SURGICAL REPAIR WAS REQUIRED, THE INVOLVED DEVICE WAS A NON-GORE DEVICE FROM ABBOTT (PERCLOSE PROGLIDE¿). THE PATIENT REQUIRED A TWO STAGE REPAIR; TEVAR AND IN SECOND STAGE AN INNER BRANCH ENDOVASCULAR AORTIC REPAIR (IBEVAR) WITH ARTRIVION (NON-GORE) DEVICE, WHICH WAS PLANNED IN 6-8 WEEKS¿ TIME. POST OPERATION, THE PATIENT HAD CHEST INFECTION TREATED WITH ANTIBIOTICS AND RECOVERED WELL AND WAS DISCHARGED. HOWEVER, 4 WEEKS LATER, THE PATIENT PRESENTED WITH SEVERE CHEST PAIN. CTA SHOWED SLIGHT INCREASE IN THE DISTAL THORACIC ANEURYSM WITH SIGNS OF SYMPTOMATIC THORACIC ANEURYSM. BLOOD RESULTS WERE NORMAL. AN IBEVAR WAS PERFORMED ON (B)(6) 2024 AS EMERGENCY. THE BRAND OF THE IMPLANTED DEVICE WAS ARTRIVION. TWO DAYS LATER, THE PATIENT DEVELOPED MEDIASTINITIS TREATED WITH ANTIBIOTICS BUT THE INFECTION CONTINUED PROGRESSING AND WAS NOT RESPONDING WITH ANTIBIOTICS. THE SURGEON STATED THAT THIS MEDIASTINITIS IS THE CAUSE OF THE GRAFT INFECTION. THE CAUSE OF THE MEDIASTINITIS IS UNKNOWN. DUE TO THIS INFECTION, AN EXPLANT TOOK PLACE (B)(6) 2025 (EXACT DATE UNKNOWN) AT (B)(6) HOSPITAL. DURING THIS PROCEDURE, THE SURGEON PROCEEDED WITH A FROZEN ELEPHANT TRUNK (FET) GRAFT. THE PATIENT UNFORTUNATELY DID NOT SURVIVE THE EXPLANT AND ADJUNCTIVE PROCEDURE. THE EXPLANT SURGEON REPORTED THAT THE CAUSE OF DEATH MIGHT BE DUE TO THROMBOTIC SHOWER. IT WAS ADDITIONALLY REPORTED, THE PATIENT HAD UNDERLYING CONDITION(S) THAT MAY HAVE TRIGGERED THE INFECTION AT LAST PROCEDURE FROM 2ND STAGE, SUCH HISTORY OF PERIODIC FEVER SYNDROME AND EMPHYSEMA. Manufacturer Narrative: UPDATED G3. B5: MORE INFORMATION IS DESCRIBED AND EVENT DESCRIPTION UPDATED. B7: UPDATED PATIENT MEDICAL HISTORY FOR PERIODIC FEVER SYND. B3: THE EVENT DATE REMAINS UNKNOWN WITH THE INFECTION ONSET DATE. NO CONCRETE EVENT DATE IS GIVEN; THUS, WE SELECTED THE DATE WHEN GORE FIRST TIME BECAME AWARE. LATER, WE COULD DETERMINE MORE INFORMATION TO THE INFECTION SO THAT THE EXACT ONSET DATE FOR THE GRAFT WAS EARLIER. HOWEVER, IT WAS KEPT THE (B)(6) 2025 AS NO EXACT DATE OF INCIDENT. THERE ARE POSSIBLE EARLIER EVENTS THAT HAVE BEEN OCCURRED AND POSSIBLY PRIOR THE IBEVAR IMPLANT DATE (BASED ON THE PATIENT HISTORY). THUS, IT CANNOT BE DETERMINED WHEN A FIRST INFECTION WAS OBSERVED, THEN WHEN MEDIASTINITIS HAS BEEN DIAGNOSED AND THEN STARTED THE REINTERVENTION BY PHYSICIAN. B2: DATE OF PATIENT DEATH, IT REMAINS UNKNOWN TO GORE WHEN THE PATIENT DIED. THE CAUSE OF DEATH WAS ONLY REPORTED AND THE DATE OF DEATH SEEMED TO BE IN (B)(6) 2025. H6, - MEDICAL DEVICE PROBLEM, CODE A1801 ADDED. A24 NO LONGER APPLICABLE. - INVESTIGATION FINDINGS, CODE C19 ADDED. C21 NO LONGER APPLICABLE. - INVESTIGATION CONCLUSIONS: CODE D12 ADDED. D16 NO LONGER APPLICABLE. - HEATH EFFECT-CLINICAL CODE: CODE E233001 AND E2326 ADDED. - HEALTH EFFECT-IMPACT CODE: CODE F02 AND F2301 ADDED. - TYPE OF INVESTIGATION: CODE B11 ADDED. THE INVESTIGATION WAS COMPLETED WITH THAT RESULT: A REVIEW OF THE MANUFACTURING & STERILIZATION RECORDS INDICATED THE LOTS MET ALL PRE-RELEASE SPECIFICATIONS. NEITHER IMAGES ENABLING DIRECT ASSESSMENT OF PRODUCT PERFORMANCE NOR THE PRODUCT ITSELF WERE RETURNED FOR EVALUATION. BASED ON THE INCIDENT DESCRIPTION AND THE SUBSEQUENT INVESTIGATION, NO FURTHER INFORMATION WAS PROVIDED TO GORE, WE ARE UNABLE TO DETERMINE THE CAUSE OF THIS INCIDENT AND ASSIGN A ROOT CAUSE. PER THE GORE® TAG® CONFORMABLE THORACIC STENT GRAFT INSTRUCTIONS FOR USE (IFU), POSSIBLE ADVERSE EVENTS AND COMPLICATIONS THAT MAY OCCUR WITH THE USE OF GORE® TAG® CONFORMABLE THORACIC STENT GRAFT INCLUDE, BUT ARE NOT LIMITED TO: DEATH, STENT GRAFT INFECTION. Manufacturer Narrative: B3/D6B: THE EXACT DATE OF EVENT IS UNKNOWN. ALSO, THE EXPLANT DATE IS NOT KNOWN. H3: THE EXPLANTED DEVICES CANNOT BE EVALUATED SINCE NOT ACCESSIBLE TO GORE. B14: PRELIMINARY RESULTS ARE NOT YET AVAILABLE. B13: INTERVIEWS WITH THE IMPLANT FACILITY WERE ONGOING AND THE EXPLANT FACILITY WAS CONTACTED FOR THE INITIAL INFORMATION FOR THE INFECTION REASON AND FOUND THE CAUSE OF THIS INFECTION. EXPLANT SURGEON IS KNOWN AND FURTHER QUESTIONS ARE PENDING. D10: CONCOMITANT MEDICAL PRODUCTS: (RELEVANT ASSOCIATED DEVICES USED WITH THE DEVICE BEING REPORTED ON): GORE® TAG® THORACIC BRANCH ENDOPROSTHESIS (AORTIC COMPONENT). NOT FILLED IN SECTION C: THEREFORE, THIS IS CBAS® HEPARIN SURFACE INCORPORATES CARMEDA HEPARIN MANUFACTURED FROM HEPARIN SODIUM API, WHICH IS COVALENTLY BOUND TO THE DEVICE SURFACE AND IS ESSENTIALLY NON-ELUTING. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. Manufacturer Narrative: UPDATED G3. H6, UPDATED IN INVESTIGATION CONCLUSIONS CODE D15. INVESTIGATION FINDINGS: CODE C19, NO DEVICE PROBLEM FOUND: A REVIEW OF THE MANUFACTURING & STERILIZATION RECORDS INDICATED THE LOTS MET ALL PRE-RELEASE SPECIFICATIONS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21728690&pc=NKE | 2017865-2025-38894 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-03-30 04:00 | NKE | QUADRA ALLURE MP CRT-P | Adverse Event Without Identified Device or Use Problem | Unspecified Infection | P030035 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT PASSED AWAY DUE TO SEPSIS AND ATRIAL FIBRILLATION. THE PATIENT'S PACEMAKER AND LEADS WERE EXPLANTED. IT IS UNKNOWN IF THE PATIENT¿S INFECTION OR DEATH WAS CAUSED OR CONTRIBUTED BY ANY OF THE PATIENT¿S ABBOTT PRODUCTS OR ANY PROCEDURE INVOLVING THE ABBOTT PRODUCTS. Manufacturer Narrative: DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES. REVIEW OF THE STERILIZATION RECORDS CONFIRMED NORMAL STERILIZATION CYCLES FOR THE PRODUCTS. THE CAUSE OF INFECTION COULD NOT BE DETERMINED. Manufacturer Narrative: DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES. REVIEW OF THE STERILIZATION RECORDS CONFIRMED NORMAL STERILIZATION CYCLES FOR THE PRODUCTS. THE CAUSE OF INFECTION COULD NOT BE DETERMINED. ANALYSIS PERFORMED INCLUDING ELECTRICAL, MECHANICAL AND ENVIRONMENTAL TESTS INDICATED THAT THE DEVICE EXHIBITED NORMAL DEVICE CHARACTERISTICS WITH NO ANOMALIES. LONGEVITY ASSESSMENT WAS PERFORMED AND DEVICE WAS IN THE NORMAL RANGE OF OPERATION WITH APPROPRIATE REMAINING LONGEVITY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21728691&pc=NVN | 2017865-2025-38895 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-03-30 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | Unspecified Infection | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT PASSED AWAY DUE TO SEPSIS AND ATRIAL FIBRILLATION. THE PATIENT'S PACEMAKER AND LEADS WERE EXPLANTED. IT IS UNKNOWN IF THE PATIENT¿S INFECTION OR DEATH WAS CAUSED OR CONTRIBUTED BY ANY OF THE PATIENT¿S ABBOTT PRODUCTS OR ANY PROCEDURE INVOLVING THE ABBOTT PRODUCTS. Manufacturer Narrative: DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES. REVIEW OF THE STERILIZATION RECORDS CONFIRMED NORMAL STERILIZATION CYCLES FOR THE PRODUCTS. THE CAUSE OF INFECTION COULD NOT BE DETERMINED. VISUAL EXAMINATION FOUND NO MALFUNCTIONS. FULL ANALYSIS NOT NEEDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21728692&pc=NVN | 2017865-2025-38896 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-03-30 04:00 | NVN | TENDRIL STS | Defective Device; Adverse Event Without Identified Device or Use Problem | Unspecified Infection | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT PASSED AWAY DUE TO SEPSIS AND ATRIAL FIBRILLATION. THE PATIENT'S PACEMAKER AND LEADS WERE EXPLANTED. IT IS UNKNOWN IF THE PATIENT¿S INFECTION OR DEATH WAS CAUSED OR CONTRIBUTED BY ANY OF THE PATIENT¿S ABBOTT PRODUCTS OR ANY PROCEDURE INVOLVING THE ABBOTT PRODUCTS. Manufacturer Narrative: DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES. REVIEW OF THE STERILIZATION RECORDS CONFIRMED NORMAL STERILIZATION CYCLES FOR THE PRODUCTS. THE CAUSE OF INFECTION COULD NOT BE DETERMINED. Manufacturer Narrative: DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES. REVIEW OF THE STERILIZATION RECORDS CONFIRMED NORMAL STERILIZATION CYCLES FOR THE PRODUCTS. THE CAUSE OF INFECTION COULD NOT BE DETERMINED. VISUAL EXAMINATION OF THE LEAD FOUND FULL BREACHED DEGRADATION BREACHING THE OUTER INSULATION AND EXPOSING THE OUTER COIL LOCATED ADJACENT TO THE CONNECTOR BOOT AT 4.5 CM FROM THE CONNECTOR PIN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21728693&pc=NIK | 2017865-2025-38897 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-03-30 04:00 | NIK | QUARTET | Adverse Event Without Identified Device or Use Problem | Unspecified Infection | P030054 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT PASSED AWAY DUE TO SEPSIS AND ATRIAL FIBRILLATION. THE PATIENT'S PACEMAKER AND LEADS WERE EXPLANTED. IT IS UNKNOWN IF THE PATIENT¿S INFECTION OR DEATH WAS CAUSED OR CONTRIBUTED BY ANY OF THE PATIENT¿S ABBOTT PRODUCTS OR ANY PROCEDURE INVOLVING THE ABBOTT PRODUCTS. Manufacturer Narrative: DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES. REVIEW OF THE STERILIZATION RECORDS CONFIRMED NORMAL STERILIZATION CYCLES FOR THE PRODUCTS. THE CAUSE OF INFECTION COULD NOT BE DETERMINED. VISUAL EXAMINATION FOUND NO MALFUNCTIONS. FULL ANALYSIS NOT NEEDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21728750&pc=NJE | 2029214-2025-00802 | 2023-11-03 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-03-30 04:00 | NJE | SOLITAIRE AB STENT | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Intracranial Hemorrhage; Ischemia Stroke |  |  | 1 | Event Description: ZHANG, J., LONG, L., LI, J., ZHANG, H., YAN, W., ABULIMITI, A., ABULAJIANG, N., LU, Q., NGUYEN, T. N., & CAI, X.; NEURORADIOLOGY; 2024; 66:1021¿1029; IMPACT OF OBESITY-RELATED INDICATORS ON FRST-PASS EFECT IN PATIENTS WITH ISCHEMIC STROKE RECEIVING MECHANICAL THROMBECTOMY; DOI.ORG/10.1007/S00234-024-03350-X LITERATURE WAS REVIEWED REGARDING: "IMPACT OF OBESITY-RELATED INDICATORS ON FIRST-PASS EFFECT IN PATIENTS WITH ISCHEMIC STROKE RECEIVING MECHANICAL THROMBECTOMY." THE OBJECTIVE IS TO EVALUATE OBESITY-RELATED INDICATORS AS POTENTIAL PREDICTORS OF THE FIRST-PASS EFFECT (FPE) IN PATIENTS WITH ACUTE ISCHEMIC STROKE RECEIVING MECHANICAL THROMBECTOMY (MT). THE TIME FRAME OF THIS STUDY WAS: JANUARY 2019 TO DECEMBER 2021. MULTIPLE MANUFACTURER¿S DEVICES WERE USED IN THE STUDY POPULATION. THE STUDY MENTIONS VARIOUS MEDTRONIC THROMBECTOMY DEVICES INCLUDING STENT RETRIEVERS LIKE SOLITAIR FR AND SOLITAIR AB AS WELL AS CONTACT ASPIRATION USING THE REACT 68 CATHETER. DEATHS OCCURRED IN THE STUDY POPULATION. THE STUDY REPORTED A TOTAL OF 2 DEATHS (1.3% OF PATIENTS), WITH ONE DEATH IN THE FPE GROUP (2.1%) AND ONE IN THE NON-FPE GROUP (1.0%). HOWEVER, THE DOCUMENT DOES NOT SPECIFY THE CAUSES OF THESE DEATHS. AMONG PATIENT ADVERSE EVENTS INCLUDED: HIGHER RATES OF SUBARACHNOID HEMORRHAGE (SAH) IN NON-FPE PATIENTS (0% VS. 11.5%, P = 0.018). NON-FPE PATIENTS HAD A HIGHER RATE OF RESCUE THERAPY (BALLOON ANGIOPLASTY OR STENT IMPLANTATION), (0% VS. 13.5%, P = 0.012). HEIDELBERG BLEEDING CLASSIFICATION: 3C: 12 (7.9%) TOTAL, WITH 0% IN FPE AND 11.5% IN NON-FPE (P = 0.018). INTRACRANIAL HEMORRHAGE: 70 (46.3%) TOTAL, WITH 23 (48.9%) IN FPE AND 47 (45.2%) IN NON-FPE (P = 0.726). SYMPTOMATIC INTRACRANIAL HEMORRHAGE: 14 (9.3%) TOTAL, WITH 5 (10.6%) IN FPE AND 9 (8.7%) IN NON-FPE (P = 0.764). ASYMPTOMATIC INTRACRANIAL HEMORRHAGE: 56 (37.1%) TOTAL, WITH 18 (38.3%) IN FPE AND 38 (36.5%) IN NON-FPE (P = 0.857). CLINICAL OUTCOMES OF INTEREST: NON-FPE PATIENTS HAD A LONGER MEDIAN PROCEDURE TIME (145 MINUTES VS. 101 MINUTES, P < 0.001), A LOWER RATE OF 24-HOUR EARLY NEUROLOGIC IMPROVEMENT (16.3% VS. 31.9%, P = 0.034), AND A HIGHER PROBABILITY OF FUNCTIONAL DECLINE AT DISCHARGE (MRS 4¿5) (57.7% VS. 38.3%, P = 0.035). NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21728752&pc=NRY | 2029214-2025-00803 | 2023-11-03 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-03-30 04:00 | NRY | SOLITAIRE RD | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Intracranial Hemorrhage; Ischemia Stroke |  |  | 1 | Event Description: ZHANG, J., LONG, L., LI, J., ZHANG, H., YAN, W., ABULIMITI, A., ABULAJIANG, N., LU, Q., NGUYEN, T. N., & CAI, X.; NEURORADIOLOGY; 2024; 66:1021¿1029; IMPACT OF OBESITY-RELATED INDICATORS ON FRST-PASS EFECT IN PATIENTS WITH ISCHEMIC STROKE RECEIVING MECHANICAL THROMBECTOMY; DOI.ORG/10.1007/S00234-024-03350-X LITERATURE WAS REVIEWED REGARDING: "IMPACT OF OBESITY-RELATED INDICATORS ON FIRST-PASS EFFECT IN PATIENTS WITH ISCHEMIC STROKE RECEIVING MECHANICAL THROMBECTOMY." THE OBJECTIVE IS TO EVALUATE OBESITY-RELATED INDICATORS AS POTENTIAL PREDICTORS OF THE FIRST-PASS EFFECT (FPE) IN PATIENTS WITH ACUTE ISCHEMIC STROKE RECEIVING MECHANICAL THROMBECTOMY (MT). THE TIME FRAME OF THIS STUDY WAS: JANUARY 2019 TO DECEMBER 2021. MULTIPLE MANUFACTURER¿S DEVICES WERE USED IN THE STUDY POPULATION. THE STUDY MENTIONS VARIOUS MEDTRONIC THROMBECTOMY DEVICES INCLUDING STENT RETRIEVERS LIKE SOLITAIR FR AND SOLITAIR AB AS WELL AS CONTACT ASPIRATION USING THE REACT 68 CATHETER. DEATHS OCCURRED IN THE STUDY POPULATION. THE STUDY REPORTED A TOTAL OF 2 DEATHS (1.3% OF PATIENTS), WITH ONE DEATH IN THE FPE GROUP (2.1%) AND ONE IN THE NON-FPE GROUP (1.0%). HOWEVER, THE DOCUMENT DOES NOT SPECIFY THE CAUSES OF THESE DEATHS. AMONG PATIENT ADVERSE EVENTS INCLUDED: HIGHER RATES OF SUBARACHNOID HEMORRHAGE (SAH) IN NON-FPE PATIENTS (0% VS. 11.5%, P = 0.018). NON-FPE PATIENTS HAD A HIGHER RATE OF RESCUE THERAPY (BALLOON ANGIOPLASTY OR STENT IMPLANTATION), (0% VS. 13.5%, P = 0.012). HEIDELBERG BLEEDING CLASSIFICATION: 3C: 12 (7.9%) TOTAL, WITH 0% IN FPE AND 11.5% IN NON-FPE (P = 0.018). INTRACRANIAL HEMORRHAGE: 70 (46.3%) TOTAL, WITH 23 (48.9%) IN FPE AND 47 (45.2%) IN NON-FPE (P = 0.726). SYMPTOMATIC INTRACRANIAL HEMORRHAGE: 14 (9.3%) TOTAL, WITH 5 (10.6%) IN FPE AND 9 (8.7%) IN NON-FPE (P = 0.764). ASYMPTOMATIC INTRACRANIAL HEMORRHAGE: 56 (37.1%) TOTAL, WITH 18 (38.3%) IN FPE AND 38 (36.5%) IN NON-FPE (P = 0.857). CLINICAL OUTCOMES OF INTEREST: NON-FPE PATIENTS HAD A LONGER MEDIAN PROCEDURE TIME (145 MINUTES VS. 101 MINUTES, P < 0.001), A LOWER RATE OF 24-HOUR EARLY NEUROLOGIC IMPROVEMENT (16.3% VS. 31.9%, P = 0.034), AND A HIGHER PROBABILITY OF FUNCTIONAL DECLINE AT DISCHARGE (MRS 4¿5) (57.7% VS. 38.3%, P = 0.035). NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21728753&pc=DQY | 2029214-2025-00804 | 2023-11-03 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-03-30 04:00 | DQY | REACT CATHETER | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Intracranial Hemorrhage; Ischemia Stroke | K180715 |  | 1 | Event Description: ZHANG, J., LONG, L., LI, J., ZHANG, H., YAN, W., ABULIMITI, A., ABULAJIANG, N., LU, Q., NGUYEN, T. N., & CAI, X.; NEURORADIOLOGY; 2024; 66:1021¿1029; IMPACT OF OBESITY-RELATED INDICATORS ON FRST-PASS EFECT IN PATIENTS WITH ISCHEMIC STROKE RECEIVING MECHANICAL THROMBECTOMY; DOI.ORG/10.1007/S00234-024-03350-X LITERATURE WAS REVIEWED REGARDING: "IMPACT OF OBESITY-RELATED INDICATORS ON FIRST-PASS EFFECT IN PATIENTS WITH ISCHEMIC STROKE RECEIVING MECHANICAL THROMBECTOMY." THE OBJECTIVE IS TO EVALUATE OBESITY-RELATED INDICATORS AS POTENTIAL PREDICTORS OF THE FIRST-PASS EFFECT (FPE) IN PATIENTS WITH ACUTE ISCHEMIC STROKE RECEIVING MECHANICAL THROMBECTOMY (MT). THE TIME FRAME OF THIS STUDY WAS: JANUARY 2019 TO DECEMBER 2021. MULTIPLE MANUFACTURER¿S DEVICES WERE USED IN THE STUDY POPULATION. THE STUDY MENTIONS VARIOUS MEDTRONIC THROMBECTOMY DEVICES INCLUDING STENT RETRIEVERS LIKE SOLITAIR FR AND SOLITAIR AB AS WELL AS CONTACT ASPIRATION USING THE REACT 68 CATHETER. DEATHS OCCURRED IN THE STUDY POPULATION. THE STUDY REPORTED A TOTAL OF 2 DEATHS (1.3% OF PATIENTS), WITH ONE DEATH IN THE FPE GROUP (2.1%) AND ONE IN THE NON-FPE GROUP (1.0%). HOWEVER, THE DOCUMENT DOES NOT SPECIFY THE CAUSES OF THESE DEATHS. AMONG PATIENT ADVERSE EVENTS INCLUDED: HIGHER RATES OF SUBARACHNOID HEMORRHAGE (SAH) IN NON-FPE PATIENTS (0% VS. 11.5%, P = 0.018). NON-FPE PATIENTS HAD A HIGHER RATE OF RESCUE THERAPY (BALLOON ANGIOPLASTY OR STENT IMPLANTATION), (0% VS. 13.5%, P = 0.012). HEIDELBERG BLEEDING CLASSIFICATION: 3C: 12 (7.9%) TOTAL, WITH 0% IN FPE AND 11.5% IN NON-FPE (P = 0.018). INTRACRANIAL HEMORRHAGE: 70 (46.3%) TOTAL, WITH 23 (48.9%) IN FPE AND 47 (45.2%) IN NON-FPE (P = 0.726). SYMPTOMATIC INTRACRANIAL HEMORRHAGE: 14 (9.3%) TOTAL, WITH 5 (10.6%) IN FPE AND 9 (8.7%) IN NON-FPE (P = 0.764). ASYMPTOMATIC INTRACRANIAL HEMORRHAGE: 56 (37.1%) TOTAL, WITH 18 (38.3%) IN FPE AND 38 (36.5%) IN NON-FPE (P = 0.857). CLINICAL OUTCOMES OF INTEREST: NON-FPE PATIENTS HAD A LONGER MEDIAN PROCEDURE TIME (145 MINUTES VS. 101 MINUTES, P < 0.001), A LOWER RATE OF 24-HOUR EARLY NEUROLOGIC IMPROVEMENT (16.3% VS. 31.9%, P = 0.034), AND A HIGHER PROBABILITY OF FUNCTIONAL DECLINE AT DISCHARGE (MRS 4¿5) (57.7% VS. 38.3%, P = 0.035). NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21727322&pc=BZD | 2518422-2025-040046 | 2023-03-28 04:00 | Death | RESPIRONICS, INC. | 2025-03-29 04:00 | BZD | REMSTAR AUTO A-FLEX | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A REMSTAR AUTO A-FLEX DEVICE. THE PATIENT HAS PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A REMSTAR AUTO A-FLEX DEVICE. THE PATIENT HAS PASSED AWAY. THERE IS NO CUSTOMER INFORMATION HENCE WE CANNOT REACH OUT TO THE CUSTOMER AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. BOX H UPDATED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21727323&pc=BZD | 2518422-2025-040024 | 2024-02-21 05:00 | Death | RESPIRONICS, INC. | 2025-03-29 04:00 | BZD | DS2ADV AUTO CPAP | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K200480 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION REGARDING A DS2ADV AUTO CPAP DEVICE. THE PATIENT'S WIFE HAS ALLEGED HER HUSBAND PASSED AWAY. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER. THE TECHNICIAN UNABLE TO CONFIRM FOAM PARTICLES IN THE AIR PATH DURING THE DEVICE EVALUATION. THE DEVICE WAS SCRAPPED. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21727330&pc=BZD | 2518422-2025-040026 | 2024-06-11 04:00 | Death | RESPIRONICS, INC. | 2025-03-29 04:00 | BZD | REP DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS REPORT OF PATIENT DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21727343&pc=MNS | 2518422-2025-040044 | 2021-05-15 04:00 | Death | RESPIRONICS, INC. | 2025-03-29 04:00 | MNS | DREAMSTATION ST30 | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K102465 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH A DREAMSTATION ST30 DEVICE PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE DEVICE WAS RETURNED. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND NO VISIBLE FOAM PARTICLES WERE OBSERVED. ADDITIONALLY, NO FAULTS WERE FOUND AND ALL TESTS WERE PASSED. Manufacturer Narrative: THE BECOME AWARENESS DATE, COMPONENT CODE GRID AND CONCLUSION CODE GRID HAS BEEN UPDATED/CORRECTED IN THIS REPORT. THE DEVICE INFORMATION HAS ALSO BEEN UPDATED/CORRECTED. IN THIS REPORT, BOX D, G, H HAS BEEN UPDATED/CORRECTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21728393&pc=OAE | 2029046-2025-00965 |  | Death | BIOSENSE WEBSTER INC | 2025-03-29 04:00 | OAE | UNK_QDOT MICRO | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Fistula; Ventricular Fibrillation | P210027 |  | 1 | Event Description: THIS COMPLAINT IS FROM A LITERATURE SOURCE AND THE FOLLOWING CITATION WAS REVIEWED: KEWCHAROEN J, SHAH K, BHARDWAJ R, TURAGAM MK, CONTRACTOR T, MANDAPATI R, LAKKIREDDY D, GARG J. CATHETER-RELATED AND CLINICAL COMPLICATIONS ASSOCIATED WITH QDOT MICROTM ABLATION CATHETER. J CARDIOVASC ELECTROPHYSIOL. 2024 SEP;35(9):1874-1877. DOI: 10.1111/JCE.16363. EPUB 2024 JUL 8. PMID: 38978283. THE FIRST ARTICLE REPORTED IN THIS COMPLAINT IS RELATED TO ANOTHER ARTICLE THAT ALSO REPORTED CATHETER RELATED AND CLINICAL COMPLICATIONS ASSOCIATED WITH QDOT ABLATION CATHETER. THE DATES OVERLAP (THE FIRST ARTICLE DID A SEARCH FROM INCEPTION OF THE CATHETER-OCTOBER 2023, WHILE THE DATES FROM THE SECOND ARTICLE SPAN FROM NOV 2022 ¿ APRIL 2024). THEREFORE, TO CONSERVATIVELY REPORT THE DEVICE ISSUES AND ADVERSE EVENTS THAT WERE NOT CAPTURED IN THE ORIGINAL ARTICLE, THE FOLLOWING ADVERSE EVENTS ARE TO BE REPORTED: ADVERSE EVENTS: - 35 PERICARDIAL EFFUSIONS REQUIRING PERICARDIOCENTESIS OR SURGERY. - 4 PATIENT HAD A STROKE, NO REPORTED INTERVENTION. - 6 PATIENTS HAD PERICARDITIS, NO REPORTED INTERVENTION. - 3 ESOPHAGEAL FISTULAS. - 6 PHRENIC NERVE PALSY, ALL PERSISTENT AT THE TIME OF THE EVENT REPORTING, NONE REQUIRED INTERVENTION. - 1 PATIENTS WITH VENTRICULAR FIBRILLATION. - 4 PATIENTS HAD SIGNIFICANT AV BLOCK THAT REQUIRED PERMANENT PACEMAKERS. - 3 PATIENTS WITH PULMONARY VEIN STENOSIS. NONSERIOUS EVENTS: - 2 PERICARDIAL EFFUSIONS THAT DID NOT REQUIRE INTERVENTION. DEATHS: - 9 DEATHS TOTAL. - 4 OTHER "ESOPHAGEAL COMPLICATIONS" (ESOPHAGEAL FISTULAS). - 4 RELATED TO CARDIAC ARREST. - 1 VENTRICULAR FIBRILLATION / CARDIAC ARREST. REPORTABLE MALFUNCTIONS: - 11 PATIENTS HAD CATHETER THROMBOSIS, NO REPORTED INTERVENTION. - 16 IRRIGATION ISSUE. MALFUNCTIONS: - 18 PATIENTS HAD CHAR FORMATION FOR WHICH 50% OF THE EVENTS REPORTED USE OF THE VHPSD STRATEGY. - 148 PHYSICAL CATHETER DAMAGE. - 106 COMMUNICATION ISSUE / DISPLAY ISSUE. - 12 SIGNAL ISSUE / ARTIFACT. - 3 INACCURATE TEMPERATURE. THIS REPORT IS BEING SENT TO ADDRESS THE "DEATHS" SECTION ONLY WITH 9 DEATHS (4 ESOPHAGEAL COMPLICATIONS, 4 CARDIAC ARRESTS, 1 VENTRICULAR FIBRILLATION). Manufacturer Narrative: AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. AN EVALUATION OF THE MANUFACTURING RECORD COULD NOT BE PERFORMED AS THE REQUIRED PRODUCT IDENTIFICATION NUMBER WAS NOT PROVIDED TO COMPLETE THE EVALUATION. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. HOWEVER, IF THE PRODUCT OR PRODUCT ID NUMBERS ARE RECEIVED AT A LATER DATE, THE INVESTIGATION WILL BE UPDATED AS APPLICABLE. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. MANUFACTURER'S REF. NO: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21716836&pc=FTL | 1213643-2025-00228 | 2025-03-06 05:00 | Death | DAVOL INC., SUB. C.R. BARD, INC. -1213643 | 2025-03-28 04:00 | FTL | MESH ¿ COMPOSIX KUGEL | Defective Device; Patient-Device Incompatibility | Pain; Insufficient Information | K061314 |  | 1 | Event Description: ATTORNEY ALLEGES THAT THE PATIENT UNDERWENT SURGERY FOR IMPLANT OF UNSPECIFIED BARD/DAVOL COMPOSIX KUGEL HERNIA PATCH ON (B)(6) 2004, (B)(6) 2005 AND (B)(6) 2006. AS REPORTED, THE PATIENT IS MAKING A CLAIM FOR AN ADVERSE PATIENT OUTCOME AGAINST ALL DEVICES. ATTORNEY ALLEGES WRONGFUL DEATH OF THE PATIENT AND GENERAL ALLEGATIONS FOR "PAST, PRESENT, AND FUTURE DAMAGES, INCLUDING BUT NOT LIMITED TO, MENTAL AND PHYSICAL PAIN AND SUFFERING FOR SEVERE AND PERMANENT PERSONAL INJURIES SUSTAINED BY THE PATIENT." IT IS ALSO ALLEGED THAT THE PATIENT EXPERIENCED EMOTIONAL DISTRESS AND THE DEVICE WAS DEFECTIVE. Manufacturer Narrative: NO CONCLUSIONS CAN BE MADE. THE PATIENT'S ATTORNEY ALLEGES ADVERSE PATIENT OUTCOME ASSOCIATED WITH THE HERNIA MESH USED TO TREAT THE PATIENT INCLUDING DEATH AND "PAST, PRESENT, AND FUTURE DAMAGES, INCLUDING BUT NOT LIMITED TO, MENTAL AND PHYSICAL PAIN AND SUFFERING FOR SEVERE AND PERMANENT PERSONAL INJURIES" SUSTAINED BY THE PLAINTIFF. NO MEDICAL RECORDS, AUTOPSY REPORT, OR DEATH CERTIFICATE HAVE BEEN PROVIDED. NO LOT NUMBER HAS BEEN PROVIDED; THEREFORE, A REVIEW OF THE MANUFACTURING RECORDS IS NOT POSSIBLE. NOTE, THE DATE OF EVENT (06-MAR-2025) AND DATE OF DEATH ((B)(6) 2025) CONSIDERED TO BE A BEST ESTIMATE. THIS EMDR REPRESENTS THE BARD/DAVOL MESH ¿ COMPOSIX KUGEL (DEVICE #1). ADDITIONAL EMDR'S WERE SUBMITTED TO REPRESENT THE BARD/DAVOL MESH ¿ COMPOSIX KUGEL (DEVICE #2, DEVICE #3 & DEVICE #4). SHOULD ADDITIONAL INFORMATION BE PROVIDED, A SUPPLEMENTAL EMDR WILL BE SUBMITTED. NOTE: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21716837&pc=FTL | 1213643-2025-00229 | 2025-03-06 05:00 | Death | DAVOL INC., SUB. C.R. BARD, INC. -1213643 | 2025-03-28 04:00 | FTL | MESH ¿ COMPOSIX KUGEL | Defective Device; Patient-Device Incompatibility | Pain; Insufficient Information | K061314 |  | 1 | Event Description: ATTORNEY ALLEGES THAT THE PATIENT UNDERWENT SURGERY FOR IMPLANT OF UNSPECIFIED BARD/DAVOL COMPOSIX KUGEL HERNIA PATCH ON (B)(6) 2004, (B)(6) 2005, AND (B)(6) 2006. AS REPORTED, THE PATIENT IS MAKING A CLAIM FOR AN ADVERSE PATIENT OUTCOME AGAINST ALL DEVICES. ATTORNEY ALLEGES WRONGFUL DEATH OF THE PATIENT AND GENERAL ALLEGATIONS FOR "PAST, PRESENT, AND FUTURE DAMAGES, INCLUDING BUT NOT LIMITED TO, MENTAL AND PHYSICAL PAIN AND SUFFERING FOR SEVERE AND PERMANENT PERSONAL INJURIES SUSTAINED BY THE PATIENT." IT IS ALSO ALLEGED THAT THE PATIENT EXPERIENCED EMOTIONAL DISTRESS AND THE DEVICE WAS DEFECTIVE. Manufacturer Narrative: NO CONCLUSIONS CAN BE MADE. THE PATIENT'S ATTORNEY ALLEGES ADVERSE PATIENT OUTCOME ASSOCIATED WITH THE HERNIA MESH USED TO TREAT THE PATIENT INCLUDING DEATH AND "PAST, PRESENT, AND FUTURE DAMAGES, INCLUDING BUT NOT LIMITED TO, MENTAL AND PHYSICAL PAIN AND SUFFERING FOR SEVERE AND PERMANENT PERSONAL INJURIES" SUSTAINED BY THE PLAINTIFF. NO MEDICAL RECORDS, AUTOPSY REPORT, OR DEATH CERTIFICATE HAVE BEEN PROVIDED. NO LOT NUMBER HAS BEEN PROVIDED; THEREFORE, A REVIEW OF THE MANUFACTURING RECORDS IS NOT POSSIBLE. NOTE, THE DATE OF EVENT ((B)(6) 2025) AND DATE OF DEATH ((B)(6) 2025) CONSIDERED TO BE A BEST ESTIMATE. THIS EMDR REPRESENTS THE BARD/DAVOL MESH ¿ COMPOSIX KUGEL (DEVICE #3). ADDITIONAL EMDR'S WERE SUBMITTED TO REPRESENT THE BARD/DAVOL MESH ¿ COMPOSIX KUGEL (DEVICE #1, DEVICE #2 & DEVICE #4). SHOULD ADDITIONAL INFORMATION BE PROVIDED, A SUPPLEMENTAL EMDR WILL BE SUBMITTED. NOTE: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21716838&pc=FTL | 1213643-2025-00227 | 2025-03-06 05:00 | Death | DAVOL INC., SUB. C.R. BARD, INC. -1213643 | 2025-03-28 04:00 | FTL | MESH ¿ COMPOSIX KUGEL | Defective Device; Patient-Device Incompatibility | Pain; Insufficient Information | K061314 |  | 1 | Event Description: ATTORNEY ALLEGES THAT THE PATIENT UNDERWENT SURGERY FOR IMPLANT OF UNSPECIFIED BARD/DAVOL COMPOSIX KUGEL HERNIA PATCH ON (B)(6) AND (B)(6) 2006. AS REPORTED, THE PATIENT IS MAKING A CLAIM FOR AN ADVERSE PATIENT OUTCOME AGAINST ALL DEVICES. ATTORNEY ALLEGES WRONGFUL DEATH OF THE PATIENT AND GENERAL ALLEGATIONS FOR "PAST, PRESENT, AND FUTURE DAMAGES, INCLUDING BUT NOT LIMITED TO, MENTAL AND PHYSICAL PAIN AND SUFFERING FOR SEVERE AND PERMANENT PERSONAL INJURIES SUSTAINED BY THE PATIENT. " IT IS ALSO ALLEGED THAT THE PATIENT EXPERIENCED EMOTIONAL DISTRESS AND THE DEVICE WAS DEFECTIVE. Manufacturer Narrative: NO CONCLUSIONS CAN BE MADE. THE PATIENT'S ATTORNEY ALLEGES ADVERSE PATIENT OUTCOME ASSOCIATED WITH THE HERNIA MESH USED TO TREAT THE PATIENT INCLUDING DEATH AND "PAST, PRESENT, AND FUTURE DAMAGES, INCLUDING BUT NOT LIMITED TO, MENTAL AND PHYSICAL PAIN AND SUFFERING FOR SEVERE AND PERMANENT PERSONAL INJURIES" SUSTAINED BY THE PLAINTIFF. NO MEDICAL RECORDS, AUTOPSY REPORT, OR DEATH CERTIFICATE HAVE BEEN PROVIDED. NO LOT NUMBER HAS BEEN PROVIDED; THEREFORE, A REVIEW OF THE MANUFACTURING RECORDS IS NOT POSSIBLE. NOTE, THE DATE OF EVENT (06-MAR-2025) AND DATE OF DEATH (B)(6) 2025) CONSIDERED TO BE A BEST ESTIMATE. THIS EMDR REPRESENTS THE BARD/DAVOL MESH ¿ COMPOSIX KUGEL (DEVICE #2). ADDITIONAL EMDR'S WERE SUBMITTED TO REPRESENT THE BARD/DAVOL MESH ¿ COMPOSIX KUGEL (DEVICE #1, DEVICE #3 & DEVICE #4). SHOULD ADDITIONAL INFORMATION BE PROVIDED, A SUPPLEMENTAL EMDR WILL BE SUBMITTED. NOTE: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21716841&pc=FTL | 1213643-2025-00230 | 2025-03-06 05:00 | Death | DAVOL INC., SUB. C.R. BARD, INC. -1213643 | 2025-03-28 04:00 | FTL | MESH ¿ COMPOSIX KUGEL | Defective Device; Patient-Device Incompatibility | Pain; Insufficient Information | K061314 |  | 1 | Event Description: ATTORNEY ALLEGES THAT THE PATIENT UNDERWENT SURGERY FOR IMPLANT OF UNSPECIFIED BARD/DAVOL COMPOSIX KUGEL HERNIA PATCH ON (B)(6) 2004, (B)(6) 2005 AND (B)(6) 2006. AS REPORTED, THE PATIENT IS MAKING A CLAIM FOR AN ADVERSE PATIENT OUTCOME AGAINST ALL DEVICES. ATTORNEY ALLEGES WRONGFUL DEATH OF THE PATIENT AND GENERAL ALLEGATIONS FOR "PAST, PRESENT, AND FUTURE DAMAGES, INCLUDING BUT NOT LIMITED TO, MENTAL AND PHYSICAL PAIN AND SUFFERING FOR SEVERE AND PERMANENT PERSONAL INJURIES SUSTAINED BY THE PATIENT." IT IS ALSO ALLEGED THAT THE PATIENT EXPERIENCED EMOTIONAL DISTRESS AND THE DEVICE WAS DEFECTIVE. THIS FILE REPRESENTS DEVICE #4. Manufacturer Narrative: NO CONCLUSIONS CAN BE MADE. THE PATIENT'S ATTORNEY ALLEGES ADVERSE PATIENT OUTCOME ASSOCIATED WITH THE HERNIA MESH USED TO TREAT THE PATIENT INCLUDING DEATH AND "PAST, PRESENT, AND FUTURE DAMAGES, INCLUDING BUT NOT LIMITED TO, MENTAL AND PHYSICAL PAIN AND SUFFERING FOR SEVERE AND PERMANENT PERSONAL INJURIES" SUSTAINED BY THE PLAINTIFF. NO MEDICAL RECORDS, AUTOPSY REPORT, OR DEATH CERTIFICATE HAVE BEEN PROVIDED. NO LOT NUMBER HAS BEEN PROVIDED; THEREFORE, A REVIEW OF THE MANUFACTURING RECORDS IS NOT POSSIBLE. NOTE, THE DATE OF EVENT (06-MAR-2025) AND DATE OF DEATH ((B)(6) 2025) CONSIDERED TO BE A BEST ESTIMATE. THIS EMDR REPRESENTS THE BARD/DAVOL MESH ¿ COMPOSIX KUGEL (DEVICE #4). ADDITIONAL EMDR'S WERE SUBMITTED TO REPRESENT THE BARD/DAVOL MESH ¿ COMPOSIX KUGEL (DEVICE #1, DEVICE #2 & DEVICE #3). SHOULD ADDITIONAL INFORMATION BE PROVIDED, A SUPPLEMENTAL EMDR WILL BE SUBMITTED. NOTE: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21717125&pc=DSP | 2248146-2025-0000183 | 2025-03-05 05:00 | Death | DATASCOPE CORP. | 2025-03-28 04:00 | DSP | SENSATION PLUS 8FR. 50CC IAB & ACCESSORIES (W/PRESSURE TUBES, NO STYLET) | Break; Appropriate Term/Code Not Available | Insufficient Information | K133074 |  | 1 | Event Description: A SUSPECTED TEAR IN THE INTRA-AORTIC BALLOON OF THE INTRA-AORTIC BALLOON PUMP (IABP) WAS REPORTED IN ASSOCIATION WITH A PATIENT'S DEATH. Manufacturer Narrative: A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF OUR INVESTIGATION. Manufacturer Narrative: UPDATED FIELDS: B4,D9, G3, G6, H2, H3,H6 (TYPE OF INVESTIGATION, INVESTIGATION FINDINGS, COMPONENT CODES, INVESTIGATION CONCLUSION), H11. CORRECTED DEVICE:D7A,H6 (MEDICAL DEVICE-PROBLEM CODE). NO DECON OR VISUAL INSPECTION PERFORMED SINCE PRODUCT WAS NOT RETURNED AND COULD NOT BE EVALUATED. THEREFORE, WE WERE UNABLE TO CONFIRM OR DETERMINE A ROOT CAUSE FOR THE REPORTED FAILURE. THE FAILURE MODE IS ADDRESSED IN THE RISK FILE AND IS OPERATING WITHIN ITS RISK PROFILE. THE IFU ADDRESSES THE REPORTED FAILURE. THERE WERE NO NCMRS IDENTIFIED WHICH COULD CAUSE OR CONTRIBUTE TO THE REPORTED FAILURE. THE INVESTIGATION DOES NOT INDICATE THAT THE DEVICE WAS INADVERTENTLY RELEASED AS NON-CONFORMING OR AN ADULTERATED PRODUCT OR WAS A COUNTERFEIT. THE COMPLAINT HISTORY REVIEW DID NOT IDENTIFY AN ADVERSE TREND (INCREASE IN NUMBER OF COMPLAINTS OVER PAST THREE (3) MONTHS). BASED ON THE RATIONAL PROVIDED ABOVE, NO ESCALATION TO THE CAPA PROCESS IS REQUIRED. COMPLAINT# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21717736&pc=FSA | 1419652-2025-00018 | 2025-01-01 05:00 | Death | ARJOHUNTLEIGH MAGOG INC. | 2025-03-28 04:00 | FSA | MAXI MOVE | Insufficient Information | Physical Entrapment |  |  | 1 | Event Description: THE EVENT OCCURRED IN JANUARY WHEN A PATIENT WAS RECOVERING AFTER ABDOMINAL SURGERY. ACCORDING TO THE CUSTOMER, A NURSE USED THE DPS FUNCTION TO LOWER THE SPREADER BAR BUT ALLEGEDLY STOPPED TOO LATE, CAUSING THE SPREADER BAR TO PRESS AGAINST THE PATIENT'S ABDOMEN. THE PATIENT PASSED AWAY 24 HOURS LATER. CURRENTLY, THERE IS NO INFORMATION INDICATING ANY FAILURES OF THE ARJO DEVICE. THE CUSTOMER INSPECTED THE DEVICE INTERNALLY AFTER THE EVENT. Manufacturer Narrative: THE DATE OF THE EVENT (SECTION B3) IS THE ESTIMATED DATE BASED ON THE INFORMATION PROVIDED BY THE CUSTOMER THAT THE EVENT OCCURRED IN (B)(6) 2025. UNIQUE IDENTIFIER (UDI) - SECTION D4: ((B)(4). THE DEVICE IS DISTRIBUTED OUTSIDE THE US AND NO GUDID INFORMATION EXISTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21717740&pc=FSA | 9681684-2025-00030 | 2024-01-01 05:00 | Death | ARJOHUNTLEIGH MAGOG INC. | 2025-03-28 04:00 | FSA | MAXI MOVE | Use of Device Problem; Insufficient Information | Physical Entrapment |  |  | 1 | Event Description: THE EVENT OCCURRED IN JANUARY WHEN A PATIENT WAS RECOVERING AFTER ABDOMINAL SURGERY. ACCORDING TO THE CUSTOMER, A NURSE USED THE DPS FUNCTION TO LOWER THE SPREADER BAR BUT ALLEGEDLY STOPPED TOO LATE, CAUSING THE SPREADER BAR TO PRESS AGAINST THE PATIENT'S ABDOMEN. THE PATIENT PASSED AWAY 24 HOURS LATER. CURRENTLY, THERE IS NO INFORMATION INDICATING ANY FAILURES OF THE ARJO DEVICE. THE CUSTOMER INSPECTED THE DEVICE INTERNALLY AFTER THE EVENT.Event Description: THE EVENT OCCURRED IN JANUARY 2024 (OVER A YEAR BEFORE ARJO WAS NOTIFIED) WHEN A PATIENT WAS RECOVERING AFTER ABDOMINAL SURGERY. ACCORDING TO THE CUSTOMER, A NURSE USED THE CONTROL BUTTON TO LOWER THE SPREADER BAR BUT ALLEGEDLY STOPPED ¿TOO LATE¿, CAUSING THE SPREADER BAR TO PRESS AGAINST THE PATIENT'S ABDOMEN. THE PATIENT PASSED AWAY 24 HOURS LATER, AND THE SURGEON CONFIRMED THAT THIS WAS A RESULT OF THE EVENT. Manufacturer Narrative: FOLLOWING A REPORTED INCIDENT INVOLVING THE MAXI MOVE LIFT, THE CUSTOMER CONDUCTED AN INTERNAL INSPECTION OF THE DEVICE. NO FAILURE OF THE ARJO DEVICE WAS IDENTIFIED. ALTHOUGH ARJO WAS INFORMED OF THE INCIDENT APPROXIMATELY ONE YEAR AFTER IT OCCURRED, AN ARJO TECHNICIAN VISITED THE CUSTOMER SITE TO PERFORM A THOROUGH INSPECTION OF THE DEVICE. THE INSPECTION CONFIRMED THAT THE LIFT WAS OPERATING CORRECTLY, WITH NO MECHANICAL OR FUNCTIONAL FAILURE DETECTED. IT REMAINS UNKNOWN WHETHER ANY REPAIRS WERE PERFORMED BY THE CUSTOMER BETWEEN THE INCIDENT AND ARJO¿S INSPECTION. THE PRODUCT INSTRUCTIONS FOR USE (IFU) FOR MAXI MOVE (001.25060.EN) STATES THE FOLLOWING: ¿TAKE CARE NOT TO LOWER THE SPREADER BAR ONTO THE PATIENT TO AVOID INJURIES.¿ IN THE COMPLAINT AT HAND, THE SPREADER BAR WAS LOWERED ONTO THE PATIENT CAUSING IT TO PRESS ON THE ABDOMEN. THE MAXI MOVE LIFT IS EQUIPPED WITH AN ANTI-CRUSH DETECTION SYSTEM, DESIGNED TO AUTOMATICALLY ENGAGE WHEN THE SPREADER BAR ENCOUNTERS AN OBSTRUCTION. ACCORDING TO THE IFU: ¿THE SYSTEM WILL ENGAGE, STOP THE MOTOR AND ALL DOWNWARD MOVEMENT WILL CEASE. IF THIS OCCURS, RELEASE THE JIB 'DOWN' BUTTON IMMEDIATELY AND PRESS THE JIB 'UP' BUTTON TO RAISE THE JIB UNTIL LIFT IS CLEAR. THEN REMOVE THE OBSTRUCTION.¿ NO ISSUES WITH THE ANTI-CRUSH FEATURE WERE REPORTED BY THE CUSTOMER OR IDENTIFIED DURING ARJO¿S EVALUATION. WHEN ANTI-CRUSH SENSES THE OBSTACLE, IT WILL STOP LOWERING. AS PER THE IFU, THE CAREGIVER THEN NEEDS TO RAISE THE JIB (MAST) USING A CONTROL BUTTON TO REMOVE THE OBSTRUCTION. IT IS UNKNOWN IF AND WHEN THE SPREADER BAR WAS REMOVED. TO SUMMARIZE, THE ARJO DEVICE USED DURING THE INCIDENT MET THE SPECIFICATIONS. THE LIFT WAS FOUND TO BE FULLY FUNCTIONAL, WITH NO EVIDENCE OF MALFUNCTION THAT COULD HAVE CONTRIBUTED TO THE PATIENT¿S DEATH. THE INCIDENT WAS ATTRIBUTED TO A FAILURE TO FOLLOW THE IFU, SPECIFICALLY THE INSTRUCTION TO AVOID LOWERING THE SPREADER BAR ONTO THE PATIENT. THE COMPLAINT WAS CLASSIFIED AS REPORTABLE DUE TO THE PATIENT¿S DEATH, WHICH OCCURRED FOLLOWING THE APPLICATION OF PRESSURE TO THE POST-OPERATIVE ABDOMINAL AREA BY THE SPREADER BAR. THIS APPEARS TO BE AN ISOLATED INCIDENT. Manufacturer Narrative: THE DATE OF EVENT IN SECTION B3 WAS CORRECTED AS THE YEAR WAS INCORRECTLY PROVIDED IN THE INITIAL REPORT. THE INVESTIGATION IS STILL ONGOING. THE RESULTS WILL BE PROVIDED IN THE NEXT REPORT. Manufacturer Narrative: THE INFORMATION GATHERING PROCESS IS ONGOING. RESULTS WILL BE PROVIDED UPON COMPLETION OF THE INVESTIGATION. THE DATE OF THE EVENT (SECTION B3) IS THE ESTIMATED DATE BASED ON THE INFORMATION PROVIDED BY THE CUSTOMER THAT THE EVENT OCCURRED IN JANUARY 2025. UNIQUE IDENTIFIER (UDI) - SECTION D4:(B)(4). THE DEVICE IS DISTRIBUTED OUTSIDE THE US AND NO GUDID INFORMATION EXISTS. EVALUATED INTERNALLY BY THE CUSTOMER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21717811&pc=PYX | 1220648-2025-27020 | 2024-12-13 05:00 | Death | ABIOMED, INC. | 2025-03-28 04:00 | PYX | IMPELLA RP FLEX | Adverse Event Without Identified Device or Use Problem | Perforation of Vessels | P170011 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA RP FLEX FOR SUPPORT OF ST ELEVATED MYOCARDIAL INFARCTION AND RIGHT VENTRICLE (RV) FAILURE. THE PATIENT WAS TAKEN TO THE OPERATING ROOM AND HAD THE IMPELLA RP FLEX PLACED BUT, THE PATIENT EXPIRED DUE TO PROLONGED DOWN TIME AND STOPPING THE RESUSCITATION EFFORTS. THE LONG-TERM OUTCOME AND QUALITY INDICATOR (LOQI) DATABASE LATER REPORTED UPON A RV PERFORATION WITH A "POSSIBLY" RELATEDNESS TO THE PUMP. THE PERFORATION WAS FOUND AFTER EMERGENCY STERNOTOMY AND THE NATIVE TISSUE WAS POOR AND UNABLE TO BE REPAIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA RP FLEX WITH SMARTASSIST SYSTEM WITH AUTOMATED IMPELLA CONTROLLER SECTION: WARNINGS & CAUTIONS: CAUTIONS ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ARRHYTHMIA, ATRIAL FIBRILLATION, BLEEDING, CARDIAC TAMPONADE, CARDIOGENIC SHOCK, DEATH, DEVICE MALFUNCTION, HEMOLYSIS, HEPATIC FAILURE, INSERTION SITE INFECTION, PERFORATION, PHLEGMASIA CERULEA DOLENS (A SEVERE FORM OF DEEP VENOUS THROMBOSIS), PULMONARY VALVE INSUFFICIENCY, RESPIRATORY DYSFUNCTION, SEPSIS, THROMBOCYTOPENIA, THROMBOTIC VASCULAR (NON-CENTRAL NERVOUS SYSTEM COMPLICATION, TRICUSPID VALVE INJURY, CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION), VENOUS THROMBOSIS, VENTRICULAR FIBRILLATION AND/OR TACHYCARDIA.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿DO NOT ADVANCE OR WITHDRAW THE IMPELLA RP FLEX WITH SMARTASSIST SYSTEM CATHETER AGAINST RESISTANCE WITHOUT USING FLUOROSCOPY TO DETERMINE THE CAUSE OF THE RESISTANCE. DOING SO COULD RESULT IN SEPARATION OF THE CATHETER OR GUIDEWIRE TIP, DAMAGE TO THE CATHETER OR VESSEL, OR PERFORATION.¿ Manufacturer Narrative: THE INVESTIGATION OF VENTRICLE PERFORATION HAS BEEN COMPLETED. THE CAUSE OF THE VENTRICLE PERFORATION WAS NOT DETERMINED SINCE PRODUCT WAS NOT RETURNED AND LACK OF CLINICAL DETAILS. B3 DATE OF EVENT WAS ERRONEOUSLY ENTERED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27020.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21717933&pc=NGV | 21717933 | 2025-03-24 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-03-28 04:00 | NGV | WATCHMAN FLX PRO |  | Insufficient Information |  |  | 1 | Event Description: 31 MM WATCHMAN FLEX PRO DEVICE WAS DEPLOYED IN THE LEFT ANTERIOR APPENDAGE AND AFTER IT WAS RELEASED, IT FLEW FROM THE APPENDAGE INTO THE LV [LEFT VENTRICLE] AND MULTIPLE ATTEMPTS TO SNARE IT WERE UNSUCCESSFUL, SO PATIENT WENT TO THE OR [OPERATING ROOM] EMERGENTLY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21718068&pc=LOX | 9612164-2025-01568 | 2025-03-23 04:00 | Death | MEDTRONIC MEXICO | 2025-03-28 04:00 | LOX | NC EUPHORA RX | Deflation Problem; Difficult to Remove; Material Deformation; Material Split, Cut or Torn | Low Blood Pressure/ Hypotension; Myocardial Infarction; Pain; Cardiogenic Shock; Obstruction/Occlusion; Vascular Dissection; Thrombosis/Thrombus; Insufficient Information | K141090 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE INVOLVING ONE NC EUPHORA RX BALLOON CATHETER, BALLOON DEFLATION DIFFICULTIES WERE ENCOUNTERED, AND THE BALLOON WOULD NOT DEFLATE AT THE LESION SITE IN THE LEFT ANTERIOR DESCENDING (LAD) ARTERY. THE ELDERLY PATIENT HAD PRESENTED WITH AN ACUTE MYOCARDIAL INFARCTION AND OCCLUSION OF THE LAD. A GUIDEWIRE, THEN A PRE-DILATION BALLOON, THEN A CORONARY STENT HAD ALL BEEN INSERTED WITHOUT DIFFICULTY PRIOR TO THE ISSUE BEING ENCOUNTERED. THE NC EUPHORA BALLOON WAS BEING USED FOR POST-DILATION AND WAS INFLATED TO 16 ATM WITH THE USUAL CONTRAST SOLUTION AND SALINE SOLUTION USED. NEGATIVE PREP WAS NOT PERFORMED. THE LESION WAS PRE-DILATED. THE BALLOON DID PASS THROUGH A PREVIOUSLY DEPLOYED STENT. RESISTANCE WAS NOT ENCOUNTERED WHEN ADVANCING THE DEVICE, AND NO EXCESSIVE FORCE WAS USED DURING DELIVERY. THE BALLOON FAILED TO DEFLATE DESPITE SEVERAL ATTEMPTS AND A CHANGE OF PUMP. THE BALLOON WAS OCCLUDING THE ARTERY RENDERING IT DIFFICULT TO REMOVE. EVENTUALLY IT WAS MANAGED TO REMOVE THE STILL-INFLATED BALLOON FROM THE PATIENT. THE PATIENT SUBSEQUENTLY DIED. Manufacturer Narrative: IMAGE REVIEW: TWO STILL IMAGES RECEIVED FOR ANALYSIS. IMAGES DEPICT AN INFLATED BALLOON LOADED IN A GUIDE CATHETER BEING HELD IN A GLOVED HAND. NECKING IS EVIDENT TO THE INFLATION/DEFLATION LUMEN PROXIMAL TO THE BALLOON. IMAGES CONFIRM THE REPORTED DEFLATION AND REMOVAL DIFFICULTIES ADDITIONAL INFORMATION: THE PATIENT HAD PRESENTED TO THE EMERGENCY DEPARTMENT BY AMBULANCE DUE TO OPPRESSIVE CHEST PAIN THAT INCREASED IN INTENSITY OVER TIME, RADIATING TO THE LEFT ARM. ECG FOUND CORONARY SYNDROME WITH ST-SEGMENT ELEVATION AND THE PATIENT WAS REFERRED TO CARDIOLOGY. THE PATIENT WAS TACHYCARDIAC AND HYPERTENSIVE WITH BORDERLINE OXIMETRY UP TO 87% ON ONE OCCASION. THE PATIENT WAS SUPPORTED WITH OXYGEN THROUGH THE NOSE AND NITROGLYCERIN INFUSION WAS STARTED AT 60 RANGES/MINUTE. A STENT WAS PLACED ACQUIRING TIMI III FLOW. THE BODY OF THE NC EUPHORA BALLOON WAS CUT IN AN ATTEMPT TO REMOVE IT, BUT WITHOUT SUCCESS. CATHETER REMOVAL WAS ATTEMPTED, WITHOUT SUCCESS. THE PATIENT PRESENTS WITH PERSISTENT PAIN AND HYPOTENSION. CODE BLUE WAS ACTIVATED, SUPPORTED WITH NORADRENALINE AND THE PATIENT WAS SEDATED. A SECOND FEMORAL APPROACH WAS PERFORMED AND 2 GUIDEWIRES WERE ATTEMPTED TO BE ADVANCED TO THE ANTERIOR DESCENDING ARTERY WITHOUT SUCCESS. USING A THIRD GUIDEWIRE IT WAS POSSIBLE TO REACH THE DISTAL ARTERY, THE ARTERY WAS DILATED THREE TIMES, THE STUCK BALLOON WAS REMOVED BY RADIAL ROUTE. WHEN OBSERVED DIRECTLY, THE BALLOON REMAINS PERMANENTLY INFLATED. THE PATIENT HAS SIGNIFICANT HEMODYNAMIC COMPROMISE, DOSAGE OF VASOACTIVE DRUGS WAS INCREASED. DISSECTION OF THE ANTERIOR DESCENDING ARTERY PROXIMAL TO STENT WAS OBSERVED WITH THE PRESENCE OF THROMBUS. A NEW STENT WAS PLACED, WITHOUT ACHIEVING REPERFUSION. ADVANCED RESUSCITATION MEASURES WERE INITIATED AND AFTER 30 MINUTES WITHOUT ACHIEVING RHYTHM OR PULSE, DEATH WAS CONFIRMED MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT ANALYSIS: THE DEVICE RETURNED LOADED IN A GUIDE CATHETER WITH A GUIDEWIRE LOADED, IT WAS NOT POSSIBLE TO REMOVE THE DEVICE FROM THE GUIDE CATHETER AS THE BALLOON WAS STILL PARTIALLY INFLATED. THE LUER HAD BEEN CUT OFF PRIOR TO DEVICE RETURN. NECKING WAS EVIDENT TO THE INFLATION/DEFLATION LUMEN. IT WAS NOT POSSIBLE TO PERFORM DEFLATION TESTING DUE TO THE CUT LUER. NO OTHER DEFORMATION EVIDENT TO THE REMAINDER OF THE DEVICE. CORRECTION: IME CODES MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: THE PATIENT HAD AN ACUTE ST-SEGMENT ELEVATION MYOCARDIAL INFARCTION (STEMI). THE LESION WAS NON-TORTUOUS AND NON-CALCIFIED. THE NC EUPHORA BALLOON WAS NOT USED TO PRE-DILATE THE LESION. THE DEVICE WAS INSPECTED BEFORE USE. IT WAS NOT DIFFICULT TO REMOVE THE PROTECTIVE SHEATH OR THE PACKAGING STYLETTE. THE DEFLATION DIFFICULTIES WERE NOTED AFTER THE FIRST INFLATION. THERE WERE NO DIFFICULTIES NOTED DURING INFLATION OF THE DEVICE. THE DEVICE WAS NOT MOVED OR REPOSITIONED WHILE INFLATED. DISSECTION OF THE ANTERIOR DESCENDING ARTERY PROXIMAL TO STENT WAS OBSERVED WITH THE PRESENCE OF A SMALL QUANTITY OF THROMBUS. THERE WAS NO EVIDENCE OF RESTENOSIS. DEATH WAS CONFIRMED DURING THE PROCEDURE. PER THE PHYSICIAN THE CAUSE OF DEATH WAS MYOCARDIAL INFARCTION, CARDIOGENIC SHOCK AND THE INABILITY TO DEFLATE THE BALLOON. RELEVANT COMORBIDITIES THAT CONTRIBUTED TO THE PATIENT'S DEATH WERE HYPERTENSION, RHEUMATOID ARTHRITIS AND FRAILTY. THE PATIENT WAS ON DAPT AT TIME OF THE EVENT. SECTION B7.RELEVANT HISTORY. CORRECTION PRODUCT ANALYSIS: KINKS WERE EVIDENT TO THE HYPO-TUBE. CORRECTION: ANNEX A CODES. ANNEX E CODES: E2330 <(>&<)> E2321. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: ANNEX D CODE. CORRECTION: CORONARY ANGIOGRAPHY WAS PERFORMED WHICH SHOWED COMPLETE OCCLUSION OF THE LEFT ANTERIOR DESCENDING ARTERY. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21718126&pc=FTL | 1213643-2025-00246 | 2025-03-06 05:00 | Death | DAVOL INC., SUB. C.R. BARD, INC. -1213643 | 2025-03-28 04:00 | FTL | BARD FLAT MESH | Defective Device; Patient-Device Incompatibility | Pain; Insufficient Information | PREAMENDMENT |  | 1 | Event Description: ATTORNEY ALLEGES THAT THE PATIENT UNDERWENT SURGERY FOR IMPLANT OF AN UNSPECIFIED BARD FLAT MESH ON (B)(6) 2001. AS REPORTED, THE PATIENT IS MAKING A CLAIM FOR AN ADVERSE PATIENT OUTCOME AGAINST THE BARD FLAT MESH. ATTORNEY ALLEGES WRONGFUL DEATH OF THE PATIENT AND GENERAL ALLEGATIONS FOR "PAST, PRESENT, AND FUTURE DAMAGES, INCLUDING BUT NOT LIMITED TO, MENTAL AND PHYSICAL PAIN AND SUFFERING FOR SEVERE AND PERMANENT PERSONAL INJURIES SUSTAINED BY THE PATIENT." IT IS ALSO ALLEGED THAT THE PATIENT EXPERIENCED EMOTIONAL DISTRESS AND THE DEVICE WAS DEFECTIVE. Manufacturer Narrative: NO CONCLUSIONS CAN BE MADE. THE PATIENT'S ATTORNEY ALLEGES ADVERSE PATIENT OUTCOME ASSOCIATED WITH THE HERNIA MESH USED TO TREAT THE PATIENT INCLUDING DEATH AND "PAST, PRESENT, AND FUTURE DAMAGES, INCLUDING BUT NOT LIMITED TO, MENTAL AND PHYSICAL PAIN AND SUFFERING FOR SEVERE AND PERMANENT PERSONAL INJURIES" SUSTAINED BY THE PLAINTIFF. NO MEDICAL RECORDS, AUTOPSY REPORT, OR DEATH CERTIFICATE HAVE BEEN PROVIDED. NO LOT NUMBER HAS BEEN PROVIDED; THEREFORE, A REVIEW OF THE MANUFACTURING RECORDS IS NOT POSSIBLE. NOTE, THE DATE OF EVENT (06-MAR-2025) AND DATE OF DEATH (B)(6)2025) CONSIDERED TO BE A BEST ESTIMATE. SHOULD ADDITIONAL INFORMATION BE PROVIDED, A SUPPLEMENTAL EMDR WILL BE SUBMITTED. NOTE: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21718132&pc=LWR | 9612164-2025-01576 |  | Death | MEDTRONIC MEXICO | 2025-03-28 04:00 | LWR | AVALUS AORTIC TISSUE VALVE | Patient Device Interaction Problem; Device Stenosis; Central Regurgitation | Aortic Valve Stenosis; Heart Failure/Congestive Heart Failure; Mitral Valve Insufficiency/ Regurgitation | P170006 |  | 1 | Event Description: MEDTRONIC RECEIVED INFORMATION 8 YEARS POST IMPLANT OF THIS 17 MM AORTIC BIOPROSTHETIC VALVE, IT WAS REPORTED THAT THE PATIENT WAS ADMITTED TO THE HOSPITAL DUE TO HEART FAILURE. THE PATIENT ALSO HAD SIGNS OF AORTIC STENOSIS AND MITRAL REGURGITATION. AFTER INTRODUCTION OF THE INTRA-AORTIC BALLOON PUMP, THE PATIENT DIED. THE CAUSE OF DEATH WAS NOT RECEIVED. THEREFORE, RELATEDNESS OF THE DEATH TO THE BIOPROSTHETIC VALVE COULD NOT BE EXCLUDED. IT IS UNKNOWN WHETHER AN AUTOPSY WAS PERFORMED. Manufacturer Narrative: B2: DATE OF DEATH (NO INFORMATION). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21718160&pc=OZD | 1220648-2025-27023 | 2025-03-02 05:00 | Death | ABIOMED, INC. | 2025-03-28 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Perforation of Vessels | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT ON IMPELLA CP SUPPORT EXPIRED AFTER THE PATIENT¿S CONDITION WORSENED DUE TO VASCULAR DAMAGE THAT OCCURRED WHEN EXTRACORPOREAL MEMBRANE OXYGENATION WAS REMOVED. VASCULAR REPAIR PROCEDURE WAS PERFORMED BUT THE PATIENT CONTINUED WITH BLEEDING AT THE SITE OF THE VASCULAR INJURY. HEPARIN WAS USED IN THE IMPELLA PURGE BUT IT COULD NOT BE DETERMINED IF THE HEPARIN IN THE PURGE EXACERBATED THE SITUATION FOR THE PATIENT. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21718252&pc=OZD | 1220648-2025-27024 | 2025-02-26 05:00 | Death | ABIOMED, INC. | 2025-03-28 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Vascular Dissection | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT HAD AN IMPELLA CP PLACED FOR SUPPORT DUE TO ACUTE MYOCARDIAL INFARCTION. THE PATIENT WAS SUPPORTED BY EXTRACORPOREAL MEMBRANE OXYGENATION ADDITIONALLY. THE PUMP WAS PLACED BUT DUE TO AORTIC DISSECTION, THE PATIENT EXPIRED. THE CONTRIBUTION OF THE IMPELLA CP TO DISSECTING THE ASCENDING AORTA IS NOT KNOWN. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE,.AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH RISK PERCUTANEOUS CORONARY INTERVENTION SECTION: WARNINGS & CAUTIONS: CAUTIONS ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH KNOWN OR SUSPECTED UNREPAIRED ABDOMINAL AORTIC ANEURYSM OR SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM OR DISSECTION OF THE ASCENDING, TRANSVERSE, OR DESCENDING AORTA.¿ CAUTION: ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS SECTION: PRE-SUPPORT EVALUATION SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER SECTION: DIRECT AORTIC INSERTION SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21718656&pc=LYJ | 1644487-2025-00338 | 2025-03-04 05:00 | Death | CYBERONICS - HOUSTON | 2025-03-28 04:00 | LYJ | PULSE GEN MODEL 103 | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P970003 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT RECENTLY PASSED AWAY. NO OTHER RELEVANT INFORMATION HAS BEEN RECEIVED TO DATE. Manufacturer Narrative: LIVANOVA USA, INC. SUBMITS THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION, BASED ON INFORMATION THAT LIVANOVA HAS OBTAINED, BUT MAY NOT HAVE BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION, OR A CONCLUSION BY FDA OR ANYONE ELSE, THAT THE DEVICE, LIVANOVA, OR LIVANOVA'S EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS" OR "MALFUNCTIONS". THESE WORDS ARE INCORPORATED INTO THE FDA 3500A MEDWATCH FORM BY THE FDA, AND LIVANOVA OBJECTS TO THEIR USE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21719078&pc=DQX | MW5168333 | 2025-03-07 05:00 | Death | OPSENS INC. | 2025-03-28 04:00 | DQX | SAVVYWIRE |  | Cardiac Arrest; Embolism/Embolus |  |  | 1 | Event Description: DURING AN ELECTIVE TAVR (TRANSCATHETER AORTIC VALVE REPLACEMENT) PROCEDURE, SAVVY WIRE TEMPORARY PACING WAS PERFORMED. A FEMORAL ACCESS WAS OBTAINED AND THE CATHETER PLACED IN THE PROPER POSITION TO CROSS THE AORTIC VALVE. A SAVVY WIRE WAS INTRODUCED AND PLACED ON THE VENTRICLE APEX. THE VENDOR REPRESENTATIVES WERE PRESENT FOR THE PROCEDURE AND VERIFIED BOTH VENTRICULAR CAPTURE AND THRESHOLD TESTING. THE VALVE WAS DEPLOYED. DURING DEPLOYMENT OF THE VALVE, VENTRICULAR PACING WAS LOST CAUSING EMBOLIZATION OF THE VALVE. THIS RESULTED IN UNSUCCESSFUL ANCHORING OF THE VALUE. AS A RESULT, THERE WAS INJURY TO THE NATIVE VALVE. THE PATIENT BECAME HYPOTENSIVE, CARDIAC ARRESTED AND WAS NOT ABLE TO BE RESUSCITATED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21719095&pc=OZD | 1220648-2025-27031 | 2025-02-13 05:00 | Death | ABIOMED, INC. | 2025-03-28 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Disseminated Intravascular Coagulation (DIC); Renal Failure | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A CP PUMP PLACED FOR 65-YEAR-OLD MALE PATIENT ADMITTED IN WITH ST-ELEVATED MYOCARDIAL INFARCTION AND SUPPORT DURING CARDIOPULMONARY RESUSCITATION (CPR). CPR WAS PERFORMED WAS FOR 90 MINUTES AND THE PATIENT WAS TAKEN FOR EXTERNAL JUGULAR VEIN CANNULATION AND PLACED ON ECMO. THE PATIENT HAD SEVERE HYPOVOLEMIA AND LACTIC ACIDOSIS WHEN MASS TRANSFUSION PROTOCOL WAS INITIATED. HOWEVER, THE PATIENT DEVELOPED DISSEMINATED INTRAVASCULAR COAGULATION (DIC) AND RENAL FAILURE AND A DECISION WAS MADE BY FAMILY TO WITHDRAW CARE. THE PATIENT SUBSEQUENTLY EXPIRED. Manufacturer Narrative: THE INVESTIGATION OF RENAL FAILURE HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. B3 DATE OF EVENT WAS ERRONEOUSLY ENTERED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27031. E4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27031 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMART ASSIST SYSTEM STATES THE FOLLOWING: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21719109&pc=FTL | 1213643-2025-00274 | 2025-03-06 05:00 | Death | DAVOL INC., SUB. C.R. BARD, INC. -1213643 | 2025-03-28 04:00 | FTL | MESH - VENTRALEX | Defective Device; Patient-Device Incompatibility | Pain; Insufficient Information | K132441 |  | 1 | Event Description: ATTORNEY ALLEGES THAT THE PATIENT UNDERWENT SURGERY FOR IMPLANT OF UNSPECIFIED BARD/DAVOL VENTRALEX ON (B)(6) 2019. AS REPORTED, THE PATIENT IS MAKING A CLAIM FOR AN ADVERSE PATIENT OUTCOME AGAINST THE VENTRALEX MESH. ATTORNEY ALLEGES WRONGFUL DEATH OF THE PATIENT AND GENERAL ALLEGATIONS FOR "PAST, PRESENT, AND FUTURE DAMAGES, INCLUDING BUT NOT LIMITED TO, MENTAL AND PHYSICAL PAIN AND SUFFERING FOR SEVERE AND PERMANENT PERSONAL INJURIES SUSTAINED BY THE PATIENT." IT IS ALSO ALLEGED THAT THE PATIENT EXPERIENCED EMOTIONAL DISTRESS AND THE DEVICE WAS DEFECTIVE. Manufacturer Narrative: NO CONCLUSIONS CAN BE MADE. THE PATIENT'S ATTORNEY ALLEGES ADVERSE PATIENT OUTCOME ASSOCIATED WITH THE HERNIA MESH USED TO TREAT THE PATIENT INCLUDING DEATH AND "PAST, PRESENT, AND FUTURE DAMAGES, INCLUDING BUT NOT LIMITED TO, MENTAL AND PHYSICAL PAIN AND SUFFERING FOR SEVERE AND PERMANENT PERSONAL INJURIES" SUSTAINED BY THE PLAINTIFF. NO MEDICAL RECORDS, AUTOPSY REPORT, OR DEATH CERTIFICATE HAVE BEEN PROVIDED. NO LOT NUMBER HAS BEEN PROVIDED; THEREFORE, A REVIEW OF THE MANUFACTURING RECORDS IS NOT POSSIBLE. NOTE, THE DATE OF EVENT (06-MAR-2025) AND DATE OF DEATH ((B)(6) 2025) CONSIDERED TO BE A BEST ESTIMATE. SHOULD ADDITIONAL INFORMATION BE PROVIDED, A SUPPLEMENTAL EMDR WILL BE SUBMITTED. NOTE: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21719199&pc=MKJ | 1220908-2025-01148 | 2025-03-18 04:00 | Death | ZOLL MEDICAL CORPORATION | 2025-03-28 04:00 | MKJ | X SERIES | Failure to Analyze Signal | Insufficient Information | P160022 |  | 1 | Event Description: COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A 66-YEAR-OLD MALE PATIENT, THE DEVICE WAS UNABLE TO ANALYZE. COMPLAINANT INDICATED THAT THE PATIENT SUBSEQUENTLY EXPIRED. Manufacturer Narrative: ZOLL MEDICAL CANADA EVALUATED THE DEVICE AND THE CUSTOMER'S REPORT WAS NOT REPLICATED OR CONFIRMED. REVIEW OF THE DEVICE LOG HISTORY SHOWED THE USER BEGAN THE CASE IN MANUAL MODE WITH CPR STAT- PADZ AND MONITORED IN LEAD II FOR 13 MINUTES BEFORE SWITCHING TO PADS. HOWEVER, THE DEVICE VIEW WAS NEVER CHANGED FROM LEAD II TO PADS. THE LOG DOES NOT APPEAR TO DETECT ECG IN PADS SUGGESTING A POOR CONNECTION. THE DEVICE AND ACCESSORIES WERE RETURNED AND FUNCTIONALLY TESTED WITH NO FAULT FOUND. THE MOST LIKELY CAUSE WAS A POOR CONNECTION BETWEEN THE ELECTRODES AND THE CPR CONNECTOR, BUT THIS COULD NOT BE CONFIRMED. THE DEVICE WAS RECERTIFIED AND RETURNED TO THE CUSTOMER. NO TREND IS ASSOCIATED WITH REPORTS OF THIS TYPE. Manufacturer Narrative: THIS DEVICE WAS MANUFACTURED BUT NOT DOMESTICALLY DISTRIBUTED; IT IS ONLY DISTRIBUTED IN THE INTENDED GEOGRAPHY. THERE IS NO EXISTING UDI REGULATION. ZOLL MEDICAL CORPORATION HAS NOT RECEIVED THE PRODUCT FOR EVALUATION AND THIS COMPLAINT IS STILL UNDER INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21719312&pc=MNS | 2518422-2025-103505 | 2024-10-26 04:00 | Death | RESPIRONICS, INC. | 2025-03-28 04:00 | MNS | DREAMSTATION ST30, GB | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K102465 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN UNSPECIFIED MALFUNCTION RESULTING IN A PATIENT DEATH. DEVICE SETTINGS WERE PROVIDED AS THE FOLLOWING: MODE ST, IPAP 30, EPAP 6, TI. 1.3, RATE 16, RISE 1, RAMP OFF, HUMIDIFICATION 3, TRIGGER AUTO TRAK, 3 L/MIN VIA O2 CONCENTRATOR. ALARMS WERE NOTED AS BEING TURNED OFF. THE CAUSE OF DEATH WAS REPORTED AS CURRENTLY AWAITING A CORONER'S REPORT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER PREVIOUSLY REPORTED INFORMATION ALLEGING AN UNSPECIFIED MALFUNCTION RESULTING IN A PATIENT DEATH. DEVICE SETTINGS WERE PROVIDED AS THE FOLLOWING: MODE ST, IPAP 30, EPAP 6, TI. 1.3, RATE 16, RISE 1, RAMP OFF, HUMIDIFICATION 3, TRIGGER AUTO TRAK, 3 L/MIN VIA O2 CONCENTRATOR. ALARMS WERE NOTED AS BEING TURNED OFF. THE CAUSE OF DEATH WAS REPORTED AS CURRENTLY AWAITING A CORONER'S REPORT. DESPITE THREE GOOD FAITH EFFORT ATTEMPTS TO CHRISTOPHER.RILEY1@NHS.NET ON 15/04/2025, 18/04/2025 AND 21/04/2025, THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. AN ADDITIONAL RESPONSE WAS SENT VIA EMAIL ON 23/04/2025 WITH NO FOLLOW-UP FROM THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE COMPLETED. IF ANY ADDITIONAL INFORMATION IS RECEIVED AT A LATER DATE, A SUPPLEMENTAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21719323&pc=NKE | 2182208-2025-01582 | 2024-04-24 04:00 | Death | MEDTRONIC, INC. | 2025-03-28 04:00 | NKE | MEDTRONIC CRT-P | Adverse Event Without Identified Device or Use Problem | Cardiomyopathy |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING HIS BUNDLE PACING (HBP) VERSUS LEFT BUNDLE BRANCH AREA PACING (LBBAP). THE AUTHORS DESCRIBED TWO PATIENT DEATHS WITHIN THIRTY DAYS OF THE IMPLANT PROCEDURE, ONE IN EACH GROUP. THE PATIENT IN THE HBP GROUP DIED DUE TO AN UNKNOWN CAUSE AND THE PATIENT IN THE LBBAP GROUP DIED DUE TO END-STAGE CARDIOMYOPATHY. THERE WERE TWO PATIENTS IN THE HBP GROUP WHO EXPERIENCED PNEUMOTHORAX WITH NO SURGICAL INTERVENTION REQUIRED AND ONE POCKET HEMATOMA WHICH REQUIRED CONSERVATIVE MEDICAL TREATMENT. LEAD REVISIONS OCCURRED IN BOTH GROUPS DUE TO A RISE IN THRESHOLDS, LOSS OF CAPTURE, EXIT BLOCK, AND DISLODGMENT. THERE WERE FIVE CASES OF EARLY ELECTIVE REPLACEMENT INDICATOR (ERI) IN DEVICES IMPLANTED IN THE HBP GROUP. LEADS IN THE HBP GROUP EXHIBITED DIMINISHED SENSING WHICH WAS MANAGED WITH REPROGRAMMING. THE STATUS OF THE DEVICES AND LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/76 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: SAFETY AND LONG-TERM PERFORMANCE OF THE MEDTRONIC 3830 LEAD IN HIS-BUNDLE VS. LEFT BUNDLE BRANCH AREA PACING: A SINGLE-CENTER 5-YEAR EXPERIENCE. MEDRXIV. 2024; 1-31. DOI: 10.1101/2024.04.23.24306255 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21719449&pc=NVN | 2182208-2025-01583 | 2024-04-24 04:00 | Death | MEDTRONIC, INC. | 2025-03-28 04:00 | NVN | SELECTSECURE MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Cardiomyopathy | P030036 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING HIS BUNDLE PACING (HBP) VERSUS LEFT BUNDLE BRANCH AREA PACING (LBBAP). THE AUTHORS DESCRIBED TWO PATIENT DEATHS WITHIN THIRTY DAYS OF THE IMPLANT PROCEDURE, ONE IN EACH GROUP. THE PATIENT IN THE HBP GROUP DIED DUE TO AN UNKNOWN CAUSE AND THE PATIENT IN THE LBBAP GROUP DIED DUE TO END-STAGE CARDIOMYOPATHY. THERE WERE TWO PATIENTS IN THE HBP GROUP WHO EXPERIENCED PNEUMOTHORAX WITH NO SURGICAL INTERVENTION REQUIRED AND ONE POCKET HEMATOMA WHICH REQUIRED CONSERVATIVE MEDICAL TREATMENT. LEAD REVISIONS OCCURRED IN BOTH GROUPS DUE TO A RISE IN THRESHOLDS, LOSS OF CAPTURE, EXIT BLOCK, AND DISLODGMENT. THERE WERE FIVE CASES OF EARLY ELECTIVE REPLACEMENT INDICATOR (ERI) IN DEVICES IMPLANTED IN THE HBP GROUP. LEADS IN THE HBP GROUP EXHIBITED DIMINISHED SENSING WHICH WAS MANAGED WITH REPROGRAMMING. THE STATUS OF THE DEVICES AND LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/76 YEARS OLD. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: SAFETY AND LONG-TERM PERFORMANCE OF THE MEDTRONIC 3830 LEAD IN HIS-BUNDLE VS. LEFT BUNDLE BRANCH AREA PACING: A SINGLE-CENTER 5-YEAR EXPERIENCE. MEDRXIV. 2024; 1-31. DOI: 10.1101/2024.04.23.24306255. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21719481&pc=OJX | 2182208-2025-01584 | 2024-04-24 04:00 | Death | MEDTRONIC, INC. | 2025-03-28 04:00 | OJX | ATTAIN PERFORMA S MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Cardiomyopathy | P080006 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING HIS BUNDLE PACING (HBP) VERSUS LEFT BUNDLE BRANCH AREA PACING (LBBAP). THE AUTHORS DESCRIBED TWO PATIENT DEATHS WITHIN THIRTY DAYS OF THE IMPLANT PROCEDURE, ONE IN EACH GROUP. THE PATIENT IN THE HBP GROUP DIED DUE TO AN UNKNOWN CAUSE AND THE PATIENT IN THE LBBAP GROUP DIED DUE TO END-STAGE CARDIOMYOPATHY. THERE WERE TWO PATIENTS IN THE HBP GROUP WHO EXPERIENCED PNEUMOTHORAX WITH NO SURGICAL INTERVENTION REQUIRED AND ONE POCKET HEMATOMA WHICH REQUIRED CONSERVATIVE MEDICAL TREATMENT. LEAD REVISIONS OCCURRED IN BOTH GROUPS DUE TO A RISE IN THRESHOLDS, LOSS OF CAPTURE, EXIT BLOCK, AND DISLODGMENT. THERE WERE FIVE CASES OF EARLY ELECTIVE REPLACEMENT INDICATOR (ERI) IN DEVICES IMPLANTED IN THE HBP GROUP. LEADS IN THE HBP GROUP EXHIBITED DIMINISHED SENSING WHICH WAS MANAGED WITH REPROGRAMMING. THE STATUS OF THE DEVICES AND LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/76 YEARS OLD. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: SAFETY AND LONG-TERM PERFORMANCE OF THE MEDTRONIC 3830 LEAD IN HIS-BUNDLE VS. LEFT BUNDLE BRANCH AREA PACING: A SINGLE-CENTER 5-YEAR EXPERIENCE. MEDRXIV. 2024; 1-31. DOI: 10.1101/2024.04.23.24306255. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21719501&pc=LWS | 2649622-2025-08544 | 2025-03-05 05:00 | Death | MPRI | 2025-03-28 04:00 | LWS | SPRINT QUATTRO SECURE S MRI SURESCAN | Over-Sensing; Under-Sensing | Ventricular Fibrillation | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IS DECEASED. ON THE DATE OF DEATH, A REMOTE TRANSMISSION WAS RECEIVED THAT INDICATED RIGHT VENTRICULAR (RV) LEAD UNDERSENSING OF AN AGONAL VENTRICULAR FIBRILLATION (VF). ADDITIONALLY, THE RV LEAD TRIGGERED A LEAD INTEGRITY ALERT (LIA) FOR HIGH RATE NON-SUSTAINED EPISODES AND SENSING INTEGRITY COUNTER (SIC). Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE LEAD WAS NOT RETURNED FOR ANALYSIS, HOWEVER, PERFORMANCE DATA COLLECTED FROM THE DEVICE WAS RECEIVED AND ANALYZED. ANALYSIS OF THE DEVICE MEMORY INDICATED RIGHT VENTRICULAR UNDERSENSING. ANALYSIS OF THE DEVICE MEMORY INDICATED THE CRITERIA FOR THE RIGHT VENTRICULAR LEAD INTEGRITY ALERT WERE MET. ANALYSIS OF THE DEVICE MEMORY INDICATED OVERSENSING DUE TO NON-PHYSIOLOGIC SIGNALS/SENSING INTEGRITY COUNTER. ANALYSIS OF THE DEVICE MEMORY INDICATED OVERSENSING ASSOCIATED WITH THE RIGHT VENTRICULAR LEAD. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21719526&pc=OZD | 1220648-2025-27035 | 2025-03-01 05:00 | Death | ABIOMED, INC. | 2025-03-28 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WAS ON IMPELLA CP SUPPORT AND DEVELOPED A HEMATOMA AND BLEEDING FROM THE IMPELLA ACCESS SITE. FEM STOP WAS PLACED BUT GIVEN THE SEVERITY OF THE HEMATOMA, THE PATIENT WAS TAKEN TO THE OPERATING ROOM AND IMPELLA CP WAS REMOVED WITH VASCULAR REPAIR. THE IMPELLA CP WAS EXCHANGED FOR IMPELLA 5.5, BUT PATIENT EXPIRED SOON AFTER. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM, SECTION: GENERAL PATIENT CARE CONSIDERATIONS: ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT, USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21719573&pc=OLO | 3005075696-2025-00239 | 2025-02-25 05:00 | Death | MAZOR ROBOTICS LTD | 2025-03-28 04:00 | OLO | MAZOR X SYSTEM | Unintended Collision; Incorrect, Inadequate or Imprecise Result or Readings; Use of Device Problem; Application Program Freezes, Becomes Nonfunctional | Insufficient Information; No Clinical Signs, Symptoms or Conditions | K182077 |  | 1 | Event Description: PATEL, N. A., BROWN, N. J., HE, C., PATEL, S., PENNINGTON, Z., GENDREAU, J., SAHYOUNI, R., <(>&<)> PHAM, M. H. (2025). CHARACTERIZING THE COMPLICATION PROFILE OF SPINAL ROBOTIC SYSTEMS: A MAUDE ANALYSIS OF DEVICE FAILURES AND ASSOCIATED COMPLICATIONS BY DEVICE MANUFACTURER AND BRAND NAME. JOURNAL OF CLINICAL NEUROSCIENCE, 135, 111149. HTTPS://DOI.ORG/10.1016/J.JOCN.2025.111149 INTRODUCTION: PEDICLE SCREW PLACEMENT IN SPINAL SURGERY HAS EVOLVED TO INCORPORATE INCREASED USE OF TECHNOLOGY FOR BOTH PREOPERATIVE PLANNING AND INTRAOPERATIVE GUIDANCE. IN RECENT YEARS, ROBOTIC GUIDANCE HAS BEEN USED TO INCREASE THE PRECISION, ACCURACY, AND OVERALL PATIENT SAFETY OF SPINAL COLUMN FIXATION VIA SCREW PLACEMENT. WITHIN THIS CONTEXT, THE PRESENT STUDY AIMS TO QUANTIFY AND CLASSIFY COMPLICATIONS OF SPINAL ROBOTIC DEVICES USING A NATIONALLY REPRESENTATIVE DATABASE. METHODS: THE MANUFACTURER AND USER FACILITY DEVICE EXPERIENCE (MAUDE) WAS QUERIED FOR ADVERSE EVENTS INVOLVING THE MOST WIDELY USED SPINAL ROBOTIC DEVICES FROM INCEPTION TO NOVEMBER 3RD, 2023. ENTRIES WERE TABULATED BY ASSIGNING CATEGORICAL VARIABLES BASED ON EVENT INFORMATION. RESULTS: THE QUERY YIELDED 339 UNIQUE ROBOTIC FAILURE EVENTS. THE MOST FREQUENT COMPLICATIONS IN ROBOTIC-ASSISTED SPINAL SURGERIES WERE IMPRECISION (73.6%), SOFTWARE ERRORS (13.7%), AND DAMAGED OR MISSING SCREWS (4.9%). MOST REPORTED EVENTS OCCURRED INTRAOPERATIVELY AFTER THE INITIAL INCISION (78.5%) [COMPARED TO PREOPERATIVELY OR POSTOPERATIVELY]. THE MAZOR X¿ SYSTEM WAS THE MOST REPRESENTED SPINAL ROBOTIC PLATFORM IN TERMS OF COMPLICATIONS, ACCOUNTING FOR 71.6% OF REPORTED ROBOTIC FAILURE EVENTS. MOST REPORTED ERRORS WERE RESOLVED INTRA-OPERATIVELY OR CONTINUED THROUGH THE EMPLOYMENT OF ALTERNATIVE IMAGING METHODS AND, OVERALL, HAD MINIMAL CLINICAL SIGNIFICANCE. WITH RESPECT TO ROBOT-ASSISTED PEDICLE SCREW PLACEMENT, MEDIAL BREACHES WERE 1.7X MORE COMMON THAN LATERAL BREACHES ACROSS ALL ROBOTIC SYSTEMS. FINALLY, REGARDING THE IMPACT OF DEVICE FAILURES ON PERIOPERATIVE TIMING, WE FOUND THAT 10.3% OF ADVERSE EVENTS RESULTED IN DELAYS GREATER THAN 1 H. CONCLUSION: ULTIMATELY, IMPRECISION AND UNRESPONSIVE SOFTWARE WERE FOUND TO BE TWO OF THE MOST FREQUENT COMPLICATIONS OF ROBOT-ASSISTED SPINE SURGERY. THOUGH MOST REPORTED ERRORS WERE CIRCUMVENTED INTRAOPERATIVELY THROUGH CONVENTIONAL, NON-ROBOTIC IMAGING/GUIDANCE METHODS, 10.3% OF ADVERSE EVENTS RESULTED IN PROCEDURAL DELAYS OF GREATER THAN 1 H. REPORTED EVENTS 241 EVENTS WERE DOCUMENTED WHILE USING A GUIDANCE SYSTEM. OF THESE EVENTS, 2 OCCURRED OUTSIDE OF THE OPERATING ROOM(OR), 2 IN THE OR BEFORE ANESTHESIA, 12 IN THE OR AFTER ANESTHESIA, 195 AFTER THE INCISION WAS MADE, AND 30 WERE POSTOPERATIVE. IF THE EVENT OCCURRED DURING A PROCEDURE, THE PROCEDURE WAS A SPINAL PROCEDURE. OF THE EVENTS THAT OCCURRED DURING A PROCEDURE, 19 CONTINUED WITHOUT CORRECTION, 82 AMELIORATED, 14 CONTINUED WITH NAVIGATION ONLY, 4 ABORTED PROCEDURE, 86 CONTINUED WITH FLUOROSCOPY, AND IN 15 IT WAS UNKNOWN HOW THE CASE WAS COMPLETED. IT WAS NOTED 37 EVENTS HAD NO DELAY, 1 HAD LESS THAN 30 MINUTES DELAY, 136 HAD LESS THAN 1 HOUR OF DELAY, 34 HAD GREATER THAN 1 HOUR OF DELAY, AND 33 IT WAS UNKNOWN HOW MUCH DELAY THERE WAS. THE LITERATURE ALSO DOCUMENTS 54 EVENTS DID NOT NEED REVISION, IN 131 THERE WAS AN INTRAOPERATIVE REVISION, IN 35 THERE WAS A POST-OPERATIVE REVISION, IN 5 THE PROCEDURE WAS RESCHEDULED, AND IN 16 IT WAS UNKNOWN IF ANY CORRECTION WAS PERFORMED. IN ONE EVENT A DEATH WAS DOCUMENTED. THERE WERE 23 EVENTS TO REPORT CLINICALLY SIGNIFICANT INJURY AND 12 REPORTED CLINICALLY INSIGNIFICANT INJURY. THERE WERE 206 IMPRECISIONS. 1 IMPRECISION WAS 1MM OFF, 5 WERE 2MM OFF, 20 WERE 3MM OFF, 19 WERE 4MM OFF, 17 WERE 5MM OFF, 2 WERE 6MM OFF, 6 WERE 7MM OFF, 1 WAS 8MM OFF, 13 WERE AT LEAST 9MM OFF, AND FOR 157 IT WAS NOT APPLICABLE OR UNKNOWN WHAT THE IMPRECISION AMOUNT WAS. OF THESE IMPRECISIONS, 18 WERE SUPERIOR, 65 WERE LATERAL, 18 WERE INFERIOR, 101 WERE MEDIAL, AND FOR 39 THE DIRECTION WAS UNKNOWN. THERE WAS 1 DEVICE COLLISION, 1 REGISTRATION ERROR, 19 EVENTS OF SOFTWARE UNRESPONSIVENESS, 5 MECHANICAL DEFORMATIONS, 2 COMPUTER HARDWARE FAILURES, AND 1 UNKNOWN ADVERSE EVENT. IN 202 EVENTS THERE WAS NO INJURY AND IN 3 EVENTS IT WAS UNKNOWN IF THERE WAS ANY IMPACT ON PATIENT OUTCOME. Manufacturer Narrative: A.1-A.5) PATIENT INFORMATION WAS NOT INCLUDED IN THE JOURNAL. B.3. PLEASE NOTE THAT THIS DATE IS BASED OFF OF THE DATE THE ARTICLE WAS AVAILABLE ONLINE AS THE EVENT DATES WERE NOT PROVIDED IN THE PUBLISHED LITERATURE. B.5. INCLUDED THE ARTICLE CITATION D.4. THE SYSTEM SERIAL NUMBER WAS NOT PROVIDED IN THE JOURNAL ARTICLE. H.3. NO EVALUATION WAS PERFORMED AS THE EVENT WAS REPORTED AS A LITERATURE ARTICLE. H.4. DEVICE MANUFACTURING DATE WAS UNAVAILABLE. H.6. MULTIPLE PATIENT CODES WERE APPLIED. BELOW LISTED WHAT EACH CORRELATES TO. E2401 - DEATH AND ALLEGED CLINICAL INJURIES E2403 - EVENTS WITH NO PATIENT HARM OR THAT OCCURRED OUTSIDE OF A PROCEDURE. MULTIPLE DEVICE CODES WERE APPLIED. BELOW LISTED WHAT EACH CORRELATES TO. A0908 - INACCURACY A051205 - DEVICE COLLISION A110201 - SOFTWARE UNRESPONSIVENESS A23 - REGISTRATION ERROR, COMPUTER HARDWARE FAILURE, AND THE UNKNOWN ADVERSE EVENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21719593&pc=PNJ | 2017865-2025-38818 | 2025-03-13 04:00 | Death | ABBOTT | 2025-03-28 04:00 | PNJ | AVEIR¿ DELIVERY CATHETER | Activation, Positioning or Separation Problem | Arrhythmia; Low Blood Pressure/ Hypotension; Cardiac Perforation; Pericardial Effusion | P150035 |  | 1 | Event Description: IT WAS REPORTED THAT AFTER FIXATION, HIGH PACING IMPEDANCE WAS OBSERVED ON THE DEVICE. SHORTLY AFTER, THE PATIENT EXPERIENCED LOW BLOOD PRESSURE AND ARRYTHMIA THAT LEAD TO CARDIAC ARREST. AN ECHOCARDIOGRAPHY WAS PERFORMED AND REVEALED AN EFFUSION DUE TO A SUSPECTED PERFORATION. RESUSCITATION AND CIRCULATORY SUPPORT WERE ADMINISTERED, HOWEVER, THE PATIENT PASSED AWAY. THE DELIVERY CATHETER AND DEVICE WERE LEFT IN THE PATIENT. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21719614&pc=PNJ | 2017865-2025-38817 | 2025-03-13 04:00 | Death | ABBOTT | 2025-03-28 04:00 | PNJ | AVEIR¿ LEADLESS PACEMAKER, RIGHT VENTRICLE | High impedance | Arrhythmia; Low Blood Pressure/ Hypotension; Cardiac Perforation; Pericardial Effusion | P150035 |  | 1 | Event Description: IT WAS REPORTED THAT AFTER FIXATION, HIGH PACING IMPEDANCE WAS OBSERVED ON THE DEVICE. SHORTLY AFTER, THE PATIENT EXPERIENCED LOW BLOOD PRESSURE AND ARRYTHMIA THAT LEAD TO CARDIAC ARREST. AN ECHOCARDIOGRAPHY WAS PERFORMED AND REVEALED AN EFFUSION DUE TO A SUSPECTED PERFORATION. RESUSCITATION AND CIRCULATORY SUPPORT WERE ADMINISTERED, HOWEVER, THE PATIENT PASSED AWAY. THE DELIVERY CATHETER AND DEVICE WERE LEFT IN THE PATIENT. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21719711&pc=DSQ | 2916596-2025-01974 | 2025-03-12 04:00 | Death | THORATEC CORPORATION | 2025-03-28 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2025 AS THE PATIENT CHOSE TO ELECTIVELY REMOVE MECHANICAL CIRCULATORY SUPPORT (MCS) DUE TO A DECREASED QUALITY OF LIFE. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED, AND THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21719984&pc=LWP | 2017865-2025-38798 | 2025-02-18 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-03-28 04:00 | LWP | ASSURITY MRI | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P140033 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT HAS DECEASED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE RELATED. THE CAUSE OF DEATH WAS UNKNOWN. NO ADDITIONAL INFORMATION WAS REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21719987&pc=LYJ | 1644487-2025-00342 | 2024-11-05 05:00 | Death | LIVANOVA USA, INC. | 2025-03-28 04:00 | LYJ | PULSE GEN MODEL 104 | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P970003 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY AND NO FURTHER DETAILS WERE PROVIDED. NO OTHER RELEVANT INFORMATION HAS BEEN RECEIVED TO DATE.Event Description: FURTHER INFORMATION WAS RECEIVED STATING THAT PER THE PHYSICIAN, THE PATIENT'S DEATH WAS NOT RELATED TO THE VNS. THE PATIENT EXPERIENCED SEIZURE REDUCTION WITH VNS AND WAS RECEIVING VNS TREATMENT AT TIME OF DEATH. NEITHER THE GENERATOR NOR THE LEAD WERE EXPLANTED AFTER DEATH. THE CAUSE OF DEATH IS RESPIRATORY FAILURE - RECURRENT PNEUMONIAS RESULTING IN RESPIRATORY FAILURE AS WITNESSED BY HOSPITAL STAFF. NO AUTOPSY WAS PERFORMED, AND NO DEATH CERTIFICATE IS AVAILABLE. Manufacturer Narrative: LIVANOVA USA, INC. SUBMITS THIS REPORT TO COMPLY WITH 21 CFR PART 803, THE MEDICAL DEVICE REPORTING REGULATION, BASED ON INFORMATION THAT LIVANOVA HAS OBTAINED, BUT MAY NOT HAVE BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION, OR A CONCLUSION BY FDA OR ANYONE ELSE, THAT THE DEVICE, LIVANOVA, OR LIVANOVA'S EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS" OR "MALFUNCTIONS". THESE WORDS ARE INCORPORATED INTO THE FDA 3500A MEDWATCH FORM BY THE FDA, AND LIVANOVA OBJECTS TO THEIR USE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21719998&pc=NVN | 2017865-2025-38799 | 2025-02-18 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-03-28 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT HAS DECEASED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE RELATED. THE CAUSE OF DEATH WAS UNKNOWN. NO ADDITIONAL INFORMATION WAS REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21720016&pc=NVN | 2017865-2025-38800 | 2025-02-18 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-03-28 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT HAS DECEASED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE RELATED. THE CAUSE OF DEATH WAS UNKNOWN. NO ADDITIONAL INFORMATION WAS REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21720575&pc=DWF | 8010762-2025-0000146 | 2025-03-17 04:00 | Death | MAQUET CARDIOPULMONARY GMBH | 2025-03-28 04:00 | DWF | HLS CANNULAE & PIK | Fluid/Blood Leak | Cardiogenic Shock; Insufficient Information | K102532 |  | 1 | Event Description: IT WAS REPORTED THAT THE EXTRACORPOREAL CIRCULATION CANNULA (WHERE THE FEMORAL ARTERY CATHETER IS CONNECTED) RUPTURES AND BLEEDS. IT IS UNKNOWN AT THIS MOMENT IF ANY IMMEDIATE ACTION WAS REQUIRED FOR PATIENT. SINCE THE FAILURE WAS OCCURRED DURING TREATMENT AND POTENTIAL HARM TO PATIENT WAS REPORTED, THE COMPLAINT IS REPORTABLE AS SERIOUS INJURY. COMPLAINT #(B)(4).Event Description: COMPLAINT #: (B)(4).Event Description: COMPLAINT # (B)(4). Manufacturer Narrative: A 57-YEAR-OLD PATIENT UNDERWENT VA-ECMO ON (B)(6) 2025, DUE TO MYOCARDIAL INFARCTION AND CARDIOGENIC SHOCK. SHORTLY AFTER ECMO INITIATION, MINOR BLOOD SEEPAGE WAS OBSERVED AT THE ARTERIAL CANNULA CONNECTOR. NO ISSUES WERE IDENTIFIED DURING PRE-USE INSPECTION. THE PATIENT¿S CONDITION DETERIORATED, AND TREATMENT WAS DISCONTINUED AT THE FAMILY'S REQUEST. THE PATIENT SUBSEQUENTLY EXPIRED. IT WAS LATER CONFIRMED BY THE SSU THAT THE BE-PAL 1723 AND PLS 2050 COMPONENTS REMAINED UNCHANGED THROUGHOUT THE TREATMENT. SAMPLE INVESTIGATION COULD NOT BE PERFORMED AS THE PRODUCT WAS DISCARDED BY CUSTOMER. HOWEVER, PHOTOGRAPHICAL EVIDENCE WAS RECEIVED FROM CUSTOMER THAT SHOWS THE CRACK CLEARLY, BASED ON THIS THE COMPLAINT COULD BE CONFIRMED. GETINGE MEDICAL AFFAIRS CONDUCTED A MEDICAL REVIEW ON 2025-05-09. WHILE THE EXACT CAUSE OF THE LEAKAGE REMAINS UNCLEAR DUE TO LIMITED EVIDENCE AND THE ABSENCE OF THE INVOLVED COMPONENTS FOR ANALYSIS, IT IS WORTH NOTING THAT HOSE TIES WERE NOT USED TO SECURE THE TUBE CONNECTIONS, AS RECOMMENDED IN THE INSTRUCTIONS FOR USE (IFU). THE IFU ADVISES: DETACHMENT OF UNSECURED TUBE CONNECTIONS CAN LEAD TO BLOOD LOSS AND AIR EMBOLISMS IN THE PATIENT. CHECK THAT TUBE CONNECTIONS ARE CORRECTLY AND SECURELY FASTENED. SECURE ALL TUBE CONNECTIONS IN THE TUBE SYSTEM WITH HOSE TIES. CONSIDERING THAT A LIMB PERFUSION LINE WAS CONNECTED, AND THE LEAK OCCURRED SHORTLY AFTER ECMO INITIATION, ONE POTENTIAL ROOT CAUSE COULD BE MECHANICAL STRESS AT THE CONNECTION POINT, POSSIBLY DUE TO HANDLING OR OVER-TIGHTENING OF THE LUER LOCK. OTHER POTENTIAL ALTERNATIVE ROOT CAUSES MAY INCLUDE INADVERTENT MECHANICAL STRESS MICROFRACTURES FROM MATERIAL FATIGUE OR AN UNDETECTED MANUFACTURING DEFECT THAT BECAME EVIDENT UNDER OPERATIONAL STRESS. HOWEVER, AS THE CANNULA AND CONNECTOR WERE DISCARDED AFTER THE INCIDENT, THEREBY PREVENTING FURTHER INVESTIGATION TO DETERMINE A DEFINITIVE ROOT CAUSE. ACCORDING TO THE CUSTOMER, THE PATIENT¿S CONDITION DETERIORATED, AND TREATMENT WAS DISCONTINUED AT THE FAMILY'S REQUEST. BASED ON THE INFORMATION PROVIDED IN CUSTOMER PRODUCT COMPLAINT SINGLE REPORT # (B)(4) (MFG REPORT #8010762-2025-0000146) AND THE COMPLETED QUESTIONNAIRE, THERE IS NO INDICATION OF A PRODUCT PERFORMANCE ISSUE. GIVEN THE LIMITED EVIDENCE AND ABSENCE OF THE INVOLVED COMPONENTS FOR FURTHER ANALYSIS, NO DEFINITIVE ROOT CAUSE FOR THE LEAKAGE CAN BE ESTABLISHED. HOWEVER, NO PRODUCT MALFUNCTION OR DIMINUTION IN PERFORMANCE IS SUSPECTED. THE PRODUCTION HISTORY RECORD (DHR) OF THE AFFECTED BE-PAL 1723 WITH LOT# 3000356409 WAS REVIEWED ON 2025-03-28. ACCORDING TO THE DHR RESULTS, THE PRODUCT BE-PAL 1723 PASSED THE DEFINED MANUFACTURING AND FINAL RELEASE SPECIFICATIONS. THERE IS NO INDICATION ON MANUFACTURING ISSUES. THUS PRODUCTION RELATED INFLUENCES ARE UNLIKELY. FURTHER, THE INCOMING INSPECTION REPORT REVIEW HAS BEEN PERFORMED FOR AFFECTED COMPONENT ¿700000285, 00285#KONNEKTOR 3/8 X 3/8 LL¿. THE INCOMING INSPECTION REPORT (BATCH # 3000276586) OF THE AFFECTED COMPONENT "700000285 / 00285#KONNEKTOR 3/8 X 3/8 LL" WAS REVIEWED ON 2025-03-27. THE CONNECTOR WERE CHECKED FOR DAMAGES, SCRATCHES, MARKS, RILLS, SINKS, STREAKS, AND CORDS VISUALLY. VISUAL TESTS WERE PASSED AS PER SPECIFICATIONS. DUE TO THIS, MATERIAL RELATED INFLUENCES ARE UNLIKELY. THE REVIEW OF THE NON-CONFORMITIES HAS BEEN PERFORMED. IT DOES NOT SHOW ANY NON-CONFORMITY IN REGARDS TO THE BATCH NUMBERS OF REPORTED COMPONENTS WITH RELATED TO THE REPORTED FAILURE. FURTHER, THE RELATED BASIC OPERATION PROCEDURE TRAINING RECORD HAS BEEN CONTROLLED FOR PRODUCTION EMPLOYEES AND THERE COULD NOT BE FOUND ANY NON-CONFORMITIES THAT COULD CAUSE TO REPORTED FAILURE. THE OCCURRENCE RATE WAS CALCULATED FOR THE REPORTED FAILURE AND PRODUCT AND IT WAS DETERMINED THAT THIS IS NOT A SYSTEMIC ISSUE. THEREFORE, NO REMEDIAL ACTION IS REQUIRED. THE OCCURRENCE RATE RELATED TO THE REPORTED ISSUE IS CURRENTLY BEING MONITORED AS PART OF MAQUET CARDIOPULMONARY¿S TRENDING PROGRAM AND ADDITIONAL INVESTIGATIONS OR CORRECTIONS WILL BE IMPLEMENTED IN CASE OF ADVERSE TRENDING. Manufacturer Narrative: A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN FURTHER INFORMATION BECOMES AVAILABLE. Manufacturer Narrative: NEW INFORMATION WAS RECEIVED ON 2025-04-17 AS THE PATIENT HAS EXPIRED DUE TO SEVERE CARDIOGENIC SHOCK. THE REPORTED FAILURE 'CRACK ON HLS CANNULAE CONNECTOR' DID NOT AFFECT THE PATIENT TREATMENT DIRECTLY. HOWEVER, THE FAMILY MEMBERS OF PATIENT VOLUNTARILY GAVE UP FROM TREATMENT. FURTHER, CUSTOMER CONFIRMS THAT THE BOTH PRODUCTS WERE NOT CHANGED DURING TREATMENT 'BE-PAL 1723 AND PLS 2050'. A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN FURTHER INFORMATION BECOMES AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21720577&pc=DSQ | 2916596-2025-01853 | 2025-03-11 04:00 | Death | THORATEC CORPORATION | 2025-03-28 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. AN AUTOPSY WAS NOT PERFORMED, AND THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6) , AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICLE ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21720751&pc=NGV | 2124215-2025-19123 | 2025-03-09 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-03-28 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Cardiac Arrest; Loss of consciousness | P130013 |  | 1 | Event Description: IT WAS REPORTED THAT DEATH OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS PERFORMED USING TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) IMAGING. A WATCHMAN TRUSTEER ACCESS SYSTEM (WAS) WAS USED, AND A 24MM WATCHMAN FLX PRO CLOSURE DEVICE WAS IMPLANTED AFTER BEING DEPLOYED ONCE, AND MEETING RELEASE CRITERIA. NO COMPLICATIONS WERE NOTED DURING THE PROCEDURE. THE PATIENT WAS DISCHARGED TO HOME THE NEXT DAY POST INDEX PROCEDURE. TWO DAYS POST INDEX PROCEDURE, THE CAREGIVER FOUND THE PATIENT UNRESPONSIVE AT HOME. EMERGENCY MEDICAL SERVICES (EMS) WAS CALLED, AND THE PATIENT WAS FOUND TO BE IN PULSELESS ELECTRICAL ACTIVITY (PEA) CARDIAC ARREST RHYTHM. THE PATIENT WAS BROUGHT TO THE HOSPITAL WHERE THE PATIENT WAS PRONOUNCED DEAD IN THE EMERGENCY ROOM. THE OFFICIAL CAUSE OF DEATH IS FULL ARREST AND PULSE-LESS ELECTRICAL ACTIVITY (PEA). THERE WAS NO FURTHER INFORMATION PROVIDED.Event Description: IT WAS REPORTED THAT DEATH OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS PERFORMED USING TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) IMAGING. A WATCHMAN TRUSTEER ACCESS SYSTEM (WAS) WAS USED, AND A 24MM WATCHMAN FLX PRO CLOSURE DEVICE WAS IMPLANTED AFTER BEING DEPLOYED ONCE, AND MEETING RELEASE CRITERIA. NO COMPLICATIONS WERE NOTED DURING THE PROCEDURE. THE PATIENT WAS DISCHARGED TO HOME THE NEXT DAY POST INDEX PROCEDURE. TWO DAYS POST INDEX PROCEDURE, THE CAREGIVER FOUND THE PATIENT UNRESPONSIVE AT HOME. EMERGENCY MEDICAL SERVICES (EMS) WAS CALLED, AND THE PATIENT WAS FOUND TO BE IN PULSELESS ELECTRICAL ACTIVITY (PEA) CARDIAC ARREST RHYTHM. THE PATIENT WAS BROUGHT TO THE HOSPITAL WHERE THE PATIENT WAS PRONOUNCED DEAD IN THE EMERGENCY ROOM. THE OFFICIAL CAUSE OF DEATH IS FULL ARREST AND PULSE-LESS ELECTRICAL ACTIVITY (PEA). THE PHYSICIAN STATED THE DEATH WAS "PROBABLY NOT" RELATED TO THE INDEX PROCEDURE OR THE IMPLANTED CLOSURE DEVICE. AN AUTOPSY WILL NOT BE PERFORMED. THERE WAS NO FURTHER INFORMATION PROVIDED OR AVAILABLE. Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI) # CORRECTED FROM (B)(4) TO (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21720757&pc=DSQ | 2916596-2025-02146 | 2024-07-28 04:00 | Death | THORATEC CORPORATION | 2025-03-28 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Hypervolemia; Gastrointestinal Hemorrhage | P160054 |  | 1 | Event Description: T WAS REPORTED THAT THE PATIENT HAD MIGRATED TO NEW ZEALAND WHERE THEY WERE ADMITTED FOR FLUID OVERLOAD AND A ONGOING GASTROINTESTINAL (GI) BLEED WITH HYPOTENSION ON (B)(6) 2025. THE PATIENT DETERIORATED THE MORNING OF (B)(6) 2025 AND THE DECISION WAS MADE TO TURN THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) OFF IN THE HOSPITAL. THE PATIENT PASSED AWAY SOON AFTER. THE DEVICE WAS NOT EXPLANTED.Event Description: IT WAS REPORTED THAT THERE WERE NO CHANGES IN ANTICOAGULATION STATUS AT THE TIME OF THE EVENT. THE PATIENT WAS ON A WARFARIN DOSE OF 0.5MG. NO DIAGNOSTIC TESTING WAS DONE AND THE BLEEDING WAS MANAGED CONSERVATIVELY. THERE WERE NO SYMPTOMS OTHER THAN HYPOTENSION AND FLUID OVERLOAD. THE DEVICE OPERATED AS EXPECTED. THE ORIGINAL BLEED WAS (B)(6) 2024.Event Description: THE BLEEDING WAS TREATED CONSERVATIVELY AND THERE WERE NO PROCEDURES OR INVASIVE MEASURES. THE SOURCE OF THE BLEEDING WAS NOT CONFIRMED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: CORRECTED DATA: SECTION H6: HEALTH EFFECT - CLINICAL CODE CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C, IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING BLEEDING AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, OUTLINES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING THE INTERNATIONAL NORMALIZED RATIO RANGE FOR PATIENTS USING THE HEARTMATE 3 LVAS, AS WELL AS THE SUGGESTED ANTICOAGULATION MODIFICATIONS IN THE EVENT THAT THERE IS A RISK OF BLEEDING. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21721139&pc=DWF | 3013876692-2025-00022 | 2025-03-17 04:00 | Death | MAQUET CARDIOPULMONARY GMBH | 2025-03-28 04:00 | DWF | HLS CANNULAE & PIK | Fluid/Blood Leak | Cardiogenic Shock; Insufficient Information | K102532 |  | 1 | Event Description: IT WAS REPORTED THAT THE EXTRACORPOREAL CIRCULATION CANNULA (WHERE THE FEMORAL ARTERY CATHETER IS CONNECTED) RUPTURES AND BLEEDS. IT IS UNKNOWN AT THIS MOMENT IF ANY IMMEDIATE ACTION WAS REQUIRED FOR PATIENT. SINCE THE FAILURE OCCURRED DURING TREATMENT AND POTENTIAL HARM TO PATIENT WAS REPORTED, THE COMPLAINT IS REPORTABLE AS SERIOUS INJURY. COMPLAINT #(B)(4). Event Description: COMPLAINT # (B)(4).Event Description: COMPLAINT#: (B)(4). Manufacturer Narrative: A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN FURTHER INFORMATION BECOMES AVAILABLE. Manufacturer Narrative: A 57-YEAR-OLD PATIENT UNDERWENT VA-ECMO ON (B)(6) 2025, DUE TO MYOCARDIAL INFARCTION AND CARDIOGENIC SHOCK. SHORTLY AFTER ECMO INITIATION, MINOR BLOOD SEEPAGE WAS OBSERVED AT THE ARTERIAL CANNULA CONNECTOR. NO ISSUES WERE IDENTIFIED DURING PRE-USE INSPECTION. THE PATIENT¿S CONDITION DETERIORATED, AND TREATMENT WAS DISCONTINUED AT THE FAMILY'S REQUEST. THE PATIENT SUBSEQUENTLY EXPIRED. IT WAS LATER CONFIRMED BY THE SSU THAT THE BE-PAL 1723 AND PLS 2050 COMPONENTS REMAINED UNCHANGED THROUGHOUT THE TREATMENT. SAMPLE INVESTIGATION COULD NOT BE PERFORMED AS THE PRODUCT WAS DISCARDED BY CUSTOMER. HOWEVER, PHOTOGRAPHICAL EVIDENCE WAS RECEIVED FROM CUSTOMER THAT SHOWS THE CRACK CLEARLY, BASED ON THIS THE COMPLAINT COULD BE CONFIRMED. GETINGE MEDICAL AFFAIRS CONDUCTED A MEDICAL REVIEW ON 2025-05-09. WHILE THE EXACT CAUSE OF THE LEAKAGE REMAINS UNCLEAR DUE TO LIMITED EVIDENCE AND THE ABSENCE OF THE INVOLVED COMPONENTS FOR ANALYSIS, IT IS WORTH NOTING THAT HOSE TIES WERE NOT USED TO SECURE THE TUBE CONNECTIONS, AS RECOMMENDED IN THE INSTRUCTIONS FOR USE (IFU). THE IFU ADVISES: - DETACHMENT OF UNSECURED TUBE CONNECTIONS CAN LEAD TO BLOOD LOSS AND AIR EMBOLISMS IN THE PATIENT. - CHECK THAT TUBE CONNECTIONS ARE CORRECTLY AND SECURELY FASTENED. - SECURE ALL TUBE CONNECTIONS IN THE TUBE SYSTEM WITH HOSE TIES. CONSIDERING THAT A LIMB PERFUSION LINE WAS CONNECTED, AND THE LEAK OCCURRED SHORTLY AFTER ECMO INITIATION, ONE POTENTIAL ROOT CAUSE COULD BE MECHANICAL STRESS AT THE CONNECTION POINT, POSSIBLY DUE TO HANDLING OR OVER-TIGHTENING OF THE LUER LOCK. OTHER POTENTIAL ALTERNATIVE ROOT CAUSES MAY INCLUDE INADVERTENT MECHANICAL STRESS MICROFRACTURES FROM MATERIAL FATIGUE OR AN UNDETECTED MANUFACTURING DEFECT THAT BECAME EVIDENT UNDER OPERATIONAL STRESS. HOWEVER, AS THE CANNULA AND CONNECTOR WERE DISCARDED AFTER THE INCIDENT, THEREBY PREVENTING FURTHER INVESTIGATION TO DETERMINE A DEFINITIVE ROOT CAUSE. ACCORDING TO THE CUSTOMER, THE PATIENT¿S CONDITION DETERIORATED, AND TREATMENT WAS DISCONTINUED AT THE FAMILY'S REQUEST. BASED ON THE INFORMATION PROVIDED IN CUSTOMER PRODUCT COMPLAINT SINGLE REPORT #: (B)(4) (MFG REPORT#: 8010762-2025-0000146) AND THE COMPLETED QUESTIONNAIRE, THERE IS NO INDICATION OF A PRODUCT PERFORMANCE ISSUE. GIVEN THE LIMITED EVIDENCE AND ABSENCE OF THE INVOLVED COMPONENTS FOR FURTHER ANALYSIS, NO DEFINITIVE ROOT CAUSE FOR THE LEAKAGE CAN BE ESTABLISHED. HOWEVER, NO PRODUCT MALFUNCTION OR DIMINUTION IN PERFORMANCE IS SUSPECTED. THE PRODUCTION HISTORY RECORD (DHR) OF THE AFFECTED BE-PAL 1723 WITH LOT#: 3000356409 WAS REVIEWED ON 2025-03-28. ACCORDING TO THE DHR RESULTS, THE PRODUCT BE-PAL 1723 PASSED THE DEFINED MANUFACTURING AND FINAL RELEASE SPECIFICATIONS. THERE IS NO INDICATION ON MANUFACTURING ISSUES. THUS, PRODUCTION RELATED INFLUENCES ARE UNLIKELY. FURTHER, THE INCOMING INSPECTION REPORT REVIEW HAS BEEN PERFORMED FOR AFFECTED COMPONENT ¿700000285, 00285# KONNEKTOR 3/8 X 3/8 LL¿. THE INCOMING INSPECTION REPORT (BATCH#: 3000276586) OF THE AFFECTED COMPONENT "700000285 / 00285# KONNEKTOR 3/8 X 3/8 LL" WAS REVIEWED ON 2025-03-27. THE CONNECTOR WERE CHECKED FOR DAMAGES, SCRATCHES, MARKS, RILLS, SINKS, STREAKS, AND CORDS VISUALLY. VISUAL TESTS WERE PASSED AS PER SPECIFICATIONS. DUE TO THIS, MATERIAL RELATED INFLUENCES ARE UNLIKELY. THE REVIEW OF THE NON-CONFORMITIES HAS BEEN PERFORMED. IT DOES NOT SHOW ANY NON-CONFORMITY IN REGARD TO THE BATCH NUMBERS OF REPORTED COMPONENTS WITH RELATED TO THE REPORTED FAILURE. FURTHER, THE RELATED BASIC OPERATION PROCEDURE TRAINING RECORD HAS BEEN CONTROLLED FOR PRODUCTION EMPLOYEES AND THERE COULD NOT BE FOUND ANY NON-CONFORMITIES THAT COULD CAUSE TO REPORTED FAILURE. THE OCCURRENCE RATE WAS CALCULATED FOR THE REPORTED FAILURE AND PRODUCT AND IT WAS DETERMINED THAT THIS IS NOT A SYSTEMIC ISSUE. THEREFORE, NO REMEDIAL ACTION IS REQUIRED. THE OCCURRENCE RATE RELATED TO THE REPORTED ISSUE IS CURRENTLY BEING MONITORED AS PART OF MAQUET CARDIOPULMONARY¿S TRENDING PROGRAM AND ADDITIONAL INVESTIGATIONS OR CORRECTIONS WILL BE IMPLEMENTED IN CASE OF ADVERSE TRENDING. Manufacturer Narrative: NEW INFORMATION WAS RECEIVED ON 2025-04-17 AS THE PATIENT HAS EXPIRED DUE TO SEVERE CARDIOGENIC SHOCK. THE REPORTED FAILURE 'CRACK ON HLS CANNULAE CONNECTOR' DID NOT AFFECT THE PATIENT TREATMENT DIRECTLY. HOWEVER, THE FAMILY MEMBERS OF PATIENT VOLUNTARILY GAVE UP FROM TREATMENT. FURTHER, CUSTOMER CONFIRMS THAT THE BOTH PRODUCTS WERE NOT CHANGED DURING TREATMENT 'BE-PAL 1723 AND PLS 2050'. A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN FURTHER INFORMATION BECOMES AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21721292&pc=BZD | 2518422-2025-103513 | 2025-03-11 04:00 | Death | RESPIRONICS, INC. | 2025-03-28 04:00 | BZD | REP DREAMSTATION AUTO CPAP, DOM - RECRT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH A REP DREAMSTATION AUTO CPAP, DOM - RECRT DEVICE HAD PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ACCORDING TO THE PATIENT'S WIFE, THE PATIENT PASSED AWAY A FEW YEARS AGO. THE PATIENT WAS IN A NURSING HOME, AND SHE DOES NOT KNOW WHAT HAPPENED WITH THE PATIENT'S RECALLED DEVICE AFTER HE PASSED AWAY. SHE WOULD LIKE TO STOP RECEIVING LETTERS BECAUSE SHE DOES NOT HAVE THE RECALLED DEVICE. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21721350&pc=LWP | 2017865-2025-38848 | 2025-03-11 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-03-28 04:00 | LWP | ASSURITY MRI | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P140033 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT DECEASED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. IT WAS REPORTED THAT THE CAUSE OF DEATH WAS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21721389&pc=NVY | 2017865-2025-38849 | 2025-03-11 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-03-28 04:00 | NVY | ISOFLEX OPTIM LEAD | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960030 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT DECEASED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. IT WAS REPORTED THAT THE CAUSE OF DEATH WAS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21721395&pc=NVN | 2017865-2025-38852 | 2025-03-11 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-03-28 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT DECEASED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. IT WAS REPORTED THAT THE CAUSE OF DEATH WAS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21721461&pc=BZD | 2518422-2025-103515 | 2025-03-05 05:00 | Death | RESPIRONICS, INC. | 2025-03-28 04:00 | BZD | PHILIPS CPAP DEVICE | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH A DREAMSTATION CPAP DEVICE. NO CLAIMS WERE MADE THAT THE DEVICE CAUSED THE HUSBAND'S DEATH. FOLLOWING THE HOMECARE PROVIDER'S ADVICE, THE PATIENT'S WIFE INDICATED THAT SHE HAD THROWN AWAY ALL THE RECALLED DEVICES USED BY HER, HER SON, AND HER HUSBAND. SHE WAS INFORMED THAT RETURNING THE DEVICES WAS UNNECESSARY DUE TO THE TIME ELAPSED SINCE THEIR USE. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21721508&pc=DSQ | 2916596-2025-02077 | 2025-01-01 05:00 | Death | THORATEC CORPORATION | 2025-03-28 04:00 | DSQ | VENTRICULAR (ASSIST) BYPASS | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE ¿ANTICOAGULATION STRATEGY WITH HEARTMATE 3¿ THAT THE HEARTMATE 3 (HM3) MIGHT BE ASSOCIATED WITH BLEEDING, STROKE, THROMBOTIC EVENTS (STROKE, SYSTEMIC EMBOLI OR DEVICE THROMBOSIS), GASTROINTESTINAL (GI) BLEEDING, AND DEATH. THIS STUDY REVIEWED CLINICAL OUTCOMES FOR 153 HM3 PATIENTS INCLUDING SEVERAL WHO DID NOT TOLERATE ANTICOAGULATION (AC) THERAPY AND HAVE BEEN OFF AC FOR A MEDIAN OF 755 DAYS. THIS WAS A SINGLE-CENTERED RETROSPECTIVE STUDY OF 153 PATIENTS WITH A HM3 DEVICE COMPARING OUTCOMES IN PATIENTS ON AC AND PATIENTS TAKEN OFF AC. 7 PATIENTS HAD BEEN OFF AC DUE TO RECURRENT BLEEDING EPISODES WHILE ON AC THERAPY. THERE WERE A TOTAL OF 7 HM3 PATIENTS WHO HAD BEEN WITHOUT AC FOR A MEDIAN 755.2 DAYS. MEAN AGE WAS 67.7 YEARS, WITH 57.1% (N=4) BEING MALE. ALL SEVEN PATIENTS¿ AC WAS DISCONTINUED DUE TO PERSISTENT GI BLEEDING. 100% (N=7) OF PATIENTS WHO WERE MANAGED WITHOUT AC SURVIVED GREATER THAN ONE YEAR. ADDITIONALLY, THIS POPULATION REMAINED FREE FROM ADVERSE EVENTS INCLUDING STROKE OR THROMBOEMBOLISM. IN COMPARISON, THERE WERE A TOTAL OF 146 HM3 PATIENTS ON ANTICOAGULATION THERAPY. THE FREEDOM FROM PUMP THROMBOSIS AND STROKE FOR HM3 PATIENTS WITH AC THERAPY WAS 95% (N=138) AS OPPOSED TO PATIENTS WITHOUT AC THERAPY WAS 100% (N=7), P=.33. THIS DATA SUGGESTS THAT HM3 PATIENTS WITHOUT AC DID NOT HAVE AN INCREASE IN THE FREQUENCY OF STROKE AND/OR THROMBOTIC EVENTS COMPARED TO HM3 PATIENTS ON AC. Manufacturer Narrative: SECTION A, D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. SECTION B3: EVENT DATE WAS ESTIMATED TO BE 01JAN2025 AS THE JOURNAL WAS PUBLISHED IN JANUARY 2025, AND NO DATES WERE MENTIONED IN THE PROVIDED TEXT. SECTION D4: DEVICE SERIAL NUMBER AND LOT NUMBER WERE NOT PROVIDED, AND EXPIRATION DATE AND MANUFACTURE DATE (H4) CANNOT BE DETERMINED. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION. JOSEPH, C., RHINEHART, G., AMPADU, J., & STRAW, L. (2025). ANTICOAGULATION STRATEGY WITH HEARTMATE3. JOURNAL OF CARDIAC FAILURE, 31(1), 299¿300. HTTPS://DOI.ORG/10.1016/J.CARDFAIL.2024.10.302. PRISMA HEALTH, COLUMBIA, SC. PRISMA MIDLANDS, COLUMBIA, SC. COLUMBIA, SC. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: CORRECTED DATA: SECTIONS E2 AND E3 CORRECTED. SECTION G2: CORRECTED. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. IT WAS REPORTED THAT THE PATIENTS WHO HAD BEEN TAKEN OFF ANTICOAGULATION HAD A 100% SURVIVAL RATE AFTER THE FIRST 48 MONTHS. THIS IS COMPARED TO AN 89% SURVIVAL RATE AFTER THE FIRST 48 MONTHS FOR PATIENTS WHO REMAINED ON ANTICOAGULATION THERAPY. THIS DATA SUGGESTS THAT HEARTMATE 3 PATIENTS WITHOUT ANTICOAGULATION HAVE AN INCREASED SURVIVAL RATE COMPARED TO HEARTMATE 3 PATIENTS ON ANTICOAGULATION. HOWEVER, THE EFFECTIVENESS AND SAFETY OF ANTICOAGULATION STRATEGIES DEMANDS FURTHER STUDY IN HEARTMATE 3 PATIENTS. NO PRODUCT WAS EVALUATED UNDER THIS COMPLAINT. THE HEARTMATE 3 DEVICE SERIAL NUMBERS, AS WELL AS OTHER SPECIFIC CASE/PATIENT INFORMATION, ARE NOT AVAILABLE. THE SERIAL NUMBER OR OTHER IDENTIFYING INFORMATION OF THE PRODUCT WAS NOT REPORTED AND WAS NOT ABLE TO BE DETERMINED DURING THE INVESTIGATION. THE HEARTMATE 3 LVAS IFU, REV. C IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21721531&pc=NPT | 2015691-2025-02377 | 2025-03-12 04:00 | Death | EDWARDS LIFESCIENCES | 2025-03-28 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Rupture | P140031 |  | 1 | Event Description: AS REPORTED BY OUR CYPRESS AFFILIATES, DURING IMPLANT OF A 23MM SAPIEN 3 ULTRA VALVE IN THE AORTIC POSITION BY TRANSFEMORAL APPROACH, WHEN THE COMMANDER DELIVERY SYSTEM WITH CRIMPED VALVE CROSSED THE NATIVE VALVE AND THE VALVE WAS EXPANDED, THE PATIENT'S PRESSURE DROPPED. DESPITE EFFORTS TO MITIGATE THE SITUATION USING SYRINGES TO PREVENT CARDIAC SWELLING, THE PATIENT'S CONDITION DETERIORATED RAPIDLY. CARDIAC SURGEONS PROCEEDED TO OPEN THE CHEST, BUT THERE WAS NO FLOW IN THE CORONARY ARTERIES, RESULTING IN THE CESSATION OF CARDIAC FUNCTION. THE PATIENT PASSED AWAY AND THE ROOT CAUSE OF THE DEATH WAS ANNULAR RUPTURE. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED FOR A COMPLETION TO THE ENGINEERING EVALUATION. THE EVENT REPORTED IS AN ANTICIPATED IN THE RISK MANAGEMENT DOCUMENTATION FOR TRANSCATHETER HEART VALVE PROCEDURES. A PREVIOUS INVESTIGATION INTO THIS TYPE OF EVENT IS CAPTURED IN AN EDWARDS LIFESCIENCES TECHNICAL SUMMARY AND APPLIES TO THIS COMPLAINT. ADDITIONAL ASSESSMENT OF THE FAILURE MODE IS NOT REQUIRED AT THIS TIME. DUE TO THE UNAVAILABILITY OF THE COMPLAINT DEVICE, ENGINEERING WAS UNABLE TO PERFORM ANY VISUAL INSPECTION, FUNCTIONAL TESTING, OR DIMENSIONAL ANALYSIS. PER THE TECHNICAL SUMMARY: THE IFU, CURRENT RISK MITIGATIONS INCLUDING DESIGN AND MANUFACTURING CONTROLS, AND TRAINING MANUALS HAVE BEEN REVIEWED, AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. ACCORDING TO THE DEVICE TRAINING MANUALS, RISK FACTORS FOR AORTIC DISSECTION, CARDIAC/AORTIC HEMATOMA, OR ANNULAR RUPTURE DURING THE TVR PROCEDURE INCLUDE SIGNIFICANT VALVE OVER SIZING, SEVERELY OBLITERATED SINUSES OF VALSALVA, PORCELAIN AORTA AND/OR PRESENCE OF BULKY CALCIFICATION, AND NARROW CALCIFIED SINOTUBULAR JUNCTION. IN ADDITION, ADVANCED AGE, FEMALE GENDER, SMALL BODY WEIGHT, AND STEROID DEPENDENCY CAN ALSO BE CONTRIBUTING FACTORS. THE SAPIEN TRANSCATHETER HEART VALVE (ALL MODELS) RELIES CALCIUM TO SECURELY ANCHOR IN THE LANDING ZONE. DESPITE THIS BENEFICIAL ASPECT OF CALCIUM, BULKY CALCIUM CAN INCREASE THE RISK OF CALCIFIC NODULE DISPLACEMENT INTO THE VASCULATURE, WHICH CAN LEAD TO VASCULAR INJURY. AT TIMES, THE EXTENT AND DISTRIBUTION OF CALCIUM CAN IMPAIR EASE OF DELIVERY OF THE VALVE, CORRECT POSITIONING OF THE VALVE, DEPLOYMENT OF THE VALVE, AND PROCEDURAL SUCCESS. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST/INDICATE PATIENT CONDITIONS (MODERATE DEGREE OF ANNULAR CALCIFICATION) CAUSED OR CONTRIBUTED TO THIS EVENT BUT A DEFINITE ROOT CAUSE WAS UNABLE TO BE DETERMINED. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME. Manufacturer Narrative: THE EVENT REPORTED IS AN ANTICIPATED IN THE RISK MANAGEMENT DOCUMENTATION FOR TRANSCATHETER HEART VALVE PROCEDURES. A PREVIOUS INVESTIGATION INTO THIS TYPE OF EVENT IS CAPTURED IN AN EDWARDS LIFESCIENCES TECHNICAL SUMMARY AND APPLIES TO THIS COMPLAINT. ADDITIONAL ASSESSMENT OF THE FAILURE MODE IS NOT REQUIRED AT THIS TIME. DUE TO THE UNAVAILABILITY OF THE COMPLAINT DEVICE, ENGINEERING WAS UNABLE TO PERFORM ANY VISUAL INSPECTION, FUNCTIONAL TESTING, OR DIMENSIONAL ANALYSIS. PER THE TECHNICAL SUMMARY: THE IFU, CURRENT RISK MITIGATIONS INCLUDING DESIGN AND MANUFACTURING CONTROLS, AND TRAINING MANUALS HAVE BEEN REVIEWED, AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. ACCORDING TO THE DEVICE TRAINING MANUALS, RISK FACTORS FOR AORTIC DISSECTION, CARDIAC/AORTIC HEMATOMA, OR ANNULAR RUPTURE DURING THE TVR PROCEDURE INCLUDE SIGNIFICANT VALVE OVER SIZING, SEVERELY OBLITERATED SINUSES OF VALSALVA, PORCELAIN AORTA AND/OR PRESENCE OF BULKY CALCIFICATION, AND NARROW CALCIFIED SINOTUBULAR JUNCTION. IN ADDITION, ADVANCED AGE, FEMALE GENDER, SMALL BODY WEIGHT, AND STEROID DEPENDENCY CAN ALSO BE CONTRIBUTING FACTORS. THE SAPIEN TRANSCATHETER HEART VALVE (ALL MODELS) RELIES CALCIUM TO SECURELY ANCHOR IN THE LANDING ZONE. DESPITE THIS BENEFICIAL ASPECT OF CALCIUM, BULKY CALCIUM CAN INCREASE THE RISK OF CALCIFIC NODULE DISPLACEMENT INTO THE VASCULATURE, WHICH CAN LEAD TO VASCULAR INJURY. AT TIMES, THE EXTENT AND DISTRIBUTION OF CALCIUM CAN IMPAIR EASE OF DELIVERY OF THE VALVE, CORRECT POSITIONING OF THE VALVE, DEPLOYMENT OF THE VALVE, AND PROCEDURAL SUCCESS. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. DUE TO A LIMITED AMOUNT OF INFORMATION RECEIVED, IT IS UNKNOWN WHAT COULD HAVE CAUSED OR CONTRIBUTED TO THE EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21721571&pc=OZD | 1220648-2025-27046 | 2025-02-26 05:00 | Death | ABIOMED, INC. | 2025-03-28 04:00 | OZD | IMPELLA CP | Restricted Flow rate; Malposition of Device | Cardiac Arrest | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WITH ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS), THE PATIENT WAS IN CRITICAL CONDITION WHEN THE DECISION WAS MADE TO IMPLANT THE IMPELLA CP DEVICE (PUMP 1). LOW PUMP FLOW AND POSITIONING ISSUE WERE ENCOUNTERED. IMPELLA CP DEVICE PUMP 1 WOULD BE EXPLANTED AND REPLACED WITH ANOTHER IMPELLA CP DEVICE (PUMP 2), WHICH EXPERIENCED LOW PUMP FLOW AS WELL. THE PATIENT WOULD EVENTUALLY PASS AWAY. THE PATIENT WAS ADMITTED TO THE HOSPITAL WITH COMPLAINTS OF CHEST PAIN AND SHORTNESS OF BREATH. THE INITIAL TEST DID NOT SHOW ANY ACUTE EVENT. THEREFORE, THE PATIENT WAS SENT TO THE UNIT AND RETURNED TO THE CATHETERIZATION LAB THE NEXT DAY. LEFT AND RIGHT HEART CATHETERIZATIONS WERE COMPLETED. THE RIGHT HEART SHOWED NORMAL PRESSURES. THE CORONARY ANGIOGRAM REVEALED A LARGE PATENT RIGHT CORONARY ARTERY. THE LEFT SIDE WAS SEVERELY CALCIFIED AND REVEALED LEFT MAIN. THE EJECTION FRACTION WAS ESTIMATED AT 25%. THE PHYSICIAN MADE THE DECISION TO TRY AND FIX THE LEFT MAIN. A GUIDING CATHETER WAS DELIVERED RADIALLY, AND AS THE PHYSICIAN ATTEMPTED TO SEAT THE GUIDE, A DISSECTION TO THE LEFT MAIN ARTERY OCCURRED DUE TO THE SEVERE CALCIFICATION. THE PATIENT RAPIDLY DECLINED, REQUIRING INTUBATION, MAXIMUM PRESSORS AND CARDIOPULMONARY RESUSCITATION (CPR). THE DECISION WAS THEN MADE TO IMPLANT AN IMPELLA CP DEVICE (PUMP 1). CPR WAS BEING DONE ON THE PATIENT FOR THE ENTIRETY OF THE PROCEDURE. THE IMPELLA CP DEVICE (PUMP 1) WAS IMPLANTED, AND SUCTION WAS IMMEDIATELY PRESENT ON THE AUTOMATED IMPELLA CONTROLLER (AIC) AS THE IMPELLA WAS STARTED IN AUTO. A BOLUS OF SALINE WAS GIVEN, AND THE PLACEMENT OF THE DEVICE WAS CONFIRMED UNDER FLUOROSCOPY. FLOWS OF 1.2 LITERS PER MINUTES (L/MIN) WERE THE BEST THAT COULD BE OBTAINED DUE TO CONSTANT SUCTION EVENTS. THE PHYSICIAN UTILIZED THE 7 FRENCH COMPANION SHEATH/SINGLE ACCESS TO DELIVER THE GUIDE AND WIRE ACROSS THE LEFT MAIN. ANGIOPLASTY AND STENTING (WITH TWO DRUG-ELUTING STENTS) OF THE LEFT MAIN AND THE LEFT MAIN OSTIUM WERE DONE. THE DISSECTION WAS NO LONGER PRESENT PER THE REPORT. HOWEVER, THE PATIENT CONTINUED TO DECLINE REQUIRING CPR. IN AN ATTEMPT TO REPOSITION THE IMPELLA CP (PUMP 1), THE CATHETER WAS UNINTENTIONALLY PULLED BACK ACROSS THE VALVE. THE DEVICE WAS EXPLANTED, REMOVED FROM THE BODY, AND THE PHYSICIAN DECIDED TO USE A NEW IMPELLA CP DEVICE (PUMP 2) TO AVOID ANY POSSIBLE THROMBUS FORMATION WITHIN THE PREVIOUS IMPELLA CP DEVICE (PUMP 1) AND HOPES OF BETTER FLOWS. THE NEW IMPELLA CP (PUMP 2) CATHETER WAS PREPPED AND PLACED INTO THE LEFT VENTRICLE IN STANDARD FASHION. AS CPR CONTINUED, THE FLOWS ON THE NEW IMPELLA CP DEVICE (PUMP 2) COULD NOT RISE ABOVE 1.2L/MIN AND CONTINUED SUCTION WAS SEEN. EVENTUALLY, AFTER ALL HEROIC EFFORTS WERE MADE, THE PATIENT DID NOT RECOVER. THE DECISION WAS MADE TO DISCONTINUE FURTHER EFFORTS AND THE PATIENT EXPIRED. MEDICAL SAFETY REVIEW OF THE EVENT NOTED THE PATIENT LIKELY EXPIRED FROM THE LEFT MAIN ARTERY DISSECTION COMPLICATION AND CARDIAC CONDITION AS THE PATIENT WAS IN A STATE OF DECLINE FROM THE VERY BEGINNING OF THE IMPELLA SUPPORT AND UNABLE TO RECOVER. HOWEVER, GIVEN THE IMPELLA CP DEVICE (PUMP 1) INABILITY TO DELIVER (WITH DELAY IN) OPTIMAL FLOW FOR SUPPORT AT ONSET AND MALPOSITIONING AS WELL, THE IMPELLA CP DEVICE (PUMP 1) CANNOT BE DISSOCIATED FOR THE EVENT OUTCOME. THE IMPACT TO THE OVERALL OUTCOME IS UNKNOWN. NOTE: IMPELLA CP DEVICE PUMP 2 WAS IMPLANTED AS CPR CONTINUED. FLOWS COULD NOT RISE ABOVE 1.2L/MIN. HOWEVER, IMPELLA CP DEVICE PUMP 1 HAS BEEN ASSOCIATED WITH THE PATIENT'S DEMISE (FOR REASON NOTED ABOVE). Manufacturer Narrative: THE IMPELLA CP DEVICE WAS DISCARDED BY THE CUSTOMER. THE CLINICAL DETAILS MENTIONED THE PUMP WAS PULLED BACK INADVERTENTLY PAST THE VALVE. THE ROOT CAUSE OF THE POSITIONING ISSUE WAS MOST LIKELY USE RELATED AS EVIDENCED BY CLINICAL DETAILS. THE DATA LOGS CONFIRMED LOW FLOWS FOR MATCHING P-LEVEL. THERE WERE SUCTION ALARMS AND AN IMPELLA POSITION IN VENTRICLE ALARM. THERE WERE NO MOTOR CURRENT SPIKES OUTSIDE P-LEVEL CHANGES. THE CLINICAL DETAILS DO NOT MENTION ANY PATIENT ANATOMY OR VOLUME COMPLICATIONS. THE ROOT CAUSE OF THE LOW PUMP FLOW WAS NOT DETERMINED AS NO PRODUCT WAS RETURNED AND INSUFFICIENT CLINICAL INFORMATION WAS PROVIDED. THIS IS ONE OF TWO DEVICES ASSOCIATED WITH THE EVENT. REFER TO MANUFACTURING REPORT NUMBER 1220648-2025-27037 FOR IMPELLA CP (PUMP 2).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21721770&pc=NPT | 2015691-2025-02385 | 2025-03-07 05:00 | Death | EDWARDS LIFESCIENCES | 2025-03-28 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE SYSTEM | Adverse Event Without Identified Device or Use Problem | Chest Pain; Rupture | P140031 |  | 1 | Event Description: AS REPORTED BY AN EDWARDS LIFESCIENCES AFFILIATE IN ISRAEL, REGARDING AN IMPLANT OF A 26 MM SAPIEN 3 ULTRA VALVE IN AORTIC POSITION BY TRANSFEMORAL APPROACH. DURING THE VALVE DEPLOYMENT, THE PATIENT COMPLAINED ABOUT CHEST PAIN. FLUOROSCOPY IMAGERY SHOWED FLUID LEAKAGE. A SIGNIFICANT AMOUNT OF BLOOD WAS DRAINED AND RESUSCITATION WAS STARTED, BUT THE PATIENT PASSED AWAY. IT WAS SUSPECTED THAT THE CAUSE OF THE LEAK WAS AN ANNULAR RUPTURE AND AORTIC DISSECTION. Manufacturer Narrative: E1: PHONE NUMBER: (B)(6). IMAGERY WAS PROVIDED FOR REVIEW. A 26 MM SAPIEN 3 ULTRA VALVE WAS IMPLANTED IN THE AORTIC POSITION VIA A TRANSFEMORAL APPROACH. PRIOR TO VALVE DEPLOYMENT, FLUOROSCOPY IMAGERY (RECEIVED ONLY MOBILE PHONE VIDEOS/SNAPSHOTS FROM FLUOROSCOPY TO REVIEW) REVELED HORIZONTAL AORTA, COMMANDER DELIVERY SYSTEM AND THE VALVE WERE NOT CENTERED OR COAXIALLY ALIGNED WITH THE ANNULUS, AND A FLAP DISSECTION WAS OBSERVED. BALLOON INFLATION APPEARED SATISFACTORY WITH NO EVIDENCE OF FRAME ASYMMETRY. HOWEVER, AN AORTIC DISSECTION WAS NOTED, ACCOMPANIED BY EVIDENT CONTRAST EXTRAVASATION. DESPITE DRAINING A SIGNIFICANT AMOUNT OF BLOOD AND INITIATING RESUSCITATION, THE PATIENT UNFORTUNATELY PASSED AWAY. WHILE THE AORTIC DISSECTION WAS CONFIRMED, ANNULAR RUPTURE COULD NOT BE VERIFIED DUE TO THE LACK OF AVAILABLE IMAGES. NOTE: 3MENSIO REPORT SHOWS ANNULUS DIMENSION 451.3 MM2 WHICH IS IN THE RECOMMENDED RANGE FOR THE USE OF 26 MM SAPIEN 3 ULTRA VALVE. USE OF 14FR E-SHEATH WAS ADEQUATE FOR THE ACCESS VESSEL DIMENSIONS PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. ACCORDING TO THE DEVICE TRAINING MANUALS, RISK FACTORS FOR AORTIC DISSECTION, CARDIAC/AORTIC HEMATOMA, OR ANNULAR RUPTURE DURING THE TVR PROCEDURE INCLUDE SIGNIFICANT VALVE OVER SIZING, SEVERELY OBLITERATED SINUSES OF VALSALVA, PORCELAIN AORTA AND/OR PRESENCE OF BULKY CALCIFICATION, AND NARROW CALCIFIED SINOTUBULAR JUNCTION. IN ADDITION, ADVANCED AGE, FEMALE GENDER, SMALL BODY WEIGHT, AND STEROID DEPENDENCY CAN ALSO BE CONTRIBUTING FACTORS. THE SAPIEN TRANSCATHETER HEART VALVE (ALL MODELS) RELIES CALCIUM TO SECURELY ANCHOR IN THE LANDING ZONE. DESPITE THIS BENEFICIAL ASPECT OF CALCIUM, BULKY CALCIUM CAN INCREASE THE RISK OF CALCIFIC NODULE DISPLACEMENT INTO THE VASCULATURE, WHICH CAN LEAD TO VASCULAR INJURY. AT TIMES, THE EXTENT AND DISTRIBUTION OF CALCIUM CAN IMPAIR EASE OF DELIVERY OF THE VALVE, CORRECT POSITIONING OF THE VALVE, DEPLOYMENT OF THE VALVE, AND PROCEDURAL SUCCESS. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST/INDICATE PATIENT FACTORS (CALCIFICATION AT THE LEVEL OF NATIVE ANNULUS) CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21722106&pc=NKM | 2135147-2025-01634 | 2016-01-01 05:00 | Death | ABBOTT MEDICAL | 2025-03-28 04:00 | NKM | MITRACLIP® SYSTEM CLIP DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | Myocardial Infarction; Unspecified Tissue Injury | P100009 |  | 1 | Event Description: THE ARTICLE "MITRAL TRANSCATHETER EDGE-TO-EDGE REPAIR AND MANDATORY MECHANICAL CIRCULATORY SUPPORT IN PATIENTS WITH STRUCTURAL SHOCK" WAS REVIEWED. THE ARTICLE PRESENTED A PROSPECTIVE SINGLE CENTER STUDY, TO EVALUATE THE PROCEDURAL SUCCESS AND SURVIVAL WITH M-TEER: PERFORMED WITH MANDATORY MCS. DEVICES MENTIONED INCLUDE MITRACLIP. THE ARTICLE CONCLUDED THAT IN CS COMPLICATED BY SIGNIFICANT MR, M-TEER WITH MANDATORY MCS CAN BE PERFORMED WITH EXCELLENT PROCEDURAL SUCCESS AND IS ASSOCIATED WITH FAVORABLE CLINICAL OUTCOMES. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS DAVID RIZIK AT HONOR HEALTH SCOTTSDALE INSTITUTION WITH CORRESPONDING EMAIL DAVIDRIZIK@AOL.COM]. THE TIME FRAME OF THE STUDY WAS JANUARY 2016 TO AUGUST 2024. A TOTAL OF 30 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH 30 RECEIVED AN ABBOTT DEVICE. AS THIS EVENT IS FROM A LITERATURE REVIEW, THERE IS NO RELEVANT PATIENT INFORMATION (DATE OF BIRTH, AGE, GENDER, WEIGHT, AND MEDICAL HISTORY) TO REPORT. COMORBIDITIES INCLUDED DIABETES, HYPERTENSION, HISTORY OF ATRIAL FIBRILLATION, PREVIOUS STROKE. (B)(6), UNK MITRACLIP: PERI- AND POST-PROCEDURAL COMPLICATIONS INCLUDED TISSUE DAMAGE, MYOCARDIAL INFARCTION, DEATH. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF MITRACLIP DEVICE WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING DIABETES, HYPERTENSION, HISTORY OF ATRIAL FIBRILLATION, PREVIOUS STROKE. COMPLICATIONS REPORTED INCLUDED TISSUE DAMAGE, MYOCARDIAL INFARCTION, DEATH; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. Manufacturer Narrative: THE DEVICE WILL NOT BE RETURNED FOR EVALUATION AS THE DEVICE WAS REPORTEDLY DISCARDED. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21722157&pc=NPT | 2015691-2025-02394 | 2025-03-05 05:00 | Death | EDWARDS LIFESCIENCES | 2025-03-28 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem |  | P140031 |  | 1 | Event Description: IT WAS REPORTED TO THE IMPLANT PATIENT REGISTRY BY NURSE AT THE HOSPITAL THAT THE PATIENT EXPIRED AFTER IMPLANT OF A 9750TF X 26 MM SAPIEN 3 ULTRA VALVE IN THE AORTIC POSITION. PER MEDICAL RECORD, THE PATIENT UNDERWENT THE TAVR SUCCESSFULLY. APPROXIMATELY 30 MINUTES AFTER THE PROCEDURE, THE PATIENT DEVELOPED PROFOUND HYPOTENSION AND EVENTUALLY CARDIAC ARREST. CPR WAS PERFORMED. CORONARY ANGIOGRAM CONFIRMED OCCLUSION OF THE OSTIUM OF THE LEFT MAIN CORONARY ARTERY. THIS WAS STENTED WITH RESTORATION OF TIMI 3 FLOW IN THE LEFT CORONARY ARTERY. HOWEVER, THE PATIENT CONTINUED TO BE IN CARDIOGENIC SHOCK AND DID NOT RESPOND TO CPR ATTEMPTS AND SUBSEQUENTLY DIED. THE CAUSE OF DEATH WAS CARDIOGENIC SHOCK AND ACUTE MYOCARDIAL INFARCTION SECONDARY TO LEFT MAIN ARTERY OCCLUSION. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), CORONARY FLOW OBSTRUCTION IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. THE IFU CAUTIONS THAT SAFETY AND EFFECTIVENESS HAVE NOT BEEN ESTABLISHED FOR PATIENTS WITH BULKY CALCIFIED AORTIC VALVE LEAFLETS NEAR CORONARY OSTIA. IN ADDITION, IT WARNS THAT CAUTION SHOULD BE EXERCISED IN IMPLANTING A BIOPROSTHESIS IN PATIENTS WITH CLINICALLY SIGNIFICANT CORONARY ARTERY DISEASE. CORONARY OBSTRUCTION CAN RESULT IN MYOCARDIAL ISCHEMIA OR INFARCTION DUE TO OBSTRUCTION OF THE CORONARY BLOOD FLOW AND MAY REQUIRE INTERVENTION (E.G., PERCUTANEOUS CORONARY INTERVENTION). MULTIPLE PATIENT FACTORS COULD PUT THE PATIENT AT RISK FOR CORONARY FLOW OBSTRUCTION DURING THE TVR PROCEDURE, INCLUDING SIGNIFICANT UNDERLYING CORONARY ARTERY DISEASE AND BULKY CALCIFICATION OF THE NATIVE LEAFLETS AND ROOT. DISPLACEMENT OF CALCIUM DEPOSITS WITH EMBOLIZATION OF DEBRIS INTO ONE OR MORE OF THE CORONARY ARTERIES, AND/OR AN AORTIC DISSECTION WITH CONTINUITY OF THE RUPTURE INTO THE INTIMA OF ONE OF THE CORONARY OSTIA, CAN RESULT TO A CORONARY OBSTRUCTION. ADDITIONALLY, MINIMAL DISTANCE BETWEEN THE NATIVE ANNULUS AND THE CORONARY OSTIA, BULKY CALCIFICATION, LONG NATIVE LEAFLETS, AND OBLITERATED CORONARY SINUSES. PROCEDURAL FACTORS SUCH AS TORN NATIVE LEAFLET DURING BALLOON VALVULOPLASTY, PLAQUE SHIFT, SIGNIFICANT VALVE OVER SIZING, AND VALVE MALPOSITION COULD ALSO CONTRIBUTE TO THIS COMPLICATION. THE DEVICE TRAINING MANUALS INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVES (ALL MODELS). TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES. THE PHYSICIAN IS INSTRUCTED TO EVALUATE THIS RISK EARLY IN THE PATIENT SCREENING PROCESS IN ALL PATIENTS. THE FOLLOWING FACTORS SHOULD BE CONSIDERED: DEGREE OF CALCIFICATION ON LEAFLETS, LANDING ZONE TO CORONARY OSTIA DISTANCE, LENGTH OF THE NATIVE VALVE LEAFLET, WIDTH OF THE VALSALVA SINUSES, MOVEMENT OF THE LEAFLETS DURING BALLOON VALVULOPLASTY (BV), PATENCY OF CORONARIES DURING BV, AND EXPANDED HEIGHT OF THE INTENDED VALVE. ASSESSMENT FOR CORONARY COMPRESSION RISK PRIOR TO IMPLANTATION IS RECOMMENDED FOR VALVES IMPLANTED IN PULMONARY POSITION. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. WITH THE LIMITED INFORMATION AVAILABLE, THE EXACT CAUSE OF THE CORONARY ARTERY OCCLUSION IS UNKNOWN BUT MAY BE RELATED TO THE MECHANISMS ABOVE. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21722347&pc=DQY | 2124215-2025-19135 | 2025-03-12 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-03-28 04:00 | DQY | WATCHMAN FXD CURVE? ACCESS SYSTEM | Adverse Event Without Identified Device or Use Problem | Air Embolism; Non specific EKG/ECG Changes; Ventricular Fibrillation | K212228 |  | 1 | Event Description: IT WAS REPORTED THAT AIR EMBOLISM OCCURRED. THE PATIENT DIED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED UNDER MONITORED ANESTHESIA CARE (MAC) SEDATION, USING INTRACARDIAC ECHOCARDIOGRAM (ICE) IMAGING. BILATERAL VENOUS ACCESS WAS OBTAINED AND ICE PROBE INSERTED, WITH A 16FR 13CM SHEATH PLACED ON THE RIGHT SIDE AS A WORKING SHEATH. TRANSSEPTAL PUNCTURE WAS COMPLETED. THE ICE PROBE WAS MOVED ON THE LEFT SIDE TO OBTAIN BASELINE IMAGES AND THE, AND A WATCHMAN FXD DOUBLE CURVE ACCESS SYSTEM (WAS) WAS INSERTED AND ADVANCED. THE DILATOR OF THE WAS REMOVED, AND NO BLOOD WAS OBSERVED. AIR WAS NOTICED ON ICE IMAGING IN THE LEFT ATRIUM AND LEFT VENTRICLE. NO AIR WAS VISUALIZED IN THE WAS. ST ELEVATION WAS IMMEDIATELY NOTED AND THE RHYTHM DEGRADED INTO VENTRICULAR FIBRILLATION. CARDIOPULMONARY RESUSCITATION (CPR) WAS INITIATED, THE PATIENT INTUBATED AND CONVERTED TO GENERAL ANESTHESIA. CPR MAINTAINED FOR 35 MINUTES WITH MULTIPLE SHOCKS PERFORMED WITH A DEFIBRILLATOR. THE PATIENT WAS PRONOUNCED DEAD. THE OFFICIAL CAUSE OF DEATH IS UNKNOWN.Event Description: IT WAS REPORTED THAT AIR EMBOLISM OCCURRED. THE PATIENT DIED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED UNDER MONITORED ANESTHESIA CARE (MAC) SEDATION, USING INTRACARDIAC ECHOCARDIOGRAM (ICE) IMAGING. BILATERAL VENOUS ACCESS WAS OBTAINED AND ICE PROBE INSERTED, WITH A 16FR 13CM SHEATH PLACED ON THE RIGHT SIDE AS A WORKING SHEATH. TRANSSEPTAL PUNCTURE WAS COMPLETED. THE ICE PROBE WAS MOVED ON THE LEFT SIDE TO OBTAIN BASELINE IMAGES AND THE, AND A WATCHMAN FXD DOUBLE CURVE ACCESS SYSTEM (WAS) WAS INSERTED AND ADVANCED. THE DILATOR OF THE WAS REMOVED, AND NO BLOOD WAS OBSERVED. AIR WAS NOTICED ON ICE IMAGING IN THE LEFT ATRIUM AND LEFT VENTRICLE. NO AIR WAS VISUALIZED IN THE WAS. ST ELEVATION WAS IMMEDIATELY NOTED AND THE RHYTHM DEGRADED INTO VENTRICULAR FIBRILLATION. CARDIOPULMONARY RESUSCITATION (CPR) WAS INITIATED, THE PATIENT INTUBATED AND CONVERTED TO GENERAL ANESTHESIA. CPR MAINTAINED FOR 35 MINUTES WITH MULTIPLE SHOCKS PERFORMED WITH A DEFIBRILLATOR. THE PATIENT WAS PRONOUNCED DEAD. THE OFFICIAL CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI) # CORRECTED FROM (B)(4) TO (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21722530&pc=NPT | 2135147-2025-01637 | 2025-03-04 05:00 | Death | ABBOTT MEDICAL | 2025-03-28 04:00 | NPT | NAVITOR TRANSCATHETER AORTIC VALVE | Patient-Device Incompatibility; Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Perforation; Pericardial Effusion; Insufficient Information | P190023 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 23 MM NAVITOR VALVE WAS IMPLANTED. ON AN UNKNOWN DATE, THE PATIENT WAS NOTED TO HAVE PASSED AWAY DUE TO AORTIC INJURY.Event Description: SUBSEQUENT TO THE PREVIOUSLY FILED REPORT, ADDITIONAL INFORMATION WAS RECEIVED: IT WAS REPORTED THAT ON (B)(6) 2025, A 23MM NAVITOR VALVE WAS IMPLANTED. THE PATIENT HAD MOSTLY FIBROTIC AORTIC STENOSIS AND OBSTRUCTIVE CARDIOMYOPATHY. THERE WAS NO DIFFICULTY NOTED DURING THE PROCEDURE, NO ALLERGY KNOWN TO METAL, NO ANGULATED OR HORIZONTAL AORTA, NO OVERSIZED PROSTHESIS NOR DILATED AORTA, AND NO PRE- OR POST-DILATION WAS REQUIRED. THREE WEEKS POST PROCEDURE, THE PATIENT BEGAN COMPLAINING OF CONTINUOUS CHEST PAIN. A CT SCAN WAS PERFORMED AND SHOWED THAT THE PATIENT HAD SUFFERED AN AORTIC WALL RUPTURE DUE TO THE PAD GOING THROUGH THE AORTIC WALL. SHORTLY AFTER THE CT SCAN, THE PATIENT WENT INTO SUDDEN CARDIAC DISSOCIATED ARREST AND AN IMMEDIATE PERICARDIOCENTESIS WAS PERFORMED, BUT THERE WAS POOR DRAINAGE. THE PATIENT PASSED AWAY DUE TO THE AORTIC INJURY. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. AN EVENT OF IMPLANT-RELATED TISSUE DAMAGE-AORTIC PERFORATION WAS REPORTED. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED, AND THE PRODUCT MET ALL SPECIFICATIONS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED IMPLANT-RELATED TISSUE DAMAGE-AORTIC PERFORATION WAS DUE TO THE PAD GOING THROUGH THE AORTIC WALL. THE CAUSE OF THE REPORTED DEATH WAS DUE TO THE AORTIC INJURY. THE REPORTED HYPOTENSION WAS A CASCADING EFFECT OF AORTIC INJURY. THE CAUSE OF THE REPORTED PERICARDIAL EFFUSION COULD NOT BE DETERMINED. THE REPORTED HOSPITALIZATION AND UNEXPECTED MEDICAL INTERVENTION WERE A RESULT OF CASE-SPECIFIC CIRCUMSTANCES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO LABELING, DESIGN, OR MANUFACTURING OF THE DEVICE. CODES REMOVED: H6 HEALTH EFFECT-CLINICAL CODE: 4580 INSUFFICIENT INFORMATION. H6 MEDICAL DEVICE PROBLEM CODE: 2993 ADVERSE EVENT WITHOUT IDENTIFIED DEVICE OR USE PROBLEM. CODES ADDED: H6 HEALTH EFFECT-CLINICAL CODE: 3271 PERICARDIAL EFFUSION, 1914 LOW BLOOD PRESSURE, 2001 PERFORATION. H6 HEALTH EFFECT - IMPACT CODE: 4607 HOSPITALIZATION, 4641 UNEXPECTED MEDICAL INTERVENTION. H6 MEDICAL DEVICE PROBLEM CODE: 2682 PATIENT-DEVICE INCOMPATABILITY. Manufacturer Narrative: N/A. IT IS UNKNOWN IF THE DEVICE IS RETURNING FOR ANALYSIS. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21723518&pc=NWX | 3004742232-2025-00066 | 2025-03-03 05:00 | Death | ORBUSNEICH MEDICAL (SHENZHEN) CO., LTD. | 2025-03-28 04:00 | NWX | SCOREFLEX NC SCORING PTCA CATHETER | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Cardiac Arrest; Cardiac Enzyme Elevation; Myocardial Infarction; Obstruction/Occlusion |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PROCEDURE WAS PERCUTANEOUS CORONARY INTERVENTION (PCI) ON OSTIAL LEFT ANTERIOR DESCENDING (LAD)/D LEFT MAIN (LM) ARTERY. UPON GETTING ON TABLE, THE PATIENT BEGAN TO HAVE CONSTRICTIVE PERICARDITIS (CP). WHEN THE PHYSICIAN WAS CANULATING VESSEL WITH A NON-ABBOTT GUIDE CATHETER, THE PATIENT BEGAN TO HAVE ST ELEVATIONS. THE VESSEL WAS WIRED WITH CONTINUING ST ELEVATIONS AND THE PHYSICIAN REQUESTED A 2.0 X 12 COMPLIANT BALLOON. HOWEVER, A 4.0 X 15 SCOREFLEX BALLOON WAS USED AND INFLATED TO 12 ATMOSPHERES. TWO INFLATIONS WERE COMPLETED AND NO INITIAL ANGIOGRAM WAS TAKEN PRIOR TO INFLATION. ANGIOGRAPHY WAS PERFORMED AFTER TWO INFLATIONS AND ALL VESSELS WERE SHUT DOWN EXCEPT FOR LM AND DIAGONAL BRANCH. THE PATIENT HAD ST ELEVATIONS, PRIOR TO ENGAGING THE VESSEL, THAT MEANT THAT THE PATIENT WAS BEGINNING TO INFARCT. PATIENT WENT INTO VENTRICULAR FIBRILLATION AND SHOCKED AT 200J. A NON-ABBOTT BALLOON WAS USED DOWN THE VESSEL. THE PHYSICIAN ATTEMPTED TO WIRE THE CIRCUMFLEX BUT WAS UNABLE TO. THE PATIENT LOST PULSE AND BEGAN TO CODE. IT WAS UNABLE TO REGAIN FLOW TO VESSEL AND OBTAIN RETURN OF SPONTANEOUS CIRCULATION (ROSC). THE PATIENT EXPIRED. IN THE OPINION OF THE PHYSICIAN, THE CAUSE OF THE DEATH WAS THROMBUS IN THE LM THAT EMBOLIZED WHEN THE VESSEL WAS CANNULATED AND THE BALLOON COULD HAVE CONTRIBUTED TO THROMBUS EMBOLIZING, AND THAT ANY BALLOON WOULD HAVE PROGRESSED THE SITUATION. Manufacturer Narrative: D4: PARTIAL UDI PROVIDED. H10: REPORT IS MANUFACTURER REPORT NUMBER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21723701&pc=NKM | 2135147-2025-01644 | 2017-01-01 05:00 | Death | ABBOTT MEDICAL | 2025-03-28 04:00 | NKM | MITRACLIP® SYSTEM CLIP DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | Renal Failure; Insufficient Information | P100009 |  | 1 | Event Description: THE ARTICLE "OUTCOMES OF OCTOGENARIANS UNDERGOING EDGE-TO-EDGE TRANSCATHETER VALVE REPAIR FOR TRICUSPID REGURGITATION: INVERSE PROPENSITY SCORE-WEIGHTED ANALYSIS" WAS REVIEWED. THE ARTICLE PRESENTED A PROSPECTIVE SINGLE CENTER STUDY, TO EVALUATE THE CHARACTERISTICS AND OUTCOMES OF TEER IN OCTOGENARIANS. COMPARED TO THOSE IN PATIENTS UNDER 80 YEARS OLD TREATED AT A REFERRAL TERTIARY TEACHING HOSPITAL. DEVICES MENTIONED INCLUDE TRICLIP AND MITRACLIP. THE ARTICLE CONCLUDED THAT *CONCLUSION*. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS FABRIZIO MONACO AT IRCCS SAN RAFFAELE SCIENTIFIC INSTITUTE, VIA OLGETTINA 60, 20132 MILAN, ITALY, WITH CORRESPONDING EMAIL MONACO.FABRIZIO@HSR.IT]. THE TIME FRAME OF THE STUDY WAS JANUARY 2017 TO SEPTEMBER 2023. A TOTAL OF 101 PATIENTS (36 GREATER THAN OR EQUAL TO 80 YEARS OLD, LESS THAN 80 YEARS OLD) WERE INCLUDED IN THIS STUDY, OF WHICH ALL RECEIVED AN ABBOTT DEVICE. MAJORITY GENDER WAS MALE. COMORBIDITIES INCLUDED DIABETES, CHRONIC RENAL FAILURE, STROKE, PREVIOUS CARDIAC SURGERY, CONCOMITANT MITRACLIP, ATRIAL FIBRILLATION, PACEMAKER LEAD, COPD. PERI AND POST-PROCEDURAL COMPLICATIONS INCLUDED DEATH, HOSPITALIZATION, MEDICATION REQUIRED, SURGICAL INTERVENTION, UNEXPECTED MEDICAL INTERVENTION, BLOOD TRANSFUSION, RENAL FAILURE. Manufacturer Narrative: IT IS UNKNOWN IF THE DEVICE IS RETURNING FOR ANALYSIS. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION ATTACHMENT: ARTICLE TITLED "OUTCOMES OF OCTOGENARIANS UNDERGOING EDGE-TO-EDGE TRANSCATHETER VALVE REPAIR FOR TRICUSPID REGURGITATION: INVERSE PROPENSITY SCORE-WEIGHTED ANALYSIS." Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF MITRACLIP WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING DIABETES, CHRONIC RENAL FAILURE, STROKE, PREVIOUS CARDIAC SURGERY, CONCOMITANT MITRACLIP, ATRIAL FIBRILLATION, PACEMAKER LEAD, COPD. COMPLICATIONS REPORTED INCLUDED DEATH, HOSPITALIZATION, MEDICATION REQUIRED, SURGICAL INTERVENTION, UNEXPECTED MEDICAL INTERVENTION, BLOOD TRANSFUSION, RENAL FAILURE; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. CODES REMOVED: HEALTH EFFECT-CLINICAL CODE: 2041 RENAL FAILURE. CODES ADDED: HEALTH EFFECT-CLINICAL CODE: 4580 INSUFFICIENT INFORMATION. B2: AS THIS IS FROM A LITERATURE REVIEW, THE EVENT DATE IS ESTIMATED. B3: AS THIS IS FROM A LITERATURE REVIEW, THE DEATH DATE IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. ATTACHMENT: ARTICLE TITLED "OUTCOMES OF OCTOGENARIANS UNDERGOING EDGE-TO-EDGE TRANSCATHETER VALVE REPAIR FOR TRICUSPID REGURGITATION: INVERSE PROPENSITY SCORE-WEIGHTED ANALYSIS"
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21723807&pc=OZD | 1220648-2025-27056 | 2024-10-28 04:00 | Death | ABIOMED, INC. | 2025-03-28 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Stroke/CVA; Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. DURING SUPPORT, THE PATIENT EXPERIENCED A DROP IN HEMOGLOBIN. THE PATIENT REPORTEDLY HAD A GASTROINTESTINAL BLEED (GIB) AND TWO UNITS OF PACKED RED BLOOD CELLS WERE ADMINISTERED TO THE PATIENT. HEPARIN WAS IN THE PURGE PRIOR TO THE GIB. A NOTIFICATION WAS RECEIVED VIA ABIOMED¿S LONG-TERM OUTCOME AND QUALITY INDICATOR (LOQI) IMPELLA REGISTRY REGARDING A PATIENT WHO EXPERIENCED VENTRICULAR TACHYCARDIA (VT) DURING IMPELLA CP SUPPORT. THE DEVICE TIP WAS FOUND TO BE SHALLOW, SO IT WAS ADVANCED IN POSITION IT WAS NOTED THAT THE PATIENT RECOVERED WITH NO SEQUELAE. ADDITIONALLY, THE PATIENT EXPERIENCED AN ACUTE LEFT ANTERIOR CEREBRAL ARTERY CEREBRAL VASCULAR ACCIDENT WITH A "UNKNOWN (UNDETERMINED)" RELATIONSHIP TO THE IMPELLA DEVICE USED. IT WAS NOTED THAT THE PATIENT WAS NOT MOVING BILATERAL LOWER EXTREMITIES AND RIGHT UPPER EXTREMITY. COMPUTED TOMOGRAPHY OF THE PATENT'S HEAD AND NECK WAS DONE WHICH REVEALED LEFT COMMON CAROTID ARTIFACT VERSUS OCCLUSION. IT WAS ALSO NOTED THAT THE PATIENT DID NOT QUALIFY FOR ACUTE STROKE THERAPY. THE PATIENT'S STROKE DEFICITS SHOWED NO IMPROVEMENTS. ADDITIONAL INFORMATION NOTED CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA DEVICE AS AN ASSOCIATED FACTOR IN THE PATIENT'S DEMISE. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA IFU: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21723830&pc=OZD | 1220648-2025-27058 | 2024-11-18 05:00 | Death | ABIOMED, INC. | 2025-03-28 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia | P140003 |  | 1 | Event Description: A NOTIFICATION WAS RECEIVED VIA ABIOMED'S LONG-TERM OUTCOME AND QUALITY INDICATOR (LOQI) IMPELLA REGISTRY REGARDING A PATIENT WHO EXPERIENCED LIMB ISCHEMIA DURING IMPELLA CP SUPPORT. THE PATIETN EXPERIENCED A LOSS OF PULSES IN THE RIGHT LOWER EXTREMITY AND IT REQUIRED EMERGENT PLACEMENT OF A DISTAL PERFUSION CANNULA. IT WAS REPORTED THAT CARE WAS EVENTUALLY WITHDRAWN DUE TO THE PATIENT'S NEURO STATUS. THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA DEVICE AS AN ASSOCIATED FACTOR IN THE PATIENT'S DEMISE. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA IFU: IMPELLA CP WITH SMARTASSIST SYSTEM: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): "ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21723993&pc=DSQ | 3007042319-2025-00703 | 2025-03-24 04:00 | Death | HEARTWARE, INC. | 2025-03-28 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Disconnection; Pumping Stopped; Power Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT A VENTRICULAR ASSIST DEVICE (VAD) PATIENT WAS FOUND DECEASED AT HOME, SITTING IN A CHAIR. THE PATIENT WAS HOLDING TWO FULLY CHARGED BATTERIES ONE IN EACH HANDS AND BOTH BATTERIES WERE DISCONNECTED FROM THE CONTROLLER. THE SHOULDER PACK WAS ON THE KNEES OF THE PATIENT. THE PATIENT WAS KNOWN TO BE COMPLIANT AND HAD NO HISTORY OF ISSUES WITH DOUBLE DISCONNECTION OF BATTERIES. THERE WAS NO INTENTION OF SUICIDE REPORTED. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR A CORRECTION TO B.3) DATE OF EVENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR ANALYSIS AND INVESTIGATION COMPLETION. ADDITIONAL PRODUCTS: CONTROLLER 2.0 (B)(6) H3: YES H6: THE CODES PRESENT IN SECTION H6 CORRESPOND TO COMPONENTS/PRODUCTS THAT COMPRISE THE REPORTED EVENT. PRODUCT EVENT SUMMARY: THE PUMP (B)(6) WAS NOT RETURNED FOR EVALUATION AS IT REMAINS IMPLANTED. ONE (1) CONTROLLER (B)(6) WAS RETURNED FOR EVALUATION. VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE RETURNED CONTROLLER IN RELATION TO THE REPORTED EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT AND ANALYSIS ARE NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. THE RETURNED CONTROLLER (B)(6) PASSED VISUAL INSPECTION AND FUNCTIONAL CHECKS. THE INTERNAL VISUAL INSPECTION DETECTED A CRACK ON THE STANDOFF POST 1. THIS IS AN ADDITIONAL FINDING NOT RELATED TO THE REPORTED EVENT. BASED ON AN INVESTIGATION, THE ROOT CAUSE OF THE STANDOFF POST CRACKS WAS DETERMINED TO BE DUE TO MINERAL OIL APPLIED TO THE CONTROLLER¿S INTERNAL MAIN GASKET AND/OR TO THE SILICONE ADHESIVE APPLIED AROUND THE CONTROLLER¿S INTERNAL BATTERY COMING INTO CONTACT WITH THE STANDOFF POST. THE MINERAL OIL AND/OR SILICONE ADHESIVE CONTRIBUTED TO ENVIRONMENTAL STRESS CRACKING. CAPA PR00517125 IS INVESTIGATING THIS ISSUE. LOG FILE ANALYSIS REVEALED (B)(6) WAS THE PRIMARY CONTROLLER IN USE DURING THE REPORTED EVENT. ANALYSIS OF THE EVENT LOG FILE ASSOCIATED WITH (B)(6) REVEALED A CONTROLLER POWER UP WITHOUT AN ASSOCIATED PUMP START EVENT LOGGED ON 2 6/MAR/2025 AT 11:51:17, INDICATING A LOSS OF POWER TO THE CONTROLLER. THE POWER SOURCES LOGGED PRIOR TO THE EVENT WERE (B)(6) IN POWER PORT ONE (PS1) WITH 40% RELATIVE STATE OF CHARGE (RSOC) AND (B)(6) ON POWER PORT TWO (PS2) WITH 22% RSOC. THE POWER SOURCES LOGGED AFTER THE EVENT WERE (B)(6) IN PS1 AND (B)(6) IN PS2. THE CONTROLLER WAS OFF FOR APPROXIMATELY 5 DAYS. NO ANOMALIES WERE RECORDED LEADING UP TO THE EVENT. LOG FILE ANALYSIS DID NOT REVEAL ANY ALARMS WITHIN THE ANALYZED PERIOD. AS A RESULT, THE REPORTED LOSS OF POWER EVENT WAS CONFIRMED. A POSSIBLE ROOT CAUSE OF THE LOSS OF POWER EVENT CAN BE ATTRIBUTED TO THE DOUBLE DISCONNECTION OF POWER SOURCES AS DESCRIBED IN THE EVENT DETAILS. CAPA PR00551638 INVESTIGATED CONTROLLER LOSSES OF POWER. EVEN THOUGH THIS CAPA IS NOW CLOSED, (B)(6) FALLS IN SCOPE OF THIS CAPA. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. ADDITIONAL PRODUCTS: D1: HEARTWARE VENTRICULAR ASSIST SYSTEM- CONTROLLER 2.0 D4: MODEL#: 1420 / CATALOG#: 1420 / EXPIRATION DATE: 30-APRIL-2021 / SERIAL#: (B)(6) D9: NO H3: NO H4: MFG DATE: 29-APRIL-2020 H5: NO H6: THE CODES PRESENT IN SECTION H6 CORRESPOND TO COMPONENTS/PRODUCTS THAT COMPRISE THE REPORTED EVENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21724337&pc=OZD | 1220648-2025-27066 | 2025-03-03 05:00 | Death | ABIOMED, INC. | 2025-03-28 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED 87-YEAR-OLD MALE PATIENT HAD NO DEPLORABLE PULSE IN THEIR RIGHT LEG FOLLOWING IMPELLA CP IMPLANT. THE PATIENT'S LEG WAS COLD TO THE TOUCH AT THE TIME OF THE OBSERVATION. THE PATIENT WAS MONITORED FOR TWO DAYS ON SUPPORT, AT WHICH POINT THE FAMILY DECIDED TO WITHDRAW CARE. THE PATIENT PASSED AWAY. THERE IS NOT ENOUGH EVIDENCE TO DISASSOCIATE THE PATIENT'S PASSING AND THE USE OF THE IMPELLA. Manufacturer Narrative: THE INVESTIGATION OF LIMB ISCHEMIA HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. AS ALL THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE LIMB ISCHEMIA WAS NOT DETERMINED. B.2 INTERVENTION WAS INCORRECTLY REPORTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27066. B.7 ADDED INFORMATION AS IT WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27066. E.1 REVISED US PHONE NUMBER AS IT WAS PREVIOUSLY ENTERED INCORRECTLY ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27066. ADDED FAX NUMBER AS IT WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27066. E.4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27066 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. G.1 ADDED REPORTING CONTACT FACILITY NAME AND FAX NUMBER AS THEY WERE INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27066. H.6 CODE 925 WAS REPORTED INCORRECTLY ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27066 AND A NEW CODE WAS ADDED. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMART ASSIST SYSTEM STATES THE FOLLOWING: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): "ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21724364&pc=OZD | 1220648-2025-27071 | 2025-03-04 05:00 | Death | ABIOMED, INC. | 2025-03-28 04:00 | OZD | IMPELLA CP | Pumping Stopped; Infusion or Flow Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT WITH ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. IT WAS REPORTED THAT THE FOLLOWING THE IMPLANT, OOZING WAS NOTED AT THE IMPELLA SITE. THE HEPARIN DRIP WAS TEMPORARILY HALTED TO MANAGE THE CONDITION. TEN DAYS INTO SUPPORT, THE MOTOR CURRENT SUDDENLY INCREASED AND SATURATED FLOW WAS ENCOUNTERED. THE PUMP STOPPED AFTER FLOWS COULD NOT BE MAINTAINED. THE PATIENT DECOMPENSATED AND THE CARE WAS WITHDRAWN AT THE REQUEST OF THE FAMILY. IT IS UNKNOWN IF THE USE OF THE IMPELLA WAS RELATED TO PATIENT DEATH, THUS THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA PUMP WAS DISCARDED. SHOULD NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. INSTRUCTIONS FOR USE AS FOLLOWS IMPELLA CP WITH SMART ASSIST SYSTEM SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH RISK PERCUTANEOUS CORONARY INTERVENTION SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA CATHETER ¿ACHIEVING AN ACT = 250 SECONDS PRIOR TO REMOVING THE DILATOR WILL HELP PREVENT A THROMBUS FROM ENTERING THE CATHETER AND CAUSING A SUDDEN STOP ON STARTUP.¿ SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA RP WITH SMARTASSIST SYSTEM CATHETER ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA RP WITH SMARTASSIST SYSTEM CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP. MONITOR MOTOR CURRENT AND CONSIDER REPLACING IMPELLA RP WITH SMARTASSIST SYSTEM CATHETER WHENEVER A RISE IN MOTOR CURRENT IS SEEN.¿ IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: TROUBLESHOOTING THE PURGE SYSTEM ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP.¿ SECTION: IMPELLA STOPPED ¿IF THE IMPELLA CATHETER HAS STOPPED SUDDENLY: 1. TRY TO RESTART THE CATHETER AT PREVIOUS P-LEVEL. 2. IF THE IMPELLA DOES NOT RESTART, TRY TO RESTART AT P-2. 3. IF THE IMPELLA DOES NOT RESTART OR STOPS AGAIN, WAIT 1 MINUTE AND TRY TO RESTART AGAIN. 4. IF THE IMPELLA RESTARTS, WEAN DOWN TO P-2 AS THE PATIENT CAN TOLERATE. UNDER THESE CIRCUMSTANCES, CATHETER FUNCTION IS NOT RELIABLE AND THE IMPELLA MAY STOP AGAIN. 5. IF THE IMPELLA DOES NOT RESTART, REMOVE THE IMPELLA FROM THE VENTRICLE AS SOON AS POSSIBLE TO AVOID AORTIC INSUFFICIENCY.¿ IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH RISK PERCUTANEOUS CORONARY INTERVENTION & IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK & IMPELLA RP SMART ASSIST SYSTEM WITH AUTOMATED IMPELLA CONTROLLER & RP FLEX WITH SMART ASSIST SYSTEM WITH AUTOMATED IMPELLA CONTROLLER SECTION: WARNINGS, CAUTIONS, AND PRECAUTIONS ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21725156&pc=NPS | 2135147-2025-01648 | 2017-01-01 05:00 | Death | ABBOTT MEDICAL | 2025-03-28 04:00 | NPS | TRICLIP DELIVERY SYSTEM (TCDS) | Adverse Event Without Identified Device or Use Problem | Renal Failure | P230007 |  | 1 | Event Description: THE ARTICLE "OUTCOMES OF OCTOGENARIANS UNDERGOING EDGE-TO-EDGE TRANSCATHETER VALVE REPAIR FOR TRICUSPID REGURGITATION: INVERSE PROPENSITY SCORE-WEIGHTED ANALYSIS" WAS REVIEWED. THE ARTICLE PRESENTED A PROSPECTIVE SINGLE CENTER STUDY, TO EVALUATE THE CHARACTERISTICS AND OUTCOMES OF TEER IN OCTOGENARIANS COMPARED TO THOSE IN PATIENTS UNDER 80 YEARS OLD TREATED AT A REFERRAL TERTIARY TEACHING HOSPITAL. DEVICES MENTIONED INCLUDE TRICLIP AND MITRACLIP. THE ARTICLE CONCLUDED THAT *CONCLUSION*. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS FABRIZIO MONACO AT IRCCS SAN RAFFAELE SCIENTIFIC INSTITUTE, VIA OLGETTINA 60, 20132 MILAN, ITALY, WITH CORRESPONDING EMAIL MONACO.FABRIZIO@HSR.IT]. THE TIME FRAME OF THE STUDY WAS (B)(6) 2023. A TOTAL OF 101 PATIENTS (36 GREATER THAN OR EQUAL TO 80 YEARS OLD, LESS THAN 80 YEARS OLD) WERE INCLUDED IN THIS STUDY, OF WHICH ALL RECEIVED AN ABBOTT DEVICE. MAJORITY GENDER WAS MALE. COMORBIDITIES INCLUDED DIABETES, CHRONIC RENAL FAILURE, STROKE, PREVIOUS CARDIAC SURGERY, CONCOMITANT MITRACLIP, ATRIAL FIBRILLATION, PACEMAKER LEAD, COPD. PERI AND POST-PROCEDURAL COMPLICATIONS INCLUDED DEATH, HOSPITALIZATION, MEDICATION REQUIRED, SURGICAL INTERVENTION, UNEXPECTED MEDICAL INTERVENTION, BLOOD TRANSFUSION, RENAL FAILURE.Event Description: THE ARTICLE "OUTCOMES OF OCTOGENARIANS UNDERGOING EDGE-TO-EDGE TRANSCATHETER VALVE REPAIR FOR TRICUSPID REGURGITATION: INVERSE PROPENSITY SCORE-WEIGHTED ANALYSIS" WAS REVIEWED. THE ARTICLE PRESENTED A PROSPECTIVE SINGLE CENTER STUDY, TO EVALUATE THE CHARACTERISTICS AND OUTCOMES OF TEER IN OCTOGENARIANS COMPARED TO THOSE IN PATIENTS UNDER 80 YEARS OLD TREATED AT A REFERRAL TERTIARY TEACHING HOSPITAL. DEVICES MENTIONED INCLUDE TRICLIP AND MITRACLIP. THE ARTICLE CONCLUDED THAT *CONCLUSION*. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS FABRIZIO MONACO AT IRCCS SAN RAFFAELE SCIENTIFIC INSTITUTE, VIA OLGETTINA 60, 20132 MILAN, ITALY, WITH CORRESPONDING EMAIL MONACO.FABRIZIO@HSR.IT]. THE TIME FRAME OF THE STUDY WAS JANUARY 2017 TO SEPTEMBER 2023. A TOTAL OF 101 PATIENTS (36 GREATER THAN OR EQUAL TO 80 YEARS OLD, LESS THAN 80 YEARS OLD) WERE INCLUDED IN THIS STUDY, OF WHICH ALL RECEIVED AN ABBOTT DEVICE. MAJORITY GENDER WAS MALE. COMORBIDITIES INCLUDED DIABETES, CHRONIC RENAL FAILURE, STROKE, PREVIOUS CARDIAC SURGERY, CONCOMITANT MITRACLIP, ATRIAL FIBRILLATION, PACEMAKER LEAD, COPD. PERI AND POST-PROCEDURAL COMPLICATIONS INCLUDED DEATH, HOSPITALIZATION, MEDICATION REQUIRED, SURGICAL INTERVENTION, UNEXPECTED MEDICAL INTERVENTION, BLOOD TRANSFUSION, RENAL FAILURE. Manufacturer Narrative: IT IS UNKNOWN IF THE DEVICE IS RETURNING FOR ANALYSIS. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. ATTACHMENT: ARTICLE TITLED "OUTCOMES OF OCTOGENARIANS UNDERGOING EDGE-TO-EDGE TRANSCATHETER VALVE REPAIR FOR TRICUSPID REGURGITATION: INVERSE PROPENSITY SCORE-WEIGHTED ANALYSIS". Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF TRICLIP WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING DIABETES, CHRONIC RENAL FAILURE, STROKE, PREVIOUS CARDIAC SURGERY, CONCOMITANT MITRACLIP, ATRIAL FIBRILLATION, PACEMAKER LEAD, COPD. COMPLICATIONS REPORTED INCLUDED DEATH, HOSPITALIZATION, MEDICATION REQUIRED, SURGICAL INTERVENTION, UNEXPECTED MEDICAL INTERVENTION, BLOOD TRANSFUSION, RENAL FAILURE; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. B2: DATE OF DEATH IS ESTIMATED. B3: DATE OF EVENT IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. ATTACHMENT: ARTICLE TITLED "OUTCOMES OF OCTOGENARIANS UNDERGOING EDGE-TO-EDGE TRANSCATHETER VALVE REPAIR FOR TRICUSPID REGURGITATION: INVERSE PROPENSITY SCORE-WEIGHTED ANALYSIS."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21725157&pc=IYN | 3023245-2025-00041 | 2025-03-18 04:00 | Death | SIEMENS HEALTHINEERS | 2025-03-28 04:00 | IYN | ACUSON P500 ULTRASOUND SYSTEM | Failure to Power Up | Insufficient Information |  |  | 1 | Event Description: IN THIS CASE, THE CUSTOMER REPORTED THAT A PATIENT WAS BROUGHT IN DURING MEDICAL EMERGENCY. WHILE ATTEMPTING TO SCAN THE PATIENT WITH THE ACUSON P500 ULTRASOUND SYSTEMS, THEY DID NOT BOOT-UP. THE PATIENT PASSED AWAY. REQUESTS HAVE BEEN SENT TO GATHER MORE INFORMATION RELATED TO THE EVENT, HOWEVER NO INFORMATION HAS BEEN RECEIVED AS OF TODAY. AN MDR IS BEING FILED CONSERVATIVELY TO REPORT ABOUT THE PATIENT DEATH. THIS REPORT WILL BE SUPPLEMENTED AFTER FURTHER INFORMATION IS RECEIVED. Manufacturer Narrative: AN INVESTIGATION IS ONGOING AT SIEMENS HEALTHINEERS TO IDENTIFY THE ROOT CAUSE OF THE ISSUE. THE REPORT WILL BE SUPPLEMENTED AFTER THE INVESTIGATION IS COMPLETED. REFERENCE # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21725182&pc=DSQ | 2916596-2025-01749 | 2025-02-12 05:00 | Death | THORATEC CORPORATION | 2025-03-28 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2025 DUE TO RIGHT VENTRICLE (RV) FAILURE AND MULTI-SYSTEM ORGAN FAILURE (MSOF). THE DEATH WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED, AND THE DEVICE OPERATED AS EXPECTED. THE PUMP WOULD NOT BE RETURNED. Manufacturer Narrative: A4: PATIENT WEIGHT WAS NOT PROVIDED. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (INCLUDING RIGHT HEART FAILURE) AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6, UNDER ¿RIGHT HEART FAILURE¿, STATES THAT PATIENTS MAY DEVELOP RIGHT VENTRICULAR FAILURE DURING OR SHORTLY AFTER IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21727158&pc=MKJ | 0003015876-2025-00629 | 2024-11-30 05:00 | Death | PHYSIO-CONTROL, INC. - 3015876 | 2025-03-28 04:00 | MKJ | LIFEPAK(R) 15 MONITOR/DEFIBRILLATOR | Failure to Deliver Shock/Stimulation | No Clinical Signs, Symptoms or Conditions | P160026 |  | 1 | Event Description: A CUSTOMER CONTACTED STRYKER TO REPORT THAT THEIR DEVICE FAILED TO DELIVER SHOCK DURING INTERVENTION ON A PATIENT. THE PATIENT ASSOCIATED WITH THE REPORTED EVENT DID NOT SURVIVE. Manufacturer Narrative: DUE TO CHARACTER LIMITATIONS SECTION E1, INITIAL REPORTER PHONE, WAS LEFT BLANK. THE INITIAL REPORTER¿S PHONE NUMBER IS (B)(6). IN COMPLIANCE WITH REGULATION (EU) 2016/679 OF THE EUROPEAN PARLIAMENT AND OF THE COUNCIL, STRYKER COLLECTS ONLY ESSENTIAL AND RELEVANT PATIENT INFORMATION NECESSARY FOR REGULATORY PURPOSES, ENSURING THAT NO DATA CAPABLE OF IDENTIFYING A PERSON, DIRECTLY OR INDIRECTLY, IS GATHERED. PATIENT FIELDS IN WHICH INFORMATION IS NOT REQUESTED OR PROVIDED WERE INTENTIONALLY LEFT BLANK. A STRYKER SERVICE REPRESENTATIVE PERFORMED AN INITIAL EVALUATION OF THE CUSTOMER¿S DEVICE AND WAS NOT ABLE TO DUPLICATE THE REPORTED ISSUE. PROPER DEVICE OPERATION WAS OBSERVED THROUGH FUNCTIONAL AND PERFORMANCE TESTING. THE DEVICE WAS SUBSEQUENTLY RETURNED TO THE CUSTOMER. STRYKER PERFORMED THE CLINICAL REVIEW OF THE REPORTED EVENT AND IT WAS CONCLUDED THAT THE DEVICE USE MAY HAVE CONTRIBUTED TO THE PATIENT'S REPORTED OUTCOME. THE DEVICE WAS TURNED ON AND IN MANUAL MODE. AT DEVICE TIME¿S 2:35:49, 2:36:27, AND2:36:51 THE PATIENT WAS IN A SHOCKABLE RHYTHM AND NO SHOCK WAS ADMINISTERED. THERE IS NO INDICATION ON THE PCO FILE THAT A CHARGE WAS COMPLETED. THE KEY LOGS PROVIDED DO NOT CONTAIN THE EVENT DATE, AND THEREFORE IT CANNOT BE CONFIRMED WHETHER THE CHARGE OR SHOCK BUTTON WAS ACTUALLY PRESSED AT THOSE TIMES. THE DEVICE WAS SWITCHED TO AED MODE AT DEVICE TIME 2:38:32, AND THE PATIENT WAS IN A NON-SHOCKABLE RHYTHM DURING ALL ANALYSIS. AFTER SWITCHING BACK TO MANUAL MODE AT DEVICE TIME 2:51:58, THE PATIENT WAS SUCCESSFULLY SHOCKED TWICE FOR VENTRICULAR FIBRILLATION. SINCE IT CANNOT BE VERIFIED WHETHER THE SHOCK/CHARGE BUTTONS WERE OR WERE NOT PRESSED, DEVICE USE MAY HAVE CONTRIBUTED TO THE PATIENT¿S REPORTED OUTCOME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21727206&pc=PRL | 2017233-2025-06001 | 2025-02-28 05:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-03-28 04:00 | PRL | GORE® VIABAHN® VBX BALLOON EXPANDABLE ENDOPROSTHESIS | Adverse Event Without Identified Device or Use Problem | Thrombosis/Thrombus | P160021 |  | 1 | Event Description: ON (B)(6) 2025, THE PATIENT UNDERWENT TREATMENT OF AN AORTIC ANEURYSM USING A GORE® EXCLUDER® THORACOABDOMINAL BRANCH ENDOPROSTHESIS (TAMBE) WHERE A GORE® VIABAHN® VBX BALLOON EXPANDABLE ENDOPROSTHESIS (VBX DEVICE) WAS USED IN THE RENAL ARTERY. IT WAS REPORTED THAT THE RIGHT RENAL ARTERY WAS CANNULATED AND STENTED WITH A 5MM X 79MM AND 7MM X 39 MM VBX DEVICES, THEN THE LEFT RENAL ARTERY WITH A 5MM X 79MM AND 7MM X 39 MM VBX DEVICES, THEN THE SUPERIOR MESENTERIC ARTERY (SMA) WITH A 6MM X 59MM AND 9MM X 39MM VBX DEVICES AND LAST THE CELIAC ARTERY WITH A 6MM X 59MM VBX DEVICE. THE PROCEDURE FOR THE TAMBE TOOK APPROXIMATELY THREE HOURS. UPON THE FINAL COMPLETION ANGIOGRAM, IT WAS NOTED THAT THE RENAL ARTERIES DID NOT HAVE FLOW. THE PHYSICIAN CANNULATED THE RENAL ARTERIES AND DISCOVERED BOTH RENAL ARTERIES HAD THROMBUS PRESENT. THE PHYSICIAN PREFORMED A THROMBECTOMY WITH PENUMBRA WHERE SHE REESTABLISHED FLOW. INTRAVASCULAR ULTRASOUND (IVUS) WAS USED TO VISUALIZE THE RENAL ARTERIES TO MAKE SURE THE STENTS WERE NOT CRUSHED. EVERYTHING LOOKED APPROPRIATE VIA IVUS. DURING THIS TIME, THE PHYSICIAN DISCOVERED THAT ANESTHESIA HAD BEEN STRUGGLING KEEPING THE ACTIVATED CLOTTING TIME (ACT) THERAPEUTIC. WITH ACT¿S BEING BELOW 200. THE PATIENT WAS SWITCHED OVER TO ARGATROBAN, AT THAT POINT THE PATIENT¿S ACT WENT TO THERAPEUTIC LEVELS. THE PATIENT¿S LABS WERE SENT IN TO DETERMINE IF THE PATIENT HAD HEPARIN-INDUCED THROMBOCYTOPENIA (HITT), THAT CAME BACK NEGATIVE. AS THE SHEATH WAS COMING OUT OF THE GROIN, THE PHYSICIAN NOTICED NO PULSES IN THE ACCESSED LEG. UPON FURTHER INVESTIGATION IT WAS DETERMINED THE LEFT LEG HAD THROMBOSED AS WELL. THE PHYSICIAN PREFORMED A CUT DOWN ON THE ACCESS AND DID A THROMBECTOMY WHERE FLOW WAS REESTABLISHED. ON (B)(6) 2025, THE FIELD SALES ASSOCIATE WAS CALLED BY THE PHYSICIAN AND INFORMED THAT THE PATIENT WAS EXPERIENCING ABDOMINAL PAIN AND HAD A LACTIC ACID OF OVER 6, PATIENT RECEIVED A CTA AND DISCOVERED ALL 4 VISCERAL VESSELS WERE THROMBOSED. THE PHYSICIAN TOOK THE PATIENT BACK TO THE OPERATING ROOM WHERE A THROMBECTOMY WAS PERFORMED OF THE CELIAC ARTERY AND SMA. FLOW WAS ESTABLISHED HOWEVER BRANCH VESSELS HAD THROMBUS STILL PRESENT. THE PATIENT ULTIMATELY WENT INTO RENAL AND LIVER FAILURE AND PASSED AWAY ON (B)(6) 2025. Manufacturer Narrative: SECTION H6 UPDATED TO REFLECT COMPLETION OF INVESTIGATION. H6 ¿ CODE C19: A REVIEW OF MANUFACTURING RECORDS VERIFIED THAT THE LOT INVOLVED IN THIS COMPLAINT MET ALL PRE-RELEASE SPECIFICATIONS. THE DEVICE REMAINS IMPLANTED AND WAS, THEREFORE, NOT AVAILABLE FOR ANALYSIS. NO CLINICAL IMAGES ENABLING DIRECT ASSESSMENT OF PRODUCT PERFORMANCE WERE RETURNED FOR EVALUATION. CAUSE OF THE REPORTED EVENT CANNOT BE ESTABLISHED BASED ON THE INFORMATION REPORTED TO GORE. FURTHER INFORMATION REGARDING THIS EVENT WAS REQUESTED BY GORE, BUT NO FURTHER INFORMATION HAS BEEN REPORTED, THEREFORE THIS INVESTIGATION IS CONSIDERED COMPLETE. Manufacturer Narrative: CBAS® HEPARIN SURFACE INCORPORATES CARMEDA HEPARIN MANUFACTURED FROM HEPARIN SODIUM API, WHICH IS COVALENTLY BOUND TO THE DEVICE SURFACE AND IS ESSENTIALLY NON-ELUTING. H6 - CODE B20: DEVICE REMAINS IMPLANTED, AND NO IMAGES WERE PROVIDED; THEREFORE, DIRECT PRODUCT ANALYSIS WAS NOT POSSIBLE. H6 - CODE C21: RESULTS PENDING COMPLETION OF INVESTIGATION. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21727210&pc=PRL | 2017233-2025-06002 | 2025-02-28 05:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-03-28 04:00 | PRL | GORE® VIABAHN® VBX BALLOON EXPANDABLE ENDOPROSTHESIS | Adverse Event Without Identified Device or Use Problem | Thrombosis/Thrombus | P160021 |  | 1 | Event Description: ON (B)(6) 2025, THE PATIENT UNDERWENT TREATMENT OF AN AORTIC ANEURYSM USING A GORE® EXCLUDER® THORACOABDOMINAL BRANCH ENDOPROSTHESIS (TAMBE) WHERE A GORE® VIABAHN® VBX BALLOON EXPANDABLE ENDOPROSTHESIS (VBX DEVICE) WAS USED IN THE RENAL ARTERY. IT WAS REPORTED THAT THE RIGHT RENAL ARTERY WAS CANNULATED AND STENTED WITH A 5MM X 79MM AND 7MM X 39 MM VBX DEVICES, THEN THE LEFT RENAL ARTERY WITH A 5MM X 79MM AND 7MM X 39 MM VBX DEVICES, THEN THE SUPERIOR MESENTERIC ARTERY (SMA) WITH A 6MM X 59MM AND 9MM X 39MM VBX DEVICES AND LAST THE CELIAC ARTERY WITH A 6MM X 59MM VBX DEVICE. THE PROCEDURE FOR THE TAMBE TOOK APPROXIMATELY THREE HOURS. UPON THE FINAL COMPLETION ANGIOGRAM, IT WAS NOTED THAT THE RENAL ARTERIES DID NOT HAVE FLOW. THE PHYSICIAN CANNULATED THE RENAL ARTERIES AND DISCOVERED BOTH RENAL ARTERIES HAD THROMBUS PRESENT. THE PHYSICIAN PREFORMED A THROMBECTOMY WITH PENUMBRA WHERE SHE REESTABLISHED FLOW. INTRAVASCULAR ULTRASOUND (IVUS) WAS USED TO VISUALIZE THE RENAL ARTERIES TO MAKE SURE THE STENTS WERE NOT CRUSHED. EVERYTHING LOOKED APPROPRIATE VIA IVUS. DURING THIS TIME, THE PHYSICIAN DISCOVERED THAT ANESTHESIA HAD BEEN STRUGGLING KEEPING THE ACTIVATED CLOTTING TIME (ACT) THERAPEUTIC. WITH ACT¿S BEING BELOW 200. THE PATIENT WAS SWITCHED OVER TO ARGATROBAN, AT THAT POINT THE PATIENT¿S ACT WENT TO THERAPEUTIC LEVELS. THE PATIENT¿S LABS WERE SENT IN TO DETERMINE IF THE PATIENT HAD HEPARIN-INDUCED THROMBOCYTOPENIA (HITT), THAT CAME BACK NEGATIVE. AS THE SHEATH WAS COMING OUT OF THE GROIN, THE PHYSICIAN NOTICED NO PULSES IN THE ACCESSED LEG. UPON FURTHER INVESTIGATION IT WAS DETERMINED THE LEFT LEG HAD THROMBOSED AS WELL. THE PHYSICIAN PREFORMED A CUT DOWN ON THE ACCESS AND DID A THROMBECTOMY WHERE FLOW WAS REESTABLISHED. ON MARCH 04, 2025, THE FIELD SALES ASSOCIATE WAS CALLED BY THE PHYSICIAN AND INFORMED THAT THE PATIENT WAS EXPERIENCING ABDOMINAL PAIN AND HAD A LACTIC ACID OF OVER 6, PATIENT RECEIVED A CTA AND DISCOVERED ALL 4 VISCERAL VESSELS WERE THROMBOSED. THE PHYSICIAN TOOK THE PATIENT BACK TO THE OPERATING ROOM WHERE A THROMBECTOMY WAS PERFORMED OF THE CELIAC ARTERY AND SMA. FLOW WAS ESTABLISHED HOWEVER BRANCH VESSELS HAD THROMBUS STILL PRESENT. THE PATIENT ULTIMATELY WENT INTO RENAL AND LIVER FAILURE AND PASSED AWAY ON (B)(6) 2025. Manufacturer Narrative: SECTION H6 UPDATED TO REFLECT COMPLETION OF INVESTIGATION. H6 ¿ CODE C19: A REVIEW OF MANUFACTURING RECORDS VERIFIED THAT THE LOT INVOLVED IN THIS COMPLAINT MET ALL PRE-RELEASE SPECIFICATIONS. THE DEVICE REMAINS IMPLANTED AND WAS, THEREFORE, NOT AVAILABLE FOR ANALYSIS. NO CLINICAL IMAGES ENABLING DIRECT ASSESSMENT OF PRODUCT PERFORMANCE WERE RETURNED FOR EVALUATION. CAUSE OF THE REPORTED EVENT CANNOT BE ESTABLISHED BASED ON THE INFORMATION REPORTED TO GORE. FURTHER INFORMATION REGARDING THIS EVENT WAS REQUESTED BY GORE, BUT NO FURTHER INFORMATION HAS BEEN REPORTED, THEREFORE THIS INVESTIGATION IS CONSIDERED COMPLETE. Manufacturer Narrative: CBAS® HEPARIN SURFACE INCORPORATES CARMEDA HEPARIN MANUFACTURED FROM HEPARIN SODIUM API, WHICH IS COVALENTLY BOUND TO THE DEVICE SURFACE AND IS ESSENTIALLY NON-ELUTING. H6 - CODE B20: DEVICE REMAINS IMPLANTED, AND NO IMAGES WERE PROVIDED; THEREFORE, DIRECT PRODUCT ANALYSIS WAS NOT POSSIBLE. H6 - CODE C21: RESULTS PENDING COMPLETION OF INVESTIGATION. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21727219&pc=PRL | 2017233-2025-06003 | 2025-02-28 05:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-03-28 04:00 | PRL | GORE® VIABAHN® VBX BALLOON EXPANDABLE ENDOPROSTHESIS | Adverse Event Without Identified Device or Use Problem | Thrombosis/Thrombus | P160021 |  | 1 | Event Description: ON (B)(6) 2025, THE PATIENT UNDERWENT TREATMENT OF AN AORTIC ANEURYSM USING A GORE® EXCLUDER® THORACOABDOMINAL BRANCH ENDOPROSTHESIS (TAMBE) WHERE A GORE® VIABAHN® VBX BALLOON EXPANDABLE ENDOPROSTHESIS (VBX DEVICE) WAS USED IN THE RENAL ARTERY. IT WAS REPORTED THAT THE RIGHT RENAL ARTERY WAS CANNULATED AND STENTED WITH A 5MM X 79MM AND 7MM X 39 MM VBX DEVICES, THEN THE LEFT RENAL ARTERY WITH A 5MM X 79MM AND 7MM X 39 MM VBX DEVICES, THEN THE SUPERIOR MESENTERIC ARTERY (SMA) WITH A 6MM X 59MM AND 9MM X 39MM VBX DEVICES AND LAST THE CELIAC ARTERY WITH A 6MM X 59MM VBX DEVICE. THE PROCEDURE FOR THE TAMBE TOOK APPROXIMATELY THREE HOURS. UPON THE FINAL COMPLETION ANGIOGRAM, IT WAS NOTED THAT THE RENAL ARTERIES DID NOT HAVE FLOW. THE PHYSICIAN CANNULATED THE RENAL ARTERIES AND DISCOVERED BOTH RENAL ARTERIES HAD THROMBUS PRESENT. THE PHYSICIAN PREFORMED A THROMBECTOMY WITH PENUMBRA WHERE SHE REESTABLISHED FLOW. INTRAVASCULAR ULTRASOUND (IVUS) WAS USED TO VISUALIZE THE RENAL ARTERIES TO MAKE SURE THE STENTS WERE NOT CRUSHED. EVERYTHING LOOKED APPROPRIATE VIA IVUS. DURING THIS TIME, THE PHYSICIAN DISCOVERED THAT ANESTHESIA HAD BEEN STRUGGLING KEEPING THE ACTIVATED CLOTTING TIME (ACT) THERAPEUTIC. WITH ACT¿S BEING BELOW 200. THE PATIENT WAS SWITCHED OVER TO ARGATROBAN, AT THAT POINT THE PATIENT¿S ACT WENT TO THERAPEUTIC LEVELS. THE PATIENT¿S LABS WERE SENT IN TO DETERMINE IF THE PATIENT HAD HEPARIN-INDUCED THROMBOCYTOPENIA (HITT), THAT CAME BACK NEGATIVE. AS THE SHEATH WAS COMING OUT OF THE GROIN, THE PHYSICIAN NOTICED NO PULSES IN THE ACCESSED LEG. UPON FURTHER INVESTIGATION IT WAS DETERMINED THE LEFT LEG HAD THROMBOSED AS WELL. THE PHYSICIAN PREFORMED A CUT DOWN ON THE ACCESS AND DID A THROMBECTOMY WHERE FLOW WAS REESTABLISHED. ON (B)(6) 2025, THE FIELD SALES ASSOCIATE WAS CALLED BY THE PHYSICIAN AND INFORMED THAT THE PATIENT WAS EXPERIENCING ABDOMINAL PAIN AND HAD A LACTIC ACID OF OVER 6, PATIENT RECEIVED A CTA AND DISCOVERED ALL 4 VISCERAL VESSELS WERE THROMBOSED. THE PHYSICIAN TOOK THE PATIENT BACK TO THE OPERATING ROOM WHERE A THROMBECTOMY WAS PERFORMED OF THE CELIAC ARTERY AND SMA. FLOW WAS ESTABLISHED HOWEVER BRANCH VESSELS HAD THROMBUS STILL PRESENT. THE PATIENT ULTIMATELY WENT INTO RENAL AND LIVER FAILURE AND PASSED AWAY ON (B)(6) 2025. Manufacturer Narrative: CBAS® HEPARIN SURFACE INCORPORATES CARMEDA HEPARIN MANUFACTURED FROM HEPARIN SODIUM API, WHICH IANTED, AND NO IMAGES WERE PROVIDED; THEREFORE, DIRECT PRODUCT ANALYSIS WAS NOT POSSIBLE. S COVALENTLY BOUND TO THE DEVICE SURFACE AND IS ESSENTIALLY NON-ELUTING. H6 - CODE B20: DEVICE REMAINS IMPL H6 - CODE C21: RESULTS PENDING COMPLETION OF INVESTIGATION. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY HAVE NOT BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. Manufacturer Narrative: SECTION H6 UPDATED TO REFLECT COMPLETION OF INVESTIGATION. H6 ¿ CODE C19: A REVIEW OF MANUFACTURING RECORDS VERIFIED THAT THE LOT INVOLVED IN THIS COMPLAINT MET ALL PRE-RELEASE SPECIFICATIONS. THE DEVICE REMAINS IMPLANTED AND WAS, THEREFORE, NOT AVAILABLE FOR ANALYSIS. NO CLINICAL IMAGES ENABLING DIRECT ASSESSMENT OF PRODUCT PERFORMANCE WERE RETURNED FOR EVALUATION. CAUSE OF THE REPORTED EVENT CANNOT BE ESTABLISHED BASED ON THE INFORMATION REPORTED TO GORE. FURTHER INFORMATION REGARDING THIS EVENT WAS REQUESTED BY GORE, BUT NO FURTHER INFORMATION HAS BEEN REPORTED, THEREFORE THIS INVESTIGATION IS CONSIDERED COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21727227&pc=PRL | 2017233-2025-06004 | 2025-02-28 05:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-03-28 04:00 | PRL | GORE® VIABAHN® VBX BALLOON EXPANDABLE ENDOPROSTHESIS | Adverse Event Without Identified Device or Use Problem | Thrombosis/Thrombus | P160021 |  | 1 | Event Description: ON (B)(6) 2025, THE PATIENT UNDERWENT TREATMENT OF AN AORTIC ANEURYSM USING A GORE® EXCLUDER® THORACOABDOMINAL BRANCH ENDOPROSTHESIS (TAMBE) WHERE A GORE® VIABAHN® VBX BALLOON EXPANDABLE ENDOPROSTHESIS (VBX DEVICE) WAS USED IN THE RENAL ARTERY. IT WAS REPORTED THAT THE RIGHT RENAL ARTERY WAS CANNULATED AND STENTED WITH A 5MM X 79MM AND 7MM X 39 MM VBX DEVICES, THEN THE LEFT RENAL ARTERY WITH A 5MM X 79MM AND 7MM X 39 MM VBX DEVICES, THEN THE SUPERIOR MESENTERIC ARTERY (SMA) WITH A 6MM X 59MM AND 9MM X 39MM VBX DEVICES AND LAST THE CELIAC ARTERY WITH A 6MM X 59MM VBX DEVICE. THE PROCEDURE FOR THE TAMBE TOOK APPROXIMATELY THREE HOURS. UPON THE FINAL COMPLETION ANGIOGRAM, IT WAS NOTED THAT THE RENAL ARTERIES DID NOT HAVE FLOW. THE PHYSICIAN CANNULATED THE RENAL ARTERIES AND DISCOVERED BOTH RENAL ARTERIES HAD THROMBUS PRESENT. THE PHYSICIAN PREFORMED A THROMBECTOMY WITH PENUMBRA WHERE SHE REESTABLISHED FLOW. INTRAVASCULAR ULTRASOUND (IVUS) WAS USED TO VISUALIZE THE RENAL ARTERIES TO MAKE SURE THE STENTS WERE NOT CRUSHED. EVERYTHING LOOKED APPROPRIATE VIA IVUS. DURING THIS TIME, THE PHYSICIAN DISCOVERED THAT ANESTHESIA HAD BEEN STRUGGLING KEEPING THE ACTIVATED CLOTTING TIME (ACT) THERAPEUTIC. WITH ACT¿S BEING BELOW 200. THE PATIENT WAS SWITCHED OVER TO ARGATROBAN, AT THAT POINT THE PATIENT¿S ACT WENT TO THERAPEUTIC LEVELS. THE PATIENT¿S LABS WERE SENT IN TO DETERMINE IF THE PATIENT HAD HEPARIN-INDUCED THROMBOCYTOPENIA (HITT), THAT CAME BACK NEGATIVE. AS THE SHEATH WAS COMING OUT OF THE GROIN, THE PHYSICIAN NOTICED NO PULSES IN THE ACCESSED LEG. UPON FURTHER INVESTIGATION IT WAS DETERMINED THE LEFT LEG HAD THROMBOSED AS WELL. THE PHYSICIAN PREFORMED A CUT DOWN ON THE ACCESS AND DID A THROMBECTOMY WHERE FLOW WAS REESTABLISHED. ON MARCH 04, 2025, THE FIELD SALES ASSOCIATE WAS CALLED BY THE PHYSICIAN AND INFORMED THAT THE PATIENT WAS EXPERIENCING ABDOMINAL PAIN AND HAD A LACTIC ACID OF OVER 6, PATIENT RECEIVED A CTA AND DISCOVERED ALL 4 VISCERAL VESSELS WERE THROMBOSED. THE PHYSICIAN TOOK THE PATIENT BACK TO THE OPERATING ROOM WHERE A THROMBECTOMY WAS PERFORMED OF THE CELIAC ARTERY AND SMA. FLOW WAS ESTABLISHED HOWEVER BRANCH VESSELS HAD THROMBUS STILL PRESENT. THE PATIENT ULTIMATELY WENT INTO RENAL AND LIVER FAILURE AND PASSED AWAY ON (B)(6) 2025. Manufacturer Narrative: CBAS® HEPARIN SURFACE INCORPORATES CARMEDA HEPARIN MANUFACTURED FROM HEPARIN SODIUM API, WHICH IS COVALENTLY BOUND TO THE DEVICE SURFACE AND IS ESSENTIALLY NON-ELUTING. H6 - CODE B20: DEVICE REMAINS IMPLANTED, AND NO IMAGES WERE PROVIDED; THEREFORE, DIRECT PRODUCT ANALYSIS WAS NOT POSSIBLE. H6 - CODE C21: RESULTS PENDING COMPLETION OF INVESTIGATION. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. Manufacturer Narrative: SECTION H6 UPDATED TO REFLECT COMPLETION OF INVESTIGATION. H6 ¿ CODE C19: A REVIEW OF MANUFACTURING RECORDS VERIFIED THAT THE LOT INVOLVED IN THIS COMPLAINT MET ALL PRE-RELEASE SPECIFICATIONS. THE DEVICE REMAINS IMPLANTED AND WAS, THEREFORE, NOT AVAILABLE FOR ANALYSIS. NO CLINICAL IMAGES ENABLING DIRECT ASSESSMENT OF PRODUCT PERFORMANCE WERE RETURNED FOR EVALUATION. CAUSE OF THE REPORTED EVENT CANNOT BE ESTABLISHED BASED ON THE INFORMATION REPORTED TO GORE. FURTHER INFORMATION REGARDING THIS EVENT WAS REQUESTED BY GORE, BUT NO FURTHER INFORMATION HAS BEEN REPORTED, THEREFORE THIS INVESTIGATION IS CONSIDERED COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21727234&pc=PRL | 2017233-2025-06005 | 2025-02-28 05:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-03-28 04:00 | PRL | GORE® VIABAHN® VBX BALLOON EXPANDABLE ENDOPROSTHESIS | Adverse Event Without Identified Device or Use Problem | Thrombosis/Thrombus | P160021 |  | 1 | Event Description: ON (B)(6) 2025, THE PATIENT UNDERWENT TREATMENT OF AN AORTIC ANEURYSM USING A GORE® EXCLUDER® THORACOABDOMINAL BRANCH ENDOPROSTHESIS (TAMBE) WHERE A GORE® VIABAHN® VBX BALLOON EXPANDABLE ENDOPROSTHESIS (VBX DEVICE) WAS USED IN THE RENAL ARTERY. IT WAS REPORTED THAT THE RIGHT RENAL ARTERY WAS CANNULATED AND STENTED WITH A 5MM X 79MM AND 7MM X 39 MM VBX DEVICES, THEN THE LEFT RENAL ARTERY WITH A 5MM X 79MM AND 7MM X 39 MM VBX DEVICES, THEN THE SUPERIOR MESENTERIC ARTERY (SMA) WITH A 6MM X 59MM AND 9MM X 39MM VBX DEVICES AND LAST THE CELIAC ARTERY WITH A 6MM X 59MM VBX DEVICE. THE PROCEDURE FOR THE TAMBE TOOK APPROXIMATELY THREE HOURS. UPON THE FINAL COMPLETION ANGIOGRAM, IT WAS NOTED THAT THE RENAL ARTERIES DID NOT HAVE FLOW. THE PHYSICIAN CANNULATED THE RENAL ARTERIES AND DISCOVERED BOTH RENAL ARTERIES HAD THROMBUS PRESENT. THE PHYSICIAN PREFORMED A THROMBECTOMY WITH PENUMBRA WHERE SHE REESTABLISHED FLOW. INTRAVASCULAR ULTRASOUND (IVUS) WAS USED TO VISUALIZE THE RENAL ARTERIES TO MAKE SURE THE STENTS WERE NOT CRUSHED. EVERYTHING LOOKED APPROPRIATE VIA IVUS. DURING THIS TIME, THE PHYSICIAN DISCOVERED THAT ANESTHESIA HAD BEEN STRUGGLING KEEPING THE ACTIVATED CLOTTING TIME (ACT) THERAPEUTIC. WITH ACT¿S BEING BELOW 200. THE PATIENT WAS SWITCHED OVER TO ARGATROBAN, AT THAT POINT THE PATIENT¿S ACT WENT TO THERAPEUTIC LEVELS. THE PATIENT¿S LABS WERE SENT IN TO DETERMINE IF THE PATIENT HAD HEPARIN-INDUCED THROMBOCYTOPENIA (HITT), THAT CAME BACK NEGATIVE. AS THE SHEATH WAS COMING OUT OF THE GROIN, THE PHYSICIAN NOTICED NO PULSES IN THE ACCESSED LEG. UPON FURTHER INVESTIGATION IT WAS DETERMINED THE LEFT LEG HAD THROMBOSED AS WELL. THE PHYSICIAN PREFORMED A CUT DOWN ON THE ACCESS AND DID A THROMBECTOMY WHERE FLOW WAS REESTABLISHED. ON MARCH 04, 2025, THE FIELD SALES ASSOCIATE WAS CALLED BY THE PHYSICIAN AND INFORMED THAT THE PATIENT WAS EXPERIENCING ABDOMINAL PAIN AND HAD A LACTIC ACID OF OVER 6, PATIENT RECEIVED A CTA AND DISCOVERED ALL 4 VISCERAL VESSELS WERE THROMBOSED. THE PHYSICIAN TOOK THE PATIENT BACK TO THE OPERATING ROOM WHERE A THROMBECTOMY WAS PERFORMED OF THE CELIAC ARTERY AND SMA. FLOW WAS ESTABLISHED HOWEVER BRANCH VESSELS HAD THROMBUS STILL PRESENT. THE PATIENT ULTIMATELY WENT INTO RENAL AND LIVER FAILURE AND PASSED AWAY ON (B)(6) 2025. Manufacturer Narrative: CBAS® HEPARIN SURFACE INCORPORATES CARMEDA HEPARIN MANUFACTURED FROM HEPARIN SODIUM API, WHICH IS COVALENTLY BOUND TO THE DEVICE SURFACE AND IS ESSENTIALLY NON-ELUTING. H6 - CODE B20: DEVICE REMAINS IMPLANTED, AND NO IMAGES WERE PROVIDED; THEREFORE, DIRECT PRODUCT ANALYSIS WAS NOT POSSIBLE. H6 - CODE C21: RESULTS PENDING COMPLETION OF INVESTIGATION. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. Manufacturer Narrative: SECTION H6 UPDATED TO REFLECT COMPLETION OF INVESTIGATION. H6 ¿ CODE C19: A REVIEW OF MANUFACTURING RECORDS VERIFIED THAT THE LOT INVOLVED IN THIS COMPLAINT MET ALL PRE-RELEASE SPECIFICATIONS. THE DEVICE REMAINS IMPLANTED AND WAS, THEREFORE, NOT AVAILABLE FOR ANALYSIS. NO CLINICAL IMAGES ENABLING DIRECT ASSESSMENT OF PRODUCT PERFORMANCE WERE RETURNED FOR EVALUATION. CAUSE OF THE REPORTED EVENT CANNOT BE ESTABLISHED BASED ON THE INFORMATION REPORTED TO GORE. FURTHER INFORMATION REGARDING THIS EVENT WAS REQUESTED BY GORE, BUT NO FURTHER INFORMATION HAS BEEN REPORTED, THEREFORE THIS INVESTIGATION IS CONSIDERED COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21727237&pc=PRL | 2017233-2025-06006 | 2025-02-28 05:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-03-28 04:00 | PRL | GORE® VIABAHN® VBX BALLOON EXPANDABLE ENDOPROSTHESIS | Adverse Event Without Identified Device or Use Problem | Thrombosis/Thrombus | P160021 |  | 1 | Event Description: ON (B)(6) 2025, THE PATIENT UNDERWENT TREATMENT OF AN AORTIC ANEURYSM USING A GORE® EXCLUDER® THORACOABDOMINAL BRANCH ENDOPROSTHESIS (TAMBE) WHERE A GORE® VIABAHN® VBX BALLOON EXPANDABLE ENDOPROSTHESIS (VBX DEVICE) WAS USED IN THE RENAL ARTERY. IT WAS REPORTED THAT THE RIGHT RENAL ARTERY WAS CANNULATED AND STENTED WITH A 5MM X 79MM AND 7MM X 39 MM VBX DEVICES, THEN THE LEFT RENAL ARTERY WITH A 5MM X 79MM AND 7MM X 39 MM VBX DEVICES, THEN THE SUPERIOR MESENTERIC ARTERY (SMA) WITH A 6MM X 59MM AND 9MM X 39MM VBX DEVICES AND LAST THE CELIAC ARTERY WITH A 6MM X 59MM VBX DEVICE. THE PROCEDURE FOR THE TAMBE TOOK APPROXIMATELY THREE HOURS. UPON THE FINAL COMPLETION ANGIOGRAM, IT WAS NOTED THAT THE RENAL ARTERIES DID NOT HAVE FLOW. THE PHYSICIAN CANNULATED THE RENAL ARTERIES AND DISCOVERED BOTH RENAL ARTERIES HAD THROMBUS PRESENT. THE PHYSICIAN PREFORMED A THROMBECTOMY WITH PENUMBRA WHERE SHE REESTABLISHED FLOW. INTRAVASCULAR ULTRASOUND (IVUS) WAS USED TO VISUALIZE THE RENAL ARTERIES TO MAKE SURE THE STENTS WERE NOT CRUSHED. EVERYTHING LOOKED APPROPRIATE VIA IVUS. DURING THIS TIME, THE PHYSICIAN DISCOVERED THAT ANESTHESIA HAD BEEN STRUGGLING KEEPING THE ACTIVATED CLOTTING TIME (ACT) THERAPEUTIC. WITH ACT¿S BEING BELOW 200. THE PATIENT WAS SWITCHED OVER TO ARGATROBAN, AT THAT POINT THE PATIENT¿S ACT WENT TO THERAPEUTIC LEVELS. THE PATIENT¿S LABS WERE SENT IN TO DETERMINE IF THE PATIENT HAD HEPARIN-INDUCED THROMBOCYTOPENIA (HITT), THAT CAME BACK NEGATIVE. AS THE SHEATH WAS COMING OUT OF THE GROIN, THE PHYSICIAN NOTICED NO PULSES IN THE ACCESSED LEG. UPON FURTHER INVESTIGATION IT WAS DETERMINED THE LEFT LEG HAD THROMBOSED AS WELL. THE PHYSICIAN PREFORMED A CUT DOWN ON THE ACCESS AND DID A THROMBECTOMY WHERE FLOW WAS REESTABLISHED. ON (B)(6) 2025, THE FIELD SALES ASSOCIATE WAS CALLED BY THE PHYSICIAN AND INFORMED THAT THE PATIENT WAS EXPERIENCING ABDOMINAL PAIN AND HAD A LACTIC ACID OF OVER 6, PATIENT RECEIVED A CTA AND DISCOVERED ALL 4 VISCERAL VESSELS WERE THROMBOSED. THE PHYSICIAN TOOK THE PATIENT BACK TO THE OPERATING ROOM WHERE A THROMBECTOMY WAS PERFORMED OF THE CELIAC ARTERY AND SMA. FLOW WAS ESTABLISHED HOWEVER BRANCH VESSELS HAD THROMBUS STILL PRESENT. THE PATIENT ULTIMATELY WENT INTO RENAL AND LIVER FAILURE AND PASSED AWAY ON (B)(6)2025. Manufacturer Narrative: SECTION H6 UPDATED TO REFLECT COMPLETION OF INVESTIGATION. H6 ¿ CODE C19: A REVIEW OF MANUFACTURING RECORDS VERIFIED THAT THE LOT INVOLVED IN THIS COMPLAINT MET ALL PRE-RELEASE SPECIFICATIONS. THE DEVICE REMAINS IMPLANTED AND WAS, THEREFORE, NOT AVAILABLE FOR ANALYSIS. NO CLINICAL IMAGES ENABLING DIRECT ASSESSMENT OF PRODUCT PERFORMANCE WERE RETURNED FOR EVALUATION. CAUSE OF THE REPORTED EVENT CANNOT BE ESTABLISHED BASED ON THE INFORMATION REPORTED TO GORE. FURTHER INFORMATION REGARDING THIS EVENT WAS REQUESTED BY GORE, BUT NO FURTHER INFORMATION HAS BEEN REPORTED, THEREFORE THIS INVESTIGATION IS CONSIDERED COMPLETE. Manufacturer Narrative: CBAS® HEPARIN SURFACE INCORPORATES CARMEDA HEPARIN MANUFACTURED FROM HEPARIN SODIUM API, WHICH IS COVALENTLY BOUND TO THE DEVICE SURFACE AND IS ESSENTIALLY NON-ELUTING. H6 - CODE B20: DEVICE REMAINS IMPLANTED, AND NO IMAGES WERE PROVIDED; THEREFORE, DIRECT PRODUCT ANALYSIS WAS NOT POSSIBLE. H6 - CODE C21: RESULTS PENDING COMPLETION OF INVESTIGATION. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21727242&pc=PRL | 2017233-2025-06007 | 2025-02-28 05:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-03-28 04:00 | PRL | GORE® VIABAHN® VBX BALLOON EXPANDABLE ENDOPROSTHESIS | Adverse Event Without Identified Device or Use Problem | Thrombosis/Thrombus | P160021 |  | 1 | Event Description: ON (B)(6) 2025, THE PATIENT UNDERWENT TREATMENT OF AN AORTIC ANEURYSM USING A GORE® EXCLUDER® THORACOABDOMINAL BRANCH ENDOPROSTHESIS (TAMBE) WHERE A GORE® VIABAHN® VBX BALLOON EXPANDABLE ENDOPROSTHESIS (VBX DEVICE) WAS USED IN THE RENAL ARTERY. IT WAS REPORTED THAT THE RIGHT RENAL ARTERY WAS CANNULATED AND STENTED WITH A 5MM X 79MM AND 7MM X 39 MM VBX DEVICES, THEN THE LEFT RENAL ARTERY WITH A 5MM X 79MM AND 7MM X 39 MM VBX DEVICES, THEN THE SUPERIOR MESENTERIC ARTERY (SMA) WITH A 6MM X 59MM AND 9MM X 39MM VBX DEVICES AND LAST THE CELIAC ARTERY WITH A 6MM X 59MM VBX DEVICE. THE PROCEDURE FOR THE TAMBE TOOK APPROXIMATELY THREE HOURS. UPON THE FINAL COMPLETION ANGIOGRAM, IT WAS NOTED THAT THE RENAL ARTERIES DID NOT HAVE FLOW. THE PHYSICIAN CANNULATED THE RENAL ARTERIES AND DISCOVERED BOTH RENAL ARTERIES HAD THROMBUS PRESENT. THE PHYSICIAN PREFORMED A THROMBECTOMY WITH PENUMBRA WHERE SHE REESTABLISHED FLOW. INTRAVASCULAR ULTRASOUND (IVUS) WAS USED TO VISUALIZE THE RENAL ARTERIES TO MAKE SURE THE STENTS WERE NOT CRUSHED. EVERYTHING LOOKED APPROPRIATE VIA IVUS. DURING THIS TIME, THE PHYSICIAN DISCOVERED THAT ANESTHESIA HAD BEEN STRUGGLING KEEPING THE ACTIVATED CLOTTING TIME (ACT) THERAPEUTIC. WITH ACT¿S BEING BELOW 200. THE PATIENT WAS SWITCHED OVER TO ARGATROBAN, AT THAT POINT THE PATIENT¿S ACT WENT TO THERAPEUTIC LEVELS. THE PATIENT¿S LABS WERE SENT IN TO DETERMINE IF THE PATIENT HAD HEPARIN-INDUCED THROMBOCYTOPENIA (HITT), THAT CAME BACK NEGATIVE. AS THE SHEATH WAS COMING OUT OF THE GROIN, THE PHYSICIAN NOTICED NO PULSES IN THE ACCESSED LEG. UPON FURTHER INVESTIGATION IT WAS DETERMINED THE LEFT LEG HAD THROMBOSED AS WELL. THE PHYSICIAN PREFORMED A CUT DOWN ON THE ACCESS AND DID A THROMBECTOMY WHERE FLOW WAS REESTABLISHED. ON (B)(6) 2025, THE FIELD SALES ASSOCIATE WAS CALLED BY THE PHYSICIAN AND INFORMED THAT THE PATIENT WAS EXPERIENCING ABDOMINAL PAIN AND HAD A LACTIC ACID OF OVER 6, PATIENT RECEIVED A CTA AND DISCOVERED ALL 4 VISCERAL VESSELS WERE THROMBOSED. THE PHYSICIAN TOOK THE PATIENT BACK TO THE OPERATING ROOM WHERE A THROMBECTOMY WAS PERFORMED OF THE CELIAC ARTERY AND SMA. FLOW WAS ESTABLISHED HOWEVER BRANCH VESSELS HAD THROMBUS STILL PRESENT. THE PATIENT ULTIMATELY WENT INTO RENAL AND LIVER FAILURE AND PASSED AWAY ON (B)(6) 2025. Manufacturer Narrative: SECTION H6 UPDATED TO REFLECT COMPLETION OF INVESTIGATION. H6 ¿ CODE C19: A REVIEW OF MANUFACTURING RECORDS VERIFIED THAT THE LOT INVOLVED IN THIS COMPLAINT MET ALL PRE-RELEASE SPECIFICATIONS. THE DEVICE REMAINS IMPLANTED AND WAS, THEREFORE, NOT AVAILABLE FOR ANALYSIS. NO CLINICAL IMAGES ENABLING DIRECT ASSESSMENT OF PRODUCT PERFORMANCE WERE RETURNED FOR EVALUATION. CAUSE OF THE REPORTED EVENT CANNOT BE ESTABLISHED BASED ON THE INFORMATION REPORTED TO GORE. FURTHER INFORMATION REGARDING THIS EVENT WAS REQUESTED BY GORE, BUT NO FURTHER INFORMATION HAS BEEN REPORTED, THEREFORE THIS INVESTIGATION IS CONSIDERED COMPLETE. Manufacturer Narrative: CBAS® HEPARIN SURFACE INCORPORATES CARMEDA HEPARIN MANUFACTURED FROM HEPARIN SODIUM API, WHICH IS COVALENTLY BOUND TO THE DEVICE SURFACE AND IS ESSENTIALLY NON-ELUTING. H6 - CODE B20: DEVICE REMAINS IMPLANTED, AND NO IMAGES WERE PROVIDED; THEREFORE, DIRECT PRODUCT ANALYSIS WAS NOT POSSIBLE. H6 - CODE C21: RESULTS PENDING COMPLETION OF INVESTIGATION. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21727252&pc=NPT | 2025587-2025-02212 | 2025-02-11 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-03-28 04:00 | NPT | EVOLUT R TRANSCATHETER AORTIC VALVE | Insufficient Information; Patient Device Interaction Problem | Cardiac Arrest | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 7 YEARS AND 10 MONTHS FOLLOWING THE IMPLANT OF THE TRANSCATHETER BIOPROSTHETIC AORTIC VALVE, THE PATIENT GRABBED THEIR CHEST AND COLLAPSED. THE DEATH WAS REPORTED AS CARDIAC RELATED, CARDIAC ARREST. AN AUTOPSY WAS NOT PERFORMED. THE PHYSICIAN INDICATED THE DEATH WAS UNLIKELY RELATED THE VALVE. Manufacturer Narrative: ESTIMATED: A2, A4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: A4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.