## Sheet1
Web Address | Report Number | Event Date | Event Type | Manufacturer | Date Received | Product Code | Brand Name | Device Problem | Patient Problem | PMA/PMN Number | Exemption Number | Number of Events | Event Text
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22824755&pc=DWF | 8010762-2025-0000357 | 2025-07-10 04:00 | Death | MAQUET CARDIOPULMONARY GMBH | 2025-08-18 04:00 | DWF | HLS CANNULA + PIK | Air/Gas in Device | Insufficient Information | K102532 |  | 1 | Event Description: IT WAS REPORTED THAT DURING PREPARATION FOR A PLANNED CT SCAN ¿ PLANNED IN-HOUSE PATIENT TRANSPORT WITH VA ECMO FROM THE INTENSIVE CARE UNIT TO THE CT ¿ THE ROTAFLOW DRIVE UNIT WAS POSITIONED IN THE PATIENT BED AS PER INSTRUCTIONS. THE DRIVE PUMP WAS RUNNING CORRECTLY AT A FLOW RATE OF 2.6 L/MIN. SUDDENLY, THE ROTAFLOW FLOW INDICATOR ALERTED, AND THE FLOW INDICATOR WAS NO LONGER READABLE. AT THE SAME TIME, A SIGNIFICANTLY LOUDER ROTATION NOISE FROM THE DRIVE PUMP WAS AUDIBLE. AN IMMEDIATE, COMPLETE INSPECTION REVEALED THAT AIR HAD ENTERED THE SYSTEM. THE ARTERIAL LINE WAS IMMEDIATELY CLAMPED, AND THE ECMO SPEED WAS COMPLETELY REDUCED. AFTER IMMEDIATELY CLAMPING THE ARTERIAL LINE AND REDUCING THE ECMO SPEED, THE PERFUSIONISTS WERE IMMEDIATELY INFORMED, WITH THE URGENT INSTRUCTION TO IMMEDIATELY TRANSFER A FILLED, READY-TO-USE SYSTEM TO THE INTENSIVE CARE UNIT. THE SUBSEQUENT INSPECTION OF THE ECMO SYSTEM TOGETHER WITH THE ATTENDING PHYSICIANS REVEALED NO VISIBLE DAMAGE TO THE ECMO. ANALYSIS: DURING THE UKR'S ANALYSIS OF THE AFFECTED PRODUCTS, NO UNDESIRABLE SIDE EFFECTS, NO MALFUNCTION, NO APPLICATION ERROR, AND NO INADEQUACY OF THE INFORMATION PROVIDED WERE IDENTIFIED. CUSTOMER MENTIONED THAT "ALTHOUGH WE HAVE EXHAUSTED OUR ANALYTICAL CAPABILITIES, WE CANNOT COMPLETELY RULE OUT THE ABOVE-MENTIONED ERROR POSSIBILITIES, AND THEREFORE WE ARE REPORTING THE FATAL PATIENT INJURY RESULTING IN DEATH". DUE TO THE POLICE INVESTIGATION, THE DISPOSABLE PRODUCTS ARE NO LONGER WITHIN OUR CONTROL. THE LOT INFORMATION FOR THE INVOLVED DISPOSABLE PRODUCTS ARE THEREFORE NOT AVAILABLE. THE OTHER HARDWARE IS SECURELY STORED AT THE UKR. AS MULTIPLE GETINGE PRODUCTS WERE USED DURING THE TREATMENT, THIS COMPLAINT WILL FOCUS ON INVESTIGATING THE SPECIFIC PRODUCT: BE-PVL 2155, MATERIAL NUMBER: 701047294. THE INVESTIGATION OF OTHER RELEVANT PRODUCTS WILL BE COVERED UNDER THE FOLLOWING COMPLAINT NUMBERS: (B)(4), MATERIAL NUMBER: 701047280: ONETRACK#:1347733, (B)(4), MATERIAL NUMBER: 701068386: ONETRACK#:1347777, ROTAFLOW: ONETRACK#:1346851. DUE TO THE REPORTED DEATH A REPORT IS REQUIRED. COMPLAINT#: (B)(4). Manufacturer Narrative: A FOLLOW UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22824756&pc=DWF | 8010762-2025-0000358 | 2025-07-10 04:00 | Death | MAQUET CARDIOPULMONARY GMBH | 2025-08-18 04:00 | DWF | HLS CANNULA + PIK | Air/Gas in Device | Insufficient Information | K102532 |  | 1 | Event Description: IT WAS REPORTED THAT DURING PREPARATION FOR A PLANNED CT SCAN ¿ PLANNED IN-HOUSE PATIENT TRANSPORT WITH VA ECMO FROM THE INTENSIVE CARE UNIT TO THE CT ¿ THE ROTAFLOW DRIVE UNIT WAS POSITIONED IN THE PATIENT BED AS PER INSTRUCTIONS. THE DRIVE PUMP WAS RUNNING CORRECTLY AT A FLOW RATE OF 2.6 L/MIN. SUDDENLY, THE ROTAFLOW FLOW INDICATOR ALERTED, AND THE FLOW INDICATOR WAS NO LONGER READABLE. AT THE SAME TIME, A SIGNIFICANTLY LOUDER ROTATION NOISE FROM THE DRIVE PUMP WAS AUDIBLE. AN IMMEDIATE, COMPLETE INSPECTION REVEALED THAT AIR HAD ENTERED THE SYSTEM. THE ARTERIAL LINE WAS IMMEDIATELY CLAMPED, AND THE ECMO SPEED WAS COMPLETELY REDUCED. AFTER IMMEDIATELY CLAMPING THE ARTERIAL LINE AND REDUCING THE ECMO SPEED, THE PERFUSIONISTS WERE IMMEDIATELY INFORMED, WITH THE URGENT INSTRUCTION TO IMMEDIATELY TRANSFER A FILLED, READY-TO-USE SYSTEM TO THE INTENSIVE CARE UNIT. THE SUBSEQUENT INSPECTION OF THE ECMO SYSTEM TOGETHER WITH THE ATTENDING PHYSICIANS REVEALED NO VISIBLE DAMAGE TO THE ECMO. ANALYSIS: DURING THE UKR'S ANALYSIS OF THE AFFECTED PRODUCTS, NO UNDESIRABLE SIDE EFFECTS, NO MALFUNCTION, NO APPLICATION ERROR, AND NO INADEQUACY OF THE INFORMATION PROVIDED WERE IDENTIFIED. CUSTOMER MENTIONED THAT "ALTHOUGH WE HAVE EXHAUSTED OUR ANALYTICAL CAPABILITIES, WE CANNOT COMPLETELY RULE OUT THE ABOVE-MENTIONED ERROR POSSIBILITIES, AND THEREFORE WE ARE REPORTING THE FATAL PATIENT INJURY RESULTING IN DEATH". DUE TO THE POLICE INVESTIGATION, THE DISPOSABLE PRODUCTS ARE NO LONGER WITHIN OUR CONTROL. THE LOT INFORMATION FOR THE INVOLVED DISPOSABLE PRODUCTS ARE THEREFORE NOT AVAILABLE. THE OTHER HARDWARE IS SECURELY STORED AT THE UKR. AS MULTIPLE GETINGE PRODUCTS WERE USED DURING THE TREATMENT, THIS COMPLAINT WILL FOCUS ON INVESTIGATING THE SPECIFIC PRODUCT BE-PAS 1515, MATERIAL NUMBER 701047280. THE INVESTIGATION OF OTHER RELEVANT PRODUCTS WILL BE COVERED UNDER THE FOLLOWING COMPLAINT NUMBERS: BE-PVL 2155, MATERIAL NUMBER 701047294: ONETRACK #(B)(4), BE-PLS 2050, MATERIAL NUMBER 701068386: ONETRACK #(B)(4), ROTAFLOW: ONETRACK #(B)(4). DUE TO THE REPORTED DEATH A REPORT IS REQUIRED. COMPLAINT #(B)(4). Manufacturer Narrative: A FOLLOW UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22825292&pc=NPT | 2015691-2025-06736 | 2025-07-30 04:00 | Death | EDWARDS LIFESCIENCES | 2025-08-18 04:00 | NPT | COMMANDER DELIVERY SYSTEM | Failure to Align | Perforation of Vessels | P140031 |  | 1 | Event Description: AS REPORTED BY THE FIELD CLINICAL SPECIALIST (FCS), DURING A RIGHT TRANSFEMORAL TAVR PROCEDURE WITH A 29MM SAPIEN 3 ULTRA RESILIA VALVE, SIGNIFICANT RESISTANCE AND DIFFICULTY WITH VALVE ALIGNMENT WAS NOTED DURING FINE ADJUSTMENT. THE TEAM ATTEMPTED TO DEPLOY THE VALVE, AND THE BALLOON "WATERMELON-SEEDED" VENTRICULAR, AND THE VALVE EMBOLIZED AORTIC. THE TEAM PULLED THE DELIVERY SYSTEM AND VALVE INTO THE DESCENDING ABDOMINAL AORTA TO ATTEMPT TO DEPLOY, BUT COULD NOT GET THE BALLOON INSIDE THE VALVE, AND SUBSEQUENTLY PERFORATED THE AORTA BELOW THE RENAL ARTERIES. A CODA BALLOON WAS INSERTED AND RESUSCITATION WAS ATTEMPTED, BUT THE PATIENT PASSED AWAY. PER REPORT, THERE WAS SIGNIFICANT TORTUOSITY IN THORACIC AORTA, VALVE ALIGNMENT WAS PERFORMED IN A STRAIGHT SECTION OF THE ANATOMY, BUT THE VALVE WAS NOT BETWEEN THE MARKERS BEFORE VALVE DEPLOYMENT (UNDER-ALIGNED). Manufacturer Narrative: INVESTIGATION IS STILL ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22825757&pc=PNJ | 9612164-2025-04032 | 2025-07-23 04:00 | Death | MEDTRONIC IRELAND | 2025-08-18 04:00 | PNJ | MICRA¿ VR2 | Appropriate Term/Code Not Available | Insufficient Information | P150033 |  | 1 | Event Description: IT WAS REPORTED BY THE PATIENT REPRESENTATIVE THAT THE PATIENT PASSED AWAY DUE TO FAILURE OF THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG). Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22826424&pc=DWF | 3013876692-2025-00070 | 2025-07-10 04:00 | Death | MAQUET CARDIOPULMONARY GMBH | 2025-08-18 04:00 | DWF | HLS CANNULA + PIK | Air/Gas in Device | Insufficient Information | K102532 |  | 1 | Event Description: IT WAS REPORTED THAT DURING PREPARATION FOR A PLANNED CT SCAN ¿ PLANNED IN-HOUSE PATIENT TRANSPORT WITH VA ECMO FROM THE INTENSIVE CARE UNIT TO THE CT ¿ THE ROTAFLOW DRIVE UNIT WAS POSITIONED IN THE PATIENT BED AS PER INSTRUCTIONS. THE DRIVE PUMP WAS RUNNING CORRECTLY AT A FLOW RATE OF 2.6 L/MIN. SUDDENLY, THE ROTAFLOW FLOW INDICATOR ALERTED, AND THE FLOW INDICATOR WAS NO LONGER READABLE. AT THE SAME TIME, A SIGNIFICANTLY LOUDER ROTATION NOISE FROM THE DRIVE PUMP WAS AUDIBLE. AN IMMEDIATE, COMPLETE INSPECTION REVEALED THAT AIR HAD ENTERED THE SYSTEM. THE ARTERIAL LINE WAS IMMEDIATELY CLAMPED, AND THE ECMO SPEED WAS COMPLETELY REDUCED. AFTER IMMEDIATELY CLAMPING THE ARTERIAL LINE AND REDUCING THE ECMO SPEED, THE PERFUSIONISTS WERE IMMEDIATELY INFORMED, WITH THE URGENT INSTRUCTION TO IMMEDIATELY TRANSFER A FILLED, READY-TO-USE SYSTEM TO THE INTENSIVE CARE UNIT. THE SUBSEQUENT INSPECTION OF THE ECMO SYSTEM TOGETHER WITH THE ATTENDING PHYSICIANS REVEALED NO VISIBLE DAMAGE TO THE ECMO. ANALYSIS: DURING THE UKR'S ANALYSIS OF THE AFFECTED PRODUCTS, NO UNDESIRABLE SIDE EFFECTS, NO MALFUNCTION, NO APPLICATION ERROR, AND NO INADEQUACY OF THE INFORMATION PROVIDED WERE IDENTIFIED. CUSTOMER MENTIONED THAT "ALTHOUGH WE HAVE EXHAUSTED OUR ANALYTICAL CAPABILITIES, WE CANNOT COMPLETELY RULE OUT THE ABOVE-MENTIONED ERROR POSSIBILITIES, AND THEREFORE WE ARE REPORTING THE FATAL PATIENT INJURY RESULTING IN DEATH". DUE TO THE POLICE INVESTIGATION, THE DISPOSABLE PRODUCTS ARE NO LONGER WITHIN OUR CONTROL. THE LOT INFORMATION FOR THE INVOLVED DISPOSABLE PRODUCTS ARE THEREFORE NOT AVAILABLE. THE OTHER HARDWARE IS SECURELY STORED AT THE UKR. AS MULTIPLE GETINGE PRODUCTS WERE USED DURING THE TREATMENT, THIS COMPLAINT WILL FOCUS ON INVESTIGATING THE SPECIFIC PRODUCT BE-PVL 2155, MATERIAL NUMBER 701047294. THE INVESTIGATION OF OTHER RELEVANT PRODUCTS WILL BE COVERED UNDER THE FOLLOWING COMPLAINT NUMBERS: BE-PAS 1515, MATERIAL NUMBER 701047280: ONETRACK #1347733, BE-PLS 2050, MATERIAL NUMBER 701068386: ONETRACK #1347777, ROTAFLOW: ONETRACK #1346851. DUE TO THE REPORTED DEATH A REPORT IS REQUIRED. COMPLAINT #(B)(4).
Manufacturer Narrative: A FOLLOW UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22826557&pc=DWF | 3013876692-2025-00071 | 2025-07-10 04:00 | Death | MAQUET CARDIOPULMONARY GMBH | 2025-08-18 04:00 | DWF | HLS CANNULA + PIK | Air/Gas in Device | Insufficient Information | K102532 |  | 1 | Event Description: IT WAS REPORTED THAT DURING PREPARATION FOR A PLANNED CT SCAN ¿ PLANNED IN-HOUSE PATIENT TRANSPORT WITH VA ECMO FROM THE INTENSIVE CARE UNIT TO THE CT ¿ THE ROTAFLOW DRIVE UNIT WAS POSITIONED IN THE PATIENT BED AS PER INSTRUCTIONS. THE DRIVE PUMP WAS RUNNING CORRECTLY AT A FLOW RATE OF 2.6 L/MIN. SUDDENLY, THE ROTAFLOW FLOW INDICATOR ALERTED, AND THE FLOW INDICATOR WAS NO LONGER READABLE. AT THE SAME TIME, A SIGNIFICANTLY LOUDER ROTATION NOISE FROM THE DRIVE PUMP WAS AUDIBLE. AN IMMEDIATE, COMPLETE INSPECTION REVEALED THAT AIR HAD ENTERED THE SYSTEM. THE ARTERIAL LINE WAS IMMEDIATELY CLAMPED, AND THE ECMO SPEED WAS COMPLETELY REDUCED. AFTER IMMEDIATELY CLAMPING THE ARTERIAL LINE AND REDUCING THE ECMO SPEED, THE PERFUSIONISTS WERE IMMEDIATELY INFORMED, WITH THE URGENT INSTRUCTION TO IMMEDIATELY TRANSFER A FILLED, READY-TO-USE SYSTEM TO THE INTENSIVE CARE UNIT. THE SUBSEQUENT INSPECTION OF THE ECMO SYSTEM TOGETHER WITH THE ATTENDING PHYSICIANS REVEALED NO VISIBLE DAMAGE TO THE ECMO. ANALYSIS: DURING THE UKR'S ANALYSIS OF THE AFFECTED PRODUCTS, NO UNDESIRABLE SIDE EFFECTS, NO MALFUNCTION, NO APPLICATION ERROR, AND NO INADEQUACY OF THE INFORMATION PROVIDED WERE IDENTIFIED. CUSTOMER MENTIONED THAT "ALTHOUGH WE HAVE EXHAUSTED OUR ANALYTICAL CAPABILITIES, WE CANNOT COMPLETELY RULE OUT THE ABOVE-MENTIONED ERROR POSSIBILITIES, AND THEREFORE WE ARE REPORTING THE FATAL PATIENT INJURY RESULTING IN DEATH". DUE TO THE POLICE INVESTIGATION, THE DISPOSABLE PRODUCTS ARE NO LONGER WITHIN OUR CONTROL. THE LOT INFORMATION FOR THE INVOLVED DISPOSABLE PRODUCTS ARE THEREFORE NOT AVAILABLE. THE OTHER HARDWARE IS SECURELY STORED AT THE UKR. AS MULTIPLE GETINGE PRODUCTS WERE USED DURING THE TREATMENT, THIS COMPLAINT WILL FOCUS ON INVESTIGATING THE SPECIFIC PRODUCT BE-PAS 1515, MATERIAL NUMBER 701047280. THE INVESTIGATION OF OTHER RELEVANT PRODUCTS WILL BE COVERED UNDER THE FOLLOWING COMPLAINT NUMBERS: BE-PVL 2155, MATERIAL NUMBER 701047294: ONETRACK #(B)(4), BE-PLS 2050, MATERIAL NUMBER 701068386: ONETRACK #(B)(4), ROTAFLOW: ONETRACK #(B)(4). DUE TO THE REPORTED DEATH A REPORT IS REQUIRED. COMPLAINT #(B)(4). Manufacturer Narrative: A FOLLOW UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22826923&pc=OZD | 1220648-2025-45751 | 2025-07-07 04:00 | Death | ABIOMED, INC. | 2025-08-18 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Respiratory Failure | P140003 |  | 1 | Event Description: A NOTIFICATION VIA ABIOMED¿S LONG-TERM OUTCOME AND QUALITY INDICATOR (LOQI) IMPELLA REGISTRY WAS RECEIVED REGARDING A RESPIRATORY FAILURE WITH RELATIONSHIP LISTED AS "POSSIBLY" RELATED. THE LOQI STATED: "INTUBATED AFTER IMPELLA PROCEDURE. AFTER 4 DAYS OF SUPPORT, FAMILY DECIDED TO WITHDREW CARE AND IMPELLA PULLED OUT OF THE LEFT VENTRICLE AND TURNED OFF. PATIENT TAKEN TO HOSPICE UNIT TO PASS COMFORTABLY. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22828025&pc=QSK | 9612515-2025-00071 | 2024-10-17 04:00 | Death | VASCUTEK LTD | 2025-08-18 04:00 | QSK | THORAFLEX HYBRID | Insufficient Information | Insufficient Information | P210006 |  | 1 | Event Description: THE PATIENT WAS ADMITTED ON (B)(6) 2024 AND DIED ON (B)(6) 2024. CAUSE OF DEATH IS CURRENTLY UNKNOWN. Manufacturer Narrative: CLINICAL CODE: 4580 - INSUFFICIENT INFORMATION: NO INFORMATION PROVIDED REGARDING CLINICAL SIGNS/ SYMPTOMS. IMPACT CODE 1802- DEATH: THE PATIENT HAD DIED ON (B)(6) 2024. MEDICAL DEVICE PROBLEM 3190 - INSUFFICIENT INFORMATION: AWAITING INFORMATION FROM THE SITE. COMPONENT CODE 4755 - PART/COMPONENT/SUB-ASSEMBLY TERM NOT APPLICABLE. TYPE OF INVESTIGATION 3331 - ANALYSIS OF PRODUCTION RECORDS: A REVIEW OF THE MANUFACTURING AND QUALITY CONTROL RECORDS FOR THIS BATCH CONFIRMED THAT THE PRODUCT WAS MANUFACTURED TO SPECIFICATION. 4110 - TREND ANALYSIS: A 4 YEAR REVIEW WAS PERFORMED FOR THORAFLEX HYBRID > MEDICAL EVENT > DEATH THIS GAVE AN OCCURRENCE RATE OF LESS THAN (B)(4). NO TREND IDENTIFIED THAT REQUIRED ACTION AT THIS TIME 4111 - COMMUNICATION/INTERVIEWS: ADDITIONAL INFORMATION HAS BEEN REQUESTED, WAITING ON INFORMATION FROM THE SITE. 4117 - DEVICE NO ACCESSIBLE FOR TESTING: THE DEVICE REMAINS IMPLANTED. INVESTIGATION FINDINGS 3233 - RESULTS PENDING COMPLETION OF INVESTIGATION. INVESTIGATION CONCLUSION 11 - CONCLUSION NOT YET AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22828237&pc=NPT | 9617601-2025-01028 | 2025-01-15 05:00 | Death | MEDTRONIC MEXICO S. DE R.L. DE CV | 2025-08-18 04:00 | NPT | EVOLUT FX VALVE | Patient Device Interaction Problem | Bacterial Infection; Stroke/CVA; Pulmonary Edema; Dyskinesia; Respiratory Tract Infection; Speech Disorder | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT 15 DAYS FOLLOWING THE IMPLANT OF A TRANSCATHETER VALVE, THE PATIENT DEVELOPED DYSKINESIA, DYSARTHRIA, AND DROOLING. SUBSEQUENTLY, THE PATIENT WAS HOSPITALIZED DUE TO A CEREBRAL INFARCTION OF UNKNOWN EMBOLIC SOURCE, PULMONARY EDEMA, POSITIVE FOR INFLUENZA, AND SUSPECTED BACTERIAL INFECTION. DESPITE TREATMENT DURING HOSPITALIZATION, THE AREAS OF INFARCTION INCREASED, THE PATIENT'S LEVELS OF CONSCIOUSNESS DECLINED, BLOOD PRESSURE WAS DIFFICULT TO MAINTAIN, AND THE PULMONARY EDEMA DID NOT IMPROVE. TWENTY DAYS FOLLOWING THE IMPLANT OF THE VALVE, THE PATIENT DIED. THE CAUSE OF DEATH WAS DUE TO WORSENED GENERAL CONDITION ASSOCIATED WITH INFLUENZA INFECTION. PER THE PHYSICIAN, THERE WAS NO RELATIONSHIP BETWEEN THE DEATH AND DEVICE, OR IMPLANT PROCEDURE. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22828595&pc=DYE | 2015691-2025-06750 | 2025-07-21 04:00 | Death | EDWARDS LIFESCIENCES | 2025-08-18 04:00 | DYE | CARPENTIER-EDWARDS PERIMOUNT MAGNA MITRAL EASE PERICARDIAL BIOPROSTHESIS | Patient Device Interaction Problem | Thrombosis/Thrombus | P860057 |  | 1 | Event Description: EDWARDS RECEIVED NOTIFICATION FROM THAILAND THAT A PATIENT WITH A SIZE VALVE MODEL 7300TFX27 IMPLANTED IN THE MITRAL POSITION WAS NOTED TO PRESENT THROMBOSIS 14 DAYS AFTER IMPLANTATION. DURING THE IMPLANT SURGERY, AFTER PROTAMINE ADMINISTRATION PATIENT'S HEMODYNAMICS RAPIDLY DETERIORATED. ADDITIONALLY TEE SHOWED NON CONTRACTILE RIGHT VENTRICULAR FAILURE. THE PATIENT REQUIRED ESCALATION OF SUPPORT WITH INTRA-AORTIC BALLOON PUMP (IABP), INOTROPIC AGENTS AND INITIATING ECMO. REPORTEDLY, ONE WEEK AFTER SURGERY, THE VALVE LEAFLETS WERE NOT FULLY OPENING AND THE GRADIENTS HAD INCREASED OVER 10MMHG. ON POST-OPERATIVE DAY 14, VALVE THROMBOSIS WAS DIAGNOSED, ALONG WITH RESTRICTED LEAFLET MOTION AND COMMISSURAL FUSION. ADDITIONAL FINDINGS INCLUDED TWO LEFT ATRIAL THROMBI, A LARGE ANTERIOR MEDIASTINAL HEMATOMA WITH ACTIVE CONTRAST EXTRAVASATION COMPRESSING THE SUPERIOR VENA CAVA (SVC) AND RIGHT ATRIUM (RA), CAUSING MODERATE LUMINAL NARROWING OF THE SVC. THE PATIENT ALSO PRESENTED WITH SEPTIC EMBOLI IN BOTH LUNGS, PULMONARY EDEMA, BILATERAL PLEURAL EFFUSION, AND ASPIRATION IN THE BASAL REGIONS OF BOTH LUNGS. THE PATIENT ENDURED PROLONGED HOSPITALIZATION WITH MULTIPLE SYSTEMIC COMPLICATIONS, ULTIMATELY THE PATIENT WAS DISCHARGED TO HOME PALLIATION AND PASSED AWAY. Manufacturer Narrative: H11. ADDITIONAL MANUFACTURER NARRATIVE: THE SUBJECT DEVICE IS NOT AVAILABLE FOR EVALUATION AS IT REMAINS IMPLANTED IN THE PATIENT. THE INVESTIGATION IS STILL IN PROGRESS; THEREFORE, A CONCLUSION HAS YET TO BE ESTABLISHED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY UPON INVESTIGATION COMPLETION. EDWARDS WILL CONTINUE TO REVIEW AND MONITOR ALL REPORTED EVENTS. TRENDS ARE MONITORED ON A MONTHLY BASIS AND IF ACTION IS REQUIRED, APPROPRIATE INVESTIGATION WILL BE PERFORMED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22828617&pc=NPT | 2015691-2025-06751 | 2025-07-10 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-08-18 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Endocarditis | P140031 |  | 1 | Event Description: AS REPORTED THROUGH THE JAPANESE TAVI REGISTRY, THE PATIENT UNDERWENT A TRANSFEMORAL TAVI AND RECEIVED A 26 MM SAPIEN 3 ULTRA RESILIA (S3UR) VALVE IN THE NATIVE AORTIC POSITION. INCREASED PRESSURE GRADIENT AND DECREASED LEAFLET MOTION WERE OBSERVED; THE EXACT OCCURRENCE DATE WAS UNKNOWN (NOT REPORTED). TWO MONTHS AFTER THE TAVI, THE PATIENT PASSED AWAY DUE TO UNCONTROLLED SEPSIS. AUTOPSY FINDINGS SHOWED SUSPICION OF VEGETATIONS ON THE S3UR VALVE. THE SOURCE OF THE SEPSIS AND THE EVENT OCCURRENCE DATE WERE UNKNOWN (NOT REPORTED). PER REPORTED MEDICAL OPINION, THE DEATH WAS RELATED TO THE VALVE. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), INFECTIONS SUCH AS SEPTICEMIA AND ENDOCARDITIS ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. ENDOCARDITIS IS AN INFECTION OF A NATIVE OR PROSTHETIC VALVE, IS TREATED WITH ANTIBIOTICS, AND MAY REQUIRE VALVE REPLACEMENT IF ANTIBIOTIC THERAPY IS NOT EFFECTIVE. CAUSES OF PROSTHETIC VALVE ENDOCARDITIS ARE WELL DOCUMENTED IN THE LITERATURE AND ARE TYPICALLY CLASSIFIED AS EARLY (<60 DAYS) OR LATE (>60 DAYS). THE TWO CATEGORIES SHOW DEFINITE DIFFERENCES IN CLINICAL FEATURES, MICROBIAL PATTERNS, AND MORTALITY. ACCORDING TO THE LITERATURE REVIEW, AND AS DOCUMENTED IN A TECHNICAL SUMMARY WRITTEN BY EW, EARLY PROSTHETIC VALVE ENDOCARDITIS IS USUALLY CAUSED BY PERIOPERATIVE BACTERIAL CONTAMINATION OF THE VALVE. IN EARLY CASES OF PROSTHETIC ENDOCARDITIS, SUBSEQUENT INFECTIONS MAY BE RELATED TO CONTAMINATION AT THE TIME OF SURGERY (POOR STERILE TECHNIQUE). ADDITIONAL CAUSES OF EARLY PVE CAN OCCUR FROM OTHER SOURCES OF INFECTION AND NOT FROM THE VALVE (E.G., DENTAL TREATMENT OR SURGICAL PROCEDURES INVOLVING THE UPPER RESPIRATORY TRACT, UTIS, PNEUMONIA). ADDITIONALLY, A NON-CONFORMANCE IN THE STERILITY OR PACKAGING PROCESSES OF THE VALVE MAY ALSO MANIFEST IN THE EARLY POST-OPERATIVE PERIOD CAUSING EARLY PVE. HOWEVER, THERE ARE SEVERAL TYPES OF BACTERIA THAT WOULD NOT SURVIVE THE EDWARDS STERILIZATION PROCESS. EDWARDS LIFESCIENCES PRODUCES AND PROVIDES STERILE TISSUE BIOPROSTHESIS TO ITS CUSTOMERS BY FOLLOWING CAREFULLY DESIGNED ROBUST STERILIZATION PROCESSES. THESE MANUFACTURING PROCESSES HAVE BEEN VALIDATED AND DEMONSTRATED TO CONSISTENTLY PROVIDE A SIGNIFICANT SAFETY FACTOR FROM WHICH MICROORGANISMS COULD NOT SURVIVE. MICROBIOLOGY AND PROCESS MONITORING IS ROUTINELY REVIEWED WITHIN QUALITY SYSTEMS TO MAINTAIN STERILITY CONTROL. VALIDATED TESTING HAS DEMONSTRATED THAT MICROORGANISMS COULD NOT SURVIVE EDWARDS' MULTI-STAGE PROCESSING WITH ENHANCED STERILANT OR HEATED GLUTARALDEHYDE TERMINAL STERILANT SOLUTION. THESE MULTIPLE, REDUNDANT MANUFACTURING CONTROLS ENSURE THE STERILITY OF EDWARDS' VALVES AS PROVIDED TO CUSTOMERS. THEREFORE, THE PROBABILITY OF ENDOCARDITIS RELATED TO EDWARDS' BIOPROSTHESIS IS REMOTE. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. WITH THE LIMITED INFORMATION PROVIDED, THE CAUSE OF THE ENDOCARDITIS IS UNKNOWN BUT MAY BE RELATED TO THE MECHANISMS DESCRIBED ABOVE. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22828890&pc=PNJ | 2017865-2025-97928 | 2025-07-01 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-08-18 04:00 | PNJ | AVEIR¿ LEADLESS PACEMAKER, RIGHT VENTRICLE | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P150035 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT RECEIVED AN AVEIR LEADLESS PACEMAKER (LP). THE PATIENT WAS FOUND TO BE DECEASED POST-PROCEDURE DUE TO UNKNOWN CAUSES. THERE WERE NO DIRECT ALLEGATIONS THAT THE PATIENT'S DEATH WAS CAUSED OR CONTRIBUTED TO BY THEIR IMPLANTED SYSTEM.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22829051&pc=DSQ | 2916596-2025-05202 | 2021-08-10 04:00 | Death | THORATEC CORPORATION | 2025-08-18 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTIORGAN FAILURE. THE OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. THE PUMP WAS WORKING AS EXPECTED AND NO PUMP RELATED ISSUES WERE PROVIDED. AN AUTOPSY WAS NOT PREFORMED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT'S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (INCLUDING RIGHT HEART, RESPIRATORY, RENAL, AND HEPATIC FAILURE) AND DEATH AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22829064&pc=NGV | 2135147-2025-04604 | 2021-08-14 04:00 | Death | ABBOTT MEDICAL | 2025-08-18 04:00 | NGV | AMPLATZER AMULET | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P200049 |  | 1 | Event Description: THE ARTICLE, "OUTCOMES FOR SINGLE ANTIPLATELET, DUAL ANTIPLATELET, OR ORAL ANTICOAGULATION AFTER AMULET: INSIGHTS FROM EMERGE LAA POST-APPROVAL STUDY", WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE, MULTICENTER STUDY ON OUTCOMES FROM PATIENTS DISCHARGED ON DUAL ANTIPLATELET THERAPY (DAPT) VERSUS SINGLE ANTIPLATELET THERAPY (SAPT) OR ORAL ANTICOAGULATION (OAC) FOLLOWING AMULET IMPLANT IN THE EMERGE LAA POST -APPROVAL STUDY. DEVICES INCLUDED IN THE STUDY WERE SOLELY AMPLATZER AMULET. THE ARTICLE CONCLUDED THAT IN THE PRESENT COHORT OF PATIENTS TREATED WITH AMULET LEFT ATRIAL APPENDAGE OCCLUSION (LAAO) THERE WERE HIGH RATES OF LAA COMPLETE CLOSURE, AND SAPT EMERGED AS A VIABLE ALTERNATIVE TO THE CURRENT DAPT REGIMEN POTENTIALLY REDUCING THE RISK OF BLEEDING COMPLICATIONS IN PATIENTS AT HIGH RISK WITHOUT COMPROMISING EFFECTIVENESS. [THE PRIMARY AND CORRESPONDING AUTHOR WAS ATMAN SHAH, THE UNIVERSITY OF CHICAGO, GOLDBLATT PAVILION, 5841 S MARYLAND AVE, CHICAGO, IL 60637, WITH CORRESPONDING EMAIL: ASHAH@BSD.UCHICAGO.EDU] THE TIME FRAME OF THE STUDY WAS 14 AUGUST 2021 TO 15 DECEMBER 2023. A TOTAL OF 11,445 PATIENTS WERE INCLUDED IN THE STUDY, OF WHICH ALL RECEIVED AN ABBOTT DEVICE. THE AVERAGE AGE WAS 77.0 YEARS AND THE MAJORITY GENDER WAS MALE. COMORBIDITIES INCLUDED ATRIAL FIBRILLATION, PRIOR STROKE, DIABETES MELLITUS, CONGESTIVE HEART FAILURE, HYPERTENSION, VASCULAR DISEASE, PRIOR BLEEDING, INTRACRANIAL BLEEDING, CORONARY ARTERY DISEASE, PRIOR LEFT ATRIAL APPENDAGE OCCLUSION INTERVENTION, AND PRIOR CARDIAC STRUCTURAL INTERVENTION. Manufacturer Narrative: IT IS UNKNOWN IF THE DEVICE IS RETURNING FOR ANALYSIS. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION B2: DATE OF DEATH WAS ESTIMATED D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: OUTCOMES FOR SINGLE ANTIPLATELET, DUAL ANTIPLATELET, OR ORAL ANTICOAGULATION AFTER AMULET: INSIGHTS FROM EMERGE LAA POST-APPROVAL STUDY
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22829195&pc=DSQ | 2916596-2025-05036 | 2025-07-19 04:00 | Death | THORATEC CORPORATION | 2025-08-18 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Disconnection; Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS FOUND DECEASED AT HOME BY FAMILY. THE PATIENT WAS CONNECTED TO THE MOBILE POWER UNIT (MPU) WHEN FOUND. ON ALARM REVIEW THE DRIVELINE WAS DISCONNECTED AT 1820 BUT NO FURTHER ALARMS UNTIL 0538 THE NEXT DAY. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22832287&pc=NPT | 2015691-2025-06773 | 2025-07-30 04:00 | Death | EDWARDS LIFESCIENCES | 2025-08-18 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem |  | P140031 |  | 1 | Event Description: AS REPORTED BY OUR UNITED KINGDOM AFFILIATES, DURING IMPLANT OF A 26MM SAPIEN 3 ULTRA RESILIA (S3UR) VALVE IN THE AORTIC POSITION, THE PATIENT EXPERIENCED SIGNIFICANT PROCEDURAL CHALLENGES AND COMPLICATIONS. DUE TO A DIFFICULT AORTIC VALVE CROSSING, MULTIPLE CATHETER AND WIRE EXCHANGES WERE REQUIRED BEFORE SUCCESSFUL PASSAGE WAS ACHIEVED. A TEMPORARY PACING WIRE WAS INSERTED DUE TO THE PATIENT'S HISTORY OF FIRST-DEGREE AV BLOCK. UPON CROSSING THE AORTIC VALVE WITH THE S3UR SYSTEM, THE PATIENT BECAME PROFOUNDLY HYPOTENSIVE (50/30 MMHG), PROMPTING THE TEAM TO PROCEED RAPIDLY WITH VALVE DEPLOYMENT, WHICH WAS COMPLETED AT NOMINAL VOLUME. DESPITE DEPLOYMENT, THE PATIENT'S BLOOD PRESSURE REMAINED CRITICALLY LOW (40/20 MMHG), REQUIRING ANESTHETIC PRESSURE SUPPORT, ALTHOUGH THE PATIENT WAS STILL RESPONSIVE. AN AORTOGRAM SHOWED GOOD VALVE POSITIONING WITH NO PARAVALVULAR LEAK OR AORTIC REGURGITATION AND ADEQUATE FILLING OF THE LEFT MAIN CORONARY ARTERY. A TRANSTHORACIC ECHOCARDIOGRAM RULED OUT PERICARDIAL EFFUSION, BUT THE PATIENT HAD NO CARDIAC OUTPUT, LEADING TO THE INITIATION OF CPR. A CORONARY ANGIOGRAM REVEALED AN OCCLUSION IN THE MID-LAD, LIKELY DUE TO EMBOLIZED CALCIUM, AND EMERGENCY CORONARY ANGIOPLASTY WITH STENTING WAS PERFORMED WHILE CONTINUOUS COMPRESSIONS WERE MAINTAINED USING A LUCAS DEVICE. ALTHOUGH FLOW TO THE LAD WAS RESTORED, THE LCX REMAINED COMPROMISED AND CARDIAC OUTPUT WAS NOT RE-ESTABLISHED. DESPITE ALL RESUSCITATIVE AND INTERVENTIONAL EFFORTS, THE TEAM ULTIMATELY MADE THE DECISION TO STOP AND THE PATIENT EXPIRED. Manufacturer Narrative: THE INVESTIGATION IS ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22832497&pc=DRE | 2124215-2025-57545 | 2025-07-29 04:00 | Death | BAYLIS MEDICALE CIE INC | 2025-08-18 04:00 | DRE | VERSACROSS LARGE ACCESS SOLUTION | Display or Visual Feedback Problem; Difficult to Advance | Low Blood Pressure/ Hypotension; Pericardial Effusion | K201288 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT DEATH OCCURRED. DURING A CONCOMITANT PULSE FIELD ABLATION PROCEDURE, A VERSACROSS LARGE ACCESS WAS SELECTED FOR USE. A BASELINE EFFUSION WAS NOTED PRIOR TO TRANSSEPTAL. ONCE GROIN ACCESS WAS OBTAINED, THERE WAS DIFFICULTY ASCENDING THE FEMORAL VEIN AND CROSSING TRANSEPTAL. TRANSEPTAL WAS SUDDEN AND VISIBILITY WAS POOR. THE PHYSICIAN FLOSSED WITH THE VERSACROSS LARGE ACCESS. HOWEVER, THE FARADRIVE SHEATH COULD NOT GO ACROSS. SO, IT WAS ATTEMPTED WITH THE FARADRIVE DILATOR AND TRUSTEER DILATOR, THEN WAS ABLE TO CROSS WITH FARADRIVE. THE PRESSURE DROPPED A FEW TIMES AT TRANSEPTAL DURING FLOSSING. THE PRESSURE WAS TREATED. THE PHYSICIAN USED INTRACARDIAC ECHOCARDIOGRAPHY (ICE) TO WATCH FOR EFFUSION WHILE ABLATION. IT STARTED GROWING WHEN ABLATING RIGHT VEINS. A PERICARDIOCENTESIS WAS DONE AFTER FINISHING ABLATION. ONE LITER WAS DRAWN AND PATIENT WAS STABILIZED. LATER, THE PATIENT PASSED AWAY IN THE INTENSIVE CARE UNIT (ICU) FROM COMPLICATIONS. THE DEVICE IS NOT EXPECTED TO BE RETURNED FOR ANALYSIS (DISPOSED). Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. GOOD FAITH EFFORT ATTEMPTS TO TRY AND RETRIEVE ADDITIONAL DETAILS REGARDING THE REPORTED EVENT ARE IN PROGRESS, BUT FURTHER INFORMATION WAS UNABLE TO BE OBTAINED YET. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22833310&pc=FRN | 3014732157-2025-00823 |  | Death | FRESENIUS KABI USA, LLC | 2025-08-18 04:00 | FRN | IVENIX INFUSION SYSTEM | Use of Incorrect Control/Treatment Settings | Insufficient Information | K183311 |  | 1 | Event Description: (B)(6) CLINICAL IMPLEMENTATION SPECIALIST REPORTED THE FOLLOWING: "I AM SUBMITTING THE REPORTING OF THIS ADVERSE EVENT THAT WAS DISCOVERED DURING AN ONSITE FOLLOW-UP VISIT THIS WEEK. IT WAS REPORTED THAT A CARE AND COMFORT ONLY PATIENT WAS INADVERTENTLY GIVEN AN INCORRECT DOSE OF FENTANYL, THAT CAUSED HARM AND PATIENT DEATH. THE NURSE DELIVERING THE CARE DID NOT ATTEND ANY FK TRAINING CLASSES THAT WERE PROVIDED. THE PATIENT WAS ON THE MEDICAL/SURGICAL FLOOR AND THE NURSE COULD NOT LOCATE THE FENTANYL IN THE M/S DRUG LIBRARY. THEY DID NOT KNOW TO LOOK IN ANOTHER CARE PROFILE TO FIND THE DRUG, SO PROCEEDED TO USE THE DRUG CALCULATOR AND PROGRAMMED THE DOSE INCORRECTLY. THE NURSE OBTAINED THE 2ND CHECK OFF FROM HIS BROTHER (WHO IS ALSO A NURSE AT THE HOSPITAL), WHO SIGNED OFF THE MEDICATION WITHOUT LOOKING AT THE DOSE THAT WAS PROGRAMMED IN THE PUMP. UNKNOWN DATE AND TIME OF EVENT. DISCOVERED DURING FOLLOW-UP VISIT ON (B)(6) 2025." BOTH PHARMA AND MEDTECH DEPARTMENTS AT (B)(6) ATTEMPTED TO OBTAIN ADDITIONAL INFORMATION FROM THE HOSPITAL ON THESE DATES: ON (B)(6) 2025. DESPITE MULTIPLE FOLLOW-UP ATTEMPTS BY (B)(6) PERSONNEL, THE HOSPITAL MAINTAINS THAT IT HAS NO ADDITIONAL INFORMATION TO PROVIDE. ALTHOUGH (B)(6) PERSONNEL HAVE DILIGENTLY SOUGHT CRITICAL DETAILS REGARDING THE CASE, INCLUDING THE EXACT DATE OF THE EVENT, THE PUMP'S SERIAL NUMBER, AND WHETHER THE DRUG WAS IN AN IV BAG OR A PRE-FILLED SYRINGE, WE HAVE BEEN UNABLE TO OBTAIN THE NECESSARY INFORMATION TO COMPLETE A THOROUGH INVESTIGATION. REPORTING DUE TO THE REFERENCED TECHNICAL ISSUE; ADVERSE EFFECT (PATIENT DEATH) REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22833417&pc=LOX | 2024168-2025-08628 | 2025-03-31 04:00 | Death | ABBOTT VASCULAR INC. | 2025-08-18 04:00 | LOX | NC TREK NEO¿ | Adverse Event Without Identified Device or Use Problem | Unspecified Heart Problem | K220634 |  | 1 | Event Description: THIS IS FILED TO REPORT PATIENT DEATH. IT WAS REPORTED THROUGH THE PMCF (POST-MARKET CLINICAL FOLLOW-UP) REPORT, THAT ABBOTT PROACTIVELY COLLECTED AND EVALUATED DATA FROM SOURCES, SUCH AS IN-HOSPITAL OUTCOMES AND PHYSICIAN SURVEYS, TO CONFIRM THE SAFETY AND PERFORMANCE OF ABBOTT'S CORONARY DILATATION CATHETERS (CDCS). KEY SAFETY OUTCOMES COLLECTED WERE DEATH, PERFORATION, DISSECTION, EMBOLISM, ARRHYTHMIAS, MYOCARDIAL INFARCTION, AND OCCLUSION. DEVICE PERFORMANCE WAS DEFINED AS SUCCESSFUL DELIVERY, INFLATION, AND DEFLATION OF THE CDCS. OFF-LABEL USE WAS REPORTED FOR USE IN CHRONIC TOTAL OCCLUSION (CTO) LESIONS FOR TREK AND NC TREK RX. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. B2: DATE OF DEATH: ESTIMATED. B3: DATE OF EVENT: ESTIMATED. D4: THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. THE ADDITIONAL PATIENT EFFECTS, DEVICE MALFUNCTIONS REPORTED IN THE PMCF ARE CAPTURED UNDER SEPARATE MEDWATCH REPORTS. AS MULTIPLE DEVICES WERE CAPTURED IN THE PMCF REPORT. ADDITIONAL REPORTS WERE FILED UNDER SEPARATE REPORT NUMBERS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22833450&pc=LOX | 2024168-2025-08624 | 2025-03-31 04:00 | Death | ABBOTT VASCULAR INC. | 2025-08-18 04:00 | LOX | MINI TREK¿ | Adverse Event Without Identified Device or Use Problem | Unspecified Heart Problem | K123279 |  | 1 | Event Description: THIS IS FILED TO REPORT PATIENT DEATH. IT WAS REPORTED THROUGH THE PMCF (POST-MARKET CLINICAL FOLLOW-UP) REPORT, THAT ABBOTT PROACTIVELY COLLECTED AND EVALUATED DATA FROM SOURCES, SUCH AS IN-HOSPITAL OUTCOMES AND PHYSICIAN SURVEYS, TO CONFIRM THE SAFETY AND PERFORMANCE OF ABBOTT'S CORONARY DILATATION CATHETERS (CDCS). KEY SAFETY OUTCOMES COLLECTED WERE DEATH, PERFORATION, DISSECTION, EMBOLISM, ARRHYTHMIAS, MYOCARDIAL INFARCTION, AND OCCLUSION. DEVICE PERFORMANCE WAS DEFINED AS SUCCESSFUL DELIVERY, INFLATION, AND DEFLATION OF THE CDCS. OFF-LABEL USE WAS REPORTED FOR USE IN CHRONIC TOTAL OCCLUSION (CTO) LESIONS FOR TREK AND NC TREK RX. DETAILS ARE LISTED IN THE ATTACHED REPORT, TITLED POST-MARKET CLINICAL FOLLOW-UP EVALUATION REPORT CORONARY DILATATION CATHETERS. PLEASE SEE THE ATTACHED POST MARKET CLINICAL FOLLOW UP EVALUATION REPORT FOR SPECIFIC INFORMATION. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. B2 - DATE OF DEATH: ESTIMATED B3 - DATE OF EVENT: ESTIMATED D4: THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. THE ADDITIONAL PATIENT EFFECTS, DEVICE MALFUNCTIONS REPORTED IN THE PMCF ARE CAPTURED UNDER SEPARATE MEDWATCH REPORTS. AS MULTIPLE DEVICES WERE CAPTURED IN THE PMCF REPORT. ADDITIONAL REPORTS WERE FILED UNDER SEPARATE REPORT NUMBERS. LITERATURE ATTACHMENT: POST-MARKET CLINICAL FOLLOW-UP (PMCF) EVALUATION REPORT CORONARY DILATATION CATHETERS FAMILY, RPT2121578 REV H.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22833462&pc=LOX | 2024168-2025-08630 | 2025-03-31 04:00 | Death | ABBOTT VASCULAR INC. | 2025-08-18 04:00 | LOX | TREK | Adverse Event Without Identified Device or Use Problem | Unspecified Heart Problem | K123279 |  | 1 | Event Description: THIS IS FILED TO REPORT PATIENT DEATH. IT WAS REPORTED THROUGH THE PMCF (POST-MARKET CLINICAL FOLLOW-UP) REPORT, THAT ABBOTT PROACTIVELY COLLECTED AND EVALUATED DATA FROM SOURCES, SUCH AS IN-HOSPITAL OUTCOMES AND PHYSICIAN SURVEYS, TO CONFIRM THE SAFETY AND PERFORMANCE OF ABBOTT'S CORONARY DILATATION CATHETERS (CDCS). KEY SAFETY OUTCOMES COLLECTED WERE DEATH, PERFORATION, DISSECTION, EMBOLISM, ARRHYTHMIAS, MYOCARDIAL INFARCTION, AND OCCLUSION. DEVICE PERFORMANCE WAS DEFINED AS SUCCESSFUL DELIVERY, INFLATION, AND DEFLATION OF THE CDCS. OFF-LABEL USE WAS REPORTED FOR USE IN CHRONIC TOTAL OCCLUSION (CTO) LESIONS FOR TREK AND NC TREK RX. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. B2: DATE OF DEATH: ESTIMATED. B3: DATE OF EVENT: ESTIMATED. D4: THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. THE ADDITIONAL PATIENT EFFECTS, DEVICE MALFUNCTIONS REPORTED IN THE PMCF ARE CAPTURED UNDER SEPARATE MEDWATCH REPORTS. AS MULTIPLE DEVICES WERE CAPTURED IN THE PMCF REPORT. ADDITIONAL REPORTS WERE FILED UNDER SEPARATE REPORT NUMBERS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22833465&pc=LOX | 2024168-2025-08632 | 2025-03-31 04:00 | Death | ABBOTT VASCULAR | 2025-08-18 04:00 | LOX | TRAVELER RX CORONARY DILATATION CATHETER | Adverse Event Without Identified Device or Use Problem | Unspecified Heart Problem | K103110 |  | 1 | Event Description: THIS IS FILED TO REPORT PATIENT DEATH. IT WAS REPORTED THROUGH THE PMCF (POST-MARKET CLINICAL FOLLOW-UP) REPORT, THAT ABBOTT PROACTIVELY COLLECTED AND EVALUATED DATA FROM SOURCES, SUCH AS IN-HOSPITAL OUTCOMES AND PHYSICIAN SURVEYS, TO CONFIRM THE SAFETY AND PERFORMANCE OF ABBOTT'S CORONARY DILATATION CATHETERS (CDCS). KEY SAFETY OUTCOMES COLLECTED WERE DEATH, PERFORATION, DISSECTION, EMBOLISM, ARRHYTHMIAS, MYOCARDIAL INFARCTION, AND OCCLUSION. DEVICE PERFORMANCE WAS DEFINED AS SUCCESSFUL DELIVERY, INFLATION, AND DEFLATION OF THE CDCS. OFF-LABEL USE WAS REPORTED FOR USE IN CHRONIC TOTAL OCCLUSION (CTO) LESIONS FOR TREK AND NC TREK RX. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. B2: DATE OF DEATH: ESTIMATED. B3: DATE OF EVENT: ESTIMATED. D4: THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. THE ADDITIONAL PATIENT EFFECTS, DEVICE MALFUNCTIONS REPORTED IN THE PMCF ARE CAPTURED UNDER SEPARATE MEDWATCH REPORTS. AS MULTIPLE DEVICES WERE CAPTURED IN THE PMCF REPORT. ADDITIONAL REPORTS WERE FILED UNDER SEPARATE REPORT NUMBERS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22833493&pc=LOX | 2024168-2025-08634 | 2025-03-31 04:00 | Death | ABBOTT VASCULAR | 2025-08-18 04:00 | LOX | NC TRAVELER CORONARY DILATATION CATHETER | Adverse Event Without Identified Device or Use Problem | Unspecified Heart Problem | K110134 |  | 1 | Event Description: THIS IS FILED TO REPORT PATIENT DEATH. IT WAS REPORTED THROUGH THE PMCF (POST-MARKET CLINICAL FOLLOW-UP) REPORT, THAT ABBOTT PROACTIVELY COLLECTED AND EVALUATED DATA FROM SOURCES, SUCH AS IN-HOSPITAL OUTCOMES AND PHYSICIAN SURVEYS, TO CONFIRM THE SAFETY AND PERFORMANCE OF ABBOTT'S CORONARY DILATATION CATHETERS (CDCS). KEY SAFETY OUTCOMES COLLECTED WERE DEATH, PERFORATION, DISSECTION, EMBOLISM, ARRHYTHMIAS, MYOCARDIAL INFARCTION, AND OCCLUSION. DEVICE PERFORMANCE WAS DEFINED AS SUCCESSFUL DELIVERY, INFLATION, AND DEFLATION OF THE CDCS. OFF-LABEL USE WAS REPORTED FOR USE IN CHRONIC TOTAL OCCLUSION (CTO) LESIONS FOR TREK AND NC TREK RX. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. B2 - DATE OF DEATH: ESTIMATED. B3 - DATE OF EVENT: ESTIMATED. D4: THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. THE ADDITIONAL PATIENT EFFECTS, DEVICE MALFUNCTIONS REPORTED IN THE PMCF ARE CAPTURED UNDER SEPARATE MEDWATCH REPORTS. AS MULTIPLE DEVICES WERE CAPTURED IN THE PMCF REPORT. ADDITIONAL REPORTS WERE FILED UNDER SEPARATE REPORT NUMBERS. LITERATURE ATTACHMENT: POST-MARKET CLINICAL FOLLOW-UP (PMCF) EVALUATION REPORT CORONARY DILATATION CATHETERS FAMILY, (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22833498&pc=LOX | 2024168-2025-08640 | 2025-03-31 04:00 | Death | ABBOTT VASCULAR INC. | 2025-08-18 04:00 | LOX | NC TREK | Adverse Event Without Identified Device or Use Problem | Unspecified Heart Problem | K110134 |  | 1 | Event Description: THIS IS FILED TO REPORT PATIENT DEATH. IT WAS REPORTED THROUGH THE PMCF (POST-MARKET CLINICAL FOLLOW-UP) REPORT, THAT ABBOTT PROACTIVELY COLLECTED AND EVALUATED DATA FROM SOURCES, SUCH AS IN-HOSPITAL OUTCOMES AND PHYSICIAN SURVEYS, TO CONFIRM THE SAFETY AND PERFORMANCE OF ABBOTT'S CORONARY DILATATION CATHETERS (CDCS). KEY SAFETY OUTCOMES COLLECTED WERE DEATH, PERFORATION, DISSECTION, EMBOLISM, ARRHYTHMIAS, MYOCARDIAL INFARCTION, OCCLUSION, AND HYPOTENSION. DEVICE PERFORMANCE WAS DEFINED AS SUCCESSFUL DELIVERY, INFLATION, AND DEFLATION OF THE CDCS. OFF-LABEL USE WAS REPORTED FOR USE IN CHRONIC TOTAL OCCLUSION (CTO) LESIONS FOR TREK AND NC TREK RX. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. B2: DATE OF DEATH: ESTIMATED. B3: DATE OF EVENT: ESTIMATED. D4: THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. THE ADDITIONAL PATIENT EFFECTS, DEVICE MALFUNCTIONS REPORTED IN THE PMCF ARE CAPTURED UNDER SEPARATE MEDWATCH REPORTS. AS MULTIPLE DEVICES WERE CAPTURED IN THE PMCF REPORT. ADDITIONAL REPORTS WERE FILED UNDER SEPARATE REPORT NUMBERS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22833503&pc=LOX | 2024168-2025-08641 | 2025-03-31 04:00 | Death | ABBOTT VASCULAR INC. | 2025-08-18 04:00 | LOX | TREK | Adverse Event Without Identified Device or Use Problem | Unspecified Heart Problem | K103110 |  | 1 | Event Description: THIS IS FILED TO REPORT PATIENT DEATH. IT WAS REPORTED THROUGH THE PMCF (POST-MARKET CLINICAL FOLLOW-UP) REPORT, THAT ABBOTT PROACTIVELY COLLECTED AND EVALUATED DATA FROM SOURCES, SUCH AS IN-HOSPITAL OUTCOMES AND PHYSICIAN SURVEYS, TO CONFIRM THE SAFETY AND PERFORMANCE OF ABBOTT'S CORONARY DILATATION CATHETERS (CDCS). KEY SAFETY OUTCOMES COLLECTED WERE DEATH, PERFORATION, DISSECTION, EMBOLISM, ARRHYTHMIAS, MYOCARDIAL INFARCTION, OCCLUSION, AND HYPOTENSION. DEVICE PERFORMANCE WAS DEFINED AS SUCCESSFUL DELIVERY, INFLATION, AND DEFLATION OF THE CDCS. OFF-LABEL USE WAS REPORTED FOR USE IN CHRONIC TOTAL OCCLUSION (CTO) LESIONS FOR TREK AND NC TREK RX. DETAILS ARE LISTED IN THE ATTACHED REPORT, TITLED POST-MARKET CLINICAL FOLLOW-UP EVALUATION REPORT CORONARY DILATATION CATHETERS. PLEASE SEE THE ATTACHED POST MARKET CLINICAL FOLLOW UP EVALUATION REPORT FOR SPECIFIC INFORMATION. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. B2 - DATE OF DEATH: ESTIMATED. B3 - DATE OF EVENT: ESTIMATED. D4: THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. THE ADDITIONAL PATIENT EFFECTS, DEVICE MALFUNCTIONS REPORTED IN THE PMCF ARE CAPTURED UNDER SEPARATE MEDWATCH REPORTS. AS MULTIPLE DEVICES WERE CAPTURED IN THE PMCF REPORT. ADDITIONAL REPORTS WERE FILED UNDER SEPARATE REPORT NUMBERS. LITERATURE ATTACHMENT: POST-MARKET CLINICAL FOLLOW-UP (PMCF) EVALUATION REPORT CORONARY DILATATION CATHETERS FAMILY, RPT2121578 REV H.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22833519&pc=LOX | 2024168-2025-08642 | 2025-03-31 04:00 | Death | ABBOTT VASCULAR INC. | 2025-08-18 04:00 | LOX | MINI TREK | Adverse Event Without Identified Device or Use Problem | Unspecified Heart Problem | K103110 |  | 1 | Event Description: THIS IS FILED TO REPORT PATIENT DEATH. IT WAS REPORTED THROUGH THE PMCF (POST-MARKET CLINICAL FOLLOW-UP) REPORT, THAT ABBOTT PROACTIVELY COLLECTED AND EVALUATED DATA FROM SOURCES, SUCH AS IN-HOSPITAL OUTCOMES AND PHYSICIAN SURVEYS, TO CONFIRM THE SAFETY AND PERFORMANCE OF ABBOTT'S CORONARY DILATATION CATHETERS (CDCS). KEY SAFETY OUTCOMES COLLECTED WERE DEATH, PERFORATION, DISSECTION, EMBOLISM, ARRHYTHMIAS, MYOCARDIAL INFARCTION, OCCLUSION, AND HYPOTENSION. DEVICE PERFORMANCE WAS DEFINED AS SUCCESSFUL DELIVERY, INFLATION, AND DEFLATION OF THE CDCS. OFF-LABEL USE WAS REPORTED FOR USE IN CHRONIC TOTAL OCCLUSION (CTO) LESIONS FOR TREK AND NC TREK RX. DETAILS ARE LISTED IN THE ATTACHED REPORT, TITLED POST-MARKET CLINICAL FOLLOW-UP EVALUATION REPORT CORONARY DILATATION CATHETERS. PLEASE SEE THE ATTACHED POST MARKET CLINICAL FOLLOW UP EVALUATION REPORT FOR SPECIFIC INFORMATION. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. B2 - DATE OF DEATH: ESTIMATED. B3 - DATE OF EVENT: ESTIMATED. D4: THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. THE ADDITIONAL PATIENT EFFECTS, DEVICE MALFUNCTIONS REPORTED IN THE PMCF ARE CAPTURED UNDER SEPARATE MEDWATCH REPORTS. AS MULTIPLE DEVICES WERE CAPTURED IN THE PMCF REPORT. ADDITIONAL REPORTS WERE FILED UNDER SEPARATE REPORT NUMBERS. LITERATURE ATTACHMENT: POST-MARKET CLINICAL FOLLOW-UP (PMCF) EVALUATION REPORT CORONARY DILATATION CATHETERS FAMILY, RPT2121578 REV H.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22833581&pc=NPT | 2015691-2025-06780 | 2025-07-15 04:00 | Death | EDWARDS LIFESCIENCES | 2025-08-18 04:00 | NPT | EDWARDS SAPIEN 3 TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem |  | P140031 |  | 1 | Event Description: AS REPORTED BY OUR VENEZUELAN AFFILIATES, DURING IMPLANT OF A 20MM SAPIEN 3 VALVE IN THE AORTIC POSITION BY TRANSFEMORAL APPROACH, THE VALVE WAS SUCCESSFULLY DEPLOYED, BUT AFTER IMPLANTATION, THE PATIENT BECAME HEMODYNAMICALLY UNSTABLE (60/30 MMHG AND 30 BPM), LEADING TO CARDIAC AND RESPIRATORY ARREST. RESUSCITATION WAS STARTED WITH INTUBATION, THORACIC COMPRESSIONS AND VASOACTIVE DRUGS, AS WELL AS THE PATIENT BEING PACED AT 80BPM THROUGH THE GUIDEWIRE. THE FIRST CORONARY ANGIO SHOT SHOWED EVIDENCE OF CORONARY PATENCY, AND THE ECHO SHOWED SEVERE SYSTOLIC DYSFUNCTION WITH SUSPECTED GLOBAL ISCHEMIA. SELECTIVE CANNULATION OF THE LEFT MAIN CORONARY ARTERY WAS PERFORMED, GUIDEWIRE WAS ADVANCE TO THE MID SEGMENT OF THE LAD, AND BALLOON DILATION OF THE LEFT MAIN OSTIUM WAS PERFORMED. A NEW ANGIOGRAPHIC ACQUISITION SHOWED TIMI I FLOW (VERY LIMITED OR ABSENT FLOW). RESUSCITATION EFFORTS WERE MAINTAINED FOR APPROXIMATELY 50 MINUTES WITH NO RETURN OF SPONTANEOUS CIRCULATION AND THE PATIENT WAS DECLARED DEAD. Manufacturer Narrative: THE EVENT REPORTED IS AN ANTICIPATED IN THE RISK MANAGEMENT DOCUMENTATION FOR TRANSCATHETER HEART VALVE PROCEDURES. A PREVIOUS INVESTIGATION INTO THIS TYPE OF EVENT IS CAPTURED IN AN EDWARDS LIFESCIENCES TECHNICAL SUMMARY AND APPLIES TO THIS COMPLAINT. ADDITIONAL ASSESSMENT OF THE FAILURE MODE IS NOT REQUIRED AT THIS TIME. DUE TO THE UNAVAILABILITY OF THE COMPLAINT DEVICE, ENGINEERING WAS UNABLE TO PERFORM ANY VISUAL INSPECTION, FUNCTIONAL TESTING, OR DIMENSIONAL ANALYSIS. PER THE TECHNICAL SUMMARY, THE IFU, CURRENT RISK MITIGATIONS INCLUDE DESIGN AND MANUFACTURING CONTROLS, AND TRAINING MANUALS HAVE BEEN REVIEWED, NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. PER THE INSTRUCTIONS FOR USE (IFU), CORONARY FLOW OBSTRUCTION IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. THE IFU CAUTIONS THAT SAFETY AND EFFECTIVENESS HAVE NOT BEEN ESTABLISHED FOR PATIENTS WITH BULKY CALCIFIED AORTIC VALVE LEAFLETS NEAR CORONARY OSTIA. IN ADDITION, IT WARNS THAT CAUTION SHOULD BE EXERCISED IN IMPLANTING A BIOPROSTHESIS IN PATIENTS WITH CLINICALLY SIGNIFICANT CORONARY ARTERY DISEASE. CORONARY OBSTRUCTION CAN RESULT IN MYOCARDIAL ISCHEMIA OR INFARCTION DUE TO OBSTRUCTION OF THE CORONARY BLOOD FLOW AND MAY REQUIRE INTERVENTION (E.G., PERCUTANEOUS CORONARY INTERVENTION). MULTIPLE PATIENT FACTORS COULD PUT THE PATIENT AT RISK FOR CORONARY FLOW OBSTRUCTION DURING THE TVR PROCEDURE, INCLUDING SIGNIFICANT UNDERLYING CORONARY ARTERY DISEASE AND BULKY CALCIFICATION OF THE NATIVE LEAFLETS AND ROOT. DISPLACEMENT OF CALCIUM DEPOSITS WITH EMBOLIZATION OF DEBRIS INTO ONE OR MORE OF THE CORONARY ARTERIES, AND/OR AN AORTIC DISSECTION WITH CONTINUITY OF THE RUPTURE INTO THE INTIMA OF ONE OF THE CORONARY OSTIA, CAN RESULT TO A CORONARY OBSTRUCTION. ADDITIONALLY, MINIMAL DISTANCE BETWEEN THE NATIVE ANNULUS AND THE CORONARY OSTIA, BULKY CALCIFICATION, LONG NATIVE LEAFLETS, AND OBLITERATED CORONARY SINUSES. PROCEDURAL FACTORS SUCH AS TORN NATIVE LEAFLET DURING BALLOON VALVULOPLASTY, PLAQUE SHIFT, SIGNIFICANT VALVE OVER SIZING, AND VALVE MALPOSITION COULD ALSO CONTRIBUTE TO THIS COMPLICATION. THE DEVICE TRAINING MANUALS INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVES (ALL MODELS). TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES. THE PHYSICIAN IS INSTRUCTED TO EVALUATE THIS RISK EARLY IN THE PATIENT SCREENING PROCESS IN ALL PATIENTS. THE FOLLOWING FACTORS SHOULD BE CONSIDERED: DEGREE OF CALCIFICATION ON LEAFLETS, LANDING ZONE TO CORONARY OSTIA DISTANCE, LENGTH OF THE NATIVE VALVE LEAFLET, WIDTH OF THE VALSALVA SINUSES, MOVEMENT OF THE LEAFLETS DURING BALLOON VALVULOPLASTY (BV), PATENCY OF CORONARIES DURING BV, AND EXPANDED HEIGHT OF THE INTENDED VALVE. ASSESSMENT FOR CORONARY COMPRESSION RISK PRIOR TO IMPLANTATION IS RECOMMENDED FOR VALVES IMPLANTED IN PULMONARY POSITION. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. DUE TO A LIMITED AMOUNT OF INFORMATION RECEIVED, IT IS UNKNOWN WHAT COULD HAVE CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22833591&pc=OZD | 1220648-2025-45805 | 2025-07-25 04:00 | Death | ABIOMED, INC. | 2025-08-18 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT IN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. ON ARRIVAL TO INTENSIVE CARE UNIT (ICU), THE PATIENT HAD PROFUSE BLEEDING FROM THE GROIN DUE TO POOR DRESSING FROM TEAM RUSHING TO GET TO THE ICU. IMPELLA POSITION WAS UNCLEAR AND THE PATIENT WAS GIVEN SEVEN UNITS OF PACKED RED BLOOD CELLS, SEVEN UNITS OF FRESH FROZEN PLASMA, AND ONE UNIT OF PLATELETS. ADDITIONALLY, MULTIPLE UNITS OF BLOOD WERE TRANSFUSED THROUGHOUT THE DAY BUT THE BLEEDING WAS UNABLE TO BE STOPPED. THE PATIENT WAS TOO UNSTABLE TO TAKE FOR A COMPUTED TOMOGRAPHY OR TO THE OPERATING ROOM FOR EXPLORATION AND RESOLUTION OF THE BLEEDING. THE SOURCE OF THE BLEED WAS UNKNOWN. HOWEVER, THE PHYSICIAN FELT IT WAS LIKELY THAT THERE MAY HAVE BEEN A RETROPERITONEAL BLEED SINCE THE IMPELLA WAS PLACED DURING CARDIOPULMONARY RESUSCITATION. THE FAMILY MADE THE DECISION TO WITHDRAW CARE AFTER EXTENSIVE DISCUSSION WITH THE MEDICAL TEAM AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22833674&pc=KWA | 0001825034-2025-02500 | 2022-10-16 04:00 | Death | ZIMMER BIOMET, INC. | 2025-08-18 04:00 | KWA | M2A MAGNUM | Material Erosion; Adverse Event Without Identified Device or Use Problem |  | K062995 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED JOINT EFFUSION, ELEVATED METAL ION LEVELS, AND EXPIRED APPROXIMATELY 11 YEARS POST-IMPLANTATION. ATTEMPTS HAVE BEEN MADE AND NO FURTHER INFORMATION HAS BEEN PROVIDED. Manufacturer Narrative: (B)(4). G2: FOREIGN ¿ CANADA. THE CUSTOMER HAS INDICATED THAT THE PRODUCT WILL NOT BE RETURNED TO ZIMMER BIOMET FOR INVESTIGATION, AS ITS LOCATION IS UNKNOWN. THE INVESTIGATION IS IN PROCESS. ONCE THE INVESTIGATION HAS BEEN COMPLETED, A FOLLOW-UP MDR WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22833676&pc=JDI | 0001825034-2025-02501 | 2022-10-16 04:00 | Death | ZIMMER BIOMET, INC. | 2025-08-18 04:00 | JDI | TAPERLOC | Material Erosion; Adverse Event Without Identified Device or Use Problem |  | K830313 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED JOINT EFFUSION, ELEVATED METAL ION LEVELS, AND EXPIRED APPROXIMATELY 11 YEARS POST-IMPLANTATION. ATTEMPTS HAVE BEEN MADE AND NO FURTHER INFORMATION HAS BEEN PROVIDED. Manufacturer Narrative: (B)(4). G2: FOREIGN ¿ CANADA. H6: SUGGESTED COMPONENT CODE: MECHANICAL (G04) ¿ STEM. THE CUSTOMER HAS INDICATED THAT THE PRODUCT WILL NOT BE RETURNED TO ZIMMER BIOMET FOR INVESTIGATION, AS ITS LOCATION IS UNKNOWN. THE INVESTIGATION IS IN PROCESS. ONCE THE INVESTIGATION HAS BEEN COMPLETED, A FOLLOW-UP MDR WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22833680&pc=LPH | 0001825034-2025-02502 | 2022-10-16 04:00 | Death | ZIMMER BIOMET, INC. | 2025-08-18 04:00 | LPH | M2A MAGNUM | Material Erosion; Adverse Event Without Identified Device or Use Problem |  | K101336 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED JOINT EFFUSION, ELEVATED METAL ION LEVELS, AND EXPIRED APPROXIMATELY 11 YEARS POST-IMPLANTATION. ATTEMPTS HAVE BEEN MADE AND NO FURTHER INFORMATION HAS BEEN PROVIDED. Manufacturer Narrative: (B)(4). G2: FOREIGN ¿ CANADA. THE CUSTOMER HAS INDICATED THAT THE PRODUCT WILL NOT BE RETURNED TO ZIMMER BIOMET FOR INVESTIGATION, AS ITS LOCATION IS UNKNOWN. THE INVESTIGATION IS IN PROCESS. ONCE THE INVESTIGATION HAS BEEN COMPLETED, A FOLLOW-UP MDR WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22833682&pc=KWA | 0001825034-2025-02503 | 2022-10-16 04:00 | Death | ZIMMER BIOMET, INC. | 2025-08-18 04:00 | KWA | M2A MAGNUM | Material Erosion; Adverse Event Without Identified Device or Use Problem |  | K042037 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED JOINT EFFUSION, ELEVATED METAL ION LEVELS, AND EXPIRED APPROXIMATELY 11 YEARS POST-IMPLANTATION. ATTEMPTS HAVE BEEN MADE AND NO FURTHER INFORMATION HAS BEEN PROVIDED. Manufacturer Narrative: (B)(4). G2: FOREIGN ¿ CANADA. H6: SUGGESTED COMPONENT CODE: MECHANICAL (G04) ¿ HEAD. THE CUSTOMER HAS INDICATED THAT THE PRODUCT WILL NOT BE RETURNED TO ZIMMER BIOMET FOR INVESTIGATION, AS ITS LOCATION IS UNKNOWN. THE INVESTIGATION IS IN PROCESS. ONCE THE INVESTIGATION HAS BEEN COMPLETED, A FOLLOW-UP MDR WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22833790&pc=KNT | 3010757606-2025-00366 | 2025-07-27 04:00 | Death | ABBVIE - MEDICAL DEVICE CENTER | 2025-08-18 04:00 | KNT | DUODOPA_DUOPA | Adverse Event Without Identified Device or Use Problem | Aspiration Pneumonitis | K142793 |  | 1 | Event Description: ON (B)(6) 2025, A PATIENT IN GERMANY UNDERWENT A PROCEDURE FOR THE PLACEMENT OF PERCUTANEOUS ENDOSCOPIC GASTROSTOMY (PEG) TUBE WITH JEJUNAL (PEG-J) TUBE. THE PATIENT INITIATED DUOPA THERAPY DUE TO GENERAL DETERIORATION, INCLUDING THE NEED TO RECEIVE ENTERAL FEEDING. THE PATIENT WAS DIAGNOSED WITH ASPIRATION PNEUMONIA, WITHIN TEN DAYS OF DEVICE IMPLANTATION. IT WAS REPORTED THAT THE PATIENT EXPIRED ON (B)(6) 2025. THE CAUSE OF DEATH IS UNKNOWN. IT IS UNKNOWN IF AN AUTOPSY WAS PERFORMED. Manufacturer Narrative: CATALOG NUMBER 062941-001 IS THE INTERNATIONAL LIST NUMBER WHICH MEETS THE REQUIREMENTS OF THE SIMILAR PRODUCT LISTED IN SECTION D4 WHICH IS THE US LIST NUMBER. THE DEVICE INVOLVED IN THE EVENT WAS NOT RETURNED; THEREFORE, A RETURN SAMPLE EVALUATION IS UNABLE TO BE PERFORMED. POST PROCEDURAL COMPLICATIONS INCLUDING ASPIRATION PNEUMONIA AND DEATH ARE KNOWN COMPLICATION OF A PEG TUBE PLACEMENT. IF ANY FURTHER RELEVANT INFORMATION IS IDENTIFIED OR OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22833889&pc=NPT | 2135147-2025-04623 | 2025-07-18 04:00 | Death | ABBOTT MEDICAL | 2025-08-18 04:00 | NPT | NAVITOR TRANSCATHETER AORTIC VALVE | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P190023 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 25MM NAVITOR VALVE WAS SELECTED FOR IMPLANT. PRIOR TO PROCEDURE, THE PATIENT WAS "FINE TO DO THE PROCEDURE." THE VALVE WAS IMPLANTED SUCCESSFULLY AND THE PATIENT WAS REPORTED TO DO "WELL IMMEDIATELY AFTER THE PROCEDURE." ON (B)(6) 2025, THE PATIENT REPORTEDLY BECAME ILL AND "EXPERIENCED CLINICAL COMPLICATIONS" AND PASSED AWAY. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22833907&pc=MIH | 2247858-2025-00192 | 2024-12-11 05:00 | Death | BOLTON MEDICAL, INC. | 2025-08-18 04:00 | MIH | RELAY PRO NBS THORACIC STENT-GRAFT SYSTEM | Adverse Event Without Identified Device or Use Problem | Pseudoaneurysm | P200045 |  | 1 | Event Description: "ON POST-OPERATIVE DAY 189 (B)(6) 2024) OF THE THORAFLEX HYBRID PLEXUS, SUBJECT 017-008 EXPERIENCED AN AE OF RIGHT COMMON FEMORAL ARTERY PSEUDOANEURYSM, WITH FURTHER DETAILS STATING "A SMALL RIGHT COMMON FEMORAL ARTERY PSEUDOANEURYSM (14MM) WITH FOCAL DISSECTION" THERE IS CURRENTLY NO INFORMATION AVAILABLE ON THE TREATMENT OR OUTCOME OF THE EVENT. AS EXPLAINED ABOVE, THERE IS NO CAUSALITY ASSESSMENT FOR THE EVENT, AND SITE HAS BEEN REQUESTED TO COMPLETE AS SOON AS POSSIBLE. MEDICAL MONITOR HAS ASSESSED THE EVENT AS POSSIBLY RELATED TO BOTH THORAFLEX HYBRID AND RELAYPRO NBS DEVICES, HENCE REPORTABLE. PATIENT DEATH: (B)(6) 2025 - CAUSE OF DEATH IS UNKNOWN, UNDETERMINED IS CAUSE OF DEATH IS RELATED TO THE DEVICE." PATIENT OUTCOME: "THERE IS CURRENTLY NO DETAILS AVAILABLE ON THE OUTCOME OF THE EVENT. 59 DAYS POST-OP THE PATIENT CHART WAS UPDATED TO DECEASED, ADDITIONAL INFORMATION PENDING PATIENT DEATH: (B)(6) 2025 - CAUSE OF DEATH IS UNKNOWN, UNDETERMINED IS CAUSE OF DEATH IS RELATED TO THE DEVICE." Manufacturer Narrative: THIS COMPLAINT WAS INVOLVED WITH TWO DEVICES. DEVICE 1 IS BEING REPORTED UNDER MDR-2247858-2025-00191, AND DEVICE 2 IS BEING REPORTED UNDER. BOLTON MEDICAL, INC. (D/B/A TERUMO AORTIC), HEREIN KNOWN AS THE "COMPANY", IS SUBMITTING THIS REPORT PURSUANT TO 21 CFR PART 803, HAS MADE REASONABLE EFFORTS TO OBTAIN COMPLETE INFORMATION, AND HAS PROVIDED AS MUCH AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT IS BASED ON INFORMATION OBTAINED BY THE COMPANY, WHICH MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE REGULATIONS. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN, UNAVAILABLE, OR NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY ANYONE THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE CERTAIN REGULATIONS, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22833908&pc=MIH | 2247858-2025-00191 | 2024-12-11 05:00 | Death | BOLTON MEDICAL, INC. | 2025-08-18 04:00 | MIH | RELAY PRO NBS THORACIC STENT-GRAFT SYSTEM | Adverse Event Without Identified Device or Use Problem | Pseudoaneurysm | P200045 |  | 1 | Event Description: "ON POST-OPERATIVE DAY 189 (B)(6) 2024) OF THE THORAFLEX HYBRID PLEXUS, SUBJECT (B)(6)EXPERIENCED AN AE OF RIGHT COMMON FEMORAL ARTERY PSEUDOANEURYSM, WITH FURTHER DETAILS STATING "A SMALL RIGHT COMMON FEMORAL ARTERY PSEUDOANEURYSM (14MM) WITH FOCAL DISSECTION". THERE IS CURRENTLY NO INFORMATION AVAILABLE ON THE TREATMENT OR OUTCOME OF THE EVENT. AS EXPLAINED ABOVE, THERE IS NO CAUSALITY ASSESSMENT FOR THE EVENT, AND SITE HAS BEEN REQUESTED TO COMPLETE AS SOON AS POSSIBLE. MEDICAL MONITOR HAS ASSESSED THE EVENT AS POSSIBLY RELATED TO BOTH THORAFLEX HYBRID AND RELAYPRO NBS DEVICES, HENCE REPORTABLE. PATIENT DEATH: (B)(6) 2025 - CAUSE OF DEATH IS UNKNOWN, UNDETERMINED IS CAUSE OF DEATH IS RELATED TO THE DEVICE." PATIENT OUTCOME: "THERE IS CURRENTLY NO DETAILS AVAILABLE ON THE OUTCOME OF THE EVENT. 59 DAYS POST-OP THE PATIENT CHART WAS UPDATED TO DECEASED, ADDITIONAL INFORMATION PENDING PATIENT DEATH: (B)(6) 2025 - CAUSE OF DEATH IS UNKNOWN, UNDETERMINED IS CAUSE OF DEATH IS RELATED TO THE DEVICE." Manufacturer Narrative: THIS COMPLAINT WAS INVOLVED WITH TWO DEVICES. DEVICE 1 IS BEING REPORTED UNDER MDR-. AND DEVICE 2 IS BEING REPORTED UNDER MDR-2247858-2025-00192. BOLTON MEDICAL, INC. (D/B/A TERUMO AORTIC), HEREIN KNOWN AS THE "COMPANY", IS SUBMITTING THIS REPORT PURSUANT TO 21 CFR PART 803, HAS MADE REASONABLE EFFORTS TO OBTAIN COMPLETE INFORMATION, AND HAS PROVIDED AS MUCH AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT IS BASED ON INFORMATION OBTAINED BY THE COMPANY, WHICH MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE REGULATIONS. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN, UNAVAILABLE, OR NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY ANYONE THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE CERTAIN REGULATIONS, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22833990&pc=QFG | 3013756811-2025-195308 | 2025-07-18 04:00 | Death | TANDEM DIABETES CARE | 2025-08-18 04:00 | QFG | T:SLIM X2 INSULIN PUMP WITH CONTROL-IQ TECHNOLOGY | Adverse Event Without Identified Device or Use Problem | Diabetic Ketoacidosis | K203234 |  | 1 | Event Description: IT WAS REPORTED THAT A CUSTOMER HAD PASSED AWAY ON (B)(6) 2025 DUE TO DIABETIC KETOACIDOSIS (DKA). CONTACT REPORTED THAT THE PUMP OR RELATED SUPPLIES DID NOT CONTRIBUTE TO THE DEATH. A REVIEW OF PUMP DATA WILL BE PERFORMED, AND THE RESULTS WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22834391&pc=QBJ | 2954323-2025-32686 | 2025-08-03 04:00 | Death | ABBOTT DIABETES CARE LTD | 2025-08-18 04:00 | QBJ | FREESTYLE LIBRE 3 | Low Readings | Diabetic Ketoacidosis | K223435 |  | 1 | Event Description: A LOW READINGS ISSUE WITH THE ABBOTT DIABETES CARE (ADC) DEVICE WAS REPORTED. THE HEALTHCARE PROFESSIONAL REPORTED THAT THEIR PATIENT RECEIVED UNSPECIFIED LOWER SENSOR SCAN RESULTS AND EXPERIENCED SEIZURE AND PAIN IN THE LEGS. THE LENGTH OF SENSOR WEAR WAS UNKNOWN AND IT IS UNKNOWN IF THE CUSTOMER NOTED A TREND ARROW ON THE DEVICE. DUE TO LOW READINGS, THE CUSTOMER REPORTED THAT THEY WERE HOSPITALIZED WHERE UNSPECIFIED READINGS WERE TAKEN ON AN HCP (HEALTHCARE PROFESSIONAL) METER FOR DIAGNOSIS OF DIABETIC KETOACIDOSIS WHICH THEY RECEIVED UNSPECIFIED TREATMENT FOR. THE CUSTOMER ALSO REPORTED THAT THEY WENT INTO DELIVERY OF THEIR BABY DUE TO MAJOR FETAL DISTRESS VIA URGENT CESAREAN SECTION AND THE BABY IS CURRENTLY IN CRITICAL CONDITION. THERE WAS NO REPORT OF DEATH OR PERMANENT IMPAIRMENT ASSOCIATED WITH THIS EVENT. Manufacturer Narrative: THE PRODUCT HAS BEEN REQUESTED BACK FOR INVESTIGATION AND THE CUSTOMER AGREED TO RETURN THE DEVICE; HOWEVER, AT THIS TIME PRODUCT HAS NOT YET BEEN RECEIVED. THE CUSTOMERS PRODUCT HAS BEEN REQUESTED FOR INVESTIGATION. AT THIS TIME PRODUCT HAS NOT YET BEEN RETURNED AND A VALID SERIAL NUMBER HAS NOT BEEN PROVIDED. AN EXTENDED INVESTIGATION HAS BEEN PERFORMED FOR THE REPORTED COMPLAINT AND IT HAS BEEN DETERMINED THAT THERE WAS NO INDICATION THAT THE PRODUCT DID NOT MEET SPECIFICATION. CLINICAL DATA WAS REVIEWED AND CONFIRMED THAT LIBRE SENSORS CONTINUE TO BE SAFE, EFFECTIVE, AND PERFORM AS INTENDED IN THE FIELD. STABILITY DATA FOR LIBRE SENSORS WAS REVIEWED AND SHOWED NO ANOMALIES OR NON-CONFORMANCE'S THAT COULD LEAD TO THE COMPLAINT. TRACKING AND TRENDING REPORTS FOR THE PAST YEAR WAS REVIEWED BASED ON THE PRODUCT LINE AND REPORTED ISSUE. THE REVIEW IDENTIFIED A TRIPPED TREND AND CORRECTIVE ACTIONS WERE TAKEN. THE MOST PROBABLE ROOT CAUSES ASSOCIATED WITH THIS FAILURE MODE ARE DISCONNECTED, FAULTY OR DAMAGED COMPONENTS, SOFTWARE/DATA CORRUPTION, OR MISUSE. HOWEVER, MITIGATION'S ARE IN PLACE TO REDUCE AND PREVENT SUCH ISSUES. ALL VITAL MANUFACTURING STEPS ARE VALIDATED, MONITORED, AND VERIFIED DURING MANUFACTURING TO ENSURE THE SYSTEM IS IN CONFORMANCE WITH THE VERIFIED DESIGN. ALL VITAL FUNCTIONS ARE MONITORED BY THE SYSTEM AND, WHEN NECESSARY, FUNCTION IS SUSPENDED TO SAFEGUARD AGAINST INACCURATE RESULTS. LABELING IS PROVIDED TO INSTRUCT THE USER ON THE INTENDED USE OF ALL VITAL PARTS OF THE SYSTEM TO MINIMIZE MISUSE. ALL COMPLAINTS AND COMPLAINT TRENDS ARE INVESTIGATED TO DETERMINE IF THERE IS A PRODUCT DEFECT/ DEFICIENCY. IF A PRODUCT DEFECT/ DEFICIENCY IS IDENTIFIED, A RISK EVALUATION IS COMPLETED AND COMPARED TO THE RISK MANAGEMENT REPORT, TO ENSURE THE RISK PROFILE HAS NOT CHANGED. ADDITIONALLY, AS A PART OF ABBOTT¿S POST-MARKET SURVEILLANCE PROCESS, ALL RISK EVALUATIONS WITH ASSOCIATED COMPLAINT DATA ARE REVIEWED ANNUALLY TO DETERMINE IF THE RISK PROFILES HAVE CHANGED AS COMPARED TO THE PRODUCT RISK MANAGEMENT REPORTS. THESE MONITORING PROCESSES ENSURE THAT ALL PRODUCT RISK PROFILES REMAIN ACCEPTABLE AND HAVE A POSITIVE BENEFIT/ RISK RATIO. IF PRODUCT IS RETURNED, A PHYSICAL INVESTIGATION WILL BE PERFORMED PER ADC'S ESTABLISHED PROCESSES AND PROCEDURES AND A REPORT WILL BE SUBMITTED UPON COMPLETION OF INVESTIGATION. ALL PERTINENT INFORMATION AVAILABLE TO ABBOTT DIABETES CARE HAS BEEN SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22834392&pc=DSQ | 2916596-2025-05201 | 2025-07-22 04:00 | Death | THORATEC CORPORATION | 2025-08-18 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Aortic Valve Insufficiency/ Regurgitation; Liver Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY WHILE INPATIENT. THE PATIENT WAS IMPLANT PRESENTED WITH SEVERE AORTIC INSUFFICIENCY AND, WORSENING LIVER CIRRHOSIS, AND ALTERED MENTAL STATUS. DUE TO THEIR FRAILTY, IT WAS DISCUSSED WITH THE FAMILY AS THERE WERE NO MEDICAL OR SURGICAL OPTIONS LEFT TO IMPROVE THEIR QUALITY OF LIFE, THE FAMILY MADE THE DECISION TO MOVE TOWARDS COMFORT MEASURES. THE PATIENT DID NOT HAVE ANY MALFUNCTIONS OF THE HEARTMATE 3 LEADING TO THE DECISION. THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) WAS WITHDRAWN PER THE REQUEST OF THE FAMILY. Manufacturer Narrative: A4 - PATIENT WEIGHT NOT PROVIDED. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22834548&pc=MFA | 3007284006-2025-00161 | 2025-08-05 04:00 | Death | SPECTRANETICS CORPORATION | 2025-08-18 04:00 | MFA | GLIDELIGHT LASER SHEATH | Adverse Event Without Identified Device or Use Problem | Perforation of Vessels; Pericardial Effusion | P960042 |  | 1 | Event Description: A LEAD EXTRACTION PROCEDURE COMMENCED TO REMOVE A RIGHT VENTRICULAR (RV) AND RIGHT ATRIAL (RA) LEAD DUE TO FRACTURE. DURING PREPARATION OF THE LEADS, INSULATION BREAKS ON BOTH LEADS WERE OBSERVED. A SPECTRANETICS LLD EZ LEAD LOCKING DEVICE (LLD EZ) WAS INSERTED ON THE RA LEAD, AND A SPECTRANETICS LLD #2 LEAD LOCKING DEVICE (LLD #2) ON THE RV LEAD TO PROVIDE TRACTION. BEGINNING WITH A SPECTRANETICS 14F GLIDELIGHT LASER SHEATH ON THE RV LEAD, ATTEMPTS WERE MADE TO EXTRACT THE LEAD, BUT RESISTANCE WAS ENCOUNTERED. THE DECISION WAS MADE TO SWITCH TO THE RA LEAD; HOWEVER, WHEN THE GLIDELIGHT WAS REMOVED, THE PATIENT¿S BLOOD PRESSURE DROPPED, AND A PERICARDIAL EFFUSION WAS DETECTED VIA TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE). RESCUE EFFORTS BEGAN, INCLUDING CPR, AND BLOOD PRODUCTS. A PERFORATION TO THE SUBCLAVIAN VEIN/INNOMINATE JUNCTION WAS DISCOVERED AND A VASCULAR SURGEON WAS CALLED INTO THE ROOM. THE PERFORATION WAS STENTED, BUT THE PATIENT¿S BLOOD PRESSURE WOULD NOT STABILIZE, AND CPR CONTINUED. DESPITE EXTENSIVE RESCUE EFFORTS, THE PATIENT DID NOT SURVIVE THE PROCEDURE. THIS REPORT CAPTURES THE 14F GLIDELIGHT IN USE WHEN THE PERFORATION OCCURRED, REQUIRING INTERVENTION, BUT RESULTING IN DEATH. THERE WAS NO ALLEGED MALFUNCTION OF ANY SPECTRANETICS DEVICES IN USE DURING THE PROCEDURE. Manufacturer Narrative: A2) PATIENT DATE OF BIRTH - NOT AVAILABLE FROM FACILITY. A3B) PATIENT GENDER - NOT AVAILABLE FROM FACILITY. A4) PATIENT WEIGHT - NOT AVAILABLE FROM FACILITY. A5) PATIENT ETHNICITY - NOT AVAILABLE FROM FACILITY. B7) OTHER RELEVANT HISTORY - NOT AVAILABLE FROM FACILITY. D4/H4) THE LOT NUMBER WAS NOT PROVIDED, THUS THE FOLLOWING INFORMATION IS UNKNOWN: DEVICE SERIAL NUMBER, UNIQUE ID, EXPIRATION DATE, AND MANUFACTURE DATE. H3) THE GLIDELIGHT WAS DISCARDED, THUS NO RETURNED PRODUCT INVESTIGATION WAS PERFORMED. H6) PERFORATION OF VESSELS AND DEATH ARE KNOWN RISKS OF COMPLICATION WITH USE OF THE GLIDELIGHT DEVICE. SUBMISSION OF THIS REPORT DOES NOT, IN ITSELF, REPRESENT A CONCLUSION BY THE MANUFACTURER AND/OR AUTHORIZED REPRESENTATIVE OR THE NATIONAL COMPETENT AUTHORITY THAT THE CONTENT OF THIS REPORT IS COMPLETE OR ACCURATE, THAT THE MEDICAL DEVICE(S) LISTED FAILED IN ANY MANNER AND/OR THAT THE MEDICAL DEVICE(S) CAUSED OR CONTRIBUTED TO AN ALLEGED DEATH OR DETERIORATION IN THE STATE OF THE HEALTH OF ANY PERSON.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22835167&pc=OZD | 1220648-2025-45756 | 2025-07-24 04:00 | Death | ABIOMED, INC. | 2025-08-18 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Ventricular Fibrillation | P140003 |  | 1 | Event Description: THE COMPLAINANT A PATIENT (UNKNOWN RACE) WITH HISTORY OF ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK REQUIRING EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) SUPPORT WAS BROUGHT TO THE OPERATING ROOM (OR) FOR ESCALATION OF VENOARTERIAL- VENOPULMONARY ECMO WITH AN IMPELLA CP (SERIAL NUMBER (B)(6)) UNLOADING TO UNLOADING WITH AN IMPELLA 5.5 (SERIAL NUMBER (B)(6)) FOR THE PURPOSE OF WEANING ECMO SUPPORT WITH THE INTENTION TO DECANNULATE THE FOLLOWING DAY. DURING THE INITIAL CUTDOWN FOR THE IMPELLA 5.5 AXILLARY GRAFT SITE, SOME BLEEDING COMPLICATIONS WERE ENCOUNTERED. THE GRAFT WAS PLACED SUCCESSFULLY, AND MULTIPLE "QUICK CLOT" METHODS WERE USED TO STITCH THE SITE. FOLLOWING IMPLANT, PROTAMINE WAS USED TO REVERSE THE EFFECTS OF HEPARIN AND BLOOD PRODUCTS WERE GIVEN TO COUNTERACT THE BLOOD LOSS. SUPPORT CONTINUED UNINTERRUPTED. THE PATIENT¿S LEFT VENTRICULAR EJECTION FRACTION RECOVERED TO 45-50%, AND THE PATIENT WAS TAKEN TO THE OR FOR THE REMOVAL OF THE IMPELLA 5.5. THE PATIENT WAS STABLE DURING THE WEAN AND THERE WERE NO ISSUES REPORTED AT EXPLANT. THE PATIENT REMAINED ON VENOPULMONARY EXTRACORPOREAL MEMBRANE OXYGENATION AFTER IMPELLA 5.5 WAS EXPLANTED. EIGHT DAYS LATER, THE PATIENT REQUIRED MECHANICAL CIRCULATORY SUPPORT TO ASSIST IN WEANING FROM ECMO SUPPORT. AN IMPELLA CP (SERIAL NUMBER (B)(6)) WAS PREPPED AND PRIMED, HOWEVER IMAGING REVEALED THE PATIENT HAD SEVERE THROMBUS AND DISSECTION AND LACK OF PERFUSION BELOW THE RIGHT ILIAC OSTIA TO THE COMMON FEMORAL ARTERY. AS THESE ARE CONTRAINDICATIONS FOR THE IMPELLA CP, THE CP WAS ABORTED, AND THE PATIENT WAS PREPARED TO BE TAKEN TO THE OR FOR AN IMPELLA 5.5 DEVICE. PRIOR TO GOING TO THE OPERATING ROOM, IT WAS NOTED THAT PATIENT¿S HEMOGLOBIN HAD DECREASED AND THROMBOELASTOGRAPHY ALSO SHOWED NEED FOR FRESH FROZEN PLASMA (FFP). ONE UNIT OF FFP AND 1 UNIT OF RED BLOOD CELLS (RBC) WERE TRANSFUSED PRIOR TO THE TRIP TO THE OR. DURING THE INITIAL CUTDOWN AND SUBSEQUENT SEWING OF THE GRAFT, THE PATIENT BECAME BRADYCARDIC, AND BLOOD PRESSURE BECAME UNREADABLE. CARDIOPULMONARY RESUSCITATION WAS INITIATED AND ADVANCED CARDIAC LIFE SUPPORT PROTOCOL ACTIVATED. THE PATIENT RECEIVED 1 ROUND OF EPINEPHRINE AND THREE AMPULES OF SODIUM BICARBONATE SOLUTION. THE RETURN OF SPONTANEOUS CIRCULATION ACHIEVED IN LESS THAN 4 MINUTES. THE PATIENT STABILIZED ON INCREASED PRESSER SUPPORT AND IMPELLA IMPLANT PROCEDURE COMPLETED FOLLOWING ABIOMED BEST PRACTICES. THE PATIENT REMAINED IN CRITICAL STATE ON MAX DOSE PRESSERS AND RECEIVED 4 MORE UNITS OF RBCS IN ATTEMPT TO CURB THE PATIENT¿S HYPOTENSIVE, VASOPLEGIC STATE. THE PATIENT WAS TRANSFERRED TO CARDIOVASCULAR ICU WITHOUT INCIDENT. IT WAS REPORTED THAT ON THE SECOND DAY OF SUPPORT WITH THE SECOND IMPELLA 5.5 DEVICE (SERIAL NUMBER (B)(6)), THE PATIENT EXPERIENCED VENTRICULAR FIBRILLATION ARREST WHILE REPOSITIONING IN BED. IT WAS NOTED THAT THE PATIENT WAS HAVING COMPLICATIONS WITH ECMO SUPPORT WITH SUSPECTED LYSIS FROM THE ECMO DEVICE. THE PATIENT CONTINUED TO DETERIORATE DESPITE INCREASING SUPPORT AND BEGAN SHOWING SIGNS OF MULTIPLE ORGAN FAILURE. THE PATIENT WAS UNABLE TO BE WEANED OFF OF ECMO SUPPORT, WAS NO LONGER RESPONDING TO MAX PRESSORS, AND WAS EXPERIENCING INTERMITTENT ARRHYTHMIAS. THE PATIENT¿S FAMILY DECIDED TO WITHDRAW CARE. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: ALTHOUGH THE PATIENT'S PRESENTING CONDITION MAY BE MORE LIKELY HAVE IMPACTED THE EVENT THE IMPELLA CANNOT BE EXCLUDED AS A POSSIBLE CONTRIBUTING FACTOR IN THE PATIENT'S DEMISE. RELATED REPORT NUMBERS: THIS MDR SPECIFICALLY ADDRESSES PUMP 4, THE IMPELLA 5.5 WITH SERIAL NUMBER (B)(6), WHICH IS ONE OF THREE PUMPS ASSOCIATED WITH COMPLAINTS INVOLVING THE REPORTED EVENT. REFER TO RELATED REPORT NUMBERS FOR THE REPORT ON THE OTHER TWO PUMPS. DEVICE STATUS: THE IMPELLA WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22835266&pc=DSQ | 2916596-2025-05469 | 2025-07-24 04:00 | Death | THORATEC CORPORATION | 2025-08-18 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Renal Failure; Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT WAS CRITICALLY ILL ON EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO), INTUBATION, DOBUTAMINE, EPINEPHRINE, LEVOPHED; VASOPRESSIN, CONTINUOUS RENAL REPLACEMENT THERAPY (CRRT), AND INFECTION WORKUP. LOG FILES WERE SUBMITTED FOR REVIEW. THE EVENT LOG CAPTURED 3 LOW FLOW ALARMS AS WELL AS BRIEF LOW FLOW ESTIMATES WHEN THE CALCULATED PULSATILITY INDEX (PI) WAS ELEVATED ON (B)(6) 2025 FROM 19:44 TO 19:59. THE ESTIMATED FLOW WAS FLUCTUATING BELOW 2.5 LPM THRESHOLD. MOST OF THESE EVENTS WERE BRIEF AND DIDN¿T GENERATE THE ALARMS (LESS THAN 10 SECONDS TO TRIGGER THE ALARMS). THE ESTIMATED FLOWS WERE AVERAGING FROM 1.2 TO 2.4 LPM AND PI WAS AVERAGING FROM 3.7 TO 8.5 DURING THESE EVENTS. THERE WERE NO UNUSUAL ROTOR FAULTS, PUMP TEMPERATURE, OR ANY OTHER ABNORMAL PUMP FAULTS WERE SEEN IN THE EVENT LOG. THE EVENT LOG PROCEEDED TO CAPTURE A LOW POWER ADVISORY (F13) ON THE BLACK POWER CABLE ON (B)(6) 2025 AT 04:41. THE EVENT LOG FILE ALSO CAPTURED FREQUENT PI EVENTS ON 26JUL2025 FROM 08:32 TO 14:39. IT WAS LATER REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTISYSTEM ORGAN FAILURE ON (B)(6) 2025. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22835373&pc=DSQ | 2916596-2025-05461 | 2025-03-09 05:00 | Death | THORATEC CORPORATION | 2025-08-18 04:00 | DSQ | HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM | Adverse Event Without Identified Device or Use Problem | Respiratory Arrest | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO RESPIRATORY ARREST. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED AND THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: SECTION D4: LOT NUMBER WAS NOT PROVIDED, AND EXPIRATION DATE AND MANUFACTURE DATE (H4) CANNOT BE DETERMINED. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22835701&pc=MKJ | 0003015876-2025-01667 | 2025-07-14 04:00 | Death | PHYSIO-CONTROL, INC. - 3015876 | 2025-08-18 04:00 | MKJ | LIFEPAK® CR2 DEFIBRILLATOR | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P170018 |  | 1 | Event Description: THE CUSTOMER CONTACTED STRYKER TO REPORT THEIR DEVICE DISPLAYED THE MESSAGE "PLACE PADS ON PATIENT" WHEN THE PADS WERE ALREADY PLACED ON THE PATIENT. IN THIS STATE THE DEVICE MAY NOT BE ABLE TO DELIVER DEFIBRILLATION THERAPY IF NEEDED. ANOTHER DEVICE WAS USED. THE PATIENT ASSOCIATED WITH THE REPORTED EVENT DID NOT SURVIVE. THERE HAS BEEN NO ALLEGED DEVICE CONTRIBUTION TO THE PATIENT OUTCOME.Event Description: THE CUSTOMER CONTACTED STRYKER TO REPORT THEIR DEVICE DISPLAYED THE MESSAGE "PLACE PADS ON PATIENT" WHEN THE PADS WERE ALREADY PLACED ON THE PATIENT. IN THIS STATE THE DEVICE MAY NOT BE ABLE TO DELIVER DEFIBRILLATION THERAPY IF NEEDED. ANOTHER DEVICE WAS USED. THE PATIENT ASSOCIATED WITH THE REPORTED EVENT DID NOT SURVIVE. THERE HAS BEEN NO ALLEGED DEVICE CONTRIBUTION TO THE PATIENT OUTCOME. Manufacturer Narrative: STRYKER EVALUATED THE CUSTOMER'S DEVICE AND WAS ABLE TO VERIFY THE REPORTED ISSUE. REVIEW OF THE SOFTWARE LOG SHOWS THE DEVICE CYCLING THROUGH THE VOICE PROMPTS INSTRUCTING THE USER TO APPLY THE PADS TO THE PATIENT'S CHEST WITH THE DEVICE BRIEFLY DETECTING A PATIENT LOAD OF 35.6 OHMS RESISTANCE AND 217.1 OHMS REACTANCE, THEN WENT BACK UP TO 123.6 OHMS RESISTANCE AND 604.9 OHMS REACTANCE, WHICH IS CONSISTENT WITH THE PADS NOT MAKING GOOD CONNECTION TO THE PATIENT'S SKIN DUE TO THE CLEAR PLASTIC FILM REMAINING ON THE PAD AS INDICATED BY THE CUSTOMER. THE ROOT CAUSE IS DETERMINED TO BE USER ERROR. THIS DEVICE WAS ARCHIVED BY STRYKER. Manufacturer Narrative: THE CUSTOMER HAS BEEN SENT A REPLACEMENT DEVICE. THE CUSTOMER REPORTED THE PADS HAD NOT BEEN OPENED PROPERLY. A CLINICAL REVIEW WAS PERFORMED AND THE DEVICE CONTRIBUTION COULD NOT BE DETERMINED. STRYKER CONTINUES TO INVESTIGATE THE REPORTED FAILURE AND WILL SUBMIT A SUPPLEMENTAL REPORT ON THIS EVENT TO THE FDA AS PROVIDED BY 21 CFR 803.56. STRYKER CONTACTED THE CUSTOMER TO REQUEST ADDITIONAL INFORMATION ON THE PATIENT. STRYKER REQUESTS PATIENT INFORMATION NECESSARY FOR REGULATORY COMPLIANCE, STRICTLY ADHERING TO HIPAA'S PRIVACY RULE. NO RESPONSE HAS BEEN RECEIVED FROM THE CUSTOMER. PATIENT FIELDS IN WHICH INFORMATION IS NOT PROVIDED WERE INTENTIONALLY LEFT BLANK.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22835841&pc=MHY | 3006630150-2025-06636 | 2025-07-24 04:00 | Death | BOSTON SCIENTIFIC NEUROMODULATION CORPORATION | 2025-08-18 04:00 | MHY | VERCISE? CARTESIA? | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P150031 |  | 1 | Event Description: IT WAS REPORTED THAT THE DEEP BRAIN STIMULATION (DBS) PATIENT UNDERWENT AN IMPLANT SURGERY. FOLLOWING THE PROCEDURE, THE PATIENT DEVELOPED A POST-OPERATIVE HEMORRHAGE. DUE TO AN ADVANCED DIRECTIVE IN PLACE, WHICH PROHIBITS LIFE-SAVING SURGICAL INTERVENTIONS, THE PATIENT UNFORTUNATELY PASSED AWAY. THE SURGEON HAS CLEARLY STATED THAT THE EVENT WAS NOT RELATED TO THE DEVICE, AND IT IS UNCERTAIN WHETHER IT WAS RELATED TO THE PROCEDURE. Manufacturer Narrative: BLOCK D.2B; ADDITIONAL APPLICABLE PRODUCT CODES: NHL, PJS. ADDITIONAL SUSPECT MEDICAL DEVICE COMPONENTS INVOLVED IN THE EVENT: PRODUCT FAMILY: DBS-LINEAR LEADS UPN: M365DB2202300 MODEL: DB-2202-30 SERIAL: (B)(6) BATCH: 7086026 UDI: (B)(4). PRODUCT FAMILY: DBS-LEAD FIXATION UPN: M365DB4600C0 MODEL: DB-4600-C BATCH: 36122989 UDI: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22835944&pc=NPU | 2015691-2025-06793 | 2025-04-10 04:00 | Death | EDWARDS LIFESCIENCES | 2025-08-18 04:00 | NPU | SAPIEN 3 TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Shock; Obstruction/Occlusion | P140031 |  | 1 | Event Description: AS REPORTED FROM PATIENT SUPPORT, ACCORDING TO THE CARE ADVOCATE, THE PATIENT UNDERWENT TRANSCATHETER MITRAL VALVE REPLACEMENT (TMVR) PROCEDURE AND "DIED AFTER THE PROCEDURE." IT WAS FURTHER EXPLAINED THAT THE PATIENT WAS KEPT ON LIFE SUPPORT FOLLOWING THE PROCEDURE UNTIL OFFICIAL DATE OF DEATH. THE EXACT CAUSE OF DEATH WAS NOT PROVIDED HOWEVER IT WAS MENTION THAT "HEART FAILURE IS NUMBER ONE". THERE WERE NO ALLEGATIONS MADE AGAINST ANY EDWARDS DEVICES. PER MEDICAL RECORD REVIEW, THE PATIENT EXPIRED 1 DAY POST IMPLANTATION OF A 29MMM S3 VALVE IN THE MITRAL POSITION VIA TRANSSEPTAL PUNCTURE AS A VALVE-IN-MITRAL ANNULAR CALCIFICATION (VIMAC). THE PATIENT UNDERWENT A SUCCESSFUL TRANSCATHETER MITRAL VALVE REPLACEMENT WITH NO DOCUMENTED COMPLICATION. POSTOPERATIVELY, THE PATIENT DEVELOPED CARDIOGENIC SHOCK SECONDARY TO NEW-ONSET LEFT VENTRICULAR OUTFLOW TRACT (LVOT) OBSTRUCTION. ECHOCARDIOGRAPHY REVEALED SEVERE LEFT VENTRICULAR SYSTOLIC DYSFUNCTION (LVEF 15% BY VISUAL ESTIMATION) AND A PEAK LVOT GRADIENT OF 70 MMHG, WITH THE PROSTHETIC VALVE ABUTTING THE INTERVENTRICULAR SEPTUM. EMERGENT ALCOHOL SEPTAL ABLATION AND INTRA-AORTIC BALLOON PUMP (IABP) PLACEMENT WERE PERFORMED, RESULTING IN A REDUCED PEAK LVOT GRADIENT OF 49 MMHG. DESPITE MAXIMAL INOTROPIC SUPPORT, THE PATIENT'S CONDITION DETERIORATED IN THE ICU AND RESULTED IN DEATH WITHIN HOURS. Manufacturer Narrative: LVOT OBSTRUCTION IN PATIENTS WHO UNDERGO TRANSCATHETER MITRAL VALVE REPLACEMENT IS A WELL-RECOGNIZED POSTOPERATIVE COMPLICATION. IN MOST CASES OF POSTOPERATIVE LVOT OBSTRUCTION, THE OBSTRUCTION RESULTS FROM THE PROTRUSION OF THE VALVE INTO THE LVOT OR FROM ABNORMAL SUBVALVULAR POSITIONING OF THE PROSTHESIS. ADDITIONALLY, LVOT OBSTRUCTION MAY OCCUR POSTOPERATIVELY IF THERE IS A NARROWED MITRAL AORTIC ANGLE, OR DUE TO A THICKENED INTERVENTRICULAR SEPTUM. OTHER POSSIBLE CAUSES INCLUDE A HYPER CONTRACTILE LEFT VENTRICLE OR ATRIAL FIBRILLATION. OBSTRUCTION OF THE LEFT VENTRICULAR OUTFLOW TRACT CAN ALSO BE CAUSED BY PATIENT FACTORS (ANTERIOR MITRAL LEAFLET PROTRUDING INTO THE LVOT) OR PROCEDURAL FACTORS (POSITIONING OF THE VALVE FRAME WITHIN THE ANNULUS). DEPENDING ON THE DEGREE OF OBSTRUCTION, CARDIAC OUTPUT AND HEMODYNAMIC STABILITY MAY BE AFFECTED. INACCURATE DEPLOYMENT IS GENERALLY A RESULT OF USE ERROR OR A COMBINATION OR PATIENT AND PROCEDURAL FACTORS. IN SOME CASES, LVOT OBSTRUCTION COULD RESULT IN CLINICALLY SIGNIFICANT HEMODYNAMIC COMPROMISE THAT MAY REQUIRE EXPLANATION OF THE THV WITH SURGICAL CORRECTION. THE THV TRAINING MANUALS INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN THV. TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS AND PROCTORED PROCEDURES. THE CORRECT ALIGNMENT AND POSITIONING OF THE DEVICE AT THE POINT OF DEPLOYMENT IS EMPHASIZED AS A KEY FACTOR TO THE PLACEMENT AND FIXATION OF THE DEVICE. OPERATORS ARE ALSO INSTRUCTED TO USE FLUOROSCOPY AS THE PRIMARY METHOD OF VISUALIZATION FOR POSITIONING AND DEPLOYMENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION THAT A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST/INDICATE SEVERE MITRAL ANNULAR CALCIFICATION, NONALCOHOLIC STEATOHEPATITIS, AND HYPOGAMMAGLOBULINEMIA COULD HAVE CONTRIBUTED TO THE LVOT OBSTRUCTION. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22813019&pc=MNS | 2518422-2025-109139 | 2025-08-12 04:00 | Death | RESPIRONICS, INC. | 2025-08-15 04:00 | MNS | DREAMSTATION ST30 | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K102465 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH A DREAMSTATION ST30 HAS PASSED AWAY. THE REPORTED EVENT MAKES NO ALLEGATION OF ANY SPECIFIC DEVICE MALFUNCTION, USE ERROR, FAILURE, LABELING, OR OTHER DEFICIENCY THAT IS RELEVANT TO THE HARMS REPORTED; THEREFORE, BASED ON INFORMATION AVAILABLE AT THIS TIME, THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO A SERIOUS INJURY OR DEATH. GOOD FAITH EFFORT TO OBTAIN ADDITIONAL INFORMATION WAS UNABLE TO BE COMPLETED AS THERE WAS NO CUSTOMER CONTACT INFORMATION PROVIDED. A FINAL REPORT IS BEING SUBMITTED. IF NEW INFORMATION BECOMES AVAILABLE FOLLOWING THE COMPLETION OF THE DEVICE REPAIR THAT CHANGES THE OUTCOME OF THE INVESTIGATION AND ASSOCIATED MEDICAL DEVICE REPORTING A FOLLOW UP/SUPPLEMENTAL REPORT WILL BE COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22813883&pc=MVK | 3008642652-2025-07878 | 2025-06-14 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-08-15 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED AN APPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 14:10:58 ON (B)(6) 2025. AT 14:33:32, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VT @ 230 BPM. AT 14:34:02, THE PATIENT RECEIVED THE APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 270 BPM WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY FOR 1 MINUTE AND 30 SECONDS BEFORE TRANSITIONING TO SINUS BRADYCARDIA/AF FROM 20-60 BPM. POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY. THE ELECTRODE BELT WAS DISCONNECTED AT 18:43:17 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22813884&pc=MVK | 3008642652-2025-07877 | 2025-07-14 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-08-15 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED AN INAPPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 15:30:55 ON (B)(6) 2025. AT 18:20:46, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH MOTION ARTIFACT. AT 18:21:56, THE PATIENT RECEIVED THE INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON LIFE-SUSTAINING RHYTHM. THE ELECTRODE BELT WAS DISCONNECTED AT 18:26:16 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22813886&pc=MVK | 3008642652-2025-07879 | 2025-07-27 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-08-15 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED THREE APPROPRIATE TREATMENTS AND AN INAPPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 20:39:53 ON (B)(6) 2025. AT 21:34:24, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF WITH CPR/MOTION ARTIFACT. AT 21:35:33, THE PATIENT RECEIVED THE FIRST APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS ASYSTOLE WITH CPR/MOTION ARTIFACT. POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY. AT 21:36:10, THE PATIENT RECEIVED THE SECOND APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS VF WITH CPR/MOTION ARTIFACT. AT 21:36:47, THE PATIENT RECEIVED THE THIRD APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT. POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY. AT 21:37:23, THE PATIENT RECEIVED THE INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS SINUS TACHYCARDIA @ 100 BPM WITH PVCS AND CPR/MOTION ARTIFACT. THE ELECTRODE BELT WAS DISCONNECTED AT 21:39:02 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR HAS BEEN COMPLETED. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22813889&pc=MVK | 3008642652-2025-07880 | 2025-07-24 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-08-15 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON 7/24/2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED AN INAPPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 13:55:16 ON 7/23/2025. AT 07:57:47 ON 7/24/2025, AN ARRHYTHMIA WAS DETECTED. ECG WAS OBSCURED DUE TO CPR/ELECTRODE LEAD FALL OFF AND INTERMITTENT CARDIAC ACTIVITY. AT 07:58:57, THE PATIENT RECEIVED THE INAPPROPRIATE TREATMENT. CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS ASYSTOLE WITH UNCONDUCTED P WAVES. THE ELECTRODE BELT WAS DISCONNECTED AT 07:59:01 ON 7/24/2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR HAS BEEN COMPLETED. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION. THE ELECTRODE BELT HAS BEEN RETURNED AND THE EVALUATION IS UNDERWAY. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22814030&pc=KFM | 8010762-2025-0000355 | 2025-07-27 04:00 | Death | MAQUET CARDIOPULMONARY GMBH | 2025-08-15 04:00 | KFM | HEART LUNG MACHINE | Battery Problem | Insufficient Information | K991864 |  | 1 | Event Description: THE EVENT OCCURRED IN CHINA. ON (B)(6) 2025, THE PATIENT WAS A 9-MONTH-OLD MALE INFANT WEIGHING 8.9KG. PATIENT WAS CONNECTED TO THE ECMO MACHINE. DURING THE TRANSFER FROM THE OPERATING ROOM TO THE CICU, THE CENTRIFUGAL PUMP TRIGGERED AN ALARM OF LOW BAT. IMMEDIATELY, THE EMERGENCY HAND CRANK WAS USED TO PUMP BLOOD AND THE TRANSFER CONTINUED. IMMEDIATELY AFTER THE CHILD WAS TRANSFERRED TO THE CICU, A NEW PUMP WAS REPLACED FOR THE CHILD, AND THE PUMP REPLACEMENT TOOK APPROXIMATELY 30 SECONDS. SUBSEQUENTLY, THE REPLACED PUMP WAS CONTINUED TO BE USED TO TREAT THE CHILD PATIENT. THE EQUIPMENT DID NOT MALFUNCTION. THE CHILD PATIENT WAS SUCCESSFULLY REMOVED FROM ECMO ON (B)(6) 2025. 48 HOURS AFTER THE WEANING, THE CHILD PATIENT DEVELOPED SEVERE COMPLICATIONS OF CENTRAL NERVOUS SYSTEM SEPTIC SHOCK. THE FAMILY GAVE UP TREATMENT AND THE PATIENT DIED. THE CUSTOMER STATES THAT THE INVOLVED BATTERY NEVER BEEN INSPECTED OR REPLACED. FURTHER THERE WAS NO RELATION BETWEEN THIS INCIDENT AND THE PATIENT OUTCOME (EXPIRATION). DUE TO THE REPORTED DEATH OF THE PATIENT A REPORT IS REQUIRED. COMPLAINT ID: (B)(4). Manufacturer Narrative: THE INVESTIGATION IS ONGOING. FURTHER INFORMATION HAS BEEN REQUESTED BUT HAS NOT YET BEEN RECEIVED. A FOLLOW UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. NOTE: THIS EVENT OCCURRED ON THE CHINESE MARKET. IT IS A SIGNIFICANTLY SIMILAR DEVICE TO "BASE UNIT, ROTAFLOW CONSOLE¿ WHICH IS SOLD IN THE USA UNDER PREMARKET SUBMISSION NUMBER K991864. FOR SECTION D4 ALL AVAILABLE IDENTIFYING INFORMATION IS FOR THE OUT OF THE US DEVICE, SUBJECTED TO THIS REPORT. THEREFORE, NO GUDID INFORMATION EXISTS. FURTHERMORE, UDI INFORMATION (PRIMARY DI NUMBER) PROVIDED IN D4 IS FOR ROTAFLOW CONSOLE WITH CATALOG NUMBER 701046405.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22814124&pc=NPT | 9617601-2025-00995 | 2024-06-13 04:00 | Death | MEDTRONIC MEXICO S. DE R.L. DE CV | 2025-08-15 04:00 | NPT | MEDTRONIC TRANSCATHETER VALVE | Inadequacy of Device Shape and/or Size; Patient Device Interaction Problem | Insufficient Information | P130021 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING MULTIMODAL IMAGING ASSESSMENT TO PREDICT PROSTHESIS-PATIENT MISMATCH (PPM) IN PATIENTS UNDERGOING VALVE-IN-VALVE TRANSCATHETER AORTIC VALVE IMPLANTATION (VIV-TAVI). THE STUDY POPULATION CONSISTED OF 40 PATIENTS WITH A DEGENERATED SURGICAL VALVE WHO ALL UNDERWENT VIV-TAVI WITH A TRANSCATHETER VALVE FROM THE MEDTRONIC EVOLUT FAMILY. FOUR OF THESE PATIENTS HAD A DEGENERATED MEDTRONIC SURGICAL VALVE (MOSAIC = 3, HANCOCK II = 1). PRIOR TO VIV-TAVI, HEART FAILURE, HIGH MEAN GRADIENTS, AND MODERATE TO SEVERE AORTIC REGURGITATION WERE NOTED AMONG THE PATIENTS IN THE STUDY. ADVERSE OUTCOMES OF VIV-TAVI INCLUDED: CORONARY OCCLUSION WITH PERCUTANEOUS INTERVENTION, VALVE EMBOLIZATION NECESSITATING THE IMPLANT OF A SECOND TRANSCATHETER VALVE, TRANSIENT ISCHEMIC ATTACK/STROKE, MODERATE TO SEVERE AORTIC REGURGITATION, AND ELEVATED MEAN GRADIENTS. ADDITIONALLY, THE AUTHORS FOUND THAT POST-PROCEDURAL SEVERE PPM OCCURRED IN 45% OF THE STUDY POPULATION AND WAS ASSOCIATED WITH A HIGHER RATE OF ALL-CAUSE DEATHS AND CARDIOVASCULAR REHOSPITALIZATIONS (VALVE- OR PROCEDURE-RELATED). Manufacturer Narrative: CITATION: BIANCHINI F, ROMAGNOLI E, AURIGEMMA C, ET AL. A MULTIMODAL APPROACH TO PREDICT PROSTHESIS-PATIENT MISMATCH IN PATIENTS UNDERGOING VALVE-IN-VALVE TRANS-CATHETER AORTIC VALVE IMPLANTATION. CARDIOVASC REVASC MED. 2025;70:41-47. DOI:10.1016/J.CARREV.2024.06.012 EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT AND DATE OF DEATH. EARLIEST APPROVED EVOLUT PRODUCT USED FOR PRODUCT CODE AND PMA#. NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22814247&pc=NSA | 3030677-2025-001867 | 2025-08-08 04:00 | Death | PHILIPS NORTH AMERICA LLC | 2025-08-15 04:00 | NSA | HEARTSTART | Failure to Deliver Shock/Stimulation | Cardiac Arrest | P160029 |  | 1 | Event Description: IT HAS BEEN REPORTED THAT THE DEVICE DID NOT SHOCK DURING A PATIENT USE CASE. THE PATIENT PASSED AWAY WITHIN 10-15 MINUTES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22814585&pc=BZD | 2518422-2025-109142 | 2025-08-12 04:00 | Death | RESPIRONICS, INC. | 2025-08-15 04:00 | BZD | DREAMSTATION HUMIDIFIER, GB-RENT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH A DREAMSTATION HUMIDIFIER HAS PASSED AWAY. THE REPORTED EVENT MAKES NO ALLEGATION OF ANY SPECIFIC DEVICE MALFUNCTION, USE ERROR, FAILURE, LABELING, OR OTHER DEFICIENCY THAT IS RELEVANT TO THE HARMS REPORTED; THEREFORE, BASED ON INFORMATION AVAILABLE AT THIS TIME, THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO A SERIOUS INJURY OR DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. Manufacturer Narrative: AFTER FURTHER REVIEW, REASSESSMENT OF THIS CASE HAS DETERMINED THIS REPORT IS CONSIDERED TO BE A DUPLICATE/CONTINUATION COMPLAINT: N/A - ISSUE WAS CONFIRMED TO BE A DUPLICATE OR CONTINUATION OF A PREVIOUS CASE. THE BASE DEVICE - DREAMSTATION ST30 - WAS REPORTED UNDER (B)(4)/RA320620251. THIS CASE IS BEING CLOSED/CANCELLED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22814739&pc=KFM | 3013876692-2025-00069 | 2025-07-27 04:00 | Death | MAQUET CARDIOPULMONARY GMBH | 2025-08-15 04:00 | KFM | HEART LUNG MACHINE | Battery Problem | Insufficient Information | K991864 |  | 1 | Event Description: THE EVENT OCCURRED IN CHINA. ON (B)(6) 2025, THE PATIENT WAS A 9-MONTH-OLD MALE INFANT WEIGHING 8.9KG. PATIENT WAS CONNECTED TO THE ECMO MACHINE. DURING THE TRANSFER FROM THE OPERATING ROOM TO THE CICU, THE CENTRIFUGAL PUMP TRIGGERED AN ALARM OF LOW BAT. IMMEDIATELY, THE EMERGENCY HAND CRANK WAS USED TO PUMP BLOOD AND THE TRANSFER CONTINUED. IMMEDIATELY AFTER THE CHILD WAS TRANSFERRED TO THE CICU, A NEW PUMP WAS REPLACED FOR THE CHILD, AND THE PUMP REPLACEMENT TOOK APPROXIMATELY 30 SECONDS. SUBSEQUENTLY, THE REPLACED PUMP WAS CONTINUED TO BE USED TO TREAT THE CHILD PATIENT. THE EQUIPMENT DID NOT MALFUNCTION. THE CHILD PATIENT WAS SUCCESSFULLY REMOVED FROM ECMO ON (B)(6) 2025. 48 HOURS AFTER THE WEANING, THE CHILD PATIENT DEVELOPED SEVERE COMPLICATIONS OF CENTRAL NERVOUS SYSTEM SEPTIC SHOCK. THE FAMILY GAVE UP TREATMENT AND THE PATIENT DIED. THE CUSTOMER STATES THAT THE INVOLVED BATTERY NEVER BEEN INSPECTED OR REPLACED. FURTHER THERE WAS NO RELATION BETWEEN THIS INCIDENT AND THE PATIENT OUTCOME (EXPIRATION). DUE TO THE REPORTED DEATH OF THE PATIENT A REPORT IS REQUIRED. COMPLAINT ID: (B)(4). Manufacturer Narrative: THE INVESTIGATION IS ONGOING. FURTHER INFORMATION HAS BEEN REQUESTED BUT HAS NOT YET BEEN RECEIVED. A FOLLOW UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. NOTE: THIS EVENT OCCURRED ON THE CHINESE MARKET. IT IS A SIGNIFICANTLY SIMILAR DEVICE TO "BASE UNIT, ROTAFLOW CONSOLE¿ WHICH IS SOLD IN THE USA UNDER PREMARKET SUBMISSION NUMBER K991864. FOR SECTION D4 ALL AVAILABLE IDENTIFYING INFORMATION IS FOR THE OUT OF THE US DEVICE, SUBJECTED TO THIS REPORT. THEREFORE, NO GUDID INFORMATION EXISTS. FURTHERMORE, UDI INFORMATION (PRIMARY DI NUMBER) PROVIDED IN D4 IS FOR ROTAFLOW CONSOLE WITH CATALOG NUMBER 701046405.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22815315&pc=BZD | 2518422-2025-109148 | 2025-08-04 04:00 | Death | RESPIRONICS, INC. | 2025-08-15 04:00 | BZD | DREAMSTATION AUTO CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT HAD STAGE 3 LUNG CANCER JUST FINISHED TREATMENT ON (B)(6) 2025, HAD RADIATION EVERY DAY AND CHEMO EVERY WEEK. WENT TO THE HOSPITAL HIS STATS WAS 70 HE WENT TO THE ER AND WAS TOLD HE HAS COMMUNITY AQUIRED PNEMONIA, HE WAS PLACED ON A VENTILATOR AND STAYED ON THE VENT UNTIL (B)(6) 2025 LABS LOOKED GREAT. PATIENT WAS HOSPITALIZED AND PASSED AWAY ON VENTILATOR SUPPORT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22815675&pc=BZD | 2518422-2025-109150 | 2025-08-12 04:00 | Death | RESPIRONICS, INC. | 2025-08-15 04:00 | BZD | REMSTAR SYSTEM ONE 60 SERIES | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K113068 |  | 1 | Event Description: MANUFACTURER RECEIVED A CALL FROM THE PATIENT'S DAUGHTER STATING THAT HER FATHER PASSED AWAY IN 2023 AND HER STEPMOM THREW THE DEVICE AWAY, SHE NEEDS TO FIND THE S/N FOR THE DEVICE. THE CALLER WAS TRANSFERRED TO PATIENT DME. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE. REPORTABLE SINCE DEVICE ISSUE/EVENT HAS OR MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22815690&pc=NGV | 2135147-2025-04552 | 2009-07-01 04:00 | Death | ABBOTT MEDICAL | 2025-08-15 04:00 | NGV | AMPLATZER AMULET | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P200049 |  | 1 | Event Description: THE ARTICLE, "DISCONTINUATION OF ANTITHROMBOTIC MANAGEMENT FOLLOWING LEFT ATRIAL APPENDAGE OCCLUSION IN HIGH BLEEDING RISK PATIENTS", WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE, SINGLE CENTER STUDY ON THE EFFICACY OF USING MAXIMUM OF 6-MONTH ANTITHROMBOTIC TREATMENT POST-LEFT ATRIAL APPENDAGE OCCLUSION (LAAO) IN HIGH BLEEDING RISK PATIENTS WITH ATRIAL FIBRILLATION. DEVICES INCLUDED IN THE STUDY WERE AMULET (ABBOTT), OMEGA (ECLIPSE MEDICAL), LAMBRE (LIFETECH SCIENTIFIC), WATCHMAN (BOSTON SCIENTIFIC), AND COHEREX (COHEREX MEDICAL INC.). THE ARTICLE CONCLUDED THAT COMPLETE CESSATION OF ANTITHROMBOTIC TREATMENT 6 MONTHS AFTER LAAO WAS ASSOCIATED WITH A LOW INCIDENCE OF SUBSEQUENT STROKE OR TIA IN PATIENTS AT HIGH BLEEDING RISK. [(B)(6)]. THE TIME FRAME OF THE STUDY WAS FROM FEBRUARY 2015 TO JANUARY 2024. A TOTAL OF 128 PATIENTS WERE INCLUDED IN THE STUDY, OF WHICH IT WAS NOT SPECIFIED HOW MANY RECEIVED AN ABBOTT DEVICE. IT WAS REPORTED AMULET, OMEGA, AND LAMBRE COMBINED ACCOUNTED FOR 101 (78.9%) PATIENTS. THE AVERAGE AGE WAS 77 YEARS AND THE MAJORITY GENDER WAS MALE. COMORBIDITIES INCLUDED PRIOR BLEEDING, PROPHYLAXIS, DYSPNEA, ATRIAL FIBRILLATION, AND CEREBROVASCULAR DISEASE (ISCHEMIC, HEMORRHAGIC). Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF DISCONTINUATION OF ANTITHROMBOTIC MANAGEMENT FOLLOWING LEFT ATRIAL APPENDAGE OCCLUSION IN HIGH BLEEDING RISK PATIENTS WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING PRIOR BLEEDING, PROPHYLAXIS, DYSPNEA, ATRIAL FIBRILLATION, AND CEREBROVASCULAR DISEASE (ISCHEMIC, HEMORRHAGIC). SOME OF THE COMPLICATIONS REPORTED WERE HEMORRHAGE, THROMBOSIS, STROKE, DEATH; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. B2: DATE OF DEATH IS ESTIMATED. B3: DATE OF EVENT IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: DISCONTINUATION OF ANTITHROMBOTIC MANAGEMENT FOLLOWING LEFT ATRIAL APPENDAGE OCCLUSION IN HIGH BLEEDING RISK PATIENTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22815722&pc=NKM | 2135147-2025-04554 | 2025-07-16 04:00 | Death | ABBOTT MEDICAL | 2025-08-15 04:00 | NKM | MITRACLIP | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage | P100009 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A MITRACLIP PROCEDURE WAS PERFORMED TO TREAT FUNCTIONAL MITRAL REGURGITATION (MR) GRADE 4+. PATIENT HAD DILATED CARDIOMYOPATHY, LEFT VENTRICULAR END-SYSTOLIC DIAMETER 70MM, AND POSTERIOR LEAFLET RESTRICTION. A XTR CLIP WAS IMPLANTED REDUCING MR TO GRADE 1+. FIVE DAYS AFTER THE PROCEDURE (B)(6) 2025, THE PATIENT DIED FROM CEREBRAL HEMORRHAGE. THERE WERE NO REPORTED DEVICE ISSUE WITH THE XTR OR THE STEERABLE GUIDE CATHETER DURING THE INITIAL PROCEDURE. POST PROCEDURE THE CLIP WAS REPORTED TO REMAIN STABLE ON BOTH LEAFLETS AND NO TISSUE DAMAGE WAS OBSERVED. THE CAUSE OF THE CEREBRAL HEMORRHAGE WAS REPORTED TO BE UNKNOWN.Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A MITRACLIP PROCEDURE WAS PERFORMED TO TREAT FUNCTIONAL MITRAL REGURGITATION (MR) GRADE 4+. PATIENT HAD DILATED CARDIOMYOPATHY, LEFT VENTRICULAR END-SYSTOLIC DIAMETER 70MM, AND POSTERIOR LEAFLET RESTRICTION. A XTR CLIP WAS IMPLANTED REDUCING MR TO GRADE 1+. FIVE DAYS AFTER THE PROCEDURE (B)(6) 2025, THE PATIENT DIED FROM CEREBRAL HEMORRHAGE. THERE WERE NO REPORTED DEVICE ISSUE WITH THE XTR OR THE STEERABLE GUIDE CATHETER DURING THE INITIAL PROCEDURE. POST PROCEDURE THE CLIP WAS REPORTED TO REMAIN STABLE ON BOTH LEAFLETS AND NO TISSUE DAMAGE WAS OBSERVED. THE CAUSE OF THE CEREBRAL HEMORRHAGE WAS REPORTED TO BE UNKNOWN. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A REVIEW OF THE LOT HISTORY RECORD REVEALED NO MANUFACTURING NONCONFORMITIES ISSUED TO THE REPORTED LOT THAT WOULD HAVE CONTRIBUTED TO THIS EVENT. BASED ON AVAILABLE INFORMATION, THE REPORTED DEATH APPEARS TO BE RELATED TO INTRACRANIAL HEMORRHAGE. HOWEVER, A CAUSE FOR THE REPORTED INTRACRANIAL HEMORRHAGE COULD NOT BE DETERMINED. THE REPORTED PATIENT EFFECT OF DEATH AND INTRACRANIAL HEMORRHAGE, AS LISTED IN THE MITRACLIP SYSTEM INSTRUCTIONS FOR USE, ARE KNOWN POSSIBLE COMPLICATION ASSOCIATED WITH MITRACLIP PROCEDURES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN, OR LABELING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22816117&pc=MCX | 3004742232-2025-00239 | 2025-07-24 04:00 | Death | CARDIOVASCULAR SYSTEMS INCORPORATED | 2025-08-15 04:00 | MCX | DIAMONDBACK CORONARY | Material Separation | Perforation of Vessels | P130005 |  | 1 | Event Description: IT WAS REPORTED THAT DURING PROCEDURE TO TREAT HEAVILY CALCIFIED LEFT ANTERIOR DESCENDING (LAD) LESION, A DIAMONDBACK 360 CORONARY PRECISION ORBITAL ATHERECTOMY DEVICE (OAD) WAS USED TO PERFORM FIRST TREATMENT ON LOW SPEED BACK AND FORTH A FEW TIMES. THE PHYSICIAN LET THE HEART REPERFUSE FOR FEW SECONDS AND THEN SECOND TREATMENT WAS PERFORMED ON HIGH SPEED AND WHEN THE PHYSICIAN REACHED THE LESION, THE CROWN JUMPED FORWARD. THE OAD WAS TURNED OFF AND REMOVED. A CONTRAST WAS USED AND PERFORATION WAS NOTED. THE VIPERWIRE CORONARY GUIDEWIRE WAS PULLED BACK AND THE DISTAL TIP COULD NOT MOVE AS THE GUIDEWIRE WAS SHEARED OFF. THE VIPERWIRE GUIDEWIRE WAS EXCHANGED FOR ANOTHER WIRE AND A BALLOON WAS INFLATED TO STOP BLEEDING. PERICARDIOCENTESIS WAS PERFORMED. COMPRESSIONS WERE PERFORMED AND AFTER THE PATIENT WAS CODED FOR 45 MINUTES, THE PATIENT EXPIRED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. THE DIAMONDBACK 360 CORONARY PRECISION ORBITAL ATHERECTOMY DEVICE REFERENCED IN B5 IS FILED UNDER SEPARATE MEDWATCH REPORT NUMBER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22816126&pc=DSQ | 3007042319-2025-01918 | 2025-08-12 04:00 | Death | HEARTWARE, INC. | 2025-08-15 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY FOLLOWING THE USE OF THE VENTRICULAR ASSIST DEVICE. THE DEVICE WAS IMPLANTED AND IS NOW OUT OF SERVICE. NO FURTHER INFORMATION REGARDING THE CAUSE OF DEATH, PATIENT SYMPTOMS, OR COMPLICATIONS WILL BE PROVIDED. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR INVESTIGATION SUMMARY. PRODUCT EVENT SUMMARY: THE VENTRICULAR ASSIST DEVICE (VAD) (B)(6) WAS NOT RETURNED FOR EVALUATION. THE COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE MANUFACTURING DOCUMENTATION WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. LOG FILE ANALYSIS WAS NOT PERFORMED SINCE LOG FILES WERE NOT AVAILABLE. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH IS UNKNOWN. BASED ON THE LIMITED INFORMATION AVAILABLE, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, DEATH IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22816142&pc=NGV | 2135147-2025-04570 | 2025-07-22 04:00 | Death | ABBOTT MEDICAL | 2025-08-15 04:00 | NGV | AMPLATZER AMULET | Patient-Device Incompatibility | Bradycardia; Low Blood Pressure/ Hypotension; Perforation; Shock; Cardiac Tamponade; Cognitive Changes; Pericardial Effusion | P200049 |  | 1 | Event Description: THE ARTICLE, "PULMONARY ARTERY PERFORATION AFTER PERCUTANEOUS LEFT ATRIAL APPENDAGE CLOSURE: TWO CASE REPORTS AND REVIEW OF LITERATURE", WAS REVIEWED. THE ARTICLE PRESENTED A CASE STUDY OF AN 83-YEAR-OLD FEMALE PATIENT WITH PAROXYSMAL ATRIAL FIBRILLATION (PAF), HYPERTENSION, DIABETES MELLITUS, NON-ISCHEMIC DILATED CARDIOMYOPATHY, AND RECURRENT MAJOR GASTROINTESTINAL (GI) BLEEDING WHILE TAKING EITHER WARFARIN OR RIVAROXABAN. IT WAS REPORTED THAT ON AN UNKNOWN DATE, A 28MM AMPLATZER AMULET WAS CHOSEN FOR IMPLANT. LEFT ATRIAL APPENDAGE (LAA) LANDING ZONE MEASURED 24MM VIA COMPUTERIZED TOMOGRAPHY (CT) AND 23MM VIA TRANSESOPHAGEAL ECHOCARDIOGRAPHY (TEE). THE DEVICE APPEARED WELL-COMPRESSED AND POSITIONED IN THE ABSENCE OF PERICARDIAL EFFUSION. IT WAS THEN REPORTED APPROXIMATELY 5-7MIN AFTER DEPLOYMENT, THE PATIENT SUDDENLY DEVELOPED HYPOTENSION, BRADYCARDIA, AND CONFUSION. A TRANSTHORACIC ECHOCARDIOGRAPHY (TTE) REVEALED CARDIAC TAMPONADE AND A DECISION WAS MADE TO PERFORM EMERGENCY PERICARDIOCENTESIS. DESPITE INTERVENTION, HEMODYNAMIC INSTABILITY PERSISTED. THE PATIENT WAS THEN REINTUBATED AND UNDERWENT EMERGENCY SURGICAL INTERVENTION. INTRAOPERATIVE FINDINGS CONFIRMED PULMONARY ARTERY PERFORATION, WHICH WAS SURGICALLY REPAIRED. HOWEVER, SPONTANEOUS CIRCULATION COULD NOT BE RESTORED AND THE PATIENT PASSED AWAY DUE TO REFRACTORY CARDIOGENIC SHOCK. [(B)(6)]. Manufacturer Narrative: IT IS UNKNOWN IF THE DEVICE IS RETURNING FOR ANALYSIS. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: PULMONARY ARTERY PERFORATION AFTER PERCUTANEOUS LEFT ATRIAL APPENDAGE CLOSURE: TWO CASE REPORTS AND REVIEW OF LITERATURE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22816168&pc=MCX | 3004742232-2025-00240 | 2025-07-24 04:00 | Death | CARDIOVASCULAR SYSTEMS INCORPORATED | 2025-08-15 04:00 | MCX | DIAMONDBACK CORONARY | Unintended System Motion | Cardiac Arrest; Perforation of Vessels | P130005 |  | 1 | Event Description: IT WAS REPORTED THAT DURING PROCEDURE TO TREAT HEAVILY CALCIFIED LEFT ANTERIOR DESCENDING (LAD) LESION, A DIAMONDBACK 360 CORONARY PRECISION ORBITAL ATHERECTOMY DEVICE (OAD) WAS USED TO PERFORM FIRST TREATMENT ON LOW SPEED BACK AND FORTH A FEW TIMES. THE PHYSICIAN LET THE HEART REPERFUSE FOR FEW SECONDS AND THEN SECOND TREATMENT WAS PERFORMED ON HIGH SPEED AND WHEN THE PHYSICIAN REACHED THE LESION, THE CROWN JUMPED FORWARD. THE OAD WAS TURNED OFF AND REMOVED. A CONTRAST WAS USED AND PERFORATION WAS NOTED. THE VIPERWIRE CORONARY GUIDEWIRE WAS PULLED BACK AND THE DISTAL TIP COULD NOT MOVE AS THE GUIDEWIRE WAS SHEARED OFF. THE VIPERWIRE GUIDEWIRE WAS EXCHANGED FOR ANOTHER WIRE AND A BALLOON WAS INFLATED TO STOP BLEEDING. PERICARDIOCENTESIS WAS PERFORMED. COMPRESSIONS WERE PERFORMED AND AFTER THE PATIENT WAS CODED FOR 45 MINUTES, THE PATIENT EXPIRED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. THE VIPERWIRE GUIDE WIRE REFERENCED IN B5 IS FILED UNDER SEPARATE MEDWATCH REPORT NUMBER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22816347&pc=DXY | 2182208-2025-04040 | 2020-11-01 04:00 | Death | MEDTRONIC, INC. | 2025-08-15 04:00 | DXY | MEDTRONIC ICD | Appropriate Term/Code Not Available | Cardiac Arrest; Unspecified Infection; Heart Failure/Congestive Heart Failure |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) IMPLANTATION IN EARLY SELECTED HIGH-RISK PATIENTS AFTER PRIMARY PERCUTANEOUS CORONARY INTERVENTION (PCI) FOR ST-SEGMENT¿ELEVATION MYOCARDIAL INFARCTION. THE AUTHORS DESCRIBED PATIENT DEATHS; HOWEVER, THE SPECIFIC CAUSES OF DEATH WERE UNKNOWN. DEATHS INCLUDED INFECTION AND CARDIAC CAUSES DEFINED AS HEART FAILURE RELATED OR SUDDEN CARDIAC DEATH AND OCCURRED IN THE FOLLOW-UP TIME FRAME OF THREE YEARS. THERE WERE PATIENTS WHO EXPERIENCED IMPLANT RELATED COMPLICATIONS WHICH INCLUDED POCKET BLEEDING, LOCAL POCKET INFECTION, AND PNEUMOTHORAX. THERE WERE OTHER PATIENTS WHO RECEIVED INAPPROPRIATE THERAPY DUE TO DETECTION OF SUPRA VENTRICULAR TACHYCARDIA (SVT), UNKNOWN LEAD FAILURES, AND T-WAVE OVERSENSING (TWOS). THE STATUS OF THE DEVICES AND LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/61 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: LONG-TERM OUTCOME OF THE RANDOMIZED DAPA TRIAL. CIRCULATION: ARRHYTHMIA AND ELECTROPHYSIOLOGY. 2020. 13: E008484. DOI: 10.1161/CIRCEP.120.008484. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22816388&pc=DXY | 2182208-2025-04041 | 2020-11-01 04:00 | Death | MEDTRONIC, INC. | 2025-08-15 04:00 | DXY | MEDTRONIC LEAD | Appropriate Term/Code Not Available | Cardiac Arrest; Unspecified Infection; Heart Failure/Congestive Heart Failure |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) IMPLANTATION IN EARLY SELECTED HIGH-RISK PATIENTS AFTER PRIMARY PERCUTANEOUS CORONARY INTERVENTION (PCI) FOR ST-SEGMENT¿ELEVATION MYOCARDIAL INFARCTION. THE AUTHORS DESCRIBED PATIENT DEATHS; HOWEVER, THE SPECIFIC CAUSES OF DEATH WERE UNKNOWN. DEATHS INCLUDED INFECTION AND CARDIAC CAUSES DEFINED AS HEART FAILURE RELATED OR SUDDEN CARDIAC DEATH AND OCCURRED IN THE FOLLOW-UP TIME FRAME OF THREE YEARS. THERE WERE PATIENTS WHO EXPERIENCED IMPLANT RELATED COMPLICATIONS WHICH INCLUDED POCKET BLEEDING, LOCAL POCKET INFECTION, AND PNEUMOTHORAX. THERE WERE OTHER PATIENTS WHO RECEIVED INAPPROPRIATE THERAPY DUE TO DETECTION OF SUPRA VENTRICULAR TACHYCARDIA (SVT), UNKNOWN LEAD FAILURES, AND T-WAVE OVERSENSING (TWOS). THE STATUS OF THE DEVICES AND LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/61 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: LONG-TERM OUTCOME OF THE RANDOMIZED DAPA TRIAL. CIRCULATION: ARRHYTHMIA AND ELECTROPHYSIOLOGY. 2020. 13: E008484. DOI: 10.1161/CIRCEP.120.008484. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22816602&pc=DXY | 2124215-2025-57160 | 2025-06-06 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-08-15 04:00 | DXY | IMPLANTABLE PACEMAKER PULSE-GENERATOR | Inappropriate/Inadequate Shock/Stimulation | Chest Pain; Anxiety; Shock from Patient Lead(s) |  |  | 1 | Event Description: IT WAS REPORTED THAT THIS PATIENT WITH AN IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) WAS ADMITTED TO HOSPITAL FOLLOWING PALPITATIONS DUE TO ATRIAL FIBRILLATION (AF). THE PHYSICIAN WAS ALSO CONCERNED REGARDING POTENTIAL INAPPROPRIATE THERAPY. UPON REVIEW, IT WAS NOTED THAT SIX SHOCKS HAD BEEN DELIVERED DUE TO FAST AF. THE PHYSICIAN WANTED TO PERFORM A CARDIOVERSION BUT WAS CONCERNED ABOUT FURTHER INAPPROPRIATE THERAPY DELIVERY. IT WAS SUGGESTED TO USE A MAGNET TO TEMPORARILY DISABLE TACHYCARDIA THERAPY, BUT THE HEALTHCARE FACILITY WAS UNWILLING TO INTERROGATE THE DEVICE. THREE DAYS LATER, A REMOTE ALERT WAS RECEIVED, INDICATING THAT THE PATIENT HAD RECEIVED 32 INAPPROPRIATE SHOCKS AND THE PATIENT COMPLAINED OF PAIN. IT WAS AGAIN RECOMMENDED TO UTILIZE A MAGNET TO TEMPORARILY DISABLE SHOCK THERAPY. TWO DAYS LATER, THE PATIENT RECEIVED ANOTHER 29 SHOCKS. THE PHYSIOLOGY TEAM DISCUSSED POTENTIAL REPROGRAMMING OPTIONS WITH BOSTON SCIENTIFIC. THE NEXT DAY, A FURTHER 22 SHOCKS WERE DELIVERED. FOUR DAYS LATER, ANOTHER NINE SHOCKS WERE DELIVERED. THEN, AN EMAIL WAS RECEIVED, INDICATING THAT THE PATIENT HAD PASSED AWAY. THE MORTUARY INDICATED THERE HAD NOT BEEN A VENTRICULAR ARRHYTHMIA AT THE TIME OF DEATH, BUT THAT THE 98 SHOCKS OVER THE COURSE OF 10 DAYS WERE EXTREMELY PAINFUL, ANXIETY-INDUCING, AND LIKELY HAD A NEGATIVE PROGNOSTIC IMPACT ON THE PATIENT'S HEALTH. AT THIS TIME, THE ICD REMAINS IMPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22816627&pc=NPT | 9617601-2025-01005 | 2025-08-13 04:00 | Death | MEDTRONIC MEXICO S. DE R.L. DE CV | 2025-08-15 04:00 | NPT | EVOLUT PRO PLUS VALVE | Difficult or Delayed Positioning; Material Invagination; Difficult to Advance; Patient Device Interaction Problem | Bradycardia; Cardiac Arrest; Low Blood Pressure/ Hypotension; Ventricular Fibrillation | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT PRIOR TO THE IMPLANT OF A 34 MILLIMETER (MM) TRANSCATHETER VALVE, THE PATIENT HAD SEVERE AORTIC STENOSIS AND COMPLEX VASCULAR ANATOMY WITH SIGNIFICANT CALCIUM DEPOSITION IN THE LEFT CORONARY CUSP, AORTIC ARCH ANGULATION OF 50 DEGREES, INITIAL VALVE MEAN GRADIENT OF 80 MMHG, AORTIC VALVE AREA OF LESS THAN 0.3 SQUARED CENTIMETERS (CM²), AND SIGNIFICANT TORTUOSITY AND KINKING IN THE ILIAC ARTERIES. THE INITIAL RIGHT FEMORAL ACCESS APPROACH REQUIRED TWO GUIDEWIRES TO STRAIGHTEN THE ACCESS IN THE AORTA. LEFT FEMORAL ACCESS WAS ALSO PERFORMED USING A NON-MEDTRONIC (LUNDERQUIST) GUIDEWIRE TO STRAIGHTEN THE ABDOMINAL AORTA. HOWEVER, MULTIPLE ATTEMPTS WERE PERFORMED TO CROSS THE TRANSCATHETER VALVE USING VARIOUS NON-MEDTRONIC CATHETERS (AL1, IM, AL2, MP, JR) AND NON-MEDTRONIC GUIDEWIRES (STRAIGHT HYDROPHILIC, STRAIGHT TEFLON-COATED, 0.018); HOWEVER, THE VALVE WAS DIFFICULT TO CROSS DUE TO THE PATIENT'S CALCIFIED/TORTUOUS ANATOMY, SIGNIFICANT CALCIFICATION, AND "ECCENTRIC" CUSP OPENINGS. PACEMAKER STIMULATION WAS PERFORMED TO REDUCE SYSTOLIC EJECTION FORCE. SUBSEQUENTLY, A TRANSAPICAL APPROACH WAS PERFORMED UNDER GENERAL ANESTHESIA, AND THE APEX OF THE LEFT VENTRICLE WAS PUNCTURED USING A 6 FRENCH (FR) INTRODUCER SHEATH, AND A HYDROPHILIC GUIDEWIRE WAS PASSED THROUGH THE ABDOMINAL AORTA. A SNARE WAS INSERTED VIA RIGHT FEMORAL ACCESS TO PERFORM AN ANTEGRADE APPROACH. AFTER VALVE CROSSING, THE INTRODUCER SHEATH WAS REPLACED BY AN 18 FR (SENTRAN) SHEATH. AT THIS TIME, THE PATIENT DEVELOPED HEMODYNAMICALLY INSTABILITY WITH HYPOTENSION AND BRADYCARDIA, AND VASOACTIVE MEDICATION WAS ADMINISTERED. A PRE-DILATION WAS PERFORMED USING A 22 MM BY 40 MM NON-MEDTRONIC (MAX LD) BALLOON WHILE THE PATIENT REMAINED HEMODYNAMICALLY UNSTABLE. DURING THE IMPLANT OF THE VALVE, AN INFOLD WAS IDENTIFIED. DUE TO THE PATIENT'S CONDITION, THE PHYSICIAN DECIDED TO PROCEED WITH THE IMPLANT OF THE VALVE. THREE DEPLOYMENT ATTEMPTS WERE PERFORMED WITH SUBSEQUENT RECAPTURES, AND ON THE THIRD DEPLOYMENT ATTEMPT, THE VALVE WAS FULLY DEPLOYED; HOWEVER, THE PATIENT WAS ALREADY IN CARDIAC ARREST. CARDIOPULMONARY RESUSCITATION (CPR) WAS INITIATED FOLLOWING THE IMPLANT OF THE VALVE, BUT THE PATIENT WAS IN VENTRICULAR FIBRILLATION. AFTER ONE HOUR OF RESUSCITATION EFFORTS WITH THORACOTOMY, THE PATIENT DIED.Event Description: IT WAS REPORTED THAT PRIOR TO THE IMPLANT OF A 34 MILLIMETER (MM) TRANSCATHETER VALVE, THE PATIENT HAD SEVERE AORTIC STENOSIS AND COMPLEX VASCULAR ANATOMY WITH SIGNIFICANT CALCIUM DEPOSITION IN THE LEFT CORONARY CUSP, AORTIC ARCH ANGULATION OF 50 DEGREES, INITIAL VALVE MEAN GRADIENT OF 80 MMHG, AORTIC VALVE AREA OF LESS THAN 0.3 SQUARED CENTIMETERS (CM²), AND SIGNIFICANT TORTUOSITY AND KINKING IN THE ILIAC ARTERIES. THE INITIAL RIGHT FEMORAL ACCESS APPROACH REQUIRED TWO GUIDEWIRES TO STRAIGHTEN THE ACCESS IN THE AORTA. LEFT FEMORAL ACCESS WAS ALSO PERFORMED USING A NON-MEDTRONIC GUIDEWIRE TO STRAIGHTEN THE ABDOMINAL AORTA. HOWEVER, MULTIPLE ATTEMPTS WERE PERFORMED TO CROSS THE TRANSCATHETER VALVE USING VARIOUS NON-MEDTRONIC CATHETERS (AL1, IM, AL2, MP, JR) AND NON-MEDTRONIC GUIDEWIRES (STRAIGHT HYDROPHILIC, STRAIGHT TEFLON-COATED, 0.018); HOWEVER, THE VALVE WAS DIFFICULT TO CROSS DUE TO THE PATIENT'S CALCIFIED/TORTUOUS ANATOMY, SIGNIFICANT CALCIFICATION, AND "ECCENTRIC" CUSP OPENINGS. PACEMAKER STIMULATION WAS PERFORMED TO REDUCE SYSTOLIC EJECTION FORCE. SUBSEQUENTLY, A TRANSAPICAL APPROACH WAS PERFORMED UNDER GENERAL ANESTHESIA, AND THE APEX OF THE LEFT VENTRICLE WAS PUNCTURED USING A 6 FRENCH (FR) INTRODUCER SHEATH, AND A HYDROPHILIC GUIDEWIRE WAS PASSED THROUGH THE ABDOMINAL AORTA. A SNARE WAS INSERTED VIA RIGHT FEMORAL ACCESS TO PERFORM AN ANTEGRADE APPROACH. AFTER VALVE CROSSING, THE INTRODUCER SHEATH WAS REPLACED BY AN 18 FR SHEATH. AT THIS TIME, THE PATIENT DEVELOPED HEMODYNAMICALLY INSTABILITY WITH HYPOTENSION AND BRADYCARDIA, AND VASOACTIVE MEDICATION WAS ADMINISTERED. A PRE-DILATION WAS PERFORMED USING A 22 MM BY 40 MM NON-MEDTRONIC BALLOON WHILE THE PATIENT REMAINED HEMODYNAMICALLY UNSTABLE. DURING THE IMPLANT OF THE VALVE, AN INFOLD WAS IDENTIFIED. DUE TO THE PATIENT'S CONDITION, THE PHYSICIAN DECIDED TO PROCEED WITH THE IMPLANT OF THE VALVE. THREE DEPLOYMENT ATTEMPTS WERE PERFORMED WITH SUBSEQUENT RECAPTURES, AND ON THE THIRD DEPLOYMENT ATTEMPT, THE VALVE WAS FULLY DEPLOYED; HOWEVER, THE PATIENT WAS ALREADY IN CARDIAC ARREST. CARDIOPULMONARY RESUSCITATION (CP R) WAS INITIATED FOLLOWING THE IMPLANT OF THE VALVE, BUT THE PATIENT WAS IN VENTRICULAR FIBRILLATION. AFTER ONE HOUR OF RESUSCITATION EFFORTS WITH THORACOTOMY, THE PATIENT DIED. ADDITIONAL INFORMATION WAS RECEIVED INDICATING THAT DURING LOADING OF THE VALVE, NO VISIBLE ERRORS WERE NOTED. THE NATIVE ANATOMY HAD SIGNIFICANT CALCIFICATION, TORTUOSITY, A UNIFORM 0.3 MM OPENING, AND REDUCED MOBILITY OF THE CUSPIDS. THESE ANATOMICAL FACTORS MADE HANDLING OF THE DELIVERY CATHETER SYSTEM (DCS) DIFFICULT. MODERATE AORTIC INSUFFICIENCY WAS ALSO OBSERVED PRIOR TO THE CROSSING OF THE VALVE. DUE TO STENOSIS, AN INITIAL ATTEMPT AT CROSSING THE AORTIC VALVE WITH THE DCS WAS UNSUCCESSFUL. MULTIPLE GUIDING CATHETERS WERE USED TO AID THE CROSSING OF THE VALVE AND THE DECISION WAS MADE TO USE AN ANTEROGRADE APPROACH BY THE TRANSAPICAL ROUTE. A SURGICAL INCISION MADE BY MINI LEFT THORACOTOMY AND PUNCTURE OF THE HEART APEX WITH THE 6 FR FEMORAL INTRODUCER SHEATH WAS PERFORMED. CROSSING OF THE VALVE WITH A NON-MEDTRONIC GUIDING CATHETER WAS CARRIED OUT. A 6 FR PIG-TAIL CATHETER WAS POSITIONED IN THE LEFT VENTRICLE RETROGRADELY THROUGH THE EXTERNALIZED GUIDEWIRE. SUBSEQUENTLY, THIS PIG-TAIL CATHETER WAS EXCHANGED FOR AN 0.035 NON-MEDTRONIC EXTRA SMALL SUPPORT WIRE. THE PHYSICIAN POSITIONED A 0.035 NON-MEDTRONIC "BUDDYWIRE" IN PARALLEL FOR LEFT CONTRALATERAL FEMORAL ACCESS. A PRE-IMPLANT BALLOON DILATION WITH A NON-MEDTRONIC 22 MM X 40 MM BALLOON, AND THE VALVE WAS IMPLANTED. THE PATIENT ENTERED CARDIOPULMONARY ARREST DURING THE FIRST ATTEMPT AT DEPLOYMENT AND INTERNAL CURVATURE IN THE VALVE FRAME WAS NOTED AT THIS TIME. THE DECISION WAS MADE TO NOT WITHDRAW THE VALVE AS THE PATIENT WAS UNSTABLE. THE VALVE WAS RECAPTURED TWICE. DURING THE FIRST TWO DEPLOYMENTS, THE VALVE WAS TOO HIGH IN THE NON-CORONARY CUSP SINUS AND TOO DEEP IN THE LEFT CORONARY CUSP SINUS. ON THE THIRD ATTEMPT AT DEPLOYMENT, THE VALVE WAS RELEASED DEEPER AND THE PATIENT'S HEART RHYTHM WAS ASYSTOLIC. PER THE PHYSICIAN, THE VALVE AND DCS DID NOT CONTRIBUTE TO THE CARDIAC ARREST OR DEATH. PER THE PHYSICIAN, THE GUIDEWIRE DID NOT CONTRIBUTE TO THE PATIENT'S DEATH. THE CAUSE OF DEATH WAS ATTRIBUTED TO THE TECHNICAL DIFFICULTY INHERENT IN THE PROCEDURE. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. SECTION D INFORMATION REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER RELEVANT DEVICE(S) ARE: PRODUCT ID: D-EVPROP34, SERIAL/LOT #: (B)(6), UBD: 24-FEB-2027, UDI#: (B)(4) THE CODES PRESENT IN SECTION H6 CORRESPOND TO MULTIPLE COMPONENTS/PRODUCTS THAT COMPRISE THIS REPORTED EVENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: A1. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. H6. ADDITIONAL CODES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22816798&pc=LWP | 2017865-2025-97828 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-08-15 04:00 | LWP | ASSURITY MRI | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P140033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DURING A CARDIAC PROCEDURE. IT IS UNKNOWN WHETHER THE PULSE GENERATOR, RIGHT VENTRICULAR (RV) LEAD OR RIGHT ATRIAL (RA) LEAD CONTRIBUTED TO THE PATIENT'S DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22816814&pc=NVN | 2017865-2025-97837 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-08-15 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DURING A CARDIAC PROCEDURE. IT IS UNKNOWN WHETHER THE PULSE GENERATOR, RIGHT VENTRICULAR (RV) LEAD OR RIGHT ATRIAL (RA) LEAD CONTRIBUTED TO THE PATIENT'S DEATH. Manufacturer Narrative: THE LEAD WAS RETURNED DUE TO PATIENT DEATH. AS RECEIVED, ONLY THE CONNECTOR PORTION OF THE LEAD WAS RETURNED IN ONE PIECE. ELECTRICAL TESTING AND X-RAY EXAMINATION DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS. VISUAL INSPECTION OF THE LEAD DID NOT FIND ANY ANOMALIES EXCEPT FOR THE CUT END WHICH IS CONSISTENT WITH PROCEDURAL DAMAGE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22816840&pc=NVN | 2017865-2025-97838 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-08-15 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DURING A CARDIAC PROCEDURE. IT IS UNKNOWN WHETHER THE PULSE GENERATOR, RIGHT VENTRICULAR (RV) LEAD OR RIGHT ATRIAL (RA) LEAD CONTRIBUTED TO THE PATIENT'S DEATH. Manufacturer Narrative: THE LEAD WAS RETURNED DUE TO PATIENT DEATH. AS RECEIVED, ONLY THE CONNECTOR PORTION OF THE LEAD WAS RETURNED IN ONE PIECE. ELECTRICAL TESTING AND X-RAY EXAMINATION DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS. VISUAL INSPECTION OF THE RETURNED PORTION DID NOT FIND ANY ANOMALIES EXCEPT FOR THE CUT END WHICH IS CONSISTENT WITH PROCEDURAL DAMAGE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22816855&pc=QLG | 2954323-2025-32434 | 2025-08-08 04:00 | Death | ABBOTT DIABETES CARE INC | 2025-08-15 04:00 | QLG | FREESTYLE LIBRE 2 | Insufficient Information | Insufficient Information | K193371 |  | 1 | Event Description: ABBOTT DIABETES CARE (ADC) RECEIVED A MEDWATCH REPORT MW5173989, WHICH INCLUDED THE FOLLOWING INFORMATION: A HEALTHCARE PROFESSIONAL (HCP) STATED THAT A CUSTOMER HAD PASSED AWAY. NO ADDITIONAL DETAILS OR CLARIFICATION WERE PROVIDED. IT IS UNKNOWN WHETHER THE CUSTOMER WAS ACTIVELY USING AN ADC DEVICE AT THE TIME OF DEATH, AND NO ALLEGATION OF DEVICE DEFICIENCY WAS MADE. THE REPORT NOTED THAT THE HEALTHCARE PROFESSIONAL DID NOT CONSENT TO FOLLOW-UP. NO CONTACT INFORMATION WAS PROVIDED TO ADC; THEREFORE, ADC IS UNABLE TO ATTEMPT ANY FOLLOW-UP ACTIVITIES RELATED TO THIS EVENT. BASED ON THE INFORMATION PROVIDED, IT CANNOT BE REASONABLY SUGGESTED THAT AN ADC PRODUCT HAS OR MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED DEATH. Manufacturer Narrative: THIS COMPLAINT WAS RECEIVED VIA USER REPORT MW5173989 AND HAS BEEN REPORTED TO FDA. ABBOTT DIABETES CARE (ADC) IS UNABLE TO FURTHER INVESTIGATE THE REPORTED COMPLAINT DUE TO THE ABSENCE OF AN ALLEGED DEVICE DEFICIENCY. ADDITIONALLY, NO CONTACT INFORMATION WAS PROVIDED TO ADC; THEREFORE, ADC IS UNABLE TO ATTEMPT ANY FOLLOW-UP ACTIVITIES RELATED TO THIS EVENT. THE DATE OF THE EVENT IS UNKNOWN. THE DATE ENTERED IN SECTION B3 IS THE DATE ABBOTT DIABETES CARE BECAME AWARE OF THE EVENT. ALL PERTINENT INFORMATION AVAILABLE TO ADC HAS BEEN SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22816993&pc=OZD | 1220648-2025-31162 | 2025-07-18 04:00 | Death | ABIOMED, INC. | 2025-08-15 04:00 | OZD | IMPELLA CP | Restricted Flow rate | Thrombosis/Thrombus | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT A PATIENT HAD AN IMPELLA CP IMPLANTED FROM AN AXILLARY APPROACH FOLLOWING AN ABORTED IMPELLA 5.5 IMPLANT DUE TO PATIENT ANATOMY. A DOUBLE BARREL TECHNIQUE WAS USED (FEMORAL IMPELLA CP WAS IN PLACE AND RUNNING AT P6-7). AXILLARY IMPELLA CP WAS RAMPED UP TO P6 AS FEMORAL IMPELLA CP WAS WEANED DOWN TO P1 WITHDRAWN INTO THE DESCENDING AORTA AND PLACED INTO SURGICAL MODE. FLOWS REACHED THREE LITERS PER MINUTE AT P6 FOR APPROXIMATELY ONE MINUTE UNTIL A SUCTION ALARM PRESENTED AND FLOWS DROPPED TO ONE LITER PER MINUTE. POSITION WAS CONFIRMED OPTIMAL AND VOLUME WAS ADMINISTERED. FOLLOWING THE EXPLANT, CLOT LIKE DEBRIS WAS VISIBLE IN THE PROXIMAL CANNULA AND UPON REVIEW, IMPELLA CONNECT THERE WAS A SPIKE IN MOTOR CURRENT. FOLLOWING LOSS OF FLOW, THE PATIENT DETERIORATED REQUIRING CHEMICAL SUPPORT AND CARDIOPULMONARY RESUSCITATION BUT ULTIMATELY EXPIRED ON SUPPORT. Manufacturer Narrative: THE IMPELLA PUMP WAS RETURNED, AND AN EVALUATION IS UNDERWAY. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION. SECTION: IMPELLA CP WITH SMART ASSIST CATHETER INSERTION. SECTION: AXILLARY INSERTION OF THE IMPELLA CP WITH SMART ASSIST CATHETER. SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES. ¿IN PATIENTS WITH TRANSCATHETER AORTIC VALVES POSITION THE IMPELLA SYSTEM CAREFULLY TO AVOID INTERACTION WITH THE TRANSCATHETER AORTIC VALVE PROSTHESIS. UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH. IN THIS SITUATION, AVOID REPOSITIONING WHILE THE DEVICE IS RUNNING; TURN THE DEVICE TO P0 DURING REPOSITIONING OR ANY MOVEMENT THAT COULD BRING THE OUTLET WINDOWS INTO PROXIMITY TO THE VALVE STENT STRUCTURES. IF THERE IS LOW FLOW OBSERVED IN A PATIENT IMPLANTED WITH A TRANSCATHETER AORTIC VALVE PROSTHESIS, CONSIDER DAMAGE OF THE IMPELLER AND REPLACE THE IMPELLA AS SOON AS POSSIBLE.¿ "UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22817058&pc=MHX | 9610816-2025-000754 | 2025-07-22 04:00 | Death | PHILIPS MEDIZIN SYSTEME BÖBLINGEN GMBH | 2025-08-15 04:00 | MHX | PATIENT INFORMATION CENTER IX | Use of Device Problem | Insufficient Information | K211900 |  | 1 | Event Description: PHILIPS RECEIVED A COMPLAINT ON THE PATIENT INFORMATION CENTER IX INDICATING THAT A PATIENT CODED AND PASSED AWAY, BUT THE USERS WERE NOT AWARE. Manufacturer Narrative: A PHILIPS REMOTE SERVICE ENGINEER (RSE) SPOKE TO THE BIOMEDICAL ENGINEER, WHO CONNECTED THE RSE WITH THE NURSE MANAGER. THE NURSE MANAGER ADVISED THAT THE PATIENT WAS TRANSFERRED FROM ANOTHER UNIT AT 12:48 PM, AND THE TELEMETRY DEVICE WAS REPLACED AT 8:55 PM; HOWEVER, THE TELEMETRY EQUIPMENT WAS NOT PAIRED WITH THE PIC IX CENTRAL STATION, AND THE TELEMETRY DEVICE WAS UNASSIGNED. THE PATIENT CODED AND PASSED AWAY WITHOUT THE USER BEING AWARE. THE CUSTOMER REQUESTED ASSISTANCE ACCESSING ANY AUDIT LOTS. IT WAS RECOMMENDED TO REVIEW THE ADMISSION/TRANSFER WORKFLOW AND ADDITIONAL EDUCATION MAY BE PROVIDED. IT WAS CLARIFIED THERE WAS NO ALLEGATION OF DEFICIENCY WITH THE DEVICE BUT THAT THE STAFF DID NOT PAIR THE TELEMETRY DEVICE. BASED ON THE INFORMATION AVAILABLE AND THE TESTING CONDUCTED, THE CAUSE OF THE REPORTED PROBLEM WAS THE USER. THE REPORTED PROBLEM WAS NOT CONFIRMED. THE RSE ADVISED THE NURSE MANAGER ON HOW TO RETRIEVE PRIOR DATA IN GENERAL REVIEW AND HOW TO SEARCH THE CLINICAL AUDIT TRAIL. THE RSE EMAILED A COPY OF THE CLINICAL AUDIT TRAIL GUIDE, THE CONFIGURATION MANUAL, A QUICK GUIDE FOR THE PIC IX, AND PAGES TO REFERENCE PRIOR DATA. THE DEVICE WAS CONFIRMED TO BE OPERATING PER SPECIFICATIONS AND NO FAILURE WAS IDENTIFIED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22817119&pc=MFE | 2029214-2025-01830 | 2025-03-26 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-08-15 04:00 | MFE | ONYX | Adverse Event Without Identified Device or Use Problem |  |  |  | 1 | Event Description: ISIDOR, J., ABO KASEM, R., SOWLAT, M.-M., CUNNINGHAM, C., LEVITT, M., MCGRATH, M., OGILVY, C. S., ALWAKAA, O., PAUL, A. R., CULLEN, M., ABECASSIS, I. J., ALHOURANI, A., FRAGATA, I., BAPTISTA, M., ALARAJ, A., TSHIBANGU, M., SIDDIQUI, F., LARSON, E., PSYCHOGIOS, M.-N., ¿ STRICKLAND, B. A. (2025). COMPARISON OF EMBOLIC AGENTS IN PREOPERATIVE EMBOLIZATION FOR INTRACRANIAL MENINGIOMAS: MULTICENTER ADJUSTED ANALYSIS OF 275 CASES. JOURNAL OF NEUROINTERVENTIONAL SURGERY. HTTPS://DOI.ORG/10.1136/JNIS-2025-023062 LITERATURE WAS REVIEWED REGARDING A STUDY THAT COMPARED THREE TYPES OF EMBOLIC MATERIALS¿PARTICLES, LIQUID AGENTS, AND COILS¿USED IN PREOPERATIVE EMBOLIZATION FOR MENINGIOMAS INVOLVING 275 PATIENTS. MULTIPLE MANUFACTURER¿S DEVICES WERE IMPLANTED IN THE STUDY POPULATION. THE FOLLOWING MEDTRONIC DEVICES WERE USED: ONYX. DEATHS OCCURRED IN THE STUDY POPULATION; HOWEVER, THERE WAS NO STATEMENT ESTABLISHING A CAUSAL OR CONTRIBUTORY RELATIONSHIP BETWEEN MEDTRONIC PRODUCT AND THE DEATHS. TWO DEATHS OCCURRED IN PATIENTS WHERE ONYX WAS USED, BUT THE CAUSE OF THEIR DEATHS WERE NOT SPECIFIED. AMONG ALL ONYX PATIENTS ADVERSE EVENTS / DEVICE PRODUCT PERFORMANCE ISSUES INCLUDED: HEMORRHAGE/EDEMA ASSOCIATED WITH EMBOLIZATION, BLOOD TRANSFUSION, POST-EMBOLIZATION STROKE, EDEMA CAUSING MASS EFFECT, CRANIAL NERVE PALSY, POSTOPERATIVE INFECTION, RETREATMENT/RESCUE SURGERY, MORTALITY. IN THE STUDY COHORT, ONYX REPRESENTED 96.6% OF THE LIQUID EMBOLIC AGENTS USED FOR PREOPERATIVE MENINGIOMA EMBOLIZATION. PATIENTS IN THE ONYX GROUP SHOWED A HIGHER PERCENTAGE OF GROSS TOTAL RESECTION, A LOWER RETREATMENT RATE, AND A STATISTICALLY INSIGNIFICANT HIGHER COMPLICATION RATE. NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. Manufacturer Narrative: LITERATURE ARTICLE IS INCLUDED IN THE ATTACHMENTS. A.2. THIS VALUE IS THE AVERAGE AGE OF THE PATIENTS REPORTED IN THE ARTICLE AS SPECIFIC PATIENTS COULD NOT BE IDENTIFIED. A.3. THIS VALUE REFLECTS THE GENDER OF THE MAJORITY OF THE PATIENTS REPORTED IN THE ARTICLE AS SPECIFIC PATIENTS COULD NOT BE IDENTIFIED. B.3. PLEASE NOTE THAT THIS DATE IS BASED OFF OF THE DATE OF ONLINE PUBLICATION OF THE ARTICLE. G4: PMA CODE MISSING AS DEVICE MODEL AND LOT NUMBERS ARE UNKNOWN. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22817285&pc=MHX | 9610816-2025-000753 | 2025-05-02 04:00 | Death | PHILIPS MEDIZIN SYSTEME BÖBLINGEN GMBH | 2025-08-15 04:00 | MHX | INTELLIVUE PATIENT MONITOR MX750 | Defective Alarm | Cardiac Arrest | K210906 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT WENT INTO CARDIAC ARREST AFTER BRADYCARDIA AND DESATURATION, BUT THE MONITOR DID NOT GENERATE ALARMS AS EXPECTED. IT WAS INDICATED THE MONITOR WAS NETWORKED WITH A CENTRAL STATION WHERE ALARMS WERE ON AND AUDIBLE. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTION WAS REPORTED; THEREFORE, GOOD FAITH EFFORTS ARE BEING PERFORMED. PREVIOUS INFORMATION OBTAINED INDICATED THE CENTRAL STATION PROVIDED ALARMS AND PATIENT HARM WAS MARKED AS NO. THE OUTCOME FOR THE PATIENT WAS INDICATED AS LIKELY BRAIN DAMAGE. THIS NEW INFORMATION INDICATED THE NURSES DID NOT RESPOND TO THE ALARMS AT THE CENTRAL STATION, WHICH RESULTED IN A DELAY IN CARE; HOWEVER, IT WAS ALSO NOTED THE CARDIAC ARREST WAS WITNESSED BY NURSING STAFF. IN ADDITION, IT REMAINED UNCLEAR WHETHER THE USER WAS SIMPLY WONDERING WHY THERE WERE NO ALARMS AT THE BEDSIDE WHEN ALARMS WERE OCCURRING AT THE CENTRAL STATION; HOWEVER, THE CUSTOMER ALSO INDICATED THE PATIENT HARM WAS 'MOST LIKELY' RELATED TO THE LACK OF BEDSIDE ALARM. BASED ON THE CONTRADICTORY INFORMATION RECEIVED, IT CANNOT BE DETERMINED WHETHER THE ALLEGED DEVICE MALFUNCTION CAUSED OR CONTRIBUTED TO THE REPORTED DEATH. ADDITIONAL DATA HAS BEEN RECEIVED FOR REVIEW. Manufacturer Narrative: PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE. REPORTING INSTITUTION PHONE # (B)(6). REPORTER PHONE # (B)(6).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22817287&pc=DYB | 2015691-2025-06688 | 2025-07-22 04:00 | Death | EDWARDS LIFESCIENCES | 2025-08-15 04:00 | DYB | EDWARDS ESHEATH+ INTRODUCER SET | Difficult to Insert; Material Puncture/Hole; Physical Resistance/Sticking | Perforation of Vessels | K200258 |  | 1 | Event Description: AS REPORTED BY THE FIELD CLINICAL SPECIALIST (FCS), DURING A TRANSFEMORAL TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) PROCEDURE WITH A 26 MM SAPIEN 3 ULTRA VALVE, IT WAS INDICATED THAT THERE HAD BEEN SOME DIFFICULTY INSERTING THE ESHEATH+ IN THE PATIENT. DURING ADVANCEMENT OF THE 26 MM COMMANDER DELIVERY SYSTEM AND CRIMPED VALVE THROUGH THE 14FR ESHEATH+, THERE WAS SLIGHT RESISTANCE ENCOUNTERED. THE TECHNICIAN NOTED AN INABILITY TO PUSH OR PULL THE WIRE. THE DELIVERY SYSTEM AND CRIMPED VALVE HAD NOT EXITED THE TIP OF THE ESHEATH+. UNDER IMAGING, THE DELIVERY SYSTEM WITH CRIMPED VALVE WAS OBSERVED TO HAVE EXITED/PROTRUDED OUT THE SIDE OF THE ESHEATH+ IN THE RIGHT ILIAC. THE VALVE WAS NOTED TO BE COMPLETELY BENT UPSIDE-DOWN, AND THE DELIVERY SYSTEM WAS BOWING INTO THE ABDOMINAL SPACE. THERE WAS DIFFICULTY REMOVING THE DELIVERY SYSTEM WITH CRIMPED VALVE AND ESHEATH+. WHEN THE DELIVERY SYSTEM WAS PULLED BACK, THE VALVE WAS CAUGHT OUTSIDE OF THE ESHEATH+ AND THE RIGHT ILIAC. THERE WAS A PERFORATION OBSERVED AT THE COMMON ILIAC AND DESCENDING AORTA, CAUSING BLEEDING. CONTRALATERAL ACCESS WAS OBTAINED AND A CODA (NON-EDWARDS) BALLOON WAS PLACED IN THE ASCENDING AORTA ABOVE THE ILIAC BIFURCATION TO OCCLUDE THE BLEEDING. CARDIOPULMONARY RESUSCITATION (CPR) WAS ALSO PERFORMED. THE PATIENT WAS TRANSFERRED TO AN OPERATING ROOM (OR) FOR OPEN VASCULAR REPAIR TO REPAIR THE TORN ILIAC AND AORTA PERFORATION. IT WAS NOTED THAT THE ESHEATH+, DELIVERY SYSTEM AND VALVE HAD REMAINED IN THE PATIENT'S BODY. DURING THE VASCULAR REPAIR, THE PATIENT EXPIRED. SUBSEQUENTLY, ADDITIONAL INFORMATION WAS PROVIDED. REGARDING THE VALVE BEING COMPLETELY BENT UPSIDE-DOWN, IT WAS CLARIFIED THAT THE VALVE ON THE DELIVERY SYSTEM DID A "LOOP-DE-LOOP" IN THE ABDOMINAL CAVITY. IT APPEARED THAT THE VALVE MAY HAVE CAUGHT ON THE EDGE OF THE SHEATH TEAR. CONTINUED FORWARD PRESSURE DURING ADVANCEMENT LIKELY CAUSED THE VALVE TO BUCKLE AND INVERT OUTSIDE OF THE SHEATH, RESULTING IN BOWING OF THE DELIVERY SYSTEM. IMAGERY WAS PROVIDED FOR EVALUATION. A REVIEW OF THE IMAGERY INDICATED THERE WERE NO ABNORMALITIES OBSERVED ON THE VALVE. THE CRIMPED VALVE WAS OBSERVED TO BE PUNCTURING THROUGH THE SIDE OF THE SHEATH SHAFT. THERE WAS ALSO POTENTIAL VALVE DIVING AND THE SILHOUETTE OF THE CATHETER DID NOT APPEAR TO BE COAXIAL WITH THE CRIMPED VALVE. Manufacturer Narrative: THE INVESTIGATION IS ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22817309&pc=OZD | 1220648-2025-45832 | 2024-10-23 04:00 | Death | ABIOMED, INC. | 2025-08-15 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Anemia; Arrhythmia; Renal Impairment | P140003 |  | 1 | Event Description: WE RECEIVED A NOTIFICATION VIA ABIOMED¿S LONG-TERM OUTCOME AND QUALITY INDICATOR (LOQI) IMPELLA REGISTRY REGARDING A PATIENT THAT ENDURED ACUTE KIDNEY INJURY WITH A "POSSIBLE" RELATIONSHIP TO THE IMPELLA DEVICE USED. THE PATIENT WAS GIVEN A BOLUS LASIX WHICH HELPED THE CONDITION. ADDITIONALLY, PATIENT ENDURED ANEMIA WITH A "POSSIBLE" RELATIONSHIP TO THE IMPELLA DEVICE USED. HEMOGLOBIN DROPPED DOWN TO 7.1 AND ONE UNIT OF BLOOD WAS ORDERED. THE URINE THEN LOOKED CLEAR WITH NO HEMATOMAS OR ECCHYMOSIS. THE CAUSE OF THE ANEMIA WAS UNCLEAR. PATIENT ALSO ENDURED ONGOING VENTRICULAR ARRYTHMIAS WITH A "PROBABLE" RELATIONSHIP TO THE IMPELLA DEVICE USED. IT WAS REPORTED THAT THE PATIENT CODED ON INDUCTION OF ANESTHESIA AND CARDIOPULMONARY RESUSCITATION WAS PERFORMED FOR HYPOTENSION FOR 5-10 MINUTES. THE PATIENT STABILIZED WITH MODERATE TO HIGH DOSES OF INOTROPES, DOBUTAMINE, EPINEPHRINE, VASOPRESSIN, AND INHALED NITRIC OXIDE. IN THE LEFT AXILLARY AN IMPELLA 5.5 WAS PLACED FOR MECHANICAL CIRCULATORY SUPPORT. THE PREDOMINANT UNDERLYING RHYTHM WAS SINUS TACHYCARDIA WITH FREQUENT VENTRICULAR ECTOPY BEING NOTED. THERE WAS AN INTRAVENTRICULAR CONDUCTION DELAY AND THE PATIENT EXPERIENCED EPISODES OF VENTRICULAR TACHYCARDIA/VENTRICULAR FIBRILLATION WITH LOSS OF CONSCIOUSNESS WITH COUGHING. THE FIRST EPISODE OCCURRED WHILE THE PATIENT WAS OUT OF BED, IN A CHAIR. THE PATIENT WAS SEDATED AND CARDIOVERTED. THE PATIENT HAD ADDITIONAL EPISODES REQUIRING SHOCK AFTER HE HAD BEEN LIFTED BACK TO BED. THE PATIENT STARTED ON AMIODARONE AND LIDOCAINES. A BEDSIDE ECHOCARDIOGRAM PERFORMED, AND THE IMPELLA WAS REPOSITIONED. THE PATIENT WAS ELECTIVELY INTUBATED. THE PATIENT NOT RECOVERED/NOT RESOLVED. THE PATIENT ULTIMATELY EXPIRED. Manufacturer Narrative: PATIENT INFORMATION: A4 PATIENT WEIGHT UNKNOWN. DEVICE STATUS: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE: AS NOTED IN THE IMPELLA IFU SECTION: POTENTIAL ADVERSE EVENTS - ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22818129&pc=NPT | 2015691-2025-06692 | 2025-07-22 04:00 | Death | EDWARDS LIFESCIENCES | 2025-08-15 04:00 | NPT | COMMANDER DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | Rupture | P140031 |  | 1 | Event Description: AS REPORTED BY AN EDWARDS THAILAND AFFILIATE, DURING A TRANSFEMORAL TAVR PROCEDURE WITH A 23 MM SAPIEN 3 TRANSCATHETER HEART VALVE, THE COMMANDER DELIVERY SYSTEM WAS ADVANCED INTO THE AORTIC ROOT; HOWEVER, CROSSING THE NATIVE AORTIC VALVE PROVED DIFFICULT DUE TO SEVERE CALCIFICATION. ADDITIONAL FLEXION WAS APPLIED, AND THE DELIVERY SYSTEM WAS ADVANCED WITH INCREASED FORCE TO CROSS THE VALVE. SHORTLY THEREAFTER, THE PATIENT EXPERIENCED A SUDDEN DROP IN BLOOD PRESSURE. A ROOT AORTOGRAM REVEALED AN AORTIC ROOT DISSECTION AND RUPTURE. AN EMERGENCY STERNOTOMY WAS PERFORMED, WHICH IDENTIFIED A 2 CM RUPTURE AT THE SINOTUBULAR JUNCTION. SURGICAL REPAIR WAS CARRIED OUT. FOLLOWING PROLONGED CARDIAC ARREST, A DISCUSSION WAS HELD WITH THE PATIENT'S FAMILY THAT DUE TO THE CLINICAL SITUATION AND POOR PROGNOSIS, END OF LIFE CARE WAS RECOMMENDED. AS PER MEDICAL OPINION, THE DELIVERY SYSTEM WAS FUNCTIONING PROPERLY AND THIS EVENT IS NOT RELATED TO ANY DEFECT IN THE DEVICE. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. ACCORDING TO THE DEVICE TRAINING MANUALS, RISK FACTORS FOR AORTIC DISSECTION, CARDIAC/AORTIC HEMATOMA, OR ANNULAR RUPTURE DURING THE TVR PROCEDURE INCLUDE SIGNIFICANT VALVE OVER SIZING, SEVERELY OBLITERATED SINUSES OF VALSALVA, PORCELAIN AORTA AND/OR PRESENCE OF BULKY CALCIFICATION, AND NARROW CALCIFIED SINOTUBULAR JUNCTION. IN ADDITION, ADVANCED AGE, FEMALE GENDER, SMALL BODY WEIGHT, AND STEROID DEPENDENCY CAN ALSO BE CONTRIBUTING FACTORS. THE SAPIEN TRANSCATHETER HEART VALVE (ALL MODELS) RELIES CALCIUM TO SECURELY ANCHOR IN THE LANDING ZONE. DESPITE THIS BENEFICIAL ASPECT OF CALCIUM, BULKY CALCIUM CAN INCREASE THE RISK OF CALCIFIC NODULE DISPLACEMENT INTO THE VASCULATURE, WHICH CAN LEAD TO VASCULAR INJURY. AT TIMES, THE EXTENT AND DISTRIBUTION OF CALCIUM CAN IMPAIR EASE OF DELIVERY OF THE VALVE, CORRECT POSITIONING OF THE VALVE, DEPLOYMENT OF THE VALVE, AND PROCEDURAL SUCCESS. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION THAT A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST PATIENT AND PROCEDURAL FACTORS (CROSSING DIFFICULTY DUE TO CALCIFICATION) CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22819121&pc=NPT | 9617601-2025-01011 | 2025-07-24 04:00 | Death | MEDTRONIC MEXICO S. DE R.L. DE CV | 2025-08-15 04:00 | NPT | EVOLUT FX PLUS VALVE | Device Dislodged or Dislocated; Patient Device Interaction Problem | Stroke/CVA; Foreign Body In Patient | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT PRIOR TO THE IMPLANT OF A TRANSCATHETER VALVE, A PRE-IMPLANT BALLOON AORTIC VALVULOPLASTY (BAV) WAS PERFORMED DUE TO A MEAN GRADIENT OF 66 MILLIMETERS OF MERCURY (MM HG). FOLLOWING THE IMPLANT OF THE VALVE, THE VALVE DISLODGED DUE TO INTERACTION WITH THE PIGTAIL CATHETER. PRIOR TO VALVE DISLODGEMENT, THE IMPLANT DEPTH ON THE NON-CORONARY CUSP (NCC) WAS 3-4 MILLIMETERS (MM), AND THE IMPLANT DEPTH ON THE LEFT CORONARY CUSP (LCC) WAS 5 MM. AFTER VALVE DISLODGEMENT, THE IMPLANT DEPTH ON THE NCC AND LCC WAS APPROXIMATELY 5 MM DISTAL TO THE INNOMINATE ARTERY. FOLLOWING DISLODGEMENT, AN AORTIC DISSECTION EXTENDING INTO THE RIGHT CAROTID ARTERY WAS OBSERVED. SUBSEQUENTLY, A SECOND TRANSCATHETER VALVE WAS IMPLANTED AND THE PATIENT WAS REFERRED TO COMFORT CARE. IT WAS NOTED THAT THE PROCEDURE WAS DELAYED BY APPROXIMATELY 1 HOUR DUE TO THE DISLODGE. FOLLOWING THE IMPLANT PROCEDURE, THE PATIENT EXPERIENCED A STROKE AND DIED. PER THE PHYSICIAN, THE AORTIC DISSECTION AND STROKE WERE SUSPECTED TO HAVE BEEN CAUSED BY THE DISLODGED VALVE. Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME: DELIV SYS D-EVOLUTFX-2329, PRODUCT ID: D-EVOLUTFX-2329, SERIAL/LOT: UNKNOWN, USE BY DATE: UNKNOWN, UDI: UNKNOWN. THE CODES PRESENT IN SECTION H6 CORRESPOND TO MULTIPLE COMPONENTS/PRODUCTS THAT COMPRISE THIS REPORTED EVENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22819164&pc=NPT | 9617601-2025-01012 | 2025-07-24 04:00 | Death | MEDTRONIC MEXICO S. DE R.L. DE CV | 2025-08-15 04:00 | NPT | EVOLUT FX PLUS VALVE | Patient Device Interaction Problem | Stroke/CVA | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT PRIOR TO THE IMPLANT OF A TRANSCATHETER VALVE, A PRE-IMPLANT BALLOON AORTIC VALVULOPLASTY (BAV) WAS PERFORMED DUE TO A MEAN GRADIENT OF 66 MILLIMETERS OF MERCURY (MM HG). FOLLOWING THE IMPLANT OF THE VALVE, THE VALVE DISLODGED DUE TO INTERACTION WITH THE PIGTAIL CATHETER. PRIOR TO VALVE DISLODGEMENT, THE IMPLANT DEPTH ON THE NON-CORONARY CUSP (NCC) WAS 3-4 MILLIMETERS (MM), AND THE IMPLANT DEPTH ON THE LEFT CORONARY CUSP (LCC) WAS 5 MM. AFTER VALVE DISLODGEMENT, THE IMPLANT DEPTH ON THE NCC AND LCC WAS APPROXIMATELY 5 MM DISTAL TO THE INNOMINATE ARTERY. FOLLOWING DISLODGEMENT, AN AORTIC DISSECTION EXTENDING INTO THE RIGHT CAROTID ARTERY WAS OBSERVED. SUBSEQUENTLY, A SECOND TRANSCATHETER VALVE WAS IMPLANTED AND THE PATIENT WAS REFERRED TO COMFORT CARE. IT WAS NOTED THAT THE PROCEDURE WAS DELAYED BY APPROXIMATELY 1 HOUR DUE TO THE DISLODGE. FOLLOWING THE IMPLANT PROCEDURE, THE PATIENT EXPERIENCED A STROKE AND DIED. PER THE PHYSICIAN, THE AORTIC DISSECTION AND STROKE WERE SUSPECTED TO HAVE BEEN CAUSED BY THE DISLODGED VALVE. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID (D-EVOLUTFX-2329); PRODUCT LOT/SERIAL NUMBER (UNKNOWN); PRODUCT TYPE: (0195-HEARTVALVES). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22819262&pc=QZI | MW5174663 | 2025-07-23 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-08-15 04:00 | QZI | FARAWAVE CATHETER - FARAWAVE PULSED FIELD ABLATION (PFA) | Insufficient Information | Arrhythmia; Cardiac Arrest; Ventricular Fibrillation; Respiratory Arrest |  |  | 1 | Event Description: EVENT DESCRIPTION: IT WAS REPORTED BY THE CLINICAL ACCOUNT SPECIALIST THAT DURING A PVI CASE, A VENTRICULAR FIBRILLATION WAS NOTICED. THE CALLER STATED THE PATIENT WENT INTO TORSADES DE POINTES. THE CALLER STATED THEY WERE ABLE TO GET THE PATIENT INTO A STABLE RHYTHM BUT THE PATIENT ENDED UP EXPIRING. THE CALLER STATED A PFA CASE WAS PERFORMED WITH A BOSTON FARAWAVE CATHETER. THE CALLER STATED AFTER THE LAST ABLATION ANESTHESIA NOTICED THEY WERE GOING INTO VFIB. THE CALLER STATED THAT WAS WHEN THE INJURY WAS DISCOVERED. THE CALLER STATED THEY DID NOT KNOWN WHAT HAD CAUSED THE INJURY. THE CALLER STATED THE INJURY WAS CONFIRMED BY THE 12 LEAD ON THE RECORDING SYSTEM. THE CALLER STATED CPR AND DEFIBRILLATION WERE PERFORMED ON THE PATIENT. THE CALLER STATED A REPROCESSED STERILMED SOUNDSTAR CATHETER, A REPROCESSED STERILMED WEBSTER CS CATHETER AND A BRAND NEW OCTARAY CATHETER WERE USED DURING THE PROCEDURE. THE CALLER STATED ALL CATHETERS USED DURING THE PROCEDURE WERE NOT AVAILABLE FOR RETURN. THE CALLER STATED THE PROCEDURE WAS SUPPOSED TO BE A PVI AND WATCHMAN CONCOMITANT CASE. THE CALLER STATED AFTER THE VFIB WAS NOTICED THEY DID NOT PROCEED WITH THE WATCHMAN PART OF THE CASE. THIS REPORT REFLECTS INFORMATION RECEIVED BY FDA IN THE FORM OF A NOTIFICATION PER 803.22 (B)(2).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22819268&pc=NPT | 9617601-2025-01014 | 2024-07-25 04:00 | Death | MEDTRONIC MEXICO S. DE R.L. DE CV | 2025-08-15 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Insufficient Information | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING THE IMPLANT OF A TRANSCATHETER VALVE, THE PATIENT DIED. THE CAUSE OF DEATH WAS UNKNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. B2: THE PATIENT'S DATE OF DEATH IS UNKNOWN. B3: EVENT DATE IS NOT KNOWN. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22819444&pc=NKM | 2135147-2025-04581 | 2025-07-25 04:00 | Death | ABBOTT MEDICAL | 2025-08-15 04:00 | NKM | MITRACLIP® SYSTEM CLIP DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Fistula; Hematoma; Hemorrhage/Bleeding; Unspecified Infection; Myocardial Infarction; Perforation; Pseudoaneurysm; Vascular Dissection; Embolism/Embolus | P100009 |  | 1 | Event Description: THE ARTICLE ¿COMPARISON OF SAFETY AND EFFICACY IN VENOUS ACCESS CLOSURE USING A DOUBLE PURSE STRING SUTURE TECHNIQUE VS. Z-SUTURE TECHNIQUE AFTER MITRACLIP PROCEDURE¿ WAS REVIEWED. THIS RESEARCH ARTICLE IS A RETROSPECTIVE SINGLE CENTER EXPERIENCE TO COMPARE ACCESS CLOSURE USING SUBCUTANEOUS ABSORBABLE DOUBLE PURSE STRING SUTURE (DPSS) AND Z-SUTURE TECHNIQUE EVALUATING VASCULAR COMPLICATIONS AND BLEEDING FOLLOWING MITRACLIP PROCEDURE. VENOUS CLOSURE WAS PERFORMED USING Z-SUTURE TECHNIQUE IN 140 PATIENTS AND DPSS TECHNIQUE IN 109 PATIENTS. VASCULAR COMPLICATIONS AND BLEEDING EVENTS WERE ASSESSED ACCORDING TO THE MITRAL VALVE ACADEMIC RESEARCH CONSORTIUM (MVARC) CRITERIA. BETWEEN JANUARY 2014 AND DECEMBER 2017, 249 PATIENTS UNDERWENT MITRACLIP PROCEDURE. THE ARTICLE CONCLUDED THAT LARGE CALIBER VENOUS ACCESS CLOSURE WITH DPSS TECHNIQUE WAS FEASIBLE, SAFE, AND EFFECTIVE TO ACHIEVE HAEMOSTASIS AFTER MITRACLIP PROCEDURE. COMPARED WITH Z-SUTURE, USE OF DPSS CLOSURE WAS ASSOCIATED WITH A LOWER RATE OF REQUIRED ACCESS RELATED SURGICAL INTERVENTION AND POSTINTERVENTIONAL AV-FISTULA. (B)(6). BETWEEN JANUARY 2014 AND DECEMBER 2017, 249 PATIENTS UNDERWENT MITRACLIP PROCEDURE. THE MEAN AGE WAS 76 YEARS AND 66% WERE MALE. AS THIS IS FROM A LITERATURE REVIEW, PATIENT WEIGHT WAS NOT PROVIDED. COMORBIDITIES INCLUDED: DEGENERATIVE MITRAL REGURGITATION, FUNCTIONAL MITRAL REGURGITATION, TRICUSPID REGURGITATION, HYPERTENSION, DIABETES MELLITUS, CHRONIC KIDNEY DISEASE, END-STAGE-RENAL DISEASE, CORONARY ARTERY DISEASE, DILATATIVE CARDIOMYOPATHY, PREVIOUS MYOCARDIAL INFARCTION (MI), PREVIOUS CORONARY ARTERY BYPASS GRAFTING (CABG), PREVIOUS STROKE, CHRONIC OBSTRUCTIVE PULMONARY DISEASE (COPD), ATRIAL FIBRILLATION, PACEMAKER/IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD), AND CARDIAC RESYNCHRONIZATION THERAPY (ICD). ADVERSE EVENTS INCLUDED: HOSPITALIZATION, VASCULAR OBSTRUCTION, VASCULAR DISSECTION, VASCULAR PERFORATION, ARTERIO AV FISTULA, PSEUDOANEURYSMS, SUPERFICIAL HEMATOMA (LARGE AND SMALL), MVARC MINOR BLEEDING, MVARC MAJOR, NON-LIFE-THREATENING BLEEDING, LIFE THREATENING/DISABLING BLEEDING, TRANSFUSION, ACCESS RELATED BLOOD TRANSFUSIONS, VASCULAR SURGERY, ENDOVASCULAR INTERVENTION, ACCESS SITE INFECTION, IN-HOSPITAL MORTALITY, IN-HOSPITAL MAJOR ADVERSE CARDIAC AND CEREBROVASCULAR EVENTS (MACCE) (MYOCARDIAL INFARCTION, STROKE, DEATH). Manufacturer Narrative: IT IS UNKNOWN IF THE DEVICE IS RETURNING FOR ANALYSIS. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22819481&pc=NGV | 2135147-2025-04583 | 2025-07-23 04:00 | Death | ABBOTT MEDICAL | 2025-08-15 04:00 | NGV | AMPLATZER AMULET | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P200049 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 28MM AMPLATZER AMULET WAS SELECTED FOR IMPLANT IN A PATIENT WITH A PAST MEDICAL HISTORY OF LOW EF, HIGH BP, PREVIOUS BI-V ICD. PRIOR TO IMPLANT, THE PATIENT WAS IN STABLE CONDITION. THE LANDING ZONE MEASUREMENTS WERE 0- 22, 45-22,90- 24, 135-22. ANGIO- 24. THE SIZING OF THE DEVICE WAS DETERMINED VIA TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE). DURING IMPLANT, ORIENTATION WAS DIFFICULT TO ACHIEVE; 2 PARTIAL RECAPTURES WERE REPORTED. HOWEVER, CLOSE CRITERIA WAS ACHIEVED AND THE AMULET WAS IMPLANTED. POST-IMPLANT, AFTER THE PATIENT WAS DISCHARGED HOME, THE PATIENT PASSED DUE TO PULSELESS ELECTRICAL ACTIVITY (PEA). Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22820684&pc=DSQ | 2916596-2025-05331 | 2024-11-10 05:00 | Death | THORATEC CORPORATION | 2025-08-15 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Renal Failure; Septic Shock; Respiratory Failure; Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED A COMPLEX MEDICAL COURSE FOLLOWING INITIAL IMPLANT. THE PATIENT DEVELOPED MULTI-SYSTEM ORGAN FAILURE, SUFFERED FROM RESPIRATORY AND RENAL FAILURE, AND THEN WENT INTO SEPTIC SHOCK. THE PATIENT WAS PUT ON COMFORT MEASURES AND THEIR LEFT VENTRICULAR ASSIST DEVICE (LVAD) WAS TURNED OFF. THE PATIENT PASSED AWAY DUE TO MULTI-SYSTEM ORGAN FAILURE, RESPIRATORY AND RENAL FAILURE, AND SEPTIC SHOCK. THE OUTCOME WAS CONSIDERED DEVICE OR DEVICE THERAPY RELATED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST DEVICE (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. NO AUTOPSY WAS PERFORMED, AND THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, REV. D IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION, INFECTION, AND DEATH AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6, ¿PATIENT CARE AND MANAGEMENT,¿ LISTS INFECTION AS A POTENTIAL LATE POST-IMPLANT COMPLICATION THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. CARE INSTRUCTIONS IN REGARD TO PREVENTING INFECTION ARE PROVIDED IN VARIOUS SECTIONS OF THE INSTRUCTIONS FOR USE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22821410&pc=PNJ | 2017865-2025-97860 | 2025-07-26 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-08-15 04:00 | PNJ | AVEIR¿ LEADLESS PACEMAKER, RIGHT VENTRICLE | Device Dislodged or Dislocated | Embolism/Embolus | P150035 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT'S RIGHT VENTRICULAR LEADLESS PACEMAKER (LP) DISLODGED AND TRAVELED TO THE CORONARY SINUS. THE LP WAS EXPLANTED. THE PATIENT WAS NOTED TO DECEASE POST-PROCEDURE. THERE WERE NO DIRECT ALLEGATIONS FROM THE PHYSICIAN THAT THE PATIENT'S DEATH WAS CAUSED BY THEIR LP SYSTEM.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22821894&pc=DSQ | 2916596-2025-05333 | 2025-08-06 04:00 | Death | THORATEC CORPORATION | 2025-08-15 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Arrhythmia; Cardiac Arrest; Dyspnea; Loss of consciousness; Gastrointestinal Hemorrhage | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD SUDDEN ONSET DYSPNEA WITH ASSOCIATED CONTINUOUS LOW FLOW ALARMS. THE PATIENT THEN FELL UNCONSCIOUS, CARDIAC ARRESTED, AND WAS NEVER ABLE TO BE RESUSCITATED AT THE OUTSIDE HOSPITAL. THE PATIENT PASSED AWAY ON (B)(6) 2025 DUE TO CARDIAC ARREST. THEY HAD A SIGNIFICANT HISTORY OF GASTROINTESTINAL BLEED (GIB) AND ARRHYTHMIAS. IT WAS UNKNOWN IF THE DEATH WAS DEVICE RELATED AS THE DEVICE WAS NOT ABLE TO BE RETRIEVED DUE TO NO AUTOPSY PERFORMED. THE DEVICE OPERATED AS EXPECTED TO THE BEST OF THEIR KNOWLEDGE. LOG FILES WERE SENT FOR REVIEW TO INVESTIGATION FOR PUMP MALFUNCTION, ALTHOUGH SUSPICION WAS LOW. FOLLOWING REVIEW, IT APPEARED THAT THE LOG BEGAN CAPTURING LOW FLOWS IN THE 2.4 LITERS PER MINUTE (LPM) RANGE BETWEEN 11:03 PM AND 11:11 PM ON (B)(6) 2025, AT 12:11 AM ON (B)(6) 2025, THE FLOWS BEGAN DROPPING BELOW 1 LPM AND CONTINUED TO DROP UNTIL THE PUMP WAS TURNED OFF AT 1:08 AM. THERE DID NOT APPEAR TO BE ANY TYPE OF MECHANICAL CIRCULATORY SUPPORT (MCS) EQUIPMENT ISSUES CAUSING THE EVENTS. THE LOW FLOW EVENTS APPEARED TO HAVE BEEN A PATIENT RELATED ISSUE, NOT AN EQUIPMENT ISSUE. THERE WERE NO OTHER UNUSUAL EVENTS RECORDED IN THE LOG FILE EVENT HISTORY. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22822313&pc=GXI | 3006630150-2025-06507 | 2025-07-23 04:00 | Death | RELIEVANT MEDSYSTEMS, INC. | 2025-08-15 04:00 | GXI | INTRACEPT RF PROBE | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K180369 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN INTRACEPT PROCEDURE AND A FIFTEEN MINUTE ABLATION WAS PERFORMED WITHOUT COMPLICATION. HOWEVER, WHEN A SECOND ABLATION WAS INITIATED, THE ANESTHESIA TEAM IDENTIFIED COMPLICATIONS AND REQUESTED THE PROCEDURE BE STOPPED. THE ABLATION WAS IMMEDIATELY HALTED AND THE ANESTHESIA TEAM INITIATED RESUSCITATIVE EFFORTS. THE EMERGENCY MEDICAL SERVICES (EMS) ARRIVED AND TRANSPORTED THE PATIENT TO THE HOSPITAL. THE HOSPITAL CALLED TIME OF DEATH UPON THE PATIENTS ARRIVAL. THE PHYSICIAN ASSESSED THAT NEITHER THE DEVICE NOR THE PROCEDURE CONTRIBUTED TO THE PATIENT COMPLICATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22822397&pc=NAY | 2955842-2025-34595 | 2025-07-19 04:00 | Death | INTUITIVE SURGICAL, INC | 2025-08-15 04:00 | NAY | DAVINCI XI | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Insufficient Information | K131861 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED PARTIAL NEPHRECTOMY PROCEDURE, THE PATIENT EXPERIENCED BLEEDING, AND THE PROCEDURE APPROACH WAS CONVERTED TO OPEN. THE CAUSE OF THE BLEEDING, THE ESTIMATED BLOOD LOSS, AND THE SPECIFIC INTERVENTIONS PERFORMED TO CONTROL THE BLEEDING REMAIN UNKNOWN. FOLLOWING THE PROCEDURE, THE PATIENT WAS TRANSFERRED TO THE ICU AND SUBSEQUENTLY EXPIRED. ADDITIONAL INFORMATION HAS BEEN REQUESTED; HOWEVER, NO RESPONSE HAS BEEN RECEIVED. Manufacturer Narrative: A REVIEW OF THE SYSTEM LOGS FOR THE REPORTED PROCEDURE DATE FOUND THAT NO SYSTEM ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. A DEVICE HISTORY RECORD (DHR) REVIEW FOUND NO NON-CONFORMANCES DOCUMENTED THAT WOULD BE RELATED TO THIS EVENT. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: 0 DEGREE ENDOSCOPE, MONOPOLAR CURVED SCISSORS, 2 LARGE NEEDLE DRIVERS, PROGRASP FORCEPS, MARYLAND BIPOLAR FORCEPS. A SITE HISTORY REVIEW FOUND NO COMPLAINTS WERE MADE FOR THE INSTRUMENTS USED DURING THIS PROCEDURE. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT DIED FOLLOWING A PARTIAL NEPHRECTOMY. HOW THE BLEEDING OCCURRED IS NOT PROVIDED. THE FINDINGS UPON OPENING ARE NOT PROVIDED NOR ARE THE EVENTS OR DETAILS OF THE POST OPERATIVE COURSE FOLLOWING THE PROCEDURE BUT THE PATIENT SUBSEQUENTLY DIED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT. .
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22822398&pc=NAY | 2955842-2025-34593 | 2025-07-21 04:00 | Death | INTUITIVE SURGICAL, INC | 2025-08-15 04:00 | NAY | DAVINCI XI | Insufficient Information | Hemorrhage/Bleeding; Insufficient Information | K131861 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A ROBOTIC ASSISTED RIGHT LOWER LOBECTOMY WITH MEDIASTINAL LYMPHADENECTOMY PROCEDURE, THERE WAS UNEXPECTED BLEEDING AND BLOOD TRANSFUSIONS WERE ADMINISTERED. THE HOSPITAL PATIENT SAFETY MANAGER REPORTED THAT THE PATIENT EXPIRED THE SAME DAY AND AN AUTOPSY IS IN PROGRESS. ATTEMPTS TO OBTAIN FURTHER INFORMATION, INCLUDING THE SURGEON¿S NAME, FROM SEVERAL SITE SOURCES WERE UNSUCCESSFUL UNTIL (B)(6) 2025. ADDITIONAL INFORMATION HAS BEEN REQUESTED FROM THE SURGEON. Manufacturer Narrative: A REVIEW OF THE SYSTEM LOGS FOUND NO ERRORS OCCURRED DURING THIS PROCEDURE. REVIEW OF THE LOGS FOR THE 5 SUBSEQUENT PROCEDURES FOUND THE SYSTEM FUNCTIONED NORMALLY WITH NO INTRAOPERATIVE EVENTS OR ISSUES FOUND. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: 0 DEGREE ENDOSCOPE PLUS, LONG BIPOLAR GRASPER, TWO TIP-UP FENESTRATED GRASPERS, CADIERE FORCEPS, VESSEL SEALER EXTEND, CURVED-TIP SUREFORM 45 STAPLER, SUCTION IRRIGATOR, LARGE CLIP APPLIER. A SITE HISTORY REVIEW FOUND NO RELATED COMPLAINTS WERE RECEIVED FOR THE INSTRUMENTS USED DURING THIS PROCEDURE. A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED FOR THE DEVICE(S) USED DURING THE PROCEDURE AND NO NON-CONFORMANCES WERE IDENTIFIED TO BE RELATED TO THIS EVENT. THE LOGS FOR THE CURVED-TIP SUREFORM 45 STAPLER SHOWED THE INSTRUMENT WAS INSTALLED 10 TIMES AND FIRED 10 RELOADS (1 WHITE, 1 GREEN, 2 WHITE, 6 BLACK RESPECTIVELY). ALL FIRINGS WERE COMPLETED PER THE LOGS, ALL BUT 3 FIRINGS HAD NO PAUSES DURING FIRING. THERE WERE NO INCOMPLETE CLAMPS OR UNCLAMP FAILURES BY THIS INSTRUMENT. THERE ARE NO STAPLER RELATED ERRORS IN THE LOGS. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT DIED DURING AN ATTEMPTED PULMONARY LOBECTOMY DUE TO BLEEDING. HOW THE BLEEDING OCCURRED IS NOT PROVIDED. NO ADDITIONAL INFORMATION IS PROVIDED REGARDING ANY ATTEMPTS TO CORRECT THE BLEEDING. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT. BLANK MDR FIELDS: THE MISSING PATIENT INFORMATION IN SECTIONS A AND B WAS EITHER UNKNOWN, UNAVAILABLE, NOT PROVIDED, OR NOT APPLICABLE. THE EXPIRATION DATE FOR SECTION D4 IS NOT APPLICABLE. FIELD D6 IS BLANK BECAUSE THE PRODUCT IS NOT IMPLANTABLE. INFORMATION FOR THE BLANK FIELDS IN SECTION E1 IS NOT AVAILABLE. FIELD E4 IS BLANK BECAUSE IT IS UNKNOWN IF THE INITIAL REPORTER SUBMITTED A REPORT TO THE FDA. FIELDS G5 AND G7 ARE NOT APPLICABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22801637&pc=EOQ | 2955842-2025-33717 | 2025-07-17 04:00 | Death | INTUITIVE SURGICAL, INC | 2025-08-14 04:00 | EOQ | ION | Adverse Event Without Identified Device or Use Problem | Bradycardia; Cardiac Arrest; Hemorrhage/Bleeding | K182188 |  | 1 | Event Description: IT WAS REPORTED THAT DURING AN ION ENDOLUMINAL LUNG BIOPSY PROCEDURE, THE PATIENT EXPERIENCED BLEEDING. THE ENDOLUMINAL SALES ASSOCIATE WAS PRESENT AT THE PROCEDURE AND REPORTED THAT THERE WERE NO OBSERVED COMPLICATIONS BEFORE OR DURING THE ION PART OF THE PROCEDURE. AFTER THE ION PORTION OF THE BIOPSY WAS COMPLETED, THE PHYSICIAN PULLED OUT THE CATHETER, PLACED A REGULAR BRONCHOSCOPE FOR SURVEILLANCE OF THE AIRWAYS, AND NOTICED SOME MINOR BLEEDING. THE BLEEDING WAS CONTROLLED WITH SUCTIONING AND TILTING THE PATIENT'S HEAD DOWN. SHORTLY AFTER, THE PATIENT¿S HEARTRATE STARTED TO DROP AND THE MEDICAL STAFF INITIATED CARDIOPULMONARY RESUSCITATION. AFTER SEVERAL MINUTES, THE PULSE RETURNED BUT ONLY FOR A FEW MINUTES AND CPR WAS RESUMED. THE PATIENT EXPIRED. MULTIPLE ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION FROM THE PHYSICIAN HAVE BEEN MADE; HOWEVER, NO RESPONSE WAS RECEIVED. Manufacturer Narrative: A SYSTEM LOG REVIEW CANNOT BE PERFORMED BECAUSE THE SYSTEM LOGS ARE NOT AVAILABLE. REVIEW OF THE DEVICE HISTORY RECORD FOUND THAT THERE WERE NO NON-CONFORMANCE'S DOCUMENTED THAT WOULD BE RELATED TO THE EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT UNDERWENT AN ION BRONCHOSCOPY. THE BIOPSY WAS COMPLETED, AND THE ION CATHETER WAS REMOVED. A MANUAL BRONCHOSCOPY WAS THEN INSERTED AND MINOR BLEEDING WAS NOTED. BLEEDING WAS CONTROLLED WITH SUCTIONING AND POSITIONING OF THE PATIENT'S HEAD. A WEAK PULSE WAS NOTED PROMPTING INITIATION OF CPR. AFTER SEVERAL MINUTES, PULSE RETURNED, TRANSIENTLY FOR A FEW MINUTES AT WHICH TIME CPR WAS AGAIN ADMINISTERED, HOWEVER, THE PATIENT DIED. THERE WAS NO ALLEGATION OF A MALFUNCTION OF THE ION SYSTEM, INSTRUMENTS, OR ACCESSORIES. BASED ON THE AVAILABLE DATA, THE REPORTED EVENTS WERE LIKELY PROCEDURE RELATED AND NOT DEVICE RELATED. BRONCHOSCOPY IS A MINIMALLY INVASIVE PROCEDURE WITH A LOW RISK PROFILE. A RETROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES REPORTED 4 ASSOCIATED DEATHS (0.02%), 5 ARRHYTHMIAS (0.02%) AND 1 MYOCARDIAL INFARCTION (0.004%). ONE PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 REPORTED 1 ASSOCIATED DEATH (0.08%) AND NO ASSOCIATED EVENTS OF MYOCARDIAL INFARCTION OR ARRHYTHMIAS. A RECENT META-ANALYSIS OF NAVIGATIONAL BRONCHOSCOPY IN 10,381 PATIENTS REPORTED AN OVERALL ADVERSE EVENT RATE OF 5.6% INCLUDING A RATE OF 0.02% OF ANY ARRHYTHMIA AND 1 DEATH. ANOTHER CASE SERIES OF ION ROBOTIC ASSISTED BRONCHOSCOPIES PUBLISHED AFTER THE META-ANALYSIS INCLUDING 415 CASES REPORTED NO CARDIAC EVENTS. FACCIOLONGO N, PATELLI M, GASPARINI S, ET AL. INCIDENCE OF COMPLICATIONS IN BRONCHOSCOPY. MULTICENTRE PROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES. MONALDI ARCHIVES FOR CHEST DISEASE. 2009. FOLCH EE, PRITCHETT MA, NEAD MA, ET AL. ELECTROMAGNETIC NAVIGATION BRONCHOSCOPY FOR PERIPHERAL PULMONARY LESIONS: ONE-YEAR RESULTS OF THE PROSPECTIVE, MULTICENTER NAVIGATE STUDY. JOURNAL OF THORACIC ONCOLOGY. 2019. KOPS SEP, HEUS P, KOREVAAR DA, ET AL. DIAGNOSTIC YIELD AND SAFETY OF NAVIGATION BRONCHOSCOPY: A SYSTEMATIC REVIEW AND META-ANALYSIS. LUNG CANCER. 2023. BROWNLEE AR, WATSON JJJ, AKHMEROV A, ET AL. ROBOTIC NAVIGATIONAL BRONCHOSCOPY IN A THORACIC SURGICAL PRACTICE: LEVERAGING TECHNOLOGY IN THE MANAGEMENT OF PULMONARY NODULES. JTCVS. 2023.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22801794&pc=OYC | 2032227-2025-235321 | 2025-03-12 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-08-14 04:00 | OYC | PARADIGM REAL-TIME REVEL INSULIN INFUSION PUMP MMT-523NAH | Adverse Event Without Identified Device or Use Problem | Dementia; Heart Failure/Congestive Heart Failure | P150019 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER WAS PASSED AWAY DUE TO HEART DISEASE AND VASCULAR DEMENTIA IN HOME ON (B)(6) 2025. THE EVENT INVOLVED PRODUCT(S) MMT-399A, MMT-332A, MMT-1880, MMT-523NAH. TROUBLESHOOTING WAS PERFORMED TO GET THE CUSTOMER DECEASED INFORMATION FROM HUSBAND. PUMP WAS NOT WORN AT THE TIME OF PASSING. MMT-1880 WAS REQUESTED AND RETURNED FOR PRODUCT ANALYSIS. MMT-523NAH WAS REQUESTED AND RETURNED FOR PRODUCT ANALYSIS. NO PRODUCT RETURN REQUIRED FOR MMT-332A. NO PRODUCT RETURN REQUIRED FOR MMT-399A. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. THE PUMP PASSED THE FUNCTIONAL TEST, INCLUDING THE SELF TEST, SLEEP CURRENT MEASUREMENT, ACTIVE CURRENT MEASUREMENT, REWIND TEST, PRIME/SEATING TEST, BASIC OCCLUSION TEST, OCCLUSION TEST, FORCE SENSOR TEST, DISPLACEMENT TEST, AND THE DAT AT 0.0871 INCHES. THE FOLLOWING WERE NOTED DURING VISUAL INSPECTION: MINOR SCRATCHED DISPLAY WINDOW, SCRATCHED CASE, SCRATCHED KEYPAD OVERLAY AND PILLOWING KEYPAD OVERLAY. TEST P-CAP AND RESERVOIR LOCKED PROPERLY INTO RESERVOIR COMPARTMENT DURING TESTING. HISTORY DOWNLOAD WAS SUCCESSFUL USING THUMP AND CARELINK UPLOAD WAS SUCCESSFUL. THE PUMP WAS RECEIVED WITHOUT A BATTERY. THE BATTERY CAP WAS RECEIVED WITH SLIGHT DAMAGED (1 BROKEN HEAT STAKE POST). UNABLE TO VERIFY DAILY TOTAL OF ALL INSULIN DELIVERED, DAILY TOTAL OF BASAL INSULIN DELIVERED AND DAILY TOTAL OF BOLUS INSULIN DELIVERED FOR EVENT DATE 12-MAR-2025 DUE TO NO DATA AVAILABLE IN THE PUMP HISTORY FILE ON THE PRIMARY SVN (B)(6). THERE WAS DATA LISTED ON (B)(6) 2025 IN THE PUMP HISTORY FILE. NO DATA AVAILABLE ON 12-MAR-2025 IN THE PUMP HISTORY FILE. PERFORM THE HISTORY REVIEW 1 WEEK PRIOR TO THE EVENT DATE 12-MAR-2025 IN THE PUMP HISTORY FILE AND FOUND 2 PUMP ERROR 23 ON (B)(6) 2025 AND 4 BATTOUTLIMIT (6) ON (B)(6) 2025 ON THE PRIMARY SVN (B)(6). PUMP ERROR 23 AND BATTERY OUT LIMIT ALARM WERE EXPECTED DUE TO THE BATTERY REMOVED FOR MORE THAN 10 MINUTES AND THE PUMP RESET. THE POWER MANAGEMENT TOOL CONFIRMED THE UNLOADED VOLTAGE (UL VLITH) AND LOADED VOLTAGE (LOADED VLITH) WERE WITHIN SPEC RANGE. PUMP WAS CUT OPEN TO PERFORM VISUAL INSPECTION AND FOUND NO EVIDENCE OF PHYSICAL OR MOISTURE DAMAGE ON THE PCBA1, PCBA2, FORCE SENSOR, MOTOR AND VIBRATOR ASSEMBLY NOTED. FORCE SENSOR ZERO OFFSET WITHIN SPECIFICATION (23.6 MV). THE PUMP PASSED THE FUNCTIONAL TESTING. FOR ADDITIONAL INFORMATION REGARDING THE TESTS PERFORMED REFER TO (B)(4) ¿ APPENDIX E. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22801796&pc=OZP | 2032227-2025-235323 | 2023-01-30 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-08-14 04:00 | OZP | MINIMED 770G US SYSTEM BLE CONNECT 3.0 MG/DL | Adverse Event Without Identified Device or Use Problem | Hyperglycemia; Hypoglycemia | P150001 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER HAD PASSED AWAY ON (B)(6) 2023, AND THE CAUSE OF DEATH WAS DIABETES-RELATED. THE CUSTOMER REPORTED THE HYPERGLYCEMIA AND HYPOGLYCEMIA. THE EVENT INVOLVED PRODUCTS UNOMEDICAL, MMT-332A, AND MMT-1880. TROUBLESHOOTING WAS PARTIALLY PERFORMED, AND IT WAS UNKNOWN WHETHER THE PUMP WAS WORN AT THE TIME OF PASSING OR NOT. NO PRODUCT RETURN IS REQUIRED FOR UNOMEDICAL, MMT-332A, AND MMT-1880. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22803129&pc=NKE | 3004209178-2025-13804 | 2025-07-17 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-08-14 04:00 | NKE | PERCEPTA QUAD CRT-P MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Erosion; Purulent Discharge; Erythema; Pain; Discomfort; Swelling/ Edema | P010015 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DEVELOPED BACTERIAL INFECTION ALONG WITH PROTRUSION NOTED AT THE IMPLANT SITE RESULTING IN REDNESS, SWELLING, PAIN, DISCOMFORT AND PUS DISCHARGE. THE ORGANISM WAS IDENTIFIED TO BE KLEBSIELLA. THE CARDIAC RESYNCHRONIZATION THERAPY P ACEMAKER (CRT-P) SYSTEM WAS EXPLANTED. ONE WEEK LATER, THE PATIENT PASSED AWAY.Event Description: IT WAS NOTED THAT THE PATIENT DEATH WAS NOT NOTED TO THE PERFORMANCE OF THE CRT-P. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22803133&pc=OJX | 2649622-2025-22050 | 2025-07-17 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-08-14 04:00 | OJX | ATTAIN PERFORMA S MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Erosion; Purulent Discharge; Erythema; Pain; Discomfort; Swelling/ Edema | P080006 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DEVELOPED BACTERIAL INFECTION ALONG WITH PROTRUSION NOTED AT THE IMPLANT SITE RESULTING IN REDNESS, SWELLING, PAIN, DISCOMFORT AND PUS DISCHARGE. THE ORGANISM WAS IDENTIFIED TO BE KLEBSIELLA. THE CARDIAC RESYNCHRONIZATION THERAPY PACEMAKER (CRT-P) SYSTEM WAS EXPLANTED. ONE WEEK LATER, THE PATIENT PASSED AWAY.Event Description: IT WAS NOTED THAT THE PATIENT DEATH WAS NOT NOTED TO THE PERFORMANCE OF THE CRT-P. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22803143&pc=DTB | 2649622-2025-22052 | 2025-07-17 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-08-14 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Erosion; Purulent Discharge; Erythema; Pain; Discomfort; Swelling/ Edema | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DEVELOPED BACTERIAL INFECTION ALONG WITH PROTRUSION NOTED AT THE IMPLANT SITE RESULTING IN REDNESS, SWELLING, PAIN, DISCOMFORT AND PUS DISCHARGE. THE ORGANISM WAS IDENTIFIED TO BE KLEBSIELLA. THE CARDIAC RESYNCHRONIZATION THERAPY PACEMAKER (CRT-P) SYSTEM WAS EXPLANTED. ONE WEEK LATER, THE PATIENT PASSED AWAY.Event Description: IT WAS NOTED THAT THE PATIENT DEATH WAS NOT NOTED TO THE PERFORMANCE OF THE CRT-P. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22803860&pc=OZP | 2032227-2025-236084 | 2025-05-19 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-08-14 04:00 | OZP | MINIMED 780G US SYSTEM BLE CONNECT 3.0 MG/DL | Adverse Event Without Identified Device or Use Problem | Hyperglycemia; Hypoglycemia; Pain; Heart Failure/Congestive Heart Failure | P160017 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY IN AN AMBULANCE ON (B)(6) 2025. THE CAUSE OF DEATH WAS CONGESTIVE HEART FAILURE AND DIABETES AND ALSO FELT PAIN IN THE HEAD. THE CUSTOMER¿S BLOOD GLUCOSE VALUE AT THE TIME OF THE INCIDENT IS UNKNOWN. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: UNOMEDICAL, MMT-7841ZN, MMT-7040A, MMT-1884, MMT-342G. TROUBLESHOOTING WAS PARTIALLY PERFORMED FOR DECEASED REPORTING. THE CUSTOMER WAS WEARING AN INSULIN PUMP AND SENSOR AT THE TIME OF PASSING. NO PRODUCT RETURN IS REQUIRED FOR UNOMEDICAL, MMT-7841ZN, MMT-7040A, MMT-1884, MMT-342G. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22804306&pc=DQY | 2124215-2025-56360 | 2025-07-22 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-08-14 04:00 | DQY | WATCHMAN FXD CURVE? ACCESS SYSTEM | Adverse Event Without Identified Device or Use Problem | Hematoma; Cardiac Tamponade; Cardiac Perforation; Pericardial Effusion | K212228 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT DEATH OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED AND A WATCHMAN FXD CURVE ACCESS SYSTEM (WAS) WAS SELECTED TO BE USED. DURING THE PROCEDURE THE PHYSICIAN STRUGGLED WITH TRANSEPTAL PUNCTURE (TSP). WITHOUT KNOWING THE EXACT LOCATION ON THE SEPTUM, THE PHYSICIAN CAUTERIZED FOR A FEW SECONDS. AN INTRACARDIAC ECHOCARDIOGRAPHY (ICE) PROBE WAS PLACED AND INTERVENTIONAL CARDIOLOGY (IC) SUPPORTED WITH TRANSESOPHAGEAL ECHOCARDIOGRAPHY (TEE) AND SUCCESSFULLY CROSSED. A PERICARDIAL EFFUSION WAS NOTED AND THE PROCEDURE WAS ABORTED. PERICARDIOCENTESIS WAS PERFORMED. THE PATIENT WENT TO THE INTENSIVE CARE UNIT AND WAS DISCHARGED ON (B)(6) 2025. THE PATIENT WENT BACK TO THE HOSPITAL AND AN OPEN HEARTH SURGERY WAS PERFORMED, BLOOD AND HEMATOMA WERE DRAINED FROM CHEST AND PERICARDIAL SAC. THE PATIENT NEVER RECOVERED AND PASSED AWAY ON (B)(6) 2025.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22804383&pc=BZD | 2518422-2025-109069 | 2024-11-05 05:00 | Death | RESPIRONICS, INC. | 2025-08-14 04:00 | BZD | DREAMSTATION AUTO CPAP | Insufficient Information | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT HAS PASSED. MEDICAL INTERVENTION WAS NOT SPECIFIED. THERE IS NO ALLEGATION THAT THE DEATH IS RELATED TO THE DEVICE. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22804809&pc=DQA | 2936999-2025-00422 | 2024-12-01 05:00 | Death | COVIDIEN LLC | 2025-08-14 04:00 | DQA | NELLCOR | Device Alarm System | Low Oxygen Saturation | K142865 |  | 1 | Event Description: IT WAS REPORTED THAT DURING USE OF THE MONITOR, A DELAYED ALARM NOTIFICATION AND DEVICE ALARM SYSTEM ISSUE OCCURRED WHILE MONITORING THE PATIENT WHO WAS RECEIVING A FULL-TIME CARE AT HOME. THE DEVICE WAS CONFIGURED TO ALERT FOR OXYGEN DESATURATION AND BRADYCARDIA. ON THE DATE OF THE EVENT, THE PATIENT EXPERIENCED BRADYCARDIA, CARDIAC ARREST, LOW OXYGEN SATURATION, RESPIRATORY ARREST, AND DEATH. IT WAS ALLEGED THAT THE DELAYED ALARM NOTIFICATION MAY HAVE AFFECTED THE ABILITY TO RESPOND QUICKLY DURING A CRITICAL DESATURATION EVENT, POTENTIALLY IMPACTING THE OPPORTUNITY TO STIMULATE AND RESUSCITATE THE PATIENT BEFORE REACHING AN UNRECOVERABLE STATE. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22804867&pc=CAW | 2518422-2025-048893 | 2021-02-25 05:00 | Death | RESPIRONICS, INC. | 2025-08-14 04:00 | CAW | SIMPLYGO | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K111885 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A DEATH INVOLVING A PHILIPS DEVICE. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT PASSED AWAY IN NOVEMBER AND THEY WOULD LIKE TO RETURN THE DEVICE AND GET PARTIAL CREDIT DUE. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE MANUFACTURER IDENTIFIED UNIT HAD 103 HOURS UPON ARRIVAL, REPLACED FRONT COVER DUE TO MECHANICAL CRACK, REPLACED BATTERY DUE TO PRECAUTION, REPLACED SIEVE CANISTER PER PM, REPLACED INLET FILTER PER PM. RUN-IN TESTING PASSED. ALL FUNCTIONAL AND FINAL TESTS PASSED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22805378&pc=MKJ | 1220908-2025-02989 | 2025-07-23 04:00 | Death | ZOLL MEDICAL CORPORATION | 2025-08-14 04:00 | MKJ | X SERIES | Failure to Deliver Shock/Stimulation | Necrosis | P160022 |  | 1 | Event Description: COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A 77-YEAR-OLD MALE PATIENT, THE DEVICE FAILED TO DISCHARGE. COMPLAINANT INDICATED THAT THE PATIENT SUBSEQUENTLY EXPIRED. Manufacturer Narrative: THE DEVICE WAS RETURNED TO ZOLL MEDICAL CORPORATION FOR EVALUATION. THE DEVICE PASSED ALL FUNCTIONAL TESTING AND WAS ABLE TO CHARGE AND DISCHARGE WITHOUT ISSUES WHEN USING TEST PADS. THE RETURNED ELECTRODES COULD NOT BE FULLY EVALUATED, AS ONE PAD WAS CUT OFF. SIGNIFICANT HAIR WAS NOTED ON THE CPR PUCK, THOUGH NO HAIR WAS PRESENT ON THE RETURNED PAD. LOG REVIEW SHOWED MULTIPLE "PADS OFF" AND "CABLE ID CHANGE" MESSAGES SUGGESTING EITHER POOR COUPLING OR AN INTERMITTENT CONNECTION TO THE DEVICE. THE LOG ALSO SUPPORTS THAT WHEN THE APPROPRIATE CRITERIA WAS MET, THE DEVICE DISCHARGED SUCCESSFULLY. OUR EVALUATION YIELDED NO FAULT WAS FOUND WITH THE DEVICE. THE DEVICE WAS RECERTIFIED AND RETURNED TO THE CUSTOMER. NO TREND IS ASSOCIATED WITH REPORTS OF THIS TYPE. Manufacturer Narrative: ZOLL MEDICAL CORPORATION HAS RECEIVED THE PRODUCT AND WILL BE PROVIDING A SUPPLEMENTAL REPORT WHEN OUR INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22805630&pc=DSQ | 2916596-2025-05272 | 2025-07-27 04:00 | Death | THORATEC CORPORATION | 2025-08-14 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Paralysis | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT CALLED THEIR PHYSICIAN TO INFORM ABOUT LOW FLOW ALARM. THE PATIENT WAS ASKED ABOUT HYDRATIONS AND BLOOD PRESSURE AND WAS ASKED TO VISIT THEIR FOLLOW UP HOSPITAL FOR FURTHER EVALUATION AND TREATMENT. THE PATIENT WAS ADMITTED TO THE NEARBY HOSPITAL AND HAD DEVELOPED PARALYSIS AND SOME SIGNS OF SUSPECTED STROKE. THE PATIENT PASSED AWAY ON (B)(6) 2025. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENT OF SUSPECTED STROKE, AS WELL AS THE PATIENT OUTCOME, COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. ADDITIONALLY, THE REPORTED LOW FLOW ALARMS COULD NOT BE CONFIRMED BASED ON THE PROVIDED INFORMATION. THE ACCOUNT INDICATED THAT HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), WILL NOT BE RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATION. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) IFU AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS STROKE AS AN ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. THIS SECTION ALSO ADDRESSES ALL PUMP PARAMETERS, INCLUDING PUMP FLOW. SECTION 4 OF THE IFU DESCRIBES PUMP FLOW AND HAZARD ALARMS. THE IFU STATES THAT PER DESIGN, WHEN THE ESTIMATED FLOW VALUE IS CALCULATED AT LESS THAN 2.5 LITERS PER MINUTE (LPM), A LOW FLOW STATUS IS POSTED TO THE LOG FILE. IF THE FLOW REMAINS BELOW 2.5 LPM FOR 10 SECONDS, A LOW FLOW HAZARD ALARM IS TRIGGERED. THIS SECTION ALSO EXPLAINS THAT CHANGES IN PATIENT CONDITION CAN RESULT IN LOW FLOW. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, PROVIDES INFORMATION REGARDING ANTICOAGULATION, INCLUDING RECOMMENDED INTERNATIONAL NORMALIZED RATIO (INR) VALUES. SECTION 7 OF THE IFU, "ALARMS AND TROUBLESHOOTING", AND SECTION 5 OF THE PATIENT HANDBOOK, "ALARMS AND TROUBLESHOOTING", ADDRESS ALL SYSTEM ALARM CONDITIONS, INCLUDING THE LOW FLOW HAZARD, AS WELL AS THE APPROPRIATE ACTIONS ASSOCIATED WITH EACH CONDITION. FURTHERMORE, SECTION 8 OF THE PATIENT HANDBOOK, ¿HANDLING EMERGENCIES¿, ALSO PROVIDES EXAMPLES OF EMERGENCIES AND THE PROPER ACTIONS TO TAKE IN THE EVENT AN EMERGENCY OCCURS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22805699&pc=OAE | 3002648230-2025-00754 | 2024-12-24 05:00 | Death | MEDTRONIC CRYOCATH LP | 2025-08-14 04:00 | OAE | ARCTIC FRONT ADVANCE PRO¿ CARDIAC CRYOABLATION CATHETER | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P100010 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING EFFECTIVENESS OF CRYOBALLOON ABLATION FOR ATRIAL FIBRILLATION IN PATIENTS WITH LEFT COMMON PULMONARY VEIN VARIANT. THE AUTHORS DESCRIBED PATIENT DEATHS DURING FOLLOW-UP; HOWEVER, THE CAUSES OF DEATH WERE UNKNOWN. THERE WERE PATIENTS WHO EXPERIENCED CARDIOVASCULAR EVENTS, THROMBOLYTIC EVENTS, VASCULAR EVENTS, NEUROLOGICAL EVENTS, PULMONARY EVENTS, ACUTE REVERSIBLE PHRENIC NERVE INJURY, TAMPONADE/PERICARDIAL EFFUSION, HEART FAILURE DECOMPENSATION, CARDIAC ARREST, AND ATRIAL FIBRILLATION RECURRENCE WITH REDO ABLATION OR REHOSPITALIZATION. THE STATUS OF THE DEVICES IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: CONTINUATION OF D10: AFAPRO28 BALLOON CATHETER; 2AF283 BALLOON CATHETER; 2AF283 BALLOON CATHETER; 990063-020 MAPPING CATHETER; 990063-020 MAPPING CATHETER THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/66 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE, EXPIRATION DATE, AND UDI CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: EFFECTIVENESS OF CRYOBALLOON ABLATION FOR ATRIAL FIBRILLATION IN PATIENTS WITH LEFT COMMON PULMONARY VEIN VARIANT. HEART RHYTHM O2. 2024. 6(3):290-298. HTTPS://DOI.ORG/10.1016/J.HROO.2024.12.002 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22805743&pc=DRF | 9617601-2025-00975 | 2024-12-24 05:00 | Death | MEDTRONIC MEXICO S. DE R.L. DE CV | 2025-08-14 04:00 | DRF | ACHIEVE MAPPING CATHETER - 20 MM | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K102588 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING EFFECTIVENESS OF CRYOBALLOON ABLATION FOR ATRIAL FIBRILLATION IN PATIENTS WITH LEFT COMMON PULMONARY VEIN VARIANT. THE AUTHORS DESCRIBED PATIENT DEATHS DURING FOLLOW-UP; HOWEVER, THE CAUSES OF DEATH WERE UNKNOWN. THERE WERE PATIENTS WHO EXPERIENCED CARDIOVASCULAR EVENTS, THROMBOLYTIC EVENTS, VASCULAR EVENTS, NEUROLOGICAL EVENTS, PULMONARY EVENTS, ACUTE REVERSIBLE PHRENIC NERVE INJURY, TAMPONADE/PERICARDIAL EFFUSION, HEART FAILURE DECOMPENSATION, CARDIAC ARREST, AND ATRIAL FIBRILLATION RECURRENCE WITH REDO ABLATION OR REHOSPITALIZATION. THE STATUS OF THE DEVICES IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: CONTINUATION OF D10: AFAPRO28 BALLOON CATHETER; AFAPRO28 BALLOON CATHETER; 2AF283 BALLOON CATHETER; 2AF283 BALLOON CATHETER; 990063-020 MAPPING CATHETER THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/66 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE, EXPIRATION DATE, AND UDI CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: EFFECTIVENESS OF CRYOBALLOON ABLATION FOR ATRIAL FIBRILLATION IN PATIENTS WITH LEFT COMMON PULMONARY VEIN VARIANT. HEART RHYTHM O2. 2024. 6(3):290-298. HTTPS://DOI.ORG/10.1016/J.HROO.2024.12.002. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22805826&pc=NPT | 2015691-2025-06640 | 2025-07-07 04:00 | Death | EDWARDS LIFESCIENCES | 2025-08-14 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA VALVE | Adverse Event Without Identified Device or Use Problem | Ischemia; Heart Block | P140031 |  | 1 | Event Description: AS REPORTED, PER ARTICLE ''NON-OCCLUSIVE MESENTERIC ISCHEMIA FOLLOWING TRANSCATHETER AORTIC VALVE REPLACEMENT: A CASE REPORT: NON-OCCLUSIVE MESENTERIC ISCHEMIA (NOMI) IS A RARE COMPLICATION AFTER TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR), WITH A POORLY UNDERSTOOD PATHOGENESIS.'' A 79-YEAR-OLD FEMALE WITH A HISTORY OF REDUCED EF, KNOWN NONOBSTRUCTIVE STENOSIS OF THE CELIAC ARTERIES WITH HIGH-GRADE STENOSIS OF THE SUPERIOR AND INFERIOR MESENTERIC ARTERIES ,PRIOR ENDOVASCULAR INTERVENTIONS (SURGICALLY REPAIRED ABDOMINAL AORTIC ANEURYSM (AAA) AND STABLE THORACIC AORTIC ANEURYSM (TAA) DEVELOPED NOMI APPROXIMATELY 1 DAY AFTER TAVR WITH A 26MM SAPIEN 3 ULTRA RESILIA VALVE, RESULTING IN EXTENSIVE BOWEL NECROSIS AND PATIENT MORTALITY. A SAPIEN ULTRA RESILIA 26 MM VALVE WAS SUCCESSFULLY IMPLANTED, VIA A 14 FRENCH E-SHEATH IN THE RIGHT FEMORAL ARTERY, WITHOUT DIFFICULTY AND WITHOUT EVIDENCE OF PARAVALVULAR LEAK OR DISSECTION ON ANGIOGRAPHY. THE PATIENT WAS HYPERTENSIVE WITH A SYSTOLIC BLOOD PRESSURE OF 148 MMHG AND DIASTOLIC PRESSURE OF 84 MMHG PRIOR TO THE PROCEDURE, AND THERE WERE NO REPORTED EPISODES OF HYPOTENSION DURING THE PROCEDURE. THE PATIENT WAS RAPIDLY PACED FOR 20 SECONDS DURING VALVE DEPLOYMENT. A TEMPORARY PACEMAKER WAS PLACED DUE TO TRANSIENT AV BLOCK. THE PATIENT WAS ADMITTED TO THE MEDICAL INTENSIVE CARE UNIT FOR POST-PROCEDURE MONITORING. TWENTY-FOUR HOURS LATER, THE PATIENT DEVELOPED PROFUSE DIARRHEA AND SEVERE LEFT-SIDED ABDOMINAL PAIN. A STAT MESENTERIC CT ANGIOGRAPHY (CTA) PERFORMED AT THAT TIME (APPROXIMATELY 25 HOURS POST PROCEDURE) SHOWED EXTENSIVE INTRAHEPATIC, GASTRIC, AND MESENTERIC PNEUMATOSIS. GENERAL AND VASCULAR SURGICAL SERVICES WERE CONSULTED AND RECOMMENDED EMERGENT LAPAROTOMY. EMERGENT LAPAROTOMY WAS PERFORMED AT 26 HOURS POST-PROCEDURE AND CONFIRMED PALE TISSUE WITH ASSOCIATED NECROSIS EXTENDING FROM THE GASTROESOPHAGEAL JUNCTION TO THE DISTAL SMALL BOWEL. NO INTRAOPERATIVE SPECIMENS OR TISSUE BIOPSIES WERE COLLECTED DUE TO THE FRANKLY ISCHEMIC APPEARANCE OF THE BOWELS. GIVEN THE EXTENT OF NECROSIS WITH OVERALL POOR PROGNOSIS, THE DECISION WAS MADE TO FORGO SURGICAL RESECTION AND INSTEAD PROCEED WITH PALLIATIVE TREATMENT. POST-LAPAROTOMY, THE PATIENT'S LACTIC ACID LEVEL ROSE TO 4.6 MMOL/L, AND WBC INCREASED TO 19.2 K/MCL. THE PATIENT BECAME PROGRESSIVELY HYPOTENSIVE WITH MEAN ARTERIAL PRESSURES BELOW 65 MMHG, NECESSITATING THE INITIATION OF NOREPINEPHRINE INFUSION AT 20 MCG/MINUTE. AT THIS TIME, A DISCUSSION WAS HAD WITH THE PATIENT'S FAMILY, WHO DECIDED THAT THEY WOULD LIKE NO FURTHER ESCALATION OF MEDICAL TREATMENTS OR INTERVENTIONS. THE PATIENT WAS STARTED ON A CONTINUOUS MORPHINE INFUSION AND EXPIRED APPROXIMATELY 36 HOURS POST-PROCEDURE. Manufacturer Narrative: THE INVESTIGATION IS ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22806649&pc=DSQ | 2916596-2025-05257 | 2025-07-31 04:00 | Death | THORATEC CORPORATION | 2025-08-14 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Vasodilatation | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO VASOPLEGIA. THE PATIENT WAS ON EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) PRE AND POST IMPLANT, AND EXPERIENCED REOPERATION DUE TO BLEEDING. THE PATIENT WAS HEMODYNAMICALLY INFAUST. THE OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED.Event Description: THE SITE PROVIDED THAT THE BLEEDING WAS GENERAL WITH NO FOCUS. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: G2 - REPORT SOURCE: CORRECTED. MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, REV. B IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING BLEEDING AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿ANTICOAGULATION¿), PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) RANGE, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22806662&pc=NIQ | 2024168-2025-08453 | 2024-12-02 05:00 | Death | ABBOTT VASCULAR | 2025-08-14 04:00 | NIQ | XIENCE V EVEROLIMUS ELUTING CORONARY STENT SYSTEM | Adverse Event Without Identified Device or Use Problem | Unspecified Heart Problem | P070015 |  | 1 | Event Description: IT WAS REPORTED IN A RESEARCH SUMMARY ARTICLE THAT THE PURPOSE OF THE ARTICLE WAS TO COMPARE LONG-TERM OUTCOMES BETWEEN MEDICAL TREATMENT AND PERCUTANEOUS CORONARY INTERVENTION (PCI) OF INTERMEDIATE LESIONS WITHOUT INVASIVE PHYSIOLOGIC ASSESSMENT. THE STUDY WAS A POST HOC ANALYSIS OF A RANDOMIZED TRIAL COMPARING DEFERRING PCI VERSUS PERFORMING PCI BASED ON ANGIOGRAPHY ALONE. A TOTAL OF 899 PATIENTS WITH INTERMEDIATE CORONARY LESIONS BETWEEN 50% AND 70% DIAMETER-STENOSIS WERE RANDOMIZED TO THE CONSERVATIVE GROUP (N=449) OR THE AGGRESSIVE GROUP (N=450). FOR INTERMEDIATE LESIONS, PCI WAS PERFORMED IN THE AGGRESSIVE GROUP, BUT WAS DEFERRED IN THE CONSERVATIVE GROUP. THE STENTS USED IN THE AGGRESSIVE GROUP WERE EITHER XIENCCE V OR XIENCE PRIME. THE PRIMARY ENDPOINT WAS MAJOR ADVERSE CARDIAC EVENTS (MACE, A COMPOSITE OF ALL-CAUSE DEATH, MYOCARDIAL INFARCTION [MI], OR ISCHEMIA-DRIVEN ANY REVASCULARIZATION) AT 3 YEARS. AT 3 YEARS, THE CLINICAL OUTCOMES INCLUDED DEATH, MYOCARDIAL INFARCTION, STENT THROMBOSIS. THE CONSERVATIVE GROUP HAD A SIGNIFICANTLY HIGHER INCIDENCE OF MACE THAN THE AGGRESSIVE GROUP. BETWEEN 1 AND 3 YEARS AFTER THE INDEX PROCEDURE, COMPARED TO THE AGGRESSIVE GROUP, THE CONSERVATIVE GROUP HAD SIGNIFICANTLY HIGHER INCIDENCE OF CARDIAC DEATH OR MI AND ISCHEMIA-DRIVEN ANY REVASCULARIZATION. IT WAS FOUND THAT FOR INTERMEDIATE LESIONS, MEDICAL THERAPY ALONE, GUIDED ONLY BY ANGIOGRAPHY, WAS ASSOCIATED WITH A HIGHER RISK OF MACE AT 3 YEARS COMPARED WITH PERFORMING PCI, MAINLY DUE TO INCREASED REVASCULARIZATION. ADDITIONAL INFORMATION IS IN THE ARTICLE "OUTCOMES OF DEFERRING PERCUTANEOUS CORONARY INTERVENTION WITHOUT PHYSIOLOGIC ASSESSMENT FOR INTERMEDIATE CORONARY LESIONS". Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORDS FOR THIS PRODUCT COULD NOT BE REVIEWED AND A SIMILAR COMPLAINT QUERY COULD NOT BE PERFORMED BECAUSE THE PART NUMBER AND/OR LOT NUMBER WAS NOT REPORTED, AND THE DEVICE WAS NOT RETURNED. A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECT, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED; HOWEVER, THE SUBSEQUENT TREATMENT APPEARS TO BE RELATED TO THE OPERATIONAL CONTEXT OF THE PROCEDURE. THE REPORTED PATIENT EFFECT OF DEATH IS LISTED IN THE XIENCE V EVEROLIMUS ELUTING CORONARY STENT SYSTEMS INSTRUCTIONS FOR USE AS A KNOWN PATIENT EFFECT OF CORONARY STENTING PROCEDURES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING; THEREFORE, NO PRODUCT-RELATED CORRECTIVE ACTION WILL BE IMPLEMENTED IN THIS CASE. B2, B3, C4, D6A: ESTIMATED DATES. D4: THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. THE ADDITIONAL PATIENT EFFECTS REPORTED IN THE ARTICLE ARE CAPTURED UNDER A SEPARATE MEDWATCH REPORT. THE ADDITIONAL DEVICE REFERENCED IN B5 IS FILED UNDER A SEPARATE MEDWATCH REPORT NUMBER. LITERATURE ATTACHMENT: ARTICLE TITLE: ¿OUTCOMES OF DEFERRING PERCUTANEOUS CORONARY INTERVENTION WITHOUT PHYSIOLOGIC ASSESSMENT FOR INTERMEDIATE CORONARY LESIONS¿.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22807542&pc=NIQ | 2024168-2025-08468 | 2012-01-01 05:00 | Death | ABBOTT VASCULAR | 2025-08-14 04:00 | NIQ | XIENCE PRIME EVEROLIMUS ELUTING CORONARY STENT SYSTEM | Adverse Event Without Identified Device or Use Problem | Unspecified Vascular Problem | P110019 |  | 1 | Event Description: IT WAS REPORTED THROUGH A RESEARCH ARTICLE REGARDING PERCUTANEOUS CORONARY INTERVENTION (PCI) WITH IMPLANTATION OF SECOND GENERATION DRUG-ELUTING STENTS (DES) HAS EMERGED AS A MAINSTAY FOR THE TREATMENT OF OBSTRUCTIVE CORONARY ARTERY DISEASE GIVEN ITS BENEFICIAL IMPACT ON CLINICAL OUTCOMES IN THESE PATIENTS. VARIOUS CLINICAL STUDIES HAVE DEMONSTRATED THE EFFICACY AND SAFETY OF XIENCE STENTS IN TREATING A BROAD RANGE OF PATIENTS, INCLUDING THOSE WITH DIABETES MELLITUS AND WOMEN, AND COMPLEX LESION MORPHOLOGIES SUCH AS LEFT MAIN DISEASE, SMALL VESSELS, LONG LESIONS AND CHRONIC TOTAL OCCLUSION. HOWEVER, CLINICAL STUDIES HAVE ALSO MENTIONED ADVERSE EFFECT OUTCOMES FOR THE FOLLOWING XIENCE STENTS; XIENCE PRIME (TARGET VESSEL FAILURE (TVF), TARGET LESION FAILURE (TLF), MYOCARDIAL INFRACTION (MI), MAJOR ADVERSE CARDIAC EVENTS (MACE), ISCHEMIA, STROKE, BLEEDING, CARDIAC DEATH, TARGE VESSEL REVASCULARIZATION (TVR)), XIENCE ALPINE (ISCHEMIA AND BLEEDING), XIENCE XPEDITION (MI, TLR, DEATH, ISCHEMIA, TVR AND BLEEDING), XIENCE PRO (ISCHEMIA AND BLEEDING), XIENCE V (MACE, STROKE, DEATH, MI, TLR, ISCHEMIA, TVR AND BLEEDING), XIENCE STENTS (DEATH, MI, STROKE, ISCHEMIA, TLR, BLEEDING, TVR), PROMUS (TLR). ANOTHER VERSION OF DES WITH BIODEGRADABLE POLYMER COATING IS EMERGING BUT THEIR ADVANTAGE OVER EVEROLIMUS-ELUTING STENTS (EES) REMAINS UNCERTAIN. SECOND-GENERATION DES, PARTICULARLY THE XIENCE FAMILY OF STENTS, HAVE EMERGED AS THE MOST UTILIZED STENTS FOR TREATMENT OF PATIENTS WITH CORONARY ARTERY DISEASE (CAD) UNDERGOING PCI. THESE STENTS ARE NOW ONE OF THE MOST WIDELY USED SECOND GENERATION DES GLOBALLY AND APPROVED BY ALL MAJOR INTERNATIONAL REGULATORY BODIES. SPECIFIC PATIENT INFORMATION IS DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE, TITLED "SECOND-GENERATION EVEROLIMUS-ELUTING INTRACORONARY STENTS: A COMPREHENSIVE REVIEW OF THE CLINICAL EVIDENCE". NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: THE DEVICES WERE NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORDS FOR THIS PRODUCT COULD NOT BE REVIEWED AND A SIMILAR COMPLAINT QUERY COULD NOT BE PERFORMED BECAUSE THE PART NUMBERS AND/OR LOT NUMBERS WERE NOT REPORTED, AND THE DEVICE WAS NOT RETURNED. A CONCLUSIVE CAUSE FOR THE REPORTED DEATH, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED; HOWEVER, THE SUBSEQUENT TREATMENT APPEARS TO BE RELATED TO THE OPERATIONAL CONTEXT OF THE PROCEDURE. THE REPORTED PATIENT EFFECT OF DEATH ARE LISTED IN THE XIENCE PRIME EVEROLIMUS ELUTING CORONARY STENT SYSTEMS INSTRUCTIONS FOR USE AS A KNOWN PATIENT EFFECT OF CORONARY STENTING PROCEDURES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING. B2, B3, B6, D6: DATES ESTIMATED D4: A PARTIAL UDI WAS PROVIDED AS THE LOT NUMBER IS NOT KNOWN. ATTACHMENT: ARTICLE TITLED, " SECOND-GENERATION EVEROLIMUS-ELUTING INTRACORONARY STENTS: A COMPREHENSIVE REVIEW OF THE CLINICAL EVIDENCE."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22807576&pc=NIQ | 2024168-2025-08467 | 2013-12-01 05:00 | Death | ABBOTT VASCULAR INC. | 2025-08-14 04:00 | NIQ | XIENCE XPEDITION | Adverse Event Without Identified Device or Use Problem | Unspecified Vascular Problem | P110019 |  | 1 | Event Description: IT WAS REPORTED THROUGH A RESEARCH ARTICLE REGARDING PERCUTANEOUS CORONARY INTERVENTION (PCI) WITH IMPLANTATION OF SECOND-GENERATION DRUG-ELUTING STENTS (DES) HAS EMERGED AS A MAINSTAY FOR THE TREATMENT OF OBSTRUCTIVE CORONARY ARTERY DISEASE GIVEN ITS BENEFICIAL IMPACT ON CLINICAL OUTCOMES IN THESE PATIENTS. VARIOUS CLINICAL STUDIES HAVE DEMONSTRATED THE EFFICACY AND SAFETY OF XIENCE STENTS IN TREATING A BROAD RANGE OF PATIENTS, INCLUDING THOSE WITH DIABETES MELLITUS AND WOMEN, AND COMPLEX LESION MORPHOLOGIES SUCH AS LEFT MAIN DISEASE, SMALL VESSELS, LONG LESIONS AND CHRONIC TOTAL OCCLUSION. HOWEVER, CLINICAL STUDIES HAVE ALSO MENTIONED ADVERSE EFFECT OUTCOMES FOR THE FOLLOWING XIENCE STENTS; XIENCE PRIME (TARGET VESSEL FAILURE (TVF), TARGET LESION FAILURE (TLF), MYOCARDIAL INFRACTION (MI), MAJOR ADVERSE CARDIAC EVENTS (MACE), ISCHEMIA, STROKE, BLEEDING, CARDIAC DEATH, TARGE VESSEL REVASCULARIZATION (TVR)), XIENCE ALPINE (ISCHEMIA AND BLEEDING), XIENCE XPEDITION (MI, TLR, DEATH, ISCHEMIA, TVR AND BLEEDING), XIENCE PRO (ISCHEMIA AND BLEEDING), XIENCE V (MACE, STROKE, DEATH, MI, TLR, ISCHEMIA, TVR AND BLEEDING), XIENCE STENTS (DEATH, MI, STROKE, ISCHEMIA, TLR, BLEEDING, TVR), PROMUS (TLR). ANOTHER VERSION OF DES WITH BIODEGRADABLE POLYMER COATING IS EMERGING BUT THEIR ADVANTAGE OVER EVEROLIMUS-ELUTING STENTS (EES) REMAINS UNCERTAIN. SECOND-GENERATION DES, PARTICULARLY THE XIENCE FAMILY OF STENTS, HAVE EMERGED AS THE MOST UTILIZED STENTS FOR TREATMENT OF PATIENTS WITH CORONARY ARTERY DISEASE (CAD) UNDERGOING PCI. THESE STENTS ARE NOW ONE OF THE MOST WIDELY USED SECOND GENERATION DES GLOBALLY AND APPROVED BY ALL MAJOR INTERNATIONAL REGULATORY BODIES. SPECIFIC PATIENT INFORMATION IS DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE, TITLED "SECOND-GENERATION EVEROLIMUS-ELUTING INTRACORONARY STENTS: A COMPREHENSIVE REVIEW OF THE CLINICAL EVIDENCE". NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: THE DEVICES WERE NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORDS FOR THIS PRODUCT COULD NOT BE REVIEWED AND A SIMILAR COMPLAINT QUERY COULD NOT BE PERFORMED BECAUSE THE PART NUMBERS AND/OR LOT NUMBERS WERE NOT REPORTED, AND THE DEVICE WAS NOT RETURNED. A CONCLUSIVE CAUSE FOR THE REPORTED DEATH AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED; HOWEVER, THE SUBSEQUENT TREATMENT APPEARS TO BE RELATED TO THE OPERATIONAL CONTEXT OF THE PROCEDURE. THE REPORTED PATIENT EFFECT OF DEATH ARE LISTED IN THE XIENCE XPEDITION EVEROLIMUS ELUTING CORONARY STENT SYSTEMS INSTRUCTIONS FOR USE AS A KNOWN PATIENT EFFECT OF CORONARY STENTING PROCEDURES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING. B2, B3, B6, D6: DATES ESTIMATED. D4: A PARTIAL UDI WAS PROVIDED AS THE LOT NUMBER IS NOT KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22807860&pc=NIQ | 2024168-2025-08470 | 2008-01-01 05:00 | Death | ABBOTT VASCULAR | 2025-08-14 04:00 | NIQ | XIENCE EVEROLIMUS ELUTING CORONARY STENT SYSTEM | Adverse Event Without Identified Device or Use Problem | Unspecified Vascular Problem | P070015 |  | 1 | Event Description: IT WAS REPORTED THROUGH A RESEARCH ARTICLE REGARDING PERCUTANEOUS CORONARY INTERVENTION (PCI) WITH IMPLANTATION OF SECOND GENERATION DRUG-ELUTING STENTS (DES) HAS EMERGED AS A MAINSTAY FOR THE TREATMENT OF OBSTRUCTIVE CORONARY ARTERY DISEASE GIVEN ITS BENEFICIAL IMPACT ON CLINICAL OUTCOMES IN THESE PATIENTS. VARIOUS CLINICAL STUDIES HAVE DEMONSTRATED THE EFFICACY AND SAFETY OF XIENCE STENTS IN TREATING A BROAD RANGE OF PATIENTS, INCLUDING THOSE WITH DIABETES MELLITUS AND WOMEN, AND COMPLEX LESION MORPHOLOGIES SUCH AS LEFT MAIN DISEASE, SMALL VESSELS, LONG LESIONS AND CHRONIC TOTAL OCCLUSION. HOWEVER, CLINICAL STUDIES HAVE ALSO MENTIONED ADVERSE EFFECT OUTCOMES FOR THE FOLLOWING XIENCE STENTS; XIENCE PRIME (TARGET VESSEL FAILURE (TVF), TARGET LESION FAILURE (TLF), MYOCARDIAL INFRACTION (MI), MAJOR ADVERSE CARDIAC EVENTS (MACE), ISCHEMIA, STROKE, BLEEDING, CARDIAC DEATH, TARGE VESSEL REVASCULARIZATION (TVR)), XIENCE ALPINE (ISCHEMIA AND BLEEDING), XIENCE XPEDITION (MI, TLR, DEATH, ISCHEMIA, TVR AND BLEEDING), XIENCE PRO (ISCHEMIA AND BLEEDING), XIENCE V (MACE, STROKE, DEATH, MI, TLR, ISCHEMIA, TVR AND BLEEDING), XIENCE STENTS (DEATH, MI, STROKE, ISCHEMIA, TLR, BLEEDING, TVR), PROMUS (TLR). ANOTHER VERSION OF DES WITH BIODEGRADABLE POLYMER COATING IS EMERGING BUT THEIR ADVANTAGE OVER EVEROLIMUS-ELUTING STENTS (EES) REMAINS UNCERTAIN. SECOND-GENERATION DES, PARTICULARLY THE XIENCE FAMILY OF STENTS, HAVE EMERGED AS THE MOST UTILIZED STENTS FOR TREATMENT OF PATIENTS WITH CORONARY ARTERY DISEASE (CAD) UNDERGOING PCI. THESE STENTS ARE NOW ONE OF THE MOST WIDELY USED SECOND GENERATION DES GLOBALLY AND APPROVED BY ALL MAJOR INTERNATIONAL REGULATORY BODIES. SPECIFIC PATIENT INFORMATION IS DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE, TITLED "SECOND-GENERATION EVEROLIMUS-ELUTING INTRACORONARY STENTS: A COMPREHENSIVE REVIEW OF THE CLINICAL EVIDENCE". NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: THE DEVICES WERE NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORDS FOR THIS PRODUCT COULD NOT BE REVIEWED AND A SIMILAR COMPLAINT QUERY COULD NOT BE PERFORMED BECAUSE THE PART NUMBERS AND/OR LOT NUMBERS WERE NOT REPORTED, AND THE DEVICE WAS NOT RETURNED. A CONCLUSIVE CAUSE FOR THE REPORTED DEATH, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED; HOWEVER, THE SUBSEQUENT TREATMENT APPEARS TO BE RELATED TO THE OPERATIONAL CONTEXT OF THE PROCEDURE. THE REPORTED PATIENT EFFECT OF DEATH ARE LISTED IN THE XIENCE V EVEROLIMUS ELUTING CORONARY STENT SYSTEMS INSTRUCTIONS FOR USE AS A KNOWN PATIENT EFFECT OF CORONARY STENTING PROCEDURES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING. B2, B3, B6, D6: DATES ESTIMATED D4: A PARTIAL UDI WAS PROVIDED AS THE LOT NUMBER IS NOT KNOWN. ATTACHMENT: ARTICLE TITLED, " SECOND-GENERATION EVEROLIMUS-ELUTING INTRACORONARY STENTS: A COMPREHENSIVE REVIEW OF THE CLINICAL EVIDENCE."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22807900&pc=NIQ | 2024168-2025-08469 | 2010-01-01 05:00 | Death | ABBOTT VASCULAR | 2025-08-14 04:00 | NIQ | XIENCE V EVEROLIMUS ELUTING CORONARY STENT SYSTEM | Adverse Event Without Identified Device or Use Problem | Unspecified Vascular Problem | P070015 |  | 1 | Event Description: IT WAS REPORTED THROUGH A RESEARCH ARTICLE REGARDING PERCUTANEOUS CORONARY INTERVENTION (PCI) WITH IMPLANTATION OF SECOND-GENERATION DRUG-ELUTING STENTS (DES) HAS EMERGED AS A MAINSTAY FOR THE TREATMENT OF OBSTRUCTIVE CORONARY ARTERY DISEASE GIVEN ITS BENEFICIAL IMPACT ON CLINICAL OUTCOMES IN THESE PATIENTS. VARIOUS CLINICAL STUDIES HAVE DEMONSTRATED THE EFFICACY AND SAFETY OF XIENCE STENTS IN TREATING A BROAD RANGE OF PATIENTS, INCLUDING THOSE WITH DIABETES MELLITUS AND WOMEN, AND COMPLEX LESION MORPHOLOGIES SUCH AS LEFT MAIN DISEASE, SMALL VESSELS, LONG LESIONS AND CHRONIC TOTAL OCCLUSION. HOWEVER, CLINICAL STUDIES HAVE ALSO MENTIONED ADVERSE EFFECT OUTCOMES FOR THE FOLLOWING XIENCE STENTS; XIENCE PRIME (TARGET VESSEL FAILURE (TVF), TARGET LESION FAILURE (TLF), MYOCARDIAL INFRACTION (MI), MAJOR ADVERSE CARDIAC EVENTS (MACE), ISCHEMIA, STROKE, BLEEDING, CARDIAC DEATH, TARGE VESSEL REVASCULARIZATION (TVR)), XIENCE ALPINE (ISCHEMIA AND BLEEDING), XIENCE XPEDITION (MI, TLR, DEATH, ISCHEMIA, TVR AND BLEEDING), XIENCE PRO (ISCHEMIA AND BLEEDING), XIENCE V (MACE, STROKE, DEATH, MI, TLR, ISCHEMIA, TVR AND BLEEDING), XIENCE STENTS (DEATH, MI, STROKE, ISCHEMIA, TLR, BLEEDING, TVR), PROMUS (TLR). ANOTHER VERSION OF DES WITH BIODEGRADABLE POLYMER COATING IS EMERGING BUT THEIR ADVANTAGE OVER EVEROLIMUS-ELUTING STENTS (EES) REMAINS UNCERTAIN. SECOND-GENERATION DES, PARTICULARLY THE XIENCE FAMILY OF STENTS, HAVE EMERGED AS THE MOST UTILIZED STENTS FOR TREATMENT OF PATIENTS WITH CORONARY ARTERY DISEASE (CAD) UNDERGOING PCI. THESE STENTS ARE NOW ONE OF THE MOST WIDELY USED SECOND GENERATION DES GLOBALLY AND APPROVED BY ALL MAJOR INTERNATIONAL REGULATORY BODIES. SPECIFIC PATIENT INFORMATION IS DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE, TITLED "SECOND-GENERATION EVEROLIMUS-ELUTING INTRACORONARY STENTS: A COMPREHENSIVE REVIEW OF THE CLINICAL EVIDENCE". NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: THE DEVICES WERE NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORDS FOR THIS PRODUCT COULD NOT BE REVIEWED AND A SIMILAR COMPLAINT QUERY COULD NOT BE PERFORMED BECAUSE THE PART NUMBERS AND/OR LOT NUMBERS WERE NOT REPORTED, AND THE DEVICE WAS NOT RETURNED. A CONCLUSIVE CAUSE FOR THE REPORTED DEATH, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED; HOWEVER, THE SUBSEQUENT TREATMENT APPEARS TO BE RELATED TO THE OPERATIONAL CONTEXT OF THE PROCEDURE. THE REPORTED PATIENT EFFECT OF DEATH IS LISTED IN THE XIENCE V EVEROLIMUS ELUTING CORONARY STENT SYSTEMS INSTRUCTIONS FOR USE AS A KNOWN PATIENT EFFECT(S) OF CORONARY STENTING PROCEDURES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING. B2, B3, B6, D6: DATES ESTIMATED. D4: A PARTIAL UDI WAS PROVIDED AS THE LOT NUMBER IS NOT KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22808575&pc=FKX | 0008030665-2025-01887 | 2025-08-12 04:00 | Death | ERIKA DE REYNOSA, S.A. DE C.V. | 2025-08-14 04:00 | FKX | LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | K181108 |  | 1 | Event Description: ON 12/AUG/2025, FRESENIUS BECAME AWARE THIS PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON CONTINUOUS CYCLIC PERITONEAL DIALYSIS [CC(PD)] UTILIZING A LIBERTY SELECT CYCLER FOR RENAL REPLACEMENT THERAPY (RRT) EXPIRED WHILE CONNECTED TO THE DEVICE. FOLLOW-UP WITH THE PATIENT¿S PD REGISTERED NURSE (PDRN) CONFIRMED THE PATIENT EXPIRED AT HOME WHILE CONNECTED TO THE LIBERTY SELECT CYCLER ON (B)(6) 2025. THE PDRN REPORTED THE PRIMARY CAUSE OF DEATH WAS CARDIAC ARREST, AND UNRELATED TO ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S) MALFUNCTION OR DEFICIENCY. ADDITIONALLY, A REVIEW OF THE TREATMENT DATA DID NOT REVEAL ANY ALARMS WERE PRESENT. Manufacturer Narrative: PLANT INVESTIGATION: THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL INVESTIGATION: A TEMPORAL RELATIONSHIP EXISTS BETWEEN CCPD THERAPY UTILIZING A LIBERTY SELECT CYCLER AND THE SERIOUS ADVERSE EVENTS OF CARDIAC ARREST AND DEATH, AS THE PATIENT WAS ACTIVELY UNDERGOING CCPD THERAPY WHEN HE EXPIRED. THE PATIENT¿S PRIMARY CAUSE OF DEATH WAS REPORTED AS CARDIAC ARREST. THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. CARDIOVASCULAR DISEASE (INCLUDING SUDDEN CARDIAC DEATH) ACCOUNTS FOR THE MOST DEATHS AMONG THE ESRD POPULATION. ADDITIONALLY, ADULTS WITH ESRD HAVE MORTALITY RATES UP TO 30-FOLD HIGHER THAN THE GENERAL POPULATION. PER THE PDRN, THE SERIOUS ADVERSE EVENTS WERE UNRELATED TO ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S) MALFUNCTION OR DEFICIENCY. BASED ON THE INFORMATION AVAILABLE, THE LIBERTY SELECT CYCLER CAN BE DISASSOCIATED FROM THE SERIOUS ADVERSE EVENTS. THERE IS NO ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) CAUSED OR CONTRIBUTED TO THE SERIOUS ADVERSE EVENTS. FURTHERMORE, THERE WAS NO REPORT A FRESENIUS PRODUCT(S) AND/OR DEVICE(S) FAILED TO MEET THE USERS¿ EXPECTATIONS AND/OR THE MANUFACTURERS¿ SPECIFICATIONS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22808696&pc=NIQ | 2024168-2025-08481 | 2024-12-02 05:00 | Death | ABBOTT VASCULAR | 2025-08-14 04:00 | NIQ | XIENCE PRIME EVEROLIMUS ELUTING CORONARY STENT SYSTEM | Adverse Event Without Identified Device or Use Problem | Unspecified Heart Problem | P110019 |  | 1 | Event Description: IT WAS REPORTED IN A RESEARCH SUMMARY ARTICLE THAT THE PURPOSE OF THE ARTICLE WAS TO COMPARE LONG-TERM OUTCOMES BETWEEN MEDICAL TREATMENT AND PERCUTANEOUS CORONARY INTERVENTION (PCI) OF INTERMEDIATE LESIONS WITHOUT INVASIVE PHYSIOLOGIC ASSESSMENT. THE STUDY WAS A POST HOC ANALYSIS OF A RANDOMIZED TRIAL COMPARING DEFERRING PCI VERSUS PERFORMING PCI BASED ON ANGIOGRAPHY ALONE. A TOTAL OF 899 PATIENTS WITH INTERMEDIATE CORONARY LESIONS BETWEEN 50% AND 70% DIAMETER-STENOSIS WERE RANDOMIZED TO THE CONSERVATIVE GROUP (N=449) OR THE AGGRESSIVE GROUP (N=450). FOR INTERMEDIATE LESIONS, PCI WAS PERFORMED IN THE AGGRESSIVE GROUP, BUT WAS DEFERRED IN THE CONSERVATIVE GROUP. THE STENTS USED IN THE AGGRESSIVE GROUP WERE EITHER XIENCCE V OR XIENCE PRIME. THE PRIMARY ENDPOINT WAS MAJOR ADVERSE CARDIAC EVENTS (MACE, A COMPOSITE OF ALL-CAUSE DEATH, MYOCARDIAL INFARCTION [MI], OR ISCHEMIA-DRIVEN ANY REVASCULARIZATION) AT 3 YEARS. AT 3 YEARS, THE CLINICAL OUTCOMES INCLUDED DEATH, MYOCARDIAL INFARCTION, STENT THROMBOSIS. THE CONSERVATIVE GROUP HAD A SIGNIFICANTLY HIGHER INCIDENCE OF MACE THAN THE AGGRESSIVE GROUP. BETWEEN 1 AND 3 YEARS AFTER THE INDEX PROCEDURE, COMPARED TO THE AGGRESSIVE GROUP, THE CONSERVATIVE GROUP HAD SIGNIFICANTLY HIGHER INCIDENCE OF CARDIAC DEATH OR MI AND ISCHEMIA-DRIVEN ANY REVASCULARIZATION. IT WAS FOUND THAT FOR INTERMEDIATE LESIONS, MEDICAL THERAPY ALONE, GUIDED ONLY BY ANGIOGRAPHY, WAS ASSOCIATED WITH A HIGHER RISK OF MACE AT 3 YEARS COMPARED WITH PERFORMING PCI, MAINLY DUE TO INCREASED REVASCULARIZATION. ADDITIONAL INFORMATION IS IN THE ARTICLE "OUTCOMES OF DEFERRING PERCUTANEOUS CORONARY INTERVENTION WITHOUT PHYSIOLOGIC ASSESSMENT FOR INTERMEDIATE CORONARY LESIONS". Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORDS FOR THIS PRODUCT COULD NOT BE REVIEWED AND A SIMILAR COMPLAINT QUERY COULD NOT BE PERFORMED BECAUSE THE PART NUMBER AND/OR LOT NUMBER WAS NOT REPORTED, AND THE DEVICE WAS NOT RETURNED. A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECTS, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED; HOWEVER, THE SUBSEQUENT TREATMENT APPEARS TO BE RELATED TO THE OPERATIONAL CONTEXT OF THE PROCEDURE. THE REPORTED PATIENT EFFECT OF DEATH IS LISTED IN THE XIENCE PRIME EVEROLIMUS ELUTING CORONARY STENT SYSTEMS INSTRUCTIONS FOR USE AS A KNOWN PATIENT EFFECT OF CORONARY STENTING PROCEDURES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING; THEREFORE, NO PRODUCT-RELATED CORRECTIVE ACTION WILL BE IMPLEMENTED IN THIS CASE. B2, B3, C4, D6A: ESTIMATED DATES D4: THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. THE ADDITIONAL PATIENT EFFECTS REPORTED IN THE ARTICLE ARE CAPTURED UNDER A SEPARATE MEDWATCH REPORT. THE ADDITIONAL DEVICE REFERENCED IN B5 IS FILED UNDER A SEPARATE MEDWATCH REPORT NUMBER. LITERATURE ATTACHMENT: ARTICLE TITLE: ¿OUTCOMES OF DEFERRING PERCUTANEOUS CORONARY INTERVENTION WITHOUT PHYSIOLOGIC ASSESSMENT FOR INTERMEDIATE CORONARY LESIONS¿.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22809410&pc=NPT | 9612164-2025-04003 | 2025-07-24 04:00 | Death | MEDTRONIC IRELAND | 2025-08-14 04:00 | NPT | EVOLUT FX DCS | Failure to Advance; Patient Device Interaction Problem |  | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE ATTEMPTED IMPLANT OF A TRANSCATHETER VALVE, UPON ADVANCING THE DELIVERY CATHETER SYSTEM (DCS) THROUGH THE ASCENDING AORTA, THE DCS WAS UNABLE TO CROSS THE AORTIC VALVE DUE TO CALCIFICATION AT THE SINOTUBULAR JUNCTION (STJ). MULTIPLE ATTEMPTS WERE PERFORMED, INCLUDING USING A BUDDY WIRE AND SWITCHING TO A NON-MEDTRONIC GUIDEWIRE, BUT THE DCS WAS STILL UNABLE TO PASS. A P RE-IMPLANT BALLOON AORTIC VALVULOPLASTY (BAV) WAS PERFORMED; HOWEVER, THE BALLOON WAS UNABLE TO PASS THROUGH THE AORTIC VALVE AS WELL. A NEW BALLOON WAS INTRODUCED TO PERFORM ANOTHER PRE-IMPLANT BAV, BUT THE BALLOON WAS ALSO UNABLE TO PASS. AT THIS POINT, THE PROCEDURE HAD BEEN DELAYED BY APPROXIMATELY 1 HOUR. THEREFORE, A NEW VALVE AND DELIVERY CATHETER SYSTEM (DCS) WERE USED. A SNARE WAS USED TO PULL THE DCS AND BALLOON THROUGH THE AORTIC VALVE. A PRE-IMPLANT BAV WAS PERFORMED TO ALLOW THE VALVE TO EXPAND IN THE PRESENCE OF SEVERE CALCIFICATION; HOWEVER, CALCIFICATION WAS OBSERVED TO HAVE PEELED OFF AT THE STJ. FOLLOWING THE IMPLANT OF THE VALVE, AN ANGIOGRAPHY WAS PERFORMED THAT REVEALED A PATENT FALSE LUMEN ORIGINATING FROM THE STJ. AN ECHOCARDIOGRAPHY WAS PERFORMED THAT CONFIRMED AN ASCENDING AORTIC DISSECTION EXTENDING TO THE ABDOMINAL REGION. SUBSEQUENTLY, OPEN-HEART SURGICAL REPAIR WAS PERFORMED TO ADDRESS THE DISSECTION. A FEW DAYS FOLLOWING THE IMPLANT OF THE TRANSCATHETER VALVE, THE PATIENT DIED. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. B2: THE PATIENT'S DATE OF DEATH IS APPROXIMATE. MONTH AND YEAR ARE CONFIRMED VALID. CONTINUATION OF D10: PRODUCT ID {EVFXPLUS-26}; PRODUCT LOT/SERIAL NUMBER (B)(6); PRODUCT TYPE: {0195-HEARTVALVES}; IMPLANT DATE {N/A}; EXPLANT DATE {N/A} MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22809434&pc=SCZ | 3030471525-2025-00003 | 2025-07-17 04:00 | Death | PI-CARDIA LTD. | 2025-08-14 04:00 | SCZ | SHORTCUT | Difficult or Delayed Positioning | Cardiac Tamponade; Cardiac Perforation | DEN240017 |  | 1 | Event Description: A 74-YEAR-OLD FEMALE WITH A HISTORY OF A 21 MM MAGNA EASE VALVE (IMPLANTED IN 2004) AND A 23 MM EVOLUT PRO VALVE (IMPLANTED IN 2017), PRESENTED WITH STRUCTURAL VALVE DETERIORATION DUE TO AORTIC REGURGITATION (AR). THE PLANNED PROCEDURE INVOLVED A DUAL LEAFLET SPLIT USING THE SHORTCUT DEVICE, FOLLOWED BY IMPLANTATION OF A 20 MM SAPIEN 3 VALVE. THE SHORTCUT WAS INTRODUCED VIA A 14F ESHEATH. A SHARP BEND IN THE AORTIC ARCH CAUSED THE SYSTEM TO BIAS TOWARD THE INNER CURVE. UPON UNSHEATHING THE POSITIONING ARM (PA) IN THE ANNULUS, THE PA WAS NOTED TO BE PARTIALLY COLLAPSED AND INTERACTING WITH THE EVOLUT FRAME. THE OPERATOR FOLLOWED STANDARD IFU-GUIDED MANIPULATION TECHNIQUES, INCLUDING WIRE AND DEVICE ADJUSTMENTS, FLEXING, AND PA ROTATION. DESPITE PROLONGED ATTEMPTS, THE PA COULD NOT BE SUCCESSFULLY POSITIONED OR ACTIVATED OVER THE LEFT CORONARY CUSP (LC), AND TWO FAILED ACTIVATION ATTEMPTS WERE RECORDED. THE TEAM RE-SHEATHED THE DEVICE AND DECIDED TO ATTEMPT LEAFLET SPLITTING AT THE RIGHT CORONARY CUSP (RC). DURING REPOSITIONING, ANESTHESIA NOTED A GRADUAL DROP IN THE PATIENT'S BLOOD PRESSURE. CARDIAC TAMPONADE WAS IDENTIFIED VIA IMAGING, ORIGINATING AROUND THE LEFT VENTRICLE AND EXTENDING TOWARD THE LEFT ATRIUM. THE DEVICE AND WIRE WERE REMOVED. UNFORTUNATELY DESPITE PERICARDIOCENTESIS WITH REMOVAL OF LARGE AMOUNT OF BLOOD AND CHEST COMPRESSIONS THERE WAS NO RETURN OF SPONTANEOUS CIRCULATION AND THE PATIENT DIED. THE PATIENT WAS NOT FELT TO BE A CANDIDATE FOR ECMO OR SURGERY. THE EVENT OCCURRED INTRA-PROCEDURALLY AND MAY BE ASSOCIATED WITH ANATOMICAL CONSTRAINTS AND INTERACTION BETWEEN THE DEVICE AND THE PRIOR IMPLANTED EVOLUT VALVE. NO DEVICE MALFUNCTION WAS IDENTIFIED DURING THE PROCEDURE, AND ALL MANIPULATIONS WERE WITHIN THE SCOPE OF THE IFU. Manufacturer Narrative: DURING A PLANNED VALVE-IN-VALVE-IN-VALVE (VIVIV) PROCEDURE IN A HIGH SURGICAL RISK PATIENT, A SHORTCUT DEVICE WAS USED FOR A DUAL-LEAFLET SPLIT DUE TO A RISK OF CORONARY OBSTRUCTION. THE PATIENT, A 73-YEAR-OLD FEMALE WITH AN ANGLED AORTIC ARCH, SMALL STJ AND ANNULUS, HYPERTROPHIED LEFT VENTRICLE WITH A SMALL CAVITY AND TWO PRIOR BIOPROSTHETIC AORTIC VALVES, EXPERIENCED A LEFT VENTRICULAR (LV) PERFORATION, RESULTING IN INTRA-PROCEDURAL DEATH. PER THE OPERATOR, THE PERFORATION WAS MOST LIKELY CAUSED BY THE GUIDEWIRE DURING CHALLENGING POSITIONING OF THE SHORTCUT DISTAL UNIT, REQUIRING MULTIPLE MANEUVERS DUE TO THE PATIENT'S COMPLEX ANATOMY. HEMODYNAMIC COLLAPSE OCCURRED DURING REPOSITIONING AT THE RIGHT CORONARY CUSP; CARDIAC TAMPONADE WAS CONFIRMED. RESUSCITATION, INCLUDING PERICARDIOCENTESIS AND CHEST COMPRESSIONS, WAS UNSUCCESSFUL. NO DEVICE MALFUNCTION OR DAMAGE WAS IDENTIFIED; POST-PROCEDURE FUNCTIONAL TESTING CONFIRMED THE DEVICE WAS FULLY OPERATIONAL. WHILE NO DEVICE MALFUNCTION WAS FOUND, THE SHORTCUT WAS IN ACTIVE USE AT THE TIME, AND A CAUSAL RELATIONSHIP TO THE OUTCOME CANNOT BE DEFINITIVELY EXCLUDED. LV PERFORATION IS A KNOWN PROCEDURAL RISK IN STRUCTURAL HEART INTERVENTIONS, INCLUDING TAVR. THE DEVICE'S DESIGN, INSTRUCTIONS FOR USE, AND PHYSICIAN TRAINING ADDRESS SUCH RISKS. AS IN TAVR PROCEDURES, CAREFUL ATTENTION TO WIRE MANIPULATION IS CRITICAL TO PREVENT COMPLICATIONS LIKE LEFT VENTRICULAR (LV) PERFORATION. BASED ON THE INVESTIGATION, THE RESIDUAL RISK REMAINS ACCEPTABLE, AND NO ADDITIONAL RISK CONTROLS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22809618&pc=KDJ | 1416980-2025-04192 |  | Death | VANTIVE US HEALTHCARE LLC | 2025-08-14 04:00 | KDJ | MINICAP WITH POVIDONE-IODINE SOLUTION | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Peritonitis | K152129 |  | 1 | Event Description: IT WAS REPORTED THAT A PERITONEAL DIALYSIS (PD) PATIENT EXPERIENCED PERITONITIS AND PASSED AWAY AFTER USING MINICAP DISCONNECT CAPS. THE CAUSE OF THE PERITONITIS WAS NOT REPORTED. THE CAREGIVER REPORTED, DURING A ROUTINE CHECK AT THE PD FACILITY, THE PATIENT WAS FOUND TO HAVE PERITONITIS. THE PATIENT RECEIVED UNSPECIFIED TREATMENT AT THE FACILITY AND WAS SENT HOME WITH UNSPECIFIED ANTIBIOTICS FOR PERITONITIS. ON THE DRIVE HOME, THE PATIENT FELT ILL AND BEGAN VOMITING. THE CAREGIVER REPORTED THEY CALLED EMERGENCY SERVICES AND SUBSEQUENTLY THE PATIENT WAS TAKEN TO THE HOSPITAL WHERE IT WAS CONFIRMED THE PATIENT HAD PERITONITIS. ACCORDING TO THE REPORTER, "OVER THE NEXT NINE MONTHS" THE PATIENT CONTINUED TO HAVE SEVERE UNSPECIFIED INFECTIONS THAT "REAPPEARED WITHIN THREE TO FOUR WEEKS FOLLOWING EACH TREATMENT". IT WAS REPORTED THAT "THE PATIENT'S CONDITION WORSENED AND REMAINED HOSPITALIZED CONTINUOUSLY". THE PATIENT SUBSEQUENTLY PASSED AWAY. THE CAUSE OF DEATH WAS REPORTED AS "A RESULT OF PERITONITIS" AND ANOTHER INDICATION. IT WAS NOT REPORTED IF AN AUTOPSY WAS PERFORMED. IT WAS REPORTED THAT PD THERAPY WAS ONGOING PRIOR TO DEATH, HOWEVER IT WAS NOT REPORTED IF PD THERAPY WAS ONGOING AT THE TIME OF DEATH. NO ADDITIONAL INFORMATION WAS AVAILABLE. Manufacturer Narrative: B3: THE PERITONITIS OCCURRED ON AN UNSPECIFIED DATE IN (B)(6) 2022 D4: THE LOT NUMBER IS UNKNOWN, THEREFORE, ONLY A PRIMARY DI NUMBER COULD BE PROVIDED. THE DEVICE WAS NOT RETURNED AND THE LOT NUMBER IS UNKNOWN; THEREFORE, A DEVICE ANALYSIS COULD NOT BE COMPLETED. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22809754&pc=LWS | 2017865-2025-97419 | 2025-07-23 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-08-14 04:00 | LWS | DURATA STS OPTIM ACTIVE FIXATION | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P950022 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT DECEASED FOLLOWING AN UNRELATED PROCEDURE. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. IT WAS REPORTED THAT THE CAUSE OF DEATH WAS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22809823&pc=NPT | 9617601-2025-00986 | 2025-07-24 04:00 | Death | MEDTRONIC MEXICO S. DE R.L. DE CV | 2025-08-14 04:00 | NPT | EVOLUT FX VALVE | Failure to Advance; Patient Device Interaction Problem | Insufficient Information | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE ATTEMPTED IMPLANT OF A TRANSCATHETER VALVE, UPON ADVANCING THE DELIVERY CATHETER SYSTEM (DCS) THROUGH THE ASCENDING AORTA, THE DCS WAS UNABLE TO CROSS THE AORTIC VALVE DUE TO CALCIFICATION AT THE SINOTUBULAR JUNCTION (STJ). MULTIPLE ATTEMPTS WERE PERFORMED, INCLUDING USING A BUDDY WIRE AND SWITCHING TO A NON-MEDTRONIC GUIDEWIRE, BUT THE DCS WAS STILL UNABLE TO PASS. A PRE-IMPLANT BALLOON AORTIC VALVULOPLASTY (BAV) WAS PERFORMED; HOWEVER, THE BALLOON WAS UNABLE TO PASS THROUGH THE AORTIC VALVE AS WELL. A NEW BALLOON WAS INTRODUCED TO PERFORM ANOTHER PRE-IMPLANT BAV, BUT THE BALLOON WAS ALSO UNABLE TO PASS. AT THIS POINT, THE PROCEDURE HAD BEEN DELAYED BY APPROXIMATELY 1 HOUR. THEREFORE, A NEW VALVE AND DELIVERY CATHETER SYSTEM (DCS) WERE USED. A SNARE WAS USED TO PULL THE DCS AND BALLOON THROUGH THE AORTIC VALVE. A PRE-IMPLANT BAV WAS PERFORMED TO ALLOW THE VALVE TO EXPAND IN THE PRESENCE OF SEVERE CALCIFICATION; HOWEVER, CALCIFICATION WAS OBSERVED TO HAVE PEELED OFF AT THE STJ. FOLLOWING THE IMPLANT OF THE VALVE, AN ANGIOGRAPHY WAS PERFORMED THAT REVEALED A PATENT FALSE LUMEN ORIGINATING FROM THE STJ. AN ECHOCARDIOGRAPHY WAS PERFORMED THAT CONFIRMED AN ASCENDING AORTIC DISSECTION EXTENDING TO THE ABDOMINAL REGION. SUBSEQUENTLY, OPEN-HEART SURGICAL REPAIR WAS PERFORMED TO ADDRESS THE DISSECTION. A FEW DAYS FOLLOWING THE IMPLANT OF THE TRANSCATHETER VALVE, THE PATIENT DIED. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. B2: THE PATIENT'S DATE OF DEATH IS APPROXIMATE. MONTH AND YEAR ARE CONFIRMED VALID. THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME: EVOLUT FX DCS; PRODUCT ID: D-EVOLUTFX-2329 (LOT: 0012659427); PRODUCT TYPE: 0195-HEART VALVES; THE CODES PRESENT IN SECTION H6 CORRESPOND TO MULTIPLE COMPONENTS/PRODUCTS THAT COMPRISE THIS REPORTED EVENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22810090&pc=QBJ | MW5174584 | 2025-05-18 04:00 | Death | DEXCOM INC. | 2025-08-14 04:00 | QBJ | DEXCOM G7 RECEIVER | Defective Device | Convulsion/Seizure; Hypoglycemic Shock |  |  | 1 | Event Description: (B)(6) USED THE DEXCOM G7. SHE USED THE G7 RECEIVER PART NUMBER MT26403-0 TO MONITOR AND TRACK HER GLUCOSE LEVELS. ON SUNDAY, (B)(6) 2025 WE FOUND HER UNRESPONSIVE. SHE WAS AIRLIFTED TO (B)(6) HOSPITAL ((B)(6)) IN (B)(6). SHE HAD SUFFERED SEIZURES WHICH LEFT HER IN THE CONDITION WE FOUND HER IN. SHE NEVER RESPONDED OR RECOVERED FROM THE BRAIN DAMAGE CAUSED BY THE SEIZURES AND PASSED AWAY ON (B)(6) 2025. THE SEIZURES WERE CAUSED BY HYPOGLYCEMIA. HER GLUCOSE WAS UNDER 40 FOR OVER 13 HOURS ONCE WE FOUND HER. HER CAUSE OF DEATH WAS METABOLIC ENCEPHALOPATHY DUE TO HYPOGLYCEMIA. AFTER SHE HAD PASSED AWAY, A LETTER FROM (B)(4) DATED (B)(6) 2025, WAS RECEIVED STATING ON (B)(6) 2025 (4 DAYS BEFORE (B)(6) HYPOGLYCEMIC EVENT OCCURRED) DEXCOM HAD NOTIFIED THEM THAT CERTAIN PART NUMBERS OF THE G7 RECEIVERS MAY EXPERIENCE AN ISSUE WHERE THE ALARM/ALERT MAY NOT PROVIDE AUDIO OUTPUT. PAMELA LIVED ALONE AND DEPENDED ON HER RECEIVER TO ALERT HER IF HER GLUCOSE WAS OUT OF RANGE. THERE WERE MANY TESTS RAN ON (B)(6) TO DETERMINE WHAT HAD CAUSED HER ISSUE. SHE HAD HER 1ST MAGNETIC RESONANCE IMAGING OF THE BRAIN SHOWING DAMAGE EITHER LATE (B)(6) 2025, OR IN THE EARLY MORNING HOURS OF (B)(6) 2025. SHE WAS THEN HOOKED UP TO THE ELECTROENCEPHALOGRAM TO DETERMINE IF SEIZURES WERE STILL OCCURRING. ANOTHER MRI WAS COMPLETED SOME DAYS AFTER THAT STILL IN (B)(6) 2025 STILL SHOWING THE DAMAGE. THE ELECTROENCEPHALOGRAM PICKED UP MORE SEIZURES AND THEN THE SLOW BRAIN ACTIVITY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22810419&pc=LWP | 2017865-2025-97524 | 2025-07-27 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-08-14 04:00 | LWP | ASSURITY MRI | Adverse Event Without Identified Device or Use Problem | Respiratory Failure | P140033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED FOR A NEW IMPLANT PROCEDURE WITH SUCCESSFUL INSERTION OF THE PACEMAKER. SYSTEM FUNCTION WAS TESTED DURING THE PROCEDURE AND FOUND TO BE NORMAL. THERE WAS NO KNOWN ALLEGATION FROM THE HEALTH CARE PROFESSIONAL SUGGESTING SYSTEM MALFUNCTION. IT IS REPORTED THAT THE PATIENT DEVELOPED RESPIRATORY FAILURE AFTER BEING EXTUBATED, INCURRED A RESPIRATORY ARREST AND DID NOT SURVIVE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22810463&pc=NVN | 2017865-2025-97525 | 2025-07-27 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-08-14 04:00 | NVN | ULTIPACE¿ LEAD, 58CM MRI | Adverse Event Without Identified Device or Use Problem | Respiratory Failure | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED FOR A NEW IMPLANT PROCEDURE WITH SUCCESSFUL INSERTION OF THE PACEMAKER. SYSTEM FUNCTION WAS TESTED DURING THE PROCEDURE AND FOUND TO BE NORMAL. THERE WAS NO KNOWN ALLEGATION FROM THE HEALTH CARE PROFESSIONAL SUGGESTING SYSTEM MALFUNCTION. IT IS REPORTED THAT THE PATIENT DEVELOPED RESPIRATORY FAILURE AFTER BEING EXTUBATED, INCURRED A RESPIRATORY ARREST AND DID NOT SURVIVE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22810571&pc=NVN | 2017865-2025-97526 | 2025-07-27 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-08-14 04:00 | NVN | ULTIPACE¿ LEAD, 52CM MRI | Adverse Event Without Identified Device or Use Problem | Respiratory Failure | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED FOR A NEW IMPLANT PROCEDURE WITH SUCCESSFUL INSERTION OF THE PACEMAKER. SYSTEM FUNCTION WAS TESTED DURING THE PROCEDURE AND FOUND TO BE NORMAL. THERE WAS NO KNOWN ALLEGATION FROM THE HEALTH CARE PROFESSIONAL SUGGESTING SYSTEM MALFUNCTION. IT IS REPORTED THAT THE PATIENT DEVELOPED RESPIRATORY FAILURE AFTER BEING EXTUBATED, INCURRED A RESPIRATORY ARREST AND DID NOT SURVIVE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22812006&pc=OZD | 1220648-2025-45807 | 2025-07-21 04:00 | Death | ABIOMED, INC. | 2025-08-14 04:00 | OZD | IMPELLA CP | Malposition of Device | Hemolysis | P140003 |  | 1 | Event Description: IT WAS REPORTED A PATIENT IN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP FOR MECHANICAL CIRCULATORY SUPPORT. HOWEVER, DURING REMOVAL OF THE IMPELLA CP INNER DILATOR, THE CLEAR LOCKING HUB DETACHED AND COULD NOT BE REMOVED. THE PUMP WAS IMPLANTED THROUGH THE DIAPHRAGM AND THE CLEAR RING WAS REMOVED WITH SURGICAL SHEARS. THE PATIENT WAS SENT TO THE UNIT ON SUPPORT. THE NEXT DAY IT WAS NOTED THE PATIENT EXPERIENCED HEMOLYSIS. THE PUMP WAS REPOSITIONED TO RESOLVE THE ISSUE. LATER, THE FOLLOWING DAY, CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: PATIENT INFORMATION: A5 AND A6 ARE UNKNOWN. PRODUCT STATUS: THE IMPELLA CP DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22812417&pc=BZD | 2518422-2025-109137 | 2025-07-31 04:00 | Death | RESPIRONICS, INC. | 2025-08-14 04:00 | BZD | DREAMSTATION | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH A DREAMSTATION DEVICE HAS PASSED AWAY. THE REPORTED EVENT MAKES NO ALLEGATION OF ANY SPECIFIC DEVICE MALFUNCTION, USE ERROR, FAILURE, LABELING, OR OTHER DEFICIENCY THAT IS RELEVANT TO THE HARMS REPORTED; THEREFORE, BASED ON INFORMATION AVAILABLE AT THIS TIME, THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO A SERIOUS INJURY OR DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22812804&pc=QAV | 9611451-2025-00774 | 2025-05-24 04:00 | Death | FISHER & PAYKEL HEALTHCARE LTD | 2025-08-14 04:00 | QAV | FISHER & PAYKEL HEALTHCARE | Insufficient Information | Insufficient Information | K221338 |  | 1 | Event Description: ON 16 JULY 2025, A HEALTHCARE FACILITY IN OREGON REPORTED VIA A FISHER & PAYKEL HEALTHCARE (F&P) REPRESENTATIVE THAT ON (B)(6) 2025, A PATIENT RECEIVING THERAPY VIA A PT301US AIRVO 3 HUMIDIFIER (AIRVO 3) DECEASED. THE HEALTHCARE FACILITY STATED THAT PRIOR TO THE REPORTED EVENT, THE PATIENT WAS ADMITTED TO HOSPITAL ON (B)(6) 2025 FOR OSTEOMYELITIS AND BACTEREMIA REQUIRING SURGICAL INTERVENTION. THE HEALTHCARE FACILITY STATED THAT ON (B)(6) 2025 AT APPROXIMATELY 03:59AM, A NASAL CANNULA USED TO DELIVER THERAPY TO THE PATIENT WAS FOUND TO BE DISCONNECTED FROM THE CIRCUIT, AND THE PATIENT'S PULSE OXIMETER PROBE WAS DETACHED FROM THE PATIENT'S FINGER. THE HEALTHCARE FACILITY STATED THAT THE SUBJECT DEVICE DID NOT ALARM. THE HEALTHCARE FACILITY FURTHER STATED THAT AT 17:35PM THE PATIENT DECEASED. F&P HAS REQUESTED FOR FURTHER INFORMATION FROM THE HEALTHCARE FACILITY, INCLUDING THE PATIENT CONDITION, DEVICE DETAILS, SEQUENCE OF EVENTS, AND THE RETURN OF THE SUBJECT DEVICE. F&P WILL PROVIDE A FOLLOW UP REPORT UPON COMPLETION OF F&P'S INVESTIGATION. Manufacturer Narrative: (B)(4) [D4, H4: FISHER & PAYKEL HEALTHCARE (F&P) HAS REQUESTED FOR DEVICE DETAILS. HOWEVER, THE DEVICE DETAILS HAVE NOT YET BEEN PROVIDED BY THE HEALTHCARE FACILITY.] F&P HAS REQUESTED FOR FURTHER INFORMATION FROM THE HEALTHCARE FACILITY, INCLUDING THE PATIENT CONDITION, SEQUENCE OF EVENTS, AND THE RETURN OF THE SUBJECT DEVICE. F&P WILL PROVIDE A FOLLOW UP REPORT UPON COMPLETION OF F&P'S INVESTIGATION. PRODUCT BACKGROUND: THE PT301US HUMIDIFIER (AIRVO 3) IS FOR THE TREATMENT OF SPONTANEOUSLY BREATHING PATIENTS WHO WOULD BENEFIT FROM RECEIVING HIGH FLOW WARMED AND HUMIDIFIED RESPIRATORY GASES. THIS INCLUDES PATIENTS WHO HAVE HAD UPPER AIRWAYS BYPASSED. THE FLOW MAY BE FROM 2 - 70 L/MIN DEPENDING ON THE PATIENT INTERFACE. THE AIRVO 3 IS FOR PATIENTS IN HOSPITALS AND SUBACUTE FACILITIES. THE AIRVO 3 IS NOT INTENDED FOR LIFE SUPPORT. THE AIRVO 3 CAN DELIVER THESE HIGH FLOW GASES THROUGH NASAL CANNULA TO AUGMENT THE BREATHING OF SPONTANEOUSLY BREATHING NEONATE, INFANT, CHILD, ADOLESCENT AND ADULT PATIENTS SUFFERING FROM RESPIRATORY DISTRESS AND/OR HYPOXEMIA IN THE HOSPITAL SETTING. THE AIRVO 3 IS NOT INTENDED TO PROVIDE TOTAL VENTILATORY REQUIREMENTS OF THE PATIENT AND IS NOT FOR USE DURING FIELD TRANSPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22812817&pc=OZD | 1220648-2025-45804 | 2025-07-22 04:00 | Death | ABIOMED, INC. | 2025-08-14 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT IMPLANTED WITH AN IMPELLA CP EXPERIENCED BLEEDING AND REQUIRED A TRANSFUSION OF RED BLOOD CELLS. THE PATIENT WAS TRANSPORTED FOR A PERIPHERAL ANGIOGRAM AND CODED UPON ARRIVAL IN THE LAB, AND EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22812855&pc=OZD | 1220648-2025-45848 | 2025-07-25 04:00 | Death | ABIOMED, INC. | 2025-08-14 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT UNDERGOING RIGHT HEART CATHETERIZATION WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. WHILE ON SUPPORT, THE PATIENT SUFFERED AN ISCHEMIC STROKE WITH MINIMAL LEFT SIDED WEAKNESS NOTED. SUPPORT CONTINUED UNINTERRUPTED FOR TWO MORE DAYS, AT WHICH POINT THE PATIENT'S FAMILY MADE THE DECISION TO WITHDRAW CARE. THE PATIENT EXPIRED FOLLOWING PUMP REMOVAL. Manufacturer Narrative: PRODUCT STATUS: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THERE IS UNCERTAINTY WHAT HAPPENED WITH THE DEVICE AFTER THE PATIENT'S DEMISE. THEREFORE, AN EVALUATION/ANALYSIS OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE SECTION: POTENTIAL ADVERSE EVENTS- ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY¿ (INCLUDING VENTRICULAR PERFORATION). SECTION: WARNINGS, CAUTIONS & PRECAUTIONS- ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22791365&pc=LJT | 3006260740-2025-05799 | 2025-07-30 04:00 | Death | C.R. BARD, INC. (BASD) -3006260740 | 2025-08-13 04:00 | LJT | POWERPORT SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 6F | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Skin Inflammation/ Irritation | K072549 |  | 1 | Event Description: IT WAS REPORTED THROUGH LITIGATION PROCESS THAT SOMETIME LATER PORT PLACEMENT PROCEDURE, THE PATIENT DEVELOPED SKIN IMPAIRMENT AND THERE IS POTENTIAL FOR INFECTION. IT WAS ALSO REPORTED THAT THE PATIENT EXPIRED, AND THE CAUSE OF DEATH WAS NOT PROVIDED. Manufacturer Narrative: H11: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WAS PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. DATE OF DEATH FOR THIS PATIENT WAS NOT PROVIDED, DATE OF DEATH UPDATED AS (B)(6) 1900 FOR THE MDR SUBMISSION REQUIREMENT. SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22791627&pc=KFM | 8010762-2025-0000348 | 2025-07-10 04:00 | Death | MAQUET CARDIOPULMONARY GMBH | 2025-08-13 04:00 | KFM | HEART LUNG MACHINE | Improper Flow or Infusion | Insufficient Information | K991864 |  | 1 | Event Description: THE EVENT OCCURRED IN GERMANY. IT WAS REPORTED THAT DURING PREPARATION FOR A PLANNED CT SCAN ¿ PLANNED IN-HOUSE PATIENT TRANSPORT WITH VA ECMO FROM THE INTENSIVE CARE UNIT TO THE CT ¿ THE ROTAFLOW DRIVE UNIT WAS POSITIONED IN THE PATIENT BED AS PER INSTRUCTIONS. THE DRIVE PUMP WAS RUNNING CORRECTLY AT A FLOW RATE OF 2.6 L/MIN. SUDDENLY, THE ROTAFLOW FLOW INDICATOR ALERTED, AND THE FLOW INDICATOR WAS NO LONGER READABLE. AT THE SAME TIME, A SIGNIFICANTLY LOUDER ROTATION NOISE FROM THE DRIVE PUMP WAS AUDIBLE. AN IMMEDIATE, COMPLETE INSPECTION REVEALED THAT AIR HAD ENTERED THE SYSTEM. THE ARTERIAL LINE WAS IMMEDIATELY CLAMPED, AND THE ECMO SPEED WAS COMPLETELY REDUCED. AFTER IMMEDIATELY CLAMPING THE ARTERIAL LINE AND REDUCING THE ECMO SPEED, THE PERFUSIONISTS WERE IMMEDIATELY INFORMED, WITH THE URGENT INSTRUCTION TO IMMEDIATELY TRANSFER A FILLED, READY-TO-USE SYSTEM TO THE INTENSIVE CARE UNIT. THE SUBSEQUENT INSPECTION OF THE ECMO SYSTEM TOGETHER WITH THE ATTENDING PHYSICIANS REVEALED NO VISIBLE DAMAGE TO THE ECMO. ANALYSIS: DURING THE UKR'S ANALYSIS OF THE AFFECTED PRODUCTS, NO UNDESIRABLE SIDE EFFECTS, NO MALFUNCTION, NO APPLICATION ERROR, AND NO INADEQUACY OF THE INFORMATION PROVIDED WERE IDENTIFIED. ALTHOUGH WE HAVE EXHAUSTED OUR ANALYTICAL CAPABILITIES, WE CANNOT COMPLETELY RULE OUT THE ABOVE-MENTIONED ERROR POSSIBILITIES, AND THEREFORE, WE ARE REPORTING THE FATAL PATIENT INJURY RESULTING IN DEATH. DUE TO THE POLICE INVESTIGATION, THE DISPOSABLE PRODUCTS ARE NO LONGER WITHIN OUR CONTROL. THE LOT INFORMATION FOR THE INVOLVED DISPOSABLE PRODUCTS ARE THEREFORE NOT AVAILABLE. THE OTHER HARDWARE IS SECURELY STORED AT THE UKR. DUE TO THE REPORTED DEATH A REPORT IS REQUIRED. COMPLAINT ID: (B)(4). Manufacturer Narrative: THE INVESTIGATION IS ONGOING. FURTHER INFORMATION HAS BEEN REQUESTED BUT HAS NOT YET BEEN RECEIVED. A FOLLOW UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. NOTE: THIS EVENT OCCURRED ON THE GERMAN MARKET. IT IS A SIGNIFICANTLY SIMILAR DEVICE TO "BASE UNIT, ROTAFLOW DRIVE¿ WHICH IS SOLD IN THE USA UNDER PREMARKET SUBMISSION NUMBER K991864. FOR SECTION D4 ALL AVAILABLE IDENTIFYING INFORMATION IS FOR THE OUT OF THE US DEVICE, SUBJECTED TO THIS REPORT. THEREFORE, NO GUDID INFORMATION EXISTS. FURTHERMORE, UDI INFORMATION (PRIMARY DI NUMBER) PROVIDED IN D4 IS FOR ROTAFLOW DRIVE WITH CATALOG NUMBER 701022161.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22792018&pc=BZD | 2518422-2025-048882 | 2022-06-02 04:00 | Death | RESPIRONICS, INC. | 2025-08-13 04:00 | BZD | DREAMSTATION CPAP | Degraded | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE ALLEGING PATIENT HAS PASSED AWAY AND NO ONE WILL USE THE DEVICE ANYMORE CALLED FOR A RETURN SHIPPING LABEL. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22792552&pc=NKM | 2135147-2025-04492 | 2025-07-21 04:00 | Death | ABBOTT MEDICAL | 2025-08-13 04:00 | NKM | MITRACLIP | Physical Resistance/Sticking | Renal Failure; Unspecified Tissue Injury; No Clinical Signs, Symptoms or Conditions | P100009 |  | 1 | Event Description: IT WAS REPORTED THIS WAS A MITRACLIP PROCEDURE TO TREAT DEGENERATIVE MITRAL REGURGITATION (MR) WITH A GRADE OF 4+. THE CLIP WAS ADVANCED INTO THE ANATOMY AND POSITIONED. THE FIRST ESTABLISHING FINAL ARM ANGLE (EFAA)WAS SUCCESSFUL, WITH THE CLIP FULLY LOCKED. THE LOCK LINE WAS RAVELED WITH NO KNOTS ON EITHER END. HOWEVER, DURING THE REMOVAL OF LOCK LINE, LOCK LINE CAN BE HEARD IN CONTACT WITH THE SYSTEM, AND SOME RESISTANCE WAS FELT. AT THIS POINT THE LOCK LINE COULD NOT BE REMOVED FURTHER, WITH APPROXIMATELY 40% OF THE LINE REMOVED. DEPLOYMENT WAS RESUMED WITH THE LOCK LINE STILL IN THE SYSTEM. SECOND EFAA WAS ACHIEVED, GRIPPER LINE WAS REMOVED SUCCESSFULLY, AND THE MANDREL WAS DETACHED. THE CLIP DELIVERY SYSTEM (CDS) WAS CONTINUED TO BE REMOVED BACK INTO THE GUIDE, HOWEVER EVEN WITH THE UNDER-STRADDLING, THE LOCK LINE COULD NOT BE REMOVED. IMAGING SHOWED TWO LINES STILL LOOPED AROUND THE CLIP, AND ANY ATTEMPTS TO REMOVED LINE FURTHER, RESULTED IN THE CLIP BEING PULLED INTO THE ATRIUM AND OPEN THE VALVE. THE PATIENT WAS SENT TO SURGERY FOR CLIP REMOVAL AND VALVE REPAIR, WHICH WAS SUCCESSFUL. IT WAS LATER REPORTED THAT THE PATIENT HAD DIED ON (B)(6) 2025. IT IS NOTED THAT THE CAUSE OF DEATH WAS RENAL FAILURE POST SURGERY DUE TO A PROLONGED PERIOD IN ICU WITH INTUBATION, AND BEING UNABLE TO RECOVER. IT WAS ALSO NOTED THAT THERE HAD BEEN LEAFLET DAMAGE AFTER SURGICAL REMOVED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22792657&pc=DSQ | 2916596-2025-05345 | 2024-05-31 04:00 | Death | THORATEC CORPORATION | 2025-08-13 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT AND FAMILY DECIDED TO WITHDRAW LEFT VENTRICULAR ASSIST DEVICE (LVAD) SUPPORT, AND THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS COMFORT CARE MEASURES. THE DEVICE WAS WORKING AS INTENDED AT THAT TIME, AND THE PATIENT'S OUTCOME WAS NOT DEVICE OR THERAPY RELATED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6) , AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICLE ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22793516&pc=DQA | 2936999-2025-00421 | 2025-07-14 04:00 | Death | COVIDIEN LLC | 2025-08-13 04:00 | DQA | LIBRA BEDSIDE | Device Alarm System | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NONE |  | 1 | Event Description: IT WAS REPORTED THAT, AT THE TIME OF THE PATIENT'S DEATH, THE ALARM SOUND OF THE MONITOR COULD NOT BE CONFIRMED. THE SAT MESSAGE AND TREND DATA INFORMATION INDICATED ALARM CONDITIONS, INCLUDING "PULSE RATE LOWER LIMIT" AND "PULSE SEARCH." HOWEVER, DURING SUBSEQUENT TESTING OF THE MONITOR, THE ALARM SOUND WAS CONFIRMED. THE DEVICE MALFUNCTION WAS NOT RELATED TO THE PATIENT'S DEATH. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22793963&pc=MPB | 22793963 | 2025-07-23 04:00 | Death | MOZARC MEDICAL | 2025-08-13 04:00 | MPB | MARHURKUR CATHETER | Loss of or Failure to Bond; Fluid/Blood Leak | Exsanguination; Hypovolemia |  |  | 1 | Event Description: PT PRESENTED FOR HIS ROUTINE HEMODIALYSIS TREATMENT ON (B)(6) 2025. PRETREATMENT VITAL SIGNS WERE BLOOD PRESSURE 95/66, PULSE 103, RESPIRATIONS 18, TEMPERATURE 96.4. AT 145 PM, TREATMENT WAS INITIATED AT A BLOOD FLOW RATE OF 400 MIS/ MIN. AT 230 PM, THE RN NOTED A POOL OF BLOOD ON THE FLOOR NEAR THE PATIENT. UPON EXAMINATION, IT WAS DISCOVERED THE VENOUS (BLUE) HUB OF THE CENTRAL VENOUS CATHETER HAD SEPARATED FROM THE LUMEN. THE HUB REMAINED ATTACHED TO THE BLOOD LINE AND THE HEMACLIP WAS PRESENT ON THE LINE. THE LUMEN WAS CLAMPED. NORMAL SALINE AND OXYGEN WERE GIVEN TO THE PATIENT. THE PATIENT WAS TAKEN TO THE HOSPITAL VIA EMS. UPON PRESENTATION TO THE EMERGENCY DEPARTMENT, THE PATIENT WAS IN CARDIAC ARREST. HEMOGLOBIN DRAWN AT THE ED RESULTED AT 3.6. MULTIPLE INTERVENTIONS WERE ATTEMPTED FOR THE PATIENT INCLUDING INTUBATION, ACLS MEDICATION ADMINISTRATION, CONTINUED CPR AND DEFIBRILLATION, AS WELL AS THE ADMINISTRATION OF BLOOD PRODUCTS. DESPITE THESE INTERVENTIONS, SPONTANEOUS CIRCULATION COULD NOT BE ACHIEVED AND TIME OF DEATH WAS CALLED AT 343 PM.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22794037&pc=NPS | 2135147-2025-04493 | 2016-03-01 05:00 | Death | ABBOTT MEDICAL | 2025-08-13 04:00 | NPS | TRICLIP DELIVERY SYSTEM (TCDS) | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P230007 |  | 1 | Event Description: THE ARTICLE "HEART FAILURE HOSPITALIZATIONS AND CLINICAL OUTCOMES IN PATIENTS UNDERGOING TRICUSPID TRANSCATHETER EDGE-TO-EDGE REPAIR: INSIGHTS FROM EUROTR" WAS REVIEWED. THE ARTICLE PRESENTED A PROSPECTIVE MULTI-CENTER STUDY, TO EVALUATE THE PREVALENCE, PROGNOSTIC SIGNIFICANCE, AND PREDICTORS OF HEART FAILURE HOSPITALIZATION (HFH) BEFORE AND AFTER TRICUSPID TRANSCATHETER EDGE-TO-EDGE REPAIR (T-TEER). DEVICES MENTIONED INCLUDE TRICLIP AND PASCAL. THE ARTICLE CONCLUDED THAT IN THE EUROTR COHORT, A HISTORY OF HFH WAS HIGHLY PREVALENT AND ASSOCIATED WITH WORSE CLINICAL OUTCOMES. AMONG HIGH-RISK PATIENTS WITH SYMPTOMATIC TR, T-TEER SIGNIFICANTLY LOWERED HFH RISK, WITH RESIDUAL TR GRADE=2 BEING THE KEY PREDICTOR FOR REDUCED HFH INCIDENCE. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS MARCO METRA AT UNIVERSITY OF BRESCIA, BRESCIA, ITALY WITH CORRESPONDING EMAIL METRAMARCO@LIBERO.IT. THE TIME FRAME OF THE STUDY WAS MARCH 2016 TO FEBRUARY 2024. A TOTAL OF 1000 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH 57% RECEIVED AN ABBOTT DEVICE. AVERAGE AGE 80 AND MAJORITY GENDER WAS FEMALE. COMORBIDITIES INCLUDED ARTERIAL HYPERTENSION, DYSLIPIDEMIA, DIABETES, MYOCARDIAL INFARCTION, CORONARY ARTERY DISEASE, PERIPHERAL ARTERY DISEASE, STROKE, ATRIAL FIBRILLATION/FLUTTER, COPD, CARDIAC SURGERY, PRIOR TRICUSPID SURGERY, RIGHT VENTRICLE LEAD, HEART FAILURE, TRICUSPID REGURGITATION. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF MITRACLIP WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING ARTERIAL HYPERTENSION, DYSLIPIDEMIA, DIABETES, MYOCARDIAL INFARCTION, CORONARY ARTERY DISEASE, PERIPHERAL ARTERY DISEASE, STROKE, ATRIAL FIBRILLATION/FLUTTER, COPD, CARDIAC SURGERY, PRIOR TRICUSPID SURGERY, RIGHT VENTRICLE LEAD, HEART FAILURE, TRICUSPID REGURGITATION. COMPLICATIONS REPORTED INCLUDED DEATH AND HEART FAILURE HOSPITALIZATION; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. B3: EVENT DATE IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: HEART FAILURE HOSPITALIZATIONS AND CLINICAL OUTCOMES IN PATIENTS UNDERGOING TRICUSPID TRANSCATHETER EDGE-TO-EDGE REPAIR: INSIGHTS FROM EUROTR.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22794337&pc=NGV | 2124215-2025-54635 | 2025-07-18 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-08-13 04:00 | NGV | SYSTEM, APPENDAGE CLOSURE, LEFT ATRIAL | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: IT WAS REPORTED A PATIENT DEATH OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS PERFORMED, AND A WATCHMAN CLOSURE DEVICE WAS IMPLANTED. THE PATIENT WAS DISCHARGED. IT WAS REPORTED THAT AT AN UNKNOWN TIME POINT AFTER DISCHARGE, THE PATIENT DIED. NO FURTHER INFORMATION WAS PROVIDED. Manufacturer Narrative: B3 DATE OF EVENT : AWARE DATE WAS USED AS EVENT DATE AS ACTUAL EVENT DATE WAS NOT KNOWN. B2 DATE OF DEATH : AWARE DATE WAS USED AS DEATH DATE AS ACTUAL DATE OF DEATH WAS NOT KNOWN. D1: BRAND NAME IS BLANK BECAUSE THE DEVICE IS KNOWN TO BE A WATCHMAN LAA CLOSURE DEVICE, BUT THE EXACT BRAND NAME IS UNKNOWN EVEN AFTER GOOD FAITH EFFORTS WERE EXHAUSTED. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22794352&pc=KFM | 3013876692-2025-00066 | 2025-07-10 04:00 | Death | MAQUET CARDIOPULMONARY GMBH | 2025-08-13 04:00 | KFM | HEART LUNG MACHINE | Improper Flow or Infusion | Insufficient Information | K991864 |  | 1 | Event Description: THE EVENT OCCURRED IN GERMANY. IT WAS REPORTED THAT DURING PREPARATION FOR A PLANNED CT SCAN ¿ PLANNED IN-HOUSE PATIENT TRANSPORT WITH VA ECMO FROM THE INTENSIVE CARE UNIT TO THE CT ¿ THE ROTAFLOW DRIVE UNIT WAS POSITIONED IN THE PATIENT BED AS PER INSTRUCTIONS. THE DRIVE PUMP WAS RUNNING CORRECTLY AT A FLOW RATE OF 2.6 L/MIN. SUDDENLY, THE ROTAFLOW FLOW INDICATOR ALERTED, AND THE FLOW INDICATOR WAS NO LONGER READABLE. AT THE SAME TIME, A SIGNIFICANTLY LOUDER ROTATION NOISE FROM THE DRIVE PUMP WAS AUDIBLE. AN IMMEDIATE, COMPLETE INSPECTION REVEALED THAT AIR HAD ENTERED THE SYSTEM. THE ARTERIAL LINE WAS IMMEDIATELY CLAMPED, AND THE ECMO SPEED WAS COMPLETELY REDUCED. AFTER IMMEDIATELY CLAMPING THE ARTERIAL LINE AND REDUCING THE ECMO SPEED, THE PERFUSIONISTS WERE IMMEDIATELY INFORMED, WITH THE URGENT INSTRUCTION TO IMMEDIATELY TRANSFER A FILLED, READY-TO-USE SYSTEM TO THE INTENSIVE CARE UNIT. THE SUBSEQUENT INSPECTION OF THE ECMO SYSTEM TOGETHER WITH THE ATTENDING PHYSICIANS REVEALED NO VISIBLE DAMAGE TO THE ECMO. ANALYSIS: DURING THE UKR'S ANALYSIS OF THE AFFECTED PRODUCTS, NO UNDESIRABLE SIDE EFFECTS, NO MALFUNCTION, NO APPLICATION ERROR, AND NO INADEQUACY OF THE INFORMATION PROVIDED WERE IDENTIFIED. ALTHOUGH WE HAVE EXHAUSTED OUR ANALYTICAL CAPABILITIES, WE CANNOT COMPLETELY RULE OUT THE ABOVE-MENTIONED ERROR POSSIBILITIES, AND THEREFORE WE ARE REPORTING THE FATAL PATIENT INJURY RESULTING IN DEATH. DUE TO THE POLICE INVESTIGATION, THE DISPOSABLE PRODUCTS ARE NO LONGER WITHIN OUR CONTROL. THE LOT INFORMATION FOR THE INVOLVED DISPOSABLE PRODUCTS ARE THEREFORE NOT AVAILABLE. THE OTHER HARDWARE IS SECURELY STORED AT THE UKR. DUE TO THE REPORTED DEATH A REPORT IS REQUIRED. COMPLAINT ID: (B)(4). Manufacturer Narrative: THE INVESTIGATION IS ONGOING. FURTHER INFORMATION HAS BEEN REQUESTED BUT HAS NOT YET BEEN RECEIVED. A FOLLOW UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. NOTE: THIS EVENT OCCURRED ON THE GERMAN MARKET. IT IS A SIGNIFICANTLY SIMILAR DEVICE TO "BASE UNIT, ROTAFLOW DRIVE¿ WHICH IS SOLD IN THE USA UNDER PREMARKET SUBMISSION NUMBER K991864. FOR SECTION D4 ALL AVAILABLE IDENTIFYING INFORMATION IS FOR THE OUT OF THE US DEVICE, SUBJECTED TO THIS REPORT. THEREFORE, NO GUDID INFORMATION EXISTS. FURTHERMORE, UDI INFORMATION (PRIMARY DI NUMBER) PROVIDED IN D4 IS FOR ROTAFLOW DRIVE WITH CATALOG NUMBER 701022161.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22794368&pc=DRA | 2124215-2025-56361 | 2025-07-18 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-08-13 04:00 | DRA | FARADRIVE | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Low Blood Pressure/ Hypotension |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED A TEAR IN THE DISTAL ILIAC ARTERY AND LOW BLOOD PRESSURE. THE PATIENT DIED. AFTER PULSE FIELD ABLATION (PFA) PROCEDURE USING A FARADRIVE SHEATH AND PRIOR TO PATIENT DISCHARGE, TREATMENT AND DIAGNOSIS FOR BLOOD LOSS AND LOW BLOOD PRESSURES WERE DONE WHILE THE PATIENT WAS IN POST-RECOVERY. THERE HAD BEEN NO INDICATIONS OF PROCEDURE COMPLICATIONS WHILE IN THE ELECTROPHYSIOLOGY LAB. THE PATIENT WAS BROUGHT TO SURGERY LATE IN THE NIGHT. IN ADDITION TO THE SURGERY MEDICAL INTERVENTIONS INCLUDED BLOOD TRANSFUSIONS AND INTUBATION; THE PATIENT SPENT TIME IN THE INTENSIVE CARE UNIT (ICU). A COMPUTED TOMOGRAPHY SCAN WAS PERFORMED AT SOME POINT. THE PATIENT DIED THE NEXT MORNING, ABOUT 17 HOURS AFTER THE PROCEDURE. THE OFFICIAL CAUSE OF DEATH WAS AS A TEAR IN THE DISTAL ILIAC ARTERY. THE DEVICE IS NOT EXPECTED TO BE RETURNED FOR ANALYSIS DUE TO DISPOSAL. Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION, BUT IT WAS NOT AVAILABLE BECAUSE THE DEVICE WAS DISCARDED. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22794441&pc=DSQ | 2916596-2025-05336 | 2025-04-10 04:00 | Death | THORATEC CORPORATION | 2025-08-13 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Unspecified Infection | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY DUE TO INFECTION. THE OUTCOME WAS CONSIDERED DEVICE OR THERAPY RELATED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: THE PUMP REMAINS IN USE SUPPORTING THE PATIENT. A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF INFECTION COULD NOT BE CONCLUSIVELY DETERMINED. INFECTION IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. INFECTION HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22794632&pc=DSQ | 2916596-2025-05393 | 2025-08-07 04:00 | Death | THORATEC CORPORATION | 2025-08-13 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE PATIENT OUTCOME, THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS NOT PROVIDED. Manufacturer Narrative: SECTION A: ADDITIONAL PATIENT INFORMATION CANNOT BE PROVIDED DUE TO PERSONAL DATA PRIVACY LEGISLATION/POLICY. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. PUMP PERFORMED AS INTENDED, AND THE DEVICE WAS NOT EXPLANTED. DUE TO PATIENT PRIVACY REGULATION NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22794641&pc=OZP | 2032227-2025-235076 | 2025-06-29 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-08-13 04:00 | OZP | PUMP MMT-1884L MM780G BLE 3.0 MG/ML | Adverse Event Without Identified Device or Use Problem | Hyperglycemia; Vomiting; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P160017 |  | 1 | Event Description: INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE CUSTOMER HAD PASSED AWAY ON (B)(6) 2025. THE BLOOD GLUCOSE VALUE OF 1000 MG/DL. THE CUSTOMER ALSO REPORTED VOMITING AND TREATED WITH EMERGENCY MEDICAL SERVICES. THE EVENT INVOLVED PRODUCTS MMT-1884L, MMT-342G, AND MMT-431AG. TROUBLESHOOTING WAS NOT PERFORMED. IT WAS UNKNOWN WHETHER THE CUSTOMER USING THE PUMP AT THE TIME OF DEATH OR NOT. NO PRODUCT RETURN IS REQUIRED FOR MMT-1884L. NO PRODUCT RETURN IS REQUIRED FOR MMT-342G. NO PRODUCT RETURN IS REQUIRED FOR MMT-431AG. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY, IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE, THEREFORE, CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22794844&pc=NPT | 9617601-2025-00942 | 2025-07-08 04:00 | Death | MEDTRONIC MEXICO S. DE R.L. DE CV | 2025-08-13 04:00 | NPT | EVOLUT FX PLUS VALVE | Patient Device Interaction Problem | Stroke/CVA; Vascular Dissection | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE IMPLANT OF A TRANSCATHETER VALVE, THE PROCEDURE WAS PERFORMED THROUGH AXILLARY ACCESS. FOLLOWING THE IMPLANT PROCEDURE, A DISSECTION WAS OBSERVED AND A STENT WAS PLACED TO ADDRESS THE DISSECTION. HOWEVER, THE PATIENT DIED DUE TO A STROKE. PER THE PHYSICIAN, THE PATIENT'S TORTUOUS ANATOMY CONTRIBUTED TO THE EVENT. Manufacturer Narrative: SECTION D REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: BRAND NAME EVOLUT FX DCS; PRODUCT ID D-EVOLUTFX-2329 (UNKNOWN SERIAL/LOT); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE N/A; EXPLANT DATE N/A. THE CODES PRESENT IN SECTION H6 CORRESPOND TO MULTIPLE COMPONENTS/PRODUCTS THAT COMPRISE THIS REPORTED EVENT. B2. THE PATIENT'S DATE OF DEATH IS UNKNOWN. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22795050&pc=MLV | 2135147-2025-04495 | 2025-07-19 04:00 | Death | ABBOTT MEDICAL | 2025-08-13 04:00 | MLV | AMPLATZER MUSCULAR VSD OCCLUDER | Adverse Event Without Identified Device or Use Problem | Hemolysis | P040040 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 6MM AMPLATZER VSD MUSCULAR OCCLUDER WAS SELECTED FOR IMPLANT. PRIOR TO VSD PROCEDURE, AN UNKNOWN SIZED TRICLIP WAS IMPLANTED WITH NO REPORTED ISSUES. THE VSD OCCLUDER WAS IMPLANTED TO CLOSE A MITRAL VALVULAR LEAK. ON (B)(6) 2025, THE PATIENT PASSED AWAY DUE TO HEMOLYSIS. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22795061&pc=FSO | 9615010-2025-00018 | 2025-07-15 04:00 | Death | CARL ZEISS MEDITEC AG (OBERKOCHEN) | 2025-08-13 04:00 | FSO | KINEVO 900 | Optical Problem | Hemorrhage/Bleeding | N/A |  | 1 | Event Description: A HEALTH CARE PROFESSIONAL (HCP) REPORTED THAT THE OPMI KINEVO 900 STOPPED FOCUSING DURING A BRAIN OPERATION AFTER BLEEDING OF THE BRAIN HAD STARTED. IT WAS FURTHER REPORTED, THAT THE PATIENT LOST 7 LITERS OF BLOOD. THE DEVICE WAS WORKING WITHOUT ANY ISSUES UPON INSPECTION; NO FURTHER REPORTED ISSUES AFTER THE INCIDENT. THE SURGEON NOTED THAT THE BLEEDING OCCURRED BEFORE THE ASSUMED ISSUE WITH THE FOCUS. THE HEMORRHAGE WAS UNDER CONTROL BEFORE THE LOSS OF FOCUS AND RESTART OF THE DEVICE. NO HARM WAS REPORTED DIRECTLY FROM THE INCIDENT. THE DEVICE PROBLEM WAS RESOLVED BY SHUTTING DOWN AND RESTARTING THE MICROSCOPE. THE PATIENT DIED THREE DAYS AFTER SURGERY. THE REASON FOR THE DEATH REMAINS UNKNOWN TO DATE. Manufacturer Narrative: THE DEVICE ITSELF AS WELL AS THE LOGFILES WERE EVALUATED AND THERE IS NO INDICATION THAT THERE WAS A MALFUNCTION OF THE OPMI KINEVO 900 THAT COULD HAVE CONTRIBUTED TO THE FOCUSING PROBLEMS AND COMPLICATIONS DURING SURGERY. THERE SEEMS TO BE NO INDICATION THAT REASONABLY SUGGESTS THAT THE DEVICE HAS CAUSED OR CONTRIBUTED TO THE PATIENT'S DEATH OR THAT A MALFUNCTION OF THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THIS INCIDENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22795836&pc=NIM | 2124215-2025-56013 | 2016-01-01 05:00 | Death | SILK ROAD MEDICAL | 2025-08-13 04:00 | NIM | ENROUTE TRANSCAROTID STENT | Adverse Event Without Identified Device or Use Problem | Pulmonary Embolism; Stroke/CVA; Myocardial Infarction; Respiratory Failure; Heart Failure/Congestive Heart Failure |  |  | 1 | Event Description: IT WAS REPORTED VIA LITERATURE THAT SOME PATIENTS WHO UNDERWENT A TRANSCAROTID ARTERY REVASCULARIZATION (TCAR) PROCEDURE EXPERIENCED AN ARRAY OF ADVERSE EVENTS LEADING TO DEATH. THIS STUDY WAS A RETROSPECTIVE CHART REVIEW OF 750 PATIENTS FROM JANUARY 2016 TO JANUARY 2022 WHO UNDERWENT TCAR. IN THE PERIOPERATIVE PERIOD, DEFINED AS 30 DAYS POST TCAR, THE DEATH RATE WAS REPORTED AS 1.3 %. REPORTED CAUSES OF DEATH INCLUDED STROKE IN 3 PATIENTS, MYOCARDIAL INFARCTION (MI) IN 1, RESPIRATORY FAILURE IN 2, AND CONGESTIVE HEART FAILURE (CHF) IN 2. IN A REVIEW OF LONG-TERM OUTCOMES, DEFINED AS GREATER THAN 30 DAYS POST TCAR, THE DEATH RATE WAS REPORTED AS 5.9%. REPORTED CAUSES OF DEATH INCLUDED RESPIRATORY FAILURE IN 8 PATIENTS, PULMONARY EMBOLUS IN 1, MESENTERIC ISCHEMIA IN 1, STROKE IN 3, VARIOUS FORMS OF CANCER IN 4, AND 18 FROM UNKNOWN CAUSES. NO FURTHER INFORMATION WAS PROVIDED TO BOSTON SCIENTIFIC. Manufacturer Narrative: B2. THE DATE OF DEATH WAS NOT PROVIDED FOR ANY OF THE REPORTED DEATHS. B3. EVENT DATE WAS ESTIMATED BASED ON THE EARLIEST PROCEDURE DATE NOTED IN THE LITERATURE ARTICLE. DETAILS OF THE EVENT, INCLUDING PRODUCT INFORMATION, WERE NOT PROVIDED TO BSC. BECAUSE THE PRODUCT LOT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) NUMBER AND OTHER SPECIFIC PRODUCT INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. GEORGE, M. J., HUSMAN, R., DAKOUR-ARIDI, H., TANAKA, A., MADISON, M., MOTAGANAHALLI, R., LECKIE, K., & WANG, S. K. (2022). DUAL INSTITUTIONAL EXPERIENCE WITH TRANSCAROTID ARTERY REVASCULARIZATION. VASCULAR AND ENDOVASCULAR SURGERY, 57(1), 35-40. HTTPS://DOI.ORG/10.1177/15385744221127846.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22795917&pc=NPT | 9617601-2025-00949 | 2025-07-17 04:00 | Death | MEDTRONIC MEXICO S. DE R.L. DE CV | 2025-08-13 04:00 | NPT | EVOLUT FX PLUS VALVE | Patient Device Interaction Problem | Myocardial Infarction | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT PRIOR TO THE IMPLANT OF A TRANSCATHETER VALVE, A PRE-IMPLANT BALLOON AORTIC VALVULOPLASTY (BAV) WAS PERFORMED, DUE TO IT BEING A STANDARD FOR THE IMPLANTING PHYSICIAN. FOLLOWING THE IMPLANT OF THE VALVE, A MYOCARDIAL INFARCTION OCCURRED AND ACORONARY OCCLUSION WAS NOTED. SUBSEQUENTLY, AN ATTEMPT WAS MADE TO ACCESS THE CORONARY ARTERY AND EXPLANT THE TRANSCATHETER VALVE. ADDITIONALLY, OPEN-HEART SURGERY WAS PERFORMED. ALTHOUGH THE VALVE WAS EXPLANTED, THE PATIENT DIED. PER THE PHYSICIAN, THE CAUSE OF DEATH WAS DUE TO THE CORONARY OCCLUSION. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID {{D-EVOLUTFX-2392}}; PRODUCT LOT/SERIAL NUMBER {{UNKNOWN}}; PRODUCT TYPE: {{0195 HEART VALVES}}; IMPLANT DATE {{NA}}; EXPLANT DATE {{NA}} MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22796257&pc=FSA | 2183887-2025-00002 | 2025-05-01 04:00 | Death | EZ WAY INC | 2025-08-13 04:00 | FSA | 500 LB PREMIER LIFT WITH SCALE | Adverse Event Without Identified Device or Use Problem | Fall |  |  | 1 | Event Description: ONE OF THE CORNERS OF THE SLING LET LOOSE AND THE RESIDENT FELL ON THE FLOOR. Manufacturer Narrative: EZ WAY REPRESENTATIVE WAS THERE ON (B)(6) 2025 FOR AN IN-SERVICE AND REVIEW ON USAGE. MET WITH EMPLOYEE. REVIEWED SLINGS AND HARNESSES. THIS FACILITY IS USING EMERALD SUPPLY UNIVERSAL HARNESSES AND SLINGS FOR ALL UNITS AND LIFTING/TRANSFERS. EZ WAY REPRESENTATIVE LET THEM KNOW THAT OUR NEWEST INSTRUCTIONS LEAVE THE CHOICE UP TO THE FACILITY, BUT WE RECOMMEND NOT TO MIX VENDOR PRODUCTS ESPECIALLY ON THE SIT TO STAND. SOME OTHER VENDOR'S SLINGS COULD BE INTERCHANGEABLE, BUT EZ WAY HAS NOT DONE TESTING TO CONFIRM THIS. THEY REVIEWED TRANSFER WITH INCIDENT. FACILITY PERSONNEL WASN'T SURE IF ISSUE WAS WITH REPOSITIONING OF HANDLES OR SIZING OF THE CHAIR. DID TRAINING WITH STAFF AND REVIEWED ALL ANGLES ON HOW TO APPROACH CHAIR. REVIEWED POSITIONING AND TECHNIQUE. STAFF SAID IT HAS BEEN A WHILE SINCE THEY HAVE BEEN USING EZ WAY ACCESSORIES. THE WHOLE FACILITY IS USING DIFFERENT SLINGS HARNESS THAT THEY SAY ARE UNIVERSAL. ALL THEIR ACCESSORIES ARE FROM EMERALD SUPPLY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22796268&pc=DQX | 9617601-2025-00952 | 2025-07-23 04:00 | Death | MEDTRONIC MEXICO S. DE R.L. DE CV | 2025-08-13 04:00 | DQX | CONFIDA BRECKER CURVE GUIDEWIRE | Patient Device Interaction Problem | Low Blood Pressure/ Hypotension; Cardiac Perforation | K132623 |  | 1 | Event Description: IT WAS REPORTED THAT PRIOR TO THE IMPLANT OF A TRANSCATHETER VALVE, A PRE-IMPLANT BALLOON AORTIC VALVULOPLASTY (BAV) WAS PERFORMED DUE TO CALCIFICATION. UPON THE FIRST DEPLOYMENT ATTEMPT, A VENTRICULAR PERFORATION WAS OBSERVED. SUBSEQUENTLY, THE VALVE WAS FULLY DEPLOYED AND THE PROCEDURE WAS CONVERTED TO OPEN-HEART SURGICAL REPAIR TO ADDRESS THE PERFORATION. A RIGHT CORONARY ARTERY (RCA) STENT WAS PLACED AFTER THE PATIENT'S CHEST WAS CLOSED; HOWEVER, THE PATIENT'S BLOOD PRESSURE DID NOT RECOVER AND THE PATIENT DIED DUE TO THE PERFORATION. PER THE PHYSICIAN, THE GUIDEWIRE CAUSED THE VENTRICULAR PERFORATION AND DEATH. PER THE PHYSICIAN, THE PATIENT'S CALCIFIED ANATOMY CONTRIBUTED TO THE EVENT. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID EVFXPLUS-26 (SERIAL: (B)(6)); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE (B)(6) 2025; EXPLANT DATE N/A. PRODUCT ID {D-EVOLUTFX-2329}; PRODUCT LOT/SERIAL NUMBER {UNKNOWN}; PRODUCT TYPE: {0195-HEARTVALVES}; IMPLANT DATE {N/A}; EXPLANT DATE {N/A}. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22796318&pc=DSQ | 2916596-2025-05234 | 2025-06-30 04:00 | Death | THORATEC CORPORATION | 2025-08-13 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Hypervolemia; Multiple Organ Failure; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS HOSPITALIZED FROM ON (B)(6) 2025 DUE TO HYPERVOLEMIA. THE PATIENT'S RIGHT VENTRICULAR FUNCTION WAS SEVERELY REDUCED. DURING THIS TIME PALLIATIVE CARE WAS CONSULTED TO TALK WITH PATIENT AND FAMILY ABOUT GOALS. ON (B)(6) 2025, PATIENT¿S CODE STATUS WAS CHANGED TO DO NOT RESUSCITATE. OUTPATIENT HOSPICE WAS CONSULTED AND SAW THE PATIENT FOR THE INITIAL VISIT ON (B)(6) 2025. THE HOSPICE NURSE NOTED THAT ON (B)(6) 2025, THE PATIENT¿S SPOUSE CONTACTED THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) TEAM. THEY STATED THAT THEY WERE ALONE WITH THE PATIENT AND SCARED. THE LVAD COULD BE HEARD BEEPING IN THE BACKGROUND. THE SPOUSE WAS ASKED IF THE PATIENT WAS BREATHING AND THEY CHECKED, THEN STATED THEY DIDN'T THINK SO. THE PATIENT PASSED AWAY.Event Description: THE PATIENT'S CAUSE OF DEATH WAS HEART FAILURE AND MULTI-SYSTEM ORGAN FAILURE. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22796321&pc=CBK | 2518422-2025-108966 | 2025-08-06 04:00 | Death | RESPIRONICS, INC. | 2025-08-13 04:00 | CBK | TRILOGY 100 | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K083526 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION FROM A CUSTOMER STATING THAT A CANNULA CAME OUT FROM A CATHETER MOUNT AND IT WAS NOT NOTICED, AND A PATIENT HAS PASSED AWAY. ON (B)(6) 2025, AT 1000 THERE WAS A CALL FROM THE POLICE. AT 1045 THE CUSTOMER CALLED BACK THE POLICE AND INFORMED THEM THAT THE CANNULA CAME OFF FROM THE CATHETER MOUNT, AND AN ALARM WENT OFF. HOWEVER, THE CARE-TAKER (THE PATIENT'S HUSBAND) FELL ASLEEP AND DID NOT HEAR IT, AND THEN THE PATIENT PASSED AWAY. IT WAS CONFIRMED BY THE POLICE AND THE PATIENT'S FAMILY THAT THE ALARM SOUNDS WORK WITHOUT ANY PROBLEMS. THE POLICE BELIEVE THAT THE EVENT DOES NOT REQUIRE POLICE INVOLVEMENT. THE PATIENT USES A TRILOGY 100 VENTILATOR DUE TO AMYOTROPHIC LATERAL SCLEROSIS (ALS). THERE WAS MEDICAL INTERVENTION, BUT IT IS UNKNOWN WHAT MEDICAL INTERVENTIONS WERE PERFORMED. IT IS NOTED THAT THE CAUSE OF DEATH IS SPECULATED TO BE FROM ASPHYXIATION. THERE IS NO ALLEGATION THAT THE TRILOGY 100 VENTILATOR CONTRIBUTED TO THE DEATH. THE CLINICAL ASSESSMENT TEAM STATES THAT BASED ON INFORMATION PROVIDED AND AVAILABLE, THIS EVENT IS LIKELY AN UNINTENTIONAL USER ERROR, FAILURE TO RESPOND TO ALARMS. DEVICE EVALUATIONS AND LOG ATTACHMENTS ARE CURRENTLY PENDING. THEREFORE, DEVICE CAUSE AND/OR CONTRIBUTION TO THE REPORTED EVENT OF DEATH WHILE ON THE DEVICE IS LIKELY CONFIRMED, BASED ON THE EVIDENCE PRESENT. FURTHER GOOD FAITH EFFORTS AND INFORMATION GATHERING ARE PENDING. THE DEATH IS ASSESSED AS A SEVERITY 4. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM ANY MALFUNCTION OF THE DEVICE. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22796362&pc=DSQ | 2916596-2025-05225 | 2024-07-15 04:00 | Death | THORATEC CORPORATION | 2025-08-13 04:00 | DSQ | HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS TRANSITIONED TO COMFORT MEASURES ONLY (CMO) FOR UNKNOWN REASONS. THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) WAS TURNED OFF, AND THE PATIENT LIVED FOR ANOTHER TWO WEEKS THEN PASSED AWAY. THE LVAD WAS WORKING AS INTENDED PRIOR TO TURNING OFF. THE PATIENT'S OUTCOME WAS NOT DEVICE OR THERAPY RELATED. THE DEVICE WILL NOT BE EXPLANTED AND WILL NOT BE RETURNED FOR EVALUATION.Event Description: THE TRANSITION TO CMO WAS NOT SECONDARY TO AN ADVERSE EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU), REV. D IS CURRENTLY AVAILABLE. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22796713&pc=NKM | 2135147-2025-04502 | 2018-04-01 04:00 | Death | ABBOTT MEDICAL | 2025-08-13 04:00 | NKM | MITRACLIP® SYSTEM CLIP DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P100009 |  | 1 | Event Description: THE ARTICLE "CLINICAL IMPACT OF EARLY CHANGES IN SERUM ALBUMIN ON PATIENT PROGNOSIS AFTER MITRAL-TRANSCATHETER EDGE-TO-EDGE REPAIR" WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE MULTI CENTER STUDY, TO EVALUATE E THE ASSOCIATION BETWEEN EARLY ALBUMIN CHANGES AND CLINICAL OUTCOMES AFTER MITRAL TRANSCATHETER EDGE-TO-EDGE REPAIR (M-TEER); AND TO IDENTIFY THE ASSOCIATED FACTORS. DEVICES MENTIONED INCLUDE MITRACLIP. THE ARTICLE CONCLUDED THAT EARLY IMPROVEMENT IN -ALBUMIN MAY SERVE AS A MEASURE OF PROCEDURAL BENEFITS AND A SURROGATE MARKER FOR PREDICTING CLINICAL OUTCOMES AFTER M-TEER.. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS MASANORI YAMAMOTO MD AT NAGOYA HEART CENTER, NAGOYA AICHI, JAPAN WITH CORRESPONDING EMAIL : MASA-NORI@NMS.AC.JP. THE TIME FRAME OF THE STUDY WAS (B)(6) 2023. A TOTAL OF 3764 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH ALL RECEIVED AN ABBOTT DEVICE. AVERAGE AGE WAS 79 AND MAJORITY SEX WAS MALE. COMORBIDITIES INCLUDED FUNCTIONAL AND DEGENERATIVE MITRAL REGURGITATION, PRIOR HEART FAILURE HOSPITALIZATION, HEART FAILURE, HYPERTENSION, DYSLIPIDEMIA, DIABETES, CHRONIC KIDNEY DISEASE, DIALYSIS DEPENDENT, ATRIAL FIBRILLATION, LIVER CIRRHOSIS, COPD, PERIPHERAL ARTERY DISEASE. PERI AND POST-PROCEDURAL COMPLICATIONS INCLUDED DEATH, HEART FAILURE HOSPITALIZATION, BLEEDING, RECURRENT MITRAL REGURGITATION. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF MITRACLIP WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING FUNCTIONAL AND DEGENERATIVE MITRAL REGURGITATION, PRIOR HEART FAILURE HOSPITALIZATION, HEART FAILURE, HYPERTENSION, DYSLIPIDEMIA, DIABETES, CHRONIC KIDNEY DISEASE, DIALYSIS DEPENDENT, ATRIAL FIBRILLATION, LIVER CIRRHOSIS, COPD, PERIPHERAL ARTERY DISEASE. COMPLICATIONS REPORTED INCLUDED DEATH, HEART FAILURE, HOSPITALIZATION, BLEEDING, RECURRENT MITRAL REGURGITATION; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. B2: DEATH DATE IS ESTIMATED B3: EVENT DATE IS ESTIMATED D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: CLINICAL IMPACT OF EARLY CHANGES IN SERUM ALBUMIN ON PATIENT PROGNOSIS AFTER MITRAL-TRANSCATHETER EDGE-TO-EDGE REPAIR.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22796759&pc=MGB | 2024168-2025-08400 | 2020-09-01 04:00 | Death | ABBOTT VASCULAR INC. | 2025-08-13 04:00 | MGB | PERCLOSE PROGLIDE¿ | Adverse Event Without Identified Device or Use Problem | Unspecified Vascular Problem | P960043 |  | 1 | Event Description: THIS STUDY SOUGHT TO DESCRIBE A NOVEL RISK SCORE (THE PASSAGE-PUNCTURE SCORE) FOR TRANSFEMORAL ACCESS USING A SINGLE SUTURE-BASED CLOSURE SYSTEM. OUTCOMES RELATED TO THE PROGLIDE DEVICE INCLUDED: DEATH, ADDITIONAL VASCULAR CLOSURE DEVICE USED, BLEEDING, RED BLOOD CELL TRANSFUSION, PROLONGED HOSPITALIZATION, STROKE, PERIPROCEDURAL MYOCARDIAL INFARCTION, RENAL DYSFUNCTION, PSEUDOANEURYSMS, HEMATOMA, ARTERIAL NARROWING/OCCLUSION, AND IMPLANTATION OF A COVERED STENT. IT WAS CONCLUDED THAT THE PASSAGE-PUNCTURE SCORE IS EFFECTIVE IN DEFINING THE OPTIMAL ACCESS SITE FOR TRANSFEMORAL TAVI. THE SYSTEMATIC EVALUATION HAS THE POTENTIAL TO FURTHER REDUCE ACCESS-SITE COMPLICATIONS. SPECIFIC PATIENT INFORMATION IS DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ARTICLE, TITLED "A NOVEL RISK SCORE FACILITATES FEMORAL ARTERY ACCESS IN TRANSCATHETER AORTIC VALVE IMPLANTATION: PASSAGE-PUNCTURE SCORE". NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: THE DEVICES WERE NOT RETURNED. A REVIEW OF THE PRODUCTION RECORD FOR THIS PRODUCT WAS NOT PERFORMED BECAUSE THE PART NUMBERS AND LOT NUMBERS WERE NOT REPORTED, AND THE DEVICES WERE NOT RETURNED. BASED ON THE ARTICLE INFORMATION, A CONCLUSIVE CAUSE FOR THE REPORTED DEATH, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED. THE REPORTED PATIENT EFFECT OF DEATH IS LISTED IN THE PERCLOSE PROGLIDE INSTRUCTIONS FOR USE AS POTENTIAL ADVERSE EVENTS OF USE OF THE DEVICE. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING. THE ADDITIONAL PATIENT EFFECTS AND MALFUNCTIONS REPORTED IN THE ARTICLE ARE CAPTURED UNDER A SEPARATE MEDWATCH REPORT. B2, B3, B6: ESTIMATED DATES. D4: THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. ATTACHMENT: ARTICLE TITLED, "A NOVEL RISK SCORE FACILITATES FEMORAL ARTERY ACCESS IN TRANSCATHETER AORTIC VALVE IMPLANTATION: PASSAGE-PUNCTURE SCORE".
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22796793&pc=NKM | 2135147-2025-04504 | 2018-01-01 05:00 | Death | ABBOTT MEDICAL | 2025-08-13 04:00 | NKM | MITRACLIP® SYSTEM CLIP DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P100009 |  | 1 | Event Description: THE ARTICLE "SEX-SPECIFIC OUTCOMES OF TRANSCATHETER EDGE-TO-EDGE REPAIR FOR DEGENERATIVE MITRAL REGURGITATION: RESULTS FROM THE CLASP IID TRIAL" WAS REVIEWED. THE ARTICLE PRESENTED A PROSPECTIVE MULTI CENTER STUDY THAT WAS USED TO EVALUATE SEX-SPECIFIC DIFFERENCES IN CONTEMPORARY MITRAL TRANSCATHETER EDGE-TO-EDGE REPAIR (M-TEER). DEVICES MENTIONED INCLUDE MITRACLIP AND PASCAL. THE ARTICLE CONCLUDED THAT BOTH SEXES EXPERIENCED HIGH SURVIVAL AND FREEDOM FROM MAJOR ADVERSE EVENTS AND HEART FAILURE HOSPITALIZATION AT 1 YEAR POST¿M-TEER, WITH COMPARABLE IMPROVEMENTS IN ECHOCARDIOGRAPHIC, FUNCTIONAL, AND QUALITY-OF-LIFE MEASURES, DESPITE DIFFERENCES IN CLINICAL PRESENTATION AND THE PROCEDURE. [(B)(6).] THE TIME FRAME OF THE STUDY WAS 2018 TO 2021. A TOTAL OF 300 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH 94 RECEIVED AN ABBOTT DEVICE. COMORBIDITIES INCLUDED ATRIAL FIBRILLATION, CARDIOMYOPATHY, CORONARY ARTERY DISEASE, RENAL INSUFFICIENCY/FAILURE, DIABETES, DYSLIPIDEMIA, HYPERLIPIDEMIA, MYOCARDIAL INFARCTION, PERIPHERAL ARTERIAL DISEASE, ANEMIA, STROKE, TRANSIENT ISCHEMIC ATTACK, COPD, PACEMAKER/ICD, PREVIOUS PERCUTANEOUS CORONARY INTERVENTION, CABG, GASTROINTESTINAL BLEEDING, HOME OXYGEN USE, PRIOR AORTIC VALVE SURGERY, PRIOR TRICUSPID VALVE SURGERY, DEGENERATIVE MITRAL REGURGITATION, MITRAL ANNULAR CALCIFICATION. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF MITRACLIP WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING ATRIAL FIBRILLATION, CARDIOMYOPATHY, CORONARY ARTERY DISEASE, RENAL INSUFFICIENCY/FAILURE, DIABETES, DYSLIPIDEMIA, HYPERLIPIDEMIA, MYOCARDIAL INFARCTION, PERIPHERAL ARTERIAL DISEASE, ANEMIA, STROKE, TRANSIENT ISCHEMIC ATTACK, COPD, PACEMAKER/ICD, PREVIOUS PERCUTANEOUS CORONARY INTERVENTION, CABG, GASTROINTESTINAL BLEEDING, HOME OXYGEN USE, PRIOR AORTIC VALVE SURGERY, PRIOR TRICUSPID VALVE SURGERY, DEGENERATIVE MITRAL REGURGITATION, MITRAL ANNULAR CALCIFICATION. COMPLICATIONS REPORTED INCLUDED DEATH, STROKE, MYOCARDIAL INFARCTION, RENAL INJURY, SEVERE BLEEDING, HEART FAILURE HOSPITALIZATION, TRANSIENT ISCHEMIC ATTACK, TRANSCATHETER REINTERVENTION, RECURRENT MITRAL REGURGITATION, SURGICAL INTERVENTION, DIFFICULTY GRASPING; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: SEX-SPECIFIC OUTCOMES OF TRANSCATHETER EDGE-TO-EDGE REPAIR FOR DEGENERATIVE MITRAL REGURGITATION: RESULTS FROM THE CLASP IID TRIAL.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22796926&pc=FPA | 1416980-2025-04166 | 2025-06-24 04:00 | Death | BAXTER HEALTHCARE CORPORATION | 2025-08-13 04:00 | FPA | CLEARLINK/CONTINU-FLO | Detachment of Device or Device Component | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K203609 |  | 1 | Event Description: IT WAS REPORTED A CONTINU-FLO SOLUTION SET SEPARATED. DURING CONTINUOUS VENO-VENOUS HEMOFILTRATION, THE CLEARLINK SET "SUDDENLY BECAME DISCONNECTED AT THE Y PORT.¿ THE PART THAT WAS DISCONNECTED ¿SHOULD NOT BE REMOVABLE AND IS NOT AN AREA THAT IS MANIPULATED IN ANY WAY BY NURSING STAFF.¿ THE PATIENT'S CENTRAL VENOUS CATHETER HAD A CARRIER LINE INFUSING NORMAL SALINE ATTACHED TO IT THAT BROKE APART. DUE TO THIS DISCONNECTION, STRAIGHT AIR WAS INFUSING RIGHT INTO THE PATIENT'S CENTRAL LINE. THE REPORTER BELIEVED THIS EVENT LIKELY CAUSED THE PATIENT TO EXPERIENCE AN AIR EMBOLISM; HOWEVER, THE ECHOCARDIOGRAM WAS NEGATIVE FOR THE PRESENCE OF AIR. THE ¿PATIENT'S CONDITION WAS DECLINING DUE TO LIVER AND RENAL FAILURE¿, AND THE PATIENT WAS TRANSITIONED TO (COMFORT MEASURES ONLY) CMO WITHIN 24 HOURS AND PASSED AWAY. IT WAS FURTHER REPORTED THE MEDICAL TEAM CANNOT SAY THE DEATH WAS DIRECTLY RELATED DISCONNECTION. NO FURTHER INFORMATION WAS AVAILABLE AT THE TIME OF THIS REPORT. Manufacturer Narrative: SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22796994&pc=CBF | MW5174457 | 2025-07-09 04:00 | Death | OXYHEALTH LLC. | 2025-08-13 04:00 | CBF | OXYHEALTH FORTIUS 420 HYPERBARIC OXYGEN CHAMBER | Crack; Material Discolored; Fire | Insufficient Information |  |  | 1 | Event Description: (B)(6) WAS FOUND DECEASED INSIDE OF A HYPERBARIC OXYGEN CHAMBER AT HIS HEALTH CLINIC, (B)(6), ON THE LATE EVENING OF (B)(6) 2025. (B)(6) WIFE HAD NOT HEARD FROM HIM, AND AFTER SEVERAL UNSUCCESSFUL ATTEMPTS TO CONTACT HIM BY PHONE, WENT TO THE HEALTH CLINIC. UPON ARRIVAL, SHE FOUND THE FACILITY FILLED WITH SMOKE. EMERGENCY RESPONDERS WERE SUBSEQUENTLY NOTIFIED, AND ARRIVED AT THE HEALTH CLINIC AT APPROXIMATELY 10:50 P.M. MST(MOUNTAIN STANDARD TIME). EMERGENCY RESPONDERS DETERMINED THAT A FLASH FIRE OCCURRED WHILE (B)(6) HAD BEEN UNDERGOING SELF-TREATMENT INSIDE THE SEALED CHAMBER, AT A PRESSURE OF 2.2 ATA(ATMOSPHERE ABSOLUTE ). REMNANTS OF SEVERAL ELECTRONIC DEVICES WERE FOUND WITHIN THE CHAMBER, INCLUDING A HANDHELD COMMUNICATIONS DEVICE, HEADSET, EARPHONES, BRAIN-TAP, BATTERY POWER BANKS, CELL PHONE, ELECTRONIC TABLET, CHARGING CABLES, FAN AND POSSIBLY A RED LIGHT THERAPY DEVICE. BULKY BLACKENED MASSES WERE PRESENT WITH SIGNS OF SWELLING AND CRACKING, WHICH ARE SUSPECTED TO BE THE LITHIUM ION BATTERY PACKS POWERING MULTIPLE OF THOSE ITEMS. IT IS SUSPECTED THAT THE LITHIUM ION BATTERIES WERE A SOURCE OF IGNITION THAT CAUSED A FLASH FIRE WITHIN THE SEALED CHAMBER. FURTHER INSPECTION OF THE INCIDENT'S PHOTOGRAPHS REVEALED THAT THE CHAMBER MAINTAINED ITS INTEGRITY, WHILE CONTAINING THE FIRE EVENT. THE CHAMBER WAS A FORTIUS 420, SERIAL NO. (B)(6), MANUFACTURED BY PROFESSIONAL PIPING SYSTEMS, LLC, 406 E. PIONEER STREET, PHOENIX, ARIZONA 85040 FOR (B)(4). THE CHAMBER RECEIVED ITS FINAL INSPECTION ON JANUARY 19, 2024. THE 510K PREMARKET APPROVAL APPLICATION FOR THE CHAMBER WAS APPROVED ON NOVEMBER 15, 2004. (B)(6) LEARNED ABOUT THE INCIDENT FROM (B)(6). MS. (B)(6) THE INTERNATIONAL HYPERBARIC ASSOCIATION. THE (B)(6) THEREAFTER DISCUSSED THE INCIDENT WITH (B)(6), MAS, MS, CHT, SAFETY DIRECTOR AT (B)(6) AND RESEARCH PHYSIOLOGIST FOR THE HYPERBARIC INSTITUTE FOR RESEARCH AND TRAINING, AND POLICE DETECTIVES/FIRE MARSHALS. (B)(6) IS UNDER THE IMPRESSION THAT THE CHAMBER IS STILL AT THE FACILITY. MR. (B)(6) WERE NOT GRANTED ACCESS TO VIEW THE DEVICE. PROFESSIONAL PIPING SYSTEMS, LLC.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22797198&pc=NGV | 2124215-2025-56203 | 2025-07-18 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-08-13 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | Air Embolism | P130013 |  | 1 | Event Description: IT WAS REPORTED PATIENT DEATH OCCURRED. A LEFT ATRIAL APPENDAGE CLOSURE (LAAC) PROCEDURE WAS BEING PERFORMED USING INTRACARDIAC ECHO (ICE) IMAGING. A WATCHMAN TRUSTEER ACCESS SYSTEM (WAS) AND A 24MM WATCHMAN FLX PRO LAA CLOSURE DEVICE & DELIVERY SYSTEM (WDS) WAS SELECTED FOR USE. THE PATIENT WAS SUCCESSFULLY IMPLANTED WITH THE 40MM CLOSURE DEVICE ON (B)(6) 2025. THE PHYSICIAN NOTED HAVING TROUBLE WAKING THE PATIENT POST PROCEDURE, AND THE STROKE TEAM WAS CALLED. COMPUTED TOMOGRAPHY AND MAGNETIC RESONANCE IMAGING (MRI) WAS PERFORMED. THE CT WAS CLEAR, HOWEVER THE MRI SHOWED AN AIR EMBOLISM OCCURRED IN THE BRAIN. IT IS UNKNOWN HOW THE AIR ENTERED THE PATIENT. THE PATIENT REMAINED IN THE HOSPITAL AND WAS REPORTED TO HAVE DIED A WEEK POST PROCEDURE, ON (B)(6) 2025. THE OFFICIAL CAUSE OF DEATH WAS NOT PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22797251&pc=DSQ | 2916596-2025-05239 | 2025-07-30 04:00 | Death | THORATEC CORPORATION | 2025-08-13 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Renal Failure; Septic Shock | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO SEPTIC SHOCK. THE DEVICE WOULD NOT RETURN.Event Description: IT WAS ADDITIONALLY REPORTED THAT THE PATIENT HAD KIDNEY FAILURE LEADING TO SEPTIC SHOCK AND WAS INTUBATED. THE DEATH WAS NOT CONSIDERED DEVICE RELATED AND THE PUMP WAS WORKING AS EXPECTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE DEVICE WILL NOT BE EXPLANTED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. B, IS CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ OF THIS DOCUMENT LISTS SEPSIS, LOCAL INFECTION, RENAL DYSFUNCTION, RESPIRATORY FAILURE, AND DEATH AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. ADDITIONALLY, SECTION 6 ¿PATIENT CARE AND MANAGEMENT¿ INCLUDES INFORMATION REGARDING HOW TO PREVENT AND CONTROL INFECTION. THE HEARTMATE 3 LVAS PATIENT HANDBOOK, REV. B, CONTAINS SEVERAL SECTIONS WITH INFORMATION ABOUT HOW TO PREVENT AND CONTROL INFECTION. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22797519&pc=MHX | 1218950-2025-000463 | 2025-07-12 04:00 | Death | PHILIPS MEDIZIN SYSTEME BÖBLINGEN GMBH | 2025-08-13 04:00 | MHX | PATIENT INFORMATION CENTER IX | Defective Alarm | Insufficient Information | K153702 |  | 1 | Event Description: IT WAS REPORTED THERE WAS A DEATH OF A PATIENT BECAUSE THE CUSTOMER DID NOT RECEIVE AN ALARM. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTION WAS REPORTED; THEREFORE, GOOD FAITH EFFORTS ARE BEING PERFORMED. Manufacturer Narrative: PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE. REPORTING INSTITUTION PHONE NUMBER: (B)(6).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22797572&pc=QBJ | 3004753838-2025-221368 | 2025-06-04 04:00 | Death | DEXCOM, INC. | 2025-08-13 04:00 | QBJ | DEXCOM G7 CONTINUOUS GLUCOSE MONITORING SYSTEM | No Audible Alarm | Hyperglycemia | K213919 |  | 1 | Event Description: IT WAS REPORTED THAT AN ALERT/NOTIFICATION SETTINGS ISSUE OCCURRED. THE PATIENT¿S SPOUSE CONTACTED DEXCOM TECHNICAL SUPPORT BECAUSE SHE RECEIVED THE FAS RECEIVER LETTER. THE PATIENT¿S RECEIVER SERIAL NUMBER MATCHED THE ONE LISTED IN THE LETTER, WHICH IS WHAT PROMPTED HER TO CONTACT DEXCOM. THE PATIENT¿S SPOUSE REPORTED THAT THE PATIENT USED THE DEXCOM RECEIVER AS HIS PRIMARY CGM DISPLAY DEVICE AND IT HAD PREVIOUSLY WORKED WELL. IT WAS ALWAYS WITH HIM IN HIS POCKET, ¿LIKE A FAMILY MEMBER.¿ SHE MENTIONED THAT HE HAD DEMENTIA, AND SHE MANAGED HIS CARE. THERE WAS ALSO A RECENTLY PURCHASED IPHONE ON WHICH THE PHARMACIST GRANDSON HAD INSTALLED THE G7 APP; HOWEVER, THIS WAS NOT HIS PRIMARY DISPLAY DEVICE. THE DAUGHTER WHO LIVES NEXT DOOR WAS A FOLLOWER THAT WAY. IN THE EARLY MORNING OF THE PATIENT¿S DEATH, THE SPOUSE INDICATED SHE DID NOT RECEIVE ANY HIGH GLUCOSE ALERT FROM THE RECEIVER. SHE HEARD NO SOUND FROM EITHER THE RECEIVER IN HIS POCKET, NOR FROM THE IPHONE. LOCATION OF THE IPHONE WAS NOT RELAYED. AROUND 0200 THE DAY OF HIS DEATH, SHE HAD BEEN AWAKE SINCE SHE DOES NOT SLEEP WELL. SHE CHECKED ON HIS RECEIVER AND NOTED THE CGM READING TO BE 375 MG/DL. SHE WAS CONCERNED BECAUSE SHE HAD RECEIVED NO AUDIO FOR THE HIGH GLUCOSE ALERT. SHE RETRIEVED INSULIN AND ADMINISTERED IT TO HIM. SHE DID CHORES FOR 2-3 HOURS AND AROUND 0445-0500 SHE CHECKED ON HIM AGAIN. AT THIS TIME, HE WAS UNRESPONSIVE AND THE CGM READING ON THE RECEIVER WAS STILL REPORTEDLY 375 MG/DL. AGAIN, SHE DID NOT RECEIVE ANY HIGH GLUCOSE ALERT FROM THE RECEIVER. SHE WENT TO RETRIEVE ANOTHER INSULIN INJECTION, BUT UPON HER RETURN, THE PATIENT WAS DECEASED. SHE DID NOT HAVE AN AUTOPSY PERFORMED. A RETURN GOODS KIT HAS BEEN SENT TO THE PATIENT¿S FAMILY TO RETURN THE RECEIVER FOR INVESTIGATION. A FOLLOW UP REPORT WILL BE SUBMITTED UPON DEVICE INVESTIGATION COMPLETION IF THE DEVICE IS RETURNED. A REVIEW OF THE PERFORMANCE DATA FROM THE MOBILE APP INDICATES THAT THE SENSOR SESSION STARTED AT 11:04 PM ON (B)(6) 2025. THE SESSION LASTED UNTIL 10:19 AM ON (B)(6) 2025. ACCORDING TO THE ALERTS SCHEDULE THE HIGH ALERT AND THE SENSOR ERROR (NO READINGS) ALERTS WERE SET TO MATCH PHONE SETTINGS. ON (B)(6) 2025 AT 11:34 PM THE USER RECEIVED A HIGH GLUCOSE ALERT WHICH REPEATED EVERY 5 MINUTES UNTIL ACKNOWLEDGED AT 12:03 PM ON (B)(6) 2025. ON (B)(6) 2025 THE USERS ESTIMATED GLUCOSE VALUES (EGVS) WERE ABOVE THEIR TARGET RANGE UP UNTIL 5:34 PM WHEN THEY DROPPED BELOW 350 MG/DL. AT 8:04 PM THE EGVS REACHED 350 MG/DL AND THE HIGH GLUCOSE ALERT OCCURRED AGAIN. THIS ALERT REPEATED EVERY 5 MINUTES UNTIL AUTO CLEARED AT 12:19 AM ON (B)(6) 2025. ON (B)(6) 2025 AT 2:29 AM THE USERS EGVS REACHED 350 MG/DL AND THE USER RECEIVED ANOTHER HIGH GLUCOSE ALERT. FROM 2:34 AM THROUGH 4:29 AM THE USERS EGVS CONTINUED TO RISE UNTIL REACHING HIGH (>400 MG/DL) AT 4:34 AM. THE EGVS THEN REMAINED HIGH (>400 MG/DL) AND THE ALERT REPEATED EVERY 5 MINUTES UNTIL 9:39 AM WHEN THE SENSOR BEGAN TO EXPERIENCE BRIEF SENSOR ISSUES. AT 10:00 AM A NO READING ALERT OCCURRED, WHICH REPEATED EVERY 5 MINUTES. THE BRIEF SENSOR ISSUES CONTINUED UNTIL 11:04 AM. SUBSEQUENTLY, SENSOR SESSION DATA ENDED FOR AN UNDISCERNIBLE REASON. NO ADDITIONAL PATIENT OR EVENT INFORMATION IS AVAILABLE.Event Description: SUBSEQUENT TO THE INITIAL MDR, ADDITIONAL INFORMATION IS REQUIRED. Manufacturer Narrative: (B)(4).
Manufacturer Narrative: (B)(4) DIABETES MELLITUS IS A KNOWN CAUSE OF HYPERGLYCEMIA.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22798001&pc=NIQ | 2024168-2025-08405 | 2025-05-29 04:00 | Death | ABBOTT VASCULAR INC. | 2025-08-13 04:00 | NIQ | XIENCE V | Adverse Event Without Identified Device or Use Problem | Unspecified Heart Problem | P070015 |  | 1 | Event Description: IT WAS REPORTED IN A RESEARCH SUMMARY ARTICLE THAT THE PURPOSE OF THE STUDY WAS TO ASSESS THE PERFORMANCE OF DIFFERENT DRUG ELUTING STENTS (DES) CURRENTLY USED IN PERCUTANEOUS CORONARY INTERVENTION (PCI) VIA A NETWORK META-ANALYSIS OF RELEVANT TRIALS. MAJOR DATABASES WERE SYSTEMATICALLY SEARCHED FOR RANDOMIZED CONTROLLED TRIALS (RCT) REPORTING 5 YEAR OUTCOMES OF CURRENTLY USED DES. TWENTY NINE RCTS INVOLVING 46,502 PATIENTS WERE INCLUDED IN THE ANALYSIS AND SIZE CURRENTLY USED DES - OSIRO, XIENCE (XIENCE V, PRIME, XPEDITION, ALPINE), RESOLUTE, NOBORI/BIOMATRIX, SYNERGY, AND PROMUS WERE ANALYZED. ALL COMPARISONS SHOWED NONSIGNIFICANT RESULTS FOR OUTCOMES AT 5-YEAR FOLLOW-UP. THE OUTCOMES INCLUDED DEFINITE/PROBABLE STENT THROMBOSIS, ALL CAUSE MORTALITY, CARDIAC DEATH, MYOCARDIAL INFARCTION, AND TARGET VESSEL REVASCULARIZATION. THE ANALYSIS REVEALED NO SIGNIFICANT DIFFERENCES IN 5-YEAR OUTCOMES AMONG THE VARIOUS DES. HOWEVER, SYNERGY AND PROMUS PERFORMED BEST FOR KEY OUTCOMES SUCH AS STENT THROMBOSIS, MORTALITY, AND MI, SUGGESTING THEIR POTENTIAL FOR FAVORABLE PERFORMANCE IN CLINICAL PRACTICE. ADDITIONAL DETAILS ARE FOUND IN THE ARTICLE "LONG-TERM (5-YEARS) OUTCOMES OF CURRENT DRUG-ELUTING STENTS IN PERCUTANEOUS CORONARY INTERVENTION: A NETWORK META-ANALYSIS OF RANDOMIZED CONTROLLED TRIALS". Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORDS FOR THIS PRODUCT COULD NOT BE REVIEWED AND A SIMILAR COMPLAINT QUERY COULD NOT BE PERFORMED BECAUSE THE PART NUMBER AND/OR LOT NUMBER WERE NOT REPORTED, AND THE DEVICE WAS NOT RETURNED. A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECT OF DEATH AND THE RELATIONSHIP TO PRODUCT, IF ANY, CANNOT BE DETERMINED. THE INFORMATION RECEIVED IN THE COMPLAINT WAS OBTAINED THROUGH A PROACTIVE LITERATURE SEARCH. SPECIFIC MODEL AND LOT NUMBERS WERE NOT AVAILABLE AND CAUSALITY BETWEEN THE DOCUMENTED DEVICE USAGE AND THE PATIENT EFFECTS COULD NOT BE ESTABLISHED. FACTORS THAT MAY CONTRIBUTE TO THE PERFORMANCE OUTCOMES LISTED IN THE ARTICLE INCLUDE, BUT ARE NOT LIMITED TO, DEVICE TO PATIENT INTERACTION, PATIENT ANATOMICAL CONDITIONS, DEVICE TO USER INTERACTION AND MANUFACTURING OR MATERIAL ISSUES. DUE TO THE LIMITED INFORMATION AVAILABLE, THE EXACT CAUSE CANNOT BE DETERMINED. THE REPORTED PRODUCT WAS AN UNSPECIFIED XIENCE. THE PRODUCT USED FOR THE COMPLAINT INVESTIGATION IS XIENCE SIERRA, AS THE ESTIMATED IMPLANT DATE WAS IN 2020, AND XIENCE SIERRA WAS THE PRIMARY XIENCE PRODUCT SOLD THAT YEAR IN THE UNITED STATES. THE REPORTED PATIENT EFFECT OF DEATH IS LISTED IN THE XIENCE SIERRA EVEROLIMUS ELUTING CORONARY STENT SYSTEMS INSTRUCTIONS FOR USE AS A KNOWN PATIENT EFFECT OF CORONARY STENTING PROCEDURES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN, OR LABELING; THEREFORE, NO PRODUCT-RELATED CORRECTIVE ACTION WILL BE IMPLEMENTED IN THIS CASE. B2, B3, C4, D6A: ESTIMATED DATES D4: THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. THE ADDITIONAL DEVICES REFERENCED IN B5 ARE FILED UNDER A SEPARATE MEDWATCH REPORT NUMBER. THE ADDITIONAL PATIENT EFFECTS REPORTED IN THE ARTICLE ARE CAPTURED UNDER A SEPARATE MEDWATCH REPORT. LITERATURE ATTACHMENT: ARTICLE TITLED, "LONG-TERM (5-YEARS) OUTCOMES OF CURRENT DRUG-ELUTING STENTS IN PERCUTANEOUS CORONARY INTERVENTION: A NETWORK META-ANALYSIS OF RANDOMIZED CONTROLLED TRIALS".
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22798046&pc=BZD | 2518422-2025-109043 | 2025-08-07 04:00 | Death | RESPIRONICS, INC. | 2025-08-13 04:00 | BZD | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | Adverse Event Without Identified Device or Use Problem | Dyspnea | K131982 |  | 1 | Event Description: THE MANUFACTURER HAS BEEN INFORMED OF A PATIENT'S DEATH WHO WAS USING A DREAMSTATION AUTO CPAP W/HUM/CELL, DOM DEVICE. THE REVIEW OF THE COMPLAINT FOUND THAT THE INFORMATION PROVIDED WAS INADEQUATE, PROMPTING THE NEED FOR A REPORT. CURRENTLY, THERE IS NO EVIDENCE TO SUGGEST THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE PATIENT'S DEATH. THE SPOUSE REPORTED THAT THE PATIENT HAD BEEN EXPERIENCING DIFFICULTY BREATHING AND WAS SEEING A PHYSICIAN FOR SLEEP APNEA. THE MANUFACTURER'S INVESTIGATION IS ONGOING, AND A FOLLOW-UP REPORT WILL BE PROVIDED UPON COMPLETION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22798134&pc=NIQ | 9612164-2025-03972 | 2025-07-23 04:00 | Death | MEDTRONIC IRELAND | 2025-08-13 04:00 | NIQ | ONYX FRONTIER | Inflation Problem | Cardiac Arrest; No Clinical Signs, Symptoms or Conditions | P160043 |  | 1 | Event Description: IT WAS REPORTED THAT INFLATION DIFFICULTIES OCCURRED DURING THE DEPLOYMENT OF ONE ONYX FRONTIER CORONARY DRUG-ELUTING STENT. UPON AT TEMPTED INFLATION, ONLY THE DISTAL EDGES OF THE BALLOON INFLATED IN THE OSTIAL LEFT ANTERIOR DESCENDING (LAD) ARTERY, AND THE STENT FAILED TO EXPAND. THE DEVICE HAD BEEN INSPECTED PRIOR TO USE, WITH NO ISSUES OBSERVED. THE LESION HAD BEEN PRE-DILATED BEFORE THE ATTEMPTED DEPLOYMENT, AND NO RESISTANCE WAS NOTED WHILE ADVANCING THE DEVICE TO THE LESION. THE SAME INFLATION DEVICE WAS USED FOR BOTH THE PRE-DILATION BALLOON AND THE STENT BALLOON. NOMINAL PRESSURE OF 12 ATM WAS APPLIED WHEN THE INFLATION DIFFICULTIES WERE IDENTIFIED. THE PATIENT EXPERIENCED CARDIAC ARREST, WAS INTUBATED, DEFIBRILLATED, AND TRANSFERRED TO THE INTENSIVE CARE UNIT (ICU). THE PHYSICIAN BELIEVES THE INFLATION DIFFICULTIES CAUSED OCCLUSION OF FLOW IN THE LAD. THE PATIENT SUBSEQUENTLY EXPIRED IN THE ICU. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22798191&pc=NIQ | 2024168-2025-08407 | 2025-05-29 04:00 | Death | ABBOTT VASCULAR | 2025-08-13 04:00 | NIQ | XIENCE V EVEROLIMUS ELUTING CORONARY STENT SYSTEM | Adverse Event Without Identified Device or Use Problem | Unspecified Heart Problem | P070015 |  | 1 | Event Description: IT WAS REPORTED IN A RESEARCH SUMMARY ARTICLE THAT THE PURPOSE OF THE STUDY WAS TO ASSESS THE PERFORMANCE OF DIFFERENT DRUG ELUTING STENTS (DES) CURRENTLY USED IN PERCUTANEOUS CORONARY INTERVENTION (PCI) VIA A NETWORK META-ANALYSIS OF RELEVANT TRIALS. MAJOR DATABASES WERE SYSTEMATICALLY SEARCHED FOR RANDOMIZED CONTROLLED TRIALS (RCT) REPORTING 5 YEAR OUTCOMES OF CURRENTLY USED DES. TWENTY NINE RCTS INVOLVING 46,502 PATIENTS WERE INCLUDED IN THE ANALYSIS AND SIZE CURRENTLY USED DES - OSIRO, XIENCE (XIENCE V, PRIME, XPEDITION, ALPINE), RESOLUTE, NOBORI/BIOMATRIX, SYNERGY, AND PROMUS WERE ANALYZED. ALL COMPARISONS SHOWED NONSIGNIFICANT RESULTS FOR OUTCOMES AT 5-YEAR FOLLOW-UP. THE OUTCOMES INCLUDED DEFINITE/PROBABLE STENT THROMBOSIS, ALL CAUSE MORTALITY, CARDIAC DEATH, MYOCARDIAL INFARCTION, AND TARGET VESSEL REVASCULARIZATION. THE ANALYSIS REVEALED NO SIGNIFICANT DIFFERENCES IN 5-YEAR OUTCOMES AMONG THE VARIOUS DES. HOWEVER, SYNERGY AND PROMUS PERFORMED BEST FOR KEY OUTCOMES SUCH AS STENT THROMBOSIS, MORTALITY, AND MI, SUGGESTING THEIR POTENTIAL FOR FAVORABLE PERFORMANCE IN CLINICAL PRACTICE. ADDITIONAL DETAILS ARE FOUND IN THE ARTICLE "LONG-TERM (5-YEARS) OUTCOMES OF CURRENT DRUG-ELUTING STENTS IN PERCUTANEOUS CORONARY INTERVENTION: A NETWORK META-ANALYSIS OF RANDOMIZED CONTROLLED TRIALS". Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORDS FOR THIS PRODUCT COULD NOT BE REVIEWED AND A SIMILAR COMPLAINT QUERY COULD NOT BE PERFORMED BECAUSE THE PART NUMBER AND/OR LOT NUMBER WAS NOT REPORTED, AND THE DEVICE WAS NOT RETURNED. A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT DEATH, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED; HOWEVER, THE SUBSEQUENT TREATMENT APPEARS TO BE RELATED TO THE OPERATIONAL CONTEXT OF THE PROCEDURE. THE REPORTED PATIENT EFFECT OF DEATH IS LISTED IN THE XIENCE V EVEROLIMUS ELUTING CORONARY STENT SYSTEMS INSTRUCTIONS FOR USE AS A KNOWN PATIENT EFFECT OF CORONARY STENTING PROCEDURES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING; THEREFORE, NO PRODUCT-RELATED CORRECTIVE ACTION WILL BE IMPLEMENTED IN THIS CASE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. B2, B3, C4, D6A: ESTIMATED DATES D4- THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. THE ADDITIONAL DEVICES REFERENCED IN B5 ARE FILED UNDER A SEPARATE MEDWATCH REPORT NUMBER. THE ADDITIONAL PATIENT EFFECTS REPORTED IN THE ARTICLE ARE CAPTURED UNDER A SEPARATE MEDWATCH REPORT. LITERATURE ATTACHMENT: ARTICLE TITLE: ¿LONG-TERM (5-YEARS) OUTCOMES OF CURRENT DRUG-ELUTING STENTS IN PERCUTANEOUS CORONARY INTERVENTION: A NETWORK META-ANALYSIS OF RANDOMIZED CONTROLLED TRIALS¿.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22798382&pc=BZD | 2518422-2025-109087 | 2022-07-04 04:00 | Death | RESPIRONICS, INC. | 2025-08-13 04:00 | BZD | DS2ADV AUTO CPAP W/HUMID CELL/BT | Adverse Event Without Identified Device or Use Problem | Respiratory Failure; Fibrosis | K200480 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION FROM THE PATIENT'S WIFE REPORTING THAT HER HUSBAND PASSED AWAY ON (B)(6) 2022, "DUE TO USE OF A CPAP DEVICE". THE PATIENT'S WIFE REPORTED THAT SHE DOES NOT KNOW IF "THE 1ST DEVICE HARMED HIM OR THE SECOND DEVICE." THE PATIENT INITIALLY HAD A RECALLED DEVICE THAT WAS "TRASHED AS ADVISED," AND THEN HE HAD A REPLACEMENT DEVICE, WHICH THEY NO LONGER HAVE. THE PATIENT'S WIFE REPORTED THAT THE DEATH CERTIFICATE STATED THAT HER HUSBAND DIED OF RESPIRATORY FAILURE AND PULMONARY FIBROSIS. THE PATIENT'S WIFE ALSO REPORTED THAT THE PATIENT USED THE DEVICE UNTIL HE WENT TO THE HOSPITAL AND THAT HE PASSED AWAY IN THE HOSPITAL THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22798393&pc=NIQ | 2024168-2025-08409 | 2025-05-29 04:00 | Death | ABBOTT VASCULAR | 2025-08-13 04:00 | NIQ | XIENCE PRIME EVEROLIMUS ELUTING CORONARY STENT SYSTEM | Adverse Event Without Identified Device or Use Problem | Unspecified Heart Problem | P110019 |  | 1 | Event Description: IT WAS REPORTED IN A RESEARCH SUMMARY ARTICLE THAT THE PURPOSE OF THE STUDY WAS TO ASSESS THE PERFORMANCE OF DIFFERENT DRUG ELUTING STENTS (DES) CURRENTLY USED IN PERCUTANEOUS CORONARY INTERVENTION (PCI) VIA A NETWORK META-ANALYSIS OF RELEVANT TRIALS. MAJOR DATABASES WERE SYSTEMATICALLY SEARCHED FOR RANDOMIZED CONTROLLED TRIALS (RCT) REPORTING 5 YEAR OUTCOMES OF CURRENTLY USED DES. TWENTY NINE RCTS INVOLVING 46,502 PATIENTS WERE INCLUDED IN THE ANALYSIS AND SIZE CURRENTLY USED DES - OSIRO, XIENCE (XIENCE V, PRIME, XPEDITION, ALPINE), RESOLUTE, NOBORI/BIOMATRIX, SYNERGY, AND PROMUS WERE ANALYZED. ALL COMPARISONS SHOWED NONSIGNIFICANT RESULTS FOR OUTCOMES AT 5-YEAR FOLLOW-UP. THE OUTCOMES INCLUDED DEFINITE/PROBABLE STENT THROMBOSIS, ALL CAUSE MORTALITY, CARDIAC DEATH, MYOCARDIAL INFARCTION, AND TARGET VESSEL REVASCULARIZATION. THE ANALYSIS REVEALED NO SIGNIFICANT DIFFERENCES IN 5-YEAR OUTCOMES AMONG THE VARIOUS DES. HOWEVER, SYNERGY AND PROMUS PERFORMED BEST FOR KEY OUTCOMES SUCH AS STENT THROMBOSIS, MORTALITY, AND MI, SUGGESTING THEIR POTENTIAL FOR FAVORABLE PERFORMANCE IN CLINICAL PRACTICE. ADDITIONAL DETAILS ARE FOUND IN THE ARTICLE "LONG-TERM (5-YEARS) OUTCOMES OF CURRENT DRUG-ELUTING STENTS IN PERCUTANEOUS CORONARY INTERVENTION: A NETWORK META-ANALYSIS OF RANDOMIZED CONTROLLED TRIALS". Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. B2, B3, C4, D6A: ESTIMATED DATES. D4: THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. THE ADDITIONAL DEVICES REFERENCED IN B5 ARE FILED UNDER A SEPARATE MEDWATCH REPORT NUMBER. THE ADDITIONAL PATIENT EFFECTS REPORTED IN THE ARTICLE ARE CAPTURED UNDER A SEPARATE MEDWATCH REPORT. LITERATURE ATTACHMENT: ARTICLE TITLE: ¿LONG-TERM (5-YEARS) OUTCOMES OF CURRENT DRUG-ELUTING STENTS IN PERCUTANEOUS CORONARY INTERVENTION: A NETWORK META-ANALYSIS OF RANDOMIZED CONTROLLED TRIALS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22798435&pc=NIQ | 2024168-2025-08411 | 2025-05-29 04:00 | Death | ABBOTT VASCULAR INC. | 2025-08-13 04:00 | NIQ | XIENCE XPEDITION | Adverse Event Without Identified Device or Use Problem | Unspecified Heart Problem | P110019 |  | 1 | Event Description: IT WAS REPORTED IN A RESEARCH SUMMARY ARTICLE THAT THE PURPOSE OF THE STUDY WAS TO ASSESS THE PERFORMANCE OF DIFFERENT DRUG ELUTING STENTS (DES) CURRENTLY USED IN PERCUTANEOUS CORONARY INTERVENTION (PCI) VIA A NETWORK META-ANALYSIS OF RELEVANT TRIALS. MAJOR DATABASES WERE SYSTEMATICALLY SEARCHED FOR RANDOMIZED CONTROLLED TRIALS (RCT) REPORTING 5 YEAR OUTCOMES OF CURRENTLY USED DES. TWENTY NINE RCTS INVOLVING 46,502 PATIENTS WERE INCLUDED IN THE ANALYSIS AND SIZE CURRENTLY USED DES - OSIRO, XIENCE (XIENCE V, PRIME, XPEDITION, ALPINE), RESOLUTE, NOBORI/BIOMATRIX, SYNERGY, AND PROMUS WERE ANALYZED. ALL COMPARISONS SHOWED NONSIGNIFICANT RESULTS FOR OUTCOMES AT 5-YEAR FOLLOW-UP. THE OUTCOMES INCLUDED DEFINITE/PROBABLE STENT THROMBOSIS, ALL CAUSE MORTALITY, CARDIAC DEATH, MYOCARDIAL INFARCTION, AND TARGET VESSEL REVASCULARIZATION. THE ANALYSIS REVEALED NO SIGNIFICANT DIFFERENCES IN 5-YEAR OUTCOMES AMONG THE VARIOUS DES. HOWEVER, SYNERGY AND PROMUS PERFORMED BEST FOR KEY OUTCOMES SUCH AS STENT THROMBOSIS, MORTALITY, AND MI, SUGGESTING THEIR POTENTIAL FOR FAVORABLE PERFORMANCE IN CLINICAL PRACTICE. ADDITIONAL DETAILS ARE FOUND IN THE ARTICLE "LONG-TERM (5-YEARS) OUTCOMES OF CURRENT DRUG-ELUTING STENTS IN PERCUTANEOUS CORONARY INTERVENTION: A NETWORK META-ANALYSIS OF RANDOMIZED CONTROLLED TRIALS". Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. B2, B3, C4, D6A: ESTIMATED DATES D4: THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. THE ADDITIONAL DEVICES REFERENCED IN B5 ARE FILED UNDER A SEPARATE MEDWATCH REPORT NUMBER. THE ADDITIONAL PATIENT EFFECTS REPORTED IN THE ARTICLE ARE CAPTURED UNDER A SEPARATE MEDWATCH REPORT. LITERATURE ATTACHMENT: ARTICLE TITLED, "LONG-TERM (5-YEARS) OUTCOMES OF CURRENT DRUG-ELUTING STENTS IN PERCUTANEOUS CORONARY INTERVENTION: A NETWORK META-ANALYSIS OF RANDOMIZED CONTROLLED TRIALS".
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22798565&pc=NIQ | 2024168-2025-08413 | 2025-05-29 04:00 | Death | ABBOTT VASCULAR INC. | 2025-08-13 04:00 | NIQ | XIENCE ALPINE | Adverse Event Without Identified Device or Use Problem | Unspecified Heart Problem | P110019 |  | 1 | Event Description: IT WAS REPORTED IN A RESEARCH SUMMARY ARTICLE THAT THE PURPOSE OF THE STUDY WAS TO ASSESS THE PERFORMANCE OF DIFFERENT DRUG ELUTING STENTS (DES) CURRENTLY USED IN PERCUTANEOUS CORONARY INTERVENTION (PCI) VIA A NETWORK META-ANALYSIS OF RELEVANT TRIALS. MAJOR DATABASES WERE SYSTEMATICALLY SEARCHED FOR RANDOMIZED CONTROLLED TRIALS (RCT) REPORTING 5 YEAR OUTCOMES OF CURRENTLY USED DES. TWENTY NINE RCTS INVOLVING 46,502 PATIENTS WERE INCLUDED IN THE ANALYSIS AND SIZE CURRENTLY USED DES - OSIRO, XIENCE (XIENCE V, PRIME, XPEDITION, ALPINE), RESOLUTE, NOBORI/BIOMATRIX, SYNERGY, AND PROMUS WERE ANALYZED. ALL COMPARISONS SHOWED NONSIGNIFICANT RESULTS FOR OUTCOMES AT 5-YEAR FOLLOW-UP. THE OUTCOMES INCLUDED DEFINITE/PROBABLE STENT THROMBOSIS, ALL CAUSE MORTALITY, CARDIAC DEATH, MYOCARDIAL INFARCTION, AND TARGET VESSEL REVASCULARIZATION. THE ANALYSIS REVEALED NO SIGNIFICANT DIFFERENCES IN 5-YEAR OUTCOMES AMONG THE VARIOUS DES. HOWEVER, SYNERGY AND PROMUS PERFORMED BEST FOR KEY OUTCOMES SUCH AS STENT THROMBOSIS, MORTALITY, AND MI, SUGGESTING THEIR POTENTIAL FOR FAVORABLE PERFORMANCE IN CLINICAL PRACTICE. ADDITIONAL DETAILS ARE FOUND IN THE ARTICLE "LONG-TERM (5-YEARS) OUTCOMES OF CURRENT DRUG-ELUTING STENTS IN PERCUTANEOUS CORONARY INTERVENTION: A NETWORK META-ANALYSIS OF RANDOMIZED CONTROLLED TRIALS". Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. B2, B3, C4, D6A: ESTIMATED DATES D4:THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. THE ADDITIONAL DEVICES REFERENCED IN B5 ARE FILED UNDER A SEPARATE MEDWATCH REPORT NUMBER. THE ADDITIONAL PATIENT EFFECTS REPORTED IN THE ARTICLE ARE CAPTURED UNDER A SEPARATE MEDWATCH REPORT. LITERATURE ATTACHMENT: ARTICLE TITLE: ¿LONG-TERM (5-YEARS) OUTCOMES OF CURRENT DRUG-ELUTING STENTS IN PERCUTANEOUS CORONARY INTERVENTION: A NETWORK META-ANALYSIS OF RANDOMIZED CONTROLLED TRIALS¿.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22781012&pc=NIQ | 9612164-2025-03928 | 2025-08-01 04:00 | Death | MEDTRONIC IRELAND | 2025-08-12 04:00 | NIQ | ONYX FRONTIER | Difficult to Remove; Device Dislodged or Dislocated; Positioning Problem | Cardiogenic Shock; Device Embedded In Tissue or Plaque | P160043 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE INVOLVING ONE ONYX FRONTIER CORONARY DRUG-ELUTING STENT, THE STENT DISLODGED FROM THE BALLOON. THE DISLODGMENT OCCURRED DURING REMOVAL OF THE DEVICE. THE DEVICE WAS INSPECTED PRIOR TO USE WITH NO ISSUES NOTED. NEGATIVE PREP WAS NOT PERFORMED. THE MID LEFT ANTERIOR DESCENDING ARTERY (LAD) HAD HIGH CALCIFICATION WITH 40% STENOSIS IN PROXIMAL AND 75% IN MID, THE LEFT MAIN (LM) HAD 60% STENOSIS WITH PROXIMAL FIBRO CALCIFICATION, THE FIRST DIAGONAL HAD 40% STENOSIS AND THE CIRCUMFLEX (CX) HAD 95% STENOSIS. THE LESION BEING TREATED WAS PRE-DILATED. THE DEVICE DID NOT PASS THROUGH A PREVIOUSLY-DEPLOYED STENT. RESISTANCE WAS ENCOUNTERED ADVANCING THE DEVICE AND NO EXCESSIVE FORCE WAS USED. ATTEMPTS TO REMOVE THE STENT BY SNARING AND USING SMALLER BALLOONS THROUGH THE STENT WERE UNSUCCESSFUL AND THE STENT WAS NOT ABLE TO BE REMOVED. NO RESISTANCE WAS ENCOUNTERED WITHDRAWING THE DEVICE AND NO EXCESSIVE FORCE WAS USED. CARDIAC SURGERY, WHICH WAS OFF-SITE WAS CALLED. CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED FOR THIRTY-FIVE (35) MINUTES AND THE PATIENT WAS RESUSCITATED TWICE. DUE TO BLOCKED ILIACS IT WAS NOT POSSIBLE TO PLACE BALLOON AND DUE TO BLOCKED SUBCLAVIAN IT WAS NOT POSSIBLE TO PLACE CENTRAL LINE. THERE WAS NO EVIDENCE RESTENOSIS OR THROMBUS. THE PATIENT WAS NOT ON DUAL ANTIPLATELET THERAPY (DAPT) AT TIME OF THE EVENT. THE PATIENT DIED. THE PHYSICIAN ASSESSED THAT THE CAUSE OF DEATH WAS CARDIOGENIC SHOCK DIRECTLY RELATED TO THE USE OF THE DEVICE. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22781711&pc=OZD | 1220648-2025-45735 | 2025-07-22 04:00 | Death | ABIOMED, INC. | 2025-08-12 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Tachycardia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT IN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA 5.5 DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. IT WAS REPORTED THAT DURING IMPELLA 5.5 SUPPORT, THE PATIENT EXPERIENCED VENTRICULAR TACHYCARDIA (VT) THAT REQUIRED DEFIBRILLATION. ON (B)(6) 2025 IT WAS NOTED THAT THE PATIENT EXPERIENCED ADDITIONAL EPISODES OF NON-SUSTAINED. THE FOLLOWING DAY, THE FAMILY DECIDED TO WITHDRAW CARE FROM THE PATIENT. CARE WAS WITHDRAWN, AND THE PATIENT EXPIRED. Manufacturer Narrative: ALTHOUGH THE PATIENT'S PRESENTING CONDITION MAY BE MORE LIKELY HAVE IMPACTED THE EVENT THE IMPELLA CANNOT BE EXCLUDED AS A POSSIBLE CONTRIBUTING FACTOR IN THE PATIENT'S DEMISE. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22781743&pc=NIK | 9614453-2025-03170 | 2025-07-03 04:00 | Death | MEDTRONIC EUROPE SARL | 2025-08-12 04:00 | NIK | COBALT¿ XT HF QUAD CRT-D MRI SURESCAN¿ | Appropriate Term/Code Not Available | Hematoma | P010031 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS PLACED ON HOSPICE APPROXIMATELY TWO WEEKS POST CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) CHANGE OUT. THE PATIENT WAS NOTED TO HAVE A HEMATOMA SINCE THE GENERATOR CHANGE. IT WAS FURTHER REPORTED THAT THE PATIENT PASSED AWAY 10 DAYS LATER. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22782088&pc=NPT | 2135147-2025-04462 | 2025-07-17 04:00 | Death | ABBOTT MEDICAL | 2025-08-12 04:00 | NPT | NAVITOR VISION | Adverse Event Without Identified Device or Use Problem |  | P190023 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 27MM NAVITOR VALVE WAS SELECTED FOR IMPLANT USING A LARGE FLEXNAV DELIVERY SYSTEM AND A LARGE NAVITOR LOADING SYSTEM. ON AN UNKNOWN DATE, THE PATIENT PASSED AWAY DUE TO OSTEAL RCA OCCLUSION. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22782185&pc=NPT | 2135147-2025-04463 | 2025-07-17 04:00 | Death | ABBOTT MEDICAL | 2025-08-12 04:00 | NPT | NAVITOR VISION | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Pericardial Effusion | P190023 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 29MM NAVITOR VALVE WAS SELECTED FOR IMPLANT USING A LARGE FLEXNAV DELIVERY SYSTEM AND A LARGE NAVITOR LOADING SYSTEM. THE NATIVE AORTIC ANNULUS MEASURED PERIMETER OF 82.2MM, LVOT 76.5MM, STJ 30.4MM, SOV'S >34MM. THE PATIENT HAD MILD LEAFLET CALCIUM; NO PERICARDIAL EFFUSION WAS OBSERVED. PRE-BAV WAS PERFORMED WITH A 24MM TRU BALLOON. DURING IMPLANT OF THE NAVITOR VALVE, THERE WAS "SOME DIFFICULTY CROSSING THE NATIVE VALVE INITIALLY WITH THE WIRE"; THERE WERE NO ISSUES CROSSING THE ROOT WITH THE DELIVERY SYSTEM. THE INITIAL DEPLOYMENT OF THE NAVITOR VALVE WAS TOO DEEP AND ONE RECAPTURE WAS PERFORMED. THE VALVE WAS SUCCESSFULLY IMPLANTED DURING SECOND DEPLOYMENT; NO LEAK WAS OBSERVED, NO GRADIENT PRESENT, AND THE DEPTH WAS 4MM ON NON AND 3MM ON LEFT. ECHOCARDIOGRAPHY WAS THEN PERFORMED, WHICH SHOWED EFFUSION IN THE LEFT VENTRICLE. THE PATIENT'S PRESSURE BEGAN TO DROP AND PERICARDIOCENTESIS WAS PERFORMED WITH NO IMPROVEMENT. THE PATIENT WAS INTUBATED AND TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) WAS PERFORMED. THE PATIENT WAS TRANSITIONED TO OPEN HEART SURGERY. THE SURGEON OBSERVED A HEMATOMA IN THE ROOT AREA. THE PATIENT PASSED DUE TO A HEMATOMA IN THE ROOT; THE CAUSE OF THE HEMATOMA WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22783064&pc=NVZ | 3008973940-2025-05979 | 2025-07-22 04:00 | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-08-12 04:00 | NVZ | AZURE XT DR MRI SURESCAN | Pacing Problem | Bradycardia; Pericardial Effusion; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P980035 |  | 1 | Event Description: IT WAS REPORTED THAT THE IMPLANTABLE PULSE GENERATOR (IPG) HAD FIXED BURST PACING. AFTER PLACEMENT OF A TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR), WHEN THE WIRE WAS REMOVED THE PATIENT PRESSURE BEGAN TO DROP AND LOSS OF CAPTURE OCCURRED. OUTPUTS WERE INCREASED AND RV CAPTURE WAS REGAINED. AN ECHOCARDIOGRAM REVEALED A PERICARDIAL EFFUSION AND THE PATIENT EXPERIENCED BRADYCARDIA. A CODE WAS CALLED AND CARDIOPULMONARY RESUSCITATION WAS PERFORMED. RESCUE WERE INEFFECTIVE AND THE PATIENT PASSED AWAY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22783067&pc=NVZ | 3004209178-2025-13641 | 2025-07-08 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-08-12 04:00 | NVZ | AZURE XT DR MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information; No Clinical Signs, Symptoms or Conditions | P980035 |  | 1 | Event Description: IT WAS REPORTED THAT THE IMPLANTABLE PULSE GENERATOR (IPG) MAY HAVE BEEN IMPLANTED AFTER THE USE BY DATE. THE IPG REMAINS IN USE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT PASSED AWAY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22783095&pc=DTB | 2649622-2025-21831 | 2025-07-22 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-08-12 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Failure to Capture | Bradycardia; Pericardial Effusion; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P930039 |  | 1 | Event Description: IIT WAS REPORTED THAT THE IMPLANTABLE PULSE GENERATOR (IPG) HAD FIXED BURST PACING. AFTER PLACEMENT OF A TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR), WHEN THE WIRE WAS REMOVED THE PATIENT PRESSURE BEGAN TO DROP AND LOSS OF CAPTURE OCCURRED. OUTPUTS WERE INCREASED AND RV CAPTURE WAS REGAINED. AN ECHOCARDIOGRAM REVEALED A PERICARDIAL EFFUSION AND THE PATIENT EXPERIENCED BRADYCARDIA. A CODE WAS CALLED AND CARDIOPULMONARY RESUSCITATION WAS PERFORMED. RESCUE WERE INEFFECTIVE AND THE PATIENT PASSED AWAY. Manufacturer Narrative: CONTINUATION OF D10: W1DR01 IPG DOI: (B)(6) 2023. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22783310&pc=DSQ | 3007042319-2025-01903 | 2025-08-07 04:00 | Death | HEARTWARE, INC. | 2025-08-12 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE VENTRICULAR ASSIST DEVICE (VAD) PATIENT EXPERIENCED HEART FAILURE AND SUBSEQUENTLY EXPIRED. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. CONTINUATION OF D10: DDBB1D4 ICD IMPLANT DATE: (B)(6) 2016. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR COMPLETION OF THE INVESTIGATION. PRODUCT EVENT SUMMARY: (B)(6) WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE MANUFACTURING DOCUMENTATION WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPERIENCED HEART FAILURE AND SUBSEQUENTLY EXPIRED. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, WORSENING HEART FAILURE AND DEATH ARE KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. BASED ON REVIEW OF PAST ADVERSE EVENTS FOR THIS PATIENT, IT WAS NOTED THAT THE PATIENT HAD A HISTORY OF WORSENING HEART FAILURE. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22783377&pc=DSQ | 2916596-2025-05324 | 2024-07-18 04:00 | Death | THORATEC CORPORATION | 2025-08-12 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Renal Failure; Septic Shock | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS SEPTIC SHOCK AND END STAGE RENAL FAILURE AT NURSING HOME. PATIENT AT NURSING HOME WITH END STAGE RENAL FAILURE, DEVICE WAS WORKING AS INTENDED AT TIME OF PATIENT DEATH. IT WAS REPORTED THAT THE PATIENT'S OUTCOME WAS NOT DEVICE OR THERAPY RELATED. THE DEVICE WOULD NOT BE EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REV. D AND THE HEARTMATE 3 PATIENT HANDBOOK, REV. A ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING SEPSIS, RENAL DYSFUNCTION, AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SEVERAL SECTIONS OF THE IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION, AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22783553&pc=DSQ | 3007042319-2025-01904 | 2025-06-19 04:00 | Death | HEARTWARE, INC. | 2025-08-12 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP/DRIVELINE | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Sepsis | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DEVELOPED A DRIVELINE INFECTION. THE PATIENT WAS ADMITTED TO THE HOSPITAL AND BECAME SEPTIC. THE PATIENT WAS TREATED WITH ANTIBIOTICS. THE PATIENT SUBSEQUENTLY EXPIRED DUE TO SEPSIS. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. D10 CONTINUED: 5076-52 LEAD IMPLANTED (B)(6) 2017. 459888 LEAD IMPLANTED (B)(6) 2017. 694765 LEAD IMPLANTED (B)(6) 2006. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22783670&pc=DSQ | 3007042319-2025-01906 | 2025-06-28 04:00 | Death | HEARTWARE, INC. | 2025-08-12 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Hemorrhagic Stroke | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WITH A VENTRICULAR ASSIST DEVICE (VAD) WAS FOUND UNRESPONSIVE AT HOME. THE INITIAL PRESENTING SYMPTOM WAS UNRESPONSIVENESS. UPON ARRIVAL AT THE EMERGENCY ROOM, A HEAD COMPUTED TOMOGRAPHY (CT) WAS CONDUCTED AND IDENTIFIED A MASSIVE HEMORRHAGIC STROKE. THE PATIENT WAS TRANSITIONED TO COMFORT MEASURES PER FAMILY WISHES AND SUBSEQUENTLY EXPIRED. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. CONTINUATION OF D10: 407652 & 6935M62 LEADS IMPLANT DATE: (B)(6) 2014. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22783822&pc=NPT | 2015691-2025-06556 | 2025-03-25 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-08-12 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Calcified; Insufficient Information | Dyspnea; Insufficient Information | P140031 |  | 1 | Event Description: AS REPORTED BY IMPLANT PATIENT REGISTRY, APPROXIMATELY 1 YEAR, 4 MONTHS POST TRANSFEMORAL TAVR PROCEDURE WITH A 20MM SAPIEN 3 ULTRA VALVE IN THE AORTIC POSITION THE VALVE WAS EXPLANTED DUE TO UNKNOWN REASONS. AN EDWARDS SURGICAL VALVE WAS IMPLANTED.Event Description: AS REPORTED BY IMPLANT PATIENT REGISTRY, APPROXIMATELY 1 YEAR, 4 MONTHS POST TRANSFEMORAL TAVR PROCEDURE WITH A 20MM SAPIEN 3 ULTRA VALVE IN THE AORTIC POSITION THE VALVE WAS EXPLANTED DUE TO UNKNOWN REASONS. AN EDWARDS SURGICAL VALVE WAS IMPLANTED. UPON REVIEW OF THE MEDICAL RECORDS RECEIVED, A REVIEW OF THE PATHOLOGY REPORT OF THE EXPLANTED TAVR REVEALED A FIBROUS VALVE WITH EXTENSIVE CALCIFIED PLAQUE. THE PATIENT HAD ENDORSED INTERMITTENT CP AND SOB. TO NOTE, ON (B)(6) 2025, THE PATIENT PRESENTED WITH AN NSTEMI. A LHC REVEALED MULTIVESSEL CORONARY DISEASE. THE PATIENT WAS STABILIZED AND WORKED UP FOR CABG AND AVR GIVEN ONGOING CONCERN FOR SEVERE BIOPROSTHETIC AS (AVA ' CONT EQU OF 0.58 CM2, AOV VMAX 3.95 M/SEC, AVMG OF 38 MMHG & DI 0.26). ON (B)(6) 2025, THE PATIENT UNDERWENT CABG X4, EXTENSIVE ENDARTERECTOMY OF AORTIC ROOT WITH AORTIC ROOT ENLARGEMENT AND AVR. DESPITE AGGRESSIVE MEDICAL MANAGEMENT, HYPOTENSION ENSUED WITH NOTED POSSIBLE VASOPLEGIC SYNDROME AND METHYLENE BLUE WAS ADMINISTRATION. THE PATIENT DEVELOPED CARDIOGENIC AND VASODILATORY SHOCK, LACTIC ACIDOSIS SECONDARY TO PROBABLE GUT ISCHEMIA, AN INCREASE IN VASOPRESSOR REQUIREMENT AND RV/LV FAILURE. ECMO CANNULATION WAS UNSUCCESSFUL DESPITE MULTIPLE ATTEMPTS DUE TO PAD/POOR ACCESS. DESPITE EXCELLENT CPR AND ACLS MANEUVERS, THE PATIENT PASSED AWAY. WHILE THE PATIENT HAD MULTIPLE CO-MORBIDITIES THAT LIKELY CAUSED OR CONTRIBUTED TO THEIR DEATH, THE MEDICAL RECORDS DISCHARGE DIAGNOSES OUTLINED THE TAVR DISEASE PROGRESSION AS A CONTRIBUTING FACTOR. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), NON-EMERGENT REOPERATION, EMERGENCY CARDIAC SURGERY, REOPERATION, DEVICE MIGRATION OR MALPOSITION REQUIRING INTERVENTION AND DEVICE EXPLANT ARE LISTED AS POTENTIAL RISKS ASSOCIATED WITH THE DEVICE AND TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. THE IFU CAUTIONS THAT LONG-TERM DURABILITY HAS NOT BEEN ESTABLISHED FOR THE VALVE. REGULAR MEDICAL FOLLOW-UP IS ADVISED TO EVALUATE VALVE PERFORMANCE. ACCELERATED DETERIORATION OF THE VALVE DUE TO CALCIFIC DEGENERATION MAY OCCUR IN CHILDREN, ADOLESCENTS, OR YOUNG ADULTS AND IN PATIENTS WITH AN ALTERED CALCIUM METABOLISM. VALVE RECIPIENTS SHOULD BE MAINTAINED ON ANTICOAGULANT/ANTIPLATELET THERAPY, EXCEPT WHEN CONTRAINDICATED, AS DETERMINED BY THEIR PHYSICIAN. THIS DEVICE HAS NOT BEEN TESTED FOR USE WITHOUT ANTICOAGULATION. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. VALVE REINTERVENTION MAY BE DUE TO VALVE DYSFUNCTION (E.G., SEVERE, OR CLINICALLY SIGNIFICANT PARAVALVULAR LEAK, CENTRAL LEAK, SIGNIFICANT MALPOSITION, EARLY/LATE ENDOCARDITIS, THROMBOSIS, STRUCTURAL VALVE DETERIORATION OR DEGENERATION) AND/OR OTHER REASONS UNRELATED TO THE EDWARDS DEVICES (E.G., TREATMENT OF OTHER PATIENT COMORBIDITIES). DURING THE MANUFACTURING PROCESS, ALL SAPIEN VALVES (ALL MODELS) ARE 100% VISUALLY INSPECTED FOR DEFECTS AND 100% TESTED UNDER PHYSIOLOGICAL BACKPRESSURE CONDITIONS PRIOR TO RELEASE FOR DISTRIBUTION. THEREFORE, IT IS HIGHLY UNLIKELY THAT A MANUFACTURING DEFECT OR DEVICE MALFUNCTION CONTRIBUTED TO THE EVENT. IN THIS CASE, DESPITE MULTIPLE INVESTIGATIONAL ATTEMPTS, IT WAS NOT POSSIBLE TO OBTAIN ADDITIONAL DETAILS REGARDING THIS EVENT. DUE TO LIMITED INFORMATION, A DEFINITIVE ROOT CAUSE WAS UNABLE TO BE DETERMINED AT THIS TIME. AT THE TIME OF THIS REPORT, THE INFORMATION AVAILABLE DOES NOT REASONABLY SUGGEST THERE WAS A MALFUNCTION OF THE EDWARDS DEVICE OR THAT THE USE OR MISS-USE OF THE DEVICE CAUSED OR CONTRIBUTED TO THE EVENT. THERE IS NO ALLEGATION OF DEFICIENCIES RELATED TO THE IDENTITY, QUALITY, DURABILITY, RELIABILITY, SAFETY, LABELING EFFECTIVENESS, OR PERFORMANCE OF THIS EDWARD DEVICE. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE INFORMATION WILL BE REVIEWED, AND THE FILE UPDATED ACCORDINGLY. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED DUE TO FILE REVIEW CORRECTION FINDINGS, SECTIONS G6, H2, B1, B2, B5, H1, H6: IMPACT CODE. INVESTIGATION IS STILL ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22783913&pc=DSQ | 2916596-2025-05205 | 2025-07-25 04:00 | Death | THORATEC CORPORATION | 2025-08-12 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Respiratory Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE SITE REQUESTED AN EVALUATION OF THE PUMP. THE CAUSE OF DEATH WAS RESPIRATORY FAILURE. THE PATIENT'S DEATH WAS NOT CONSIDERED DEVICE RELATED. THE LEFT VENTRICULAR ASSIST SYSTEM (LVAS) FUNCTIONED AS INTENDED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22784030&pc=NPT | 9617601-2025-00914 | 2024-10-11 04:00 | Death | MEDTRONIC MEXICO S. DE R.L. DE CV | 2025-08-12 04:00 | NPT | EVOLUT R TRANSCATHETER AORTIC VALVE | Gradient Increase; Insufficient Information; Patient Device Interaction Problem; Device Stenosis | Aortic Valve Stenosis; Chest Pain; Myocardial Infarction | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 7 YEARS FOLLOWING THE IMPLANT OF THE TRANSCATHETER BIOPROSTHETIC AORTIC VALVE A TRANSTHORACIC ECHOCARDIOGRAM (TTE) IDENTIFIED A MEAN GRADIENT OF 39 MILLIMETERS OF MERCURY (MM HG) AND A VMAX OF 4.06 METERS PER SECOND (M/SEC). AORTIC REGURGITATION WAS NOT PRESENT AT THAT TIME. APPROXIMATELY 8 MONTHS LATER CHEST PAIN OCCURRED. THE PATIENT WAS ADMITTED TO CARDIOLOGY DUE TO A NON-ST-ELEVATION MYOCARDIAL INFARCTION (NSTEMI). A CARDIAC CATHETERIZATION WAS PERFORMED. AS REPORTED, THE PATIENT TOLERATED AN AORTIC BALLOON VALVULOPLASTY. HOWEVER, THE PATIENT REQUIRED THE PLACEMENT OF A VENTRICULAR ASSISTED DEVICE AS THE PATIENT RAPIDLY DECOMPENSATED. THE DECISION WAS MADE TO WITHDRAW CARE. THE PATIENT DIED 2 DAYS FOLLOWING ONSET OF THE NSTEMI FOLLOWING THE CATHETERIZATION PROCEDURE. THE PHYSICIAN INDICATED THE EVENT WAS UNRELATED TO THE VALVE. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22784421&pc=DSQ | 2916596-2025-05332 | 2025-08-01 04:00 | Death | THORATEC CORPORATION | 2025-08-12 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY OF UNKNOWN CAUSES. DUE TO HOSPITAL POLICY, NO FURTHER INFORMATION WAS PROVIDED ON CAUSE OF DEATH. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. WHETHER THE DEVICE WOULD BE EXPLANTED WAS UNKNOWN. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT'S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REV. D IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22784437&pc=DRE | MW5174380 | 2025-03-04 05:00 | Death | BAYLIS MEDICAL COMPANY, INC. | 2025-08-12 04:00 | DRE | VERSACROSS CONNECT | Insufficient Information | Insufficient Information |  |  | 1 | Event Description: WATCHMAN IMPLANT PROCEDURE ATTEMPTED ON (B)(6) 2025. DEVICE CODE: 3190. PATIENT CODE: 4580. REFERENCE REPORT: MW5174379.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22784609&pc=NGV | 2124215-2025-55872 | 2025-04-28 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-08-12 04:00 | NGV | WATCHMAN FLX LEFT ATRIAL APPENDAGE CLOSURE DEVICE WITH DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: REPORTED VIA LITERATURE ARTICLE: IT WAS REPORTED PATIENT DEATH(S) OCCURRED. THIS RETROSPECTIVE STUDY INVESTIGATES THE SAFETY AND EFFECTIVENESS OF A SIMPLIFIED, NON-COMPLEX LEFT ATRIAL APPENDAGE OCCLUSION (LAAO) PROCEDURE IN ELDERLY PATIENTS AGED 80 YEARS AND OLDER, REFERRED TO AS THE "OLDEST-OLD," COMPARED TO YOUNGER PATIENTS. CONDUCTED AT A SINGLE CENTER, THE RETROSPECTIVE ANALYSIS INCLUDED 230 PATIENTS WITH NON-VALVULAR ATRIAL FIBRILLATION AND CONTRAINDICATIONS TO LONG-TERM ORAL ANTICOAGULATION. OF THESE, 43% WERE OLDEST-OLD. THE PROCEDURE UTILIZED VARIOUS OCCLUSION DEVICES, INCLUDING WATCHMAN? AND WATCHMAN FLX?, WITH A HIGH OVERALL TECHNICAL SUCCESS RATE OF 97.4%, SHOWING NO SIGNIFICANT DIFFERENCE BETWEEN AGE GROUPS. PROCEDURAL METRICS SUCH AS DURATION, FLUOROSCOPY TIME, AND RADIATION EXPOSURE WERE SIMILAR ACROSS GROUPS, AND COMPLICATIONS LIKE DEVICE EMBOLIZATION AND PERICARDIAL EFFUSION WERE INFREQUENT. OF THE 230 PATIENTS THAT UNDERWENT THE LAAO PROCEDURE, INTRAPROCEDURAL EVENTS OCCURRED AS FOLLOWS: INTRAPROCEDURAL DEVICE EMBOLIZATION, OCCURRED IN 3 PATIENTS INTRAPROCEDURAL PERICARDIAL EFFUSION IN 11 PATIENTS. THE MEDIAN HOSPITAL STAY WAS TWO DAYS FOR BOTH GROUPS. COMPLICATIONS REQUIRING IMMEDIATE MEDICAL INTERVENTION OCCURRED IN 8.1 VS. 3.8% OF PATIENTS. INTRAHOSPITAL MORTALITY (1.7 % OF CASES) WAS NOT STATISTICALLY SIGNIFICANTLY DIFFERENT BETWEEN THE TWO GROUPS. NO STROKES OR SYSTEMIC EMBOLISMS OCCURRED AFTER LAAO. MAJOR BLEEDING OCCURRED IN 3.0 VS. 3.1% OF PATIENTS UNTIL HOSPITAL DISCHARGE. THE VAST MAJORITY OF PATIENTS RECEIVED A COMBINATION OF ASA AND CLOPIDOGREL AFTER LAAO OF THE 230 PATIENTS THAT UNDERWENT THE LAAO PROCEDURE, INTRA HOSPITAL EVENTS OCCURRED AS FOLLOWS: ALL-CAUSE DEATH OCCURRED IN 4 PATIENTS PERICARDIAL EFFUSION IN 8 PATIENTS, WHICH INCLUDED PERICARDIAL TAMPONADE IN 3 PATIENTS ACUTE HEART FAILURE IN 1 PATIENT. DEVICE EMBOLIZATION IN 3 PATIENTS VENOUS THROMBOEMBOLISM IN 1 PATIENT LAAO REMOVAL IN 3 PATIENTS ACCESS SITE COMPLICATIONS: HEMATOMA IN 23 PATIENTS, PROLONGED BLEEDING IN 16 PATIENTS. ANEURYSM SPURIUM IN 2 PATIENTS. OF THE 230 PATIENTS THAT UNDERWENT THE LAAO PROCEDURE, THE INTRA HOSPITAL FOLLOW-UP EVENTS ARE AS FOLLOWS: ALL CAUSE DEATH IN 39 PATIENTS. STROKE IN 15 PATIENTS. SYSTEMIC EMBOLISM 5. PERICARDIAL EFFUSION IN 9 PATIENT, WHICH INCLUDED PERICARDIAL TAMPONADE IN 4 PATIENTS BLEEDING EVENTS IN 19 PATIENTS. ACUTE RENAL FAILURE IN 29 PATIENTS. DEVICE EMBOLISM IN 4 PATIENTS. LAAO REMOVAL IN 4 PATIENTS. DEVICE RELATED THROMBUS IN 4 PATIENTS. MYOCARDIAL INFARCTION IN 7 PATIENT.S TRANSIENT ISCHEMIC ATTACK IN 4 PATIENTS. VENOUS THROMBOEMBOLISM, IN 2 PATIENTS. THE STUDY CONCLUDES THAT A NON-COMPLEX LAAO APPROACH, INCLUDING THE USE OF WATCHMAN DEVICES, IS SAFE AND TECHNICALLY FEASIBLE FOR OLDEST-OLD PATIENTS. THE NON-COMPLEX LAAO PROCEDURE DID NOT SHOW INCREASED RATES OF UNSTABLE DEVICE ANCHORING. IT ALLOWS IMPLANTATION TO BE PERFORMED SAFELY AND IN A RESOURCE-SAVING MANNER. OLDEST-OLD PATIENTS MIGHT BENEFIT FROM THIS LAAO APPROACH, AS IT HELPS TO AVOID STRESSFUL PREPROCEDURAL EXAMINATIONS OR LONG HOSPITAL STAYS. HOWEVER, A LOWER COMBINED EFFECTIVENESS RATE WAS FOUND IN THE LONG-TERM FOLLOW-UP AFTER HOSPITAL DISCHARGE IN THE OLDEST-OLD PATIENTS, DRIVEN BY EXCESS MORTALITY. THESE PATIENTS SHOULD THEREFORE BE SCREENED FOR ROBUSTNESS AND LIFE EXPECTANCY USING SIMPLE GERIATRIC ASSESSMENT TOOLS PRIOR TO THE PROCEDURE. Manufacturer Narrative: B2 DATE OF DEATH: UNKNOWN AS NOT PROVIDED IN LITERATURE ARTICLE. B3 DATE OF EVENT: ESTIMATED AS 04/28/2025 AS THE ONLINE PUBLISHED DATE FOR THE ARTICLE IS NOTED AS APRIL 28, 2025. D4 UNIQUE IDENTIFIER (UDI) #: DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BASED ON THE NATURE OF THE INFORMATION PROVIDED TO BSC, IT IS NOT POSSIBLE TO PERFORM A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER PRODUCT SPECIFIC INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. LITERATURE CITATION: LEFT ATRIAL APPENDAGE OCCLUSION IN VULNERABLE OLDEST-OLD PATIENTS-EFFECTIVENESS AND SAFETY IN A LARGE SINGLE-CENTER REGISTRY FASTNER, CHRISTIAN ET AL. JOURNAL OF THE NEUROLOGICAL SCIENCES, VOLUME 473, 123521.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22784912&pc=FBN | 9610595-2025-17491 | 2025-06-01 04:00 | Death | AIZU OLYMPUS CO., LTD. | 2025-08-12 04:00 | FBN | OES CHOLEDOCHOFIBERSCOPE | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K912120 |  | 1 | Event Description: OLYMPUS REVIEWED THE FOLLOWING LITERATURE TITLED "TRANSCYSTIC LAPAROSCOPIC COMMON BILE DUCT EXPLORATION: WHEN TO BAIL." BACKGROUND: TRANSCYSTIC LAPAROSCOPIC COMMON BILE DUCT EXPLORATION (LCBDE) IS A PROCEDURE CONSIDERED IN THE MANAGEMENT OF COMMON BILE DUCT STONES. IN MANY WAYS IT IS SUPERIOR TO ALTERNATIVES SUCH AS ENDOSCOPIC RETROGRADE CHOLANGIOPANCREATOGRAPHY (ERCP); HOWEVER, SURGEONS WHO HAVE LIMITED EXPERIENCE IN CBDE ARE OFTEN RELUCTANT TO PERSIST IN DIFFICULT CASES WITH CONCERNS REGARDING INCREASING COMPLICATION RATES AND WASTE OF THEATER TIME. THIS STUDY AIMS TO PROVIDE AN EVIDENCE-BASED APPROACH TO IDENTIFY POINTS TO AID EARLY ABANDONMENT (¿BAIL¿). METHODS: REVIEW OF ALL LCBDE PERFORMED IN A SINGLE CENTER FROM SEPTEMBER 2008 TO SEPTEMBER 2022 WAS PERFORMED. STATISTICAL ANALYSIS WAS PERFORMED ON SUCCESS AND FAILURE GROUPS, WITH RELEVANT UNDESIRABLE OUTCOMES CHOSEN FOR FURTHER ANALYSIS TO IDENTIFY FACTORS TO BE USED AS A GUIDE TO BAIL. RESULTS: A TOTAL OF 952 PATIENTS WERE IDENTIFIED FOR ANALYSIS. FEMALES REPRESENTED 63.8% (609) OF THE COHORT. SUCCESS WAS REPORTED IN 89.2% (849) OF PROCEDURES. THOSE IN WHOM THE CYSTIC DUCT COULD NOT BE CANNULATED WITH THE CHOLEDOCHOSCOPE, THOSE THAT PROGRESSED TO CHOLEDOCHOTOMY, THOSE WITH A PROLONGED OPERATIVE TIME, AND THOSE WHO HAD ADVERSE OUTCOMES WERE SELECTED AS UNDESIRED OUTCOMES. FACTORS OF AGE, HIGHER ASA, PREOPERATIVE ERCP, AND THOSE WITH PREOPERATIVELY IDENTIFIED STONES OR LARGER STONES AT OPERATION WERE ASSOCIATED WITH HIGHER RATES OF AN UNDESIRED OUTCOME. CONCLUSION: OLDER AND MORE COMORBID PATIENTS, THOSE WHO UNDERWENT PREOPERATIVE ERCP, AND THOSE WITH PREOPERATIVELY OR OPERATIVELY IDENTIFIED LARGE STONES ARE FACTORS THAT SHOULD PROMPT THOSE SURGEONS WHO ARE DEVELOPING THEIR LCBDE EXPERIENCE TO CONSIDER BAILING EARLY. TYPE OF ADVERSE EVENTS/NUMBER OF PATIENTS. MORTALITY (4 PATIENTS). MORTALITY WAS NOTED ACROSS BOTH SUCCESS AND FAILURE GROUPS. A TOTAL OF 4 DEATHS (3 EMERGENCY, 1 ELECTIVE) WERE RECORDED, A RATE OF 0.4%. THE ELECTIVE CASE DOCUMENTED CAUSE OF DEATH WAS A POSTOPERATIVE CARDIAC ARRHYTHMIA, WHICH LED TO CARDIAC ARREST. NO FURTHER DETAILS REGARDING THE PATIENT DEATH WERE PROVIDED WITHIN THE LITERATURE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION IDENTIFIED IN THE ARTICLE.Event Description: NEW INFORMATION RECEIVED FROM THE AUTHOR STATED THAT ALL MORTALITIES WERE NOT RELATED TO THE ACTUAL OLYMPUS SCOPE AND NOT RELATED TO DEVICE MALFUNCTION. Manufacturer Narrative: THIS REPORT IS RELATED TO THE FOLLOWING LINKED PATIENT IDENTIFIER: (B)(6). THE INVESTIGATION IS ONGOING. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETED OR IF ADDITIONAL INFORMATION BECOMES AVAILABLE. Manufacturer Narrative: THIS SUPPLEMENTAL REPORT IS BEING SUBMITTED TO PROVIDE A CORRECTION, ADDITIONAL INFORMATION FROM THE AUTHOR AND THE RESULTS OF THE COMPLAINT INVESTIGATION. UPDATED FIELDS: B5, H2 AND H6. CORRECTIONS TO B1, B2 AND H1. THIS EVENT WAS INITIALLY CONSERVATIVELY REPORTED AS DEATH; HOWEVER, UPON FURTHER FOLLOW-UP, THE AUTHOR CONFIRMED THAT MORTALITIES WERE NOT RELATED TO THE ACTUAL OLYMPUS SCOPE AND NOT RELATED TO DEVICE MALFUNCTION. THUS CORRECTING B1, B2 AND H1: B1 AND B2 SHOULD NOT BE MARKED AT ALL. H1 SHOULD ALSO NOT BE MARKED. THE DEVICE WAS NOT RETURNED TO OLYMPUS FOR INSPECTION. BASED ON THE RESULTS OF THE INVESTIGATION AND THE AUTHOR'S CONFIRMATION, IT WAS DETERMINED THAT THE OLYMPUS DEVICE DID NOT CAUSE OR CONTRIBUTE TO THE MORTALITIES REPORTED IN THE LITERATURE. ADDITIONALLY, THERE WAS NO DOCUMENTED DEVICE MALFUNCTION IN THE LITERATURE. OLYMPUS WILL CONTINUE TO MONITOR FIELD PERFORMANCE FOR THIS DEVICE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22784931&pc=NGV | 2124215-2025-55862 | 2025-04-28 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-08-12 04:00 | NGV | WATCHMAN LAA CLOSURE DEVICE & DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: REPORTED VIA LITERATURE ARTICLE. IT WAS REPORTED PATIENT DEATH(S) OCCURRED. THIS RETROSPECTIVE STUDY INVESTIGATES THE SAFETY AND EFFECTIVENESS OF A SIMPLIFIED, NON-COMPLEX LEFT ATRIAL APPENDAGE OCCLUSION (LAAO) PROCEDURE IN ELDERLY PATIENTS AGED 80 YEARS AND OLDER, REFERRED TO AS THE "OLDEST-OLD," COMPARED TO YOUNGER PATIENTS. CONDUCTED AT A SINGLE CENTER, THE RETROSPECTIVE ANALYSIS INCLUDED 230 PATIENTS WITH NON-VALVULAR ATRIAL FIBRILLATION AND CONTRAINDICATIONS TO LONG-TERM ORAL ANTICOAGULATION. OF THESE, 43% WERE OLDEST-OLD. THE PROCEDURE UTILIZED VARIOUS OCCLUSION DEVICES, INCLUDING WATCHMAN? AND WATCHMAN FLX?, WITH A HIGH OVERALL TECHNICAL SUCCESS RATE OF 97.4%, SHOWING NO SIGNIFICANT DIFFERENCE BETWEEN AGE GROUPS. PROCEDURAL METRICS SUCH AS DURATION, FLUOROSCOPY TIME, AND RADIATION EXPOSURE WERE SIMILAR ACROSS GROUPS, AND COMPLICATIONS LIKE DEVICE EMBOLIZATION AND PERICARDIAL EFFUSION WERE INFREQUENT. OF THE 230 PATIENTS THAT UNDERWENT THE LAAO PROCEDURE, INTRAPROCEDURAL EVENTS OCCURRED AS FOLLOWS: INTRAPROCEDURAL DEVICE EMBOLIZATION, OCCURRED IN 3 PATIENTS. INTRAPROCEDURAL PERICARDIAL EFFUSION IN 11 PATIENTS. THE MEDIAN HOSPITAL STAY WAS TWO DAYS FOR BOTH GROUPS. COMPLICATIONS REQUIRING IMMEDIATE MEDICAL INTERVENTION OCCURRED IN 8.1 VS. 3.8% OF PATIENTS. INTRAHOSPITAL MORTALITY (1.7 % OF CASES) WAS NOT STATISTICALLY SIGNIFICANTLY DIFFERENT BETWEEN THE TWO GROUPS. NO STROKES OR SYSTEMIC EMBOLISMS OCCURRED AFTER LAAO. MAJOR BLEEDING OCCURRED IN 3.0 VS. 3.1% OF PATIENTS UNTIL HOSPITAL DISCHARGE. THE VAST MAJORITY OF PATIENTS RECEIVED A COMBINATION OF ASA AND CLOPIDOGREL AFTER LAAO. OF THE 230 PATIENTS THAT UNDERWENT THE LAAO PROCEDURE, INTRA HOSPITAL EVENTS OCCURRED AS FOLLOWS: ALL-CAUSE DEATH OCCURRED IN 4 PATIENTS. PERICARDIAL EFFUSION IN 8 PATIENTS, WHICH INCLUDED PERICARDIAL TAMPONADE IN 3 PATIENTS. ACUTE HEART FAILURE IN 1 PATIENT. DEVICE EMBOLIZATION IN 3 PATIENTS. VENOUS THROMBOEMBOLISM IN 1 PATIENT. LAAO REMOVAL IN 3 PATIENTS. ACCESS SITE COMPLICATIONS: HEMATOMA IN 23 PATIENTS. PROLONGED BLEEDING IN 16 PATIENTS. ANEURYSM SPURIUM IN 2 PATIENTS. OF THE 230 PATIENTS THAT UNDERWENT THE LAAO PROCEDURE, THE INTRA HOSPITAL FOLLOW-UP EVENTS ARE AS FOLLOWS: ALL CAUSE DEATH IN 39 PATIENTS. STROKE IN 15 PATIENTS. SYSTEMIC EMBOLISM 5. PERICARDIAL EFFUSION IN 9 PATIENT, WHICH INCLUDED PERICARDIAL TAMPONADE IN 4 PATIENTS. BLEEDING EVENTS IN 19 PATIENTS. ACUTE RENAL FAILURE IN 29 PATIENTS. DEVICE EMBOLISM IN 4 PATIENTS. LAAO REMOVAL IN 4 PATIENTS. DEVICE RELATED THROMBUS IN 4 PATIENTS. MYOCARDIAL INFARCTION IN 7 PATIENTS. TRANSIENT ISCHEMIC ATTACK IN 4 PATIENTS. VENOUS THROMBOEMBOLISM, IN 2 PATIENTS. THE STUDY CONCLUDES THAT A NON-COMPLEX LAAO APPROACH, INCLUDING THE USE OF WATCHMAN DEVICES, IS SAFE AND TECHNICALLY FEASIBLE FOR OLDEST-OLD PATIENTS. THE NON-COMPLEX LAAO PROCEDURE DID NOT SHOW INCREASED RATES OF UNSTABLE DEVICE ANCHORING. IT ALLOWS IMPLANTATION TO BE PERFORMED SAFELY AND IN A RESOURCE-SAVING MANNER. OLDEST-OLD PATIENTS MIGHT BENEFIT FROM THIS LAAO APPROACH, AS IT HELPS TO AVOID STRESSFUL PREPROCEDURAL EXAMINATIONS OR LONG HOSPITAL STAYS. HOWEVER, A LOWER COMBINED EFFECTIVENESS RATE WAS FOUND IN THE LONG-TERM FOLLOW-UP AFTER HOSPITAL DISCHARGE IN THE OLDEST-OLD PATIENTS, DRIVEN BY EXCESS MORTALITY. THESE PATIENTS SHOULD THEREFORE BE SCREENED FOR ROBUSTNESS AND LIFE EXPECTANCY USING SIMPLE GERIATRIC ASSESSMENT TOOLS PRIOR TO THE PROCEDURE. Manufacturer Narrative: B2 DATE OF DEATH: UNKNOWN AS NOT PROVIDED IN LITERATURE ARTICLE. B3 DATE OF EVENT: ESTIMATED AS 04/28/2025 AS THE ONLINE PUBLISHED DATE FOR THE ARTICLE IS NOTED AS APRIL 28, 2025. D4 UNIQUE IDENTIFIER (UDI) #: DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BASED ON THE NATURE OF THE INFORMATION PROVIDED TO BSC, IT IS NOT POSSIBLE TO PERFORM A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER PRODUCT SPECIFIC INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. LITERATURE CITATION: LEFT ATRIAL APPENDAGE OCCLUSION IN VULNERABLE OLDEST-OLD PATIENTS-EFFECTIVENESS AND SAFETY IN A LARGE SINGLE-CENTER REGISTRY FASTNER, CHRISTIAN ET AL. JOURNAL OF THE NEUROLOGICAL SCIENCES, VOLUME 473, 123521.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22786082&pc=DSQ | 2916596-2025-05370 | 2025-06-17 04:00 | Death | THORATEC CORPORATION | 2025-08-12 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT'S FAMILY CALLED THE SITE AT APPROXIMATELY 2:15 AM ON (B)(6) 2025 AND BELIEVED THAT THE PATIENT HAD PASSED AWAY. THE PATIENT WAS NOT BREATHING OR MOVING AND THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) WAS ALARMING FOR LOW FLOWS. IT WAS CONFIRMED THAT THE FLOW AT THE TIME WAS 0.7 LITERS PER MINUTE (LPM). AT 3 AM, THE HOSPICE NURSE CALLED AND CONFIRMED THAT THE PATIENT HAD PASSED AWAY AS THERE WAS NO SPONTANEOUS BREATHING AND PUPILS WERE FIXED AND DILATED. THE PUMP WAS NOT EXPLANTED. NO AUTOPSY WOULD BE PERFORMED. IT WAS ADDITIONALLY STATED THAT THE DEVICE FLOW, SPEED AND POWER HAD BEEN WITHIN NORMAL LIMITS FOR THE PATIENT. Manufacturer Narrative: SECTION A: PATIENT WEIGHT NOT YET AVAILABLE. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22786227&pc=DSQ | 2916596-2025-05341 | 2024-08-15 04:00 | Death | THORATEC CORPORATION | 2025-08-12 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Bacterial Infection; Chest Pain; Purulent Discharge; Dyspnea; Fall | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD CHRONIC DRIVELINE INFECTIONS THAT LED TO THE PATIENT EVENTUALLY PASSING AWAY AT HOME. THE ONSET OF THE PATIENT'S INFECTION WAS NOT KNOWN AS THEY HAD MULTIPLE INFECTIONS WITH MULTIPLE ORGANISMS. THEIR LAST DRIVELINE CULTURE ON (B)(6) 2024 WAS POSITIVE FOR SERRATIA MARCESCENS AND CORYNEBACTERIUM. THE PATIENT COMPLAINED OF FEELING UNWELL WITH SHORTNESS OF BREATH, INCREASED DRIVELINE DRAINAGE, AND LEFT SIDED CHEST PAIN PER DOCUMENTATION FROM (B)(6) 2024. THEY WERE ASKED TO PRESENT TO THE EMERGENCY ROOM WITH CONCERN FOR SEPSIS VS ABSCESS, BUT THEY DID NOT. THEY REPORTED FEELING BETTER ON (B)(6) 2024. THERE WAS DOCUMENTATION OF INCREASING VENTRICULAR TACHYCARDIA BURDEN IN THE DAYS LEADING UP TO THEIR DEATH, AND THE PATIENT REPORTED A FALL. THEY WANTED TO COME TO THE (B)(6) HOSPITAL AT THAT TIME BUT DID NOT HAVE A RIDE. THE PATIENT HAD SYMPTOMS OF SHORTNESS OF BREATH AND CHEST PAIN DUE TO THE ARRHYTHMIA. IT WAS UNKNOWN IF THE VENTRICULAR TACHYCARDIA WAS SUSTAINED. IT WAS UNKNOWN OF THE PATIENT HAD A HISTORY OF ARRHYTHMIA, HOWEVER AN ICD WAS IMPLANTED PRIOR TO VAD. IT WAS UNKNOWN IF THE ARRHYTHMIA WAS DEVICE OR THERAPY RELATED. IT WAS UNKNOWN IF THE PATIENT'S FALL WAS DEVICE OR THERAPY RELATED. THE OUTCOME WAS NOTED TO BE DEVICE OR THERAPY RELATED. NO AUTOPSY WAS PERFORMED AND THE DEVICE WAS NOT EXPLANTED. NO ALARMS WERE REPORTED. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A SPECIFIC CAUSE FOR THE PATIENT¿S DRIVELINE INFECTION COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. FURTHER, A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REV. D AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK, REV. A ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING INFECTION (LOCAL, DRIVELINE, AND PUMP POCKET), SEPSIS, AND ARRHYTHMIA, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, LISTS INFECTION AND ARRHYTHMIA AS POTENTIAL LATE POSTIMPLANT COMPLICATIONS. FURTHERMORE, SEVERAL SECTIONS OF THE HM3 IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: OTHER APPLICABLE ANNEX E CODES: INFECTIONS E03: SEPSIS E0306. INFECTIONS E17: ABSCESS E1720. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22786265&pc=DSQ | 3007042319-2025-01907 | 2025-06-25 04:00 | Death | HEARTWARE, INC. | 2025-08-12 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Tachycardia; Ventricular Fibrillation | P100047 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING A META-ANALYSIS REVIEW OF VENTRICULAR ARRHYTHMIAS FOLLOWING LEFT VENTRICULAR ASSIST DEVICE (VAD) IMPLANTATION. THE AUTHORS DESCRIBED PATIENT DEATHS; HOWEVER, THE SPECIFIC CAUSES OF DEATH WERE UNKNOWN AND DESCRIBED AS ALL-CAUSE MORTALITY. THERE WERE PATIENTS WHO EXPERIENCED EARLY AND LATE ONSET VENTRICULAR ARRHYTHMIAS WHICH INCLUDED ELECTRICAL STORM, AND PATIENTS WITH PUMP THROMBOSIS. THE STATUS OF THE DEVICES IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/54 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. REFERENCED ARTICLE: VENTRICULAR ARRHYTHMIAS FOLLOWING LEFT VENTRICULAR ASSIST DEVICE IMPLANTATION: A SYSTEMATIC REVIEW AND META-ANALYSIS OF INCIDENCE AND CLINICAL IMPACT. EUROPACE. 2025. 27, EUAF129. DOI: 10.1093/EUROPACE/EUAF129 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22786576&pc=DSQ | 3004582654-2025-00045 | 2025-04-05 04:00 | Death | BERLIN HEART GMBH | 2025-08-12 04:00 | DSQ | EXCOR® | Adverse Event Without Identified Device or Use Problem | Hemorrhagic Stroke | P160035 |  | 1 | Event Description: BERLIN HEART GMBH CLINICAL AFFAIRS WAS INFORMED BY THE CLINIC ON 2025-07-17 THAT A PATIENT SUFFERED AN HEMORRHAGIC STROKE AFTER RECEIVING FIBRINOLYTICS FOR A CLOT IN ONE OF THE CANNULAE. AS A RESULT, THE PATIENT WAS DECLARED BRAIN DEAD. CARE WAS WITHDRAWN FROM THE PATIENT ON (B)(6) 2025. THE PATIENT WAS BEING SUPPORTED WITH THE EXCOR SYSTEM IN BI-VAD CONFIGURATION. DEPOSITS WERE NOTED IN THE EXCOR CANNULAE. Manufacturer Narrative: THE PATIENT WAS IMPLANTED WITH A TOTAL OF FOUR CANNULAE ON (B)(6) 2025 AS FOLLOWS: THE LVAD EXCOR APEX CANNULA FOR CHILDREN Ø 12/9 MM (C22A-004); L 27 CM; LOT NO. 00145787. THE LVAD ARTERIAL CANNULA FOR SMALL CHILDREN Ø 6 MM; L 25 CM (C80G-021M); LOT NO. 00125255. THE RVAD ATRIAL CANNULA Ø 12/9 MM; L 32 CM; HEAD 23 MM (C23V-004M); LOT NO. 00145788. THE RVAD ARTERIAL CANNULA FOR SMALL CHILDREN Ø 6 MM; L 25 CM (C80G-021M); LOT NO. 00148388. THE CANNULAE WERE IN USE ON THE PATIENT FROM ON (B)(6) 2025 UNTIL THEY WERE REMOVED FROM THE PATIENT ON (B)(6) 2025. THE EVENT OCCURRED ON 2025-04-05, WHICH IS (87 DAYS) AFTER THE CANNULAE WERE PLACED ON THE PATIENT. THE DETAILS OF THE LVAD APEX CANNULA LOT NO. 00145787 ARE PROVIDED IN SECTION D4. THE INFORMATION OF THE OTHER THREE CANNULAE USED ON THE PATIENT ARE AS FOLLOWS: THE CANNULA C23V-004M WITH LOT NO. 00145788: THE MANUFACTURE DATE FOR THIS LOT: 2024-08-29, THE EXPIRATION DATE FOR THIS LOT: 2027-07-25, UDI: (B)(4). THE CANNULA C80G-021M WITH LOT NO. 00125255: THE MANUFACTURE DATE FOR THIS LOT: 2022-11-03, THE EXPIRATION DATE FOR THIS LOT: 2025-09-30, UDI: (B)(4). THE CANNULA C80G-021M WITH LOT NO. 00148388: THE MANUFACTURE DATE FOR THIS LOT: 2024-10-25, THE EXPIRATION DATE FOR THIS LOT: 2027-10-02, UDI: (B)(4). WE HAVE REVIEWED THE PRODUCTION RECORDS OF THE EXCOR CANNULAE AND THE CANNULAE WERE PRODUCED ACCORDING TO OUR SPECIFICATIONS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22786636&pc=OZD | 1220648-2025-31157 | 2025-07-23 04:00 | Death | ABIOMED, INC. | 2025-08-12 04:00 | OZD | IMPELLA 5.5 | Increase in Pressure | Thrombosis/Thrombus | P140003 |  | 1 | Event Description: THE COMPLAINANT PLACED A IMPELLA 5.5 TO SUPPORT THE PATIENT WITH CARDIOGENIC SHOCK AND KNOWN ARRHYTHMIA. THE PUMP WAS SUPPORTING WHEN AFTER 3 DAYS THERE WAS THROMBUS AND PURGE PRESSURE/PURGE FLOW ISSUES. THE TEAM CONSULTED WITH ABIOMED MEDICAL TEAM AND THE PURGE HAD TISSUE PLASMINOGEN ACTIVATOR WAS ADDED. PURGE PRESSURE AND PURGE FLOW IN NORMAL RANGES WITHIN A FEW HOURS OF THE INFUSION BEING INITIATED. THE PUMP REMAINED ON FOR SUPPORT FOR JUST UNDER 9 DAYS. THERE WAS NO SUCCESS WITH ABLATION AND THERAPY WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION. SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER. SECTION: DIRECT AORTIC INSERTION. SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES "IN PATIENTS WITH TRANSCATHETER AORTIC VALVES POSITION THE IMPELLA SYSTEM CAREFULLY TO AVOID INTERACTION WITH THE TRANSCATHETER AORTIC VALVE PROSTHESIS. UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH." . IN THIS SITUATION, AVOID REPOSITIONING WHILE THE DEVICE IS RUNNING; TURN THE DEVICE TO P0 DURING REPOSITIONING OR ANY MOVEMENT THAT COULD BRING THE OUTLET WINDOWS INTO PROXIMITY TO THE VALVE STENT STRUCTURES. IF THERE IS LOW FLOW OBSERVED IN A PATIENT IMPLANTED WITH A TRANSCATHETER AORTIC VALVE PROSTHESIS, CONSIDER DAMAGE OF THE IMPELLER AND REPLACE THE IMPELLA AS SOON AS POSSIBLE.¿ SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA CATHETER ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP. CLINICIANS SHOULD MONITOR MOTOR CURRENT AND CONSIDER REPLACING THE IMPELLA CATHETER WHENEVER A RISE IN MOTOR CURRENT IS SEEN.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22786823&pc=DSQ | 2916596-2025-05391 | 2025-08-01 04:00 | Death | THORATEC CORPORATION | 2025-08-12 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE CUSTOMER COMMUNICATED THAT NO ADDITIONAL INFORMATION WILL BE PROVIDED DUE TO PATIENT PRIVACY LAWS. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. B, IS CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22786869&pc=DSQ | 2916596-2025-05351 | 2025-06-21 04:00 | Death | THORATEC CORPORATION | 2025-08-12 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Aneurysm; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MYCOTIC ANEURYSM, PSEUDOMONAS BACTEREMIA, RIGHT VENTRICULAR FAILURE, INPATIENT HOSPICE. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22787096&pc=MHX | 9610816-2025-000732 | 2025-08-02 04:00 | Death | PHILIPS MEDIZIN SYSTEME BÖBLINGEN GMBH | 2025-08-12 04:00 | MHX | INTELLIVUE MX450 PATIENT MONITOR | Defective Alarm | No Clinical Signs, Symptoms or Conditions | K130849 |  | 1 | Event Description: IT WAS REPORTED THE DEVICE DID NOT GENERATE AND ALARM AND THE PATIENT PASSED AWAY. FOLLOWING THE PATIENT'S DEATH, THE CUSTOMER REQUESTED ASSISTANCE PULLING LOGS AND WAVEFORMS TO CONFIRM THE DEVICE WAS WORKING AS EXPECTED. THE PATIENT WAS MONITORED FOR SPO2 AND NIBP AT A MINIMUM FROM WHAT THE USER REMEMBERS. IT WAS INDICATED THE EMERGENCY DEPARTMENT DOES NOT ADMIT THE PATIENTS TO THE CENTRAL STATION, SO INFORMATION AVAILABLE WAS LIMITED; HOWEVER, THE CUSTOMER WAS ABLE TO OBTAIN THE LOGS AND REPORTS, WHICH WAS SUFFICIENT. ADDITIONAL INFORMATION WAS NOT PROVIDED; THEREFORE, GOOD FAITH EFFORTS ARE BEING PERFORMED. Manufacturer Narrative: A PHILIPS TECHNICAL CONSULTANT (TC) WENT TO THE CUSTOMER SITE. WHILE THE TC WAS ONSITE WITH THE CUSTOMER BIOMEDICAL ENGINEER (BIOMED) AND EMERGENCY DEPARTMENT (ED) PHYSICIAN, THEY WERE ABLE TO CLARIFY THAT THERE WAS NO ALLEGED FAILURE OF THE DEVICE; WHAT THEY WANTED WAS TO REVIEW THE WAVE STRIP DURING FOR (B)(6) 2025. IN THE ED DEPARTMENT THEY DO NOT ADMIT PATIENT'S TO THE PATIENT INFORMATION CENTER IX (PIC IX) SO THAT INFORMATION WAS NOT AVAILABLE. THE TC ASKED IF FURTHER INQUIRY WAS NECESSARY AND THEY DECLINED. BASED ON THE INFORMATION AVAILABLE, THERE WAS NO ALLEGED FAILURE. THE REPORTED PROBLEM WAS NOT CONFIRMED. THE DEVICE WAS CONFIRMED TO BE OPERATING PER SPECIFICATIONS; THE CUSTOMER ONLY NEEDED ASSISTANCE WITH RETROSPECTIVE DATA. THE INVESTIGATION CONCLUDES THAT NO FURTHER ACTION IS REQUIRED AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED THE COMPLAINT FILE WILL BE REOPENED. Manufacturer Narrative: PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22787379&pc=KRH | 9612164-2025-03946 | 2025-06-04 04:00 | Death | MEDTRONIC MEXICO | 2025-08-12 04:00 | KRH | RING 620RG33 DURAN ANCORE ANNULOPLASTY C | Insufficient Information; Patient Device Interaction Problem | Bacterial Infection; Cardiac Arrest; Embolism/Embolus | K032810 |  | 1 | Event Description: MEDTRONIC RECEIVED INFORMATION THAT AFTER AN UNKNOWN DURATION POST IMPLANT OF THIS 33MM MITRAL ANNULOPLASTY RING, IT WAS REPORTED THAT THE VALVE WAS INFECTED, AND AN INTERVENTION TO REPLACE THE INFECTED VALVE AND VEGETATION FROM WAS SCHEDULED. THE MORNING OF THE PROCEDURE, THE VEGETATION ON THE VALVE EMBOLIZED CAUSING A FATAL HEART ATTACK OR CARDIAC ARREST. RESUSCITATION WAS PERFORMED AND THE PATIENT WAS PLACED ON EXTRA CORPOREAL MEMBRANE OXYGENATION (ECMO). THE PATIENT WAS THEN TAKEN OFF OF LIFE SUPPORT AND PASSED AWAY. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22787457&pc=DSQ | 2916596-2025-05329 | 2025-05-19 04:00 | Death | THORATEC CORPORATION | 2025-08-12 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Multiple Organ Failure; Heart Failure/Congestive Heart Failure; No Clinical Signs, Symptoms or Conditions | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO RIGHT VENTRICULAR FAILURE, MULTI SYSTEM ORGAN FAILURE, AND NON-EXTUBATED. THIS WAS CONSIDERED DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED AND WOULD NOT RETURN FOR EVALUATION.Event Description: THE PATIENT DID NOT HAVE RIGHT HEART FAILURE PRE-LVAD. THE LVAD/LVAD THERAPY CONTRIBUTED TO THE PATIENT'S DECLINE IN RIGHT VENTRICULAR FUNCTION. THE MULTI SYSTEM ORGAN FAILURE WAS NOT DUE TO PROGRESSION OF THE PATIENT'S HEART FAILURE. IT WAS REPORTED THAT THE PHYSICIAN HAD ISSUES CLOSING THE PURPLE LOCKING MECHANISM WHILE USING A MINI APICAL CUFF. IT WAS NOTED THERE WERE MANY PLEDGETS USED TO STOP BLEEDING AROUND THE SEWING RING AND THAT THERE WAS PUCKERING OF TISSUE THAT MIGHT HAVE INTERFERED WITH THE LOCKING MECHANISM. THE SURGEON REPLACED THE MINI APICAL CUFF WITH A REGULAR SEWING RING WHEN THEY WERE UNABLE TO ATTACH THE MINI APICAL CUFF. THE SEWING RING WAS ABLE TO BE ATTACHED, HOWEVER THERE WAS AN INCOMPLETE SEAL RESULTING IN THE SITE REQUESTING A PUMP EXCHANGE (2916596-2025-02801). THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: THIS EVENT IS SUPPLEMENTAL AND INVESTIGATION FINDINGS WILL BE SUBMITTED UNDER MFR # 2916596-2025-02801. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22787557&pc=DSQ | 2916596-2025-05354 | 2024-12-19 05:00 | Death | THORATEC CORPORATION | 2025-08-12 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Respiratory Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE RESPIRATORY FAILURE. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22787753&pc=OZD | 1220648-2025-45738 | 2025-07-25 04:00 | Death | ABIOMED, INC. | 2025-08-12 04:00 | OZD | IMPELLA CP | Malposition of Device | Hemolysis; Hemorrhage/Bleeding; Ischemia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOLLOWED BY AN IMPELLA RP FLEX DEVICE FOR BIPELLA MECHANICAL CIRCULATORY SUPPORT. ONE DAY AFTER START OF THE BIPELLA SUPPORT, THE PATIENT LOST PULSES IN THE FOOT. UPGRADE TO AN IMPELLA 5.5 FROM THE IMPELLA CP DEVICE WAS BEING CONSIDERED. THE FOLLOWING DAY HEMOLYSIS WAS NOTED. THE PHYSICIAN DECIDED TO REMOVE THE RP FLEX DEVICE AT THE BEDSIDE. AN ECHOCARDIOGRAM COMPLETED AND SHOWED IMPELLA CP A APPROXIMATELY 5CM AND WAS, THEREFORE, REPOSITIONED TO MEASURE 3.4 CM. THE NEXT DAY THEREAFTER, IT WAS NOTED THE PATIENT HAD A GASTROINTESTINAL BLEED; BLOOD PRODUCTS WERE GIVEN. THE BLEED WORSENED THE FOLLOWING DAY AND HEPARIN DRIP WAS SUSPENDED. THIS DAY IT WAS ALSO NOTED POOR TO ABSENT PULSES HAD PERSISTED OVER THE WEEKEND. THE IMPELLA SIDE LEG LOOKED ISCHEMIC, AND THE PATIENT REQUIRED AMPUTATION IN RESPONSE TO THE ISCHEMIA. THE PATIENT'S FAMILY DECIDED TO PROVIDE COMFORT CARE ONLY UNTIL THE PATIENT EXPIRED. Manufacturer Narrative: BASED ON THE DETAILS OF THE EVENTS DURING SUPPORT, THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA CP DEVICE AS AN ASSOCIATED FACTOR TO THE PATIENT'S DEMISE OUTCOME. THE IMPELLA RP FLEX WAS EXPLANTED, PATIENT REMAINED ON IMPELLA CP SUPPORT AND EXPERIENCED FURTHER COMPLICATIONS ONE LEADING TO AMPUTATION. PRODUCT STATUS: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES)-¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY¿ (INCLUDING VENTRICULAR PERFORATION). SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE. ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS - ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS.¿ ¿PATIENT CONDITIONS - INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES - MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION - ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22787756&pc=NIK | 2017865-2025-97426 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-08-12 04:00 | NIK | GALLANT HF | Adverse Event Without Identified Device or Use Problem | Unspecified Infection | P030054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAS DECEASED ON (B)(6) 2025. THE CAUSE OF DEATH WAS CARDIOGENIC SHOCK AND SEPTIC SHOCK. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS RELATED TO THEIR IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) SYSTEM. Manufacturer Narrative: THE STERILIZATION RECORDS WERE REVIEWED AND NO EVIDENCE OF ABNORMAL STERILIZATION CYCLE WAS FOUND. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INFECTION COULD NOT BE CONCLUSIVELY DETERMINED. Manufacturer Narrative: THE LEAD WAS RETURNED DUE TO PATIENT DECEASED. AS RECEIVED, ONLY THE CONNECTOR PORTION OF THE LEAD WAS RETURNED IN ONE PIECE. ELECTRICAL TESTING DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS. VISUAL INSPECTION OF THE LEAD DID NOT FIND ANY ANOMALIES EXCEPT FOR THE CUT END WHICH IS CONSISTENT WITH PROCEDURAL DAMAGE. Manufacturer Narrative: CORRECTION- PLEASE DISREGARD PREVIOUS ADDITIONAL REPORT AS EVALUATION WAS FOR THE ASSOCIATED LEAD RATHER THAN THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD). PLEASE SEE DEVICE EVALUATION BELOW: INTERROGATION OF THE DEVICE REVEALED IT WAS ABOVE ERI WHEN RECEIVED. BASED ON THIS INFORMATION, THE DEVICE WAS FOUND TO COMMUNICATE APPROPRIATELY WITH A MERLIN PROGRAMMER AND HAS NOT REACHED THE ELECTIVE REPLACEMENT INDICATOR (ERI).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22787761&pc=NIK | 2017865-2025-97429 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-08-12 04:00 | NIK | QUARTET | Adverse Event Without Identified Device or Use Problem | Unspecified Infection | P030054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAS DECEASED ON (B)(6) 2025. THE CAUSE OF DEATH WAS CARDIOGENIC SHOCK AND SEPTIC SHOCK. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS RELATED TO THEIR IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) SYSTEM. Manufacturer Narrative: THE LEAD WAS RETURNED DUE TO PATIENT DECEASED. AS RECEIVED, ONLY THE CONNECTOR PORTION OF THE LEAD WAS RETURNED IN ONE PIECE. ELECTRICAL TESTING DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS. VISUAL INSPECTION OF THE LEAD DID NOT FIND ANY ANOMALIES EXCEPT FOR THE CUT END WHICH IS CONSISTENT WITH PROCEDURAL DAMAGE. THE STERILIZATION RECORDS WERE REVIEWED AND NO EVIDENCE OF ABNORMAL STERILIZATION CYCLE WAS FOUND. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INFECTION COULD NOT BE CONCLUSIVELY DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22787797&pc=NVZ | 2017865-2025-97418 | 2025-07-23 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-08-12 04:00 | NVZ | ELLIPSE VR | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P910023 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT DECEASED FOLLOWING AN UNRELATED PROCEDURE. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. IT WAS REPORTED THAT THE CAUSE OF DEATH WAS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22788393&pc=DXF | 2124215-2025-56193 | 2025-07-23 04:00 | Death | BAYLIS MEDICALE CIE INC | 2025-08-12 04:00 | DXF | VERSACROSS CONNECT LAAC ACCESS SOLUTION | Adverse Event Without Identified Device or Use Problem | Cardiac Tamponade; Pericardial Effusion | K150709 |  | 1 | Event Description: IT WAS REPORTED A PATIENT DEATH OCCURRED. DURING A LEFT ATRIAL APPENDAGE CLOSURE (LAAC) PROCEDURE, A VERSACROSS CONNECT LAAC WAS SELECTED FOR USE. THE PATIENT HAD MARFAN SYNDROME. A SMALL EFFUSION WAS NOTED ON THE RIGHT VENTRICLE PRIOR TO THE START OF THE PROCEDURE. AFTER THE TRANSEPTAL PUNCTURE (TSP), THE PATIENT DEVELOPED LARGER EFFUSION AROUND THE WHOLE PERICARDIUM ALONG WITH CARDIAC TAMPONADE. CARDIOPULMONARY RESUSCITATION (CPR) WAS STARTED. A PERICARDIAL TAP WAS ATTEMPTED TO TREAT THE EFFUSION BUT WAS UNSUCCESSFUL. RESUSCITATION WAS DONE FOR THIRTY (30) MINUTES, HOWEVER THE PATIENT EXPIRED. THE DEVICE IS NOT EXPECTED TO BE RETURNED FOR ANALYSIS (DISPOSED). Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION, BUT IT WAS NOT AVAILABLE. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22788401&pc=QEY | 2124215-2025-55908 | 2025-07-14 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-08-12 04:00 | QEY | EKOS ENDOVASCULAR DEVICE, 106X12CM | Leak/Splash | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT DEATH OCCURRED. A 106CM X 12CM EKOSONIC ENDOVASCULAR DEVICE WAS SELECTED FOR USE IN A BILATERAL CASE; BOTH WINGS OF THE LUNG WERE AFFECTED, SO TWO CATHETERS WERE USED. SHORTLY AFTER STARTING THERAPY IN THE ICU, FLUID AND BLOOD BEGAN LEAKING FROM THE WHITE CONNECTOR OF THIS INFUSION CATHETER. AS A RESULT, THE PATIENT WAS BROUGHT BACK TO THE LAB FOR CATHETER REPLACEMENT. THIS OCCURRED APPROXIMATELY 45 MINUTES AFTER THE PATIENT HAD BEEN TRANSFERRED FROM THE LAB. THERAPY THEN CONTINUED FOR ABOUT TWO HOURS, AFTER WHICH THE PATIENT REQUIRED RESUSCITATION AND THEN PASSED AWAY. Manufacturer Narrative: B2: DATE OF DEATH AND B3: DATE OF EVENT ESTIMATED USING AWARE DATE. E1: INITIAL REPORTER PHONE NUMBER: (B)(6). GOOD FAITH EFFORT ATTEMPTS WERE MADE TO TRY AND RETRIEVE ADDITIONAL DETAILS REGARDING THE REPORTED EVENT AND CAUSE OF DEATH, BUT FURTHER INFORMATION WAS UNABLE TO BE OBTAINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22789573&pc=QLG | 2954323-2025-31853 | 2025-07-14 04:00 | Death | ABBOTT DIABETES CARE INC | 2025-08-12 04:00 | QLG | FREESTYLE LIBRE 2 | Insufficient Information | Insufficient Information | K193371 |  | 1 | Event Description: ABBOTT DIABETES CARE (ADC) RECEIVED MEDWATCH REPORT MW5172077, WHICH INCLUDED THE FOLLOWING INFORMATION: ON (B)(6) 2025, AN UNIDENTIFIED REPORTER STATED THAT A CUSTOMER HAD PASSED AWAY. NO ADDITIONAL DETAILS OR CLARIFICATION WERE PROVIDED. THE REPORTER ALSO REQUESTED DELETION OF A PRIOR REQUEST DUE TO THE PATIENT'S PASSING; HOWEVER, THE NATURE OF THAT REQUEST WAS NOT SPECIFIED. IT IS UNKNOWN WHETHER THE CUSTOMER WAS ACTIVELY USING AN ADC DEVICE AT THE TIME OF DEATH, AND NO ALLEGATION OF DEVICE DEFICIENCY WAS MADE. THE REPORT NOTED THAT THE HEALTHCARE PROFESSIONAL DID NOT CONSENT TO FOLLOW-UP. NO CONTACT INFORMATION WAS PROVIDED TO ADC; THEREFORE, ADC IS UNABLE TO ATTEMPT ANY FOLLOW UP ACTIVITIES RELATED TO THIS EVENT. BASED ON THE INFORMATION PROVIDED, IT CANNOT BE REASONABLY SUGGESTED THAT AN ADC PRODUCT HAS OR MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED DEATH. Manufacturer Narrative: THIS COMPLAINT WAS RECEIVED VIA USER REPORT MW5172077 AND HAS BEEN REPORTED TO FDA. ABBOTT DIABETES CARE (ADC) IS UNABLE TO FURTHER INVESTIGATE THE REPORTED COMPLAINT DUE TO THE ABSENCE OF AN ALLEGED DEVICE DEFICIENCY. ADDITIONALLY, NO CONTACT INFORMATION WAS PROVIDED TO ADC; THEREFORE, ADC IS UNABLE TO ATTEMPT ANY FOLLOW UP ACTIVITIES RELATED TO THIS EVENT. THE DATE OF EVENT IS UNKNOWN. THE DATE ENTERED IN SECTION B3 IS THE DATE ABBOTT DIABETES CARE BECAME AWARE OF THE EVENT. ALL PERTINENT INFORMATION AVAILABLE TO ADC HAS BEEN SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22789938&pc=GZF | 3010676138-2025-00161 | 2025-07-16 04:00 | Death | CURONIX LLC | 2025-08-12 04:00 | GZF | FREEDOM PERIPHERAL NERVE STIMULATOR | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K171366 |  | 1 | Event Description: THE COMMERCIAL TEAM MEMBER WAS INFORMED THE PATIENT PASSED AWAY. HOWEVER, THE EXACT DATE WAS NOT PROVIDED. THERE IS NO ALLEGED CONNECTION OR DEFICIENCY ASSOCIATED WITH THE CURONIX DEVICE. Manufacturer Narrative: THE OTHER ADVERSE EVENTS ISSUES QUESTIONNAIRE WAS REVIEWED. THE PATIENT WAS ELDERLY AND IN HOSPICE. THE CAUSE OF THE PATIENT'S DEATH IS UNRELATED TO THE CURONIX DEVICE AND THERE WERE NO ALLEGED DEFICIENCIES. THE TRANSMITTERS ASSOCIATED WITH THE COMPLAINT WERE NOT SENT TO CURONIX HEADQUARTERS FOR ENGINEERING INVESTIGATION. ATTEMPTS TO RECOVER THE DEVICE WERE UNSUCCESSFUL, PREVENTING ENGINEERING FROM PERFORMING INVESTIGATION. A CURONIX REPRESENTATIVE CONDUCTED A REVIEW OF STERILIZATION AND PACKAGING RECORDS FOR THE RESPECTIVE PRODUCT LOT; CURONIX HAS CONFIRMED THAT THE PRODUCT WAS DELIVERED STERILE, VALIDATED STERILIZATION PARAMETERS WERE USED, AND STERILE BARRIERS WERE VERIFIED TO BE INTACT FOLLOWING PACKAGING. THE STIMULATOR IS USED TO TREAT PAIN. THE CAUSE OF THE PATIENT'S DEATH IS UNRELATED TO THE CURONIX DEVICE. THE PROVIDED INFORMATION DOES NOT EVIDENCE THAT THE CURONIX DEVICE CONTRIBUTED TO THE ISSUE (NO FAULT FOUND). RATES REVIEWED AT THE MOST RECENT COMPLAINT TRENDING MEETING DO NOT INDICATE A SIGNIFICANT INCREASE IN OTHER ADVERSE EVENTS ISSUES. OTHER ADVERSE EVENTS ISSUE RATES REMAIN ACCEPTABLY LOW; THUS, A CAPA IS NOT REQUIRED. OTHER ADVERSE EVENTS ISSUE RATES WILL CONTINUE TO BE TRACKED AND TRENDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22790033&pc=NPT | 2135147-2025-04490 | 2018-11-01 04:00 | Death | ABBOTT MEDICAL | 2025-08-12 04:00 | NPT | NAVITOR TRANSCATHETER AORTIC VALVE | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P190023 |  | 1 | Event Description: THE ARTICLE, "TRANSCATHETER AORTIC VALVE REPLACEMENT WITH INTRA-ANNULAR SELF-EXPANDING OR BALLOON-EXPANDABLE VALVES THE MULTICENTER INTERNATIONAL NAVULTRA REGISTRY", WAS REVIEWED. THE ARTICLE PRESENTED A PROSPECTIVE, MULTICENTER STUDY TO COMPARE THE CLINICAL AND ECHOCARDIOGRAPHIC OUTCOMES AT 30 DAYS AND 1-YEAR OF THE NAVITOR VS THE ULTRA TRANSCATHETER HEART VALVE (THV) IN A PROPENSITY SCORE (PS)¿MATCHED POPULATION. DEVICES INCLUDED IN THE STUDY WERE NAVITOR AND SAPIEN 3 ULTRA. THE ARTICLE CONCLUDED BOTH INTRA-ANNULAR THVS WERE ASSOCIATED WITH SIMILAR 1-YEAR CLINICAL OUTCOMES; HOWEVER, DIFFERENCES WERE OBSERVED IN SECONDARY CLINICAL ENDPOINTS AND VALVE HEMODYNAMIC PERFORMANCE. [PRIMARY AND CORRESPONDING AUTHOR WAS STEFANO CANNATA, IRCCS-ISMETT (MEDITERRANEAN INSTITUTE FOR TRANSPLANTATION AND ADVANCED SPECIALIZED THERAPIES), CARIM (CARDIOVASCULAR RESEARCH INSTITUTE MAASTRICHT), VIA E. TRICOMI 5 90127 PALERMO, ITALY, WITH CORRESPONDING E-MAIL: STEFANOCANN@GMAIL.COM] THE TIME FRAME OF THE STUDY WAS FROM NOVEMBER 2018 TO APRIL 2024. A TOTAL OF 4878 PATIENTS WERE INCLUDED IN THE STUDY, OF WHICH 1746 (35.8%) RECEIVED AN ABBOTT DEVICE. THE AVERAGE AGE WAS 80.3 YEARS, AND THE MAJORITY GENDER WAS FEMALE. COMORBIDITIES INCLUDED HYPERTENSION, DIABETES MELLITUS, CHRONIC OBSTRUCTIVE PULMONARY DISORDER, SEVERE LIVER DISEASE, PORCELAIN AORTA, ATRIAL FIBRILLATION, PRIOR PERCUTANEOUS CORONARY INTERVENTION, PERIPHERAL VASCULAR DISEASE, CORONARY ARTERY DISEASE, PRIOR MYOCARDIAL INFARCTION, PRIOR CORONARY ARTERY BYPASS GRAFT, PRIOR CARDIAC SURGERY, DIALYSIS, SEVERE PULMONARY HYPERTENSION, PRIOR PACEMAKER, HEART BLOCK, BICUSPID AORTIC VALVE. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF TRANSCATHETER AORTIC VALVE REPLACEMENT WITH INTRA-ANNULAR SELF-EXPANDING OR BALLOON-EXPANDABLE VALVES THE MULTICENTER INTERNATIONAL NAVULTRA REGISTRY WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING HYPERTENSION, DIABETES MELLITUS, CHRONIC OBSTRUCTIVE PULMONARY DISORDER, SEVERE LIVER DISEASE, PORCELAIN AORTA, ATRIAL FIBRILLATION, PRIOR PERCUTANEOUS CORONARY INTERVENTION, PERIPHERAL VASCULAR DISEASE, CORONARY ARTERY DISEASE, PRIOR MYOCARDIAL INFARCTION, PRIOR CORONARY ARTERY BYPASS GRAFT, PRIOR CARDIAC SURGERY, DIALYSIS, SEVERE PULMONARY HYPERTENSION, PRIOR PACEMAKER, HEART BLOCK, BICUSPID AORTIC VALVE. SOME OF THE COMPLICATIONS REPORTED WERE HEMORRHAGE, ATRIAL FIBRILLATION, HEART BLOCK, STROKE, RENAL FAILURE, HEART FAILURE, DEATH; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. PLEASE NOTE THAT PER THE INSTRUCTION FOR USE, ¿THE NAVITOR¿ VALVE IS INDICATED FOR PATIENTS WITH SYMPTOMATIC SEVERE NATIVE AORTIC STENOSIS WHO ARE CONSIDERED HIGH OR EXTREME RISK FOR SURGICAL AORTIC VALVE REPLACEMENT (AVR)." SINCE THIS WAS REPORTED THROUGH LITERATURE AND THERE IS NO EVIDENCE THE OFF-LABEL USE CAUSED OR CONTRIBUTED TO THE REPORTED ISSUES, A LETTER WILL NOT BE SENT. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO LABELING DESIGN OR MANUFACTURING OF THE DEVICE. B2: DEATH DATE IS ESTIMATED. B3: EVENT DATE IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: TRANSCATHETER AORTIC VALVE REPLACEMENT WITH INTRA-ANNULAR SELF-EXPANDING OR BALLOON-EXPANDABLE VALVES: THE MULTICENTER INTERNATIONAL NAVULTRA REGISTRY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22790550&pc=OZD | 1220648-2025-45798 | 2025-07-26 04:00 | Death | ABIOMED, INC. | 2025-08-12 04:00 | OZD | IMPELLA CP | Malposition of Device; Patient Device Interaction Problem | Hemolysis; Ischemia; Hematuria | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT UNDERGOING HIGH-RISK PERCUTANEOUS CORONARY INTERVENTION WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. WHILE ON SUPPORT, THE PATIENT'S URINE TURNED DARK RED AND LABS WERE HEMOLYZED. AN ECHOCARDIOGRAM WAS COMPLETED AND THE IMPELLA CP WAS MEASURED DEEPLY AT 4.7 CENTIMETERS. A NURSE PRACTITIONER REPOSITIONED THE IMPELLA CP FOR A FINAL MEASUREMENT OF 3.5 CENTIMETERS. THE PATIENT'S URINE WAS CLEAR AFTER REPOSITIONING. THE LEFT LOWER EXTREMITY WAS NOTED AS COLD AND MOTTLED. THE CLINICAL TEAM WAS AWARE. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: PRODUCT STATUS: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THERE IS UNCERTAINTY WHAT HAPPENED WITH THE DEVICE AFTER THE PATIENT'S DEMISE. THEREFORE, AN EVALUATION/ANALYSIS OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: SECTION: POTENTIAL ADVERSE EVENTS ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY¿ (INCLUDING VENTRICULAR PERFORATION). SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22790583&pc=OZP | 2032227-2025-234166 | 2024-07-08 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-08-12 04:00 | OZP | PUMP MMT-1780KPK 670G PATHWAY BLACK MG | Adverse Event Without Identified Device or Use Problem | Hyperglycemia; Diabetic Ketoacidosis | P150001 |  | 1 | Event Description: INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE CUSTOMER PASSED AWAY AT THE HOSPITAL ON (B)(6)2024. THE CUSTOMER WAS ADMITTED TO A HOSPITAL PRIOR TO THE REPORTED INCIDENT. THE CAUSE OF DEATH WAS HYPERGLYCEMIA, DIABETIC KETOACIDOSIS AND HEART ENLARGEMENT TREATED WITH IV INSULIN DRIP (INTRAVENOUS INSULIN INFUSION). THE CUSTOMER¿S BLOOD GLUCOSE WAS 900 MG/DL. THE EVENT INVOLVED PRODUCT(S) UNOMEDICAL, MMT-1780KPK, UNK_RESERVOIR. TROUBLESHOOTING WAS PERFORMED. THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP AT THE TIME OF DEATH. THE CUSTOMER WAS NOT USING SENSORS. THE INSULIN PUMP WAS LAST WORN ON (B)(6) 2024. NO PRODUCT RETURN IS REQUIRED FOR UNOMEDICAL, MMT-1780KPK, UNK_RESERVOIR. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22790615&pc=NVZ | 2017865-2025-96735 | 2025-07-14 04:00 | Death | ABBOTT | 2025-08-12 04:00 | NVZ | ENDURITY¿ CORE DR, JPD | Premature Discharge of Battery | Arrhythmia | P880086 |  | 1 | Event Description: THE PATIENT PRESENTED TO EMERGENCY SERVICE WITH NO HEART RATE. THE PATIENT WAS BEING PACED WITH AN EXTERNAL DEFIBRILLATOR AND PASSED AWAY DUE TO THE EVENT. THE HEALTHCARE PROVIDER SUSPECTED PREMATURE BATTERY DEPLETION. HOWEVER, TECHNICAL SUPPORT REVIEWED THE RECORDS AND DID NOT OBSERVE A BATTERY ANOMALY IN THE RECORDS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22790619&pc=NVN | 2017865-2025-96736 | 2025-07-14 04:00 | Death | ABBOTT | 2025-08-12 04:00 | NVN | TENDRIL STS | Over-Sensing | Arrhythmia | P960013 |  | 1 | Event Description: THE PATIENT PRESENTED TO EMERGENCY SERVICE WITH NO HEART RATE. THE PATIENT WAS BEING PACED WITH AN EXTERNAL DEFIBRILLATOR AND PASSED AWAY DUE TO THE EVENT. TECHNICAL SUPPORT REVIEWED THE RECORDS FROM MONTHS PRIOR TO THE DEATH OF THE PATIENT, AS NO RECORDS WERE PROVIDED FROM THE DAY THE PATIENT PASSED. FROM THE PRIOR RECORDS, NOISE RESULTING IN OVERSENSING AND PACING INHIBITION WERE OBSERVED ON THE RIGHT VENTRICULAR (RV) LEAD. THE HEALTH CARE PROVIDER WAS UNSURE IF THE RV LEAD CAUSED OR CONTRIBUTED TO THE DEATH THE OF THE PATIENT. Manufacturer Narrative: D4 - PRIMARY UNIQUE DEVICE IDENTIFICATION (UDI) NUMBER CANNOT BE PROVIDED AS A PRODUCT IDENTIFIER COULD NOT BE OBTAINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22768759&pc=LJT | 3006260740-2025-05677 | 2025-07-17 04:00 | Death | C.R. BARD, INC. (BASD) -3006260740 | 2025-08-11 04:00 | LJT | POWERPORT ISP M.R.I. IMPLANTABLE PORT, GROSHONG SINGLE-LUMEN, 8F | Adverse Event Without Identified Device or Use Problem | Abscess; Unspecified Infection; Insufficient Information | K063377 |  | 1 | Event Description: IT WAS REPORTED THROUGH LITIGATION PROCESS THAT SOMETIME LATER PORT PLACEMENT PROCEDURE, THE PATIENT DEVELOPED WITH UNSPECIFIED INFECTION. LATER PATIENT WAS DIAGNOSED WITH AORTIC VALVE INFECTION AND ENDOCARDITIS. ALSO, CONDITIONS LIKE CEREBRAL AND RENAL ABSCESS WERE OBSERVED. REPORTEDLY, PATIENT WAS EXPIRED. Manufacturer Narrative: H11: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. DATE OF DEATH FOR THIS PATIENT WAS NOT PROVIDED, DATE OF DEATH UPDATED AS (B)(6) 1900 FOR THE MDR SUBMISSION REQUIREMENT. SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22769482&pc=OZD | 1220648-2025-31131 | 2025-07-23 04:00 | Death | ABIOMED, INC. | 2025-08-11 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemolysis | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT FOLLOWING IMPELLA CP IMPLANT, THE PATIENT'S LABS HEMOLYZED AND THEIR URINE TURNED DARK BROWN. THE P-LEVEL WAS REDUCED AND THE PUMP WAS REPOSITIONED VIA ECHOCARDIOGRAM IN RESPONSE TO THE CONDITION. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22769581&pc=OZD | 1220648-2025-31133 | 2025-07-25 04:00 | Death | ABIOMED, INC. | 2025-08-11 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Ischemia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS ON IMPELLA 5.5 SUPPORT AND DURING SUPPORT, THE PATIENT DEVELOPED LIMB ISCHEMIA. THE PATIENT LOST PULSES AND DEVELOPED A COLD AND MODELED LEG. THE PEEL AWAY SHEATH WAS LEFT IN PLACE. THE PATIENT RECEIVED A FASCIOTOMY. LATER, THE PATIENT WAS NEVER ABLE TO BE WEANED FROM MAX DOSES OF EPINEPHRINE, VASOPRESSIN, LEVOPHED, AND MILRINONE AND CONTINUOUS RENAL REPLACEMENT THERAPY. THE PATIENT'S LIVER AND KIDNEYS BEGAN TO FAIL AND SEVERE ACIDOSIS SET IN LEADING TO AN EVENTUAL CODE WHERE THE PATIENT EXPIRED ON SUPPORT. RELATIONSHIP BETWEEN IMPELLA AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22769781&pc=MVK | 3008642652-2025-07693 | 2025-06-23 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-08-11 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED AN APPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 08:55:16 ON (B)(6) 2025. AT 21:13:19, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF. AT 21:13:54, THE PATIENT RECEIVED THE APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS ASYSTOLE. POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY. THE ELECTRODE BELT WAS DISCONNECTED AT 21:55:01 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR HAS BEEN COMPLETED. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION. THE ELECTRODE BELT HAS NOT BEEN RECOVERED. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22769860&pc=DSQ | 2916596-2025-05277 | 2025-07-31 04:00 | Death | THORATEC CORPORATION | 2025-08-11 04:00 | DSQ | HEARTMATE II LEFT VENTRICULAR ASSIST DEVICE | Adverse Event Without Identified Device or Use Problem | Intra-Abdominal Hemorrhage | P060040 |  | 1 | Event Description: IT WAS REPORTED THROUGH PATIENT OUTCOME THAT THE PATIENT PASSED AWAY DUE TO MULTIPLE ABDOMINAL HEMORRHAGES. THE DEATH WAS NOT CONSIDERED DEVICE RELATED, THE DEVICE WAS NOTED TO HAVE OPERATED AS EXPECTED AND WOULD NOT BE RETURNING FOR EVALUATION. IT WAS NOTED THAT DIAGNOSTIC TESTING WAS USED TO CONFIRM BLEEDING, BUT IT WAS UNKNOWN WHAT DIAGNOSTIC TESTING WAS PERFORMED. IT WAS NOTED THAT THE PATIENT AND THEIR FAMILY WERE NOT INTERESTED IN TREATMENT, AND PURSUED COMFORT CARE ELECTIVELY DEACTIVATING THE LEFT VENTRICULAR ASSIST DEVICE (LVAD). Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED BLEEDING AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. (B)(6) THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE II LVAS IFU REV. A AND THE HEARTMATE II PATIENT HANDBOOK REV. A ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING BLEEDING, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿ANTICOAGULATION¿), PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) RANGE, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22770145&pc=OZD | 1220648-2025-31141 | 2025-07-26 04:00 | Death | ABIOMED, INC. | 2025-08-11 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemolysis | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED AN IMPELLA CP WAS PLACED TO SUPPORT A PATIENT ADMITTED IN WITH CARDIAC HISTORY, CARDIOMYOPATHY AND HYPOXIA. THE TEAM OBSERVED SIGNS OF HEMOLYSIS. THE URINE COLOR CHANGED AND LABS HEMOLYZED. THE TEAM ASSESSED PUMP POSITIONING BUT THE PUMP WAS APPROPRIATELY PLACED WITHIN THE LEFT VENTRICLE. THE PUMP REMAINED ON WITH THE HEMOLYSIS AND FAMILY DECISION NOT TO ESCALATE. THE CARE WAS WITHDRAWN AFTER UNDER 76 HOURS. THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): "ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)." SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER: "EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT." IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE: "OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE." SECTION: HEMOLYSIS: "HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS." "PATIENT CONDITIONS: INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES: MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS." SECTION: SUCTION: "SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22770210&pc=NAY | 2955842-2025-32636 | 2025-02-13 05:00 | Death | INTUITIVE SURGICAL, INC | 2025-08-11 04:00 | NAY | DAVINCI XI | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Insufficient Information | K131861 |  | 1 | Event Description: A LITERATURE ARTICLE DESCRIBED A RETROSPECTIVE SINGLE-CENTER COHORT ANALYSIS STUDY THAT AIMED TO OPTIMIZE THE SURGICAL APPROACH TO ROBOTIC RADICAL NEPHROURETERECTOMY (RRNU) FOR PATIENTS WITH UROTHELIAL CARCINOMA IN THEIR KIDNEYS AND URETERS. THE APPROACH INVOLVED REDUCING THE NUMBER OF PORTS USED AND IMPROVING VISUALIZATION DURING THE PROCEDURE. THE MEDICAL RECORD REVIEW INVOLVED 15 RRNU PROCEDURES WITH BLADDER CUFF EXCISION (BCE) BETWEEN 2019 AND 2024 PERFORMED BY A SINGLE SURGEON USING A ONE-DOCKING, 3-PORT TECHNIQUE. THE ARTICLE NOTED ONE PATIENT EXPIRED AFTER A CARDIAC ARREST ON POSTOPERATIVE DAY (POD) 2. RESUSCITATION WAS INITIATED IMMEDIATELY, FOLLOWED BY AN EMERGENCY CORONARY ANGIOGRAPHY; THE PATIENT DEVELOPED IRREVERSIBLE CARDIAC ARRHYTHMIAS AND DID NOT RESPOND TO FURTHER RESUSCITATION. NO DEVICE ISSUES OR INTRAOPERATIVE COMPLICATIONS WERE NOTED IN THE ARTICLE. ADDITIONAL INFORMATION HAS BEEN REQUESTED FROM THE DESIGNATED AUTHOR WITH NO RESPONSE RECEIVED. Manufacturer Narrative: NO INVESTIGATION COULD BE PERFORMED AS THE ARTICLE DID NOT PROVIDE ANY SPECIFIC INFORMATION REGARDING THE PROCEDURE DATE OR DA VINCI SYSTEM IDENTIFIERS. NO DEVICE ISSUE, INTRAOPERATIVE COMPLICATIONS, OR ANY INDICATION THAT A DA VINCI PRODUCT CONTRIBUTED TO THE REPORTED EVENT WAS NOTED IN THE ARTICLE. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT A PATIENT FROM A LITERATURE ARTICLE DIED FROM A CARDIAC EVENT 2 DAYS AFTER A ROBOTIC NEPHROURETERECTOMY PROCEDURE. NO ALLEGATION HAS BEEN MADE AGAINST ANY INTUITIVE SURGICAL PRODUCT. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, NO INTUITIVE SURGICAL PRODUCT OR INSTRUMENT CONTRIBUTED TO THIS EVENT. SECTION B3 EVENT DATE IS SELECTED AS THE DATE THIS ARTICLE IS PUBLISHED: 13 FEBRUARY 2025. SECTION A2 PATIENT AGE WAS NOT SPECIFIED. THE STUDY AGE RANGE WAS 55¿86 YEARS. SECTION A3 PATIENT GENDER WAS NOT SPECIFIED. THE STUDY INCLUDED 12 MALE, 3 FEMALE PATIENTS. CITATION: ZOHAR, Y.; HEFER, B.; VAZANA, I.; JABAREEN, ET.AL. MINIMALLY INVASIVE ONE-DOCKING, TWO-TARGET, AND THREE-PORT ROBOTIC-ASSISTED NEPHROURETERECTOMY: REDEFINING SURGICAL APPROACH. CANCERS 2025, 17, 627. HTTPS://DOI.ORG/10.3390/CANCERS17040627.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22770452&pc=PCU | 3005099803-2025-03922 | 2020-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-08-11 04:00 | PCU | AXIOS STENT AND ELECTROCAUTERY ENHANCED DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K150692 |  | 1 | Event Description: BOSTON SCIENTIFIC CORPORATION BECAME AWARE OF AN EVENT INVOLVING AXIOS STENT AND ELECTROCAUTERY ENHANCED DELIVERY SYSTEMS THROUGH THE ARTICLE TITLED "OUTCOMES OF EUS-GUIDED GALLBLADDER DRAINAGE IN ACUTE CHOLECYSTITIS WITH CONTAINED PERFORATION: A PROSPECTIVE COHORT", BY GIUSEPPE VANELLA ET AL. PER THE ARTICLE, BETWEEN 2020 AND 2024, 48 PATIENTS SUFFERING FROM ACUTE CHOLECYSTITIS AND CONTAINED PERFORATION UNDERWENT EUS-GUIDED GALLBLADDER DRAINAGE (EUS-GBD) USING AXIOS AS LUMEN-APPOSING METAL STENTS (LAMS). ONE MORTALITY OCCURRED DURING AN ATTEMPTED DRAINAGE OF A LIVER ABSCESS ADJACENT TO THE GALLBLADDER. PLEASE SEE THE REFERENCE ARTICLE FOR FULL DETAILS. Manufacturer Narrative: BLOCK B2: THE EXACT DATE OF THE PATIENT'S DEATH WAS NOT REPORTED. THE RETROSPECTIVE STUDY START DATE OF JANUARY 1, 2020, IS USED FOR THE ESTIMATED DATE OF DEATH. BLOCK B3: THE EXACT DATE OF THE EVENT WAS NOT REPORTED. THE RETROSPECTIVE STUDY DATE OF (B)(6) 2020, IS USED FOR THE ESTIMATED DATE OF EVENT BETWEEN (B)(6) 2020, AND (B)(6) 2024. BLOCK D4, H4: DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WERE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER PRODUCT SPECIFIC INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. BLOCK G2: LITERATURE SOURCE: GIUSEPPE VANELLA; FRANCESCO FRIGO; ROBERTO LEONE; MATEO TACELLI; NICOLE LISCIA; LUIGI LETA; NICOLO PECORELLI; GIULIO BELFIORI, MICHELE RENE; MASSIMO FALCONI; GABRIELE CAPURSO; PAOLO G. ARCIDIACONO. "OUTCOMES OF EUS-GUIDED GALLBLADDER DRAINAGE IN ACUTE CHOLECYSTITIS WITH CONTAINED PERFORATION: A PROSPECTIVE COHORT" EUROPEAN SOCIETY OF GASTROINTESTINAL ENDOSCOPY, HTTPS://DOI.ORG/10.1016/J.GIE.2025.03.934. BLOCK H6: IMDRF IMPACT CODE F02 CAPTURES THE PATIENT'S DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22770852&pc=NAY | 2955842-2025-33374 | 2025-07-16 04:00 | Death | INTUITIVE SURGICAL, INC | 2025-08-11 04:00 | NAY | DAVINCI XI | Adverse Event Without Identified Device or Use Problem | Pulmonary Embolism; Respiratory Arrest | K131861 |  | 1 | Event Description: IT WAS REPORTED THAT AFTER AN UNEVENTFUL DA VINCI-ASSISTED SLEEVE GASTRECTOMY SURGICAL PROCEDURE, THE PATIENT EXPIRED IN THE POST ANESTHESIA CARE UNIT (PACU). THE SITE OPERATING ROOM DIRECTOR INFORMED THE INTUITIVE CLINICAL SALES REPRESENTATIVE THAT DURING RECOVERY IN THE POST-ANESTHESIA CARE UNIT THE PATIENT EXPERIENCED A PULMONARY EMBOLISM. CARDIOPULMONARY RESUSCITATION MEASURES WERE PERFORMED BUT WERE UNSUCCESSFUL. THE OR DIRECTOR PROVIDED THAT THE ROBOTIC SLEEVE GASTRECTOMY PROCEDURE WAS COMPLETED WITHOUT ANY COMPLICATIONS OR ANY DA VINCI-RELATED ISSUES. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: INSUFFICIENT INFORMATION IS AVAILABLE FOR PROCEDURE VERIFICATION; THEREFORE, NO DAVINCI SYSTEM OR INSTRUMENT LOGS ARE AVAILABLE FOR REVIEW. THE OR DIRECTOR REPORTED THAT THE PATIENT EXPIRED FROM A POST-PROCEDURE PULMONARY EMBOLISM. ATTEMPTS TO SPEAK WITH THE SURGEON AND THE SITE RISK MANAGER WERE NOT SUCCESSFUL. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT DIED FROM A PULMONARY EMBOLISM FOLLOWING AN UNREMARKABLE SLEEVE GASTRECTOMY. NO ALLEGATION HAS BEEN MADE AGAINST ANY INTUITIVE SURGICAL PRODUCT. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, NO INTUITIVE SURGICAL PRODUCT OR INSTRUMENT CONTRIBUTED TO THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22770941&pc=QFG | 3013756811-2025-189384 | 2025-07-19 04:00 | Death | TANDEM DIABETES CARE | 2025-08-11 04:00 | QFG | T:SLIM X2 INSULIN PUMP WITH CONTROL-IQ TECHNOLOGY | Adverse Event Without Identified Device or Use Problem | Diabetic Ketoacidosis | K203234 |  | 1 | Event Description: IT WAS REPORTED THAT A CUSTOMER HAD PASSED AWAY ON (B)(6) 2025, WITH THE CAUSE OF DEATH ATTRIBUTED TO DIABETIC KETOACIDOSIS, ASPIRATION PNEUMONIA, AND CANCER. THERE WAS NO ALLEGATION THAT THE PUMP OR RELATED SUPPLIES CONTRIBUTED TO THE DEATH, ALTHOUGH THE CUSTOMER WAS WEARING THE PUMP AT THE TIME. A REVIEW OF PUMP DATA WILL BE PERFORMED, AND THE RESULTS WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22770971&pc=NPT | 9617601-2025-00891 | 2025-08-06 04:00 | Death | MEDTRONIC MEXICO S. DE R.L. DE CV | 2025-08-11 04:00 | NPT | EVOLUT FX VALVE | Device Difficult to Setup or Prepare; Insufficient Information; Patient Device Interaction Problem | Stroke/CVA; Paresis | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE IMPLANT PROCEDURE, THE VALVE WAS MISLOADED BY THE MEDTRONIC REPRESENTATIVE DUE TO FOLDING/OVERLAP. A NEW VALVE WAS LOADED ONTO A DIFFERENT DELIVERY CATHETER SYSTEM (DCS) AND DEPLOYED WITH NO ISSUES. FOLLOWING THE PROCEDURE, THE PATIENT SUFFERED A STROKE IN RECOVERY AND SUBSEQUENTLY DIED OVERNIGHT IN THE INTENSIVE CARE UNIT (ICU).Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE IMPLANT PROCEDURE, THE VALVE WAS MISLOADED BY THE MEDTRONIC REPRESENTATIVE DUE TO FOLDING/OVERLAP. A NEW VALVE WAS LOADED ONTO A DIFFERENT DELIVERY CATHETER SYSTEM (DCS) AND DEPLOYED WITH NO ISSUES. FOLLOWING THE PROCEDURE, THE PATIENT SUFFERED A STROKE IN RECOVERY AND SUBSEQUENTLY DIED OVERNIGHT IN THE INTENSIVE CARE UNIT (ICU). ADDITIONAL INFORMATION WAS RECEIVED THAT THE STROKE WAS CONFIRMED WITH A COMPUTED TOMOGRAPHY (CT) SCAN. AS A RESULT OF THE STROKE, THE PATIENT WAS STRUGGLING TO WAKE UP IN RECOVERY, WAS NOT ABLE TO MOVE THEIR LEFT ARM AND HAD LEFT-SIDED FACIAL DROOPING. UNSPECIFIED TREATMENT WAS PROVIDED IN THE ICU. PER THE PHYSICIAN, THE CAUSE OF THE STROKE WAS UNKNOWN AND DIFFICULT TO DETERMINE, BUT IT IS A KNOWN RISK FOR ANY INTERVENTION AND ANY EQUIPMENT USED COULD HAVE CONTRIBUTED TO THE STROKE. THE PHYSICIAN DID NOT ATTRIBUTE THE STROKE TO THE DCS. THE MISLOAD OR FOLDING WAS NOT DETECTED ON THE SCREENING, BUT WHEN THE VALVE WAS DEPLOYED UP UNTIL 80% MARK, THE INFOLDWAS OBSERVED EXTENDING ALL THE WAY UP THE VALVE. THE VALVE WAS QUICKLY RECAPTURED AND THE NEW VALVE LOADED. Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME: EVOLUT FX DCS, PRODUCT ID: D-EVOLUTFX-34, LOT: 0012425596, USE BY DATE: 2026-09-03, UDI: (B)(4). CONTINUATION OF D10: PRODUCT ID: D-EVOLUTFX-34 (LOT: 0012425596); PRODUCT TYPE: 0195-HEART VALVES; PRODUCT ID: EVOLUTFX-34 (SERIAL: (B)(6); PRODUCT TYPE: 0195-HEART VALVES; PRODUCT ID: L-EVOLUTFX-34 (LOT: 0012210393); PRODUCT TYPE: 0195-HEART VALVES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED: B2, B5, H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22771095&pc=DTB | 3008973940-2025-05868 | 2020-09-28 04:00 | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-08-11 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Failure to Capture; High impedance; Under-Sensing; Device Dislodged or Dislocated | Cardiac Arrest; Hypoxia; Coma; Bronchopneumonia | P930039 |  | 1 | Event Description: IT WAS REPORTED APPROXIMATELY ONE MONTH POST IMPLANT THAT THE PATIENT EXPERIENCED CARDIAC ARREST AND HAD CARDIOPULMONARY RESUSCITATION ADMINISTERED. THE RIGHT VENTRICULAR (RV) LEAD "WAS NOT IN PLACE" AND EXHIBITED FAILURE TO CAPTURE, INCREASED IMPEDANCE AND NO SENSING. THE RV LEAD WAS REPOSITIONED. THE PATIENT HAD ENTERED A COMATOSE STATE DUE TO HYPOXIA AND ALSO EXPERIENCED BRONCHOPNEUMONIA. THE PATIENT IS DECEASED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22771653&pc=DSQ | 3007042319-2025-01897 | 2025-06-13 04:00 | Death | HEARTWARE, INC. | 2025-08-11 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO UNKNOWN CAUSES. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22772111&pc=DYE | 2015691-2025-06518 | 2025-06-05 04:00 | Death | EDWARDS LIFESCIENCES | 2025-08-11 04:00 | DYE | CARPENTIER-EDWARDS PERIMOUNT MAGNA MITRAL EASE PERICARDIAL BIOPROSTHESIS | Entrapment of Device | No Clinical Signs, Symptoms or Conditions | P860057 |  | 1 | Event Description: EDWARDS RECEIVED NOTIFICATION FROM INDIA THAT A 7300TFX25 VALVE IMPLANTED IN MITRAL POSITION WAS EXPLANTED AT IMPLANT DUE TO A SUTURE LOOP LEADING TO LEAFLET MOBILITY ISSUES AND SEVERE REGURGITATION. AS REPORTED, A POST-IMPLANTATION TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) REVEALED THAT ONE LEAFLET WAS STUCK AND NOT MOVING AND THE REGURGITATION WAS NOTED AFTER PROTAMINE ADMINISTRATION. THE VALVE WAS EXPLANTED AND A 25MM MECHANICAL VALVE WAS IMPLANTED IN REPLACEMENT. AS REPORTED, THE PATIENT WAS STILL IN CARDIOPULMONARY BYPASS AT THE TIME THE ISSUE WAS OBSERVED AND THE EXPLANT OF THE VALVE WAS DECIDED. THE PATIENT PASSED AWAY AND THE CAUSE OF DEATH WAS REPORTED TO BE A LOW CARDIAC OUTPUT FAILURE, CONSEQUENCE OF THE SIGNIFICANTLY PROLONGED PROCEDURE. Manufacturer Narrative: H11: ADDITIONAL MANUFACTURER NARRATIVE: H3: DEVICE EVALUATION: REPORT OF "LEAFLET MOBILITY ISSUES LEADING TO REGURGITATION" WAS CONFIRMED THROUGH OBSERVED SUTURE LOOP. X-RAY DEMONSTRATED WIREFORM INTACT. AS RECEIVED, A GAP WAS VISIBLE IN THE CENTRAL COAPTATION REGION, LEAFLETS 2 AND 3 WERE OBSERVED TO BE WAVY ALONG THE FREE MARGIN. WHAT APPEARED TO BE SUTURE TRACK INDENTATION WAS OBSERVED AT THE FREE MARGINS ON LEAFLET 2 AND 3 NEAR COMMISSURE 3 ON THE OUTFLOW ASPECT; INDICATING A SUTURE WAS LOOPED AROUND COMMISSURE 3. SUTURE HOLES WERE VISIBLE AROUND THE SEWING RING. SEWING RING CLOTH HAD MULTIPLE CUTS AROUND THE VALVE AND EXPOSED THE SILICONE. LAB FINDINGS WERE ALIGNED TO THE CUSTOMER PHOTOS. THE INVESTIGATION IS STILL IN PROGRESS; THEREFORE, A CONCLUSION HAS YET TO BE ESTABLISHED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY UPON INVESTIGATION COMPLETION. EDWARDS WILL CONTINUE TO REVIEW AND MONITOR ALL REPORTED EVENTS. TRENDS ARE MONITORED ON A MONTHLY BASIS AND IF ACTION IS REQUIRED, APPROPRIATE INVESTIGATION WILL BE PERFORMED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22772405&pc=MMX | 2017865-2025-97269 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-08-11 04:00 | MMX | AVEIR¿ RETRIEVAL SYSTEM CATHETER | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K213494 |  | 1 | Event Description: IT WAS REPORTED THAT IN THE LITERATURE REVIEW ENTITLED "SUCCESSFUL DEBULKING OF VEGETATION EXTENDING INTO THE RIGHT VENTRICLE USING A STEERABLE SHEATH AND LEADLESS RETRIEVAL SYSTEM", A PATIENT PASSED AWAY DUE TO AN UNKNOWN CAUSE. THERE WERE NO ALLEGATIONS MADE THAT THE PATIENT'S DEATH WAS CAUSED OR CONTRIBUTED TO BY THEIR IMPLANTED SYSTEM OR SYSTEM RELATED PROCEDURE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22772406&pc=PNJ | 2017865-2025-97270 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-08-11 04:00 | PNJ | AVEIR¿ LEADLESS PACEMAKER, RIGHT VENTRICLE | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P150035 |  | 1 | Event Description: IT WAS REPORTED THAT IN THE LITERATURE REVIEW ENTITLED "SUCCESSFUL DEBULKING OF VEGETATION EXTENDING INTO THE RIGHT VENTRICLE USING A STEERABLE SHEATH AND LEADLESS RETRIEVAL SYSTEM", A PATIENT PASSED AWAY DUE TO AN UNKNOWN CAUSE. THERE WERE NO ALLEGATIONS MADE THAT THE PATIENT'S DEATH WAS CAUSED OR CONTRIBUTED TO BY THEIR IMPLANTED SYSTEM OR SYSTEM RELATED PROCEDURE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22772426&pc=MGB | 2024168-2025-08124 | 2025-01-01 05:00 | Death | ABBOTT VASCULAR INC. | 2025-08-11 04:00 | MGB | PERCLOSE PROGLIDE¿ | Patient Device Interaction Problem | Hemorrhage/Bleeding; Ischemia; Myocardial Infarction; Shock | P960043 |  | 1 | Event Description: IT WAS REPORTED THAT THIS WAS A CLOSURE WITH FIVE PROGLIDE DEVICES. REPORTEDLY, THE FIVE SUTURES WERE PLACED, HOWEVER BLEEDING CONTINUED, RESULTING IN HEMORRHAGIC SHOCK. AFTER BEING CONVERTED TO OPEN SURGERY, THE PATIENT WAS WEANED OFF THE MACHINE AND LATER DIED OF MYOCARDIAL ISCHEMIA DUE TO MASSIVE BLOOD LOSS. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORDS FOR THIS PRODUCT COULD NOT BE REVIEWED AND A SIMILAR COMPLAINT QUERY COULD NOT BE PERFORMED BECAUSE THE PART NUMBER AND/OR LOT NUMBER WAS NOT REPORTED, AND THE DEVICE WAS NOT RETURNED. BASED ON THE INFORMATION PROVIDED, A DEFINITIVE CAUSE FOR THE REPORTED ISSUE COULD NOT BE DETERMINED. FACTORS THAT CONTRIBUTE TO FAILURE TO ACHIEVE HEMOSTASIS INCLUDE, BUT NOT LIMITED TO, POOR OR PARTIAL TISSUE CAPTURE, AIR KNOT. THE PATIENT EFFECTS OF HEMORRHAGE, ISCHEMIA AND DEATH ARE LISTED IN THE PROGLIDE INSTRUCTIONS FOR USE AS POTENTIAL ADVERSE EVENTS ASSOCIATED WITH USE OF THE VESSEL CLOSURE DEVICES. MEDICAL REVIEW CONCLUDED THAT BASED ON THE INFORMATION PROVIDED, IT IS POSSIBLE THAT THE USE OF THE PROGLIDE MAY HAVE CONTRIBUTED TO THE PATIENT¿S CLINICAL COMPLICATION(S), HOWEVER THE DATA FROM CASE #3 LACKS SUFFICIENT EVIDENCE TO DEFINITIVELY LINK DEVICE PERFORMANCE TO PATIENT DEATH. A MORE DEFINITIVE ASSESSMENT MAY BE COMPLETED PENDING RECEIPT OF FURTHER CASE INFORMATION. THE REPORTED PATIENT EFFECT OF HEMORRHAGE, ISCHEMIA, DEATH ARE KNOWN INHERENT RISKS OF THE PROCEDURE. A DEFINITIVE CAUSE FOR THE REPORTED PATIENT EFFECTS OF MYOCARDIAL INFARCTION AND SHOCK, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED. THE SUBSEQUENT TREATMENT APPEARS TO BE RELATED TO PROCEDURE CIRCUMSTANCE. BASED ON THE RESULTS OF THE COMPLAINT INVESTIGATION, THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO THE DESIGN, MANUFACTURE, OR LABELING OF THE DEVICE. B2, B3: DATES ESTIMATED. D4: THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. THE ADDITIONAL DEVICES REFERENCED IN B5 ARE BEING FILED UNDER SEPARATE MEDWATCH REPORT NUMBERS. ATTACHMENT: ARTICLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22772614&pc=NVN | 1028232-2025-03867 |  | Death | BIOTRONIK SE & CO. KG | 2025-08-11 04:00 | NVN | DEXTRUS 60 MODEL 4137 US 644137-102 | Over-Sensing | No Clinical Signs, Symptoms or Conditions | P950037 |  | 1 | Event Description: BOSTON SCIENTIFIC PROVIDED THE FOLLOWING INFORMATION: THIS LEAD WAS EXPLANTED DUE TO OVERSENSING. THE PATIENT EXPIRED AFTER LASER LEAD REMOVAL. SHOULD ADDITIONAL INFORMATION BE RECEIVED, THIS FILE WILL BE UPDATED. Manufacturer Narrative: COMBINATION PRODUCT: YES, THE LEAD IS CURRENTLY NOT AVAILABLE FOR ANALYSIS. NO CONCLUSION CAN BE DRAWN BASED ON AVAILABLE INFORMATION AT THE MOMENT. THE FILE IS CLOSED. THE INVESTIGATION WILL BE RE-OPENED SHOULD ADDITIONAL DATA BECOME AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22772886&pc=PNY | 2024168-2025-08297 | 2012-11-04 04:00 | Death | ABBOTT VASCULAR | 2025-08-11 04:00 | PNY | ABSORB BIORESORBABLE VASCULAR SCAFFOLD SYSTEM | Adverse Event Without Identified Device or Use Problem | Unspecified Vascular Problem | P150023 |  | 1 | Event Description: BIORESORBABLE SCAFFOLDS HAVE BEEN ASSOCIATED WITH INFERIOR OUTCOMES COMPARED TO CONTEMPORARY PERMANENT METALLIC DRUG-ELUTING STENTS (DES) FOR PERCUTANEOUS CORONARY INTERVENTION, PARTICULARLY WITHIN THE INITIAL YEARS AFTER IMPLANTATION; HOWEVER, THEIR LONG-TERM PERFORMANCE REMAINS UNCERTAIN. THIS STUDY AIMED TO EVALUATE THE LONG-TERM OUTCOMES OF SWEDISH PATIENTS TREATED WITH ABSORB BIORESORBABLE SCAFFOLDS (ABBOTT) VS CONTEMPORARY DES (BIOMATRIX, XIENCE PRIME, PROMUS ELEMENT, RESOLUTE INTEGRITY, ORSIRO, PROMUS ELEMENT PLUS, XIENCE XPEDITION, PROMUS PREMIER, SYNERGY, ULTIMASTER, BIOFREEDOM, RESOLUTE ONYX, XIENCE PRO X) IMPLANTED FROM NOVEMBER 4, 2011 TO MARCH 2, 2018, ASSESSING DEVICE-RELATED COMPLICATIONS AND EXAMINING POTENTIAL LATE BENEFITS. ADVERSE EVENTS ASSOCIATED WITH THE ABSORB SCAFFOLD INCLUDED THE FOLLOWING: STENT THROMBOSIS, IN-STENT RESTENOSIS, DEATH, MYOCARDIAL INFARCTION, AND TARGET LESION REVASCULARIZATION. THE ARTICLE CONCLUDED ABSORB BIORESORBABLE SCAFFOLDS SHOWED INFERIOR EARLY CLINICAL PERFORMANCE COMPARED WITH CONTEMPORARY DES, BUT AFTER 3 YEARS, DEVICE-RELATED OUTCOMES WERE SIMILAR, WHILE MYOCARDIAL INFARCTION RATES FAVORED ABSORB BIORESORBABLE SCAFFOLDS. THESE FINDINGS SUGGEST A COMPLEX TRADE-OFF BETWEEN EARLY DEVICE-RELATED EVENTS AND POTENTIAL LONG-TERM BENEFITS OF BIORESORBABLE SCAFFOLD-MEDIATED VASCULAR RESTORATION. DETAILS ARE LISTED IN THE ATTACHED ARTICLE, TITLED " LONG-TERM OUTCOME OF PERCUTANEOUS CORONARY INTERVENTION USING ABSORB BIORESORBABLE SCAFFOLD: A SCAAR STUDY.¿ NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORDS FOR THIS PRODUCT COULD NOT BE REVIEWED AND A SIMILAR COMPLAINT QUERY COULD NOT BE PERFORMED BECAUSE THE PART NUMBER AND/OR LOT NUMBER WAS NOT REPORTED, AND THE DEVICE WAS NOT RETURNED. THE REPORTED PATIENT EFFECT OF DEATH IS LISTED IN THE ABSORB BIORESORBABLE VASCULAR SCAFFOLD SYSTEMS INSTRUCTIONS FOR USE AS A KNOWN PATIENT EFFECT OF CORONARY STENTING PROCEDURES. A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECT, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING. LITERATURE ATTACHMENT: ARTICLE TITLE: "LONG-TERM OUTCOME OF PERCUTANEOUS CORONARY INTERVENTION USING ABSORB BIORESORBABLE SCAFFOLD: A SCAAR STUDY" B2: DATE OF DEATH ESTIMATED AS ON (B)(6) 2012. B3: DATE OF EVENT ESTIMATED AS ON (B)(6) 2012. D4: THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. THE ADDITIONAL PATIENT EFFECTS REPORTED IN THE ARTICLE ARE CAPTURED UNDER SEPARATE MEDWATCH REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22773417&pc=NIQ | 2124215-2025-55800 | 2022-03-24 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-08-11 04:00 | NIQ | SYNERGY? | Off-Label Use; Defective Device | Cardiac Arrest | P150003 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. TWO SYNERGIES 3.50 X 8 MM STENTS WERE IMPLANTED FOR TREATMENT OF A PATENT DUCTUS ARTERIOSUS (PDA). HOWEVER, DURING THE PROCEDURE, THE DEVICE FAILED TO FUNCTION AS INTENDED WHICH LED TO CARDIAC ARREST AND THE PATIENT DIED. Manufacturer Narrative: DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22773435&pc=DSQ | 2916596-2025-05177 | 2025-07-24 04:00 | Death | THORATEC CORPORATION | 2025-08-11 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Ischemia; Respiratory Failure; Multiple Organ Failure; Drug Resistant Bacterial Infection | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD A LONG BATTLE WITH CHRONIC DRIVELINE INFECTION AND PRESENTED TO A LOCAL EMERGENCY ROOM ON (B)(6) 2025 WITH RESPIRATORY FAILURE AND WAS INTUBATED PRIOR TO TRANSFER TO THE HOSPITAL. UPON ARRIVAL, THEY WERE FOUND TO BE BACTEREMIC WITH METHICILLIN-RESISTANT STAPHYLOCOCCUS AUREUS AND AN ELEVATED LACTIC OF 6.3. THEY ALSO HAD A SUPRATHERAPEUTIC INTERNATIONAL NORMALIZED RATIO (INR) OF 6.0. THEY HAD MINIMAL RESPONSIVENESS WITH WEANING SEDATION, SO A COMPUTED TOMOGRAPHY SCAN OF THE HEAD WAS OBTAINED ON (B)(6) 2025 WHICH SHOWED "MULTIFOCAL DEVELOPING HYPODENSITIES IN BILATERAL PARIETAL, OCCIPITAL LOBES, AND CEREBELLAR HEMISPHERES CONCERNING FOR ACUTE TO SUBACUTE ISCHEMIA, WITH MILD MASS EFFECT ON THE 4TH VENTRICLE." THE DECISION WAS MADE TO MOVE TOWARD COMFORT CARE MEASURES. CARE WAS WITHDRAWN. THE PATIENT PASSED AWAY IN THE INTENSIVE CARE UNIT (ICU), FAMILY AT BEDSIDE ON (B)(6) 2025. HISTORICALLY RELATED MFR COVERING THE ONSET OF THE CHRONIC DRIVELINE INFECTION: 2916596-2024-05295. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. A SPECIFIC CAUSE FOR THE PATIENT¿S INFECTIONS COULD NOT BE CONCLUSIVELY DETERMINED. MULTIPLE REQUESTS FOR ADDITIONAL INFORMATION REGARDING THIS EVENT WERE SUBMITTED TO THE ACCOUNT; HOWEVER, NO RESPONSE HAS BEEN RECEIVED AT THIS TIME. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. D, IS CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ OF THIS DOCUMENT LISTS DRIVELINE INFECTION, LOCALIZED INFECTION, STROKE, AND DEATH AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. THE IFU ALSO LISTS MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RIGHT HEART FAILURE, RESPIRATORY FAILURE, RENAL DYSFUNCTION, AND HEPATIC DYSFUNCTION) AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. ADDITIONALLY, SECTION 6 ¿PATIENT CARE AND MANAGEMENT¿ INCLUDES INFORMATION FOR CARING FOR THE DRIVELINE EXIT SITE AS WELL AS INFORMATION REGARDING HOW TO PREVENT AND CONTROL INFECTION. SECTION 6 PROVIDES INFORMATION REGARDING ANTICOAGULATION, INCLUDING RECOMMENDED INR (INTERNATIONAL NORMALIZED RATIO) RANGE, AS WELL AS CONSIDERATIONS FOR WHEN THERE IS A RISK OF BLEEDING. THE HEARTMATE 3 LVAS PATIENT HANDBOOK, REV. A, CONTAINS SEVERAL SECTIONS WITH INFORMATION ABOUT HOW TO PREVENT AND CONTROL INFECTION AS WELL AS INFORMATION FOR CARING FOR THE DRIVELINE EXIT SITE. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22773960&pc=LWS | 3004209178-2025-13584 | 2025-06-02 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-08-11 04:00 | LWS | VISIA AF MRI XT VR SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P980016 |  | 1 | Event Description: IT WAS REPORTED THAT THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) WAS REMOVED DUE TO AN UNRELATED INFECTION, AND THE PATIENT LATER EXPIRED. A CAUSE OF DEATH OR DATE OF EXPLANT WAS UNABLE TO BE OBTAINED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22775156&pc=DSQ | 2916596-2025-05216 | 2025-07-08 04:00 | Death | THORATEC CORPORATION | 2025-08-11 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN.Event Description: IT WAS REPORTED THAT THE PATIENT WAS ADMITTED TO THE HOSPITAL ON (B)(6) 2025 WITH A DEVASTATING INTRACEREBRAL HEMORRHAGE CONFIRMED BY COMPUTED TOMOGRAPHY (CT). THE PATIENT WAS NOT A SURGICAL CANDIDATE AND WAS TRANSFERRED TO NUCLEAR MEDICINE ON (B)(6) 2025, WHERE IMAGING CONFIRMED THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) WAS TRIGGERING THE VENTRICULAR CONUNDRUM WITH ABSENCE OF CEREBRAL BLOOD FLOW, AND WAS PRONOUNCED BRAIN DEAD VIA A NUCLEAR MEDICINE. PATIENT SUPPORT WAS DISCONTINUED SOON AFTER. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THIS EVENT HAD PREVIOUSLY BEEN REPORTED UNDER MFR 2916596-2025-05241. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22776412&pc=MKJ | 1218058-2025-00135 | 2025-07-21 04:00 | Death | BIO-DETEK INCORPORATED | 2025-08-11 04:00 | MKJ | ELECTRODES, ONESTEP COMPLETE, SINGLE, R SERIES | Unable to Obtain Readings | Insufficient Information | P160022 |  | 1 | Event Description: COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO DEFIBRILLATE A (B)(6) PATIENT (GENDER UNKNOWN), THE ASSOCIATED DEFIBRILLATOR WAS UNABLE TO OBTAIN AN ECG SIGNAL USING THESE ELECTRODE PADS. COMPLAINANT INDICATED THAT THE CLINICIAN OBTAINED ANOTHER SET OF ELECTRODE PADS TO CONTINUE TREATING THE PATIENT. COMPLAINANT INDICATED THAT THE PATIENT SUBSEQUENTLY EXPIRED. Manufacturer Narrative: JUSTIFICATION FOR NO UDI: THIS DEVICE HAS A UDI; SOME INFORMATION WAS NOT PROVIDED AND IS UNKNOWN; THEREFORE, A COMPLETE UDI CANNOT BE PROVIDED. ZOLL MEDICAL CORPORATION HAS NOT RECEIVED THE PRODUCT FOR EVALUATION AND THIS COMPLAINT IS STILL UNDER INVESTIGATION. Manufacturer Narrative: THE ONESTEP ELECTRODES (LOT # 0925J) WERE NOT RETURNED TO ZOLL MEDICAL CORPORATION FOR EVALUATION AND NO DEVICE LOGS OR PHOTOGRAPHS WERE PROVIDED. THE CUSTOMER REPORTED THAT THE PATIENT DEATH WAS UNRELATED TO THE LACK OF ECG SIGNAL FROM THE ONESTEP ELECTRODES. NO SHOCKS WERE DELIVERED DURING THE EVENT AND THE ELECTRODE PADS WERE PLACED DURING THE EVENT WHICH APPEARED TO RESOLVE THE ISSUE. A DEVICE HISTORY RECORD (DHR) REVIEW SHOWED NO DISCREPANCIES WITH THE LOT IN QUESTION. ALL IN PROCESS AND PRE-EVALUATION TESTING PERFORMED ON RETAINED SAMPLE PASSED. NO TREND IS ASSOCIATED WITH REPORTS OF THIS TYPE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22777025&pc=MNS | 2518422-2025-048868 | 2021-09-03 04:00 | Death | RESPIRONICS, INC. | 2025-08-11 04:00 | MNS | RENTAL DREAMSTATION AUTOSV | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K090539 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION REGARDING A RENTAL, DREAMSTATION, AUTOSV DEVICE. THE PATIENT WAS PASSED AWAY. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURE'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22777307&pc=NIK | 9614453-2025-03160 | 2025-07-25 04:00 | Death | MEDTRONIC EUROPE SARL | 2025-08-11 04:00 | NIK | COBALT¿ XT HF CRT-D MRI SURESCAN¿ | Failure to Deliver Shock/Stimulation | Insufficient Information | P010031 |  | 1 | Event Description: IT WAS REPORTED THAT THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) DID NOT SHOCK THE PATIENT IN AN EPISODE. IT WAS NOTED THAT AFTER ANTI-TACHYCARDIA PACING (ATP) WAS DELIVERED, THE RHYTHM SLOWED AND NEVER MET REDETECTION. SUBSEQUENT EPISODES NEVER REACHED VENTRICULAR FIBRILLATION (VF) NUMBER OF INTERVALS DETECTED. THE PATIENT ULTIMATELY PASSED AWAY. Manufacturer Narrative: CONTINUATION OF D10: 5076-45 LEAD; 383069 LEAD IMPLANTED: (B)(6) 2025. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22777673&pc=DSQ | 2916596-2025-05179 | 2025-07-25 04:00 | Death | THORATEC CORPORATION | 2025-08-11 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE CUSTOMER COMMUNICATED THAT NO ADDITIONAL INFORMATION WILL BE PROVIDED DUE TO PATIENT PRIVACY LAWS. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22778066&pc=NVZ | 2017865-2025-97296 | 2024-09-12 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-08-11 04:00 | NVZ | ELLIPSE DR | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P910023 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. IT WAS REPORTED THAT THE CAUSE OF DEATH IS CONGESTIVE HEART FAILURE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22778091&pc=NVY | 2017865-2025-97297 | 2024-09-12 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-08-11 04:00 | NVY | DURATA STS OPTIM ACTIVE FIXATION, DF-4 CONNECTOR | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P950022 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. IT WAS REPORTED THAT THE CAUSE OF DEATH IS CONGESTIVE HEART FAILURE. Manufacturer Narrative: ANALYSIS WAS NORMAL. NO ANOMALIES WERE NOTED. THE REPORTED EVENT WAS PATIENT DEATH. AS RECEIVED, ONLY THE CONNECTOR PORTION OF THE LEAD WAS RETURNED FOR ANALYSIS. ELECTRICAL TESTING THAT COULD BE MEASURED DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS. VISUAL INSPECTION OF THE LEAD DID NOT FIND ANY ANOMALIES WITH THE EXCEPTION OF PROCEDURAL DAMAGE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22778220&pc=NVN | 2017865-2025-97298 | 2024-09-12 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-08-11 04:00 | NVN | OPTISENSE | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. IT WAS REPORTED THAT THE CAUSE OF DEATH IS CONGESTIVE HEART FAILURE. Manufacturer Narrative: ANALYSIS WAS NORMAL. NO ANOMALIES WERE NOTED. THE REPORTED EVENT WAS PATIENT DEATH. AS RECEIVED, ONLY THE CONNECTOR PORTION OF THE LEAD WAS RETURNED FOR ANALYSIS. ELECTRICAL TESTING THAT COULD BE MEASURED DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS. VISUAL INSPECTION OF THE LEAD DID NOT FIND ANY ANOMALIES WITH THE EXCEPTION OF PROCEDURAL DAMAGE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22779448&pc=BZD | 2518422-2025-108980 | 2021-12-08 05:00 | Death | RESPIRONICS, INC. | 2025-08-11 04:00 | BZD | DREAMSTATION | Degraded | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT THE PATIENT PASSED AWAY FROM A HEART ATTACK. THE DEVICE WAS TURNED ON, BUT THE MASK WAS ON THE GROUND. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22779737&pc=DSQ | 2916596-2025-05326 | 2024-07-20 04:00 | Death | THORATEC CORPORATION | 2025-08-11 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN AT THIS TIME, FOUND AT HOME DECEASED. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. THE DEVICE WOULD NOT BE EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT'S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICLE ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REVISION, REV. D IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22767973&pc=FKX | 0008030665-2025-01823 | 2025-07-24 04:00 | Death | ERIKA DE REYNOSA, S.A. DE C.V. | 2025-08-10 04:00 | FKX | LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K181108 |  | 1 | Event Description: IT WAS REPORTED THAT THIS FEMALE PERITONEAL DIALYSIS (PD) PATIENT PASSED AWAY DURING SLEEP WHILE CONNECTED TO THE LIBERTY SELECT CYCLER. ADDITIONAL INFORMATION WAS OBTAINED THROUGH FOLLOW-UP WITH THE CLINIC. THE PATIENT¿S SPOUSE FOUND THE PATIENT DECEASED AT APPROXIMATELY 4:00AM ON (B)(6) 2025. THE PATIENT WAS STILL CONNECTED TO THE LIBERTY SELECT CYCLER. 911 WAS CALLED, HOWEVER; NO RESUSCITATIVE EFFORTS WERE MADE, AND THE PATIENT WAS NOT TRANSPORTED TO THE HOSPITAL. IT IS UNKNOWN WHAT PHASE AND CYCLE OF TREATMENT THE PATIENT WAS IN AT THE TIME OF DEATH AS THE FAMILY DID NOT SUBMIT THE TREATMENT FLOWSHEETS. THERE WERE NO REPORTED ISSUES WITH THE LIBERTY SELECT CYCLER PRIOR TO THE DEATH. A CAUSE OF DEATH WAS NOT PROVIDED, BUT IT WAS DOCUMENTED THAT THE PATIENT HAD OTHER COMORBIDITIES INCLUDING LEUKEMIA. THE PATIENT¿S DEATH WAS NOT RELATED TO ANY FRESENIUS DEVICE(S) OR PRODUCT(S) AND/OR THEIR DIALYSIS THERAPY. Manufacturer Narrative: CLINICAL REVIEW: THERE IS A TEMPORAL RELATIONSHIP BETWEEN PERITONEAL DIALYSIS UTILIZING THE LIBERTY SELECT CYCLER AND THE PATIENT EVENT OF DEATH. HOWEVER, THERE IS NO DOCUMENTATION OF A CAUSAL RELATIONSHIP BETWEEN THE PATIENT DEATH AND USE OF THE LIBERTY SELECT CYCLER. ADDITIONALLY, THERE WAS NO ALLEGATION OF A LIBERTY SELECT CYCLER MALFUNCTION OR DEFICIENCY. ALTHOUGH NO CAUSE OF DEATH WAS KNOWN, IT WAS REPORTED THAT THIS PATIENT HAD OTHER COMORBIDITIES (UNSPECIFIED) INCLUDING LEUKEMIA. ¿DIALYSIS PATIENTS ARE AT EXTRAORDINARILY HIGH RISK FOR DEATH WITH CARDIAC DISEASE BEING THE MAJOR CAUSE OF DEATH AND THE SINGLE, LARGEST, SPECIFIC CAUSE OF DEATH BEING ATTRIBUTED TO ARRHYTHMIC MECHANISMS OR SUDDEN CARDIAC ARREST (SCA)¿. BASED ON THE AVAILABLE INFORMATION AND NO ALLEGATION OR EVIDENCE OF MALFUNCTION OR DEFICIENCY, THE LIBERTY SELECT CYCLER CAN BE EXCLUDED AS THE CAUSE OF THE PATIENT¿S DEATH. THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22768252&pc=DSQ | 2916596-2025-05297 | 2025-07-16 04:00 | Death | THORATEC CORPORATION | 2025-08-10 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTI-ORGAN FAILURE. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED AND IT WAS NOTED THAT THE DEATH WAS RELATED TO A HEALTH CONDITION. THE DEVICE WAS WORKING AS INTENDED. NO AUTOPSY WAS PERFORMED, AND THE DEVICE WOULD NOT BE RETURNED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RESPIRATORY FAILURE, RIGHT HEART FAILURE, RENAL FAILURE, HEPATIC DYSFUNCTION) AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22768319&pc=DQY | 9681834-2025-00143 | 2025-07-10 04:00 | Death | TERUMO CORPORATION, ASHITAKA | 2025-08-10 04:00 | DQY | NAVICROSS | Break; Difficult to Remove | Hemorrhage/Bleeding; Low Blood Pressure/ Hypotension | K173799 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT DURING A RIGHT LOWER LIMB ANGIOGRAPHY FOR SUPERFICIAL FEMORAL ARTERY CHRONIC TOTAL OCCLUSION, ACCESS WAS OBTAINED VIA THE LEFT COMMON FEMORAL ARTERY USING A SIX FRENCH COOK ANSEL SHEATH. A 90-CENTIMETER QUICK-CROSS CATHETER AND A 260-CENTIMETER ANGLED GLIDEWIRE WERE USED TO CROSS THE CHRONIC TOTAL OCCLUSION, BUT THE ATTEMPT WAS UNSUCCESSFUL. SUBSEQUENTLY, ACCESS WAS OBTAINED VIA THE RIGHT ANTERIOR TIBIAL ARTERY USING AN 0.018-INCH, 300-CENTIMETER TRACK WIRE AND A 0.018-INCH, 135-CENTIMETER NAVICROSS CATHETER. THE NAVICROSS CATHETER WAS ADVANCED THROUGH THE ANTERIOR TIBIAL ARTERY, POPLITEAL ARTERY, AND INTO THE CHRONIC TOTAL OCCLUSION CAP. THE TRACK WIRE AND NAVICROSS CATHETER WERE ADVANCED INTO THE SIX FRENCH ANSEL SHEATH. THE WIRE WAS FLOSSED THROUGH THE SHEATH, WITH THE TRACK WIRE EXTENDING THROUGH THE SYSTEM. THE PHYSICIAN ATTEMPTED TO WITHDRAW THE NAVICROSS CATHETER VIA THE PEDAL ACCESS. DURING REMOVAL, THE NAVICROSS CATHETER DID NOT ADVANCE OUT AS EXPECTED. ONE PHYSICIAN HELD THE WIRE AT THE GROIN SHEATH, WHILE A FELLOW MANAGED THE PEDAL WIRE, AND ANOTHER PHYSICIAN ATTEMPTED TO PULL THE NAVICROSS CATHETER. DURING THIS PROCESS, THE CATHETER SEPARATED. THE BROKEN PORTION WAS REMOVED, BUT ANOTHER SEGMENT REMAINED ATTACHED TO THE TRACK WIRE. THE TRACK WIRE WAS CUT TO FACILITATE REMOVAL OF THE REMAINING CATHETER SEGMENT. THE BROKEN NAVICROSS CATHETER AND THE SEGMENT ATTACHED TO THE WIRE WERE RETAINED. THE DISTAL END OF THE TRACK WIRE WAS DISCARDED. ALL COMPONENTS OF THE NAVICROSS CATHETER AND TRACK WIRE WERE SUCCESSFULLY REMOVED FROM THE PATIENT. A SHEATH WAS PLACED IN THE PEDAL ACCESS. A QUICK-CROSS CATHETER WAS USED FROM THE GROIN TO ACCESS THE TIBIAL ARTERY. A SPARTACORE WIRE WAS ADVANCED DOWN THE LEG, AND A HAWK DEVICE WAS USED FOR ATHERECTOMY. AFTER SEVERAL PASSES WITH THE HAWK DEVICE, THE PATIENT DEVELOPED TACHYCARDIA AND HYPOTENSION. ANGIOGRAPHIC IMAGING REVEALED EXTRAVASATION IN THE ILIAC ARTERY AND AORTA. A CODA BALLOON WAS DEPLOYED, AND AORTIC/ILIAC STENTS WERE PLACED. DESPITE THESE INTERVENTIONS, THE PATIENT EXPIRED. THE PROCEDURE PERFORMED WAS A LEFT LEG ANGIOGRAPHY. THE EVENT OCCURRED INTRA-OPERATIVELY. THE ESTIMATED BLOOD LOSS WAS LESS THAN 250 CCS. ADDITIONAL INFORMATION WAS RECEIVED ON 14 JUL 2025: THE SAMPLE SHOWED BLOOD AND FLUID CONTAMINATION BUT NO PRESENCE OF ANTI-CANCER AGENTS. A PHYSICIAN INVOLVED IN THE CASE STATED THAT THERE WAS NO INDICATION THAT TERUMO PRODUCTS CAUSED OR CONTRIBUTED TO THE PATIENT'S DEATH. Manufacturer Narrative: A1: PATIENT IDENTIFIER: REQUESTED, NOT PROVIDED. A2: DATE OF BIRTH: REQUESTED, NOT PROVIDED . A4: WEIGHT: REQUESTED, NOT PROVIDED. D6A: IMPLANTED DATE: DEVICE WAS NOT IMPLANTED. D6B: EXPLANTED DATE: DEVICE WAS NOT EXPLANTED. THE ACTUAL DEVICE HAS NOT BEEN RETURNED FOR EVALUATION. THE INVESTIGATION IS CURRENTLY ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE. A HISTORICAL INVESTIGATION WAS CONDUCTED FOR THE PRODUCT ASSOCIATED WITH THE SPECIFIED PRODUCT CODE AND LOT NUMBER. NO ANOMALIES WERE IDENTIFIED IN THE MANUFACTURING OR SHIPPING INSPECTION RECORDS. ADDITIONALLY, NO SIMILAR REPORTS WERE FOUND IN THE PREVIOUS COMPLAINT FILES. REFER TO SECTION H10 FOR THE IMPORTER'S REPORT AND TO THE SECOND REPORT FOR DETAILS REGARDING THE OTHER DEVICE USED DURING THE REPORTED EVEN. TERUMO MEDICAL CORPORATION (TMC) (IMPORTER) REGISTRATION NO. (B)(4) IS SUBMITTING THIS REPORT ON BEHALF OF ASHITAKA FACTORY OF TERUMO CORPORATION (MANUFACTURER) REGISTRATION NO. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22768337&pc=DQY | 2243441-2025-00041 | 2025-07-10 04:00 | Death | TERUMO CORPORATION, ASHITAKA | 2025-08-10 04:00 | DQY | NAVICROSS | Break; Difficult to Remove | Hemorrhage/Bleeding; Low Blood Pressure/ Hypotension |  |  | 1 | Event Description: THE USER FACILITY REPORTED THAT DURING A RIGHT LOWER LIMB ANGIOGRAPHY FOR SUPERFICIAL FEMORAL ARTERY CHRONIC TOTAL OCCLUSION, ACCESS WAS OBTAINED VIA THE LEFT COMMON FEMORAL ARTERY USING A SIX FRENCH COOK ANSEL SHEATH. A 90-CENTIMETER QUICK-CROSS CATHETER AND A 260-CENTIMETER ANGLED GLIDEWIRE WERE USED TO CROSS THE CHRONIC TOTAL OCCLUSION, BUT THE ATTEMPT WAS UNSUCCESSFUL. SUBSEQUENTLY, ACCESS WAS OBTAINED VIA THE RIGHT ANTERIOR TIBIAL ARTERY USING AN 0.018-INCH, 300-CENTIMETER TRACK WIRE AND A 0.018-INCH, 135-CENTIMETER NAVICROSS CATHETER. THE NAVICROSS CATHETER WAS ADVANCED THROUGH THE ANTERIOR TIBIAL ARTERY, POPLITEAL ARTERY, AND INTO THE CHRONIC TOTAL OCCLUSION CAP. THE TRACK WIRE AND NAVICROSS CATHETER WERE ADVANCED INTO THE SIX FRENCH ANSEL SHEATH. THE WIRE WAS FLOSSED THROUGH THE SHEATH, WITH THE TRACK WIRE EXTENDING THROUGH THE SYSTEM. THE PHYSICIAN ATTEMPTED TO WITHDRAW THE NAVICROSS CATHETER VIA THE PEDAL ACCESS. DURING REMOVAL, THE NAVICROSS CATHETER DID NOT ADVANCE OUT AS EXPECTED. ONE PHYSICIAN HELD THE WIRE AT THE GROIN SHEATH, WHILE A FELLOW MANAGED THE PEDAL WIRE, AND ANOTHER PHYSICIAN ATTEMPTED TO PULL THE NAVICROSS CATHETER. DURING THIS PROCESS, THE CATHETER SEPARATED. THE BROKEN PORTION WAS REMOVED, BUT ANOTHER SEGMENT REMAINED ATTACHED TO THE TRACK WIRE. THE TRACK WIRE WAS CUT TO FACILITATE REMOVAL OF THE REMAINING CATHETER SEGMENT. THE BROKEN NAVICROSS CATHETER AND THE SEGMENT ATTACHED TO THE WIRE WERE RETAINED. THE DISTAL END OF THE TRACK WIRE WAS DISCARDED. ALL COMPONENTS OF THE NAVICROSS CATHETER AND TRACK WIRE WERE SUCCESSFULLY REMOVED FROM THE PATIENT. A SHEATH WAS PLACED IN THE PEDAL ACCESS. A QUICK-CROSS CATHETER WAS USED FROM THE GROIN TO ACCESS THE TIBIAL ARTERY. A SPARTACORE WIRE WAS ADVANCED DOWN THE LEG, AND A HAWK DEVICE WAS USED FOR ATHERECTOMY. AFTER SEVERAL PASSES WITH THE HAWK DEVICE, THE PATIENT DEVELOPED TACHYCARDIA AND HYPOTENSION. ANGIOGRAPHIC IMAGING REVEALED EXTRAVASATION IN THE ILIAC ARTERY AND AORTA. A CODA BALLOON WAS DEPLOYED, AND AORTIC/ILIAC STENTS WERE PLACED. DESPITE THESE INTERVENTIONS, THE PATIENT EXPIRED. THE PROCEDURE PERFORMED WAS A LEFT LEG ANGIOGRAPHY. THE EVENT OCCURRED INTRA-OPERATIVELY. THE ESTIMATED BLOOD LOSS WAS LESS THAN 250 CCS. ADDITIONAL INFORMATION WAS RECEIVED ON 14 JUL 2025: THE SAMPLE SHOWED BLOOD AND FLUID CONTAMINATION BUT NO PRESENCE OF ANTI-CANCER AGENTS. A PHYSICIAN INVOLVED IN THE CASE STATED THAT THERE WAS NO INDICATION THAT TERUMO PRODUCTS CAUSED OR CONTRIBUTED TO THE PATIENT'S DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22766632&pc=MCX | 3004742232-2025-00234 | 2025-06-26 04:00 | Death | CARDIOVASCULAR SYSTEMS INCORPORATED | 2025-08-09 04:00 | MCX | DIAMONDBACK CORONARY | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Vascular Dissection | P130005 |  | 1 | Event Description: IT WAS REPORTED THAT THE PROCEDURE WAS TO TREAT THE RIGHT CORONARY ARTERY WITH HEAVY CALCIFICATION, HEAVY TORTUOSITY AND 80% STENOSIS. THE DRAGONFLY OPSTAR IMAGING CATHETER WAS ADVANCED TO THE LESION WITH RESISTANCE NOTED BUT WAS ABLE TO CROSS THE LESION AND PERFORM PULLBACK. A DISSECTION WAS SUSPECTED TO BE CAUSED BY THE DRAGONFLY AND A DIAMONDBACK 360 CORONARY ORBITAL ATHERECTOMY DEVICE (OAD) WAS USED TO PERFORM ORBITAL ATHERECTOMY. THERE WAS A CLINICALLY SIGNIFICANT DELAY IN THE PROCEDURE DUE TO THE DISSECTION AND UNSPECIFIED TREATMENT AND THE PATIENT WAS UNSTABLE AND SYMPTOMATIC THROUGHOUT THE PROCEDURE. THE PATIENT EXPERIENCED CARDIAC ARREST AND EXPIRED. IT WAS INDICATED DRAGONFLY OPSTAR IMAGING CATHETER WAS ADVANCED TO THE LESION WITH RESISTANCE NOTED WITH THE ANATOMY BUT WAS ABLE TO CROSS THE LESION AND PERFORM PULLBACK. A DISSECTION WAS SUSPECTED TO BE CAUSED BY THE DRAGONFLY. THE DELAY IN THE PROCEDURE WAS DUE TO THE DISSECTION AND TREATMENT WITH BALLOONING AND STENTING. THE CARDIAC ARREST WAS TREATED WITH MANUAL CARDIOPULMONARY RESUSCITATION AND DEFIBRILLATION AND THE PATIENT WAS UNSTABLE AND SYMPTOMATIC THROUGHOUT THE PROCEDURE. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. PRODUCTION RECORD AND CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A QUERY OF THE COMPLAINT HANDLING DATABASE FOR THE REPORTED LOT REVEALED THERE IS NO INDICATION OF A LOT SPECIFIC ISSUE. BASED ON THE INFORMATION RECEIVED, THE INVESTIGATION DETERMINED THAT THE REPORTED PATIENT HARMS INCLUDING CARDIAC ARREST, VASCULAR DISSECTION, DELAY TO TREATMENT AND DEATH APPEAR TO BE RELATED TO OPERATIONAL CONTEXT. IN THIS CASE, IT IS POSSIBLE THAT THE CARDIAC ARREST, VASCULAR DISSECTION, DELAY TO TREATMENT AND DEATH ARE DUE TO DEVICE PLACEMENT OR USE TECHNIQUES EMPLOYED; HOWEVER, SINCE THE DEVICE WAS NOT RETURNED THIS COULD NOT BE CONFIRMED. BASED ON THE RESULTS OF THE COMPLAINT INVESTIGATION, THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO THE DESIGN, MANUFACTURE, OR LABELING OF THE DEVICE. H10: THE DRAGONFLY OPSTAR IMAGING CATHETER REFERENCED IN B5 IS FILED UNDER SEPARATE MEDWATCH REPORT NUMBER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22746479&pc=DXO | 2024168-2025-08248 | 2025-07-16 04:00 | Death | LIGHTLAB IMAGING, INC. | 2025-08-08 04:00 | DXO | PRESSUREWIRE¿ | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Ventricular Fibrillation | K180558 |  | 1 | Event Description: IT WAS REPORTED THAT THE PROCEDURE WAS PERFORMED IN THE LEFT ANTERIOR DESCENDING (LAD) ARTERY. THE PRESSUREWIRE X WIRELESS GUIDEWIRE WAS SUCCESSFULLY USED AND THE DECISION WAS NOT TO STENT. TWO MINUTES AFTER THE PRESSURE WIRE WAS REMOVED FROM THE BODY, THE PATIENT PASSED AWAY. THE PATIENT HAD SUDDEN SEVERE HYPOTENSION AT THE END OF THE CORONARY ANGIOGRAPHY/FFR MEASUREMENT, FOLLOWED BY PULSELESS ELECTRICAL ACTIVITY (PEA). NITRATE AND PAPAVERINE MEDICATIONS WERE GIVEN FOR THE MEASUREMENTS THAT IN THE PHYSICIAN'S OPINION, MAY HAVE CAUSED OR CONTRIBUTED TO THE PATIENT'S DEATH. REANIMATION WAS PERFORMED FOR 30 MINUTES WHICH INCLUDED CARDIOPULMONARY RESUSCITATION (CPR), ADRENALINE SEVERAL TIMES (UP TO 10 MG), NORADRENALINE, 2-3 TIMES DC SHOCK DUE TO VFIB (OCCURRING LATER DURING REANIMATION NOT IN THE BEGINNING). THE PATIENT HAD SEVERE VASOPLEGIA, IT WAS IMPOSSIBLE TO GET ANY SPONTANEOUS PRESSURE BACK AFTER THE EPISODE OF HYPOTENSION. DURING CPR, THE PRESSURE WAS GOOD, BUT EVERY TIME CPR WAS STOPPED, THERE WAS TOTALLY NO PRESSURE AT ALL, DESPITE ELECTRICAL ACTIVITY. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORDS FOR THIS PRODUCT COULD NOT BE REVIEWED AND A SIMILAR COMPLAINT QUERY COULD NOT BE PERFORMED BECAUSE THE LOT NUMBER WAS NOT REPORTED, AND THE DEVICE WAS NOT RETURNED. IN THIS CASE, THERE WAS NO REPORTED MALFUNCTION ASSOCIATED WITH THE PRESSUREWIRE X DEVICE. THE REPORTED PATIENT EFFECTS OF HYPOTENSION, VENTRICULAR FIBRILLATION, AND DEATH ARE LISTED IN THE PRESSUREWIRE INSTRUCTION FOR USE AS KNOWN POTENTIAL COMPLICATIONS WHICH MAY BE ENCOUNTERED DURING ALL CATHETERIZATION PROCEDURES. IN THIS CASE, A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECTS, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED. A DEFINITIVE CAUSE COULD NOT BE DETERMINED FOR THE REPORTED PATIENT EFFECT OF HYPOTENSION WHICH RESULTED IN VENTRICULAR FIBRILLATION, MEDICATION REQUIRED, UNEXPECTED MEDICAL INTERVENTION, AND DEATH. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN, OR LABELING. D4: A PARTIAL UDI WAS PROVIDED AS THE LOT NUMBER IS NOT KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22746604&pc=OZD | 1220648-2025-45707 | 2025-07-22 04:00 | Death | ABIOMED, INC. | 2025-08-08 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT IN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. AFTER SUPPORT WAS INITIATED, IT WAS REPORTED THAT THE PATIENT EXPERIENCED ABSENT PEDAL AND POSTERIOR TIBIAL PULSES TO THE DISTAL EXTREMITY OF THE IMPELLA LEG, BUT THE MEDICAL STAFF WAS ABLE TO OBTAIN POPLITEAL PULSES WITH DOPPLER. IT WAS NOTED THAT THE PATIENT CONTINUED TO EXPERIENCE SELF-TERMINATING VENTRICULAR TACHYCARDIA WITH MINIMAL PULSATILITY AND THE PATIENT¿S NEURO PROGNOSIS APPEARED TO BE POOR WITH SEIZURE ACTIVITY NOTED, AND OVERALL PROGNOSIS GUARDED. THE PATIENT¿S NEURO STATUS REMAINED UNCHANGED IN THE FOLLOWING DAYS AND NOTED THAT THE PATIENT DID NOT RESPOND TO PAIN STIMULUS. THE PATIENT¿S FAMILY DECIDED TO WITHDRAW CARE, CARE WAS WITHDRAWN AND THE PATIENT EXPIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22746926&pc=DTB | 2649622-2025-21517 | 2018-03-05 05:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-08-08 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD EXPIRED AND A CAUSE OF DEATH WAS NOT PROVIDED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22746931&pc=NVZ | 3004209178-2025-13460 | 2018-03-05 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-08-08 04:00 | NVZ | ADVISA MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information; No Clinical Signs, Symptoms or Conditions | P980035 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD EXPIRED AND A CAUSE OF DEATH WAS NOT PROVIDED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22746932&pc=DTB | 2649622-2025-21518 | 2018-03-05 05:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-08-08 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD EXPIRED AND A CAUSE OF DEATH WAS NOT PROVIDED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22747633&pc=OZD | 1220648-2025-31105 | 2025-07-11 04:00 | Death | ABIOMED, INC. | 2025-08-08 04:00 | OZD | IMPELLA 5.5 | Restricted Flow rate | Fever; Hemorrhage/Bleeding; Ventricular Fibrillation; Convulsion/Seizure; Unspecified Vascular Problem | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WITH POSTCARDIOTOMY CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA 5.5 FOR MECHANICAL CIRCULATORY SUPPORT AND EXPERIENCED A NUMBER OF COMPLICATIONS AND WOULD SUBSEQUENTLY EXPIRE. THE PATIENT WAS ADMITTED AFTER COMING TO THE EMERGENCY DEPARTMENT WITH COMPLAINTS OF LEFT-SIDED CHEST PAIN AND SHORTNESS OF BREATH FOR SEVERAL DAYS. THE PATIENT UNDERWENT A LEFT HEART CATHETERIZATION AND WAS DIAGNOSED WITH MULTI-VESSEL DISEASE. SURGICAL CONSULT WAS INITIATED. THE DECISION WAS MADE TO IMPLANT AN IMPELLA 5.5 WHICH WAS SUCCESSFULLY COMPLETED. AN AXILLARY CUT DOWN MADE AND A 10 X 30 MILLIMETER HEMASHIELD PLATINUM WAS ANASTOMOSED. THE LEFT VENTRICLE WAS ACCESSED, AND A GUIDEWIRE EXCHANGE OCCURRED. THE IMPELLA 5.5 WAS ADVANCED ACROSS THE AORTIC VALVE. THE WIRE WAS PULLED AND SUPPORT WAS INITIATED AT PERFORMANCE LEVEL P-2. SUPPORT SLOWLY RAMPED UP AND THEN HAD TO GO BACK ONTO BYPASS. DURING THE PROCEDURE, WHEN TRYING TO SEPARATE FROM PUMP, THE PATIENT BECAME PULSELESS, DEFIBRILLATED MULTIPLE TIMES AND CRASHED BACK ONTO THE PUMP IN WHICH THE PREVIOUSLY PLACED GRAFT HAD TO BE REPAIRED. THE PATIENT RECEIVED TEN UNITS OF PACKED RED BLOOD CELLS, TWO UNITS OF PLATELETS, SIX UNITS OF FRESH FROZEN PLASMA AND TWO UNITS OF CRYOPRECIPITATE. AN IMPELLA RP FLEX WAS PLACED AND THEN FLOWS TITRATED UP. ECHOCARDIOGRAM MEASURED MID FLOW 4.3 CENTIMETERS FROM THE AORTIC VALVE ANNULUS. THE GRAFT WAS CUT DOWN AND SUTURED INTO PLACE. THE IMPELLA WAS SECURED AND POCKET PACKED AND CHOSEN TO BE LEFT OPEN. THE PATIENT HAD TO BE DEFIBRILLATED TWO MORE TIMES BEFORE LEAVING THE OPERATING ROOM. THE PATIENT WAS PLANNED FOR A TRANSFER TO A DIFFERENT FACILITY. THE FOLLOWING DAY, THE PATIENT BEGAN TO HAVE SEIZURE LIKE ACTIVITY AND SPIKED A TEMPERATURE. THE PATIENT WAS HAVING MULTIPLE SUCTION EVENTS AND DEVICES PERFORMANCE LEVEL WAS DECREASED. THE PATIENT WAS GETTING ALBUMIN AND DISCUSSION WAS MADE FOR THE NEED FOR VOLUME VERSE INCREASING VASOPRESSORS AND INOTROPES. THE PATIENT, AT THIS TIME, WAS TOLERATING IMPELLA RP FLEX AT P-7 WITH 2.9 LITERS PER MINUTE (L/MIN) OF FLOW AND IMPELLA 5.5 IS AT P-6 WITH 3.5 L/MIN OF FLOW. THE PATIENT CONTINUED ON SUPPORT WITH PLANS REMAINING FOR TRANSFER TO AN OUTSIDE FACILITY. THE PATIENT REMAINED INTUBATED AND SEDATED TOLERATING P-9 ON THE IMPELLA RP FLEX WITH 3.1 L/MIN OF FLOW AND IMPELLA 5.5 AT P-9 WITH 5.1 L/MIN OF FLOW. ON DAY THREE OF SUPPORT, OVERNIGHT AND IN THE MORNING, THE PATIENT HAD TO BE SHOCKED MULTIPLE TIMES DUE TO GOING INTO VENTRICULAR FIBRILLATION (VF). BEDSIDE, ECHOCARDIOGRAM WAS PERFORMED. EJECTION FRACTION WAS APPROXIMATELY TEN PERCENT AND THE RIGHT VENTRICLE WAS NOT REALLY MOVING. THE POSITION WAS VERIFIED AT 5.15 CENTIMETERS. A NOTED PLAN WAS TO DO A BEDSIDE WASHOUT AND AN ELECTROENCEPHALOGRAM TO DETERMINE NEUROLOGY STATUS. THE FOLLOWING DAY, THE FAMILY WAS CALLED AND DISCUSSED WITHDRAWAL OF CARE WITHIN THE NEXT DAY OR SO AS THE PHYSICIAN NOTED SUSPICION OF A MASSIVE NEUROLOGICAL INJURY. THE PATIENT WAS TOO UNSTABLE FOR SCANS. THE PHYSICIAN WAS UNSURE IF IT WAS A BLEED VERSUS ISCHEMIC. CORE TEMPERATURE WAS 40 CELSIUS, LEFT VENTRICULAR EJECTION FRACTION WAS FIVE PERCENT AND RIGHT VENTRICULAR EJECTION FRACTION WAS ¿WORSE.¿ THE FOLLOWING DAY CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. ADDITIONAL INFORMATION WAS RECEIVED. THE REASON FOR THE VENTRICULAR FIBRILLATION WAS NOT ASSOCIATED WITH THE PUMP BUT WAS DUE TO POOR CARDIAC OUTPUT/PERFUSION ALONG WITH ELECTROLYTE IMBALANCES. Manufacturer Narrative: THE MEDICAL SAFETY OFFICER REVIEWED THE EVENT AND DETERMINED THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA 5.5 AS AN ASSOCIATED FACTOR IN THE PATIENT'S DEMISE OUTCOME (GIVEN THIS DEVICE-RELATED ADVERSE EVENTS OF VASCULAR/DAMAGE BLEEDING, TEMPERATURE, VENTRICULAR FIBRILLATION AND SEIZURE). EVENTS (TEMPERATURE AND VENTRICULAR FIBRILLATION) OCCURRING WITH THE IMPELLA RP FLEX ARE TO BE REPORTED AS SERIOUS INJURY. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿ SECTION: TABLE 3.2 IMPELLA CATHETER COMPONENTS ¿THE INFUSION FILTER PREVENTS BACTERIAL CONTAMINATION AND AIR FROM ENTERING THE PURGE LUMEN.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22747794&pc=OZD | 1220648-2025-31106 | 2025-07-07 04:00 | Death | ABIOMED, INC. | 2025-08-08 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Atrial Fibrillation; Thrombocytopenia; Aortic Valve Insufficiency/ Regurgitation | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT A PATIENT WAS ON IMPELLA 5.5 SUPPORT AND EXPERIENCED THROMBOCYTOPENIA. ON DAY 14 OF IMPELLA SUPPORT, THE PATIENT STARTED DECOMPENSATING AND HAD ATRIAL FIBRILLATION WITH RAPID VENTRICULAR RESPONSE. ADDITIONALLY, THAT PATIENT HAD SEVERE AORTIC INSUFFICIENCY (AI) WITH A POSSIBLE LEAFLET TEAR NOTED ON TRANSESOPHAGEAL ECHOCARDIOGRAM. THE PHYSICIAN FELT IT MAY HAVE STARTED AS A SMALL PERFORATION AS A RESULT OF WIRE INSERTION DURING IMPELLA IMPLANT AND HAD GROWN LARGER OVER TIME. HOWEVER, WE WILL NOT KNOW FOR SURE UNTIL AORTIC VALVE REPLACEMENT (AVR) IS COMPLETE AND THE VALVE CAN BE INSPECTED. THE IMPELLA INSERTION AND LEFT VENTRICLE ACCESS WAS SMOOTH WITHOUT REPEATED ATTEMPTS TO CROSS THE ATRIOVENTRICULAR WITH STIFF OR MULTIPLE WIRES. IT WAS FURTHER REPORTED THAT THE PATIENT WAS NOT DEEMED A CANDIDATE FOR AORTIC VALVE REPLACEMENT AVR AND CARE WAS WITHDRAWN. THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK, SECTION: POTENTIAL ADVERSE EVENTS (UNITES STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: AXILLARY INSERTION OF IMPELLA 5.5 CATHETER: ¿USE FLUOROSCOPY FOR PLACEMENT IMPELLA CATHETER PERFORMANCE WILL BE COMPROMISED IF CORRECT PLACEMENT CANNOT BE CONFIRMED. WHILE OTHER IMAGING TECHNIQUES, SUCH AS TRANSESOPHAGEAL ECHOCARDIOGRAPHY (TEE), CAN HELP CONFIRM THE POSITION OF THE IMPELLA CATHETER AFTER PLACEMENT, TEE DOES NOT ALLOW VISUALIZATION OF THE ENTIRE CATHETER ASSEMBLY AND IS INADEQUATE FOR RELIABLY PLACING THE IMPELLA CATHETER ACROSS THE AORTIC VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION, SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER, SECTION: DIRECT AORTIC INSERTION, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: ¿IN PATIENTS WITH TRANSCATHETER AORTIC VALVES POSITION THE IMPELLA SYSTEM CAREFULLY TO AVOID INTERACTION WITH THE TRANSCATHETER AORTIC VALVE PROSTHESIS ¿UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH.¿ IN THIS SITUATION, AVOID REPOSITIONING WHILE THE DEVICE IS RUNNING; TURN THE DEVICE TO P0 DURING REPOSITIONING OR ANY MOVEMENT THAT COULD BRING THE OUTLET WINDOWS INTO PROXIMITY TO THE VALVE STENT STRUCTURES. IF THERE IS LOW FLOW OBSERVED IN A PATIENT IMPLANTED WITH A TRANSCATHETER AORTIC VALVE PROSTHESIS, CONSIDER DAMAGE OF THE IMPELLER AND REPLACE THE IMPELLA AS SOON AS POSSIBLE.¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF IMPELLA CATHETER: ¿IF THE INLET AREA OF THE IMPELLA CATHETER IS TOO CLOSE TO OR ENTANGLED IN THE PAPILLARY MUSCLE AND/OR SUBANNULAR STRUCTURES SURROUNDING THE MITRAL VALVE, IT CAN AFFECT MITRAL VALVE FUNCTION AND NEGATIVELY IMPACT CATHETER FLOW.¿ SECTION: UNDERSTANDING AND MANAGING IMPELLA CATHETER POSITION ALARMS. ¿IF THE IMPELLA CATHETER IS COMPLETELY OUT OF THE VENTRICLE, DO NOT ATTEMPT TO REPOSITION THE CATHETER ACROSS THE VALVE WITHOUT A GUIDEWIRE.¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22748411&pc=OZD | 1220648-2025-31112 | 2025-07-25 04:00 | Death | ABIOMED, INC. | 2025-08-08 04:00 | OZD | IMPELLA CP | Positioning Failure | Hemolysis; Thrombosis/Thrombus | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT THE HAD A CP PUMP PLACED FOR A CRITICAL PATIENT WITH PRIOR CORONARY ARTERY BYPASS GRAFT (CABG), MYOCARDIAL INFARCTION (MI), ST-SEGMENT ELEVATION MYOCARDIAL INFARCTION (STEMI), AND HYPOTENSION WITH VENTRICULAR FIBRILLATION AND DEFIBRILLATION. THE PUMP WAS PLACED AND THE PUMP WAS PERSISTENTLY IN POOR POSITION AND INTERFERING WITH THE MITRAL VALVE. THE TEAM ALSO OBSERVED SIGNS OF HEMOLYSIS. THE PUMP REPOSITIONING AND OTHER TROUBLESHOOTING MEASURES FAILED. THE LABS WERE HEMOLYZED. WHEN IMAGING FOR THE MULTIPLE ORGAN FAILURE THE TEAM ALSO OBSERVED THROMBUS AND EMBOLIC EVENT TO THE SUPERIOR MESENTERIC ARTERY. THE CARE WAS WITHDRAWN AFTER 2 DAYS OF SUPPORT. Manufacturer Narrative: THE IMPELLA DEVICE HAS NOT BEEN RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE IFU STATES: IMPELLA CP WITH SMART ASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22748543&pc=OZD | 1220648-2025-45726 | 2025-07-11 04:00 | Death | ABIOMED, INC. | 2025-08-08 04:00 | OZD | IMPELLA 5.5 | Restricted Flow rate | Bacterial Infection; Hemorrhage/Bleeding; Sepsis | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WITH CARDIOMYOPATHY WAS IMPLANTED WITH AN IMPELLA 5.5 DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS). THE PATIENT WAS ESCALATED TO AN IMPELLA 5.5 WHILE AWAITING ORTHOTOPIC HEART TRANSPLANT WHICH WAS SUCCESSFULLY COMPLETED. THE PATIENT WAS SENT THE UNIT, REST AND RECOVER, ON IMPELLA MCS. HOWEVER, LATER THAT EVENING INTERMITTENT SUCTION AT PERFORMANCE LEVEL P-6 WAS ENCOUNTERED. BEDSIDE ECHO TO BE COMPLETED. INTERMITTENT SUCTION CONTINUED OVERNIGHT, RESOLVED AT P-4. REMAINS INTUBATED AND CONTINUED IMPELLA MCS. FIVE UNITS OF BLOOD WERE GIVEN OVERNIGHT. THE FOLLOWING MORNING THE PATIENT PLANNED FOR COLONIC EMBOLIZATION IN INTERVENTIONAL RADIOLOGY FOR GASTROINTESTINAL BLEED. AT THIS TIME PATIENT WAS STATUS 7 FOR TRANSPLANT. TWO DAYS LATER, THE PATIENT REMAINED SEDATED AND INTUBATED WITH INHALED VELETRI AND OPEN ABDOMEN WITH PLANS FOR POTENTIAL ABDOMINAL CLOSURE THE FOLLOWING DAY. IMPELLA 5.5 DEVICE AT P-4 WITH UNSUCCESSFUL ATTEMPTS TO INCREASE THE PERFORMANCE LEVEL DUE TO SUCTION EACH TIME. THERE WERE NO PLANS TO CHANGE THE SUPPORT UNTIL AFTER ABDOMINAL CLOSURE. THE NEXT AFTERNOON THE ABDOMINAL CLOSURE WAS COMPLETED AND AN ATTEMPT TO INCREASE PERFORMANCE LEVEL TO P-5 WAS UNSUCCESSFUL AS THERE WAS, AGAIN, IMMEDIATE SUCTION EVENTS. POINT OF CARE ULTRASOUND/FORMAL ECHOCARDIOGRAM WAS ORDERED. OVERNIGHT OXYGENATION ISSUES WERE ENCOUNTERED AND IMPROVED WITH DIURESIS AND VENT MANAGEMENT. THE PATIENT WAS STILL STATUS 7 FOR TRANSPLANT. HOWEVER, THERE WERE NOW NOTED CONCERNS FOR BACTEREMIA. THE PATIENT WAS PLACED ON ANTIBIOTICS. THE CURRENT PLAN CHANGED TO POTENTIALLY WEAN PRESSORS AS TOLERATED AND POTENTIALLY START WEANING SUPPORT OF IMPELLA AS PATIENT IS POTENTIALLY CHANGING TO BRIDGE TO RECOVERY AS OPPOSED TO HEART TRANSPLANT WITH POOR PROGNOSIS. AN EMERGENT CHOLE AT BEDSIDE THE NEXT MORNING DUE TO SLUDGE NOTED ON CT SCAN. TREATMENT FOR POSSIBLE SEPSIS RELATED TO GASTROINTESTINAL BLEED WAS NOTED. THE PATIENT WAS DELISTED FROM HEART TRANSPLANT LIST AND NOTED TO BE CRITICALLY ILL WITH A RECENT EPISODE OF PROFOUND HYPOTENSION REQUIRING CARDIOPULMONARY RESUSCITATION. CARE GOALS WERE DISCUSSED WITH THE FAMILY, THE NEXT DAY CARE WAS WITHDRAWN, AND THE PATIENT EXPIRED. Manufacturer Narrative: MEDICAL SAFETY REVIEWED THE EVENT AND DETERMINED THAT THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA 5.5 DEVICE AS AN ASSOCIATED FACTOR IN THE PATIENT¿S DEMISE OUTCOME. PRODUCT STATUS: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, AN EVALUATION/ANALYSIS OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: IMPELLA 5.5 WITH SMARTASSIST SECTION: TABLE 3.2 IMPELLA CATHETER COMPONENTS ¿THE INFUSION FILTER PREVENTS BACTERIAL CONTAMINATION AND AIR FROM ENTERING THE PURGE LUMEN.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22748799&pc=MKJ | 2112020-2025-00532 |  | Death | CARDIAC SCIENCE CORPORATION | 2025-08-08 04:00 | MKJ | ST JOHN AMB G5 AED, AUTO, ICPR PADS | Failure to Analyze Signal | No Clinical Signs, Symptoms or Conditions | P160033 |  | 1 | Event Description: COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A PATIENT (AGE AND GENDER UNKNOWN), THE DEVICE EXPERIENCED EXTENDED ANALYSIS TIME. COMPLAINANT INDICATED THAT THE PATIENT SUBSEQUENTLY EXPIRED. Manufacturer Narrative: THIS DEVICE WAS MANUFACTURED BUT NOT DOMESTICALLY DISTRIBUTED; IT IS ONLY DISTRIBUTED IN THE INTENDED GEOGRAPHY. THERE IS NO EXISTING UDI REGULATION. ZOLL MEDICAL CORPORATION HAS NOT RECEIVED THE DEVICE FOR EVALUATION AND THIS COMPLAINT IS STILL UNDER INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22748824&pc=DYE | 2025587-2025-05477 | 2022-06-30 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-08-08 04:00 | DYE | HANCOCK II AORTIC BIOPROSTHETIC HEART VALVE | Insufficient Information | Insufficient Information | P980043 |  | 1 | Event Description: MEDTRONIC RECEIVED INFORMATION THAT 10 YEARS POST IMPLANT OF THIS 23MM AORTIC BIOPROSTHETIC VALVE, THE PATIENT DIED. THE CAUSE OF DEATH WAS NOT RECEIVED. THEREFORE, RELATEDNESS OF THE DEATH TO THE VALVE COULD NOT BE EXCLUDED. IT IS UNKNOWN WHETHER AN AUTOPSY WAS PERFORMED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22749362&pc=OZD | 1220648-2025-31121 | 2025-07-25 04:00 | Death | ABIOMED, INC. | 2025-08-08 04:00 | OZD | IMPELLA 5.5 | Deformation Due to Compressive Stress | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT THE PATIENT HAD A 5.5 PUMP PLACED FOR THE PATIENT WHO WAS DETERIORATING AFTER THE INTRA-AORTIC BALLOON PUMP WAS PLACED, PATIENT WAS CARDIOVERTED AND WAS UNSTABLE WITH BLOOD PRESSURES HYPOTENSIVE AT 50S/30S. THE TEAM ESCALATED TO THE 5.5 AND EXTRACORPOREAL MEMBRANE OXYGENATION. AT THE INSERTION THE DELIVERY WAS COMPLICATED AND THE KINKING CAUSED THE PUMP TO "JUMP" AND CAUSE SUBCLAVIAN ARTERY TEARING. THE TEAM GAVE 8 UNITS OF BLOOD, VASCULAR SURGERY INTERVENED, AND THEN FAMILY MADE DECISION TO WITHDRAW CARE AFTER JUST OVER 4 HOURS OF SUPPORT. Manufacturer Narrative: THE IMPELLA DEVICE HAS NOT BEEN RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK STATES THE FOLLOWING: SECTION: WARNINGS & CAUTIONS: CAUTIONS ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH KNOWN OR SUSPECTED UNREPAIRED ABDOMINAL AORTIC ANEURYSM OR SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM OR DISSECTION OF THE ASCENDING, TRANSVERSE, OR DESCENDING AORTA.¿ CAUTION: ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22749612&pc=PZS | 3011468686-2025-00069 | 2025-01-16 05:00 | Death | MC3 INC. | 2025-08-08 04:00 | PZS | MC3 | Adverse Event Without Identified Device or Use Problem | Perforation; Septic Shock; Viral Infection; Respiratory Insufficiency | K203409 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING: EARLY USE EXPERIENCE OF THE CRESCENT RIGHT ATRIAL CANNULA FOR NEONATAL VENO-VENOUS EXTRACORPOREAL MEMBRANE OXYGENATION. THE TIME FRAME OF THIS STUDY WAS: SEPTEMBER 2021 THROUGH AUGUST 2023. THE FOLLOWING MEDTRONIC DEVICES WERE USED: THE CRESCENT RIGHT ATRIAL (RA) CANNULA. THE MAJORITY OF PATIENTS WERE CANNULATED WITH THE 13 FR CANNULA, HOWEVER THE 15 FR CANNULA WAS ALSO USED. THERE WERE NO IMMEDIATE COMPLICATIONS DURING CANNULATION. DEATHS OCCURRED IN THE STUDY POPULATION. BASED ON THE AVAILABLE INFORMATION, THESE DEATHS MAY HAVE BEEN ATTRIBUTED TO MEDTRONIC PRODUCT. THE CAUSES OF DEATH WERE: PERFORATION OF THE INFERIOR VENA CAVA (IVC) (LEADING TO HEMOPERICARDIUM, CARDIAC ARREST), DISSEMINATED HERPES SIMPLEX VIRUS (HSV), SARS-COV-2, NEONATAL RESPIRATORY DISTRESS SYNDROME, SEPTIC SHOCK, AND PULMONARY HEMORRHAGE. AMONG PATIENT ADVERSE EVENTS INCLUDED: PERFORATION, CANNULA MALPOSITION, REQUIRING SURGICAL REPOSITIONING, AND CLINICALLY SIGNIFICANT RECIRCULATION. BASED ON THE AVAILABLE INFORMATION, THESE ADVERSE EVENTS MAY HAVE BEEN ATTRIBUTED TO MEDTRONIC PRODUCT. AMONG ALL PATIENTS, DEVICE MALFUNCTIONS INCLUDED: FLOW ISSUES (DEFINED AS SUBOPTIMAL ECMO FLOW), CANNULA MALPOSITION. MEDTRONIC RECEIVED ADDITIONAL INFORMATION THAT THE ADVERSE EVENTS WERE RELATED DIRECTLY TO THE MEDTRONIC DEVICES. THERE WAS A CARDIAC PERFORATION AT THE RA/IVC JUNCTION, DUE TO THE 13 FR CRESCENT RA CANNULA. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID 70415 (UNKNOWN SERIAL/LOT); PRODUCT TYPE: 0183-CANNULA; IMPLANT DATE N/A; EXPLANT DATE N/A PRODUCT ID 70413 (UNKNOWN SERIAL/LOT); PRODUCT TYPE: 0183-CANNULA; IMPLANT DATE N/A; EXPLANT DATE N/A PRODUCT ID 70415 (UNKNOWN SERIAL/LOT); PRODUCT TYPE: 0183-CANNULA; IMPLANT DATE N/A; EXPLANT DATE N/A PRODUCT ID 70413 (UNKNOWN SERIAL/LOT); PRODUCT TYPE: 0183-CANNULA; IMPLANT DATE N/A; EXPLANT DATE N/A PRODUCT ID 70415 (UNKNOWN SERIAL/LOT); PRODUCT TYPE: 0183-CANNULA; IMPLANT DATE N/A; EXPLANT DATE N/A CITATION: MICHAEL SOBOLIC, YUJIN PARK, SHARADA H. GOWDA. EARLY USE EXPERIENCE OF THE CRESCENT RIGHT ATRIAL CANNULA FOR NEONATAL VENO-VENOUS EXTRACORPOREAL MEMBRANE OXYGENATION. ASAIO JAN 16 2025. 10.1097/MAT.0000000000002371 EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT. NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22749886&pc=PZS | 3011468686-2025-00070 | 2025-01-16 05:00 | Death | MC3 INC. | 2025-08-08 04:00 | PZS | MC3 | Adverse Event Without Identified Device or Use Problem | Perforation; Septic Shock; Viral Infection; Respiratory Insufficiency | K203409 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING: EARLY USE EXPERIENCE OF THE CRESCENT RIGHT ATRIAL CANNULA FOR NEONATAL VENO-VENOUS EXTRACORPOREAL MEMBRANE OXYGENATION. THE TIME FRAME OF THIS STUDY WAS: SEPTEMBER 2021 THROUGH AUGUST 2023. THE FOLLOWING MEDTRONIC DEVICES WERE USED: THE CRESCENT RIGHT ATRIAL (RA) CANNULA. THE MAJORITY OF PATIENTS WERE CANNULATED WITH THE 13 FR CANNULA, HOWEVER THE 15 FR CANNULA WAS ALSO USED. THERE WERE NO IMMEDIATE COMPLICATIONS DURING CANNULATION. DEATHS OCCURRED IN THE STUDY POPULATION. BASED ON THE AVAILABLE INFORMATION, THESE DEATHS MAY HAVE BEEN ATTRIBUTED TO MEDTRONIC PRODUCT. THE CAUSES OF DEATH WERE: PERFORATION OF THE INFERIOR VENA CAVA (IVC) (LEADING TO HEMOPERICARDIUM, CARDIAC ARREST), DISSEMINATED HERPES SIMPLEX VIRUS (HSV), SARS-COV-2, NEONATAL RESPIRATORY DISTRESS SYNDROME, SEPTIC SHOCK, AND PULMONARY HEMORRHAGE. AMONG PATIENT ADVERSE EVENTS INCLUDED: PERFORATION, CANNULA MALPOSITION, REQUIRING SURGICAL REPOSITIONING, AND CLINICALLY SIGNIFICANT RECIRCULATION. BASED ON THE AVAILABLE INFORMATION, THESE ADVERSE EVENTS MAY HAVE BEEN ATTRIBUTED TO MEDTRONIC PRODUCT. AMONG ALL PATIENTS, DEVICE MALFUNCTIONS INCLUDED: FLOW ISSUES (DEFINED AS SUBOPTIMAL ECMO FLOW), CANNULA MALPOSITION. MEDTRONIC RECEIVED ADDITIONAL INFORMATION THAT THE ADVERSE EVENTS WERE RELATED DIRECTLY TO THE MEDTRONIC DEVICES. THERE WAS A CARDIAC PERFORATION AT THE RA/IVC JUNCTION, DUE TO THE 13 FR CRESCENT RA CANNULA. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID 70413 (UNKNOWN SERIAL/LOT); PRODUCT TYPE: 0183-CANNULA; IMPLANT DATE N/A; EXPLANT DATE N/A PRODUCT ID 70413 (UNKNOWN SERIAL/LOT); PRODUCT TYPE: 0183-CANNULA; IMPLANT DATE N/A; EXPLANT DATE N/A PRODUCT ID 70415 (UNKNOWN SERIAL/LOT); PRODUCT TYPE: 0183-CANNULA; IMPLANT DATE N/A; EXPLANT DATE N/A PRODUCT ID 70413 (UNKNOWN SERIAL/LOT); PRODUCT TYPE: 0183-CANNULA; IMPLANT DATE N/A; EXPLANT DATE N/A PRODUCT ID 70415 (UNKNOWN SERIAL/LOT); PRODUCT TYPE: 0183-CANNULA; IMPLANT DATE N/A; EXPLANT DATE N/A CITATION: MICHAEL SOBOLIC, YUJIN PARK, SHARADA H. GOWDA. EARLY USE EXPERIENCE OF THE CRESCENT RIGHT ATRIAL CANNULA FOR NEONATAL VENO-VENOUS EXTRACORPOREAL MEMBRANE OXYGENATION. ASAIO JAN 16, 2025. 10.1097/MAT.0000000000002371 EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT. NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22749918&pc=DSQ | 2916596-2025-05269 | 2025-07-16 04:00 | Death | THORATEC CORPORATION | 2025-08-08 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS IMPLANTED ON (B)(6) 2025. THEIR POST OPERATION COURSE CONSISTED OF MULTISYSTEM ORGAN FAILURE, AND CARE WAS WITHDRAWN ON (B)(6) 2025. THE VENTRICULAR ASSIST DEVICE (VAD) WAS FUNCTIONING AS INTENDED. NO AUTOPSY WOULD BE DONE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REV. D IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RESPIRATORY FAILURE, RIGHT HEART FAILURE, RENAL FAILURE, HEPATIC DYSFUNCTION), AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22750341&pc=DSQ | 2916596-2025-05276 | 2025-08-04 04:00 | Death | THORATEC CORPORATION | 2025-08-08 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Cardiomyopathy; Tachycardia; Cardiogenic Shock | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO CARDIOGENIC SHOCK DUE TO VENTRICULAR TACHYCARDIA FROM CARDIOMYOPATHY. THE VENTRICULAR TACHYCARDIA LEADING TO DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED, THE DEVICE OPERATED AS EXPECTED AND WOULD NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) REV. D, IS CURRENTLY AVAILABLE. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING CARDIAC ARRHYTHMIA AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, LISTS ARRHYTHMIA AS A POTENTIAL LATE POST-IMPLANT COMPLICATION THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22750443&pc=OZD | 1220648-2025-45742 | 2025-07-11 04:00 | Death | ABIOMED, INC. - 1220648 | 2025-08-08 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: US COMPLAINANT REPORTED THAT THE IMPELLA 5.5 IMPLANTED FOR CARDIOGENIC SHOCK. DURING 5.5 SUPPORT IT WAS REPORTED THAT THE PATIENT HAD SIGNIFICANT BLOOD LOSS FROM JACKSON-PRATT DRAIN. RECEIVED 2 PACKED RED BLOOD CELLS, 2 FRESH FROZEN PLASMA AND 2 CRYOPRECIPITATE AND HEPARIN WAS STOPPED. BLEEDING HAD SUBSIDED WITH STABILIZATION OF HEMOGLOBIN AND HEMATOCRIT. ADDITIONALLY, IT WAS NOTED THAT THE PUMP WAS WEANED TO P-2 WITH SMALL COLLAPSED LEFT VENTRICLE VISUALIZED ON ECHOCARDIOGRAM. ACTIVE UPPER GASTROINTESTINAL BLEED WAS PRESENT. BLOOD WAS TRANSFUSING. THE PATIENT REMAINED ON SUPPORT WITH A DURATION OF 599.90 HOURS. ULTIMATELY CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA IFU: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22750582&pc=NPT | 2015691-2025-06468 | 2025-07-17 04:00 | Death | EDWARDS LIFESCIENCES | 2025-08-08 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Calcified | Dyspnea | P140031 |  | 1 | Event Description: AS REPORTED BY THE FIELD CLINICAL SPECIALIST (FCS), APPROXIMATELY 4 YEARS, 3 MONTHS, AND 10 DAYS POST TRANSFEMORAL TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) PROCEDURE WITH A 23 MM SAPIEN 3 ULTRA VALVE, THE PATIENT PRESENTED WITH PROSTHETIC VALVE STENOSIS. THE MEAN PRESSURE GRADIENT (MPG) WAS 46 MMHG, AND THE MAXIMUM VELOCITY (VMAX) WAS 4.5 M/S. THE PATIENT WAS NOTED TO BE AN INPATIENT AND WAS AWAITING REVIEW OF THE COMPUTED TOMOGRAPHY (CT). IMAGERY WAS PROVIDED FOR EVALUATION, AND CALCIUM WAS VISUALIZED ON ALL THREE LEAFLETS. SUBSEQUENTLY, ADDITIONAL INFORMATION WAS PROVIDED INDICATING THAT THE PATIENT PRESENTED WITH DYSPNEA AND CHEST PRESSURE. THE PATIENT'S TROPONIN LEVEL WAS MILDLY ELEVATED AND TRENDED DOWN. THE PATIENT ALSO APPEARED HYPERVOLEMIC AND THERE WAS POSSIBLE REOCCURRING PNEUMONIA. THE AORTIC VALVE AREA (AVA) WAS 0.46 CM2. THE PATIENT WAS PLANNED FOR A TAVR VALVE-IN-VALVE PROCEDURE; HOWEVER, APPROXIMATELY 4 YEARS, 3 MONTHS, AND 20 DAYS POST TAVR PROCEDURE WITH THE 23 MM SAPIEN 3 ULTRA VALVE, THE PATIENT PASSED AWAY AT THEIR HOME. THE CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR CORRECTION AND INCLUDES ADDITIONAL INFORMATION FROM THE INVESTIGATION. THE FOLLOWING SECTIONS OF THIS REPORT HAVE BEEN CORRECTED/UPDATED: H6 (TYPE OF INVESTIGATION, INVESTIGATION FINDINGS, INVESTIGATION CONCLUSIONS) AND H11 (ADDITIONAL NARRATIVE/DATA). THE EVENT REPORTED IS ANTICIPATED IN THE RISK MANAGEMENT DOCUMENTATION FOR TRANSCATHETER HEART VALVE PROCEDURES. A PREVIOUS INVESTIGATION INTO THIS TYPE OF EVENT IS CAPTURED IN AN EDWARDS LIFESCIENCES TECHNICAL SUMMARY AND APPLIES TO THIS COMPLAINT. ADDITIONAL ASSESSMENT OF THE FAILURE MODE IS NOT REQUIRED AT THIS TIME. THE DEVICE WAS NOT RETURNED FOR EVALUATION AS IT REMAINED IMPLANTED. AS A RESULT, NO VISUAL INSPECTION, FUNCTIONAL TESTING, OR DIMENSIONAL ANALYSIS COULD BE PERFORMED. THERE WAS IMAGERY PROVIDED FOR EVALUATION AND CALCIUM WAS VISUALIZED ON ALL THREE LEAFLETS. PER THE TECHNICAL SUMMARY, THE INSTRUCTIONS FOR USE (IFU), TRAINING MANUALS AND CURRENT RISK MITIGATIONS HAVE BEEN REVIEWED. NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. THROUGH EXTENSIVE COMPLAINT INVESTIGATIONS, STRUCTURAL VALVE DEGENERATION RELATED TO CALCIFICATION FOR THE SAPIEN 3 ULTRA VALVES HAVE NOT BEEN ASSOCIATED WITH DEVICE MALFUNCTIONS OR MANUFACTURING NONCONFORMANCES. RATHER, THE ROOT CAUSE FOR THESE EVENTS HAS HISTORICALLY BEEN DUE TO PATIENT FACTORS (AGE, DISEASE STATE, PATIENT CO-MORBIDITIES, AND/OR PHARMACOLOGICAL INTERVENTION). IN THIS CASE, THE VALVE CALCIFICATION WAS CONFIRMED BASED ON THE IMAGERY PROVIDED FOR EVALUATION. A SURGICAL INTERVENTION WAS PLANNED, BUT THE PATIENT PASSED AWAY PRIOR TO THE SURGICAL INTERVENTION. THE AVAILABLE INFORMATION SUGGESTS PATIENT FACTORS (END STAGE RENAL DISEASE) LIKELY CONTRIBUTED TO THE EVENT AS A PATIENT MEDICAL HISTORY OF END STAGE RENAL DISEASE (ESRD) ON DIALYSIS WAS REPORTED. THE PROCESS OF CALCIFICATION INVOLVES THE REACTION OF CALCIUM-CONTAINING EXTRACELLULAR FLUID WITH MEMBRANE-ASSOCIATED PHOSPHORUS, CAUSING CALCIFICATION OF THE CELLS. MANY PATIENT-RELATED FACTORS CAN CONTRIBUTE TO THE ONSET AND PROPAGATION OF CALCIFICATION, AND CONSEQUENTLY, STRUCTURAL DEGENERATION OF THE THV. THESE FACTORS INCLUDE AGE, DISEASE STATE, PATIENT CO-MORBIDITIES, AND/OR PHARMACOLOGICAL INTERVENTION, SUCH AS BUT NOT LIMITED TO RENAL FAILURE, HEMODIALYSIS, HYPERCHOLESTEROLEMIA, HYPERLIPIDEMIA, HYPOTHYROIDISM, AND DIABETES MELLITUS. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. SINCE NO EDWARDS PRODUCT DEFECTS OR LABELING DEFICIENCIES WERE IDENTIFIED, NO CORRECTIVE OR PREVENTATIVE ACTION IS REQUIRED. Manufacturer Narrative: THE INVESTIGATION IS ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22750936&pc=NPT | 2015691-2025-06474 | 2025-07-15 04:00 | Death | EDWARDS LIFESCIENCES | 2025-08-08 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE SYSTEM | Adverse Event Without Identified Device or Use Problem |  | P140031 |  | 1 | Event Description: AS REPORTED BY AN EDWARDS LIFESCIENCES AFFILIATE IN CANADA, REGARDING AN IMPLANT OF A 23MM SAPIEN 3 ULTRA RESILIA IN AORTIC POSITION BY TRANSCAROTID APPROACH. THE PATIENT PRESENTED A SEVERE ANNULAR CALCIFICATION WITH A BICUSPID SEVERE STENOTIC VALVE. THE VALVE WAS IMPLANTED SUCCESSFULLY, AND IT WAS THEN DECIDED TO POST DILATE THE VALVE AT NOMINAL VOLUME. DURING THE POST DILATION, PATIENT BLOOD PRESSURE DROPPED AND ECHO SHOWED A LARGE DISSECTION IN THE ASCENDING AORTA AND THE PATIENT PASSED AWAY. AS PER MEDICAL OPINION, THE ROOT CAUSE OF THE DISSECTION WAS THE CALCIFIED BICUSPID VALVE. Manufacturer Narrative: E1 PHONE NUMBER (B)(6). PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. ACCORDING TO THE DEVICE TRAINING MANUALS, RISK FACTORS FOR AORTIC DISSECTION, CARDIAC/AORTIC HEMATOMA, OR ANNULAR RUPTURE DURING THE TVR PROCEDURE INCLUDE SIGNIFICANT VALVE OVER SIZING, SEVERELY OBLITERATED SINUSES OF VALSALVA, PORCELAIN AORTA AND/OR PRESENCE OF BULKY CALCIFICATION, AND NARROW CALCIFIED SINOTUBULAR JUNCTION. IN ADDITION, ADVANCED AGE, FEMALE GENDER, SMALL BODY WEIGHT, AND STEROID DEPENDENCY CAN ALSO BE CONTRIBUTING FACTORS. THE SAPIEN TRANSCATHETER HEART VALVE (ALL MODELS) RELIES ON CALCIUM TO SECURELY ANCHOR IN THE LANDING ZONE. DESPITE THIS BENEFICIAL ASPECT OF CALCIUM, BULKY CALCIUM CAN INCREASE THE RISK OF CALCIFIC NODULE DISPLACEMENT INTO THE VASCULATURE, WHICH CAN LEAD TO VASCULAR INJURY. AT TIMES, THE EXTENT AND DISTRIBUTION OF CALCIUM CAN IMPAIR EASE OF DELIVERY OF THE VALVE, CORRECT POSITIONING OF THE VALVE, DEPLOYMENT OF THE VALVE, AND PROCEDURAL SUCCESS. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST/INDICATE PATIENT FACTORS AND PROCEDURAL FACTORS (SEVERE ANNULAR CALCIFICATION/POST DILATION) CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22751051&pc=NPT | 9617601-2025-00874 | 2025-07-31 04:00 | Death | MEDTRONIC MEXICO S. DE R.L. DE CV | 2025-08-08 04:00 | NPT | EVOLUT PRO PLUS VALVE | Perivalvular Leak; Insufficient Information; Patient Device Interaction Problem | Cardiac Arrest; Valvular Insufficiency/ Regurgitation | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT A TRANSCATHETER AORTIC VALVE IMPLANTATION (TAVI) PROCEDURE USING A SELF-EXPANDING/EXPANDABLE VALVE (34 MILLIMETER (MM)) WAS PERFORMED IN A PATIENT WITH A NATIVE BICUSPID VALVE, A DILATED ASCENDING AORTA, AND SIGNIFICANT VALVE CALCIFICATION. PRE-DILATION WAS CONDUCTED WITH A 26 × 40 MILLIMETER MM BALLOON, FOLLOWED BY IMPLANTATION OF THE TAVI PROSTHESIS. POST-DILATION WITH A 26 × 40 MM BALLOON WAS PERFORMED DUE TO A MODERATE PARAVALVULAR LEAK, WHICH WAS CORRECTED AFTER THE INTERVENTION. THE PATIENT INITIALLY RECOVERED AS EXPECTED AND WAS AWAKE AND RESPONSIVE UPON LEAVING THE OPERATING ROOM. ON THE FOLLOWING MORNING, IN THE INTENSIVE CARE UNIT, THE PATIENT EXPERIENCED SUDDEN CARDIOPULMONARY ARREST. CARDIOPULMONARY RESUSCITATION WAS INITIATED IMMEDIATELY, AND THE PATIENT WAS URGENTLY TRANSFERRED TO THE OPERATING ROOM FOR EXPLORATORY STERNOTOMY. DURING SURGERY, AN ASCENDING AORTIC DISSECTION WAS IDENTIFIED. THE PATIENT DID NOT TOLERATE THE SURGICAL INTERVENTION AND DIED INTRAOPERATIVELY/IMMEDIATELY POSTOPERATIVELY.Event Description: ADDITIONAL INFORMATION WAS RECEIVED INDICATING THAT THE CAUSE OF DEATH WAS THE DISSECTION IN THE ASCENDING AORTA. PER THE PHYSICIAN, THE PATIENT HAD A BICUSPID NATIVE AND DILATED ASCENDING AORTA, WHICH MAY HAVE CONTRIBUTED TO THE EVENT. PER THE PHYSICIAN, THE VALVE CAN BE EXCLUDED AS A CONTRIBUTING FACTOR TO THE PATIENT¿S DEATH. Manufacturer Narrative: UPDATED: B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID D-EVPROP34 (LOT: 0012561866); PRODUCT TYPE: 0195-HEART VALVES; PRODUCT ID L-EVPROP34 (LOT: 0012526092); PRODUCT TYPE: 0195-HEART VALVES; SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22751156&pc=FKX | 0008030665-2025-01812 | 2025-07-27 04:00 | Death | ERIKA DE REYNOSA, S.A. DE C.V. | 2025-08-08 04:00 | FKX | LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K222318 |  | 1 | Event Description: ON (B)(6) 2025, DURING A FOLLOW-UP, A PERITONEAL DIALYSIS REGISTERED NURSE [(PD)RN] REPORTED TO FRESENIUS THIS PD PATIENT ON CONTINUOUS CYCLIC PD (CCPD) THERAPY UTILIZING THE LIBERTY SELECT CYCLER WAS PLACED ON HOSPICE AND EXPIRED DURING A DIALYSIS SESSION. THERE WAS NO SPECIFIC ALLEGATION THAT THIS EVENT WAS DUE TO A DEFICIENCY OR MALFUNCTION OF ANY FRESENIUS PRODUCT(S) OR DEVICE(S) IN THE INITIAL REPORTING. UPON FOLLOW-UP WITH THE PATIENT¿S PDRN, IT WAS REPORTED THIS PATIENT EXPIRED AT HOME WHILE ON HOSPICE ON (B)(6) 2025. IT WAS EXPLAINED THAT THE PATIENT WAS PLACED ON HOSPICE ON (B)(6) 2025 FOR AN UNREPORTED REASON BUT REMAINED ON CCPD THERAPY ON THE SAME LIBERTY SELECT CYCLER FOR THERAPEUTIC COMFORT. THE PATIENT EXPIRED DURING SLEEP IN THE EVENING ON (B)(6) 2025 DURING A CCPD TREATMENT. IT WAS CONFIRMED, THOUGH THE EXACT REASON FOR THE PATIENT¿S HOSPICE PLACEMENT AND THE CAUSE OF HER SUBSEQUENT DEATH REMAINS UNKNOWN, THERE WAS NO INDICATION THESE EVENTS WERE RELATED TO PD THERAPY OR DUE TO A DEFICIENCY OR MALFUNCTION OF ANY FRESENIUS PRODUCT(S) OR DEVICE(S). Manufacturer Narrative: PLANT INVESTIGATION: THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL INVESTIGATION: A TEMPORAL RELATIONSHIP EXISTS BETWEEN CCPD THERAPY UTILIZING THE LIBERTY SELECT CYCLER AND THE PATIENT¿S DEATH. THE CAUSE OF THIS PATIENT¿S DEATH CANNOT BE DETERMINED; HOWEVER, THERE WAS NO INDICATION THE PATIENT¿S DEATH WAS DUE TO A DEFICIENCY OR MALFUNCTION OF ANY FRESENIUS PRODUCT(S) OR DEVICE(S) AS REPORTED BY A MEDICAL PROFESSIONAL. IT IS WELL KNOWN THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. THEREFORE, THE LIBERTY SELECT CYCLER CAN BE EXCLUDED AS A ROOT CAUSE OR CONTRIBUTOR TO THIS PATIENT¿S ADVERSE EVENT. BASED ON THE AVAILABLE INFORMATION, THERE IS NO SPECIFIC ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) OR PRODUCT(S) DEFICIENCY OR MALFUNCTION, CAUSED OR CONTRIBUTED TO THE PATIENT¿S ADVERSE EVENTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22751536&pc=NPT | 2015691-2025-06482 | 2025-07-17 04:00 | Death | EDWARDS LIFESCIENCES | 2025-08-08 04:00 | NPT | EDWARDS SAPIEN 3 TRANSCATHETER HEATH VALVE | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P140031 |  | 1 | Event Description: EDWARDS RECEIVED NOTIFICATION FROM OUR AFFILIATE IN AUSTRIA. AS REPORTED, THIS WAS A CASE OF AN IMPLANT OF A 29MM SAPIEN 3 IN AORTIC POSITION BY LEFT TRANSAXILLARY APPROACH. DURING VALVE DEPLOYMENT, AN ANNULAR RUPTURE WAS OBSERVED. NO ACTIONS WERE TAKEN BUT UNFORTUNATELY, THE PATIENT PASSED AWAY. AS PER MEDICAL OPINION, THE ROOT CAUSE OF THE RUPTURE WAS SEVERE ANNULAR CALCIFICATION. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. ACCORDING TO THE DEVICE TRAINING MANUALS, RISK FACTORS FOR AORTIC DISSECTION, CARDIAC/AORTIC HEMATOMA, OR ANNULAR RUPTURE DURING THE TVR PROCEDURE INCLUDE SIGNIFICANT VALVE OVER SIZING, SEVERELY OBLITERATED SINUSES OF VALSALVA, PORCELAIN AORTA AND/OR PRESENCE OF BULKY CALCIFICATION, AND NARROW CALCIFIED SINOTUBULAR JUNCTION. IN ADDITION, ADVANCED AGE, FEMALE GENDER, SMALL BODY WEIGHT, AND STEROID DEPENDENCY CAN ALSO BE CONTRIBUTING FACTORS. THE SAPIEN TRANSCATHETER HEART VALVE (ALL MODELS) RELIES CALCIUM TO SECURELY ANCHOR IN THE LANDING ZONE. DESPITE THIS BENEFICIAL ASPECT OF CALCIUM, BULKY CALCIUM CAN INCREASE THE RISK OF CALCIFIC NODULE DISPLACEMENT INTO THE VASCULATURE, WHICH CAN LEAD TO VASCULAR INJURY. AT TIMES, THE EXTENT AND DISTRIBUTION OF CALCIUM CAN IMPAIR EASE OF DELIVERY OF THE VALVE, CORRECT POSITIONING OF THE VALVE, DEPLOYMENT OF THE VALVE, AND PROCEDURAL SUCCESS. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST THAT PATIENT RELATED FACTORS (SEVERE ANNULAR CALCIFICATION) CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22751599&pc=DSQ | 2916596-2025-05306 | 2025-08-03 04:00 | Death | THORATEC CORPORATION | 2025-08-08 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY DUE TO CEREBRAL BLEED.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT HAD HIGH INTERNATIONAL NORMALIZED RATIO (INR) OF 5. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE OPERATED AS EXPECTED. NO OTHER DETAILS WERE PROVIDED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: THE PUMP WAS NOT RETURNED. A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF GASTROINTESTINAL (GI) BLEEDING COULD NOT BE CONCLUSIVELY DETERMINED. BLEEDING IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. GI BLEEDING HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF STROKE COULD NOT BE CONCLUSIVELY DETERMINED. STROKE IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. STROKE HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22751772&pc=DSQ | 2916596-2025-05230 | 2025-06-04 04:00 | Death | THORATEC CORPORATION | 2025-08-08 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. Manufacturer Narrative: SECTION E1; REPORTER EMAIL WAS NOT AVAILABLE MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE DEATH WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. DUE TO PATIENT PRIVACY LAWS THE MANAGING HOSPITAL DID NOT COMMUNICATE PATIENT OUTCOME INFORMATION. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), IS CURRENTLY AVAILABLE. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22751817&pc=DRA | 2124215-2025-55437 | 2025-07-22 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-08-08 04:00 | DRA | FARADRIVE STEERABLE SHEATH | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Cardiac Arrest; Hemothorax; Ventricular Fibrillation |  |  | 1 | Event Description: DURING A PULSED FIELD ABLATION FOR ATRIAL FIBRILLATION/RADIOFREQUENCY ABLATION FOR ATRIAL FLUTTER A FARAWAVE CATHETER AND FARADRIVE SHEATH WERE SELECTED FOR USE. VASCULAR ACCESS WAS OBTAINED AT 12:00 VIA THE RIGHT GROIN, AND A BOSTON SCIENTIFIC EP CORONARY SINUS CATHETER ALONG WITH A NON-BOSTON SCIENTIFIC ULTRASOUND CATHETER WERE INTRODUCED. A TRANSSEPTAL PUNCTURE WAS SUCCESSFULLY PERFORMED AT 12:13 USING THE VERSACROSS LARGE ACCESS SYSTEM, WHICH WAS THEN EXCHANGED FOR THE FARADRIVE SHEATH. AT 12:20, A NON-BOSTON SCIENTIFIC MAPPING CATHETER WAS ADVANCED THROUGH THE FARADRIVE SHEATH, AND ELECTROANATOMIC MAPPING WAS CONDUCTED UNTIL 12:26. THE NON-BOSTON SCIENTIFIC MAPPING CATHETER WAS THEN REMOVED, AND THE FARAWAVE PULSED FIELD ABLATION (PFA) CATHETER WAS INTRODUCED. ABLATION BEGAN AT 12:27, BEGINNING WITH THE LEFT SUPERIOR PULMONARY VEIN (LSPV) USING THE FLOWER CONFIGURATION. THIS INCLUDED FOUR FLOWER APPLICATIONS, FOLLOWED BY TWO ANCHOR AND FOUR BASKET APPLICATIONS. THE SAME SEQUENCE WAS APPLIED TO THE LEFT INFERIOR PULMONARY VEIN (LIPV). THEN THE POSTERIOR WALL WAS ABLATED WITH A TOTAL OF 12 PFA APPLICATIONS. THE PROCEDURE THEN PROGRESSED TO THE RIGHT SUPERIOR PULMONARY VEIN (RSPV), WHERE ABLATION WAS INITIATED IN THE FLOWER CONFIGURATION. AFTER THE SECOND APPLICATION IN FLOWER CONFIGURATION, THE PHYSICIAN WAS ROTATING THE FLOWER AND OBSERVED THE PATIENT WAS IN VENTRICULAR TACHYCARDIA (VT), WHICH ESCALATED INTO VENTRICULAR FIBRILLATION (VF). AN EXTERNAL DEFIBRILLATION SHOCK WAS DELIVERED AT 200 JOULES; HOWEVER, THIS WAS AN ERROR, AS THE INTENDED DOSE WAS 360 JOULES. A SECOND SHOCK AT 360 JOULES WAS ADMINISTERED WITHOUT SUCCESS, AND THE INITIATION OF CARDIOPULMONARY RESUSCITATION (CPR) STARTED. MULTIPLE ROUNDS OF EPINEPHRINE WERE ADMINISTERED. AFTER APPROXIMATELY 20 MINUTES THE ANESTHESIA NOTED THE PRESENCE OF BLOOD IN THE AIRWAY TUBE. IT WAS DETERMINED THAT THE PATIENT LUNG WAS INTERNALLY BLEEDING LIKELY DUE TO A PUNCTURE FROM THE CHEST COMPRESSIONS. THROUGHOUT THE ABLATION PROCEDURE, THE PATIENT EXPERIENCED THREE EPISODES OF SUPRAVENTRICULAR TACHYCARDIA (SVT), WHICH THE PHYSICIAN SUSPECTED TO BE ATRIOVENTRICULAR NODAL REENTRANT TACHYCARDIA (AVNRT). THE FIRST EPISODE SELF-TERMINATED, WHILE THE LAST TWO RESOLVED DURING PFA APPLICATION. THE PATIENT EXPIRED ON (B)(6) 2025. THE OFFICIAL CAUSE OF DEATH HAS NOT YET BEEN DETERMINED AND IS PENDING THE RESULTS OF A MEDICAL EXAMINER'S AUTOPSY. THE DEVICE IS NOT EXPECTED TO BE RETURNED, AS IT IS CURRENTLY BEING RETAINED BY THE FACILITY. Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC, AND AN ADDITIONAL COMPLAINT WAS CREATED FOR THE FARADRIVE SHEATH. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22751821&pc=QZI | 2124215-2025-55438 | 2025-07-22 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-08-08 04:00 | QZI | FARAWAVE | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Cardiac Arrest; Hemothorax; Ventricular Fibrillation |  |  | 1 | Event Description: DURING A PULSED FIELD ABLATION FOR ATRIAL FIBRILLATION/RADIOFREQUENCY ABLATION FOR ATRIAL FLUTTER A FARAWAVE CATHETER AND FARADRIVE SHEATH WERE SELECTED FOR USE. VASCULAR ACCESS WAS OBTAINED AT 12:00 VIA THE RIGHT GROIN, AND A BOSTON SCIENTIFIC EP CORONARY SINUS CATHETER ALONG WITH A NON-BOSTON SCIENTIFIC ULTRASOUND CATHETER WERE INTRODUCED. A TRANSSEPTAL PUNCTURE WAS SUCCESSFULLY PERFORMED AT 12:13 USING THE VERSACROSS LARGE ACCESS SYSTEM, WHICH WAS THEN EXCHANGED FOR THE FARADRIVE SHEATH. AT 12:20, A NON-BOSTON SCIENTIFIC MAPPING CATHETER WAS ADVANCED THROUGH THE FARADRIVE SHEATH, AND ELECTROANATOMIC MAPPING WAS CONDUCTED UNTIL 12:26. THE NON-BOSTON SCIENTIFIC MAPPING CATHETER WAS THEN REMOVED, AND THE FARAWAVE PULSED FIELD ABLATION (PFA) CATHETER WAS INTRODUCED. ABLATION BEGAN AT 12:27, BEGINNING WITH THE LEFT SUPERIOR PULMONARY VEIN (LSPV) USING THE FLOWER CONFIGURATION. THIS INCLUDED FOUR FLOWER APPLICATIONS, FOLLOWED BY TWO ANCHOR AND FOUR BASKET APPLICATIONS. THE SAME SEQUENCE WAS APPLIED TO THE LEFT INFERIOR PULMONARY VEIN (LIPV). THEN THE POSTERIOR WALL WAS ABLATED WITH A TOTAL OF 12 PFA APPLICATIONS. THE PROCEDURE THEN PROGRESSED TO THE RIGHT SUPERIOR PULMONARY VEIN (RSPV), WHERE ABLATION WAS INITIATED IN THE FLOWER CONFIGURATION. AFTER THE SECOND APPLICATION IN FLOWER CONFIGURATION, THE PHYSICIAN WAS ROTATING THE FLOWER AND OBSERVED THE PATIENT WAS IN VENTRICULAR TACHYCARDIA (VT), WHICH ESCALATED INTO VENTRICULAR FIBRILLATION (VF). AN EXTERNAL DEFIBRILLATION SHOCK WAS DELIVERED AT 200 JOULES; HOWEVER, THIS WAS AN ERROR, AS THE INTENDED DOSE WAS 360 JOULES. A SECOND SHOCK AT 360 JOULES WAS ADMINISTERED WITHOUT SUCCESS, AND THE INITIATION OF CARDIOPULMONARY RESUSCITATION (CPR) STARTED. MULTIPLE ROUNDS OF EPINEPHRINE WERE ADMINISTERED. AFTER APPROXIMATELY 20 MINUTES THE ANESTHESIA NOTED THE PRESENCE OF BLOOD IN THE AIRWAY TUBE. IT WAS DETERMINED THAT THE PATIENT LUNG WAS INTERNALLY BLEEDING LIKELY DUE TO A PUNCTURE FROM THE CHEST COMPRESSIONS. THROUGHOUT THE ABLATION PROCEDURE, THE PATIENT EXPERIENCED THREE EPISODES OF SUPRAVENTRICULAR TACHYCARDIA (SVT), WHICH THE PHYSICIAN SUSPECTED TO BE ATRIOVENTRICULAR NODAL REENTRANT TACHYCARDIA (AVNRT). THE FIRST EPISODE SELF-TERMINATED, WHILE THE LAST TWO RESOLVED DURING PFA APPLICATION. THE PATIENT EXPIRED ON (B)(6) 2025. THE OFFICIAL CAUSE OF DEATH HAS NOT YET BEEN DETERMINED AND IS PENDING THE RESULTS OF A MEDICAL EXAMINER'S AUTOPSY. THE DEVICE IS NOT EXPECTED TO BE RETURNED, AS IT IS CURRENTLY BEING RETAINED BY THE FACILITY. Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22752331&pc=NIQ | 2024168-2025-08281 | 2020-12-01 05:00 | Death | ABBOTT VASCULAR INC. | 2025-08-08 04:00 | NIQ | XIENCE V | Adverse Event Without Identified Device or Use Problem | Unspecified Heart Problem | P070015 |  | 1 | Event Description: IT WAS REPORTED IN A SUMMARY ARTICLE THE OF A RANDOMIZED CLINICAL TRIAL: STOPDAPT-2 (SHORT AND OPTIMAL DURATION OF DUAL ANTIPLATELET THERAPY 2). THE TRIAL SUMMARIZED COMPARISON PATIENTS WHO UNDERWENT PERCUTANEOUS CORONARY INTERVENTION (PCI) WITH UNSPECIFIED XIENCE STENT WHO HAD DUAL ANTIPLATELET THERAPY (DAPT) WITH EITHER CLOPIDOGREL GROUP OR ASPIRIN MONOTHERAPY AFTER PCI FOR 5 YEARS. THE PRIMARY ENDPOINT WAS A COMPOSITE OF CARDIOVASCULAR OUTCOMES (CARDIOVASCULAR DEATH, MYOCARDIAL INFARCTION, STROKE, OR DEFINITE STENT THROMBOSIS) OR MAJOR BLEEDING (TIMI MAJOR OR MINOR BLEEDING). CORONARY ARTERY BYPASS GRAFTING (CABG) AND CORONARY REVASCULARIZATION WERE PERFORMED. THE STUDY FOUND THAT THERE WAS A BORDERLINE ISCHEMIC BENEFIT WITH THE USE OF CLOPIDOGREL THAN ASPIRIN BEYOND 1 YEAR AFTER PCI. ADDITIONAL INFORMATION CAN BE FOUND IN THE SUMMARY ARTICLE, "CLOPIDOGREL VS ASPIRIN MONOTHERAPY BEYOND 1 YEAR AFTER PERCUTANEOUS CORONARY INTERVENTION". Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORDS FOR THIS PRODUCT COULD NOT BE REVIEWED AND A SIMILAR COMPLAINT QUERY COULD NOT BE PERFORMED BECAUSE THE PART NUMBER AND/OR LOT NUMBER WERE NOT REPORTED, AND THE DEVICE WAS NOT RETURNED. A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECTS OF DEATH AND THE RELATIONSHIP TO PRODUCT, IF ANY, CANNOT BE DETERMINED; HOWEVER, THE SUBSEQUENT SERIOUS INJURY/ ILLNESS/ IMPAIRMENT APPEARS TO BE RELATED TO THE OPERATIONAL CONTEXT OF THE PROCEDURE. THE INFORMATION RECEIVED IN THE COMPLAINT WAS OBTAINED THROUGH A PROACTIVE LITERATURE SEARCH. SPECIFIC MODEL AND LOT NUMBERS WERE NOT AVAILABLE AND CAUSALITY BETWEEN THE DOCUMENTED DEVICE USAGE AND THE PATIENT EFFECTS COULD NOT BE ESTABLISHED. FACTORS THAT MAY CONTRIBUTE TO THE PERFORMANCE OUTCOMES LISTED IN THE ARTICLE INCLUDE, BUT ARE NOT LIMITED TO, DEVICE TO PATIENT INTERACTION, PATIENT ANATOMICAL CONDITIONS, DEVICE TO USER INTERACTION AND MANUFACTURING OR MATERIAL ISSUES. DUE TO THE LIMITED INFORMATION AVAILABLE, THE EXACT CAUSE CANNOT BE DETERMINED. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN, OR LABELING; THEREFORE, NO PRODUCT-RELATED CORRECTIVE ACTION WILL BE IMPLEMENTED IN THIS CASE. ATTACHMENT ARTICLE TITLED: "CLOPIDOGREL VS ASPIRIN MONOTHERAPY BEYOND 1 YEAR AFTER PERCUTANEOUS CORONARY INTERVENTION". B2, B3, D6A: ESTIMATED DATES D4- THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. THE ADDITIONAL PATIENT EFFECTS S REPORTED IN THE ARTICLE ARE CAPTURED UNDER A SEPARATE MEDWATCH REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22754406&pc=MAF | 2024168-2025-08283 | 2025-07-17 04:00 | Death | ABBOTT VASCULAR INC. | 2025-08-08 04:00 | MAF | GRAFTMASTER | Patient-Device Incompatibility | Unspecified Heart Problem | H000001 |  | 1 | Event Description: IT WAS REPORTED THAT THE PROCEDURE WAS TO TREAT A PERFORATION IN THE LEFT MAIN CORONARY ARTERY. THE 4X16 MM GRAFTMASTER COVERED STENT WAS IMPLANTED BUT DID NOT SEAL THE PERFORATION. THE PATIENT WAS UNABLE TO BE STABILIZED AND EXPIRED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22754705&pc=BZD | 2518422-2025-048855 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-08-08 04:00 | BZD | PHILIPS CPAP DEVICE | Degraded | Cancer; Unspecified Respiratory Problem |  |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGES NOSE IRRITATION, CANCER, DEATH, INVASIVE NEUROENDOCRINE CARCINOMA OF RIGHT SINUS AND CHRONIC SINUSITIS. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22754749&pc=LOX | 2024168-2025-08287 | 2025-07-18 04:00 | Death | ABBOTT VASCULAR INC. | 2025-08-08 04:00 | LOX | NC TREK NEO¿ | Improper or Incorrect Procedure or Method | Perforation of Vessels; Cardiac Tamponade | K220634 |  | 1 | Event Description: IT WAS REPORTED THAT THE PROCEDURE WAS TO TREAT A LESION IN THE DISTAL LEFT ANTERIOR DESCENDING (DLAD). ON (B)(6) 2025, AN UNSPECIFIED STENT WAS IMPLANTED. THEN 2.5X8MM NC TREK NEO BALLOON DILATATION CATHETER (BDC) WAS ADVANCED TO THE DISTAL END OF THE STENT WITHOUT ISSUE FOR POST DILATATION. THE BALLOON WAS INFLATED ABOVE THE RATED BURST PRESSURE OF 18 THREE TIMES TO 18, 23 AND 25 ATMOSPHERES (ATM). THE BDC WAS REMOVED WITH NO IMAGING PERFORMED. ANOTHER UNSPECIFIED BALLOON WAS ADVANCED FOR POST DILATION FOR THE PROXIMAL END OF THE STENT. THEN IMAGING WAS PERFORMED WHEN A PERFORATION WAS NOTED AT THE DISTAL END OF THE IMPLANTED STENT AND CARDIAC TAMPONADE WAS OBSERVED. THEREFORE, 2 COVERED UNSPECIFIED STENTS WERE PLACED TO TREAT THE PERFORATION, A STERNOTOMY WAS PERFORMED, AN IMPELLA DEVICE WAS PLACED AND PERICARDIOCENTESIS WAS PERFORMED, WHICH RESULTED IN ADDITIONAL HOSPITALIZATION AS THIS WAS INITIALLY AN OUT PATIENT PROCEDURE. ON (B)(6) 2025, THE PATIENT DIED. IN THE PHYSICIANS OPINION, THE AGGRESSIVE INFLATION OF THE NC TREK CAUSED BOTH THE PERFORATION AND TAMPONADE WHICH LED TO THE PATIENT DEATH. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. H6: MEDICAL DEVICE PROBLEM CODE 2017: ABOVE RBP.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22755160&pc=NIK | 9614453-2025-03141 | 2025-04-25 04:00 | Death | MEDTRONIC EUROPE SARL | 2025-08-08 04:00 | NIK | COBALT¿ XT HF CRT-D MRI SURESCAN¿ | Adverse Event Without Identified Device or Use Problem | Atrial Fibrillation; Cardiogenic Shock; Heart Failure/Congestive Heart Failure | P010031 |  | 1 | Event Description: IT WAS REPORTED THAT TEN DAYS POST CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) GENERATOR REPLACEMENT. THE CAUSE OF DEATH WAS CARDIOGENIC SHOCK, ACUTE ON CHRONIC SYSTOLIC HEART FAILURE, ATRIAL FIBRILLATION (AF), AND OBESITY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22755274&pc=NPT | 9617601-2025-00885 | 2025-08-05 04:00 | Death | MEDTRONIC MEXICO S. DE R.L. DE CV | 2025-08-08 04:00 | NPT | EVOLUT PRO PLUS VALVE | Patient Device Interaction Problem | Cardiac Arrest | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) PROCEDURE, THE "BASILICA" TECHNIQUE WAS PERFORMED TO LACERATE THE VALVE LEAFLET DUE TO THE LEFT CORONARY ARTERY BEING LESS THAN 10 MILLIMETERS (MM) IN DIAMETER AND A SMALL LEFT SINUS. DURING THE IMPLANT OF THE VALVE, THE PATIENT WENT INTO CARDIAC ARREST. THE VALVE WAS NOTED TO BE IMPLANTED. SUBSEQUENTLY, CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED YET WAS UNSUCCESSFUL AND CONTRAST IMAGING REVEALED A BLOCKAGE IN THE LEFT CORONARY ARTERY. AN ATTEMPT WAS MADE TO PLACE A STENT BUT THE PATIENT ULTIMATELY DIED. Manufacturer Narrative: CONTINUATION OF D10: SECTION D INFORMATION REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER RELEVANT DEVICE(S) ARE: PRODUCT ID: D-EVPROP23-29, SERIAL/LOT #: (B)(6), UBD: 05-MAR-2027, UDI#: (B)(4) H6. THE CODES PRESENT IN SECTION H6 CORRESPOND TO MULTIPLE COMPONENTS/PRODUCTS THAT COMPRISE THIS REPORTED EVENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22755818&pc=DRT | 8030229-2025-05271 | 2025-07-11 04:00 | Death | NIHON KOHDEN CORPORATION | 2025-08-08 04:00 | DRT | ZM-531PA | Labelling, Instructions for Use or Training Problem; Protective Measures Problem | Arrhythmia; Sepsis; Asystole | K043517 |  | 1 | Event Description: THE BIOMEDICAL ENGINEER (BME) REPORTED THAT THE PATIENT PASSED AWAY WHILE BEING MONITORED ON THE ZM TRANSMITTER. IT WAS REPORTED THAT THE CENTRAL NURSE'S STATION (CNS) DID NOT ALARM FOR A BRADYCARDIA WHEN THE PATIENT DROPPED FROM 80 TO BELOW 20BPM. HOWEVER, A REVIEW OF THE DEVICE LOGS VERIFIED THAT THE DEVICE DID ALARM BRADYCARDIA AT 03:24 AND THE ALARM ENDED AT 03:31. THE LOGS ALSO REVEALED AN EXCESSIVE AMOUNT OF "ELECTRODE OFF" ERROR MESSAGES.Event Description: THE BIOMEDICAL ENGINEER (BME) REPORTED THAT THE PATIENT PASSED AWAY WHILE BEING MONITORED ON THE ZM TRANSMITTER. IT WAS REPORTED THAT THE CENTRAL NURSE'S STATION (CNS) DID NOT ALARM FOR A BRADYCARDIA WHEN THE PATIENT DROPPED FROM 80 TO BELOW 20BPM. HOWEVER, A REVIEW OF THE DEVICE LOGS VERIFIED THAT THE DEVICE DID ALARM BRADYCARDIA AT 03:24 AND THE ALARM ENDED AT 03:31. THE LOGS ALSO REVEALED AN EXCESSIVE AMOUNT OF "ELECTRODE OFF" ERROR MESSAGES. Manufacturer Narrative: DETAILS OF COMPLAINT: THE BIOMEDICAL ENGINEER (BME) REPORTED THAT THE PATIENT PASSED AWAY WHILE BEING MONITORED ON THE ZM TRANSMITTER. IT WAS REPORTED THAT THE CENTRAL NURSE'S STATION (CNS) DID NOT ALARM FOR A BRADYCARDIA WHEN THE PATIENT DROPPED FROM 80 TO BELOW 20BPM. HOWEVER, A REVIEW OF THE DEVICE LOGS VERIFIED THAT THE DEVICE DID ALARM BRADYCARDIA AT 03:24 AND THE ALARM ENDED AT 03:31. THE LOGS ALSO REVEALED AN EXCESSIVE AMOUNT OF "ELECTRODE OFF" ERROR MESSAGES. NO PATIENT HARM WAS REPORTED. INVESTIGATION SUMMARY: THE REPORTED ALARM FAILURE WAS NOT CONFIRMED. LOG ANALYSIS REVEALED THAT THE DEVICE FUNCTIONED CORRECTLY, TRIGGERING A BRADYCARDIA ALARM AT 03:24, WHICH ENDED AT 03:31. HOWEVER, THE SECONDARY ISSUE WAS IDENTIFIED AS FREQUENT "ELECTRODE OFF" ERROR MESSAGES DURING THE SAME PERIOD. THE ROOT CAUSE OF THESE ERROR MESSAGES REMAINS UNDETERMINED. NK PROVIDED THE CUSTOMER WITH THE LOG ANALYSIS RESULTS AS WELL AS EDUCATION AND RECOMMENDATIONS ON TROUBLESHOOTING THE SECONDARY ISSUE. ENSURING PROPER PLACEMENT, USING NEW, PROPERLY STORED AND PROPERLY CONNECTED ELECTRODES, IN ADDITION TO VERIFYING THE DEVICE SETTINGS CAN HELP PREVENT "ELECTRODE OFF" ERROR MESSAGES FROM BEING DISPLAYED IN THE FUTURE. NK WILL CONTINUE TO MONITOR AND TREND. ADDITIONAL DEVICE INFORMATION: D10 CONCOMITANT MEDICAL DEVICE: THE FOLLOWING DEVICE WAS USED IN CONJUNCTION WITH THE ZM TRANSMITTER: CENTRAL NURSE'S STATION (CNS): MODEL #: PU-681RA. SERIAL #: (B)(6). DEVICE MANUFACTURER DATA: 06/12/2018. UNIQUE IDENTIFIER (UDI) #: (B)(4). RETURNED TO NIHON KOHDEN: NA. ADDITIONAL INFORMATION: A1 - A6 PATIENT INFORMATION. B4 DATE OF THIS REPORT. B5 DESCRIBE EVENT OR PROBLEM. B6 RELEVANT TESTS/LABORATORY DATA. G3 DATE RECEIVED BY MANUFACTURER. G6 TYPE OF REPORT. H2 IF FOLLOW-UP, WHAT TYPE? H6 EVENT PROBLEM AND EVALUATION CODES. H11 ADDITIONAL MANUFACTURER NARRATIVE. Manufacturer Narrative: DETAILS OF COMPLAINT: THE BIOMEDICAL ENGINEER (BME) REPORTED THAT THE PATIENT PASSED AWAY WHILE BEING MONITORED ON THE ZM TRANSMITTER. IT WAS REPORTED THAT THE CENTRAL NURSE'S STATION (CNS) DID NOT ALARM FOR A BRADYCARDIA WHEN THE PATIENT DROPPED FROM 80 TO BELOW 20BPM. HOWEVER, A REVIEW OF THE DEVICE LOGS VERIFIED THAT THE DEVICE DID ALARM BRADYCARDIA AT 03:24 AND THE ALARM ENDED AT 03:31. THE LOGS ALSO REVEALED AN EXCESSIVE AMOUNT OF "ELECTRODE OFF" ERROR MESSAGES. NO PATIENT HARM WAS REPORTED. INVESTIGATION SUMMARY: THE REPORTED ALARM FAILURE WAS NOT CONFIRMED. LOG ANALYSIS REVEALED THAT THE DEVICE FUNCTIONED CORRECTLY, TRIGGERING A BRADYCARDIA ALARM AT 03:24, WHICH ENDED AT 03:31. HOWEVER, THE SECONDARY ISSUE WAS IDENTIFIED AS FREQUENT "ELECTRODE OFF" ERROR MESSAGES DURING THE SAME PERIOD. THE ROOT CAUSE OF THESE ERROR MESSAGES REMAINS UNDETERMINED. NK PROVIDED THE CUSTOMER WITH THE LOG ANALYSIS RESULTS AS WELL AS EDUCATION AND RECOMMENDATIONS ON TROUBLESHOOTING THE SECONDARY ISSUE. ENSURING PROPER PLACEMENT, USING NEW, PROPERLY STORED AND PROPERLY CONNECTED ELECTRODES, IN ADDITION TO VERIFYING THE DEVICE SETTINGS CAN HELP PREVENT "ELECTRODE OFF" ERROR MESSAGES FROM BEING DISPLAYED IN THE FUTURE. NK WILL CONTINUE TO MONITOR AND TREND. ADDITIONAL DEVICE INFORMATION: D10 CONCOMITANT MEDICAL DEVICE: THE FOLLOWING DEVICE WAS USED IN CONJUNCTION WITH THE ZM TRANSMITTER: CENTRAL NURSE'S STATION (CNS): MODEL #: PU-681RA, SERIAL #: (B)(6), DEVICE MANUFACTURER DATA: 06/12/2018, UNIQUE IDENTIFIER (UDI) #: (B)(4), RETURNED TO NIHON KOHDEN: NA.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22755985&pc=EOQ | 3021325287-2025-00017 | 2025-07-15 04:00 | Death | NOAH MEDICAL INC. | 2025-08-08 04:00 | EOQ | GALAXY SYSTEM | Adverse Event Without Identified Device or Use Problem | Bronchial Hemorrhage | K223144 |  | 1 | Event Description: A PATIENT'S DEATH WAS REPORTED 2 HOURS AFTER UNDERGOING A GALAXY SYSTEM-ASSISTED BIOPSY OF A RIGHT LOWER LOBE LESION. WITHIN FIVE TO TEN MINUTES POST-PROCEDURE, SIGNIFICANT BLEEDING WAS NOTED AND APPROXIMATELY 100 ML OF BLOOD WAS SUCTIONED IN THE POST-ANESTHESIA CARE UNIT. AN EMERGENCY BRONCHOSCOPY FAILED TO CONTROL THE BLEEDING, AND THE PATIENT CODED THREE TIMES BEFORE PASSING AWAY. THE PATIENT HAD NO KNOWN COAGULOPATHIES OR ANTICOAGULANT THERAPY BUT DID HAVE A HISTORY OF CANCER. NO DEVICE MALFUNCTIONS WERE REPORTED. Manufacturer Narrative: SYSTEM AND BRONCHOSCOPE MANUFACTURING RECORDS WERE REVIEWED AND FOUND NO ISSUES ASSOCIATED WITH THIS CASE. FAILURE ANALYSIS OF THE SCOPE COULD NOT BE PERFORMED AS THE SCOPE WAS DISCARDED. FINAL INSPECTION AND TEST RECORDS CONFIRM THAT THE SCOPE USED IN THE CASE PASSED EACH TEST. REVIEW OF THE PROCEDURAL VIDEO SHOWS THAT AT 3:26, THE SCOPE WAS ADVANCED INTO THE LUNG PARENCHYMA, PRECIPITATING BLEEDING IN THE RIGHT MIDDLE LOBE AT 4:22. IRRIGATION CONTROLLED THE INITIAL BLEED AT 5:53, BUT BY 12:35 A PARENCHYMAL TEAR WAS VISIBLE UPON SCOPE WITHDRAWAL. BLEEDING RECURRED IN THE DISTAL AIRWAY AT 27:38, MANAGED AGAIN WITH IRRIGATION AT 28:14. DESPITE PERSISTENT BLEEDING OBSERVED AT 38:50, THE USER PROCEEDED WITH A FINAL BIOPSY AT 43:03, THEN WEDGED THE SCOPE TO TAMPONADE THE BLEED BEFORE CONCLUDING THE PROCEDURE. NO USE ERRORS WERE IDENTIFIED. THE PHYSICIAN WAS UNABLE TO DETERMINE A DEFINITIVE CAUSE OF INJURY DUE TO THE VOLUME OF BLOOD AND DID NOT ATTRIBUTE THE EVENT TO THE GALAXY SYSTEM. EFFORTS TO OBTAIN AN AUTOPSY REPORT, DEFINITIVE CAUSE OF DEATH AND TO CLARIFY DNR STATUS ARE ONGOING. IN CONCLUSION, IT IS UNLIKELY THAT THE GALAXY SYSTEM CAUSED OR CONTRIBUTED TO THE INJURY OR THE PATIENT'S DEATH, THE DEATH IS LIKELY A RESULT OF THE PATIENT'S CONDITION. NO MALFUNCTIONS WERE REPORTED AND NO USE ERRORS WERE IDENTIFIED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22761410&pc=NAY | 2955842-2025-33371 | 2025-06-06 04:00 | Death | INTUITIVE SURGICAL, INC | 2025-08-08 04:00 | NAY | DAVINCI XI | Adverse Event Without Identified Device or Use Problem | Respiratory Failure | K131861 |  | 1 | Event Description: IT WAS REPORTED THAT AFTER A DA VINCI-ASSISTED CHOLECYSTECTOMY SURGICAL PROCEDURE, THE PATIENT EXPIRED. THE INTUITIVE CLINICAL SALES REPRESENTATIVE WAS PRESENT AT THE PROCEDURE. DURING THE SURGERY, THE SURGEON STATED THAT THE CHOLECYSTECTOMY WAS DIFFICULT AND LONG DUE TO THE PATIENT¿S ANATOMY. THE PROCEDURE WAS COMPLETED ROBOTICALLY WITH NO INTRAOPERATIVE COMPLICATIONS AND NO DA VINCI-RELATED ERRORS OR ISSUES. THE PATIENT WAS MEDICALLY APPROVED FOR SURGERY; HOWEVER, THE UNSPECIFIED ANATOMY CHALLENGES RESULTED IN PROLONGED ANESTHESIA TIME. THE ELDERLY PATIENT REMAINED HOSPITALIZED AND INTUBATED IN THE INTENSIVE CARE UNIT FOR AN EXTENDED PERIOD OF TIME DUE TO UNSPECIFIED PULMONARY ISSUES RELATED TO THE ANESTHESIA TIME AND OTHER UNSPECIFIED COMORBIDITIES. THE PATIENT ULTIMATELY EXPIRED ON AN UNKNOWN DATE APPROXIMATELY ONE MONTH LATER. Manufacturer Narrative: REVIEW OF THE SYSTEM LOGS FOUND NO ERRORS WERE LOGGED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM SHOWED THE SYSTEM FUNCTIONED NORMALLY WITH NO INTRAOPERATIVE EVENTS OR ISSUES FOUND. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THE PROCEDURE: ENDOSCOPE PLUS 30 DEGREE, CADIERE FORCEPS X2, PERMANENT CAUTERY HOOK, PROGRASP FORCEPS, FENESTRATED BIPOLAR FORCEPS, LARGE CLIP APPLIER X2, MARYLAND BIPOLAR FORCEPS, SUREFORM 60 STAPLER, SUCTION IRRIGATOR. A SITE HISTORY REVIEW FOUND NO COMPLAINTS HAVE BEEN REPORTED FOR ANY OF THESE PRODUCTS. DEVICE HISTORY RECORD (DHR) REVIEW FOR THE INSTRUMENTS FOUND THAT THERE WERE NO RELATED NON-CONFORMANCES DOCUMENTED. SUREFORM 60 STAPLER LOGS SHOW THE INSTRUMENT WAS INSTALLED ON THE SYSTEM 2 TIMES AND FIRED 2 RELOADS (1 BLUE, FOLLOWED BY 1 WHITE). ON INSTALL 1, MULTIPLE INCOMPLETE CLAMP ATTEMPTS AND USER ABORTED CLAMPS OCCURRED, FOLLOWED BY SUCCESSFUL CLAMPING AND FIRING WITH NO PAUSES FOR COMPRESSION. ON INSTALL 2, THE FIRST CLAMP WAS SUCCESSFUL, AND THE FIRING WAS COMPLETED WITH 2 PAUSES FOR COMPRESSION. THERE WERE NO STAPLER RELATED ERRORS IN THE LOGS. AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) REVIEW OF THE EVENT CONCLUDED THAT THE PATIENT IN THIS REPORT DIED FOLLOWING A CHOLECYSTECTOMY. ALTHOUGH THERE ARE NO ALLEGATIONS AGAINST ANY INTUITIVE SURGICAL PRODUCT OR INSTRUMENT, AND THE PATIENTS AGE AND COMORBIDITIES MAY HAVE BEEN A CONTRIBUTING FACTOR, THE SPECIFIC CAUSE OF DEATH AND THE EVENTS THAT LED TO THE DEATH ARE NOT PROVIDED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT. SECTION B2: PATIENT DATE OF DEATH IS ESTIMATED. THE EVENT OCCURRED ON (B)(6) 2025; SURGEON ESTIMATED DATE OF DEATH AS "THE BEGINNING OF (B)(6)".
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22761646&pc=FKX | 0008030665-2025-01818 | 2025-07-19 04:00 | Death | ERIKA DE REYNOSA, S.A. DE C.V. | 2025-08-08 04:00 | FKX | LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | Adverse Event Without Identified Device or Use Problem | High Blood Pressure/ Hypertension; Renal Failure; Heart Failure/Congestive Heart Failure | K222318 |  | 1 | Event Description: A PERITONEAL DIALYSIS REGISTERED NURSE (PDRN) REPORTED TO FRESENIUS CUSTOMER SERVICE THAT A PERITONEAL DIALYSIS (PD) PATIENT PASSED AWAY WHILE CONNECTED TO THEIR LIBERTY SELECT CYCLER. A PATIENT CONTACT TOLD THE PDRN THAT IT LOOKED LIKE THE PATIENT ATTEMPTED TO GET UP TO USE THE RESTROOM. ADDITIONAL INFORMATION WAS OBTAINED THROUGH FOLLOW-UP WITH THE PDRN. THE PATIENT WAS FOUND DECEASED BY THE PATIENT CONTACT ON (B)(6) 2025. THE PATIENT WAS CONNECTED TO THE CYCLER AT THE TIME. THE PATIENT CONTACT CALLED 911. EMERGENCY MEDICAL SERVICES (EMS) ARRIVED, BUT NO RESUSCITATIVE EFFORTS WERE PERFORMED AS THE PATIENT WAS ALREADY DECEASED. THE PATIENT WAS NOT TRANSPORTED TO THE HOSPITAL. THERE WERE NO REPORTED ISSUES WITH THE CYCLER OR PD THERAPY PER THE PATIENT CONTACT. IT IS UNKNOWN WHAT PHASE OR CYCLE OF TREATMENT THE PATIENT WAS IN AS THE PATIENT WAS UTILIZING FLOWSHEETS. THE CONTACT REPORTED THAT THE FLOWSHEETS WERE IN THE PATIENT¿S ROOM, AND THEY BELIEVE THAT THE EMT TOOK THEM. THE CORONER¿S OFFICE REPORTED THAT THE CAUSE OF DEATH WAS DOCUMENTED ON THE DEATH CERTIFICATE AS RENAL FAILURE, CONGESTIVE HEART FAILURE, HISTORY OF STROKE, AND HYPERTENSION. THE PDRN STATED IT IS UNKNOWN IF THE DEATH WAS RELATED TO PD THERAPY AND/OR THE USE OF ANY FRESENIUS DEVICE(S) OR PRODUCT(S). THE CYCLER WAS REPORTEDLY PICKED UP AT THE PATIENT'S PD CLINIC TO BE RETURNED TO THE MANUFACTURER. Manufacturer Narrative: THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL INVESTIGATION: THERE IS A TEMPORAL RELATIONSHIP BETWEEN PERITONEAL DIALYSIS AND THE UTILIZATION OF THE LIBERTY SELECT CYCLER AND THE PATIENT EVENT OF DEATH. HOWEVER, THERE IS NO DOCUMENTATION IN THE COMPLAINT FILE TO SHOW A CAUSAL RELATIONSHIP BETWEEN THE PATIENT DEATH AND USE OF THE CYCLER. ADDITIONALLY, THERE IS NO ALLEGATION OF A DEVICE MALFUNCTION OR DEFICIENCY REPORTED FOR THE EVENT. ALTHOUGH THE PDRN STATED IT WAS UNKNOWN IF THE DEATH WAS RELATED TO ANY FRESENIUS DEVICE(S) OR PRODUCT(S) AND/OR THE PATIENT¿S DIALYSIS THERAPY, THE PATIENT CONTACT STATED THERE WERE NO ISSUES WITH THE LIBERTY SELECT CYCLER PRIOR TO THE DEATH. FURTHERMORE, THE DEATH CERTIFICATE INDICATED THE CAUSE OF DEATH AS RENAL FAILURE, CONGESTIVE HEART FAILURE, HISTORY OF STROKE, AND HYPERTENSION. DIALYSIS PATIENTS ARE AT EXTRAORDINARILY HIGH RISK FOR DEATH WITH CARDIAC DISEASE BEING THE MAJOR CAUSE OF DEATH AND THE SINGLE, LARGEST, SPECIFIC CAUSE OF DEATH BEING ATTRIBUTED TO ARRHYTHMIC MECHANISMS OR SUDDEN CARDIAC ARREST (SCA). BASED ON THE AVAILABLE INFORMATION AND NO ALLEGATION OR EVIDENCE OF DEVICE MALFUNCTION OR DEFICIENCY, THE LIBERTY SELECT CYCLER CAN BE EXCLUDED AS THE CAUSE OF THE PATIENT¿S DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22765423&pc=FRN | 3005053560-2025-00007 | 2025-07-08 04:00 | Death | IRADIMED CORPORATION | 2025-08-08 04:00 | FRN | MRIDUIM 3860 SERIES MRI INFUSION PUMP/MONITOR SYSTEM | Use of Device Problem | Insufficient Information | K143369 |  | 1 | Event Description: IT WAS REPORTED THAT THE CUSTOMER EXPERIENCED A POSSIBLE UNDER-INFUSION EVENT. ON (B)(6) 2025, IRADIMED'S QUALITY MANAGER RECEIVED INFORMATION FROM AN IRADIMED SALES REPRESENTATIVE THAT THE DIRECTOR OF RADIOLOGY AND IMAGING AT LCMC CHILDREN'S HOSPITAL REPORTED THAT AN IRADIMED 3860 INFUSION PUMP SHUT OFF DURING AN ACTIVE INFUSION AND THAT A PATIENT CODED IN THE MRI. THE DIRECTOR FURTHER REPORTED THAT THE PATIENT WAS INITIALLY REVIVED BUT SUBSEQUENTLY PASSED AWAY. Manufacturer Narrative: THE EVENT WAS FIRST REPORTED BY THE DIRECTOR OF RADIOLOGY AND IMAGING AT THE FACILITY TO OUR LOCAL SALES REPRESENTATIVE. THE COMPLAINT WAS THEN ESCALATED INTERNALLY TO IRADIMED REGULATORY AFFAIRS. IRADIMED HAS PERFORMED DUE DILIGENCE TO GATHER MORE INFORMATION. OUR INVESTIGATION HAS INCLUDED: PHONE CALLS: ATTEMPTS WERE MADE TO CONTACT RACHEL SHERMAN, THE DIRECTOR OF RADIOLOGY AND IMAGING, ON (B)(6) 2025, BUT NO CONTACT WAS ESTABLISHED. EMAIL CORRESPONDENCE: AN EMAIL REQUESTING ADDITIONAL DETAILS WAS SENT ON (B)(6) 2025, TO THIS SAME DIRECTOR, WHICH RECEIVED NO RESPONSE. ON-SITE VISIT: ON (B)(6) 2025, IRADIMED SALES PERSONNEL MET WITH THE VICE PRESIDENT OF OPERATIONS, THE MRI MANAGER, AND THE CHIEF OF ANESTHESIA AT (B)(6) TO GATHER MORE DETAILS. HOWEVER, NO FURTHER DETAILS OF THIS INCIDENT WERE PROVIDED. THIS INVESTIGATION IS ONGOING. DESPITE SEVERAL DUE DILIGENCE ATTEMPTS TO OBTAIN THE DEVICE INFORMATION, SERIAL NUMBERS, OR HISTORY LOGS, THE CUSTOMER HAS MADE NO RESPONSE. HOWEVER, IRADIMED IS FILING THIS INITIAL REPORT BECAUSE THE EVENT INVOLVED A PATIENT DEATH. A SUPPLEMENT TO THIS REPORT WILL BE SUBMITTED UPON THE CONCLUSION OF OUR INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22766219&pc=OZD | 1220648-2025-45711 | 2025-07-13 04:00 | Death | ABIOMED, INC. | 2025-08-08 04:00 | OZD | IMPELLA CP | Positioning Failure; Restricted Flow rate | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT (UNKNOWN RACE) WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT AND DURING SUPPORT, PERSISTING POSITIONING AND LOW PUMP FLOW ISSUES WERE ENCOUNTERED. IT WAS REPORTED THAT THE TEAM HAD ATTEMPTED REPOSITIONING OF IMPELLA BUT HAD FAILED DUE TO LACK OF VOLUME, ENLARGED LEFT VENTRICLE MUSCLE, AND DIMINISHED CAVITY. APPROXIMATELY THREE DAYS LATER, THE PATIENT TRANSPORTED FOR TRANSFER AND DURING AUTOMATED IMPELLA CONTROLLER (AIC)- TO-AIC TRANSFER, IT WAS NOTED WHEN SWITCHING TO ANOTHER CONTROLLER, A MESSAGE THAT CATHETER WAS CONNECTED AND RESUME SETTINGS WAS NOT RECEIVED. THE TEAM INDICATED CONSOLES REMAINED ON DURING TRANSFER AND THE ATTEMPT TO SWITCH BACK TO THE ORIGINAL AS WELL AS START A NEW CONSOLE FOR A SWITCH WAS WITHOUT SUCCESS AS THE IMPELLA CATHETER WAS NOT RECOGNIZED. IT WAS STATED THE PATIENT HAD BEEN JUST CANNULATED FOR VA ECMO AND ON P6. THE TEAM PLANNED TO PULL THE CATHETER BACK. HOWEVER, IT APPEARS AN AIC WORKED WITH THE IMPELLA CATHETER AS THE PATIENT CONTINUED ON IMPELLA MECHANICAL CIRCULATORY SUPPORT (MCS). THEREAFTER, A FEW SUCTION ALARMS WERE ENCOUNTERED, WHICH WERE RESOLVED AFTER ADMINISTERING THE PATIENT ONE UNIT OF ALBUMIN AND PACKED RED BLOOD CELLS (PRBC'S). SUPPORT CONTINUED. IT WAS NOTED THERE WERE NO ALARMS OR ISSUES NOTED, AND SUPPORT WAS DECREASED TO P-3 FROM P-5 WITH PLANS TO STOP HEPARIN THE FOLLOWING MORNING. A COUPLE OF DAYS LATER, THERE WAS BLEEDING IN THE ABDOMEN. AN EXPLORATORY LAPAROSCOPY WAS STARTED AT BEDSIDE BUT STOPPED DUE TO THE BLEEDING. THE NEXT DAY THE PATIENT UNDERWENT EXPLORATORY BOWEL SURGERY DUE TO A HIGH CONCERN OF ACUTE BOWEL ISCHEMIA. THE ABDOMEN WAS LEFT OPEN; HOWEVER, IT WAS CLOSED THE NEXT DAY. ONGOING DISCUSSION REGARDING CARE GOALS WERE MADE WITH PLANS OF WEANING THE IMPELLA PUMP AS THE PATIENT WAS NOT A CANDIDATE FOR A DURABLE LEFT VENTRICLE ASSIST DEVICE. THE FAMILY DECIDED TO TRANSITION THE PATIENT TO COMFORT CARE, AND THE PATIENT WOULD EXPIRE AFTER A TOTAL OF 21 DAYS ON IMPELLA MCS. Manufacturer Narrative: THE PATIENT WAS ON HEPARIN FOR THE USE OF THE IMPELLA AND IT CANNOT BE DETERMINED IF THE USE OF HEPARIN IMPACTED/EXACERBATED THE OUTCOME OF THE BLEED FOR THE PATIENT. MEDICAL SAFETY REVIEWED THE EVENT AND DETERMINED THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA PUMP AS AN ASSOCIATED FACTOR IN THE PATIENT'S DEMISE OUTCOME GIVEN ITS MALFUNCTION AND RELATED ADVERSE EVENT. THE AUTOMATED IMPELLA CONTROLLERS INVOLVED IS A PRODUCT MALFUNCTION TYPE OF REPORTABLE EVENT. PRODUCT STATUS: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, AN EVALUATION/ANALYSIS OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22735542&pc=PNJ | 9612164-2025-03873 | 2025-07-12 04:00 | Death | MEDTRONIC IRELAND | 2025-08-07 04:00 | PNJ | MICRA | High Capture Threshold | Insufficient Information | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) EXHIBITED HIGH THRESHOLDS. IT WAS REPORTED THAT THE PATIENT HAD PASSED AWAY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22736237&pc=OZD | 1220648-2025-31087 | 2025-07-23 04:00 | Death | ABIOMED, INC. | 2025-08-07 04:00 | OZD | IMPELLA 5.5 | Failure to Advance; Patient-Device Incompatibility | Ventricular Fibrillation | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WAS ADMITTED IN WITH A PLANNED HIGH RISK SURGERY OF THE ROOT ENLARGEMENT/AORTIC VALVE REPLACEMENT/SEPTAL MYECTOMY. THE IMPELLA 5.5 FAILED TO DELIVER ACROSS THE VALVE AS EXPECTED. WHILE TROUBLESHOOTING THE DELIVERY, THE PATIENT DEVELOPED VENTRICULAR FIBRILLATION AND WAS SHOCKED. THE WIRE WAS REPLACED AND THE PUMP DELIVERY WAS ATTEMPTED AGAIN. THE SECOND DELIVERY ATTEMPT WAS SUCCESSFUL AND THE IMPELLA 5.5 SUPPORTED THE PATIENT FOR JUST UNDER EIGHT HOURS UNTIL CARE WAS WITHDRAWN. THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: WARNINGS & CAUTIONS: WARNINGS ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION. [ADDITION FOR 5.0, 5.5 ONLY] IN INSTANCES WHERE THE IMPELLA PUMP HAS BEEN PLACED PRIOR TO PERFORMING CARDIAC SURGERY WITH AORTIC CROSS CLAMPING AND CARDIOPLEGIC ARREST, CARE SHOULD BE TAKEN WHEN MANIPULATING THE HEART WHEN THE PUMP POSITION IS FIXED WITH APPLICATION OF THE AORTIC CROSS CLAMP ACROSS THE PUMP CATHETER.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: CAUTIONS ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22736292&pc=OZD | 1220648-2025-31088 | 2025-07-23 04:00 | Death | ABIOMED, INC. | 2025-08-07 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Tachycardia | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT A PATIENT ON IMPELLA CP SUPPORT HAD MULTIPLE EPISODES OF VENTRICULAR TACHYCARDIA (VT) OVERNIGHT WITH SIGNIFICANT DROP IN MEAN ARTERIAL PRESSURES DURING EPISODES. THE PATIENT RECEIVED AN AMIODARONE BOLUS AND REMAINED ON AN INFUSION. IT WAS FURTHER REPORTED THAT THE PATIENT EXPIRED ON SUPPORT. ADDITONAL INFORMATION WAS RECEIVED. UPON REVIEW OF THE EPISODES OF VT, THEY APPEARED TO BE SPONTANEOUS AND UNRELATED TO IMPELLA POSITION OR IMPELLA RELATED EVENTS AND COULD BE ATTRIBUTABLE TO THE PATIENT¿S UNDERLYING PATHOLOGY AND CONTINUED HYPOPERFUSION/ISCHEMIA. ECHOCARDIOGRAM WAS REVIEWED WITH BEDSIDE NURSE AND ADVISED THAT THE IMPELLA POSITION BE RE-ASSESSED TO ENSURE THAT WAS NOT THE CAUSE. IT APPEARED IN GOOD POSITION ON ECHOCARDIGRAM PER PROVIDERS. ECHOCARDIGRAM IMAGES WERE NOT THE BEST DUE TO PATIENT¿S BODY HABITUS. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22736391&pc=OZD | 1220648-2025-31089 | 2025-07-24 04:00 | Death | ABIOMED, INC. | 2025-08-07 04:00 | OZD | IMPELLA CP | Positioning Failure | Insufficient Information | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT A PATIENT HAD AN IMPELLA CP PUMP PLACED TO SUPPORT A PATIENT ADMITTED IN WITH CARDIOGENIC SHOCK/ACUTE MYOCARDIAL INFARCTION. THE PUMP WAS PLACED BUT THE PUMP CAME OUT OF POSITION AND WITH THE MISPOSITIONING, THE PUMP WAS REMOVED AND THE PATIENT WAS MEDICALLY MANAGED. THE PATIENT DID EXPIRE AFTER 41 HOURS. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22736415&pc=NIQ | 2124215-2025-54992 | 2025-07-01 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-08-07 04:00 | NIQ | SYNERGY | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT DEATHS OCCURRED. INFORMATION WAS OBTAINED AT THE 33RD ANNUAL MEETING OF THE JAPANESE SOCIETY FOR CARDIOVASCULAR INTERVENTION AND TREATMENT IN 2025. ONE-MONTH DUAL ANTIPLATELET THERAPY (DAPT) IS CONSIDERED APPROPRIATE FOR PATIENTS WITH HIGH BLEEDING RISK (HBR), BUT IT WAS UNCLEAR WHETHER A 1-MONTH DAPT IS APPLICABLE TO ALL PATIENTS WITH OR WITHOUT HBR. THE PURPOSE OF THIS STUDY IS TO EVALUATE THE SAFETY AND EFFECTIVENESS OF P2Y12 INHIBITOR MONOTHERAPY AFTER 1-MONTH DAPT. DATA FROM ONE PCI CENTER AND NINE HOSPITALS IN NORTHERN JAPAN WERE USED. THE SUBJECTS WERE 1,202 PATIENTS WHO SUCCESSFULLY UNDERWENT PERCUTANEOUS CORONARY ANGIOPLASTY (PCI) USING A BIODEGRADABLE POLYMER DRUG ELUTING STENTS (DES) (SYNERGY: 400 CASES, NON-BOSTON SCIENTIFIC (NON-BSC): 401 CASES, NON-BSC: 401 CASES). DAPT FOR 1 MONTH WAS APPLIED IN ALL CASES, FOLLOWED BY A P2Y12 INHIBITOR (PRASUGREL 3.75 MG/DAY OR CLOPIDOGREL 75 MG/DAY) MONOTHERAPY. THE PRIMARY ENDPOINT (COMBINED CARDIOVASCULAR AND BLEEDING EVENTS) OCCURRED IN 12 MONTHS, WHICH WAS SIGNIFICANTLY BELOW THE PRE-SPECIFIED NON-INFERIORITY TARGET VALUE. THERE WAS ONLY ONE CASE OF STENT THROMBOSIS, WHICH OCCURRED ON THE 258TH DAY. THERE WERE 6 CASES OF STROKE AND 22 CASES OF DEATH. IT WAS SUGGESTED THAT THE TREATMENT STRATEGY OF PERFORMING P2Y12 INHIBITOR MONOTHERAPY AFTER 1-MONTH DAPT IS SAFE FOR BOTH HBR PATIENTS AND NON-HBR PATIENTS IN DAILY MEDICAL CARE. IT COULD NOT BE COMPARED DIRECTLY WITH THE LONG-TERM DAPT BECAUSE IT IS A SINGLE-GROUP DESIGN, BUT THE CLINICAL SIGNIFICANCE OF THIS STRATEGY WAS SHOWN. Manufacturer Narrative: B3 DATE OF EVENT: ESTIMATED BASED ON AWARE DATE AS THE ACTUAL DATE WAS NOT PROVIDED. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22736526&pc=OZD | 1220648-2025-31090 | 2025-07-24 04:00 | Death | ABIOMED, INC. | 2025-08-07 04:00 | OZD | IMPELLA 5.5 | Pumping Stopped | Insufficient Information | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT A PATIENT HAD AN IMPELLA 5.5 PUMP PLACED FOR ESCALATED CHEST PAIN AND EXTRACORPOREAL MEMBRANE OXYGENATION. THE IMPELLA 5.5 WAS SUPPORTING AND ALLOWED FOR EXTRACORPOREAL MEMBRANE OXYGENATION TO BE REMOVED AND PLAN FOR THE LEFT VENTRICULAR ASSIST DEVICE IN THE FUTURE. THE IMPELLA 5.5 WAS SUPPORTING FOR WEEKS WHEN THERE WERE COMPLICATIONS OF DEEP VEIN THROMBOSIS AND INFECTION AS WELL AS NEUROLOGICAL CHANGES. THE PUMP REMAINED ON DESPITE THESE ISSUES UNTIL DAY 22. THE PUMP STOPPED AND WAS NOT ABLE TO BE RESTARTED. THE STOPPED PUMP ALSO HAD MULTIPLE ALARMS. THE FAMILY KNEW THE PATIENT WAS NOT A CANDIDATE TO ESCALATE TO OTHER MEASURES AND SO THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA PUMP AND DATA STICK WERE RETURNED, AND AN EVALUATION IS UNDERWAY. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: DIRECT AORTIC INSERTION ¿MAINTAINING ACT AT OR ABOVE 250 SECONDS WILL HELP PREVENT A THROMBUS FROM ENTERING THE CATHETER AND CAUSING A SUDDEN STOP ON STARTUP.¿ SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA CATHETER ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP. CLINICIANS SHOULD MONITOR MOTOR CURRENT AND CONSIDER REPLACING THE IMPELLA CATHETER WHENEVER A RISE IN MOTOR CURRENT IS SEEN.¿ SECTION: WARNINGS, CAUTIONS, AND PRECAUTIONS ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22737377&pc=OZD | 1220648-2025-31095 | 2025-07-24 04:00 | Death | ABIOMED, INC. | 2025-08-07 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Ischemia | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT THE PATIENT WITH ACUTE MYOCARDIAL INFARCTION COMPLICATED BY CARDIOGENIC SHOCK WAS ESCALATED FROM IMPELLA CP TO IMPELLA 5.5. IN INTENSIVE CARE UNIT THE PATIENT HAD NO PULSE ON THE RIGHT FOOT. THE MEDICAL DOCTOR REQUESTED AN ARTERIAL DOPPLER FOR RIGHT LEG AND FOOT. THE NEXT DAY, THERE WAS STILL NO PULSE IN THE PATIENT¿S RIGHT FOOT. VASCULAR WAS NOTIFIED. UNKNOWN INTERVENTION. IT WAS FURTHER REPORTED THAT THE PATIENT DEVELOPED ISCHEMIC BOWEL AND THE FAMILY DECIDED TO WITHDRAW CARE. THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE HAS NOT BEEN RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE IFU STATES: IMPELLA CP WITH SMARTASSIST SYSTEM. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22737612&pc=LWS | 2017865-2025-97038 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-08-07 04:00 | LWS | ABBOTT ICD | Failure to Deliver Shock/Stimulation; Failure to Sense | Arrhythmia; Low Blood Pressure/ Hypotension |  |  | 1 | Event Description: VOLUNTARY MEDWATCH RECEIVED STATED THAT THE PATIENT PRESENTED IN VENTRICULAR FIBRILLATION (VF). IN HOSPITAL AN INTERROGATION OF THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) REVEALED THAT THE DEVICE FAILED TO SENSE EPISODES OF VENTRICULAR TACHYCARDIA (VT) AND VF; AND SUBSEQUENTLY FAILED TO DELIVER HIGH VOLTAGE (HV) THERAPY. AS A RESULT, THE PATIENT WAS TAKEN TO SURGERY FOR IMPLANT OF TEMPORARY PACING LEADS. DURING THE PROCEDURE THE PATIENT CODED, BECAME HYPOTENSIVE AND SUBSEQUENTLY EXPIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22737794&pc=DSQ | 3007042319-2025-01888 | 2025-07-03 04:00 | Death | HEARTWARE, INC. | 2025-08-07 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WITH A MEDICAL HISTORY OF CONGESTIVE HEART FAILURE WHO HAD A VENTRICULAR ASSIST DEVICE IMPLANTED PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. Manufacturer Narrative: A SUPPLEMENTAL IS BEING SUBMITTED FOR COMPLETION OF THE INVESTIGATION. PRODUCT EVENT SUMMARY: (B)(6) WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. OF NOTE, IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO UNKNOWN CAUSES. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, DEATH IS A KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22737852&pc=OZD | 1220648-2025-31097 | 2025-07-24 04:00 | Death | ABIOMED, INC. | 2025-08-07 04:00 | OZD | IMPELLA CP | Positioning Failure | Hemolysis; Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT THE PATIENT EXPERIENCED ACCESS SITE BLEEDING A COUPLE OF HOURS AFTER BEING IN THE INTENSIVE CARE UNIT. MANUAL PRESSURE WAS HELD, AND HEPARIN WAS STOPPED. ADDITIONALLY, THE PATIENT DEVELOPED RED URINE. THE IMPELLA CP WAS TURNED DOWN TO P7 FROM P8. UNDER ECHOCARDIOGRAM IT LOOKED LIKE THE IMPELLA WAS TOO DEEP, SO IT WAS REPOSITIONED. IT WAS FURTHER REPORTED THAT CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE HAS NOT BEEN RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMART ASSIST SYSTEM STATES THE FOLLOWING: SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22737885&pc=LKN | 1722028-2025-00173 | 2025-07-14 04:00 | Death | TERUMO BCT | 2025-08-07 04:00 | LKN | SPECTRA OPTIA | Insufficient Information | Local Reaction | K183081 |  | 1 | Event Description: THE CUSTOMER REPORTED A PATIENT DEATH AFTER A RED BLOOD CELL EXCHANGE (RBCX) PROCEDURE. PER CUSTOMER "THE PATIENT WAS RESTLESS AND ALL OVER THE BED AND ALSO HAD A DIAGNOSIS OF DEMENTIA. THE PATIENT WAS INTUBATED AND ON 3 PRESSORS AND WAS IN THE MEDICAL ICU." THE CUSTOMER DOES NOT SUSPECT THE DEVICE CAUSED THE DEATH AND INDICATED THEY BELIEVED IT COULD HAVE BEEN A TRANSFUSION REACTION. NO MEDICATIONS GIVEN AS PART OF THE TREATMENT. THE COLLECTION SET IS NOT AVAILABLE FOR RETURN BECAUSE IT WAS DISCARDED BY THE CUSTOMER. THIS REPORT IS BEING FILED DUE TO PATIENT DEATH, ALTHOUGH PER CURRENT INFORMATION THERE IS NO DETECTABLE MALFUNCTION WITH THE TERUMO BCT DEVICE OR ALLEGATION OF A MALFUNCTION. Manufacturer Narrative: INVESTIGATION: THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. REVIEW OF THE RUN DATA FILE SHOWED THAT THE SYSTEM TO BE OPERATING AS INTENDED WITH NO SIGNALS OR ALARMS TO INDICATE UNEXPECTED DEVICE BEHAVIOR. SIGNALS IN THE RUN DATA FILE DO NOT SHOW ANY ISSUES WITH THE DEVICE. ALL SAFETY SYSTEMS REMAINED IN PLACE AND THE CORRECT ALARMS WERE RAISED. THE DEVICE WAS FOUND TO BE PERFORMING AS INTENDED. INVESTIGATION IS IN PROCESS, A FOLLOW-UP REPORT WILL BE PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22737909&pc=OZD | 1220648-2025-31098 | 2024-06-01 04:00 | Death | ABIOMED, INC. | 2025-08-07 04:00 | OZD | IMPELLA CP | Malposition of Device | Arrhythmia | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT IMPLANTED WITH AN IMPELLA CP EXPERIENCED MULTIPLE ¿PLACEMENT SIGNAL LOW¿ ALARMS. AN ECHOCARDIOGRAM SHOWED THE IMPELLA WAS POSITIONED DEEP SO THE PUMP WAS REPOSITIONED. THE PATIENT EXPERIENCED VENTRICULAR TACHYCARDIA REQUIRING CARDOVERSION AND INCREASED VASOACTIVE REQUIREMENTS. CARE WAS ULTIMATELY WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: A4 AND A5, ETHNICITY, ARE UNKNOWN. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMARTASSIST SYSTEM LISTS VENTRICULAR ARRYTHMIA AS AN ADVERSE EVENT THAT OCCURRED WHILE IN USE WITH THE IMPELLA DEVICE. IT STATES THAT "ONLY A FRACTION OF THESE RATES WERE ATTRIBUTED TO THE IMPELLA DEVICE AND THE EVENTS RESOLVED WITHOUT RESIDUAL EFFECT IN MOST OF THE CASES, UNLESS THE EVENT OF DEATH OCCURRED. OVERALL, THE RESULTS DID NOT SHOW ANY EVIDENCE OF INCREASED MORBIDITY ASSOCIATED WITH THE IMPELLA SUPPORT IN CARDIOMYOPATHY, MYOCARDITIS, AND PERIPARTUM CARDIOMYOPATHY (PPCM) PATIENTS."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22738004&pc=NVZ | 2182208-2025-03872 | 2025-03-07 05:00 | Death | MEDTRONIC, INC. | 2025-08-07 04:00 | NVZ | MEDTRONIC IPG | Adverse Event Without Identified Device or Use Problem | Wound Dehiscence; Pulmonary Embolism; Bacterial Infection; Erosion; Endocarditis; Fever; Hemorrhage/Bleeding; Intracranial Hemorrhage; Sepsis; Fluid Discharge; Pericardial Effusion; Swelling/ Edema |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING TRANSVENOUS LEAD EXTRACTION. THE AUTHORS DESCRIBED FIFTY-ONE PATIENTS WHO UNDERWENT LEAD AND IMPLANTABLE PULSE GENERATOR (IPG) EXTRACTIONS DUE TO INFECTION. CLINICAL PRESENTATIONS OF INFECTION INCLUDED FEVER, LOCAL SWELLING, LOCAL DISCHARGE, DEVICE EXTRUSION, CUT LEAD EXTRUSION, AND LEAD ENDOCARDITIS. UPPER LIMB VENOGRAPHY SHOWED COMPLETE OBSTRUCTION OF AXILLARY/SUBCLAVIAN VEIN IN SOME PATIENTS ON THE IPSILATERAL SIDE AND SOME ON THE CONTRALATERAL SIDE. THERE WERE SIXTEEN PATIENTS WITH POSITIVE CULTURES, AND NO ORGANISM WAS ISOLATED IN THIRTY-FIVE CASES. ORGANISMS ISOLATED FROM THE POSITIVE CULTURES WERE STAPHYLOCOCCUS AUREUS, PSEUDOMONAS, STREPTOCOCCUS BOVIS, COAGULASE-NEGATIVE STAPHYLOCOCCI (CONS), AND ENTEROCOCCUS. SEVEN PATIENTS HAD COMPLICATIONS FROM THE EXTRACTION PROCEDURE; FOUR PATIENTS HAD LOCAL BLEED WHICH REQUIRED ELECTROCAUTERY AND PRESSURE DRESSING, AND THREE PATIENTS EXPERIENCED PERICARDIAL EFFUSION WITHOUT THE NEED FOR PERICARDIOCENTESIS. THERE WERE THREE PATIENT DEATHS; ONE PATIENT DIED DUE TO INTRACRANIAL HEMORRHAGE, ONE DIED DUE TO PULMONARY THROMBOEMBOLISM, AND ONE DIED DUE TO SEPSIS WITH COLON CANCER. THE STATUS OF THE DEVICES AND LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/63 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: TRANSVENOUS LEAD EXTRACTION: EXPERIENCE FROM A NORTHERN STATE OF INDIA - THE SRINAGAR EXTRACTION REGISTRY. INDIAN PACING AND ELECTROPHYSIOLOGY JOURNAL. 2025. 25; 58¿66. DOI: 10.1016/J.IPEJ.2025.03.002. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22738030&pc=DXY | 2182208-2025-03873 | 2025-03-07 05:00 | Death | MEDTRONIC, INC. | 2025-08-07 04:00 | DXY | MEDTRONIC LEAD | Adverse Event Without Identified Device or Use Problem | Wound Dehiscence; Pulmonary Embolism; Bacterial Infection; Erosion; Endocarditis; Fever; Hemorrhage/Bleeding; Intracranial Hemorrhage; Sepsis; Fluid Discharge; Pericardial Effusion; Swelling/ Edema |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING TRANSVENOUS LEAD EXTRACTION. THE AUTHORS DESCRIBED FIFTY-ONE PATIENTS WHO UNDERWENT LEAD AND IMPLANTABLE PULSE GENERATOR (IPG) EXTRACTIONS DUE TO INFECTION. CLINICAL PRESENTATIONS OF INFECTION INCLUDED FEVER, LOCAL SWELLING, LOCAL DISCHARGE, DEVICE EXTRUSION, CUT LEAD EXTRUSION, AND LEAD ENDOCARDITIS. UPPER LIMB VENOGRAPHY SHOWED COMPLETE OBSTRUCTION OF AXILLARY/SUBCLAVIAN VEIN IN SOME PATIENTS ON THE IPSILATERAL SIDE AND SOME ON THE CONTRALATERAL SIDE. THERE WERE SIXTEEN PATIENTS WITH POSITIVE CULTURES, AND NO ORGANISM WAS ISOLATED IN THIRTY-FIVE CASES. ORGANISMS ISOLATED FROM THE POSITIVE CULTURES WERE STAPHYLOCOCCUS AUREUS, PSEUDOMONAS, STREPTOCOCCUS BOVIS, COAGULASE-NEGATIVE STAPHYLOCOCCI (CONS), AND ENTEROCOCCUS. SEVEN PATIENTS HAD COMPLICATIONS FROM THE EXTRACTION PROCEDURE; FOUR PATIENTS HAD LOCAL BLEED WHICH REQUIRED ELECTROCAUTERY AND PRESSURE DRESSING, AND THREE PATIENTS EXPERIENCED PERICARDIAL EFFUSION WITHOUT THE NEED FOR PERICARDIOCENTESIS. THERE WERE THREE PATIENT DEATHS; ONE PATIENT DIED DUE TO INTRACRANIAL HEMORRHAGE, ONE DIED DUE TO PULMONARY THROMBOEMBOLISM, AND ONE DIED DUE TO SEPSIS WITH COLON CANCER. THE STATUS OF THE DEVICES AND LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/63 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: TRANSVENOUS LEAD EXTRACTION: EXPERIENCE FROM A NORTHERN STATE OF INDIA - THE SRINAGAR EXTRACTION REGISTRY. INDIAN PACING AND ELECTROPHYSIOLOGY JOURNAL. 2025. 25; 58¿66. DOI: 10.1016/J.IPEJ.2025.03.002. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22738143&pc=OZD | 1220648-2025-45729 | 2024-01-26 05:00 | Death | ABIOMED, INC. | 2025-08-07 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Low Blood Pressure/ Hypotension | P140003 |  | 1 | Event Description: A NOTIFICATION VIA ABIOMED¿S LONG-TERM OUTCOME AND QUALITY INDICATOR (LOQI) IMPELLA REGISTRY WAS RECEIVED REGARDING A PATIENT THAT ENDURED DEATH WITH A "UNKNOWN (UNDETERMINED)" RELATIONSHIP TO THE IMPELLA DEVICE USED. IT WAS FOUND THAT DURING SUPPORT, THE PATIENT DEVELOPED SUSTAINED VENTRICULAR TACHYCARDIA (VT) AND WAS CARDIOVERTED MULTIPLE TIMES WITH DEFIBRILLATION SHOCK, HOWEVER, THE PATIENT KEPT CONVERTING BACK TO VT. LIDOCAINE DRIP WAS STARTED FOR ARRHYTHMIA. DESPITE MORE STABLE RHYTHM, PATIENT REMAINED HYPOTENSIVE. EVENTUALLY PATIENT CARE WAS WITHDRAWN AS WAS NOT A CANDIDATE FOR ANY OTHER ADVANCE THERAPIES. PATIENT PASSED AWAY. PER THE REGISTRY, THE PATIENT WAS DISCHARGED HOME POST PROCEDURE AND EXPIRED AT HOME. Manufacturer Narrative: ADDITIONAL INFORMATION IS BEING REQUESTED FOR CLARIFICATION REGARDING THE LOCATION AND DATE OF DEATH. SHOULD ANY CLARIFYING OR NEW RELEVANT INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE SUBMITTED ACCORDINGLY. PRODUT STATUS: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE AS NOTED IN THE IMPELLA: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22738367&pc=KNS | 3005099803-2025-03893 | 2025-07-16 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-08-07 04:00 | KNS | AXIOS? | Difficult to Advance; Positioning Problem | Abdominal Pain; Abscess; Fever; Hematoma; Hemorrhage/Bleeding; Unspecified Infection; Perforation; Insufficient Information | K233318 |  | 1 | Event Description: IT WAS REPORTED TO BOSTON SCIENTIFIC CORPORATION THAT AN AXIOS STENT AND ELECTROCAUTERY-ENHANCED DELIVERY SYSTEM WAS IMPLANTED TO TREAT A GASTRIC BILIARY OBSTRUCTION DURING A PROCEDURE PERFORMED ON (B)(6) 2025. DURING THE PROCEDURE, A TYPE 2 MISPLACEMENT OF THE AXIOS STENT WAS IDENTIFIED. A STENT-IN-STENT PROCEDURE WAS THEN PERFORMED TO CORRECT STENT PLACEMENT. HOWEVER, TWO HOURS POST STENT-IN STENT PROCEDURE, THE PATIENT EXPERIENCED ABDOMINAL PAIN. THE PATIENT WAS SENT TO SURGERY, AND A PERFORATION WAS NOTED. THE STENT REMAINS IMPLANTED. PROPHYLACTIC ANTIBIOTICS WERE ADMINISTERED, AND THE PATIENT WAS HOSPITALIZED. Manufacturer Narrative: BLOCK E1: INITIAL REPORTER'S SECOND PHONE NUMBER IS (B)(6); REPORTED HERE AS PHONE NUMBER EXCEEDED CHARACTER LIMIT FOR DESIGNATED FIELD. BLOCK H6: IMDRF DEVICE CODE A1502 CAPTURES THE REPORTABLE EVENT OF STENT POSITIONING ISSUE. IMDRF PATIENT CODE E2114 CAPTURES THE REPORTABLE EVENT OF PERFORATION. IMDRF PATIENT CODE E1002 CAPTURES THE REPORTABLE EVENT OF ABDOMINAL PAIN. IMDRF IMPACT CODE F08 CAPTURES THE REPORTABLE EVENT OF PATIENT WAS HOSPITALIZED. IMDRF IMPACT CODE F19 CAPTURES THE REPORTABLE EVENT OF PATIENT WAS SENT TO SURGERY. IMDRF IMPACT CODE F2301 CAPTURES THE REPORTABLE EVENT OF STENT-IN-STENT PROCEDURE. IMDRF IMPACT CODE F2303 CAPTURES THE REPORTABLE EVENT OF THE PROPHYLACTIC ANTIBIOTICS WERE ADMINISTERED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22739217&pc=MHX | 9610816-2025-000715 | 2025-07-15 04:00 | Death | PHILIPS MEDIZIN SYSTEME BÖBLINGEN GMBH | 2025-08-07 04:00 | MHX | INTELLIVUE MP5 | Insufficient Information | Insufficient Information | K150310 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT'S CONDITION DETERIORATED AND THEY SUBSEQUENTLY PASSED AWAY. THE CUSTOMER INDICATED THIS EVENT WAS RELATED TO THE STAFF NOT PAYING ATTENTION TO ALARMS. BASED ON THIS INFORMATION, THERE WAS NO ALLEGATION OF MALFUNCTION WHICH CAUSED OR CONTRIBUTED TO THE REPORTED DEATH; HOWEVER, A USE ERROR IN NOT RESPONDING TO ALARMS WAS A FACTOR. Manufacturer Narrative: PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22739614&pc=NPT | 9617601-2025-00853 | 2025-08-01 04:00 | Death | MEDTRONIC MEXICO S. DE R.L. DE CV | 2025-08-07 04:00 | NPT | EVOLUT FX PLUS VALVE | Insufficient Information; Central Regurgitation | Pulmonary Edema; Aortic Valve Insufficiency/ Regurgitation | P130021 |  | 1 | Event Description: IT WAS REPORTED DURING A TRANSCATHETER AORTIC BIOPROSTHETIC VALVE REPLACEMENT PROCEDURE, IN A PATIENT WITH A NATIVE ANNULUS PERIMETER MEASURING 102 MM, A 32 MM NON-MEDTRONIC (MERIL MYVAL) VALVE WAS IMPLANTED. FOLLOWING IMPLANT, SEVERE AORTIC INSUFFICIENCY WAS OBSERVED. A DECISION WAS MADE TO IMPLANT A 32MM MEDTRONIC EVFXPLUS-34 VALVE. THE MEDTRONIC VALVE WAS SUCCESSFULLY IMPLANTED VALVE-IN-VALVE. HOWEVER, THE PATIENT DEVELOPED PULMONARY EDEMA AND COULD NOT RECOVER FROM THE LONG DURATION OF AORTIC INSUFFICIENCY WHICH LED TO EXITUS.Event Description: IT WAS REPORTED DURING A TRANSCATHETER AORTIC BIOPROSTHETIC VALVE REPLACEMENT PROCEDURE, IN A PATIENT WITH A NATIVE ANNULUS PERIMETER MEASURING 102MM, A 32MM NON-MEDTRONIC (MERIL MYVAL) VALVE WAS IMPLANTED. FOLLOWING IMPLANT, SEVERE AORTIC INSUFFICIENCY WAS OBSERVED. A DECISION WAS MADE TO IMPLANT A 32MM MEDTRONIC EVFXPLUS-34 VALVE. THE MEDTRONIC VALVE WAS SUCCESSFULLY IMPLANTED VALVE-IN-VALVE. HOWEVER, THE PATIENT DEVELOPED PULMONARY EDEMA AND COULD NOT RECOVER FROM THE LONG DURATION OF AORTIC INSUFFICIENCY WHICH LED TO EXITUS. ADDITIONAL INFORMATION WAS RECEIVED TO CLARIFY THE AORTIC INSUFFICIENCY WAS CENTRAL REGURGITATION AND THE MEDTRONIC VALVE RESOLVED THE REGURGITATION. HOWEVER, THE MEDTRONIC VALVE WAS NOT PLACED IN TIME; THE PATIENT ENDURED SEVERE CENTRAL AORTIC REGURGITATION FOR TOO LONG AND COULD NOT RECOVER. PER THE PHYSICIAN, THE MEDTRONIC CANNOT BE EXCLUDED AS A CONTRIBUTING CAUSE TO THE AORTIC INSUFFICIENCY, BUT CAN BE EXCLUDED AS A CONTRIBUTING FACTOR TO THE PULMONARY EDEMA AND DEATH. THE PHYSICIAN FURTHER NOTED THE PATIENT'S UNDERLYING CONDITION OF AORTIC INSUFFICIENCY FROM THE NON-MEDTRONIC (MERIL MYVAL) VALVE WAS THE MAIN CAUSE OF DEATH.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE MEDTRONIC VALVE COULD BE EXCLUDED AS A CONTRIBUTING CAUSE TO THE AORTIC INSUFFICIENCY. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED B.5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. A SECOND PARAGRAPH WAS ADDED. H6. ADDITIONAL CODES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22739776&pc=DSQ | 3007042319-2025-01891 | 2025-08-05 04:00 | Death | HEARTWARE, INC. | 2025-08-07 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Pumping Problem | Pneumonia; Sepsis; Low Oxygen Saturation; Abdominal Distention | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS HOSPITALIZED FOR SEPSIS CAUSED BY PNEUMONIA. ABNORMALITIES WERE NOTED IN THE INTESTINAL TRACT AND ABDOMINAL DISTENSION WAS OBSERVED. DETAILED DIAGNOSIS FOR THE INTESTINAL LESIONS WAS NOT PROVIDED, HOWEVER THEY ARE SUSPECTED TO BE INFECTION RELATED. REVIEW OF LOG FILES DATA REVEALED MULTIPLE LOW FLOW ALARMS. THE PATIENT WAS TREATED WITH ANTIBIOTICS AND FLUID REPLACEMENT THERAPY. A DECREASE IN OXYGEN SATURATION WAS OBSERVED. IT WAS NOTED THAT THE PATIENT HAD A DO NOT RESUSCITATE ORDER, SO THE PATIENT RECEIVED OXYGEN VIA MASK. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN USE.Event Description: IT WAS FURTHER REPORTED THAT THE VAD EXHIBITED BELOW NORMAL POWER AND THAT THE PATIENT DIED DUE TO THE SEPSIS HOSPITALIZATION. THE VAD WAS EXPLANTED. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED AS ADDITIONAL INFORMATION HAS BEING RECEIVED FOR THIS EVENT AND SECTIONS HAVE BEEN UPDATED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR INVESTIGATION SUMMARY. PRODUCT EVENT SUMMARY: THE VENTRICULAR ASSIST DEVICE (VAD) (B)(6) WAS NOT RETURNED FOR EVALUATION. A REVIEW OF THE MANUFACTURING DOCUMENTATION WAS NOT PER FORMED AS THE REPORTED EVENT AND ANALYSIS ARE NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. REVIEW OF THE AVAILABLE AUTOLOGS REPORTS REVEALED INTERMITTENT DECREASES IN POWER CONSUMPTION AND ESTIMATED FLOWS, FOLLOWED BY A SHARP DECREASE IN POWER CONSUMPTION AND ESTIMATED FLOWS, THROUGHOUT THE ANALYZED PERIOD, LEADING TO PARAMETERS BELOW NORMAL OPERATING RANGE. FOUR (4) LOW FLOW ALARMS WERE LOGGED SINCE (B)(6) 2025. ADDITIONALLY, REVIEW OF AUTOLOGS REVEALED TWO (2) MOTOR START EVENTS RECORDED ON (B)(6) 2021 AT 15:45:53 AND (B)(6) 2023 AT 09:14:45, IN WHICH PARAMETERS WERE ATYPICAL. THIS IS AN ADDITIONAL FINDING, NOT RELATED TO THE REPORTED EVENT. AS A RESULT, THE REPORTED EVENT WAS CONFIRMED. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. BASED ON THE RISK DOCUMENTATION, POSSIBLE CAUSES OF THE REPORTED EVENT MAY BE ATTRIBUTED TO MULTIPLE FACTORS INCLUDING BUT NOT LIMITED TO THROMBUS AT THE INFLOW CANNULA/OUTFLOW GRAFT, CONSTRICTION AT THE OUTFLOW GRAFT, POOR VAD FILLING, AND/OR INAPPROPRIATE PUMP ROTATIONAL SPEED. BASED ON RISK DOCUMENTATION, THE MOST LIKELY ROOT CAUSE OF ATYPICAL MOTOR START PARAMETERS MAY BE ATTRIBUTED, BUT NOT LIMITED, TO OUTER SHROUD CONTACT THAT CREATES MORE FRICTION AT THE HOUSING TO IMPELLER INTERFACE DURING PUMP STARTUP. PER THE INSTRUCTIONS FOR USE, INFECTION AND DEATH ARE KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. BASED ON REVIEW OF PAST ADVERSE EVENTS FOR THIS PATIENT, IT WAS NOTED THAT THE PATIENT HAD A HISTORY OF INFECTION EVENTS. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS, AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22740079&pc=MIH | MW5174189 | 2025-08-02 04:00 | Death | ENDOLOGIX LLC. | 2025-08-07 04:00 | MIH | ENDOLOGIX AFX GRAFT | Leak/Splash | Insufficient Information |  |  | 1 | Event Description: LEAK WITHIN AFX ENDOGRAFT RESULTING IN PATIENT DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22740638&pc=KOO | 1530449-2025-00001 | 2024-06-23 04:00 | Death | PROCTER & GAMBLE MANUFACTURING CO. | 2025-08-07 04:00 | KOO | FIXODENT ADHESIVE CREAM, VERSION UNKNOWN | Insufficient Information | Airway Obstruction; Aspiration/Inhalation; Dyspnea; Hypoxia; Obstruction/Occlusion; Choking; Foreign Body In Patient | K945200 |  | 1 | Event Description: HYPOXIC BRAIN INJURY [BRAIN HYPOXIA], BLOCKED, EXTENDED INTO LEFT BRONCHUS - FIXODENT [BRONCHIAL OBSTRUCTION], THICK ADHESIVE GEL INSIDE MOUTH [FOREIGN BODY IN MOUTH], THICK ADHESIVE GEL INSIDE PHARYNX [FOREIGN BODY IN THROAT], BLOCKED LARYNX - FIXODENT [LARYNGEAL OBSTRUCTION], BLOCKED TRACHEA - FIXODENT [TRACHEAL OBSTRUCTION], UPPER AIRWAY OBSTRUCTION BY THICK ADHESIVE GEL [UPPER AIRWAY OBSTRUCTION], CHOKING [CHOKING], DENTURE ADHESIVE IN EAR [EXPOSURE VIA DIRECT CONTACT], ADULT WITH DEMENTIA FOUND WITH DENTURE ADHESIVE IN MOUTH, (EARS AND NOSE) [ACCIDENTAL DEVICE INGESTION], ADULT WITH DEMENTIA FOUND WITH DENTURE ADHESIVE IN (MOUTH), EARS AND NOSE [ACCIDENTAL EXPOSURE TO PRODUCT BY ELDERLY PERSON], DENTURE ADHESIVE IN MOUTH [EXPOSURE VIA INGESTION], INHALATION OF ADHESIVE GEL - FIXODENT [EXPOSURE VIA INHALATION], BREATHING LABORED [DYSPNOEA], DENTURE ADHESIVE GEL IN EARS [FOREIGN BODY IN EAR], THICK ADHESIVE GEL INSIDE PHARYNX, COMPLETELY BLOCKED LARYNX AND TRACHEA, EXTENDED INTO LEFT BRONCHUS [FOREIGN BODY IN RESPIRATORY TRACT], DENTURE ADHESIVE GEL IN MOUTH, NOSE LARYNX, TRACHEA, BRONCHUS PHARYNX [EXPOSURE VIA MUCOSA], DENTURE ADHESIVE GEL IN NOSE AND EARS [WRONG TECHNIQUE IN DEVICE USAGE PROCESS]. CASE NARRATIVE: A CORONER REPORTED VIA EMAIL ON 27-FEB-2025 THAT A (B)(6)-YEAR-OLD MALE PATIENT WITH DEMENTIA (DIAGNOSED ON (B)(6) 2022) THAT LIVED IN A CARE HOME WAS FOUND BY A MEMBER OF STAFF ON (B)(6) 2024 WITH FIXODENT DENTURE CARE DENTURE ADHESIVE IN HIS MOUTH, EARS, AND NOSE. UNSPECIFIED ATTEMPTS WERE MADE TO REMOVE THE SUBSTANCE FROM HIS MOUTH. HIS BREATHING BECAME LABORED (BEGINNING ON (B)(6) 2024). PARAMEDICS ATTENDED TO HIM AND TRIED TO REMOVE THE GEL WITH A SUCTION MACHINE BUT WERE UNSUCCESSFUL, THEY TOOK HIM TO THE HOSPITAL WHERE HIS HEALTH DETERIORATED, AND HE DIED ON (B)(6) 2024. MEDICAL CAUSES OF DEATH: HYPOXIC BRAIN INJURY, CHOKING, AND INHALATION OF ADHESIVE GEL. DURING THE AUTOPSY THERE WAS EVIDENCE OF UPPER AIRWAY OBSTRUCTION, THICK ADHESIVE GEL WAS SEEN INSIDE THE MOUTH AND PHARYNX, AND A COMPLETELY BLOCKED THE LARYNX AND TRACHEA THAT EXTENDED TO THE LEFT BRONCHUS. THE PRODUCT WAS USED ON A DAILY BASIS TO SECURE UPPER AND LOWER DENTURES. ADDITIONAL PRODUCT USAGE DETAIL WERE NOT PROVIDED AND IT WAS UNKNOWN IF THE EXACT PRODUCT WAS USED PREVIOUSLY. CONCOMITANT PRODUCT(S): NONE REPORTED. THE EVENT AND CASE OUTCOMES WERE FATAL. NO FURTHER INFORMATION WAS REPORTED. Manufacturer Narrative: THERE IS INSUFFICIENT INFORMATION TO PERFORM AN INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22741474&pc=MHX | 1218950-2025-000453 | 2025-07-09 04:00 | Death | PHILIPS MEDICAL SYSTEMS | 2025-08-07 04:00 | MHX | PATIENT INFORMATION CENTER IX | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K153702 |  | 1 | Event Description: PHILIPS RECEIVED A COMPLAINT ON THE PATIENT INFORMATION CENTER IX INDICATING THAT THE PATIENT WAS IN THE TELEMETRY UNIT AND HAD RED ALARMS UNTIL TRANSFERRED TO THE ICU. THE PATIENT EVENTUALLY CODED AND EXPIRED IN THE ICU. THE CUSTOMER CONTACT DID NOT REPORT ANY DEFECTS OR A DEVICE MALFUNCTION. Manufacturer Narrative: A PHILIPS REMOTE SERVICE ENGINEER (RSE) REACHED OUT AND WAS ABLE TO RETRIEVE DATA LOGS THAT WERE ANALYZED BY PRODUCT SERVICE ENGINEER (PSE). THE PSE NOTED THAT DURING THE EVENT TIMEFRAME, THERE WAS AN ACTIVE RA LEAD OFF TECHNICAL INOP MESSAGE FOR APPROXIMATELY 1 HOUR AND 15 MINUTES, FROM 2027 TO 2217. THIS CONDITION WOULD HAVE ALLOWED FOR ECG MONITORING; HOWEVER, DURING THIS SAME TIMEFRAME, OTHER TECHNICAL INOPS WERE ALSO PRESENT, WHICH WOULD NOT HAVE ALLOWED SUCCESSFUL ECG MONITORING. A 'CANNOT ANALYZE ECG' INOP WAS GENERATED 25 TIMES DURING FROM 2040 TO 2143, FOR APPROXIMATELY 14 TOTAL MINUTES. THERE WAS AN EXTENDED 'ECG LEADS OFF' INOP AT THE END OF THE EVENT TIMEFRAME, STARTING AT APPROXIMATELY 2143, WHICH WAS ACKNOWLEDGED AT 2146 BUT THE CONDITION CONTINUED UNTIL 2217. THERE WERE REMINDERS FOR THIS INOP EVERY 3 MINUTES FOR 30 MINUTES DURING THIS TIME. THERE WERE RED ECG ALARMS FOR EXTREME TACHYCARDIA AND VENTRICULAR TACHYCARDIA FOR APPROXIMATELY 15 MINUTES THROUGHOUT THE TIMEFRAME AND YELLOW ALARMS FOR AFIB, HIGH HEART RATE, PVCS AND PACER NOT PACING FOR ONE HOUR AND FIFTEEN MINUTES DURING THE EVENT TIMEFRAME. BASED ON THIS INFORMATION, THE DEVICE PROVIDED ALARMS THROUGHOUT THE EVENT TIMEFRAME WITH NO APPARENT MALFUNCTION; HOWEVER, THE CAUSE OF THE PATIENT'S DEATH WAS UNKNOWN, THOUGH IT MAY HAVE BEEN RELATED TO ALARM MANAGEMENT OR CLINICAL WORKFLOW. THE DEVICE WAS CONFIRMED TO BE OPERATING PER SPECIFICATIONS AND NO FAILURE WAS IDENTIFIED. THE INVESTIGATION CONCLUDES THAT NO FURTHER ACTION IS REQUIRED AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED THE COMPLAINT FILE WILL BE REOPENED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22741780&pc=NIP | 2024168-2025-08200 | 2021-01-01 05:00 | Death | ABBOTT VASCULAR | 2025-08-07 04:00 | NIP | SUPERA | Adverse Event Without Identified Device or Use Problem | Unspecified Vascular Problem | P120020 |  | 1 | Event Description: THIS IS FILED TO REPORT PATIENT DEATH. IT WAS REPORTED THROUGH THE PMCF REPORT, THAT THE SUPERA STENT MAY BE RELATED TO THE ADVERSE PATIENT EFFECTS OF DEATH, REVASCULARIZATION, UNPLANNED AMPUTATION / SURGICAL INTERVENTION, AND RE-HOSPITALIZATION. DETAILS ARE LISTED IN THE ATTACHED REPORT, TITLED POST-MARKET CLINICAL FOLLOW-UP EVALUATION REPORT PERIPHERAL SELF-EXPANDING STENTS. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEW WAS PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORDS FOR THIS PRODUCT COULD NOT BE REVIEWED AND A SIMILAR COMPLAINT QUERY COULD NOT BE PERFORMED BECAUSE THE PART NUMBER AND LOT NUMBER WAS NOT REPORTED, AND THE DEVICE WAS NOT RETURNED. THE REPORTED PATIENT EFFECT OF DEATH IS LISTED IN THE SUPERA PERIPHERAL STENT SYSTEMS INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT. A CONCLUSIVE CAUSE FOR THE REPORTED DEATH AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED. BASED ON THE REPORTED INFORMATION AND RESULTS OF THE COMPLAINT INVESTIGATION THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO THE DESIGN, MANUFACTURE, OR LABELING OF THE DEVICE. B2 - DATE OF DEATH: ESTIMATED. B3 - DATE OF EVENT: ESTIMATED. D4 - THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. D6A - IMPLANT DATE: ESTIMATED. THE ADDITIONAL PATIENT EFFECTS REPORTED IN THE PMCF ARE CAPTURED UNDER A SEPARATE MEDWATCH REPORT. LITERATURE ATTACHMENT: ARTICLE TITLE "POST MARKET CLINICAL FOLLOW-UP EVALUATION REPORT PERIPHERAL SELF-EXPANDING STENTS, REVISION E".
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22742088&pc=MHX | 9610816-2025-000716 | 2025-07-17 04:00 | Death | PHILIPS MEDIZIN SYSTEME BÖBLINGEN GMBH | 2025-08-07 04:00 | MHX | PATIENT INFORMATION CENTER IX | Defective Alarm | No Clinical Signs, Symptoms or Conditions | K211900 |  | 1 | Event Description: IT WAS REPORTED THE DEVICE DID NOT GENERATE AN ALARM FOR A DESATURATION AND THE PATIENT PASSED AWAY. IT WAS INDICATED THE ALARM MAY HAVE BEEN MISSED OR NOT HEARD BY THE STAFF AND THE PATIENT PASSED WAY AT APPROXIMATELY 9:42AM. BIOMED PERFORMED SIMULATED TESTING ON THE MX40 TELEMETRY DEVICE, CONFIRMING THE ALARM FUNCTIONED AND APPEARED ON BOTH THE MX40 AND THE PIC IX CENTRAL STATION. BIOMED ALSO REVIEWED THE CLINICAL AUDIT LOGS, REVEALING A DESATURATION ALARM WAS GENERATED AT 9:20AM FOR AN OXYGEN SATURATION OF 75% AND AT 9:40AM FOR AN OXYGEN SATURATION OF 74%. THE PATIENT DEATH IS BEING ADDRESSED IN RELATED MFR REPORT NUMBER 1218950-2025-000417. Manufacturer Narrative: A PHILIPS PRODUCT SUPPORT ENGINEER (PSE) WAS PROVIDED SYSTEM LOGS FOR REVIEW OF THIS ISSUE. THE PSE FOUND THE FOLLOWING IN THE 27 MINUTE TIME PERIOD FROM 9:15 UNTIL 9:42: 97 SPO2 INOPERATIVE MESSAGES (INOP TECHNICAL ALARMS) OCCURRED, INCLUDING: SPO2T SENSOR OFF (87), SPO2T NO PULSE (6), SPO2T NO SENSOR (2), SPO2T NOISY SIGNAL (2). THESE INOPERATIVE CONDITIONS PREVENTED MONITORING AND ALARMING FOR A TOTAL DURATION OF 18 MINUTES. THE MONITORING TIME BETWEEN INOPS WAS VERY SHORT (AN AVERAGE OF ABOUT 6 SECONDS BETWEEN INOPS) WHICH PREVENTED THE DEVICE FROM BEING ABLE TO COMPLETE A MEASUREMENT AND ALARM. THE INOP MESSAGES WERE ACKNOWLEDGED FROM THE PIC IX CENTRAL STATION "PICIXVCCOV7" 16 TIMES DURING THE SAME TIME PERIOD. PER THE INSTRUCTIONS FOR USE (IFU) FO THE MX40 DEVICE, "A CONFIGURABLE AVERAGING FILTER IS USED IN THE CALCULATION OF THE RESULT. DISPLAYED RESULTS ARE TYPICALLY UPDATED EVERY SECOND, BUT THE UPDATE PERIOD CAN BE AUTOMATICALLY DELAYED BY UP TO 30 SECONDS IN THE PRESENCE OF NOISE", AND "THE ALARM DELAY TIME IS 0 TO 30 SECONDS (ADJUSTABLE IN ONE SECOND STEPS) WITH A DEFAULT TIME FOR DESAT OF 20 SECONDS AND A DEFAULT TIME FOR HIGH/LOW ALARM OF 10 SECONDS. THE DELAY TIME FOR ALARM AVAILABILITY ON THE NETWORK IS LESS THAN FIVE SECONDS". IT WAS ALSO CONFIRMED THAT THE SPEAKER WAS WORKING AS EXPECTED AND THAT THE VOLUME WAS SET TO A LEVEL DETERMINED BY HOSPITAL POLICY. THE VOLUME WAS SET TO 1. THE PSE PROVIDED THE FOLLOWING INFORMATION FROM THE MX40 IFU: ALARM TEXT: <SPO2 LABEL> SENSOR OFF (NOTE ¿ THE ABILITY OF THE ALGORITHM TO DETECT THIS CONDITION DEPENDS ON THE SENSOR TYPE IN USE.) CONDITION: THE ALGORITHM HAS DETERMINED THAT A SENSOR IS CONNECTED, BUT NOT PROPERLY APPLIED TO THE PATIENT. WHAT TO DO: APPLY THE SENSOR ACCORDING TO THE MANUFACTURER'S INSTRUCTIONS. IF THE CONDITION PERSISTS, RELOCATE THE SENSOR TO A DIFFERENT SITE ON THE PATIENT. ALARM TEXT: <SPO2 LABEL> NO SENSOR (WARNING: SILENCING THIS TECHNICAL ALARM TURNS OFF THE SPO2 MEASUREMENT ON THE MX40 AND AT THE INFORMATION CENTER WHEN OPERATING IN MANUAL OR CONTINUOUS MODE. IT DOES NOT TURN THE MEASUREMENT OFF WHEN OPERATING IN AUTO MODE.) SOURCE - MX40 CONDITION: NO SENSOR ATTACHED TO SPO2 DEVICE. WHAT TO DO: ATTACH SPO2 SENSOR. ALARM TEXT: <SPO2 LABEL> NO PULSE (NOTE ¿ FOR FAST SPO2 ONLY, THIS INOP MAY ALSO BE CONFIGURED TO DISPLAY AS A RED OR YELLOW TECHNICAL ALARM.) SOURCE - MX40 CONDITION: PULSE IS TOO WEAK OR NOT DETECTABLE. WHAT TO DO: CHECK CONNECTION TO PATIENT. BASED ON THE INFORMATION AVAILABLE AND THE TESTING CONDUCTED, THE CAUSE OF THE REPORTED PROBLEM WAS RELATED TO SPO2 SENSORS NOT BEING ABLE TO READ THE SIGNAL DUE TO EITHER THE SENSOR BEING OFF, OR BEING CAUSED BY INTERFERENCE (NOISY SIGNAL). THE REPORTED PROBLEM OF AN ALARM FAILURE WAS NOT CONFIRMED, AS NUMEROUS INOP ALARMS WERE SEEN, AND THE USERS HAD SILENCED ALARMS MULTIPLE TIMES. THE DEVICE WAS CONFIRMED TO BE OPERATING PER SPECIFICATIONS AND NO FAILURE WAS IDENTIFIED. THE INVESTIGATION CONCLUDES THAT NO FURTHER ACTION IS REQUIRED AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED THE COMPLAINT FILE WILL BE REOPENED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22742335&pc=DSQ | 2916596-2025-04983 | 2025-07-02 04:00 | Death | THORATEC CORPORATION | 2025-08-07 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY AT HOME. THE CAUSE OF DEATH WAS UNKNOWN. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22743025&pc=FPO | 0001831750-2025-00718 | 2025-04-25 04:00 | Death | STRYKER MEDICAL-KALAMAZOO | 2025-08-07 04:00 | FPO | POWER-PRO TL COT | Insufficient Information | Insufficient Information |  |  | 1 | Event Description: IT WAS REPORTED THAT THE STAFF WERE TRYING TO RAISE THE COT UP TO THE CATH LAB TABLE. THE COT LOOKED AS IF IT WAS RECEIVING POWER AND THE BATTERY APPEARED TO BE CORRECTLY DOCKED BUT THE COT WOULD NOT RAISE. THE STAFF DID NOT REALIZE THAT THERE WAS A MANUAL OVERRIDE TO RAISE THE STRETCHER SO DID NOT ATTEMPT THIS. ONE OF THE PARAMEDICS WENT BACK TO THE AMBULANCE TO COLLECT THE SPARE BATTERY WHICH TOOK APPROXIMATELY 10 MINUTES. DURING THIS TIME, THE PATIENT ARRESTED AND LATER DIED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22743592&pc=NIQ | 2024168-2025-08217 | 2012-06-01 04:00 | Death | ABBOTT VASCULAR | 2025-08-07 04:00 | NIQ | XIENCE EVEROLIMUS ELUTING CORONARY STENT SYSTEM | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P070015 |  | 1 | Event Description: IT WAS REPORTED IN AN ARTICLE THAT NEXT, (NOBORI BIOLIMUS-ELUTING VERSUS XIENCE/PROMUS EVEROLIMUS-ELUTING STENT TRIAL), IS A PROSPECTIVE, MULTI-CENTRE, RANDOMISED, NON-INFERIORITY TRIAL COMPARING BIODEGRADABLE POLYMER DRUG-ELUTING STENT (BP-BES) WITH DURABLE POLYMER DRUG-ELUTING STENTS (DP-EES) IN JAPAN. WRITTEN INFORMED CONSENT WAS OBTAINED FROM ALL THE STUDY PATIENTS. THE STUDY WAS REGISTERED AT CLINICAL TRIALS.GOV (NCT01303640). THE EXTENDED FOLLOW-UP STUDY OF NEXT WAS DESIGNED WITH PLANNED FOLLOW-UP UP TO 10 YEARS. ALL THE CENTERS WERE INVITED TO PARTICIPATE IN THE EXTENDED STUDY, BUT 20 CENTERS REFUSED TO PARTICIPATE IN THE EXTENDED STUDY. SAFETY AND EFFICACY OUTCOMES OF NOBORI BP-BES WERE NON-INFERIOR TO THOSE OF XIENCE/PROMUS DP-EES. COMPLETE FIVE-YEAR FOLLOW-UP WAS ACHIEVED IN 2,408 PATIENTS. THE PATIENT EFFECTS MENTIONED IN THE ARTICLE WERE DEATH OR MYOCARDIAL INFARCTION, TARGET LESION REVASCULARISATION, TARGET VESSEL REVASCULARISATION, CORONARY REVASCULARISATION, BLEEDING, STENT THROMBOSIS. DETAILS ARE LISTED IN THE ATTACHED ARTICLE, TITLED "FIVE-YEAR OUTCOME OF A RANDOMISED TRIAL COMPARING SECONDGENERATION DRUG-ELUTING STENTS USING EITHER BIODEGRADABLE POLYMER OR DURABLE POLYMER: THE NOBORI BIOLIMUS-ELUTING VERSUS XIENCE/PROMUS EVEROLIMUS-ELUTING STENT TRIAL (NEXT)". NO OTHER INFORMATION WAS PROVIDED. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORDS FOR THIS PRODUCT COULD NOT BE REVIEWED AND A SIMILAR COMPLAINT QUERY COULD NOT BE PERFORMED BECAUSE THE PART NUMBER AND/OR LOT NUMBER WAS NOT REPORTED, AND THE DEVICE WAS NOT RETURNED. THE REPORTED PATIENT EFFECT(S) OF DEATH IS LISTED IN THE XIENCE V EVEROLIMUS ELUTING CORONARY STENT SYSTEMS INSTRUCTIONS FOR USE AS A KNOWN PATIENT EFFECT(S) OF CORONARY STENTING PROCEDURES. A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT DEATH, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED; HOWEVER, THE SUBSEQUENT TREATMENT APPEARS TO BE RELATED TO THE OPERATIONAL CONTEXT OF THE PROCEDURE. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING; THEREFORE, NO PRODUCT-RELATED CORRECTIVE ACTION WILL BE IMPLEMENTED IN THIS CASE. LITERATURE ATTACHMENT: ARTICLE TITLE: ¿FIVE-YEAR OUTCOME OF A RANDOMISED TRIAL COMPARING SECONDGENERATION DRUG-ELUTING STENTS USING EITHER BIODEGRADABLE POLYMER OR DURABLE POLYMER: THE NOBORI BIOLIMUS-ELUTING VERSUS XIENCE/PROMUS EVEROLIMUS-ELUTING STENT TRIAL (NEXT).¿ B3: DATE OF DEATH IS ESTIMATED. B3, B6: DATES ESTIMATED. D6: IMPLANT DATE IS ESTIMATED. D4: THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. THE ADDITIONAL DEVICE AND PATIENT EFFECTS THAT ARE REFERENCED IN B5 ARE FILED UNDER SEPARATE MEDWATCH REPORT NUMBER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22743895&pc=NIQ | 2024168-2025-08220 | 2012-06-01 04:00 | Death | ABBOTT VASCULAR | 2025-08-07 04:00 | NIQ | PROMUS EVEROLIMUS ELUTING CORONARY STENT SYSTEM | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P070015 |  | 1 | Event Description: IT WAS REPORTED IN AN ARTICLE THAT NEXT, (NOBORI BIOLIMUS-ELUTING VERSUS XIENCE/PROMUS EVEROLIMUS-ELUTING STENT TRIAL), IS A PROSPECTIVE, MULTI-CENTRE, RANDOMISED, NON-INFERIORITY TRIAL COMPARING BIODEGRADABLE POLYMER DRUG-ELUTING STENT (BP-BES) WITH DURABLE POLYMER DRUG-ELUTING STENTS (DP-EES) IN JAPAN. WRITTEN INFORMED CONSENT WAS OBTAINED FROM ALL THE STUDY PATIENTS. THE STUDY WAS REGISTERED AT CLINICAL TRIALS.GOV (NCT01303640). THE EXTENDED FOLLOW-UP STUDY OF NEXT WAS DESIGNED WITH PLANNED FOLLOW-UP UP TO 10 YEARS. ALL THE CENTERS WERE INVITED TO PARTICIPATE IN THE EXTENDED STUDY, BUT 20 CENTERS REFUSED TO PARTICIPATE IN THE EXTENDED STUDY. SAFETY AND EFFICACY OUTCOMES OF NOBORI BP-BES WERE NON-INFERIOR TO THOSE OF XIENCE/PROMUS DP-EES. COMPLETE FIVE-YEAR FOLLOW-UP WAS ACHIEVED IN 2,408 PATIENTS. THE PATIENT EFFECTS MENTIONED IN THE ARTICLE WERE DEATH OR MYOCARDIAL INFARCTION, TARGET LESION REVASCULARISATION, TARGET VESSEL REVASCULARISATION, CORONARY REVASCULARISATION, BLEEDING, STENT THROMBOSIS. DETAILS ARE LISTED IN THE ATTACHED ARTICLE, TITLED "FIVE-YEAR OUTCOME OF A RANDOMISED TRIAL COMPARING SECONDGENERATION DRUG-ELUTING STENTS USING EITHER BIODEGRADABLE POLYMER OR DURABLE POLYMER: THE NOBORI BIOLIMUS-ELUTING VERSUS XIENCE/PROMUS EVEROLIMUS-ELUTING STENT TRIAL (NEXT)". NO OTHER INFORMATION PROVIDED. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORDS FOR THIS PRODUCT COULD NOT BE REVIEWED AND A SIMILAR COMPLAINT QUERY COULD NOT BE PERFORMED BECAUSE THE PART NUMBER AND/OR LOT NUMBER WAS NOT REPORTED, AND THE DEVICE WAS NOT RETURNED. THE REPORTED PATIENT EFFECT(S) OF DEATH IS LISTED IN THE PROMUS EVEROLIMUS ELUTING CORONARY STENT SYSTEMS INSTRUCTIONS FOR USE AS A KNOWN PATIENT EFFECT(S) OF CORONARY STENTING PROCEDURES. A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECT(S), AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED; HOWEVER, THE SUBSEQUENT TREATMENT APPEARS TO BE RELATED TO THE OPERATIONAL CONTEXT OF THE PROCEDURE. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING; THEREFORE, NO PRODUCT-RELATED CORRECTIVE ACTION WILL BE IMPLEMENTED IN THIS CASE. LITERATURE ATTACHMENT: ARTICLE TITLE: ¿FIVE-YEAR OUTCOME OF A RANDOMISED TRIAL COMPARING SECONDGENERATION DRUG-ELUTING STENTS USING EITHER BIODEGRADABLE POLYMER OR DURABLE POLYMER: THE NOBORI BIOLIMUS-ELUTING VERSUS XIENCE/PROMUS EVEROLIMUS-ELUTING STENT TRIAL (NEXT)¿ B3: DATE OF DEATH IS ESTIMATED B3, B6: DATE OF EVENT IS ESTIMATED D6: IMPLANT DATE IS ESTIMATED D4: THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. THE ADDITIONAL DEVICE AND PATIENT EFFECTS THAT ARE REFERENCED IN B5 ARE FILED UNDER SEPARATE MEDWATCH REPORT NUMBER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22744944&pc=QFG | 3019004087-2025-00799 | 2025-07-20 04:00 | Death | BETA BIONICS ORGANIZATION | 2025-08-07 04:00 | QFG | ILET BIONIC PANCREAS | Obstruction of Flow; Device Dislodged or Dislocated; Unintended Movement | Dyspnea; Hyperglycemia; Vomiting; Diabetic Ketoacidosis; Confusion/ Disorientation | K223846 |  | 1 | Event Description: IT WAS REPORTED THAT THE USER WAS HOSPITALIZED ON (B)(6) 2025 FOLLOWING THE ACCIDENTAL DISLODGEMENT OF THEIR INFUSION SET, WHICH WAS NOT NOTICED UNTIL THE FOLLOWING DAY. THE USER EXPERIENCED ELEVATED BLOOD GLUCOSE LEVELS AND WAS TREATED IN THE HOSPITAL FOR DIABETIC KETOACIDOSIS (DKA) WITH INSULIN INJECTIONS AND INTRAVENOUS FLUIDS. THE ILET DEVICE WAS NOT IN USE AT THE TIME OF THE EVENT. THE USER WAS ADMITTED TO THE HOSPITAL AND SUBSEQUENTLY PASSED AWAY ON (B)(6) 2025. THE CAREGIVER INDICATED THAT THE USER HAD PRE-EXISTING MEDICAL CONDITIONS, INCLUDING MULTIPLE ORGAN FAILURE, HIGH BLOOD PRESSURE, HEART FAILURE, AND AN INTELLECTUAL DISABILITY. NO CONCERNS WERE RAISED REGARDING THE ILET CONTRIBUTING TO THE EVENT. THE DEVICE IS BEING RETURNED FOR EVALUATION. Manufacturer Narrative: NO PRODUCT WAS RETURNED FOR EVALUATION. THE DEVICE HISTORY RECORD (DHR) REVIEW WAS COMPLETED, AND THIS DEVICE PASSED ALL MANUFACTURING RELEASE CRITERIA FOR DISTRIBUTION. THERE WERE NO ISSUES IDENTIFIED THAT WOULD HAVE IMPACTED THIS EVENT. AT THE TIME OF THIS REPORT THE ILET EVALUATION IS STILL IN PROGRESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22745205&pc=QEW | 3020347218-2025-00050 | 2025-07-10 04:00 | Death | INARI MEDICAL, INC. | 2025-08-07 04:00 | QEW | CLOTTRIEVER XL CATHETER | Adverse Event Without Identified Device or Use Problem | Pulmonary Embolism | K242557 |  | 1 | Event Description: A 58-YEAR-OLD PATIENT WITH LEFT ILIAC DEEP VEIN THROMBOSIS THAT EXTENDED INTO THE INFERIOR VENA CAVA (IVC), UNDERWENT A THROMBECTOMY WITH THE CLOTTRIEVER XL CATHETER (CTXL). THE PATIENT WAS PLACED PRONE AND EXPRESSED THAT THEY COULD NOT BREATH WELL. THE PATIENT'S O2 WAS MEASURED TO BE 89% AT THIS TIME, SO THE PROCEDURE WAS CONTINUED. DUE TO THE PATIENT'S ELEVATED CREATINE LEVELS, NO COMPUTED TOMOGRAPHY SCAN WAS PERFORMED PRIOR TO THE PROCEDURE, AND NO CONTRAST WAS USED. INTRAVENOUS ULTRASOUND WAS USED AND SHOWED SIGNIFICANT CLOT FROM THE ILIAC INTO THE IVC. NO CLOT CAPTURE WAS USED AND THE CTXL WAS USED TO PERFORM 2 PASSES. BETWEEN THE PASSES, THE PATIENT BEGAN SAYING THEY COULD NOT BREATHE DESPITE THEIR O2 BEING 88%. THE PATIENT WAS PLACED ON A NON-REBREATHER WHICH HELPED IMPROVE THEIR O2. HOWEVER, THE PATIENT WAS ALSO TACHYCARDIC AND BEGAN EXPRESSING CONCERNS OVER BREATHING AGAIN. THE PATIENT ALSO BEGAN TO PUSH HIMSELF UP FROM THE TABLE. THE PATIENT WAS FLIPPED OVER BUT DURING THIS PROCESS, THE PATIENT'S HEART RATE DROPPED AND STOPPED. CARDIOPULMONARY RESUSCITATION WAS PERFORMED FOR 40 MINUTES BUT UNFORTUNATELY, A PULSE WAS NEVER REGAINED. Manufacturer Narrative: THE INARI DEVICES WERE DISCARDED BY THE USER FACILITY AND ARE NOT AVAILABLE FOR EVALUATION. THE DEVICE IDENTIFIERS WERE PROVIDED, AND THE LOT HISTORY RECORDS WERE REVIEWED. THERE WERE NO DISCREPANCIES OR UNUSUAL FINDINGS. THERE WAS NO REPORT OF ANY DEVICE MALFUNCTION. BASED ON THE INFORMATION REVIEWED, THERE IS NO EVIDENCE TO INDICATE THE PRESENCE OF A POTENTIAL QUALITY ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. THE CASE WAS REVIEWED BY INARI STRYKER MEDICAL AFFAIRS, WHO CONCLUDED THAT THE PATIENT'S DEATH WAS POTENTIALLY THE RESULT OF INADVERTENT CLOT EMBOLISM. MANUFACTURER REFERENCE #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22745281&pc=QBJ | 2954323-2025-31217 | 2025-06-30 04:00 | Death | ABBOTT DIABETES CARE LTD | 2025-08-07 04:00 | QBJ | FREESTYLE LIBRE 3 | Insufficient Information | Insufficient Information | K223435 |  | 1 | Event Description: ABBOTT DIABETES CARE RECEIVED A MEDWATCH REPORT # MW5172345, WHICH REPORTED THE FOLLOWING INFORMATION PERTAINING TO AN ABBOTT DIABETES CARE (ADC) DEVICE: A HEALTHCARE PROFESSIONAL (HCP) REPORTED ON JULY 1, 2025, THAT A PATIENT PASSED AWAY AND INDICATED THAT NO FURTHER INFORMATION OR CLARIFICATION IS AVAILABLE. IT IS UNKNOWN IF THE PATIENT WAS ACTIVELY USING AN ADC DEVICE DURING THE TIME OF THE PATIENT'S PASSING AND NO ALLEGATION OF A DEVICE DEFICIENCY WAS MADE. THE REPORTER DECLINED TO PROVIDE CONTACT INFORMATION TO ADC THEREFORE, ADC IS UNABLE TO ATTEMPT ANY FOLLOW UP ACTIVITIES RELATED TO THIS EVENT. BASED ON THE INFORMATION PROVIDED, IT CANNOT BE REASONABLY SUGGESTED THAT THE ADC PRODUCT HAS OR MAY HAVE CAUSED OR CONTRIBUTED TO THE DEATH. Manufacturer Narrative: THIS COMPLAINT WAS RECEIVED VIA USER REPORT # MW5172345. ABBOTT DIABETES CARE (ADC) IS UNABLE TO FURTHER INVESTIGATE THE REPORTED COMPLAINT DUE TO THE ABSENCE OF AN ALLEGED PRODUCT ISSUE. ADDITIONALLY, NO CONTACT INFORMATION WAS PROVIDED IN THE USER REPORT, RENDERING ADC UNABLE TO CONDUCT ANY FOLLOW-UP ACTIVITIES. ALL PERTINENT INFORMATION AVAILABLE TO ADC HAS BEEN SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22745544&pc=MIH | 2017233-2025-06492 | 2025-08-01 04:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-08-07 04:00 | MIH | GORE® TAG® THORACIC BRANCH ENDOPROSTHESIS | Off-Label Use; Migration | Hyperemia; High Blood Pressure/ Hypertension; Vascular Dissection; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P210032 |  | 1 | Event Description: ON (B)(6) 2025, THE PATIENT SUSTAINED AN AORTIC DISSECTION DURING A TRANSCATHETER AORTIC VALVE IMPLANTATION (TAVI). GORE® TAG® THORACIC BRANCH ENDOPROSTHESIS AORTIC EXTENDER DEVICES WERE USED TO COVER THE ENTRY TEAR WHICH APPEARED TO BE ~2.5CM DISTAL TO THE CORONARY ARTERIES. THE WIRE ACCESS WAS GAINED AND CONFIRMED TO BE IN THE TRUE LUMEN WITH A TRANS-OESOPHAGEAL ECHOCARDIOGRAM. THE TE3742E AORTIC EXTENDER WAS ADVANCED TO AND DEPLOYED AT THE ATTENDED POSITION. REPORTEDLY, WHEN THE DELIVERY CATHETER WAS RETRIEVED, THE AORTIC EXTENDER STARTED TO MIGRATE DISTALLY (BEST ESTIMATE ~6MM). EVENTUALLY THEN FORWARD PRESSURE WAS ADDED TO THE WIRE, THE AORTIC EXTENDER STOPPED MOVING AND THE CATHETER WAS RETRIEVED. NO VESSELS¿ COVERAGE WERE NOTICED DURING THE DISTAL DEVICE MIGRATION. THE PATIENT THEN STARTED TO LOSE CARDIAC OUTPUT, AND IT WAS DETERMINED THERE WAS LIKELY A RETROGRADE PERICARDIAL BLEED. A LINE WAS INSERTED INTO THIS SPACE AND BLOOD RETRIEVED. IT WAS DETERMINED THAT AN ADDITIONAL AORTIC EXTENDER (TE4043) WAS REQUIRED. THE DEVICE THEN ALSO MOVED DISTALLY. HOWEVER, THE SURGEON REPORTED THAT IT APPEARED TO HAPPEN WITHOUT PRESSURE ON THE CATHETER. THE DEVICE WAS THEN ATTEMPTED TO BE RETRIEVED. THIS AGAIN IMMEDIATELY CAUSED THE EXTENDER TO MODE MORE DISTALLY. THE PHYSICIAN STARTED TO TWIST THE CATHETER. THE CATHETER WAS THEN ABLE TO BE RETRIEVED AS THE SLEEVE HAD PULLED BACK FAR ENOUGH TO NOT CATCH ON ANYTHING. ON (B)(6) 2025, DURING THE PROCEDURE, THE PATIENT EXPIRED DUE TO THE BLEED INTO THE PERITONEAL SPACE. THE PHYSICIAN DID NOT MENTION THAT THE PATIENT¿S DEATH WAS DUE TO THE GORE DEVICE USED. THE PHYSICIAN BELIEVED THAT THE AORTIC EXTENDER PULLED BACK NOT ONLY DUE TO THE SLEEVE PULLING ON THE DEVICE BUT ALSO DUE TO THE REVERSE TAPERED ANATOMY (OUTSIDE OF THE INSTRUCTIONS FOR USE) AND THE HIGH PRESSURE OF BLOOD. ACCORDING TO THE PHYSICIAN, THE PATIENT¿S AORTA WAS FRAGILE AND UNSTABLE. IT WAS REPORTED THAT THE AORTIC EXTENDER DEVICES WERE OFF LABEL USED AS THERE WERE NO OTHER POSSIBLE COVERED STENTS AVAILABLE TO COVER THE ENTRY TEAR WITHOUT COVERING THE INNOMINATE AND THEY WERE NOT WANTING TO PERFORM BYPASS SURGERY AS THERE WERE DISSECTIONS INTO THE HEAD VESSELS. Manufacturer Narrative: H6. CODE C21: A REVIEW OF MANUFACTURING RECORDS OF THE DEVICES IS GOING TO BE CONDUCTED. THE INVESTIGATION IS IN PROCESS. THE DEVICES REMAIN IMPLANTED AND ARE NOT AVAILABLE FOR ANALYSIS. ALSO WAS IMPLANTED: TE4043E/28601663 UDI# (B)(4). IMAGES WERE PROVIDED FOR INVESTIGATION. FURTHER INFORMATION WILL BE PROVIDED. CODE E2402- WAS USED TO COVER THAT THE PATIENT EXPIRED DUE TO THE BLEED INTO THE PERITONEAL SPACE. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY HAVE NOT BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. Manufacturer Narrative: B6: IMAGING EVALUATION WAS ATTACHED. CANNOT CONFIRM A DISSECTION OR PRIOR DISSECTION/TEAR IN THE ASCENDING THORACIC AORTA WITH IMAGES AVAILABLE FOR EVALUATION. CANNOT CONFIRM MIGRATION WITH ONE EVALUABLE JPEG IMAGE. ADDITIONAL IMAGES WERE REQUESTED, HOWEVER, WERE NOT PROVIDED. H6. CODE C19: A REVIEW OF THE MANUFACTURING RECORDS OF THE DEVICES INDICATED THE LOTS MET ALL THE PRE-RELEASE SPECIFICATIONS. THE DEVICES REMAIN IMPLANTED AND ARE NOT AVAILABLE FOR ANALYSIS. ALSO WAS IMPLANTED: TE4043E/(B)(6) UDI# (B)(4). CODE E2402- WAS USED TO COVER THAT THE PATIENT EXPIRED DUE TO THE BLEED INTO THE PERITONEAL SPACE. THE PHYSICIAN BELIEVED THAT THE AORTIC EXTENDER PULLED BACK NOT ONLY DUE TO THE SLEEVE PULLING ON THE DEVICE BUT ALSO DUE TO THE REVERSE TAPERED ANATOMY (OUTSIDE OF THE INSTRUCTIONS FOR USE) AND THE HIGH PRESSURE OF BLOOD. ACCORDING TO THE PHYSICIAN, THE PATIENT¿S AORTA WAS FRAGILE AND UNSTABLE. IT WAS REPORTED THAT THE AORTIC EXTENDER DEVICES WERE OFF LABEL USED AS THERE WERE NO OTHER POSSIBLE COVERED STENTS AVAILABLE TO COVER THE ENTRY TEAR WITHOUT COVERING THE INNOMINATE AND THEY WERE NOT WANTING TO PERFORM BYPASS SURGERY AS THERE WERE DISSECTIONS INTO THE HEAD VESSELS. ACCORDING TO THE GORE® TAG® THORACIC BRANCH ENDOPROSTHESIS INSTRUCTIONS FOR USE (IFU), THE GORE® TAG® THORACIC BRANCH ENDOPROSTHESIS IS A MODULAR DEVICE CONSISTING OF THE AORTIC COMPONENT, THE SIDE BRANCH (SB) COMPONENT, AND AN OPTIONAL AORTIC EXTENDER AND THESE COMPONENTS MAY BE USED TOGETHER AS A STAND-ALONE DEVICE OR IN CONJUNCTION WITH THE GORE® TAG® CONFORMABLE THORACIC STENT GRAFT. THE AORTIC EXTENDER IS INTENDED TO BE USED TO IMPROVE SEALING OF THE AORTIC COMPONENT. PER IFU, ADDITIONAL CONSIDERATIONS FOR PATIENT SELECTION INCLUDE, BUT ARE NOT LIMITED TO: PATIENT¿S ANATOMICAL SUITABILITY FOR ENDOVASCULAR REPAIR. ADDITIONALLY, THE IFU STATES: USE OF THE GORE® TAG® THORACIC BRANCH ENDOPROSTHESIS OUTSIDE OF THE RECOMMENDED ANATOMICAL SIZING GUIDELINES MAY RESULT IN POTENTIALLY SERIOUS DEVICE-RELATED EVENTS (MIGRATION). THE SIZING MEASUREMENTS ARE CRITICAL TO THE SUCCESSFUL PERFORMANCE OF THE ENDOVASCULAR REPAIR. ACCORDING TO THE IFU, POSSIBLE ADVERSE EVENTS AND COMPLICATIONS THAT MAY OCCUR WITH THE USE OF GORE® TAG® THORACIC BRANCH ENDOPROSTHESIS INCLUDE, BUT ARE NOT LIMITED TO DEPLOYMENT DIFFICULTIES/FAILURES, BLEEDING (PROCEDURAL AND POST-TREATMENT), CARDIAC COMPLICATIONS (HYPERTENSION), ENDOPROSTHESIS MIGRATION AND DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22745590&pc=NVY | 2017865-2025-97095 | 2025-07-16 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-08-07 04:00 | NVY | OPTISURE ACTIVE FIXATION, DF-4 CONNECTOR | Failure to Capture; Adverse Event Without Identified Device or Use Problem | Insufficient Information | P950022 |  | 1 | Event Description: IT WAS REPORTED THAT LOSS OF CAPTURE WAS EXHIBITED BY THE PATIENT'S IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) SYSTEM. THE PATIENT HAD PASSED AWAY FOR UNKNOWN CAUSES AND IT WAS UNKNOWN IF IT WAS RELATED TO ITS ICD SYSTEM.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22745591&pc=NIK | 2017865-2025-96638 | 2025-07-16 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-08-07 04:00 | NIK | GALLANT HF | Failure to Capture; Adverse Event Without Identified Device or Use Problem | Insufficient Information | P030054 |  | 1 | Event Description: IT WAS REPORTED THAT LOSS OF CAPTURE WAS EXHIBITED BY THE PATIENT'S IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) SYSTEM. THE PATIENT HAD PASSED AWAY FOR UNKNOWN CAUSES AND IT WAS UNKNOWN IF IT WAS RELATED TO ITS ICD SYSTEM.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22745595&pc=NIK | 2017865-2025-97098 | 2025-07-16 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-08-07 04:00 | NIK | QUARTET | Failure to Capture; Adverse Event Without Identified Device or Use Problem | Insufficient Information | P030054 |  | 1 | Event Description: IT WAS REPORTED THAT LOSS OF CAPTURE WAS EXHIBITED BY THE PATIENT'S IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) SYSTEM. THE PATIENT HAD PASSED AWAY FOR UNKNOWN CAUSES AND IT WAS UNKNOWN IF IT WAS RELATED TO ITS ICD SYSTEM.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22745597&pc=NVN | 2017865-2025-97097 | 2025-07-16 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-08-07 04:00 | NVN | TENDRIL SDX LEAD | Failure to Capture; Adverse Event Without Identified Device or Use Problem | Insufficient Information | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT LOSS OF CAPTURE WAS EXHIBITED BY THE PATIENT'S IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) SYSTEM. THE PATIENT HAD PASSED AWAY FOR UNKNOWN CAUSES AND IT WAS UNKNOWN IF IT WAS RELATED TO ITS ICD SYSTEM.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22745711&pc=QLG | 2954323-2025-31222 | 2025-07-09 04:00 | Death | ABBOTT DIABETES CARE INC | 2025-08-07 04:00 | QLG | FREESTYLE LIBRE 2 | Insufficient Information | No Clinical Signs, Symptoms or Conditions | K193371 |  | 1 | Event Description: ABBOTT DIABETES CARE (ADC) RECEIVED A MEDWATCH REPORT # MW5172344 WHICH REPORTED THE FOLLOWING INFORMATION: A HEALTHCARE PROFESSIONAL (HCP) REPORTED ON 30-JUN-25 THAT A PATIENT WHO WAS A USER OF THE ADC DEVICE ¿PASSED BUT WAS UNABLE TO CONFIRM IF ON MEDICATION AT TIME OF PASSING." IT IS UNKNOWN IF THE PATIENT WAS ACTIVELY USING AN ADC DEVICE DURING THE TIME OF THE PATIENT'S PASSING BUT NO ALLEGATION OF DEVICE DEFICIENCY WAS MADE. THE REPORTER DECLINED TO PROVIDE CONTACT INFORMATION TO ADC; THEREFORE, ADC IS UNABLE TO ATTEMPT ANY FOLLOW UP ACTIVITIES RELATED TO THIS EVENT. BASED ON THE INFORMATION PROVIDED, IT CANNOT BE REASONABLY SUGGESTED THAT THE ADC PRODUCT HAS OR MAY HAVE CAUSED OR CONTRIBUTED TO THE DEATH. Manufacturer Narrative: THIS COMPLAINT WAS RECEIVED VIA USER REPORT AND HAS BEEN REPORTED TO FDA. A VALID SERIAL NUMBER HAS NOT BEEN PROVIDED. AN INVESTIGATION HAS BEEN PERFORMED FOR THE REPORTED COMPLAINT AND THERE WAS NO INDICATION THAT THE PRODUCT DID NOT MEET SPECIFICATION. TRACKING AND TRENDING REPORTS FOR THE PAST YEAR WAS REVIEWED BASED ON THE PRODUCT LINE AND REPORTED ISSUE. NO TRIP WAS IDENTIFIED AND NO FURTHER ACTION WAS REQUIRED. THE MOST PROBABLE ROOT CAUSES ASSOCIATED WITH THIS FAILURE MODE ARE DISCONNECTED, FAULTY OR DAMAGED COMPONENTS OR MISUSE. HOWEVER, MITIGATIONS ARE IN PLACE TO REDUCE AND PREVENT SUCH ISSUES. ALL VITAL MANUFACTURING STEPS ARE VALIDATED, MONITORED, AND VERIFIED DURING MANUFACTURING TO ENSURE THE SYSTEM IS IN CONFORMANCE WITH THE VERIFIED DESIGN. LABELING IS PROVIDED TO INSTRUCT THE USER ON THE INTENDED USE OF ALL VITAL PARTS OF THE SYSTEM TO MINIMIZE MISUSE. ALL COMPLAINTS AND COMPLAINT TRENDS ARE INVESTIGATED TO DETERMINE IF THERE IS A PRODUCT DEFECT/ DEFICIENCY. IF A PRODUCT DEFECT/ DEFICIENCY IS IDENTIFIED, A RISK EVALUATION IS COMPLETED AND COMPARED TO THE RISK MANAGEMENT REPORT, TO ENSURE THE RISK PROFILE HAS NOT CHANGED. ADDITIONALLY, AS A PART OF ABBOTT¿S POST-MARKET SURVEILLANCE PROCESS, ALL RISK EVALUATIONS WITH ASSOCIATED COMPLAINT DATA ARE REVIEWED ANNUALLY TO DETERMINE IF THE RISK PROFILES HAVE CHANGED AS COMPARED TO THE PRODUCT RISK MANAGEMENT REPORTS. THESE MONITORING PROCESSES ENSURE THAT ALL PRODUCT RISK PROFILES REMAIN ACCEPTABLE AND HAVE A POSITIVE BENEFIT/ RISK RATIO. IF PRODUCT IS RETURNED, A PHYSICAL INVESTIGATION WILL BE PERFORMED PER ADC'S ESTABLISHED PROCESSES AND PROCEDURES AND A REPORT WILL BE SUBMITTED UPON COMPLETION OF INVESTIGATION. ALL PERTINENT INFORMATION AVAILABLE TO ABBOTT DIABETES CARE HAS BEEN SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22745919&pc=NPT | 9617601-2025-00863 | 2017-03-10 05:00 | Death | MEDTRONIC MEXICO S. DE R.L. DE CV | 2025-08-07 04:00 | NPT | EVOLUT R TRANSCATHETER AORTIC VALVE | Fluid/Blood Leak; Perivalvular Leak; Appropriate Term/Code Not Available; Patient Device Interaction Problem | Arrhythmia; Atrial Flutter; Bacterial Infection; Cardiac Arrest; Dyspnea; Cardiac Enzyme Elevation; Septic Shock; Cardiogenic Shock; Respiratory Failure; Aortic Valve Insufficiency/ Regurgitation | P130021 |  | 1 | Event Description: MEDTRONIC RECEIVED INFORMATION THAT TWO DAYS AFTER THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, AN ECHOCARDIOGRAM INDICATED TRIVIAL PARAVALVULAR LEAK (PVL) AND REGURGITATION. THE PATIENT WAS ASYMPTOMATIC. NO TREATMENT WAS PRESCRIBED. NO ADVERSE PATIENT EFFECTS WERE REPORTED. ADDITIONAL INFORMATION RECEIVED INDICATED THAT APPROXIMATELY 8 YEARS AND 4 MONTHS FOLLOWING THE TRANSCATHETER VALVE IMPLANT THE PATIENT DEVELOPED SUDDEN ONSET DYSPNEA ON EXERTION AFTER ALCOHOL CONSUMPTION OVER A FEW DAYS. THE PATIENT PRESENTED AS AN OUTPATIENT TO THE CARDIOLOGIST 3 DAYS LATER FOR AN EVALUATION. AN ELECTROCARDIOGRAM (ECG) IDENTIFIED ATRIAL FLUTTER WITH A RAPID VENTRICULAR RESPONSE (RVR). BLOOD WORK WAS DRAWN AT THAT VISIT AND THE RESULTS OBTAINED THE FOLLOWING DAY NOTED AN INITIAL HIGH SENSITIVITY ELEVATED TROPONIN OF 798 NG/L. THE PATIENT WAS REFERRED TO THE EMERGENCY ROOM THAT SAME DAY AND ADMITTED TO THE HOSPITAL FOR EVALUATION. A REPEAT TROPONIN MEASURED 699 NG/L. THE TROPONIN WAS BELIEVED TO BE RELATED TO THE ATRIAL FLUTTER WITH THE RVR. HOWEVER, A LEFT HEART CATHETERIZATION WAS PLANNED FOR THE FOLLOWING DAY. A CALCIUM CHANNEL BLOCKER, BETA BLOCKER, AND AN ANTICOAGULANT WERE ADMINISTERED. A TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) AND CARDIOVERSION WERE PLANNED AFTER THE LEFT HEART CATHETERIZATION. THE HEART CATHETERIZATION NOTED NO SIGNIFICANT OBSTRUCTIVE/SIGNIFICANT CORONARY ARTERY DISEASE (CAD). TWO DAYS FOLLOWING THE ADMISSION THE TEE AND CARDIOVERSION WERE PLANNED. THE TEE IDENTIFIED AN UNSPECIFIED VALVE FAILURE WITH SEVERE AORTIC REGURGITATION. THE PATIENT WAS REFERRED TO CARDIAC SURGERY FOR A TRANSCATHETER VALVE INTERVENTION. OF NOTE, THE VALVE FAILURE AND REGURGITATION WERE NOT DEMONSTRATED ON THE TRANSTHORACIC ECHOCARDIOGRAM¿S (TTE) PERFORMED 3 DAYS PRIOR. THE CARDIOVERSION WAS NOT PERFORMED DUE TO THE FAILURE OF THE TRANSCATHETER VALVE. THE HEART RATE WAS REPORTED AS ELEVATED. TWO DAYS LATER, THE PATIENT¿S CONDITION DECOMPENSATED WITH ACUTE RESPIRATORY FAILURE AND SUPRAVENTRICULAR TACHYCARDIA (SVT) WHICH RESULTED IN CODING TWICE. THE PHYSICIAN INDICATED THIS WAS RELATED TO THE POOR PERFUSION AS RESULT OF THE FAILING TRANSCATHETER VALVE WITH LIKELY CARDIOGENIC SHOCK. THE PATIENT WAS INTUBATED AND REPORTED AS OBTUNDED. THE PATIENT WAS MAXIMIZED ON PRESSORS AND INOTROPES WITH THE MEAN ATRIAL PRESSURE (MAP) THAT CONTINUED TO BE LOW. ADDITIONAL MEDICATIONS INCLUDED INTRAVENOUS STEROIDS, DIURETIC, ADRENERGIC RECEPTOR AGONISTS, ALPHA-2 ADRENERGIC RECEPTOR AGONIST, CATECHOLAMINE AND A BETA-ADRENERGIC AGONIST, EPINEPHRINE, METHYLENE BLUE, AND SODIUM BICARBONATE GTT. THE PREVIOUS PLANNED SURGICAL INTERVENTION FOR THE TRANSCATHETER VALVE WAS CANCELED. THE PATIENT REMAINED A RICHMOND AGITATION SEDATION SCORE (RASS) OF -5 AND A GLASGOW COMA SCALE OF 3 WITH NO WITHDRAWAL FROM PAINFUL STIMULI OR OTHER REFLEXES. BLOOD CULTURES WERE OBTAINED AFTER THE CODE EPISODES AND RETURNED LATER IN THE DAY WITH POSITIVE CULTURES IN 3 OF 4 OF STAPHYLOCOCCUS AUREUS. THE LEFT ANTECUBITAL (AC) BLOOD DRAW HAD TWO POSITIVE RESULTS AND THE RIGHT ARM HAD 1 OF 2 POSITIVE RESULTS FOR SAME BACTERIA. THE PREVIOUS TWO TTES AND THE TEE DID NOTDEMONSTRATE ANY VALVE VEGETATIONS OR SOURCE OF INFECTION IN THE HEART. URINALYSIS COMPLETED. PER THE FAMILY REQUEST, THE PATIENT STATUS WAS DNR (DO NOT RESUSCITATE). THE PATIENT DIED AS RESULT OF MIXED CARDIOGENIC AND SEPTIC SHOCK. THE CARDIOGENIC SHOCK WAS REPORTED SECONDARY TO THE FAILED TRANSCATHETER VALVE. THE SEPTIC SHOCK WAS REPORTED SECONDARY TO STAPHYLOCOCCUS AUREUS LIKELY FROM URINARY ORIGIN. THE DEATH WAS CLASSIFIED AS CARDIAC. AN AUTOPSY WAS NOT PERFORMED. THE EVENTS WERE REPORTED RELATED TO THE VALVE.Event Description: ADDITIONAL INFORMATION RECEIVED INDICATED THAT AT THE 8 YEAR AND 4 MONTH MARKER PARAVALVULAR LEAK (PVL) WAS NOT IDENTIFIED ON THE TRANSTHORACIC ECHOCARDIOGRAM (TTE) NOR THE TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE). HOWEVER, THE TEE IDENTIFIED THE SEVERE CENTRAL REGURGITATION. AS REPORTED, THE AORTA REVERSE TO FORWARD VELOCITY-TIME-INTEGRAL (VTI) RATIO WAS INCREASED IN THE DESCENDING THORACIC AORTA WHICH WAS CONSISTENT WITH SEVERE AORTIC REGURGITATION/INCREASED REGURGITATION FRACTION AND VOLUME. THE PHYSICIANS INDICATED THE VALVE FAILURE LED TO THE REGURGITATION AND SUBSEQUENT DEATH. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5, H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22723992&pc=NRY | 3014644549-2025-00001 | 2025-07-05 04:00 | Death | WALLABY MEDICAL | 2025-08-06 04:00 | NRY | ESPERANCE 3+ ASPIRATION CATHETER SYSTEM | Detachment of Device or Device Component | Hemorrhage/Bleeding | K240917 |  | 1 | Event Description: EMERGENT THROMBECTOMY PERFORMED. DURING THE THIRD PASS FOR THROMBECTOMY, PART OF THE ESPERANCE ASPIRATION CATHETER BROKE OFF INTO PATIENT'S RIGHT MIDDLE CEREBRAL ARTERY. THE PIECE WAS RETRIEVED USING A STENT RETRIEVER; HOWEVER, SIGNIFICANT BLEEDING WAS NOTED. ESPERANCE CATHETER: MANUFACTURER:WALLABY MEDICAL LOT#: ASP00572, REF#: ASP3F120KIT, EXP: 2025.12.06 .THE WIRE USED WITH THE CATHETER:DRIVEWIRE24. MANUFACTURER: RAPID MEDICAL LOT#: 250301DW01, REF#: DRPP2411, EXP: 01.31.2026. Manufacturer Narrative: AS NO PRODUCT HAS BEEN RETURNED FOR INVESTIGATION, WE ARE UNABLE TO VERIFY THE SPECIFIC FAILURE MODE AND THUS CANNOT DRAW CONCLUSIONS REGARDING THE ROOT CAUSE. HOWEVER, A REVIEW OF MANUFACTURING DOCUMENTATION FOR LOT ASP00572 CONFIRMED THAT THE LOT WAS MANUFACTURED, INSPECTED, AND RELEASED IN FULL COMPLIANCE WITH ALL REQUIRED SPECIFICATIONS. NO DESIGN OR LABELING MODIFICATIONS THAT COULD IMPACT THE DEVICE'S SAFETY HAVE BEEN IMPLEMENTED, AND TO DATE, NO OTHER COMPLAINTS OR ADVERSE EVENTS RELATED TO THIS LOT HAVE BEEN REPORTED. NO OTHER COMPLAINTS OR ADVERSE EVENTS WITH A SIMILAR FAILURE OR PATIENT IMPACT HAVE BEEN REPORTED FOR ANY OTHER ESPERANCE 3F+ PRODUCT. THE DEVICE LABELING INCLUDES APPROPRIATE WARNINGS AND USER INSTRUCTIONS FOR PROPER OPERATION. ADDITIONALLY, THE RISK ASSOCIATED WITH THE REPORTED EVENT IS DOCUMENTED IN THE RISK MANAGEMENT FILES, AND CURRENT CONTROL MEASURES ARE IN PLACE TO MITIGATE SUCH RISKS. FURTHERMORE, THE COMMUNICATION WITH PHYSICIAN INDICATED THAT THE EVENT WAS MORE LIKELY CAUSED BY THE DRIVE WIRE "I DID NOT TELL YOU CAUSE IT WAS THE DRIVE WIRE THAT WAS THE ISSUE. THE WIRE GOT STUCK, IT IS THE SECOND TIME IT HAS HAPPENED TO ME. IT IS A PROBLEM WITH THE WIRE. THE ACTUATOR BROKE AND IT DAMAGED AND BROKE THE CATHETER. I THEN WENT UP AND USED A STENT RETRIEVER TO PULL OUT THE BROKEN TIP".
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22724170&pc=OZP | 2032227-2025-228431 | 2025-07-10 04:00 | Death | MEDTRONIC MINIMED | 2025-08-06 04:00 | OZP | UNKNOWN NGP | Appropriate Term/Code Not Available | Insufficient Information |  |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY ON AN UNKNOWN DATE. THE CUSTOMER'S DAUGHTER REPORTED AN INSULIN PUMP MALFUNCTION CAUSED THE CUSTOMER'S DEATH. NO OTHER DETAILS WERE PROVIDED. THE EVENT INVOLVED PRODUCT(S) UNK_NGP, MMT-332A & UNOMEDICAL. TROUBLESHOOTING WAS NOT DONE. IT WAS UNKNOWN IF THE INSULIN PUMP WAS USED WITHIN 48 HOURS OF THE REPORTED EVENT OR IF THE AUTOMODE/SMART GUARD FEATURE WAS ACTIVE. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED. THE PRODUCT(S) UNK_NGP, MMT-332A & UNOMEDICAL WILL NOT BE RETURNED FOR ANALYSIS. Manufacturer Narrative: CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22724256&pc=QZI | 2124215-2025-54630 | 2025-07-07 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-08-06 04:00 | QZI | FARAWAVE | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING A PULSE FIELD ABLATION (PFA) PROCEDURE USING A FARAWAVE PFA CATHETER THE PATIENT PASSED AWAY. THE PFA PROCEDURE WAS CONCOMITANT WITH A WATCHMAN LEFT ATRIAL APPENDAGE IMPLANT. A SMALL EFFUSION WAS IDENTIFIED PRIOR TO THE PROCEDURE AND DIDN'T CHANGE THROUGHOUT. THE PROCEDURE WENT SMOOTHLY, AND NO COMPLICATIONS WERE NOTED AT THAT TIME OR DURING THE LIMITED PRE-DISCHARGE ECHO. THE PATIENT WAS DISCHARGED AS EXPECTED FOR THE PROCEDURE. THE FOLLOWING MORNING THE PATIENT WAS FOUND DECEASED IN THEIR HOME. THE CAUSE OF DEATH IS NOT YET KNOWN. THE CATHETER WAS DISPOSED BY THE FACILITY AND WILL NOT BE RETURNED AS THE DEVICE PERFORMED AS EXPECTED DURING THE PROCEDURE AND THERE WERE NO REPORTS OF ANY PERFORMANCE CONCERNS. Manufacturer Narrative: GOOD FAITH EFFORT ATTEMPTS WERE MADE TO TRY AND RETRIEVE ADDITIONAL DETAILS REGARDING THE REPORTED EVENT, SUCH AS THE CAUSE OF DEATH AND SPECIFIC PRODUCT INFORMATION (UDI), BUT FURTHER INFORMATION WAS UNABLE TO BE OBTAINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22724305&pc=OZP | 2032227-2025-228512 | 2025-07-10 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-08-06 04:00 | OZP | PUMP MMT-1780KPK 670G PATHWAY BLACK MG | Adverse Event Without Identified Device or Use Problem | Hyperglycemia; Hypoglycemia; Malaise | P150001 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY ABOUT 3 YEARS AGO (UNABLE TO VERIFY INCIDENT DATE). THE CAUSE OF DEATH WAS HYPERGLYCEMIA, HYPOGLYCEMIA AND THE CUSTOMER WAS SICK. IT WAS UNKNOWN WHETHER CUSTOMER INSULIN PUMP WAS WORN AT THE TIME OF PASSING. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: MMT-332A, MMT-1780KPK, UNOMEDICAL. TROUBLESHOOTING WAS NOT PERFORMED FOR DECEASED REPORTING. IT WAS UNKNOWN IF THE CUSTOMER HAD USED THE INSULIN PUMP SYSTEM WITHIN 48 HOURS OF THE REPORTED EVENT. IT WAS UNKNOWN WHETHER THE CUSTOMER WAS USING THE SMARTGUARD AUTO MODE OF THE INSULIN PUMP. NO PRODUCT RETURN IS REQUIRED FOR MMT-332A. PRODUCT RETURN REQUESTED FOR MMT-1780KPK. CUSTOMER DECLINED TO RETURN OR IS UNABLE TO RETURN. NO PRODUCT RETURN IS REQUIRED FOR UNOMEDICAL. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22724750&pc=LJT | 3006260740-2025-05587 | 2025-07-14 04:00 | Death | C.R. BARD, INC. (BASD) -3006260740 | 2025-08-06 04:00 | LJT | M.R.I. IMPLANTABLE PORT, HICKMAN SINGLE-LUMEN, 9.6F | Fracture | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K873213 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE LITIGATION PROCESS THAT SOMETIME AFTER PORT PLACEMENT PROCEDURE, THE CATHETER HAD ALLEGEDLY FRACTURED. REPORTEDLY, THE PATIENT EXPIRED, AND THE CAUSE OF DEATH WAS NOT PROVIDED. Manufacturer Narrative: H11: DATE OF DEATH FOR THIS PATIENT WAS NOT PROVIDED, DATE OF DEATH UPDATED AS (B)(6) 1900 FOR THE MDR SUBMISSION REQUIREMENT. AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22724801&pc=DSP | 2248146-2025-0000459 | 2025-07-20 04:00 | Death | DATASCOPE CORP. | 2025-08-06 04:00 | DSP | SENSATION 7FR. 34CC IAB | Gas/Air Leak | Tachycardia; Cough | K063525 |  | 1 | Event Description: IT WAS REPORTED DURING INTRA-AORTIC BALLOON THERAPY THAT THE PATIENT ON CARDIOSAVE INTRA-AORTIC BALLOON PUMP (IABP) HAD PASSED AWAY. DURING USE, THE PUMP WENT INTO STANDBY MODE. THE ICU RN, INITIALLY COULD NOT IDENTIFY THE ALARM BUT LATER REPORTED A RECENT GAS LOSS ALARM. SHE PERFORMED TROUBLESHOOTING BY PRESSING THE START AND IAB FILL BUTTONS, AND THERAPY WAS RESTARTED. NO BLOOD, DISCOLORATION, OR PARTICLES WERE OBSERVED IN THE HELIUM TUBING, AND ALL CONNECTIONS WERE SECURE. A SCREENSHOT OF THE IABP MONITOR SHOWED APPROPRIATE WAVEFORMS AND BLOOD PRESSURE INDICES. THE PATIENT WAS REPORTEDLY ON A VENTILATOR, COUGHING DURING SUCTIONING, AND IN SINUS TACHYCARDIA, WHICH WERE CONSIDERED CONTRIBUTING CLINICAL FACTORS TO THE ALARM. Manufacturer Narrative: A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF OUR INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22725245&pc=OZD | 1220648-2025-31041 | 2025-07-23 04:00 | Death | ABIOMED, INC. | 2025-08-06 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT A PATIENT ADMITTED IN WITH POSTCARDIOTOMY CARDIOGENIC SHOCK HAD AN IMPELLA CP AND IMPELLA RP FLEX PLACED FOR BIPELLA SUPPORT. THE PATIENT WAS ESCALATED FROM AN INTRA-AORTIC BALLOON PUMP. THE IMPELLA CP WAS SUPPORTING WHEN THERE WAS LIMB ISCHEMIA NOTED ON DAY TWO. THE LEG WAS REVASCULARIZED IN THE CARDIAC CATHETERIZATION LABORATORY. THE PUMP REMAINED ON RATHER THAN ESCALATING AND EXPLANTING DUE TO THE PATIENTS CRITICAL NATURE AND DISCUSSIONS ONGOING WITH THE FAMILY. THE TEAM WITHDREW CARE AND PATIENT EXPIRED AFTER 61 HOURS. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22725419&pc=OZD | 1220648-2025-31043 | 2025-07-20 04:00 | Death | ABIOMED, INC. | 2025-08-06 04:00 | OZD | IMPELLA 5.5 | Device Sensing Problem | Arrhythmia; Septic Shock | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT THE PLACEMENT SIGNAL BECAME ERRATIC DURING IMPELLA 5.5 SUPPORT. PLACEMENT MONITORING WAS DISABLED AS A RESULT OF PERSISTENT IMPELLA IN AORTA ALARMS DESPITE CONFIRMED PROPER POSITIONING. BOTH THE FLOW RATE AND MOTOR CURRENT REMAINED WITHIN THE EXPECTED LIMITS AT THE TIME OF THE ALARM. THE FOLLOWING DAY, THE PATIENT ENDURED RUNS OF VENTRICULAR TACHYCARDIA. SUPPORT CONTINUED UNINTERRUPTED DESPITE THE NOTED CONDITION. THE PATIENT WAS SHOWING SIGNS OF SEPTIC SHOCK. IT IS UNKNOWN IF THE SEPSIS WAS IMPELLA RELATED. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK: SECTION: TABLE 3.2 IMPELLA CATHETER COMPONENTS: "THE INFUSION FILTER PREVENTS BACTERIAL CONTAMINATION AND AIR FROM ENTERING THE PURGE LUMEN." SECTION: WARNINGS & CAUTIONS: WARNINGS: "TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22725765&pc=OZD | 1220648-2025-31049 | 2025-07-20 04:00 | Death | ABIOMED, INC. | 2025-08-06 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT THE PATIENT ARRIVED AT THE HOSPITAL AFTER MULTIPLE CODES. THE PATIENT WAS BROUGHT TO CARDIAC CATHETERIZATION LABORATORY AND IMPELLA WAS PLACED. PERCUTANEOUS CORONARY INTERVENTION WHILE CODE COMPLETED. THE GROIN WAS BLEEDING FROM THE IMPELLA ACCESS SITE. THE IMPELLA WAS REMOVED, AND A COVERED STENT WAS PLACED. THE PATIENT WAS SENT TO THE UNIT WHERE THE PATIENT EXPIRED SOON AFTER ARRIVAL. Manufacturer Narrative: THE IMPELLA DEVICE HAS NOT BEEN RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMART ASSIST SYSTEM STATES THE FOLLOWING: SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT. USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22725952&pc=LWS | 3004209178-2025-13294 | 2025-08-01 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-08-06 04:00 | LWS | EVERA¿ XT VR | Failure to Charge; Failure to Deliver Shock/Stimulation; Defibrillation/Stimulation Problem; Device Sensing Problem | Tachycardia; Ventricular Fibrillation | P980016 |  | 1 | Event Description: IT WAS REPORTED THAT IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) INITIALLY INAPPROPRIATELY WITHHELD THERAPY FOR VENTRICULAR TACHYCARDIA (VT). DETECTION FOR VT WAS MET AND ANTI-TACHYCARDIA PACING (ATP) WAS DELIVERED, BUT THE VT ACCELERATED TO VENTRICULAR FIBRILLATION (VF). THE ICD TRIGGERED SHORT CIRCUIT PROTECTION DURING THE SECOND HIGH VOLTAGE THERAPY AND LESS THAN PROGRAMMED HIGH VOLTAGE ENERGY WAS DELIVERED FOR VF ARREST. A THIRD SHOCK WAS NEVER DELIVERED FOR VF DUE TO A CHARGE CIRCUIT TIME OUT. THE PATIENT PASSED AWAY. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE DEVICE WAS NOT RETURNED FOR ANALYSIS, HOWEVER, PERFORMANCE DATA COLLECTED FROM THE DEVICE WAS RECEIVED AND ANALYZED. ANALYSIS OF THE DEVICE MEMORY INDICATED THAT THE VENTRICULAR TACHYARRHYTHMIA THERAPY ENERGY WAS LOW DUE TO AN UNINTENDED CURRENT PATHWAY. ANALYSIS OF THE DEVICE MEMORY INDICATED CHARGE CIRCUIT TIMEOUT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22726399&pc=DSQ | 3007042319-2025-01858 | 2025-07-26 04:00 | Death | HEARTWARE, INC. | 2025-08-06 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP/DRIVELINE | Electrical /Electronic Property Problem; Pumping Problem | Intracranial Hemorrhage; Tachycardia; Loss of consciousness | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS ADMITTED FROM THE EMERGENCY ROOM DUE TO THE VENTRICULAR ASSIST DEVICE (VAD) EXHIBITING HIGH POWER ASSOCIATED WITH HIGH WATT ALARMS AND AN ELECTRICAL FAULT ALARM. IT WAS NOTED THAT THE PATIENT REMAINED HEMODYNAMICALLY STABLE UNTIL THE FOLLOWING DAY WHEN THE PATIENT STARTED HAVING RUNS OF VENTRICULAR TACHYCARDIA (VT) WHICH CAUSED THEIR BLOOD PRESSURE TO DROP. THERE WERE INITIALLY CONCERNS FOR POSSIBLE THROMBOSIS DUE TO THE HIGH WATT ALARMS, BUT THE PATIENTS' LACTATE DEHYDROGENASE (LDH) WAS STABLE AT 220 AND THERE WERE NO OTHER SIGNS OF THROMBOSIS. LOG FILES WERE REVIEWED AND IT WAS CONFIRMED THAT THE VAD WAS RUNNING ON ONE MOTOR STATOR DUE TO A SHORT IN THE REAR MOTOR STATOR. THE DRIVELINE WAS CHECKED FOR DAMAGE AND TO MAKE SURE IT WAS FULLY SEATED AND ENGAGED IN THE CONTROLLER. AS THE ENGINEER AT THE HOSPITAL WAS REMOVING TAPE COVERING THE DRIVELINE TO ASSESS FOR DAMAGE, IT WAS NOTED THAT THE PATIENT WASN'T AS TALKATIVE AS NORMAL. THE PATIENT ASKED FOR WATER, AND WHEN THE ENGINEER RETURNED WITH THE WATER, THE PATIENT WASN'T RESPONDING. THE PATIENT WAS RUSHED FOR A COMPUTERIZED TOMOGRAPHY (CT) SCAN WHERE IT WAS DETERMINED THAT THE PA TIENT HAD HAD A MASSIVE BRAIN BLEED. THE PATIENT WAS NOT A CANDIDATE FOR SURGERY AND SUBSEQUENTLY PASSED AWAY. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22726563&pc=MOM | 3004936110-2025-01517 | 2025-07-25 04:00 | Death | ABBOTT MEDICAL | 2025-08-06 04:00 | MOM | HF SENSOR DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Hemoptysis; Low Oxygen Saturation; Cough | P100045 |  | 1 | Event Description: DURING THE CARDIOMEMS IMPLANT PROCEDURE, THE PHYSICIAN WAS ATTEMPTING TO INTRODUCE THE 0.18 PLATINUM PLUS GUIDEWIRE INTO THE IMPLANT TARGET IMPLANT SITE WHEN THE PATIENT COUGHED AND NOTED HEMOPTYSIS. THE AIRWAY WAS SUCTIONED AND PER THE PHYSICIAN A REPEAT ANGIOGRAM WAS DONE, AND A PERFORATION WAS NOTED. THE PATIENT WENT INTO CARDIAC ARREST, CPR WAS STARTED, RETURN OF SPONTANEOUS CIRCULATION, AND THEN THE PATIENT WAS CANNULATED FOR ECMO. THE PATIENT PASSED AWAY LATER THAT EVENING IN THE CVICU. PER THE PHYSICIAN, THE PERFORATION WAS CAUSED BY THE 0.18 PLATINUM PLUS GUIDEWIRE, NOT THE CARDIOMEMS DELIVERY SYSTEM. THE CARDIOMEMS SENSOR WAS NOT USED OR IMPLANTED. PENDING REQUESTED INFORMATION FROM THE CLINIC. Manufacturer Narrative: THE INVESTIGATION CODES ALONG WITH INVESTIGATION RESULTS WILL BE PROVIDED IN A SUBSEQUENT SUBMISSION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22726642&pc=MCX | 3004742232-2025-00222 | 2025-07-03 04:00 | Death | CARDIOVASCULAR SYSTEMS INCORPORATED | 2025-08-06 04:00 | MCX | DIAMONDBACK CORONARY | Unexpected Shutdown | Myocardial Infarction; Perforation of Vessels; Cardiac Tamponade; Vascular Dissection | P130005 |  | 1 | Event Description: IT WAS REPORTED THE DIAMONDBACK 360 CORONARY ORBITAL ATHERECTOMY DEVICE (OAD) WAS BEING USED TO TREAT A HEAVILY CALCIFIED OSTIAL LESION IN A VERY TORTUOUS CIRCUMFLEX. TREATMENT WAS BEING PERFORMED IN AN ANTEGRADE DIRECTION WHEN THE OAD STALLED. THE PHYSICIAN BACKED THE OAD OUT OF THE LESION AND ATTEMPTED TO TREAT ANTEGRADE AGAIN. DURING THIS ATTEMPTED TREATMENT, A SIGNIFICANT DISSECTION WAS OBSERVED AT THE PROXIMAL CIRCUMFLEX BACKING INTO THE LEFT MAIN ARTERY. THE OAD WAS REMOVED AND A COVERED STENT WAS PLACED AT THE DISSECTION SITE. THE PATIENT WAS SENT FOR BYPASS SURGERY, HOWEVER, THE PATIENT EXPIRED. IN THE PHYSICIAN'S OPINION, ORBITAL ATHERECTOMY CAUSED THE DISSECTION WITH EXTREME ANGULATION AND CALCIFICATION OF THE VESSEL CONTRIBUTING. IT WAS INDICATED CARDIAC TAMPONADE AND/OR MYOCARDIAL INFARCTION WAS THE PRIMARY CAUSE OF DEATH, WITH CORONARY PERFORATION THE SECONDARY CAUSE OF THE DEATH. THE CAUSE OF THE OAD STALL WAS DUE TO THE DEVICE BEING ANGLED GREATER THAN 90 DEGREES WHILE TREATING.Event Description: IT WAS REPORTED THE DIAMONDBACK 360 CORONARY ORBITAL ATHERECTOMY DEVICE (OAD) WAS BEING USED TO TREAT A HEAVILY CALCIFIED OSTIAL LESION IN A VERY TORTUOUS CIRCUMFLEX. TREATMENT WAS BEING PERFORMED IN AN ANTEGRADE DIRECTION WHEN THE OAD STALLED. THE PHYSICIAN BACKED THE OAD OUT OF THE LESION AND ATTEMPTED TO TREAT ANTEGRADE AGAIN. DURING THIS ATTEMPTED TREATMENT, A SIGNIFICANT DISSECTION WAS OBSERVED AT THE PROXIMAL CIRCUMFLEX BACKING INTO THE LEFT MAIN ARTERY. THE OAD WAS REMOVED AND A COVERED STENT WAS PLACED AT THE DISSECTION SITE. THE PATIENT WAS SENT FOR BYPASS SURGERY, HOWEVER, THE PATIENT EXPIRED. NO ADDITIONAL INFORMATION WAS PROVIDED. ADDITIONAL INFORMATION WAS RECEIVED INDICATING IN THE PHYSICIAN'S OPINION, ORBITAL ATHERECTOMY CAUSED THE DISSECTION WITH EXTREME ANGULATION AND CALCIFICATION OF THE VESSEL CONTRIBUTING. IT WAS INDICATED CARDIAC TAMPONADE AND/OR MYOCARDIAL INFARCTION WAS THE PRIMARY CAUSE OF DEATH, WITH CORONARY PERFORATION THE SECONDARY CAUSE OF THE DEATH. THE CAUSE OF THE OAD STALL WAS DUE TO THE DEVICE BEING ANGLED GREATER THAN 90 DEGREES WHILE TREATING.Event Description: IT WAS REPORTED THE DIAMONDBACK 360 CORONARY ORBITAL ATHERECTOMY DEVICE (OAD) WAS BEING USED TO TREAT A HEAVILY CALCIFIED OSTIAL LESION IN A VERY TORTUOUS CIRCUMFLEX. TREATMENT WAS BEING PERFORMED IN AN ANTEGRADE DIRECTION WHEN THE OAD STALLED. THE PHYSICIAN BACKED THE OAD OUT OF THE LESION AND ATTEMPTED TO TREAT ANTEGRADE AGAIN. DURING THIS ATTEMPTED TREATMENT, A SIGNIFICANT DISSECTION WAS OBSERVED AT THE PROXIMAL CIRCUMFLEX BACKING INTO THE LEFT MAIN ARTERY. THE OAD WAS REMOVED AND A COVERED STENT WAS PLACED AT THE DISSECTION SITE. THE PATIENT WAS SENT FOR BYPASS SURGERY, HOWEVER, THE PATIENT EXPIRED. NO ADDITIONAL INFORMATION WAS PROVIDED. ADDITIONAL INFORMATION WAS RECEIVED INDICATING IN THE PHYSICIAN'S OPINION, ORBITAL ATHERECTOMY CAUSED THE DISSECTION WITH EXTREME ANGULATION AND CALCIFICATION OF THE VESSEL CONTRIBUTING. IT WAS INDICATED CARDIAC TAMPONADE AND/OR MYOCARDIAL INFARCTION WAS THE PRIMARY CAUSE OF DEATH, WITH CORONARY PERFORATION THE SECONDARY CAUSE OF THE DEATH. THE CAUSE OF THE OAD STALL WAS DUE TO THE DEVICE BEING ANGLED GREATER THAN 90 DEGREES WHILE TREATING. Manufacturer Narrative: CORRECTION: H6 (CODES 3221 AND 4315 REMOVED). Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. PRODUCTION RECORD AND CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A QUERY OF THE COMPLAINT HANDLING DATABASE FOR THE REPORTED LOT REVEALED THERE IS NO INDICATION OF A LOT SPECIFIC ISSUE. BASED ON THE INFORMATION RECEIVED, THE INVESTIGATION DETERMINED THAT THE REPORTED UNEXPECTED SHUTDOWN - STALL, CARDIAC TAMPONADE, MYOCARDIAL INFARCTION, PERFORATION OF VESSELS, VASCULAR DISSECTION, PATIENT DEATH, SURGICAL INTERVENTION, AND UNEXPECTED MEDICAL INTERVENTION APPEARS TO BE RELATED TO OPERATIONAL CONTEXT. IN THIS CASE IT IS POSSIBLE THAT THE REPORTED UNEXPECTED SHUTDOWN - STALL, CARDIAC TAMPONADE, MYOCARDIAL INFARCTION, PERFORATION OF VESSELS, VASCULAR DISSECTION, PATIENT DEATH, SURGICAL INTERVENTION, AND UNEXPECTED MEDICAL INTERVENTION ARE A RESULT OF THE PATIENT¿S DISEASE SEVERITY AND/OR USE TECHNIQUES EMPLOYED; HOWEVER, SINCE THE DEVICE WAS NOT RETURNED THIS COULD NOT BE CONFIRMED. BASED ON THE RESULTS OF THE COMPLAINT INVESTIGATION, THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO THE DESIGN, MANUFACTURE, OR LABELING OF THE DEVICE. UPDATED: B5, H6 (CODES 4118, 3233, AND 11 NO LONGER APPLY). Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22726697&pc=QZI | 2124215-2025-54577 | 2025-07-15 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-08-06 04:00 | QZI | FARAWAVE | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Ventricular Fibrillation |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED VENTRICULAR TACHYCARDIA (VT) AND VENTRICULAR FIBRILLATION (VF). THE PATIENT DIED. DURING A PULSE FIELD ABLATION (PFA) PROCEDURE TO TREAT ATRIAL FIBRILLATION A FARAWAVE CATHETER WAS USED. THE PATIENT HAD A BOSTON SCIENTIFIC IMPLANTABLE CARDIOVERTER-DEFIBRILLATOR (ICD). THERE WAS NO EFFUSION PRE-PROCEDURE. PULMONARY VEIN ISOLATION (PVI) WAS CONDUCTED NORMALLY AND COMPLETED WITH THE FARAWAVE CATHETER, AND IT WAS REMOVED FROM THE PATIENT AFTERWARDS. POST-PVI THERE WAS NO EFFUSION, AND NO POST ELECTROPHYSIOLOGY (EP) STUDY WAS PERFORMED. THE INTRACARDIAC ECHOCARDIOGRAPHY (ICE) CATHETER WAS REMOVED FOLLOWED BY THE NON-BOSTON SCIENTIFIC RIGHT VENTRICLE (RV) CATHETER. UPON REMOVING THE RV CATHETER, THE PATIENT HAD A SHORT BURST OF VT THAT QUICKLY DEGRADED INTO VF. THE PATIENT RECEIVED MULTIPLE SHOCKS FROM THE ICD AND FROM EXTERNAL CARDIOVERSION; THE MEDICAL STAFF WATCHED THE ELECTROGRAMS (EKGS) TO MAKE SURE AN EXTERNAL SHOCK WAS NOT GIVEN AT THE SAME TIME AS AN ICD SHOCK. THE SHOCKS FAILED TO RESOLVE THE ARRYTHMIAS; THE PATIENT CONTINUED TO GO BACK INTO VT/VF AS SOON AS THEY CAME OUT OF IT. CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED AND ADVANCED CARDIAC LIFE SUPPORT (ACLS) MEDICATIONS WERE ADMINISTERED. THE PATIENT EVENTUALLY CAME BACK AND WAS TRANSFERRED TO THE INTENSIVE CARE UNIT (ICU) BUT CONTINUED TO CODE. A DO NOT RESUSCITATE (DNR) ORDER WAS GIVEN BY THE PATIENT'S FAMILY, AND THE PATIENT DIED IN THE NIGHT. IT WAS NOTED THAT THE VT WAS FROM A PREVIOUS CHRONIC CONDITION OF THE PATIENT, HAVING HAD MULTIPLE CARDIAC PROCEDURES SINCE 1998; THE PHYSICIAN AND STAFF STATED THAT "THE PATIENT'S HEART JUST GAVE OUT." THE FARAWAVE CATHETER IS NOT EXPECTED TO BE RETURNED FOR ANALYSIS DUE TO DISPOSAL. Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION, BUT IT WAS UNABLE TO BE OBTAINED. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22726756&pc=OZD | 1220648-2025-31053 | 2025-07-21 04:00 | Death | ABIOMED, INC. | 2025-08-06 04:00 | OZD | IMPELLA CP | Positioning Failure | Cardiovascular Insufficiency | P140003 |  | 1 | Event Description: THE COMPLAINANT HAD A IMPELLA CP PLACED TO SUPPORT A PATIENT ADMITTED IN FROM THE COMMUNITY TO THE TERTIARY CENTER. CARE WAS ESCALATED FROM AN INTRA-AORTIC BALLOON PUMP TO AN IMPELLA CP BUT THE PUMP CAME OUT OF POSITION AND COULD NOT BE REDELIVERED AND SO IT WAS EXPLANTED AND REPLACED. THE PATIENT EXPIRED ON THE SECOND PUMP. HOWEVER, WE CANNOT DISASSOCIATE THE PATIENTS OUTCOME FROM THE FIRST PUMP. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22726960&pc=NPT | 2015691-2025-06368 | 2025-07-09 04:00 | Death | EDWARDS LIFESCIENCES | 2025-08-06 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA VALVE, 29MM | Difficult to Insert | Rupture | P140031 |  | 1 | Event Description: AS REPORTED BY THE EDWARDS FIELD CLINICAL SPECIALIST, THE PATIENT PRESENTED WITH SEVERE AORTIC STENOSIS AND EXTENSIVE CALCIFICATION, CLASSIFIED AS HIGH RISK. INITIAL PRE-DILATION OF THE HEAVILY CALCIFIED AORTIC ANNULUS WAS PERFORMED USING A 16 MM Z-MED BALLOON; HOWEVER, THE DELIVERY SYSTEM WAS UNABLE TO CROSS THE VALVE AND WAS SUBSEQUENTLY REMOVED. A SECOND PRE-DILATION WAS PERFORMED USING A 23 MM Z-MED BALLOON, ALLOWING SUCCESSFUL INSERTION AND DEPLOYMENT OF A 29 MM SAPIEN 3 ULTRA RESILIA VALVE. POST-DEPLOYMENT, AN EFFUSION NEAR THE AORTIC VALVE AND AORTIC DISSECTION WERE NOTED. HEMODYNAMIC STABILITY WAS MAINTAINED, AND SURGICAL REPAIR WAS INITIATED. THE PATIENT WAS ALIVE AT THE END OF THE PROCEDURE. NO DEVICE MALFUNCTION OR PRODUCT-RELATED ALLEGATION WAS REPORTED. THE PERCEIVED ROOT CAUSE WAS ATTRIBUTED TO PATIENT ANATOMY. THE PERCEIVED ROOT CAUSE WAS VALVE CROSSING WAS UNSUCCESSFUL DESPITE INITIAL PRE-DILATION WITH Z-MED BALLOON.Event Description: UPDATED INFORMATION REVEALED THAT DURING VALVE DEPLOYMENT THE INJURY OCCURRED. THE DEVICE WAS DISCARDED AT THE HOSPITAL. THERE WAS NO DAMAGE TO ANY OF THE EDWARDS DEVICES. Manufacturer Narrative: THE INVESTIGATION IS ONGOING. Manufacturer Narrative: UPDATE TO A2 (AGE), B2 (OUTCOME ATTRIBUTING TO ADVERSE EVENT), B5 (DESCRIBE EVENT OR PROBLEM), B7 ( OTHER RELEVANT HISTORY, INCLUDING PREEXISTING MEDICAL CONDITIONS), D1 (BRAND NAME), D4 (SERIAL NUMBER, EXPIRATION DATE, AND UDI#), H1 (TYPE OF REPORTABLE EVENT), H2 (IF FOLLOW UP, WHAT TYPE), H4 (DEVICE MANUFACTURER DATE), AND H6 (COMPONENT CODE, IMPACT AND CLINICAL), H11 (PROVIDE NARRATIVE /DATA) BASED ON RECEIVED INFORMATION. ADDITIONAL INFORMATION OBTAINED THROUGH EDWARDS FIELD CLINICAL SPECIALIST CLARIFIED THAT THE DISSECTION OCCURRED DURING VALVE DEPLOYMENT. AS SUCH THIS EVENT IS NOW BEING REPORTED AGAINST THE 29MM SAPIEN 3 ULTRA RESILIA VALVE. IT WAS REPORTED THAT THE PATIENT EXPIRED. INVESTIGATION OF THIS EVENT IS ONGOING PENDING ENGINEERING INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22726982&pc=OZD | 1220648-2025-31055 | 2025-07-09 04:00 | Death | ABIOMED, INC. | 2025-08-06 04:00 | OZD | IMPELLA 5.5 | Use of Device Problem; Failure to Advance; Patient-Device Incompatibility |  | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THE INSERTION OF IMPELLA 5.5 DEVICE TO A PATIENT IN POSTCARDIOTOMY CARDIOGENIC SHOCK WAS UNSUCCESSFUL. THE PATIENT UNDERWENT PLANNED CORONARY ARTERY BYPASS GRAFT SURGERY (X2 VESSELS) AND WAS WEANED INITIALLY OFF BYPASS. HOWEVER, WHILE CLOSING CHEST THE PATIENT CONTINUED TO DECOMPENSATE AND WAS PLACED BACK ON AND OFF PUMP FOUR TIMES BEFORE DECISION MADE TO PLACE AN INTRA-AORTIC BALLOON PUMP. THE PHYSICIAN AGAIN ATTEMPTED TO WEAN OFF BYPASS AND AGAIN THE PATIENT CONTINUED TO DECOMPENSATE WITH ST CHANGES NOTED ON EKG. THE DECISION WAS THEN MADE TO PLACE AN IMPELLA 5.5 DEVICE VIA DIRECT AORTA APPROACH. THE 10MM GRAFT WAS SEWN AND TUNNELED RIGHT SUPRACLAVICULAR INTRODUCER/GRAFT LOCKS IN PLACE. A 0.035 GUIDEWIRE AND AL1 WERE USED TO CROSS AORTIC VALVE UNDER FLUOROSCOPY AND THEN THE GUIDEWIRE WAS EXCHANGED FOR AN 0.018 GUIDEWIRE. THE IMPELLA WIRE BACKLOADED ONTO IMPELLA 5.5 DEVICE AND ATTEMPTED TO PLACE INTO LEFT VENTRICLE; HOWEVER, THE DEVICE WAS UNABLE TO PASS THE ANASTOMOSIS. A DECISION MADE TO EXCHANGE WIRE. THE IMPELLA WAS, THEREFORE, REMOVED AND WHILE ATTEMPTING TO EXCHANGE WIRES THE DEVICE BECAME UNSTERILE. ANOTHER IMPELLA 5.5 DEVICE WAS USED TO SUPPORT THE PATIENT. HOWEVER, AFTER APPROXIMATELY 14 HOURS, THE PATIENT EXPIRED. NO FURTHER DETAILS WERE NOTED. Manufacturer Narrative: THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE INITIALLY ATTEMPTED IMPELLA 5.5 DEVICE AS AN ASSOCIATED FACTOR TO THE OUTCOME GIVEN THE DELAY IN THERAPY FOR THE DECOMPENSATING PATIENT. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, AN EVALUATION/ANALYSIS OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) - ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS & CAUTIONS: WARNINGS - ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION. [ADDITION FOR 5.0, 5.5 ONLY] IN INSTANCES WHERE THE IMPELLA PUMP HAS BEEN PLACED PRIOR TO PERFORMING CARDIAC SURGERY WITH AORTIC CROSS CLAMPING AND CARDIOPLEGIC ARREST, CARE SHOULD BE TAKEN WHEN MANIPULATING THE HEART WHEN THE PUMP POSITION IS FIXED WITH APPLICATION OF THE AORTIC CROSS CLAMP ACROSS THE PUMP CATHETER.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22727231&pc=OZD | 1220648-2025-31056 | 2025-07-23 04:00 | Death | ABIOMED, INC. | 2025-08-06 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Cardiac Perforation | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT DURING PLACEMENT OF AN IMPELLA 5.5 IN THE PATIENT FOR CARDIOGENIC SHOCK BAILOUT POST TRANSCATHETER AORTIC VALVE REPLACEMENT, A LEFT VENTRICULAR PERFORATION OCCURRED. DUE TO PROLONGED LEFT MAIN (CORONARY ARTERY) ISCHEMIA AND LEFT VENTRICLE (LV) FRIABILITY, A DECISION WAS MADE NOT TO REPAIR AND TO CEASE LIFESAVING EFFORTS. THEREFORE, THE PATIENT EXPIRED. ADDITIONAL INFORMATION WAS RECEIVED. THE PHYSICIAN FEELS THAT THE LV PERFORATION WAS FROM EXTENSIVE CARDIAC MASSAGE PERFORMED AT VARIOUS TIMES AFTER THE CHEST WAS OPENED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿ SECTION: WARNINGS & CAUTIONS: WARNINGS SECTION: PRE-SUPPORT EVALUATION SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER SECTION: DIRECT AORTIC INSERTION SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION. [ADDITION FOR 5.0, 5.5 ONLY] IN INSTANCES WHERE THE IMPELLA PUMP HAS BEEN PLACED PRIOR TO PERFORMING CARDIAC SURGERY WITH AORTIC CROSS CLAMPING AND CARDIOPLEGIC ARREST, CARE SHOULD BE TAKEN WHEN MANIPULATING THE HEART WHEN THE PUMP POSITION IS FIXED WITH APPLICATION OF THE AORTIC CROSS CLAMP ACROSS THE PUMP CATHETER.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE. ¿TO REDUCE THE RISK OF CARDIAC INJURY (INCLUDING VENTRICULAR PERFORATION), PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY IN THE SETTINGS OF ACUTE INFARCTION WITH TISSUE NECROSIS.¿ ¿TO REDUCE THE RISK OF VASCULAR INJURY, PHYSICIANS SHOULD EXERCISE CAUTION WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX PERIPHERAL VASCULAR ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: UNREPAIRED ABDOMINAL AORTIC ANEURYSM, SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM, DISSECTION OF THE ASCENDING/ TRANSVERSE/DESCENDING AORTA, CHRONIC ANATOMICAL CHANGES IN THE RELATIONSHIP OF THE AORTA/AORTIC VALVE/VENTRICULAR ALIGNMENT, SIGNIFICANT MOBILE ATHEROMATOUS DISEASE IN THE THORACIC OR ABDOMINAL AORTA OR PERIPHERAL VESSELS.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING A RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY IN THE LEFT VENTRICLE AFTER CPR WITH ECHOCARDIOGRAPHY GUIDANCE. SECTION: WARNINGS & CAUTIONS: CAUTIONS ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22728269&pc=NPT | 2015691-2025-06376 | 2025-07-14 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-08-06 04:00 | NPT | EDWARDS SAPIEN 3 TRANSCATHETER HEART VALVE | Central Regurgitation | Heart Failure/Congestive Heart Failure | P140031 |  | 1 | Event Description: AS REPORTED BY EDWARDS JAPANESE AFFILIATE, APPROXIMATELY 8 YEARS AND 7 MONTHS POST TAVR, MODERATE TO SEVERE CENTRAL REGURGITATION WAS REPORTED. THE PATIENT UNDERWENT A TRANSFEMORAL TAVR PROCEDURE AND RECEIVED A 26 MM SAPIEN 3 VALVE IN THE NATIVE AORTIC POSITION. APPROXIMATELY 8 YEARS AND 7 MONTHS POST TAVR, MODERATE TO SEVERE CENTRAL REGURGITATION WAS OBSERVED INSIDE OF THE LEAFLET OF NON-CORONARY CUSP (NCC) SIDE. SYMPTOMS OF HEART FAILURE AND DECLINE IN CARDIAC FUNCTION WERE ALSO NOTED. A TAV IN TAV IS PLANNED AT THIS TIME. PER ADDITIONAL INFORMATION ON (B)(6) 2025, THE PATIENT SUDDENLY DETERIORATED AND PASSED AWAY RECENTLY. TAV IN TAV WAS NOT PERFORMED. THE DETAILS AND THE EXACT DATE OF DEATH WAS NOT PROVIDED. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: INVESTIGATION IS ONGOING. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED DUE TO ENGINEERING EVALUATION FINDINGS, TYPE OF INVESTIGATION, INVESTIGATION FINDINGS, INVESTIGATION CONCLUSIONS AND H11 HAS BEEN ADDED. THE EVENT(S) REPORTED IS/ARE ANTICIPATED IN THE RISK MANAGEMENT DOCUMENTATION FOR TRANSCATHETER HEART VALVE PROCEDURES. ADDITIONAL ASSESSMENT OF THE FAILURE MODE(S) IS NOT REQUIRED AT THIS TIME. THE DEVICE WAS NOT RETURNED FOR EVALUATION. DUE TO THE UNAVAILABILITY OF THE COMPLAINT DEVICE, ENGINEERING WAS UNABLE TO PERFORM ANY VISUAL INSPECTION, FUNCTIONAL TESTING, OR DIMENSIONAL ANALYSIS. THE COMPLAINT FOR CENTRAL REGURGITATION WAS CONFIRMED BY THE INFORMATION PROVIDED FROM FCS (FIELD CLINICAL SPECIALIST). THE REPORTED EVENT DOES NOT ALLEGE A MALFUNCTION THAT COULD BE RELATED TO AN EDWARDS MANUFACTURING DEFICIENCY. A REVIEW OF THE IFU REVEALED NO DEFICIENCIES. DURING THE MANUFACTURING PROCESS, ALL SAPIEN 3 VALVES ARE 100% VISUALLY INSPECTED FOR DEFECTS AND 100% TESTED FOR PROPER COAPTATION UNDER PHYSIOLOGICAL BACKPRESSURE CONDITIONS PRIOR TO RELEASE FOR DISTRIBUTION. THEREFORE, IT IS HIGHLY UNLIKELY THAT A MANUFACTURING DEFECT OR DEVICE MALFUNCTION CONTRIBUTED TO THE EVENT. AS REPORTED, 'THE PATIENT UNDERWENT A TRANSFEMORAL TAVR PROCEDURE AND RECEIVED A 26 MM SAPIEN 3 VALVE IN THE NATIVE AORTIC POSITION. APPROXIMATELY 8 YEARS AND 7 MONTHS POST TAVR, MODERATE TO SEVERE CENTRAL REGURGITATION WAS OBSERVED INSIDE OF THE LEAFLET OF NON-CORONARY CUSP (NCC) SIDE.' PER THE INSTRUCTIONS FOR USE (IFU), VALVE REGURGITATION IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH BIOPROSTHETIC HEART VALVES AND THE TRANSCATHETER VALVE REPLACEMENT (THV) PROCEDURE. REGURGITATION WHICH DEVELOPS PROGRESSIVELY OVER TIME CAN BE DUE TO A NUMBER OF ISSUES INCLUDING PATIENT RELATED FACTORS OR STRUCTURAL VALVE DETERIORATION, INCLUDING CALCIFICATION, NON-CALCIFIC DEGENERATION, LEAFLET THICKENING OR FIBROSIS, OR A COMBINATION OF THESE. REGURGITATION MAY ALSO DEVELOP PROGRESSIVELY IF HOST FIBROTIC TISSUE, OR PANNUS, GROWS ONTO THE BIOPROSTHETIC VALVE. PANNUS, A CAUSE OF NONSTRUCTURAL DYSFUNCTION, MAY INTERFERE WITH FUNCTIONALITY OF THE DEVICE BY RESTRICTING THE LEAFLET MOTION LEADING TO ABNORMAL COAPTATION. DUE TO LIMITED INFORMATION WAS PROVIDED, A DEFINITIVE ROOT CAUSE IS UNABLE TO BE DETERMINED AT THIS TIME. THE EVENT OF CENTRAL REGURGITATION HAS BEEN CONFIRMED BASED ON THE INFORMATION AVAILABLE TO DATE. ANY UPDATES OR ADDITIONAL FINDINGS WILL BE SUBMITTED IN ACCORDANCE WITH FDA REPORTING REQUIREMENTS. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. SINCE NO EDWARDS PRODUCT DEFECTS OR LABELING DEFICIENCIES WERE IDENTIFIED, NO CORRECTIVE OR PREVENTATIVE ACTION IS REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22728549&pc=LWS | 9614453-2025-03108 | 2025-06-26 04:00 | Death | MEDTRONIC EUROPE SARL | 2025-08-06 04:00 | LWS | COBALT¿ XT DR MRI SURESCAN¿ | Appropriate Term/Code Not Available | Myocardial Infarction; Heart Failure/Congestive Heart Failure | P980016 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IS DECEASED. THE PATIENT'S FAMILY MEMBER INDICATED THAT THE PATIENT HAD "AN INCIDENT" AND THE FAMILY MEMBER WAS CONCERNED THAT THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) WAS NOT WORKING SINCE THE PATIENT DIED OF A HEART CONDITION WHEN THEY HAD THE ICD FOR THE HEART CONDITION. THE CAUSE OF DEATH WAS NOTED TO BE MYOCARDIAL INFARCTION AND HEART FAILURE. A REQUEST FOR ADDITIONAL INFORMATION REGARDING DEVICE FUNCTION AT THE TIME OF DEATH WAS REQUESTED BUT WAS NOT RECEIVED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22728710&pc=MKJ | 3030677-2025-001799 | 2025-07-30 04:00 | Death | PHILIPS NORTH AMERICA LLC | 2025-08-06 04:00 | MKJ | PHILIPS HEARTSTART FR2+ | Failure to Deliver Shock/Stimulation | Cardiac Arrest | K051632 |  | 1 | Event Description: IT HAS BEEN REPORTED THAT THE DEVICE DID NOT RECOMMEND A SHOCK ON A POTENTIALLY SHOCKABLE RHYTHM. THE PATIENT DID NOT SURVIVE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22728845&pc=ODG | 3002808148-2025-12317 | 2024-02-13 05:00 | Death | SHIRAKAWA OLYMPUS CO., LTD. | 2025-08-06 04:00 | ODG | EVIS EXERA II ULTRASOUND GASTROVIDEOSCOPE | Adverse Event Without Identified Device or Use Problem | Abdominal Pain; Laceration(s); Nausea; Laceration(s) of Esophagus; Gastrointestinal Hemorrhage; Hematemesis | K093395 |  | 1 | Event Description: IT WAS REPORTED THAT BETWEEN APPROXIMATELY 12:15 P.M. AND 12:45 P.M., A PATIENT WITH A PANCREATIC TUMOR UNDERWENT A THICK-NEEDLE BIOPSY OF THE PANCREATIC TUMOR UNDER ENDOSCOPIC ULTRASOUND GUIDANCE (EUS) UNDER INTRAVENOUS ANESTHESIA. AFTER THE PROCEDURE, THE PATIENT WAS TRANSFERRED TO THE INTERNAL MEDICINE WARD, WHERE SHE WAS MONITORED FOR 2 HOURS AFTER THE PROCEDURE AND REMAINED HEMODYNAMICALLY STABLE. AT APPROXIMATELY 3:00 P.M., THE PATIENT REPORTED ABDOMINAL PAIN AND NAUSEA. AT AROUND 4:00 P.M., THE PATIENT VOMITED BLOOD, AND AT 6:14 P.M., THE PATIENT DIED. THE AUTOPSY REVEALED TWO LONGITUDINAL TEARS OF THE MUCOSA, EACH 5 CM LONG, IN THE ESOPHAGUS IN THE SUPRA-CARDIA REGION AND UPPER HALF OF THE STOMACH, WITH ACUTE HEMORRHAGES IN THE MUCOSA AND SUBMUCOSA, AND THE PRESENCE OF 1100 ML OF CLOTS AND LIQUID BLOOD IN THE STOMACH AND DUODENUM. PANCREATIC CANCER WAS CONFIRMED BY A BIOPSY OF THE PANCREATIC TUMOR UNDER EUS GUIDANCE AND BY POSTMORTEM EXAMINATION. ADDITIONAL INFORMATION RECEIVED FROM THE CUSTOMER INDICATED THAT THE DEVICE WAS CHECKED IN ACCORDANCE WITH THE MANUFACTURER'S INSTRUCTIONS PRIOR TO THE PROCEDURE. THERE WERE NO ABNORMALITIES OBSERVED DURING USE OF THE DEVICE AND DURING CLEANING OR PREPARATION OF THE DEVICE. THERE WAS ALSO NO VISIBLE DAMAGE TO THE DEVICE. THE CAUSE OF DEATH WAS REPORTED TO BE BLEEDING INTO THE UPPER GASTROINTESTINAL TRACT, AND THE CAUSE OF THE BLEEDING WAS TEARS TO THE ESOPHAGEAL MUCOSA. A DAY BEFORE THE PROCEDURE, THE PATIENT HAD PRECUT, SPHINCTEROTOMY, AND IMPLANTATION OF A PARTIALLY COATED SELF-EXPANDING COOK EVOLUTION 60 MM X 10 MM METAL BILIARY STENT. THE CUSTOMER FURTHER STATED THAT IT WAS SUSPECTED THAT THE SUBJECT DEVICE CAUSED OR CONTRIBUTED TO THE ADVERSE EVENTS DESCRIBED AS THE SYMPTOMS APPEARED SEVERAL HOURS AFTER THE EUS-GUIDED PANCREATIC TUMOR BIOPSY. DAMAGE TO THE ESOPHAGUS MAY HAVE OCCURRED DURING INSERTION OR REMOVAL OF THE ENDOSCOPE FROM THE GASTROINTESTINAL TRACT. IT IS ALSO POSSIBLE THAT SPONTANEOUS TEARING OF THE MUCOSA OCCURRED AS A RESULT OF NAUSEA, GAGGING, AND VOMITING. THE CUSTOMER NOTED THAT A PANCREATIC TUMOR BIOPSY WAS PERFORMED VIA THE DUODENUM; THE AUTOPSY REPORT DID NOT DESCRIBE ANY PANCREATIC FLUID LEAKAGE OR BLEEDING IN THE AREA OF THE BIOPSY SITE. BLOOD AND BLOOD CLOTS WERE FOUND IN THE STOMACH, BLOOD CLOTS IN THE DUODENUM, AND BLOOD AND CLOTS IN THE INITIAL SECTION OF THE SMALL INTESTINE. A NON-OLYMPUS NEEDLE WAS USED FOR THE BIOPSY. THREE BIOPSIES WERE TAKEN THROUGH THE DUODENAL WALL USING THE FAN METHOD AND NO DAMAGE TO THE NEEDLE WAS FOUND DURING OR AFTER THE PROCEDURE. Manufacturer Narrative: THE INVESTIGATION IS ONGOING. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETED OR IF ADDITIONAL INFORMATION BECOMES AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22728913&pc=NPT | 2015691-2025-06379 | 2010-03-01 05:00 | Death | EDWARDS LIFESCIENCES | 2025-08-06 04:00 | NPT | SAPIEN XT TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Cardiac Perforation | P130009 |  | 1 | Event Description: AS REPORTED IN THE JOURNAL ARTICLE "CARDIAC TAMPONADE COMPLICATING TRANSCATHETER AORTIC VALVE REPLACEMENT: INSIGHTS FROM A SINGLE CENTER REGISTRY" FROM ISRAEL, A RETROSPECTIVE ANALYSIS WAS CONDUCTED ON PATIENTS WHO UNDERWENT THE TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) PROCEDURE FOR SEVERE SYMPTOMATIC AORTIC STENOSIS BETWEEN MARCH 2010 AND JUNE 2024. THIS COMPLAINT INVOLVES A PATIENT WHO UNDERWENT A TAVR PROCEDURE WITH A 26 MM SAPIEN XT VALVE. APPROXIMATELY ONE (1) HOUR POST VALVE IMPLANTATION, THE PATIENT PRESENTED WITH A LEFT VENTRICULAR (LV) PERFORATION ASSOCIATED WITH A STIFF GUIDEWIRE. PERCUTANEOUS PERICARDIOCENTESIS AND CARDIAC SURGERY WERE PERFORMED. SUBSEQUENTLY, IT WAS NOTED THAT THE PATIENT PASSED AWAY DURING HOSPITALIZATION. THE DATE OF DEATH IS UNKNOWN. Manufacturer Narrative: THIS IS ONE OF THREE MANUFACTURER REPORTS BEING SUBMITTED FOR THIS CASE. PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. THERE ARE SEVERAL POTENTIAL ETIOLOGIES FOR VENTRICULAR PERFORATION DURING A TVR PROCEDURE, INCLUDING PERFORATION BY THE GUIDEWIRE, THE DELIVERY SYSTEM, OR THE TRANSVENOUS PACER (TVP) LEAD. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVE (THV). TRAINING INCLUDES PROPER GUIDEWIRE POSITIONING, FIXATION OF THE TVP TO PREVENT VENTRICLE PERFORATION, AND CAREFUL MANIPULATION OF DEVICES. PER THE PROCEDURE DIDACTIC, PATIENTS WITH SMALL VENTRICLES ARE AT PARTICULARLY HIGH RISK FOR VENTRICULAR PERFORATION. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST/INDICATE PROCEDURAL FACTORS (LV PERFORATION CAUSED BY THE STIFF GUIDEWIRE) MAY HAVE CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED TO INCLUDE ADDITIONAL INFORMATION, REFERENCE THE JOURNAL ARTICLE AND DOCUMENT THE RELATED MANUFACTURER REPORT NUMBERS IN SECTION H10 (RELATED REPORT NUMBERS). THIS IS ONE OF THREE MANUFACTURER REPORTS BEING SUBMITTED FOR THIS CASE. PLEASE SEE MANUFACTURER REPORT NUMBERS 2015691-2025-06380 AND 2015691-2025-06381 FOR DETAILS OF THE OTHER EVENTS. REFERENCE FOR JOURNAL ARTICLE: NAOUM I, EITAN A, SLIMAN H, SHIRAN A, ADAWI S, ASMER I, ZISSMAN K, JAFFE R. CARDIAC TAMPONADE COMPLICATING TRANSCATHETER AORTIC VALVE REPLACEMENT: INSIGHTS FROM A SINGLE-CENTER REGISTRY. CJC OPEN. 2024 NOV 14;7(2):153-160.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22729017&pc=NPT | 9612164-2025-03863 | 2025-08-01 04:00 | Death | MEDTRONIC IRELAND | 2025-08-06 04:00 | NPT | EVOLUT FX DCS | Failure to Advance; Insufficient Information; Patient Device Interaction Problem | Low Blood Pressure/ Hypotension; Rupture | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE PROCEDURE INVOLVING THE TRANSCATHETER AORTIC VALVE EVFXPLUS-29, THERE WAS NODE OVERLAP ACROSS 4 IN THE FLUOROSCOPIC CHECK. THE VALVE WAS NOT USED AND A NEW DEVICES WERE USED TO LOAD A NEW VALVE. AFTER POSITIONING A NON-MEDTRONIC GUIDEWIRE (LUNDERQUIST DC) IN THE LEFT VENTRICLE, THE IMPLANTER ATTEMPTED TO CROSS THE AORTIC ARCH BUT WAS UNABLE TO. A LONG 64 CENTIMETER SHEATH WAS PREPARED, AND UPON THE DELIVERY CATHETER SYSTEM (DCS) RETRIEVAL, THE PATIENT'S BLOOD PRESSURE DROPPED IMMEDIATELY. AORTIC ANGIOGRAPHY IDENTIFIED AN AORTIC RUPTURE AND A TYPE B AORTIC DISSECTION DIRECTLY DISTAL TO THE OSTIUM OF THE LEFT SUBCLAVIAN ARTERY. TEMPORARY BUT INSUFFICIENT STABILIZATION WAS ATTEMPTED WITH A 24 MILLIMETER (MM) NON-MEDTRONIC (VACS2) BALLOON. CARDIOPULMONARY RESUSCITATION AND BLOOD TRANSFUSION WERE REQUIRED. FOLLOWING A PHYSICIAN TEAM DISCUSSION, AN AORTIC STENT AND A NON-MEDTRONIC AORTIC GRAFT STENT (GORE TAG 31MM/15CM) WERE IMPLANTED IN THE DISTAL AORTIC ARCH. THE HEART-LUNG MACHINE WAS ESTABLISHED, A CHANGE TO A CARDIOHELP SYSTEM WAS MADE, AND THE PATIENT WAS TRANSFERRED TO THE INTENSIVE CARE UNIT. SUBSEQUENTLY, THE PATIENT DIED AND THE CAUSE OF DEATH WAS NOT REPORTED.Event Description: PER THE PHYSICIAN, THE CAUSE OF THE INJURY WAS THE PATIENT'S AORTIC DISEASE, TORTUOSITY, CALCIFICATION OF THE AORTA, PREVIOUSLY IMPLANTED VALVES IN THE AORTIC (PERIMOUNT 27 MM) AND MITRAL (EDWARDS) POSITIONS AND THE RIGID DELIVERY CATHETER SYSTEM (DCS). PER THE PHYSICIAN, THE CAUSE OF DEATH WAS AORTIC RUPTURE. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID EVFXPLUS-29; PRODUCT LOT/SERIAL NUMBER (B)(6); PRODUCT TYPE: HEART VALVES; IMPLANT DATE; EXPLANT DATE; PRODUCT ID EVFXPLUS-29; PRODUCT LOT/SERIAL NUMBER K122311; PRODUCT TYPE: HEART VALVES; IMPLANT DATE; EXPLANT DATE PRODUCT ID D-EVOLUTFX-2329; PRODUCT LOT/SERIAL NUMBER (B)(6); PRODUCT TYPE: HEART VALVES; IMPLANT DATE; EXPLANT DATE PRODUCT ID L-EVOLUTFX-2329; PRODUCT LOT/SERIAL NUMBER (B)(6); PRODUCT TYPE: HEART VALVES; IMPLANT DATE; EXPLANT DATE PRODUCT ID LUNDERQUIST DC WIRE; PRODUCT LOT/SERIAL NUMBER N/A; PRODUCT TYPE: GUIDEWIRE; IMPLANT DATE; EXPLANT DATE PRODUCT ID 64 CM SHEATH; PRODUCT LOT/SERIAL NUMBER N/A; PRODUCT TYPE: SHEATH; IMPLANT DATE; EXPLANT DATE PRODUCT ID 24 MM VACS2 BALLOON; PRODUCT LOT/SERIAL NUMBER N/A; PRODUCT TYPE: DILATION BALLOON; IMPLANT DATE; EXPLANT DATE PRODUCT ID AORTIC GRAFT STENT GORE TAG 31MM/15CM; PRODUCT LOT/SERIAL NUMBER N/A; PRODUCT TYPE: STENT GRAFT; IMPLANT DATE (B)(6) 2025; EXPLANT DATE, SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5 CONTINUATION OF D10: PRODUCT ID 27 MM PERIMOUNT; PRODUCT LOT/SERIAL NUMBER N/A; PRODUCT TYPE: HEART VALVE; IMPLANT DATE (B)(6) 2010; PRODUCT ID EDWARDS MITRAL VALVE; PRODUCT LOT/SERIAL NUMBER N/A; PRODUCT TYPE: HEART VALVE; IMPLANT DATE (B)(6) 2010; H6 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22729024&pc=PNJ | 2017865-2025-96956 | 2025-07-21 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-08-06 04:00 | PNJ | AVEIR¿ LEADLESS PACEMAKER, RIGHT VENTRICLE | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P150035 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT DECEASED DUE TO AN UNKNOWN CAUSE THE DAY AFTER VENTRICULAR LEADLESS PACEMAKER (LP) IMPLANT. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22729201&pc=NPT | 9617601-2025-00837 | 2025-07-15 04:00 | Death | MEDTRONIC MEXICO S. DE R.L. DE CV | 2025-08-06 04:00 | NPT | EVOLUT FX PLUS VALVE | Patient Device Interaction Problem | Rupture | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT PRIOR TO THE IMPLANT OF A TRANSCATHETER VALVE, A PRE-IMPLANT BALLOON AORTIC VALVULOPLASTY (BAV) WAS PERFORMED AND AN ANNULAR RUPTURE ALONG WITH AN ASCENDING AORTIC DISSECTION OCCURRED. THE VALVE WAS THEN NOTED TO BE IMPLANTED; HOWEVER, THE PROCEDURE WAS THEN CONVERTED TO OPEN-HEART SURGERY WITH THE INTENT TO REPLACE THE ASCENDING AORTA, BUT THE PATIENT DIED. PER THE PHYSICIAN, THE CAUSE OF DEATH WAS AN ANNULAR RUPTURE AND ASCENDING AORTIC DISSECTION CAUSED BY THE PRE-BAV. Manufacturer Narrative: CONTINUATION OF D10: SECTION D INFORMATION REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER RELEVANT DEVICE(S) ARE: PRODUCT ID: D -EVOLUTFX-2329, SERIAL/LOT #: UNKNOWN, H6. THE CODES PRESENT IN SECTION H6 CORRESPOND TO MULTIPLE COMPONENTS/PRODUCTS THAT COMPRISE THIS REPORTED EVENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22729504&pc=NPT | 2015691-2025-06381 | 2010-03-01 05:00 | Death | EDWARDS LIFESCIENCES | 2025-08-06 04:00 | NPT | SAPIEN 3 TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Rupture | P140031 |  | 1 | Event Description: AS REPORTED IN THE JOURNAL ARTICLE "CARDIAC TAMPONADE COMPLICATING TRANSCATHETER AORTIC VALVE REPLACEMENT: INSIGHTS FROM A SINGLE CENTER REGISTRY" FROM ISRAEL, A RETROSPECTIVE ANALYSIS WAS CONDUCTED ON PATIENTS WHO UNDERWENT THE TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) PROCEDURE FOR SEVERE SYMPTOMATIC AORTIC STENOSIS BETWEEN MARCH 2010 AND JUNE 2024. THIS COMPLAINT INVOLVES A PATIENT WHO UNDERWENT A TAVR PROCEDURE WITH A 26 MM SAPIEN 3 VALVE. THE VALVE WAS IMPLANTED, AND A POST-DILATION WAS PERFORMED USING A 23 MM BALLOON. DURING THE PROCEDURE, THE PATIENT PRESENTED WITH AN ANNULAR RUPTURE. PERCUTANEOUS PERICARDIOCENTESIS WAS PERFORMED. SUBSEQUENTLY, IT WAS NOTED THAT THE PATIENT PASSED AWAY DURING HOSPITALIZATION. THE DATE OF DEATH IS UNKNOWN. Manufacturer Narrative: REFERENCE FOR JOURNAL ARTICLE: NAOUM I, EITAN A, SLIMAN H, SHIRAN A, ADAWI S, ASMER I, ZISSMAN K, JAFFE R. CARDIAC TAMPONADE COMPLICATING TRANSCATHETER AORTIC VALVE REPLACEMENT: INSIGHTS FROM A SINGLE-CENTER REGISTRY. CJC OPEN. 2024 NOV 14;7(2):153-160. THIS IS ONE OF THREE MANUFACTURER REPORTS BEING SUBMITTED FOR THIS CASE. PLEASE SEE MANUFACTURER REPORT NUMBERS 2015691-2025-06379 AND 2015691-2025-06380 FOR DETAILS OF THE OTHER EVENTS. PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. ACCORDING TO THE DEVICE TRAINING MANUALS, RISK FACTORS FOR AORTIC DISSECTION, CARDIAC/AORTIC HEMATOMA, OR ANNULAR RUPTURE DURING THE TVR PROCEDURE INCLUDE SIGNIFICANT VALVE OVER SIZING, SEVERELY OBLITERATED SINUSES OF VALSALVA, PORCELAIN AORTA AND/OR PRESENCE OF BULKY CALCIFICATION, AND NARROW CALCIFIED SINOTUBULAR JUNCTION. IN ADDITION, ADVANCED AGE, FEMALE GENDER, SMALL BODY WEIGHT, AND STEROID DEPENDENCY CAN ALSO BE CONTRIBUTING FACTORS. THE SAPIEN TRANSCATHETER HEART VALVE RELIES CALCIUM TO SECURELY ANCHOR IN THE LANDING ZONE. DESPITE THIS BENEFICIAL ASPECT OF CALCIUM, BULKY CALCIUM CAN INCREASE THE RISK OF CALCIFIC NODULE DISPLACEMENT INTO THE VASCULATURE, WHICH CAN LEAD TO VASCULAR INJURY. AT TIMES, THE EXTENT AND DISTRIBUTION OF CALCIUM CAN IMPAIR EASE OF DELIVERY OF THE VALVE, CORRECT POSITIONING OF THE VALVE, DEPLOYMENT OF THE VALVE, AND PROCEDURAL SUCCESS. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. ALTHOUGH A DEFINITIVE ROOT CAUSE WAS UNABLE TO BE DETERMINED, THE EVENT COULD BE RELATED TO THE MECHANISMS DESCRIBED ABOVE. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22729655&pc=OZD | 1220648-2025-31069 | 2025-07-23 04:00 | Death | ABIOMED, INC. | 2025-08-06 04:00 | OZD | IMPELLA CP | Failure to Advance; Deformation Due to Compressive Stress | Hematoma; Ventricular Fibrillation; Asystole | P140003 |  | 1 | Event Description: THE COMPLAINANT HAD AN IMPELLA CP BEING PLACED TO SUPPORT A NON-ST-SEGMENT ELEVATION MYOCARDIAL INFARCTION PATIENT WITH MULTIVESSEL CORONARY ARTERY DISEASE AND ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK. THE PRIMARY REGISTERED NURSE REPORTED A LARGE HEMATOMA. MANUAL PRESSURE WAS HELD BUT THE HEMATOMA RETURNED. A LARGE UNRESOLVABLE HEMATOMA IN THE RIGHT GROIN WAS REPORTED. ADDITIONALLY, THE IMPELLA POSITION WAS UNABLE TO BE OPTIMIZED. THE PATIENT WAS INTUBATED FOR AIRWAY PROTECTION AND WAS PLACED BACK ON THE CATHETERIZATION LAB TABLE. A REPO UNIT WAS SEEN TO HAVE NOT BE ABLE TO REACH THE VESSEL. THE DOCTOR ATTEMPTED TO ADVANCE BUT RESISTANCE WAS BEING MET ON ADVANCEMENT OF THE IMPELLA. INLET AT THE LEVEL OF AORTA VALVE IN THE VENTRICLE. VISIBLE KINK FROM REPO AT THE LEVEL OF THE GROIN. THE MEDICAL DOCTOR ATTEMPTED TO ADVANCE AGAIN BUT WAS UNSUCCESSFUL. THE DOCTOR REWIRED THE REACCESS SHEATH TO ATTEMPT REINSERTION. HOWEVER, THE IMPELLA CP WAS REMOVED AND EXPLANTED DUE TO THE COMPLICATION. AN ADDITIONAL IMPELLA WAS USED. A NEW INTRODUCER SHEATH WAS PLACED. THE PATIENT CODED AND THE MEDICAL DOCTOR TRIED TO ACCESS THE LEFT GROIN. ADVANCED CARDIAC LIFE SUPPORT MEASURES WERE CONTINUED FOR SEVERAL MINUTES. RETURN OF SPONTANEOUS CIRCULATION (ROSC) ACHIEVED AFTER SEVERAL ROUNDS OF CARDIOPULMONARY RESUSCITATION. MASSIVE BLOOD TRANSFUSION WAS CALLED. THE PATIENT CONTINUED TO GO INTO VENTRICULAR FIBRILLATION. DRIPS WERE STARTED AND MAXED OUT. THE PATIENT CONTINUED TO BE IN ASYSTOLE AND IMPELLA CP FLOWS WERE UNABLE TO REACH GREATER THAN ONE LITER PER MINUTE. THE PATIENT EXPIRED. ADDITIONAL INFORMATION FROM THE REPRESENTATIVE WAS PROVIDED. "THEY MAINLY ATTRIBUTE THE DEATH TO THE INTERNAL BLEED FROM THE REPO SHEATH NOT REACHING THE VESSEL. THE MD ATTEMPTED TO ADVANCE THE SHEATH BACK BUT WAS UNSUCCESSFUL DUE TO STENOSIS, SO THE ONLY OTHER OPTION WAS TO REINSERT IT AGAIN. HE TRIED GETTING CONTRALATERAL ACCESS BUT UNFORTUNATELY UNSUCCESSFUL DUE TO STENOSIS ALSO IN THE LEFT CFA. THE PATIENT DIDN'T TOLERATE THE REMOVE AND REINSERTION OF THE IMPELLA. THE BLOOD LOSS, SEVERE DISEASE, AND PAUSE OF SUPPORT ALL CONTRIBUTED TO THE DEATH OF THE PATIENT." Manufacturer Narrative: THE IMPELLA PUMP AND PURGE CASSETTE WERE RETURNED, AND AN EVALUATION IS UNDERWAY. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM. SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT. USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER - ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH RISK PERCUTANEOUS CORONARY INTERVENTION. SECTION: WARNINGS & CAUTIONS: CAUTIONS - ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS - ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22729861&pc=NGV | 2124215-2025-54563 | 2025-04-10 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-08-06 04:00 | NGV | SYSTEM, APPENDAGE CLOSURE, LEFT ATRIAL | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: IT WAS REPORTED VIA JOURNAL ARTICLE THAT DEATHS OCCURRED. THE FOLLOWING EVENT WAS OBTAINED VIA A RETROSPECTIVE ANALYSIS FOLLOWING 103 PATIENTS WHO UNDERWENT A A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WHERE WATCHMAN CLOSURE DEVICES WERE SUCCESSFULLY IMPLANTED AND THE PATIENTS WERE FOLLOWED. IT WAS REPORTED THAT OUT OF 103 PATIENTS, THE FOLLOWING DEATHS OCCURRED: 29 PATIENTS DIED DURING FOLLOW UP FROM VARYING CAUSES. Manufacturer Narrative: B3: USED ARTICLE PUBLISH DATE AS EVENT DATE, AS IT IS UNKNOWN. D1: BRAND NAME IS BLANK BECAUSE THE DEVICE IS KNOWN TO BE A WATCHMAN LAA CLOSURE DEVICE, BUT THE EXACT BRAND NAME IS UNKNOWN. D4: DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BASED ON THE NATURE OF THE INFORMATION PROVIDED TO BSC, IT IS NOT POSSIBLE TO PERFORM A GOOD FAITH EFFORT TO OBTAIN ADDITIONAL INFORMATION. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER PRODUCT SPECIFIC INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. LITERATURE CITATION: LIU, M. ET AL (APRIL 10, 2025). THE SAFETY AND EFFICACY OF LEFT ATRIAL APPENDAGE CLOSURE DEVICES IN PATIENTS WITH NON-TRAUMATIC INTRACRANIAL HEMORRHAGE. 473. JOURNAL OF THE NEUROLOGICAL SCIENCES. HTTPS://DOI.ORG/10.1016/J.JNS.2025.123490.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22730029&pc=DSQ | 2916596-2025-05172 | 2025-07-25 04:00 | Death | THORATEC CORPORATION | 2025-08-06 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTI SYSTEM ORGAN FAILURE. THE DEATH WAS NOT CONSIDERED DEVICE OR THERAPY RELATED WAS DUE TO THE PATIENT BEING TOO SICK WITH MULTI SYSTEM ORGAN FAILURE. THE DEVICE WAS NOT EXPLANTED AND WOULD NOT RETURN FOR EVALUATION.Event Description: THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED EVENT AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION AND DEATH AS POTENTIAL ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROIVDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22730217&pc=QFG | 3013756811-2025-186489 | 2025-07-06 04:00 | Death | TANDEM DIABETES CARE | 2025-08-06 04:00 | QFG | T:SLIM X2 INSULIN PUMP WITH CONTROL-IQ TECHNOLOGY | Adverse Event Without Identified Device or Use Problem | Diabetic Ketoacidosis | K203234 |  | 1 | Event Description: IT WAS REPORTED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2025. THE CONTACT MENTIONED THERE WAS SUSPICION THAT ALCOHOL AND DIABETIC KETOACIDOSIS (DKA) MIGHT HAVE BEEN INVOLVED. NO ADDITIONAL INFORMATION WAS PROVIDED. A REVIEW OF PUMP DATA WILL BE PERFORMED, AND THE RESULTS WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22730225&pc=DSQ | 2916596-2025-05098 | 2025-07-21 04:00 | Death | THORATEC CORPORATION | 2025-08-06 04:00 | DSQ | THORATEC® HEARTMATE®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO HEART FAILURE. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM IFU, IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22730294&pc=OZD | 1220648-2025-45722 | 2025-07-26 04:00 | Death | ABIOMED, INC. | 2025-08-06 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA 5.5 DEVICE AS ELECTIVE PLACEMENT FOR MECHANICAL CIRCULATORY SUPPORT AND THEN SUBSEQUENTLY EXPERIENCED BLEEDING AND EXPIRED. THE PATIENT WAS A TRANSFER IN FOR ADVANCE CARE AND FOUND TO BE IN CARDIOGENIC SHOCK WITH SEVERE AORTIC INSUFFICIENCY, MITRAL REGURGITATION, TRICUSPID REGURGITATION, AND PULMONARY STENOSIS. THE PATIENT UNDERWENT AORTIC VALVE REPLACEMENT (AVR) AND PULMONARY VALVE REPLACEMENT (PVR). DURING THIS SURGERY, SURGEONS ALSO FOUND A LEFT MAIN STENT WAS IN THE AORTA AND WAS SUBSEQUENTLY REMOVED. AFTER THIS STENT REMOVAL, THE LEFT MAIN WAS FOUND TO BE PATENT, AND CORONARY ARTERY BYPASS GRAFT SURGERY WAS NOT REQUIRED. POST AVR AND PVR, THE IMPELLA 5.5 DEVICE WAS SUCCESSFULLY IMPLANTED VIA THE RIGHT AXILLARY ARTERY. THE IMPELLA SUPPORT WAS INITIATED ON PERFORMANCE LEVEL P-2, THE PATIENT SLOWLY WEANED FROM CARDIOPULMONARY BYPASS (CPB), AND THE IMPELLA WAS THEN DIALED UP TO P-8. FLOWS WERE APPROXIMATELY 4.7LPM WITH INLET PLACED APPROXIMATELY 5.2 CM FROM THE AORTIC VALVE ANNULUS AND MIDVENTRICULAR. SHORTLY AFTER COMING OFF CPB, IMPELLA SUCTIONING OCCURRED. RIGHT VENTRICULAR FAILURE WAS NOTED ON ECHOCARDIOGRAM, AND THERE WAS A CONCERN FOR FAILED PULMONARY VALVE REPLACEMENT. A DECISION MADE TO RECONFIGURE TO VENOUS-ARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION (VA ECMO). THE PATIENT CONTINUED TO HAVE EXCESSIVE BLEEDING AND REQUIRED MULTIPLE BLOOD PRODUCTS. THE PATIENT WAS EVENTUALLY TRANSFERRED TO THE UNIT ON IMPELLA AND ECMO SUPPORT. HOWEVER, APPROXIMATELY ONE DAY LATER, THE PATIENT WAS NOTED TO HAVE EXPIRED. CARE WAS WITHDRAWN. Manufacturer Narrative: ADDITIONAL COMMENTS BY THE PHYSICIAN STATED THE BLEEDING IS RELATED TO EXTRACORPOREAL MEMBRANE OXYGENATION AND THE PATIENT¿S PREDISPOSED COAGULOPATHY ISSUES WITH UNDERLYING CONGENITAL CONDITION. HOWEVER, THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA DEVICE AS A CONTRIBUTING FACTOR IN THE PATIENT¿S OUTCOME. PRODUCT STATUS: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THERE IS UNCERTAINTY WHAT HAPPENED WITH THE DEVICE AFTER THE PATIENT'S DEMISE. THEREFORE, AN EVALUATION/ANALYSIS OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22733496&pc=QFG | 3019004087-2025-00770 | 2025-07-05 04:00 | Death | BETA BIONICS, INC | 2025-08-06 04:00 | QFG | ILET BIONIC PANCREAS | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K231485 |  | 1 | Event Description: ON (B)(6) 2025, A FRIEND OF THE USER REPORTED THAT THE USER PASSED AWAY AT HOME ON (B)(6) 2025. THE FRIEND STATED THAT THE CAUSE OF DEATH WAS UNKNOWN AND THAT THE USER HAD REPORTEDLY NOT BEEN FEELING WELL PRIOR TO PASSING. THE USER WAS REPORTEDLY WEARING THE ILET AT THE TIME. NO ADDITIONAL DETAILS WERE PROVIDED, AND NO CONCERNS SPECIFIC TO THE DEVICE WERE RAISED. Manufacturer Narrative: THE DEVICE HISTORY RECORD (DHR) REVIEW CONFIRMED THAT THE ILET PASSED ALL MANUFACTURING SPECIFICATIONS PRIOR TO RELEASE. ENGINEERING ANALYSIS OF THE DEVICE LOG FILE COULD NOT BE COMPLETED AS THE DEVICE DATA LOGS HAVE NOT BEEN SYNCED AND HAS NOT YET BEEN RETURNED. A REQUEST FOR DEVICE RETURN HAS BEEN MADE. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22733637&pc=OZD | 1220648-2025-45715 | 2025-07-22 04:00 | Death | ABIOMED, INC. | 2025-08-06 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT THE PATIENT HAD A GASTROINTESTINAL BLEED (GI BLEED) DURING IMPELLA 5.5 SUPPORT. THE PATIENT HAD RECURRENT GI BLEEDING REQUIRING BLOOD TRANSFUSION. IN THE SETTING OF MIXED SHOCK PICTURE AND SIGNIFICANT GI BLEEDING, THE PATIENT WAS SEDATED, EXTUBATED, VASOACTIVE DRIPS WERE STOPPED AND THE IMPELLA WAS TAPERED OFF AND STOPPED. AS SUCH, THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22734121&pc=NPT | 2015691-2025-06399 | 2025-07-15 04:00 | Death | EDWARDS LIFESCIENCES | 2025-08-06 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Migration or Expulsion of Device | Foreign Body In Patient; Foreign Body Embolism | P140031 |  | 1 | Event Description: EDWARDS RECEIVED NOTIFICATION FROM OUR AFFILIATE IN SWITZERLAND. AS REPORTED, THIS WAS A CASE OF AN IMPLANT OF A 26MM SAPIEN 3 ULTRA RESILIA VALVE, IN THE AORTIC POSITION BY TRANSFEMORAL APPROACH. THE NATIVE AORTIC VALVE WAS NOT CALCIFIED (PURE AORTIC INSUFFICIENCY) AND AFTER VALVE DEPLOYMENT, IT EMBOLIZED. MANY ATTEMPTS WERE PERFORMED TO REPOSITION THE VALVE FROM THE ASCENDING AORTA TOWARD THE AORTIC ARCH. FINALLY, THE VALVE WAS DEPLOYED IN THE AORTIC ARCH BUT THE MANEUVERS RESULTED IN AN AORTIC DISSECTION. THE PATIENT WAS INITIALLY NOTED TO BE STABLE AFTER THE PROCEDURE. HOWEVER, THE PATIENT EXPIRED SEVERAL HOURS LATER. Manufacturer Narrative: THE DEVICE WAS USED IN OFF-LABEL IMPLANTATION IN THE AORTIC POSITION FOR PURE AORTIC INSUFFICIENCY. AS THERE ARE NO SPECIFIC IFU OR TRAINING MATERIALS RELATED TO THESE PROCEDURES, THE AVAILABLE TRAINING MATERIALS WERE REVIEWED ONLY FOR INFORMATION POTENTIALLY RELEVANT TO THE DEVICE USE. PER THE INSTRUCTIONS FOR USE (IFU), VALVE EMBOLIZATION IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. THERE ARE MULTIPLE PATIENT AND PROCEDURAL FACTORS THAT ALONE OR IN COMBINATION CAN CAUSE OR CONTRIBUTE TO EMBOLIZATION OF THE DEVICE, INCLUDING IMPROPER POSITIONING BEFORE DEPLOYMENT , POOR IMAGE INTENSIFIER ANGLE, POOR COAXIAL ALIGNMENT OF THE VALVE AND DELIVERY SYSTEM, SEVERE SEPTAL HYPERTROPHY, MINIMALLY OR BULKY/SEVERELY CALCIFIED LEAFLETS, LOSS OF PACING CAPTURE, A NARROW SINOTUBULAR JUNCTION (IN AORTIC POSITION PROCEDURES), RAPID DEPLOYMENT, THE RELEASE OF STORED TENSION DURING DEPLOYMENT, INACCURATE MEASUREMENT OF THE LANDING ZONE, AND MOVEMENT OF THE DELIVERY SYSTEM BY THE OPERATOR. THE IFU CAUTIONS THAT INCORRECT SIZING OF THE LANDING ZONE MAY LEAD TO PARAVALVULAR LEAK, MIGRATION, OR EMBOLIZATION. THE DEVICE TRAINING MANUALS INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVES (ALL MODELS). TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES. THE CORRECT ALIGNMENT AND POSITIONING OF THE DEVICE AT THE POINT OF DEPLOYMENT IS EMPHASIZED AS A KEY FACTOR IN THE PLACEMENT AND FIXATION OF THE DEVICE. OPERATORS ARE ALSO INSTRUCTED TO USE FLUOROSCOPY AS THE PRIMARY METHOD OF VISUALIZATION FOR POSITIONING AND DEPLOYMENT. IN PATIENTS WITH HIGH-RISK ANATOMICAL FEATURES FOR EMBOLIZATION (I.E., MINIMAL LEAFLET CALCIFICATION, SEVERE SEPTAL HYPERTROPHY), BALLOON VALVULOPLASTY MAY INDICATE POTENTIAL BALLOON MOVEMENT DURING VALVE DEPLOYMENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST THAT PATIENT RELATED FACTORS (LACK OF CALCIUM) AND PROCEDURAL RELATED FACTORS (OFF LABEL USE AND REPOSITIONING THE VALVE FROM THE ASCENDING AORTA TOWARD THE AORTIC ARCH) CAUSED OR CONTRIBUTED TO THE EVENT OF EMBOLIZATION AND SUBSEQUENT AORTIC DISSECTION. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22710833&pc=GAW | 2210968-2025-09025 | 2025-07-07 04:00 | Death | ETHICON INC. | 2025-08-05 04:00 | GAW | PRLNE HS BLU 30IN 4-0 D/A C-1 | Break; Adverse Event Without Identified Device or Use Problem | Failure to Anastomose; Hemorrhage/Bleeding; Unspecified Tissue Injury | K133356 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A CABG W/ PFO CLOSURE VIA RIGHT ATRIOTOMY ON (B)(6) 2025 DUE TO CORONARY ARTERY DISEASE W/ CRYPTOGENIC STROKE IN (B)(6). AND SUTURE WAS USED. THE SURGEON WAS SUPPOSED TO CLOSE THE RIGHT ATRIUM AND THE PFO. HOWEVER, THE RIGHT ATRIUM REPORTEDLY BLEW, POTENTIALLY DUE TO A SUTURE FAILURE, CAUSING THE PATIENT TO PASS AWAY. THE CAUSE OF DEATH WAS THAT THE PATIENT ARRESTED IN THE ICU DUE TO EXSANGUINATING HEMORRHAGE THAT HAPPENED AND THE PATIENT HAD SUDDEN BLEEDING OUT OF CHEST TUBES. THE SURGEON OPENED THE CHEST IN THE ICU, REPAIRED THE RIGHT ATRIUM. THE PATIENT WAS NOT ABLE TO BE REVIVED. SURGEON EXPLORED THE SITE OF INJURY/REPAIR AND WAS NOT ABLE TO FIND THE SUTURE. NOT TISSUE WEAKNESS OBSERVED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON INC, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON INC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. H6 COMPONENT CODE: G07002 DEVICE NOT RETURNED. ADDITIONAL INFORMATION: D4, H4 - THE ACTUAL DEVICE BATCH NUMBER ASSOCIATED WITH THIS EVENT IS NOT KNOWN. THREE POSSIBLE BATCH NUMBERS ARE REPORTED AS FOLLOWS: UBMSBZ - EXPIRATION DATE: JAN/31/2029 DATE OF MFG.: FEB/29/2024 TLMQKM, UCMCKL. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE UBMSBZ BATCH NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. D4: FULL UDI CURRENTLY UNAVAILABLE SINCE THE EXACT LOT OF THE SUSPECTED DEVICE CANNOT BE DETERMINED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE PRODUCT UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE PRODUCT EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. ADDITIONAL INFORMATION WAS REQUESTED, AND THE FOLLOWING WAS OBTAINED: 1. CAN YOU PLEASE PROVIDE THE PATIENT'S DEMOGRAPHIC INFORMATION, INCLUDING AGE, GENDER, WEIGHT, AND BMI AT THE TIME OF THE INDEX PROCEDURE? 1. 50 YO FEMALE, 37 BMI, 90 KG. 2. WHAT WAS THE DIAGNOSIS AND INDICATION FOR THE INDEX SURGICAL PROCEDURE? 1. CORONARY ARTERY DISEASE W/ CRYPTOGENIC STROKE IN (B)(6). 3. WERE THERE ANY OTHER RELEVANT PATIENT HISTORY OR CONCOMITANT MEDICATIONS? 1. OBESITY, DIABETIC ON INSULIN PUMP. SURGICAL PROCEDURE DETAILS. 6. WHAT WAS THE NAME OF THE INDEX SURGICAL PROCEDURE? 1. CABG W/ PFO CLOSURE VIA RIGHT ATRIOTOMY. 7. WHAT WAS THE TISSUE CONDITION AT THE TIME OF THE PROCEDURE (NORMAL, THIN, CALCIFIED, FRAGILE, DISEASED)? 1. NORMAL. 8. HOW WAS THE SUTURE PLACED (INTERRUPTED OR CONTINUOUS)? 1. SINGLE CONTINUOUS STITCH. 9. HOW WAS THE SUTURE ORIGINALLY TIED (MULTIPLE KNOTS, SQUARE KNOT, ETC.)? 1. MULTIPLE SQUARE KNOTS. 10. DID YOU OBSERVE ANY SUTURE DEFICIENCIES OR ANOMALIES BEFORE, DURING, OR AFTER THE SUTURE PLACEMENT OR DURING ANY RE-OPERATION? 1. NONE. EVENT DETAILS: 10. WHAT WAS THE CAUSE OF DEATH? 1. ARRESTED IN THE ICU DUE TO EXSANGUINATING HEMORRHAGE. 11. WAS A CAUSE OF DEATH REPORTED WITHIN THE MEDICAL RECORD OR DESCRIBED IN AN AUTOPSY. REPORT? IF SO, COULD YOU SPECIFY? 1. ARRESTED IN THE ICU DUE TO EXSANGUINATING HEMORRHAGE. 12. WHAT SYMPTOMS, IF ANY, DID THE PATIENT EXPERIENCE FOLLOWING THE INDEX SURGICAL. PROCEDURE? WHAT WAS THE ONSET DATE OF THOSE SYMPTOMS? 1. EXTUBATED AND NORMAL POST-OP. SUDDEN BLEEDING OUT OF CHEST TUBES. 13. WHAT WAS THE BLEEDING SOURCE AND TRIGGERING EVENT? WHEN DID THE BLEEDING OCCUR? 1. RIGHT ATRIAL BLEEDING. 14. PLEASE DESCRIBE ANY MEDICAL OR SURGICAL INTERVENTIONS REQUIRED FOR THIS SUTURE EVENT, INCLUDING DATES AND RESULTS. 1. SUTURE OPENED THE CHEST IN THE ICU, REPAIRED THE RIGHT ATRIUM. THE PATIENT WAS NOT ABLE TO BE REVIVED. 15. DO YOU BELIEVE THAT THE PROLENE SUTURE CAUSED OR CONTRIBUTED TO THE PATIENT¿S DEATH? 1. SURGEON EXPLORED THE SITE OF INJURY/REPAIR AND WAS NOT ABLE TO FIND THE SUTURE. NOT TISSUE WEAKNESS OBSERVED. 16. PLEASE DESCRIBE THE SUTURE FAILURE. ADDITIONAL INFORMATION: 16. CAN YOU CONFIRM THAT ONE SUTURE WAS RELATED TO THIS EVENT, INCLUDING TWO POSSIBLE LOT NUMBERS (UBMSBZ, TLMQKM)? 1. YES, LOT NUMBERS CAN¿T BE CONFIRMED. 17. WHAT IS YOUR OPINION REGARDING THE ETIOLOGY OF OR CONTRIBUTING FACTORS TO THIS EVENT? 1. N/A. 18. WOULD YOU BE WILLING TO SPEAK WITH ETHICON MEDICAL SAFETY AND ENGINEERING VIA A SCHEDULED CONFERENCE CALL REGARDING THE PRODUCT INVOLVED IN THIS EVENT? 1. SURGEON DOES NOT FIND THIS NECESSARY AT THIS TIME. Manufacturer Narrative: PRODUCT COMPLAINT#: (B)(4). THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON INC, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON INC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. ADDITIONAL INFORMATION: D4, H4: THE ACTUAL DEVICE BATCH NUMBER ASSOCIATED WITH THIS EVENT IS NOT KNOWN. THREE POSSIBLE BATCH NUMBERS ARE REPORTED AS FOLLOWS: UBMSBZ - EXPIRATION DATE: JAN/31/2029 DATE OF MFG.: FEB/29/2024. TLMQKM - EXPIRATION DATE: SEP/30/2028 DATE OF MFG.: OCT/20/2023. UCMCKL - EXPIRATION DATE: FEB/28/2029 DATE OF MFG.: MAR/07/2024. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE POSSIBLE FINISHED DEVICE LOTS, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. 1. INVESTIGATION SUMMARY: THE PRODUCT WAS RETURNED TO ETHICON FOR EVALUATION. THREE REPRESENTATIVE UNOPENED SAMPLES WERE RECEIVED FOR ANALYSIS. PRODUCT CODE: 9706H, LOT: UCMCKL. THE COMPLAINT SAMPLE WAS NOT RECEIVED FOR EVALUATION. IN ORDER TO EVALUATE THE CONDITION OF THE RETURNED SAMPLES, THE PACKETS WERE OPENED, AND THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG THE STRAND TO DETECT ANY ISSUE RELATED TO THE CUSTOMER COMPLAINT AND NO DEFECTS WERE OBSERVED DURING EVALUATION. THE FUNCTIONAL TEST WAS PERFORMED USING INSTRON EQUIPMENT AND THE TENSILE FORCE WAS ABOVE THE MINIMUM REQUIREMENTS. THE EVENT DESCRIBED COULD NOT BE CONFIRMED AS THE DEVICE PERFORMED WITHOUT ANY DEFECT NOTED. AS PART OF THE ETHICON QUALITY PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND RELEASED TO APPROVED SPECIFICATIONS. THE EVENT DESCRIBED COULD NOT BE CONFIRMED AS NO PRODUCT DEFECT WAS IDENTIFIED, THERE MAY HAVE BEEN OTHER CIRCUMSTANCES OR ISSUES THAT OCCURRED DURING THE USE OF THE DEVICE THAT COULD NOT BE REPLICATED DURING THE LABORATORY ANALYSIS. ADDITIONAL COMPLAINT INFORMATION MONITORING FOR POTENTIAL SAFETY SIGNALS IS CONDUCTED THROUGH COMPLAINT TRENDING AS PART OF POST-MARKET SURVEILLANCE., 2. INVESTIGATION SUMMARY: THE PRODUCT WAS RETURNED TO ETHICON FOR EVALUATION. THIRTEEN REPRESENTATIVE UNOPENED SAMPLES WERE RECEIVED FOR ANALYSIS. PRODUCT CODE: 9706H, LOT: TLMQKM. THE COMPLAINT SAMPLE WAS NOT RECEIVED FOR EVALUATION. IN ORDER TO EVALUATE THE CONDITION OF THE RETURNED SAMPLES, THE PACKETS WERE OPENED, AND THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG THE STRAND TO DETECT ANY ISSUE RELATED TO THE CUSTOMER COMPLAINT AND NO DEFECTS WERE OBSERVED DURING EVALUATION. THE FUNCTIONAL TEST WAS PERFORMED USING INSTRON EQUIPMENT AND THE TENSILE FORCE WAS ABOVE THE MINIMUM REQUIREMENTS. THE EVENT DESCRIBED COULD NOT BE CONFIRMED AS THE DEVICE PERFORMED WITHOUT ANY DEFECT NOTED. AS PART OF THE ETHICON QUALITY PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND RELEASED TO APPROVED SPECIFICATIONS. THE EVENT DESCRIBED COULD NOT BE CONFIRMED AS NO PRODUCT DEFECT WAS IDENTIFIED, THERE MAY HAVE BEEN OTHER CIRCUMSTANCES OR ISSUES THAT OCCURRED DURING THE USE OF THE DEVICE THAT COULD NOT BE REPLICATED DURING THE LABORATORY ANALYSIS. ADDITIONAL COMPLAINT INFORMATION MONITORING FOR POTENTIAL SAFETY SIGNALS IS CONDUCTED THROUGH COMPLAINT TRENDING AS PART OF POST-MARKET SURVEILLANCE., 3. INVESTIGATION SUMMARY: THE PRODUCT WAS RETURNED TO ETHICON FOR EVALUATION. ONE SEALED BOX WITH THIRTY-SIX REPRESENTATIVE UNOPENED SAMPLES AND FOUR OPEN BOXES WITH SIXTEEN, TWENTY-EIGHT, THIRTEEN AND EIGHTEEN REPRESENTATIVE UNOPENED SAMPLES EACH WERE RECEIVED FOR ANALYSIS. PRODUCT CODE: 9706H, LOT: UBMSBZ. THE COMPLAINT SAMPLE WAS NOT RECEIVED FOR EVALUATION. UPON INITIAL INSPECTION OF THE SAMPLES, NO EXTERNAL DAMAGES WERE OBSERVED ON THE EXTERNAL PACKETS. FOLLOWING THE SAMPLING PLAN, A VISUAL INSPECTION WAS CONDUCTED ON THIRTY-TWO SAMPLES. THE PACKETS WERE OPENED, AND THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG THE STRAND TO DETECT ANY ISSUE RELATED TO THE CUSTOMER COMPLAINT AND NO DEFECTS WERE OBSERVED DURING EVALUATION. A FUNCTIONAL TEST WAS PERFORMED USING INSTRON EQUIPMENT AND TENSILE FORCE WAS ABOVE THE MINIMUM REQUIREMENTS. THE EVENT DESCRIBED COULD NOT BE CONFIRMED AS THE DEVICE PERFORMED WITHOUT ANY DEFECT NOTED. AS PART OF THE ETHICON QUALITY PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND RELEASED TO APPROVED SPECIFICATIONS. THE EVENT DESCRIBED COULD NOT BE CONFIRMED AS NO PRODUCT DEFECT WAS IDENTIFIED, THERE MAY HAVE BEEN OTHER CIRCUMSTANCES OR ISSUES THAT OCCURRED DURING THE USE OF THE DEVICE THAT COULD NOT BE REPLICATED DURING THE LABORATORY ANALYSIS. ADDITIONAL COMPLAINT INFORMATION MONITORING FOR POTENTIAL SAFETY SIGNALS IS CONDUCTED THROUGH COMPLAINT TRENDING AS PART OF POST-MARKET SURVEILLANCE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22711269&pc=OZP | 2032227-2025-227063 | 2024-12-27 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-08-05 04:00 | OZP | 670G INSULIN PUMP MMT-1780KL | Adverse Event Without Identified Device or Use Problem | Dehydration; Hyperglycemia; Renal Failure | P150001 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY ON THE (B)(6) 2024. THE CAUSE OF DEATH WAS REPORTED AS TYPE-1 DIABETES, HIGH BLOOD PRESSURE, LEUKEMIA, PARKINSONS, KIDNEY FAILURE & DEHYDRATION. THE EVENT INVOLVED PRODUCT(S) MMT-1780KL, MMT-332A & MMT-377A. TROUBLESHOOTING WAS PARTIALLY DONE AND FOUND THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP AT THE TIME OF DEATH AS THE BLOOD GLUCOSE VALUE WAS NOT CATCHING UP AND HAD TO BE TREATED WITH PEN INJECTIONS. IT WAS UNKNOWN IF THE AUTOMODE/SMART GUARD FEATURE WAS ACTIVE. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED. THE PRODUCT MMT-1780KL WILL BE RETURNED FOR ANALYSIS BUT MMT-332A & MMT-377A WILL NOT BE RETURNED. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. THE DEVICE WILL BE RETURNED FOR ANALYSIS AND FURTHER INFORMATION WILL FOLLOW ONCE THE ANALYSIS HAS BEEN COMPLETED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22713290&pc=DSQ | 2916596-2025-05048 | 2025-07-26 04:00 | Death | THORATEC CORPORATION | 2025-08-05 04:00 | DSQ | HEARTMATE II LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Decreased Pump Speed; Pumping Stopped; Infusion or Flow Problem | Respiratory Failure; No Clinical Signs, Symptoms or Conditions | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD NON-GROUND CABLE SINCE 2020. THE PATIENT CALLED THE HOSPITAL ON (B)(6) 2025 DUE TO ALARM. IT WAS NOTED THAT THE PATIENT WAS IN THE SHOWER AND ON BATTERIES. THE PATIENT CAME TO THE HOSPITAL AND LOG FILES WERE DOWNLOADED THAT SHOWED THE PUMP STOPPED AND SPEED WAS ON A NORMAL RANGE. OTHER ALARMS NOTED WERE LOW SPEED ADVISORY, LOW SPEED HAZARD, AND LOW FLOW ALARMS. THE TEAM DECIDED TO KEEP THE PATIENT IN THE HOSPITAL, AND THEY WERE TO DECIDE THE NEXT STEP. THE PATIENT WAS CLINICALLY DOING WELL. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22714087&pc=QBJ | MW5173989 |  | Death | ABBOTT DIABETES CARE INC. | 2025-08-05 04:00 | QBJ | FREESTYLE LIBRE 2 SENSOR | Insufficient Information | Cardiac Arrest; Respiratory Arrest |  |  | 1 | Event Description: PATIENT PASSED AWAY. REPORTED BY HCP (HEALTHCARE PROFESSIONAL) WHO DID NOT CONSENT TO FOLLOW UP. NO FURTHER INFORMATION PROVIDED OR AVAILABLE REGARDING THIS REPORT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22714114&pc=MGB | MW5173990 | 2025-05-23 04:00 | Death | ABBOTT VASCULAR | 2025-08-05 04:00 | MGB | ABBOTT PERCLOSE | Therapeutic or Diagnostic Output Failure | Ischemia; Obstruction/Occlusion |  |  | 1 | Event Description: PATIENT HAD TEVAR (THORACIC ENDOVASCULAR AORTIC REPAIR) PROCEDURE DONE FOR SACCULAR ANEURYSM OF AORTIC ARCH UNRUPTURED. AFTER PROCEDURE SURGEON UTILIZED ABBOT PER CLOSE DEVICE. ON (B)(6) 2025 CVA SHOWS LEFTY EXTERNAL ILIAC ARTERY OCCLUSION SECONDARY TO PERCLOSE. REQUIRED EMERGENT OPEN REPAIR FOR ACUTE LIMB ISCHEMIA DUE TO FAILURE OF PERCLOSE DEVICE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22714138&pc=IQK | MW5173991 |  | Death | OTTO BOCK HEALTHCARE LP | 2025-08-05 04:00 | IQK | OTTOBOCK 50C80=MINERVA ORTHOSIS | Device Dislodged or Dislocated | Aspiration/Inhalation; Pneumonia; Pressure Sores; Respiratory Failure; Unspecified Tissue Injury |  |  | 1 | Event Description: A 35-YEAR-OLD MALE WHO SUSTAINED AN UNSTABLE RIGHT C4 - C5 FACET DISLOCATION WITH SPONDYLOLISTHESIS AND CERVICAL SPINAL CORD INJURY AFTER A FALL DOWNSTAIRS ON (B)(6) 2025. HE WAS ADMITTED TO (B)(6) HOSPITAL AND UNDERWENT C3-C6 POSTERIOR CERVICAL FUSION, C3-C5 LAMINECTOMY AND C6 HEMILAMINECTOMY AND WAS DISCHARGED TO REHABILITATION HOSPITAL WITH AN ASPEN VISTAL CERVICAL COLLAR ON (B)(6) 2025. HE WAS READMITTED TO (B)(6) HOSPITAL ON (B)(6) 2025 BECAUSE OF DISLODGMENT OF C3 AND C4 SCREWS AND RE SUBLUXATION AT C4-5. ON (B)(6) 2025, HE UNDERWENT C3-6 REVISION OF FUSION WITH EXTENSION TO C2 WITH PLANS TO PERFORM ADDITIONAL ANTERIOR CERVICAL SPINE STABILIZATION THREE DAYS LATER. HE WAS PLACED ON OTTOBOCK 50C80 = MINERVA ORTHOSIS WHILE WAITING FOR COMPLETION OF UPCOMING SURGICAL INTERVENTION. HE DEVELOPED ASPIRATION INTO THE LOWER RESPIRATORY TRACT WITH ACUTE RESPIRATORY FAILURE REQUIRING INTUBATION AND MECHANICAL VENTILATORY SUPPORT. HE DEVELOPED PNEUMONIA AND HIS SURGICAL INTERVENTION WAS POSTPONED INDEFINITELY. HE WAS THEREFORE MAINTAINED ON THE OTTOBOCK 50C80 = MINERVA ORTHOSIS. ON (B)(6) 2025, HE WAS FOUND TO HAVE DEVELOPED PRESSURE INJURY TO BOTH SIDES OF THE MANDIBLE THAT OVER THE COURSE OF HIS HOSPITALIZATION UNTIL HIS DEATH ON (B)(6) 2025; PROGRESSED TO FULL THICKNESS LOSS OF COVERAGE OF BOTH SIDES OF THE MANDIBLE WITH EXPOSED HORIZONTAL MANDIBULAR RAMI BILATERALLY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22714196&pc=QFG | 3013756811-2025-184579 | 2025-05-14 04:00 | Death | TANDEM DIABETES CARE | 2025-08-05 04:00 | QFG | TANDEM MOBI INSULIN PUMP WITH CONTROL-IQ TECHNOLOGY | Adverse Event Without Identified Device or Use Problem | Diabetic Ketoacidosis | K233044 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT DIED ON (B)(6) 2025, DUE TO DIABETIC KETOACIDOSIS (DKA). A REVIEW OF PUMP DATA WILL BE PERFORMED, AND THE RESULTS WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22714388&pc=DSQ | 2916596-2025-05160 | 2025-07-26 04:00 | Death | THORATEC CORPORATION | 2025-08-05 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. THERE WERE NO DEVICE ISSUES AT THE TIME OF THE OUTCOME. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT'S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICLE ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22714682&pc=NVZ | 2182208-2025-03833 | 2024-01-01 05:00 | Death | MEDTRONIC, INC. | 2025-08-05 04:00 | NVZ | MEDTRONIC IPG | Appropriate Term/Code Not Available | Insufficient Information |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING CONDUCTION SYSTEM PACING FOR BRADYCARDIA. THE AUTHORS DESCRIBED PATIENT DEATHS; THE CAUSES OF DEATH WERE DESCRIBED AS ALL-CAUSE AS CARDIOVASCULAR, NON-CARDIOVASCULAR, OR UNKNOWN CAUSES. THERE WERE PATIENTS WHO EXPERIENCED PNEUMOTHORAX, POCKET INFECTIONS, ENDOCARDITIS, PERICARDIAL EFFUSION, INTERVENTRICULAR SEPTAL INJURY, SEVERE TRICUSPID REGURGITATION, VASCULAR EVENTS, PACING-INDUCED CARDIOMYOPATHY, HEART FAILURE HOSPITALIZATIONS, AND WOUND DEFECTS. THERE WERE LEADS WHICH EXHIBITED INCREASED AND HIGH THRESHOLDS, LOSS OF CAPTURE, AND LEAD DISLODGMENTS. SOME PATIENTS UNDERWENT REOPERATIONS FOR COMPLICATIONS. THE STATUS OF THE DEVICES AND LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/74 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: FEASIBILITY, SAFETY AND OUTCOMES OF CONDUCTION SYSTEM PACING FOR BRADYCARDIA AMONGST THE VERY ELDERLY. SCIENTIFIC REPORTS. 2024. 14:18755. DOI: 10.1038/S41598-024-69388-2 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22714698&pc=NVN | 2182208-2025-03834 | 2024-01-01 05:00 | Death | MEDTRONIC, INC. | 2025-08-05 04:00 | NVN | SELECTSECURE MRI SURESCAN | Appropriate Term/Code Not Available | Insufficient Information | P030036 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING CONDUCTION SYSTEM PACING FOR BRADYCARDIA. THE AUTHORS DESCRIBED PATIENT DEATHS; THE CAUSES OF DEATH WERE DESCRIBED AS ALL-CAUSE AS CARDIOVASCULAR, NON-CARDIOVASCULAR, OR UNKNOWN CAUSES. THERE WERE PATIENTS WHO EXPERIENCED PNEUMOTHORAX, POCKET INFECTIONS, ENDOCARDITIS, PERICARDIAL EFFUSION, INTERVENTRICULAR SEPTAL INJURY, SEVERE TRICUSPID REGURGITATION, VASCULAR EVENTS, PACING-INDUCED CARDIOMYOPATHY, HEART FAILURE HOSPITALIZATIONS, AND WOUND DEFECTS. THERE WERE LEADS WHICH EXHIBITED INCREASED AND HIGH THRESHOLDS, LOSS OF CAPTURE, AND LEAD DISLODGMENTS. SOME PATIENTS UNDERWENT REOPERATIONS FOR COMPLICATIONS. THE STATUS OF THE DEVICES AND LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/74 YEARS OLD. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: FEASIBILITY, SAFETY AND OUTCOMES OF CONDUCTION SYSTEM PACING FOR BRADYCARDIA AMONGST THE VERY ELDERLY. SCIENTIFIC REPORTS. 2024. 14:18755. DOI: 10.1038/S41598-024-69388-2 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22715045&pc=NIQ | 2024168-2025-08010 | 2016-04-01 04:00 | Death | ABBOTT VASCULAR | 2025-08-05 04:00 | NIQ | PROMUS EVEROLIMUS ELUTING CORONARY STENT SYSTEM | Adverse Event Without Identified Device or Use Problem | Unspecified Vascular Problem | P070015 |  | 1 | Event Description: THIS IS BEING FILED TO REPORT DEATH. ANGINA PECTORIS CAN OCCUR IN UP TO 40% OF PATIENTS FOLLOWING PERCUTANEOUS CORONARY INTERVENTION (PCI). THERE IS LIMITED DATA ASSESSING WHETHER THE TYPE OF STENT IMPLANTED DURING REVASCULARIZATION CAN PREDICT POST-PCI ANGINA SYMPTOMS. IN THIS STUDY, DATA REGARDING REVASCULARIZATION CHARACTERISTICS INCLUDING THE STENT TYPE IN PATIENTS ADMITTED FOR PERCUTANEOUS CORONARY INTERVENTION (PCI) WAS COLLECTED. PROSPECTIVE DATA INCLUDING OCCURRENCE OF ANGINA AND THE PRESENTING CLASS, NEW ONSET ST-SEGMENT ELEVATION MYOCARDIAL INFARCTION (STEMI), AND OTHER CLINICAL OUTCOMES WERE COLLECTED AT 1, 3, AND 6-MONTH FOLLOW-UP INTERVALS. A TOTAL OF 787 PATIENTS (64.5% MALES) UNDERGOING PCI WITH THREE STENT TYPES (ORSIRO, PROMUS, AND XIENCE) WERE INCLUDED IN THE STUDY. ADVERSE PATIENT EFFECTS RELATED TO THE PROMUS STENT INCLUDED: ANGINA, MYOCARDIAL INFARCTION, STROKE, REPEAT VASCULARIZATION, HOSPITALIZATION, AND DEATH. ADVERSE PATIENT EFFECTS RELATED TO THE XIENCE STENT INCLUDE: ANGINA, MYOCARDIAL INFARCTION, REPEAT VASCULARIZATION, AND HOSPITALIZATION. THE ARTICLE CONCLUDED THAT ALTHOUGH THE TYPE OF STENT MAY NOT HAVE AN ASSOCIATION WITH THE POST-PCI ANGINA, OTHER PREDICTORS SUCH AS DYSLIPIDEMIA AND PREVIOUS CAD AND HOSPITALIZATION MAY PREDICT RECURRENCE OF CARDIAC ANGINA. DETAILS ARE LISTED IN THE ARTICLE, TITLED ¿THE ASSOCIATION BETWEEN STENT TYPE AND DEVELOPING ANGINA PECTORIS FOLLOWING PERCUTANEOUS CORONARY INTERVENTION.¿ NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORDS FOR THIS PRODUCT COULD NOT BE REVIEWED AND A SIMILAR COMPLAINT QUERY COULD NOT BE PERFORMED BECAUSE THE PART NUMBER AND/OR LOT NUMBER WAS NOT REPORTED, AND THE DEVICE WAS NOT RETURNED. THE REPORTED PATIENT EFFECT OF DEATH IS LISTED IN THE PROMUS EVEROLIMUS ELUTING CORONARY STENT SYSTEMS INSTRUCTIONS FOR USE AS A KNOWN PATIENT EFFECT OF CORONARY STENTING PROCEDURES. A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECT, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING. LITERATURE ATTACHMENT: ARTICLE TITLE: ¿THE ASSOCIATION BETWEEN STENT TYPE AND DEVELOPING ANGINA PECTORIS FOLLOWING PERCUTANEOUS CORONARY INTERVENTION¿ B3: DATE OF DEATH ESTIMATED AS (B)(6) 2016 B3: DATE OF PROCEDURE ESTIMATED AS (B)(6) 2016 D4: THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. D6A: DATE OF IMPLANT ESTIMATED AS (B)(6) 2016 THE ADDITIONAL PATIENT EFFECTS REPORTED IN THE ARTICLE ARE CAPTURED UNDER SEPARATE MEDWATCH REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22715087&pc=NIQ | 2024168-2025-08012 | 2015-01-01 05:00 | Death | ABBOTT VASCULAR | 2025-08-05 04:00 | NIQ | XIENCE EVEROLIMUS ELUTING CORONARY STENT SYSTEM | Adverse Event Without Identified Device or Use Problem | Unspecified Vascular Problem | P070015 |  | 1 | Event Description: DRUG-COATED BALLOONS (DCBS) HAVE BECOME A NEW TREATMENT STRATEGY FOR CORONARY ARTERY DISEASE, AND THE DRUGS USED IN THE COATING AND THE COATING TECHNOLOGY HAVE PROGRESSED IN THE PAST YEARS. WITHOUT PERMANENT IMPLANT, A DCB DELIVERS ANTIPROLIFERATIVE DRUGS RAPIDLY AND UNIFORMLY INTO THE VESSEL WALL VIA THE EXCIPIENT DURING A SINGLE BALLOON DILATION. THIS ARTICLE IS META-ANALYSIS INVOLVING A TOTAL OF 39 STUDIES, WITH 9 STUDIES REPORTING THE XIENCE STENT. THE XIENCE STUDIES WILL BE CAPTURED FOR THE MENTIONED OUTCOMES OF THROMBOSIS, RESTENOSIS, TARGET LESION REVASCULARIZATION (TLR), MAJOR ADVERSE CARDIAC EVENT (MACE, INCLUDING A COMBINATION OF CARDIOVASCULAR EVENTS LIKE MYOCARDIAL INFARCTION (MI), STROKE, AND CARDIOVASCULAR DEATH AND/OR DEATH), AND DEATH. THE ARTICLE CONCLUDED DCB MAY BECOME AN IMPORTANT MEASURE FOR INTERVENTIONAL TREATMENT OF DE NOVO SVD¿S. DETAILS ARE LISTED IN THE ATTACHED ARTICLE, TITLED "CHINESE EXPERT CONSENSUS ON THE CLINICAL APPLICATION OF DRUGCOATED BALLOON (2ND EDITION).¿ NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORDS FOR THIS PRODUCT COULD NOT BE REVIEWED AND A SIMILAR COMPLAINT QUERY COULD NOT BE PERFORMED BECAUSE THE PART NUMBER AND/OR LOT NUMBER WAS NOT REPORTED, AND THE DEVICE WAS NOT RETURNED. THE REPORTED PATIENT EFFECT OF DEATH IS LISTED IN THE XIENCE V EVEROLIMUS ELUTING CORONARY STENT SYSTEMS INSTRUCTIONS FOR USE AS A KNOWN PATIENT EFFECT(S) OF CORONARY STENTING PROCEDURES. A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECT, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING. LITERATURE: ARTICLE TITLE: "CHINESE EXPERT CONSENSUS ON THE CLINICAL APPLICATION OF DRUGCOATED BALLOON (2ND EDITION)". B3: DATE OF EVENT/DEATH ESTIMATED AS (B)(6) 2015. D4: THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. D6A: DATE OF PROCEDURE/IMPLANT WILL BE ESTIMATED AS (B)(6) 2014. THE ADDITIONAL PATIENT EFFECTS REPORTED IN THE ARTICLE ARE CAPTURED UNDER SEPARATE MEDWATCH REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22715241&pc=NPT | 2025587-2025-05467 | 2024-11-14 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-08-05 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Insufficient Information | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 4 MONTHS FOLLOWING THE IMPLANT OF A TRANSCATHETER VALVE, THE PATIENT DIED DUE TO AN UNKNOWN REASON. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22717375&pc=DSQ | 2916596-2025-05134 | 2021-09-09 04:00 | Death | THORATEC CORPORATION | 2025-08-05 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTI-ORGAN FAILURE. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. THE PUMP WAS WORKING AS EXPECTED. NO PUMP RELATED ISSUES WERE PROVIDED. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. AN AUTOPSY WAS NOT PERFORMED, AND THE PUMP WAS NOT EXPLANTED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REV. C IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS, INCLUDING VARIOUS TYPES OF ORGAN FAILURES/DYSFUNCTIONS AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22717667&pc=DRE | 2124215-2025-54426 | 2025-07-14 04:00 | Death | BAYLIS MEDICALE CIE INC | 2025-08-05 04:00 | DRE | VERSACROSS CONNECT LAAC ACCESS SOLUTION | Positioning Problem | Hematoma; Low Blood Pressure/ Hypotension; Pain; Cardiac Perforation; Thrombosis/Thrombus | K220414 |  | 1 | Event Description: IT WAS REPORTED THAT A CARDIAC PERFORATION OCCURRED. THE PROCEDURE WAS CANCELLED. DURING A LEFT ATRIAL APPENDAGE CLOSURE (LAAC) PROCEDURE, A VERSACROSS CONNECT LAAC WAS SELECTED FOR USE. THE PHYSICIAN WAS HAVING DIFFICULTY TO PERFORM TRANSEPTAL AND STRUGGLED TO ENGAGE SEPTUM WITHIN THE FOSSA. IT WAS ATTEMPTED MULTIPLE TIMES BY READVANCING UP THE SUPERIOR VENA CAVA (SVC) WITH THE WIRE LEADING AND COMING DOWN TO ENGAGE THE FOSSA LIKE NORMAL. THERE WAS A FAIR AMOUNT OF DIFFICULTY VISUALIZING THE SEPTUM CLEARLY AND CONSISTENTLY BUT NO DIFFICULTY VISUALIZING THE EQUIPMENT. THE DILATOR WAS REPOSITIONED UP AND DOWN THE SEPTUM MANY TIMES TO TRY AND GET POSITIONED. THE PHYSICIAN PULLED OUT THE SHEATH AND RESHAPED IT TO GIVE IT GREATER BEND. AFTER ANOTHER FEW ATTEMPTS LIKE BEFORE, THE SHEATH WAS PULLED OUT ONE MORE TIME TO ADD MORE BEND PROXIMALLY. THIS TIME THE PHYSICIAN WAS ABLE TO ENGAGE THE SEPTUM AND GET ACROSS NORMALLY. HOWEVER ONCE ACROSS AND WERE ATTEMPTING TO MOVE THE VERSACROSS WIRE, THE ANESTHESIOLOGIST NOTICED SOMETHING IN THE RIGHT ATRIUM (RA). A LARGE FLUID SAC IN THE RA WAS NOTED AND THE DECISION WAS MADE TO ABORT THE CASE. THE SHEATH WAS WITHDRAWN, FLUID AND PRESSORS WERE GIVEN BY ANESTHESIA TO HELP INCREASE BLOOD PRESSURE, AN ARTERIAL LINE WAS STARTED, AND THEN THE PATIENT WAS TRANSFERRED TO THE INTENSIVE CARE UNIT (ICU) FOR MONITORING. CARDIOTHORACIC SURGERY WAS CONSULTED HOWEVER OPERATION WAS NOT RECOMMENDED DUE TO THE PATIENT AGE. IN THE PHYSICIAN OPINION, WHILE TRYING TO GO TRANSEPTAL THE PHYSICIAN INADVERTENTLY CAUSES SOME SORT OF INJURY TO RA WALL WHICH CAUSED BLEEDING AND A DISSECTION OF THE RA WALL, CREATING A SUB-LUMINAL HEMATOMA. IT WAS NOT A PERICARDIAL PERFUSION PER ANESTHESIA (TEE IMAGER) AND THE PHYSICIAN, BUT RATHER A TEAR IN THE RA WHICH FILLED WITH BLOOD AND EXPANDED INTO THE ATRIAL SPACE. THIS WAS SEEN ON ECHO AND NO OTHER CONFIRMATORY IMAGING WAS DONE TO OUR KNOWLEDGE. THE PATIENT WAS DOING OK, ACTUALLY AND THE THROMBUS WAS RECEDING IN SIZE, OFF PRESSORS, ON ANTICOAGULATION AND ANTIPLATELET THERAPY. THE FAMILY CHOSE FOR HIM TO GO COMFORT CARE. THE DEVICE IS NOT EXPECTED TO BE RETURNED FOR ANALYSIS (DISPOSED). Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. A SEPARATED REPORT FOR VERSACROSS RF WIRE IS BEING SUBMITTED. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION, BUT IT WAS NOT AVAILABLE. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION. GOOD FAITH EFFORT ATTEMPTS TO TRY AND RETRIEVE ADDITIONAL DETAILS REGARDING THE REPORTED EVENT ARE IN PROGRESS, BUT FURTHER INFORMATION WAS UNABLE TO BE OBTAINED YET. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22718900&pc=OZD | 1220648-2025-45716 | 2025-06-12 04:00 | Death | ABIOMED, INC. | 2025-08-05 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia | P140003 |  | 1 | Event Description: A NOTIFICATION VIA ABIOMED'S LONG-TERM OUTCOME AND QUALITY INDICATOR (LOQI) IMPELLA REGISTRY REGARDING A PATIENT (UNKNOWN ETHNICITY) EXPERIENCED LIMB ISCHEMIA ON THE IMPELLA REPLACEMENT DEVICE LEADING TO EXPLANT OF THE DEVICE. THE LIMB ISCHEMIA HAD A PROBABLE RELATIONSHIP TO THE REPLACEMENT IMPELLA CP DEVICE. FURTHER INFORMATION IN ADDITION TO THE STUDY NOTIFICATION NOTED THE PATIENT SUBSEQUENTLY EXPIRED. THE PATIENT WAS ON A NON-MONITORED FLOOR PREPARING FOR DISCHARGE AND HAD A SYNCOPAL EPISODE. THE PATIENT HIT HER HEAD (DEVELOPED A SUBARACHNOID HEMORRHAGE) AND ARRESTED. MULTIPLE ROUNDS OF CARDIOPULMONARY RESUSCITATION TOOK PLACE. COMPLETE HEART BLOCK WAS FOUND, AND A TRANSVENOUS PACER AND SWAN-GANZ CATHETER WERE PLACED. THE FIRST IMPELLA CP DEVICE WAS INSERTED, AND LOW PUMP FLOW WAS ENCOUNTERED. CONSEQUENTLY, THIS INITIALLY IMPLANTED IMPELLA CP DEVICE WAS EXPLANTED AND THE REPLACEMENT IMPELLA CP DEVICE WAS IMPLANTED. CORONARIES ARTERIES WERE ASSESSED AND NOTED TO BE CLEAR. THE PATIENT WAS TRANSFERRED TO THE INTENSIVE CARE UNIT FOR FURTHER MONITORING AND CARE. PATIENT UPDATES NOTED THERE WAS TROUBLESHOOTING FOR THROMBOSED SWAN CATHETER. SUBARACHNOID HEMORRHAGE WAS BEING CLOSELY MONITOR PENDING INITIATION OF HEPARIN. ALSO, THIS DAY IN THE MORNING, THE NURSE NOTICED THE PATIENT'S RIGHT LOWER EXTREMITY WAS COOL, AND THE RIGHT FOOT WAS WITHOUT PEDAL/POST TIBIAL PULSES. THE PHYSICIAN WAS NOTIFIED. THE RIGHT LOWER EXTREMITY BECAME DUSKIER AND COLDER IN THE AFTERNOON, AND THE PHYSICIAN WAS THEN CALLED TO THE BEDSIDE. CONSEQUENTLY, THE REPLACEMENT IMPELLA WAS REMOVED DUE TO CONCERNS FOR ARTERIAL OBSTRUCTION. POST-REMOVAL THE PATIENT HAD IMPROVED PERFUSION TO RIGHT LOWER EXTREMITY WITH 1+ DOPPLER PEDAL PULSE. PER THE REGISTRY, THE PATIENT RECOVERED WITH NO SEQUELAE. PATIENT UPDATES FURTHER NOTED THE PATIENT WAS STILL 100% ON FIO2. LACTATE WAS CLEAR, PRESSOR/INOTROPE REQUIREMENTS WERE STILL HIGH. ARTERIAL BLOOD GAS, CHEST X-RAY AND CONTINUOUS RENAL REPLACEMENT THERAPY WERE PENDING. HOWEVER, LATER THIS SAME DAY CARE WAS WITHDRAWN, AND THE PATIENT EXPIRED. Manufacturer Narrative: ALTHOUGH, THE PATIENT EXPERIENCED SERIOUS ADVERSE EVENT WITH SWAN CATHETER AND CARE WAS WITHDRAWN, THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE REPLACEMENT IMPELLA CP DEVICE FROM THE PATIENT'S OUTCOME GIVEN THE IMPELLA-RELATED LIMB ISCHEMIA EVENT. PRODUCT STATUS: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, AN EVALUATION/ANALYSIS OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: IMPELLA CP WITH SMARTASSIST SYSTEM. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): "ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY" (INCLUDING VENTRICULAR PERFORATION).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22718914&pc=DSQ | 2916596-2025-05241 | 2025-07-10 04:00 | Death | THORATEC CORPORATION | 2025-08-05 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) WAS TRIGGERING THE VENTRICULAR CONUNDRUM AND WAS PRONOUNCED BRAIN DEAD VIA A NUCLEAR MEDICINE SCAN. Manufacturer Narrative: SECTION A: PATIENT INFORMATION NOT YET AVAILABLE. SECTION D4: DEVICE SERIAL NUMBER WAS NOT PROVIDED, SO THE EXPIRATION DATE AND MANUFACTURE DATE (H4) CANNOT BE DETERMINED. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22718944&pc=DXF | 2124215-2025-54427 | 2025-07-14 04:00 | Death | BAYLIS MEDICALE CIE INC | 2025-08-05 04:00 | DXF | VERSACROSS CONNECT LAAC ACCESS SOLUTION | Adverse Event Without Identified Device or Use Problem | Hematoma; Low Blood Pressure/ Hypotension; Pain; Cardiac Perforation; Thrombosis/Thrombus | K150709 |  | 1 | Event Description: IT WAS REPORTED THAT A CARDIAC PERFORATION OCCURRED. THE PROCEDURE WAS CANCELLED. DURING A LEFT ATRIAL APPENDAGE CLOSURE (LAAC) PROCEDURE, A VERSACROSS CONNECT LAAC WAS SELECTED FOR USE. THE PHYSICIAN WAS HAVING DIFFICULTY TO PERFORM TRANSEPTAL AND STRUGGLED TO ENGAGE SEPTUM WITHIN THE FOSSA. IT WAS ATTEMPTED MULTIPLE TIMES BY READVANCING UP THE SUPERIOR VENA CAVA (SVC) WITH THE WIRE LEADING AND COMING DOWN TO ENGAGE THE FOSSA LIKE NORMAL. THERE WAS A FAIR AMOUNT OF DIFFICULTY VISUALIZING THE SEPTUM CLEARLY AND CONSISTENTLY BUT NO DIFFICULTY VISUALIZING THE EQUIPMENT. THE DILATOR WAS REPOSITIONED UP AND DOWN THE SEPTUM MANY TIMES TO TRY AND GET POSITIONED. THE PHYSICIAN PULLED OUT THE SHEATH AND RESHAPED IT TO GIVE IT GREATER BEND. AFTER ANOTHER FEW ATTEMPTS LIKE BEFORE, THE SHEATH WAS PULLED OUT ONE MORE TIME TO ADD MORE BEND PROXIMALLY. THIS TIME THE PHYSICIAN WAS ABLE TO ENGAGE THE SEPTUM AND GET ACROSS NORMALLY. HOWEVER ONCE ACROSS AND WERE ATTEMPTING TO MOVE THE VERSACROSS WIRE, THE ANESTHESIOLOGIST NOTICED SOMETHING IN THE RIGHT ATRIUM (RA). A LARGE FLUID SAC IN THE RA WAS NOTED AND THE DECISION WAS MADE TO ABORT THE CASE. THE SHEATH WAS WITHDRAWN, FLUID AND PRESSORS WERE GIVEN BY ANESTHESIA TO HELP INCREASE BLOOD PRESSURE, AN ARTERIAL LINE WAS STARTED, AND THEN THE PATIENT WAS TRANSFERRED TO THE INTENSIVE CARE UNIT (ICU) FOR MONITORING. CARDIOTHORACIC SURGERY WAS CONSULTED HOWEVER OPERATION WAS NOT RECOMMENDED DUE TO THE PATIENT AGE. IN THE PHYSICIAN OPINION, WHILE TRYING TO GO TRANSEPTAL THE PHYSICIAN INADVERTENTLY CAUSES SOME SORT OF INJURY TO RA WALL WHICH CAUSED BLEEDING AND A DISSECTION OF THE RA WALL, CREATING A SUB-LUMINAL HEMATOMA. IT WAS NOT A PERICARDIAL PERFUSION PER ANESTHESIA (TEE IMAGER) AND THE PHYSICIAN, BUT RATHER A TEAR IN THE RA WHICH FILLED WITH BLOOD AND EXPANDED INTO THE ATRIAL SPACE. THIS WAS SEEN ON ECHO AND NO OTHER CONFIRMATORY IMAGING WAS DONE TO OUR KNOWLEDGE. THE PATIENT WAS DOING OK, ACTUALLY AND THE THROMBUS WAS RECEDING IN SIZE, OFF PRESSORS, ON ANTICOAGULATION AND ANTIPLATELET THERAPY. THE FAMILY CHOSE FOR HIM TO GO COMFORT CARE. THE DEVICE IS NOT EXPECTED TO BE RETURNED FOR ANALYSIS (DISPOSED). Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. A SEPARATED REPORT FOR VERSACROSS DILATOR IS BEING SUBMITTED. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION, BUT IT WAS NOT AVAILABLE. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION. GOOD FAITH EFFORT ATTEMPTS TO TRY AND RETRIEVE ADDITIONAL DETAILS REGARDING THE REPORTED EVENT ARE IN PROGRESS, BUT FURTHER INFORMATION WAS UNABLE TO BE OBTAINED YET. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22719130&pc=QSK | 1063481-2025-00054 | 2025-03-27 04:00 | Death | ASCYRUS MEDICAL GMBH | 2025-08-05 04:00 | QSK | AMDS DISSECTION STENT - PROX 40-DIS 40 | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest |  |  | 1 | Event Description: THIS EVENT IS RELATED TO A POST-MARKET CLINICAL STUDY "PROTECT" AND REPORTED RETROSPECTIVELY. ACCORDING TO REPORTS FROM THE PROTECT STUDY, PATIENT EXPERIENCED A SERIOUS ADVERSE EVENT (SAE). AMDS 40-40, LOT# C2460019B, PROCEDURE DATE: (B)(6) 2023, SERIOUS ADVERSE EVENT (SAE), EVENT ONSET DATE ¿ 27MAR2025, EVENT END DATE ¿ 18JUN2025. SEQUENCE NUMBER: 6, EVENT ONSET DATE: 27MAR2025, IS THE EVENT ONGOING: NO, EVENT END DATE: 18JUN2025, WAS THIS EVENT EXPECTED: NO. EVENT: VASCULAR SPECIFY OTHER: PROGRESSION OF THE AORTIC DIAMETER DUE TO PERSISTING PERFUSION OF THE FALSE LUMEN DESCRIBE EVENT: PROGRESSION OF THE AORTIC DIAMETER DUE TO PERSISTING PERFUSION OF THE FALSE LUMEN FROM ZONE 4 ONWARDS. INDICATE RELATION TO AMDS DEVICE: POSSIBLY . INDICATE RELATION TO AMDS IMPLANT PROCEDURE: POSSIBLY. DID A PRE-EXISTING CONDITION OR THE ACUTE DISEASE CAUSE OR CONTRIBUTE TO THE EVENT: YES PLEASE SPECIFY PRE-EXISTING DISEASE: DEBAKEY TYPE A DISSECTION. DID THE EVENT LEAD TO A SERIOUS DETERIORATION IN HEALTH THAT RESULTED IN ANY OF THE FOLLOWING: DEATH, LIFE-THREATENING ILLNESS OR INJURY, PERMANENT IMPAIRMENT OF A BODY STRUCTURE OR FUNCTION, IN-PATIENT HOSPITALIZATION OR PROLONGED EXISTING HOSPITALIZATION, MEDICAL OR SURGICAL INTERVENTION TO PREVENT IMPAIRMENT OF A BODY STRUCTURE OR FUNCTION? YES PLEASE SPECIFY: DEATH LIFE-THREATENING ILLNESS OR INJURY PERMANENT IMPAIRMENT OF A BODY STRUCTURE OR FUNCTION. IS THE SUBJECT HOSPITALIZED DUE TO THIS AE: NO. DID THE DEVICE MALFUNCTION OR DETERIORATE IN CHARACTERISTICS OR PERFORMANCE: NO COULD THIS EVENT HAVE LED TO DEATH OR SERIOUS DETERIORATION IN HEALTH HAD SUITABLE INTERVENTION NOT OCCURRED: YES. DID THE SUBJECT EXIT THE STUDY: YES EVENT OUTCOME ¿ DEATH WAS THERE ANY TREATMENT FOR THE EVENT? NO. THIS EVENT IS RELEGATED TO AMDS 40-40, LOT# C2460019B FOR PROGRESSION OF THE AORTIC DIAMETER DUE TO PERSISTING PERFUSION OF THE FALSE LUMEN FROM ZONE 4 ONWARDS AND DEATH. ADDITIONAL INFO RECEIVED. EXACT CAUSE OF DEATH IS UNKNOWN, BUT THE FOLLOWING CONTRIBUTING FACTORS ARE MENTIONED: PNEUMONIA . POSSIBLE AORTIC RUPTURE, AS PATIENT HAD ABDOMINAL PAIN AND THEN SUDDENLY PULSELESS ELECTRICAL ACTIVITY. MEDICAL/SURGICAL/PHARMACEUTICAL INTERVENTIONS: NO INTERVENTIONS WERE PERFORMED FOR THE PROGRESSION OF THE AORTIC DIAMETER. PER THE MEDICAL RECORDS, DEATH DUE TO CARDIAC ARREST. Manufacturer Narrative: PLEASE NOTE THE HDE NUMBER FOR THIS DEVICE - H230007. THIS INVESTIGATION IS CURRENTLY ONGOING. ANY ADDITIONAL INFORMATION WILL BE PROVIDED IN THE FOLLOW-UP REPORT. THIS REPORT IS BEING SUBMITTED AS REQUIRED BY FEDERAL REGULATIONS AND DOES NOT CONSTITUTE AN ADMISSION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. FURTHERMORE, THIS REPORT REFLECTS THE EVENT AS ALLEGED BY THE COMPLAINANT AND DOES NOT IMPLY THAT THE INFORMATION REPORTED TO ARTIVION ¿ FORMERLY CRYOLIFE/JOTEC IS ACCURATE OR HAS BEEN CONFIRMED BY ARTIVION ¿ FORMERLY CRYOLIFE/JOTEC.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22719185&pc=OZD | 1220648-2025-31033 | 2025-07-21 04:00 | Death | ABIOMED, INC. | 2025-08-05 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia; Septic Shock | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT ON IMPELLA CP SUPPORT LOST PEDAL PULSE AND A DISTAL PERFUSION CATHETER WAS PLACED. THE PATIENT ALSO WAS IN SEPTIC SHOCK. THE FAMILY DECIDED TO WITHDRAW CARE AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22719188&pc=MIH | 2247858-2025-00185 | 2025-07-11 04:00 | Death | BOLTON MEDICAL, INC. | 2025-08-05 04:00 | MIH | RELAY PRO THORACIC STENT-GRAFT SYSTEM | Adverse Event Without Identified Device or Use Problem; Insufficient Information | Insufficient Information | P200045 |  | 1 | Event Description: "DEATH SERIOUS AE · UNDETERMINED AS RELATED TO PROCEDURE · UNDETERMINED AS RELATED TO DEVICE · UNDETERMINED AS RELATED TO PREEXISTING CONDITION. ADDITIONAL INFORMATION RECEIVED ON JULY 21, 2025: WE DON'T HAVE SPECIFIC INFORMATION ABOUT THE CAUSE OF DEATH OR ADVERSE EVENT BECAUSE THE PARTICIPANT WAS TREATED AND PASSED AWAY IN ANOTHER HOSPITAL (I'M NOT EVEN SURE WHICH HOSPITAL BECAUSE THERE ISN'T ANY INFORMATION ON RECORDS EITHER)." PATIENT OUTCOME: "DEATH SERIOUS AE UNANTICIPATED AE · UNDETERMINED AS RELATED TO PROCEDURE. · UNDETERMINED AS RELATED TO DEVICE. · UNDETERMINED AS RELATED TO PREEXISTING CONDITION". Manufacturer Narrative: BOLTON MEDICAL, INC. (D/B/A TERUMO AORTIC), HEREIN KNOWN AS THE "COMPANY", IS SUBMITTING THIS REPORT PURSUANT TO 21 CFR PART 803, HAS MADE REASONABLE EFFORTS TO OBTAIN COMPLETE INFORMATION, AND HAS PROVIDED AS MUCH AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT IS BASED ON INFORMATION OBTAINED BY THE COMPANY, WHICH MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE REGULATIONS. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN, UNAVAILABLE, OR NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY ANYONE THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE CERTAIN REGULATIONS, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22719211&pc=KNS | 3005099803-2025-03853 | 2025-07-07 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-08-05 04:00 | KNS | AXIOS? | Adverse Event Without Identified Device or Use Problem | Cough | K233318 |  | 1 | Event Description: IT WAS REPORTED TO BOSTON SCIENTIFIC CORPORATION THAT AN AXIOS STENT AND ELECTROCAUTERY ENHANCED DELIVERY SYSTEM WAS IMPLANTED TRANSGASTRIC TO THE GALLBLADDER DURING A GALLBLADDER DRAINAGE (GBD) PROCEDURE PERFORMED ON (B)(6) 2025. DURING THE PROCEDURE, THE AXIOS STENT WAS SUCCESSFULLY POSITIONED WITHOUT ANY ISSUES. HOWEVER, IT WAS OBSERVED THAT THE PATIENT EXPERIENCED FREQUENT COUGHING DURING THE PROCEDURE, AND MAINTAINING SCOPE POSITIONING WAS PARTICULARLY DIFFICULT. ON (B)(6) 2025, THE PATIENT PASSED AWAY. ACCORDING TO THE PHYSICIAN'S ASSESSMENT, THE PATIENT'S POOR GENERAL CONDITION, ALONG WITH THE MEDICATIONS ADMINISTERED DURING THE ENDOSCOPIC PROCEDURE, MAY HAVE PLACED A SIGNIFICANT PHYSIOLOGICAL STRESS ON THE BODY, POTENTIALLY CONTRIBUTING TO THE ADVERSE OUTCOME. Manufacturer Narrative: BLOCK D4, H4: THE COMPLAINANT WAS UNABLE TO PROVIDE THE SUSPECT DEVICE LOT NUMBER. THEREFORE, THE MANUFACTURE AND EXPIRATION DATES ARE UNKNOWN. THE COMPLAINANT WAS UNABLE TO PROVIDE THE COMPLAINT DEVICE LOT NUMBER. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER PRODUCT SPECIFIC INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. BLOCK H6: IMDRF IMPACT CODE F02 CAPTURES THE REPORTABLE EVENT OF THE PATIENT PASSED AWAY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22719692&pc=OZD | 1220648-2025-31035 | 2025-07-22 04:00 | Death | ABIOMED, INC. | 2025-08-05 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hematoma | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT ON IMPELLA CP SUPPORT DEVELOPED A HEMATOMA AT THE IMPELLA ACCESS SITE. PRESSURE WAS APPLIED, HEPARIN WAS WITHHELD, AND THE DRESSING WAS RE-DRESSED WITH HIGHER ANGLE MATCHING TO RESOLVED THE ISSUE. LATER, CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM: SECTION: GENERAL PATIENT CARE CONSIDERATIONS: "ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA." SECTION: ENTERING CARDIAC OUTPUT: USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: "REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE." SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): "ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22719738&pc=LWP | 2124215-2025-54493 | 2025-07-08 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-08-05 04:00 | LWP | ACCOLADE MRI EL DR | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT IMPLANT OF A PACEMAKER SYSTEM ON (B)(6) 2025, FOR THIRD DEGREE ATRIOVENTRICULAR BLOCK (ESCAPE RATE OF APPROXIMATELY 35 BEATS/MINUTE). THE PATIENT WAS DISCHARGED FROM THE HOSPITAL ON THE AFTERNOON OF (B)(6) 2025, AND WAS SUBSEQUENTLY FOUND DECEASED. THE TIME OF DEATH WAS ESTIMATED TO BE BETWEEN 6 PM AN 10 PM THE EVENING OF THE (B)(6). THE PATIENT WAS FOUND WITH BLOOD VISIBLY EXITING THEIR MOUTH. IT WAS PLANNED TO EXPLANT THE PACEMAKER ON (B)(6) 2025 AND INTERROGATE THE DEVICE. FOLLOW-UP WITH THE FIELD INDICATED THAT THE HOSPITAL REGISTRAR REQUESTED THAT THE CORONER CHECK THE DEVICE AT EXPLANT. NO FURTHER INFORMATION HAS BEEN REPORTED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22719753&pc=LWP | 2124215-2025-54510 | 2025-06-30 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-08-05 04:00 | LWP | ACCOLADE MRI DR | Defective Device | No Clinical Signs, Symptoms or Conditions | P150012/S000 |  | 1 | Event Description: IT WAS REPORTED BY THE PATIENT'S SON THAT THE PATIENT PASSED AWAY ON (B)(6) 2025. THE CAUSE OF DEATH WAS REPORTED BY THE CORONER TO BE AORTA INFLAMMATION, KIDNEY FAILURE AND PACEMAKER FAILURE. THE PACEMAKER WAS EXPLANTED BY THE FUNERAL HOME AND DISCARDED IN PREPARATION FOR THE PATIENT'S CREMATION. THE LAST LATITUDE TRANSMISSION HAD BEEN IN (B)(6) 2025, AND CURRENT DEVICE DATA HAD NOT BEEN PROVIDED TO BOSTON SCIENTIFIC FOR EVALUATION. NO FURTHER INFORMATION WAS PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22720028&pc=LWS | 2649622-2025-21267 | 2025-05-25 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-08-05 04:00 | LWS | SPRINT QUATTRO SECURE S MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY TWENTY SIX DAYS AFTER AN UPGRADE FROM IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM TO CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) SYSTEM. THE CAUSE OF DEATH WAS REQUESTED AND HAS NOT BEEN RECEIVED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTION: A1, A2, A3. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22720047&pc=NIK | 9614453-2025-03101 | 2025-05-25 04:00 | Death | MEDTRONIC EUROPE SARL | 2025-08-05 04:00 | NIK | COBALT¿ XT HF CRT-D MRI SURESCAN¿ | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P010031 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY TWENTY SIX DAYS AFTER AN UPGRADE FROM IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM TO CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) SYSTEM. THE CAUSE OF DEATH WAS REQUESTED AND HAS NOT BEEN RECEIVED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTION: A1, A2, A3. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22720059&pc=NVN | 2649622-2025-21269 | 2025-05-25 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-08-05 04:00 | NVN | SELECTSECURE MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P030036 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY TWENTY SIX DAYS AFTER AN UPGRADE FROM IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM TO CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) SYSTEM. THE CAUSE OF DEATH WAS REQUESTED AND HAS NOT BEEN RECEIVED. Manufacturer Narrative: CORRECTION: A1, A2, A3. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22720086&pc=DSQ | 2916596-2025-05164 | 2025-07-24 04:00 | Death | THORATEC CORPORATION | 2025-08-05 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2025. DEVICE PARAMETERS WERE NOT WITHIN NORMAL LIMITS FOR THIS PATIENT, BUT THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. AN AUTOPSY WAS NOT PERFORMED, AND THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22720296&pc=NPS | 2135147-2025-04352 | 2016-04-01 04:00 | Death | ABBOTT MEDICAL | 2025-08-05 04:00 | NPS | TRICLIP DELIVERY SYSTEM (TCDS) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P230007 |  | 1 | Event Description: THE ARTICLE "BEYOND 2-DIMENSIONAL ECHOCARDIOGRAPHY: A NOVEL MULTIPARAMETRIC ASSESSMENT OF RIGHT VENTRICULAR DYSFUNCTION IN TRANSCATHETER TRICUSPID VALVE REPAIR" WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE SINGLE CENTER STUDY, TO EVALUATE RVD BY A NOVEL RVPAC PARAMETER, INCLUDING THE 3 IMPORTANT DRIVERS OF RVD, FOR AN IMPROVED PREDICTION OF 1-YEAR MORTALITY AFTER T-TEER. DEVICES MENTIONED INCLUDE MITRACLIP, TRICLIP AND NON-ABBOTT DEVICE. THE ARTICLE CONCLUDED THAT THE NOVEL RVPAC PARAMETER, INTEGRATING RV FUNCTION, VOLUME STRESS, AND PRESSURE STRESS IS A POWERFUL METRIC FOR RV FAILURE AND A SUPERIOR PREDICTOR FOR SURVIVAL POST-T-TEER. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS PHILIPP M. DOLDI AT KLINIKUM DER UNIVERSITÄT MÜNCHEN INSTITUTION WITH CORRESPONDING EMAIL PHILIPP.DOLDI@MED.UNI-MUENCHEN.DE]. THE TIME FRAME OF THE STUDY WAS APRIL 2016 TO FEBRUARY 2022. A TOTAL OF 262 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH AN UNKNOWN AMOUNT RECEIVED AN ABBOTT DEVICE. AS THIS EVENT IS FROM A LITERATURE REVIEW, THERE IS NO RELEVANT PATIENT INFORMATION (DATE OF BIRTH, AGE, GENDER, WEIGHT, AND MEDICAL HISTORY) TO REPORT. COMORBIDITIES INCLUDED CORONARY ARTERY DISEASE, ATRIAL FIBRILLATION, PREVIOUS PCI. (B)(4), UNK MITRACLIP. PERI- AND POST-PROCEDURAL COMPLICATIONS INCLUDED DEATH. (B)(4), UNK TRICLIP. PERI- AND POST-PROCEDURAL COMPLICATIONS INCLUDED DEATH. Manufacturer Narrative: IT IS UNKNOWN IF THE DEVICE IS RETURNING FOR ANALYSIS. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22720298&pc=NKM | 2135147-2025-04351 | 2016-04-01 04:00 | Death | ABBOTT MEDICAL | 2025-08-05 04:00 | NKM | MITRACLIP® SYSTEM CLIP DELIVERY SYSTEM | Off-Label Use | No Clinical Signs, Symptoms or Conditions | P100009 |  | 1 | Event Description: THE ARTICLE "BEYOND 2-DIMENSIONAL ECHOCARDIOGRAPHY: A NOVEL MULTIPARAMETRIC ASSESSMENT OF RIGHT VENTRICULAR DYSFUNCTION IN TRANSCATHETER TRICUSPID VALVE REPAIR" WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE SINGLE CENTER STUDY, TO EVALUATE RVD BY A NOVEL RVPAC PARAMETER, INCLUDING THE 3 IMPORTANT DRIVERS OF RVD, FOR AN IMPROVED PREDICTION OF 1-YEAR MORTALITY AFTER T-TEER. DEVICES MENTIONED INCLUDE MITRACLIP, TRICLIP AND NON-ABBOTT DEVICE. THE ARTICLE CONCLUDED THAT THE NOVEL RVPAC PARAMETER, INTEGRATING RV FUNCTION, VOLUME STRESS, AND PRESSURE STRESS IS A POWERFUL METRIC FOR RV FAILURE AND A SUPERIOR PREDICTOR FOR SURVIVAL POST-T-TEER. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS PHILIPP M. DOLDI AT KLINIKUM DER UNIVERSITÄT MÜNCHEN INSTITUTION WITH CORRESPONDING EMAIL PHILIPP.DOLDI@MED.UNI-MUENCHEN.DE]. THE TIME FRAME OF THE STUDY WAS APRIL 2016 TO FEBRUARY 2022. A TOTAL OF 262 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH AN UNKNOWN AMOUNT RECEIVED AN ABBOTT DEVICE. AS THIS EVENT IS FROM A LITERATURE REVIEW, THERE IS NO RELEVANT PATIENT INFORMATION (DATE OF BIRTH, AGE, GENDER, WEIGHT, AND MEDICAL HISTORY) TO REPORT. COMORBIDITIES INCLUDED CORONARY ARTERY DISEASE, ATRIAL FIBRILLATION, PREVIOUS PCI. (B)(4), UNK MITRACLIP PERI- AND POST-PROCEDURAL COMPLICATIONS INCLUDED DEATH. (B)(4), UNK TRICLIP PERI- AND POST-PROCEDURAL COMPLICATIONS INCLUDED DEATH. Manufacturer Narrative: IT IS UNKNOWN IF THE DEVICE IS RETURNING FOR ANALYSIS. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22720301&pc=NIQ | 2024168-2025-08075 | 2015-10-30 04:00 | Death | ABBOTT VASCULAR INC. | 2025-08-05 04:00 | NIQ | XIENCE SKYPOINT¿ | Adverse Event Without Identified Device or Use Problem | Unspecified Heart Problem | P110019 |  | 1 | Event Description: IT WAS REPORTED IN A RESEARCH SUMMARY ARTICLE THAT THE STOPDAPT TRIAL IS A PROSPECTIVE MULTI-CENTER SINGLE-ARM STUDY EVALUATING 3-MONTH DUAL ANTIPLATELET THERAPY (DAPT) DAPT DURATION AFTER COBALT-CHROMIUM EVEROLIMUS-ELUTING STENT (COCR-EES) (XIENCE SKYPOINT) IMPLANTATION. THE PRIMARY ENDPOINT WAS A COMPOSITE OF CARDIOVASCULAR DEATH, MYOCARDIAL INFARCTION (MI), STROKE, DEFINITE STENT THROMBOSIS (ST) AND TIMI MAJOR/MINOR BLEEDING AT 1 YEAR. BETWEEN SEPTEMBER 2012 AND OCTOBER 2013, A TOTAL OF 1525 PATIENTS WERE ENROLLED FROM 58 JAPANESE CENTERS, WITH COMPLETE 1-YEAR FOLLOW-UP IN 1519 PATIENTS (99.6 %). RESULTS INCLUDED DEATH, MYOCARDIAL INFARCTION, STROKE, BLEEDING, POSSIBLE STENT THROMBOSIS, TARGET LESION REVASCULARIZATION AND CORONARY ARTERY BYPASS GRAFTING (CABG). THE RESULTS FROM THE STOPDAPT TRIAL WERE COMPARED TO THE RESET TRIAL AS A HISTORICAL COMPARISON GROUP. THE (RANDOMIZED EVALUATION OF SIROLIMUS-ELUTING VERSUS EVEROLIMUS-ELUTING STENT TRIAL) RESET TRIAL IS A RANDOMIZED CONTROLLED TRIAL COMPARING COCR-EES WITH SIROLIMUS-ELUTING STENT CONDUCTED BY THE SAME STUDY GROUP IN 2010). RESULTS FROM THE RESET TRIAL INCLUDED DEATH, MYOCARDIAL INFARCTION, STROKE, BLEEDING, STENT THROMBOSIS, TARGET LESION REVASCULARIZATION, AND CABG. IT WAS FOUND THAT STOPPING DAPT AT 3 MONTHS IN SELECTED PATIENTS AFTER COCR- EES IMPLANTATION WAS AT LEAST AS SAFE AS THE PROLONGED DAPT REGIMEN ADOPTED IN THE HISTORICAL CONTROL GROUP. ADDITIONAL INFORMATION CAN BE FOUND IN THE ARTICLE "ONE-YEAR OUTCOME OF A PROSPECTIVE TRIAL STOPPING DUAL ANTIPLATELET THERAPY AT 3 MONTHS AFTER EVEROLIMUS-ELUTING COBALT-CHROMIUM STENT IMPLANTATION: SHORTT AND OPTIMAL DURATION OF DUAL ANTIPLATELET THERAPY AFTER EVEROLIMUS-ELUTING COBALT-CHROMIUM STENT (STOPDAPT) TRIAL." Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORDS FOR THIS PRODUCT COULD NOT BE REVIEWED AND A SIMILAR COMPLAINT QUERY COULD NOT BE PERFORMED BECAUSE THE PART NUMBER AND/OR LOT NUMBER WERE NOT REPORTED, AND THE DEVICE WAS NOT RETURNED. A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECT OF DEATH AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED. THE INFORMATION RECEIVED IN THE COMPLAINT WAS OBTAINED THROUGH A PROACTIVE LITERATURE SEARCH. SPECIFIC MODEL AND LOT NUMBERS WERE NOT AVAILABLE AND CAUSALITY BETWEEN THE DOCUMENTED DEVICE USAGE AND THE PATIENT EFFECTS COULD NOT BE ESTABLISHED. FACTORS THAT MAY CONTRIBUTE TO THE PERFORMANCE OUTCOMES LISTED IN THE ARTICLE INCLUDE, BUT ARE NOT LIMITED TO, DEVICE TO PATIENT INTERACTION, PATIENT ANATOMICAL CONDITIONS, DEVICE TO USER INTERACTION AND MANUFACTURING OR MATERIAL ISSUES. DUE TO THE LIMITED INFORMATION AVAILABLE, THE EXACT CAUSE CANNOT BE DETERMINED. THE REPORTED PATIENT EFFECT OF DEATH, IS LISTED IN THE XIENCE SKYPOINT, EVEROLIMUS ELUTING CORONARY STENT SYSTEM (EECSS), ELECTRONIC INSTRUCTIONS FOR USE (EIFU) AS A KNOWN PATIENT EFFECTS OF CORONARY STENTING PROCEDURES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN, OR LABELING; THEREFORE, NO PRODUCT-RELATED CORRECTIVE ACTION WILL BE IMPLEMENTED IN THIS CASE. B2, B3: ESTIMATED DATE. D4: THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. D6A: ESTIMATED DATE. THE ADDITIONAL PATIENT EFFECTS REPORTED IN THE ARTICLE ARE CAPTURED UNDER A SEPARATE MEDWATCH REPORT. ATTACHMENT ARTICLE TITLED: "ONE-YEAR OUTCOME OF A PROSPECTIVE TRIAL STOPPING DUAL ANTIPLATELET THERAPY AT 3 MONTHS AFTER EVEROLIMUS-ELUTING COBALT-CHROMIUM STENT IMPLANTATION: SHORTT AND OPTIMAL DURATION OF DUAL ANTIPLATELET THERAPY AFTER EVEROLIMUS-ELUTING COBALT-CHROMIUM STENT (STOPDAPT) TRIAL."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22720774&pc=DSQ | 2916596-2025-05242 | 2024-04-23 04:00 | Death | THORATEC CORPORATION | 2025-08-05 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS NOT PROVIDED. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. DUE TO PATIENT PRIVACY LAWS THE MANAGING HOSPITAL DID NOT COMMUNICATE PATIENT OUTCOME INFORMATION. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6) , AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. AN AUTOPSY WAS NOT PERFORMED, AND THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22720820&pc=DRA | 2135147-2025-04357 | 2025-07-11 04:00 | Death | ABBOTT MEDICAL | 2025-08-05 04:00 | DRA | MITRACLIP | Improper or Incorrect Procedure or Method | Cardiac Arrest; Hemorrhage/Bleeding; Low Blood Pressure/ Hypotension; Tachycardia; Vascular Dissection; Swelling/ Edema | K190167 |  | 1 | Event Description: IT WAS REPORTED ON (B)(6) 2025 THAT THE PATIENT PRESENTED WITH GRADE 4 FUNCTIONAL MITRAL REGURGITATION (MR), OXYGEN, MONITOR, INFUSION(OMI) AND LOW EJECTION FRACTION (EF) FOR A MITRACLIP PROCEDURE. DURING THE PROCEDURE, ONE MITRACLIP WAS IMPLANTED SUCCESSFULLY. DURING STEERABLE GUIDE CATHETER(SGC) REMOVAL, BLOOD PRESSURE DECREASED AND HEMODYNAMIC DISRUPTION WAS OBSERVED. THERE WAS A GRADUAL DECREASE IN BLOOD PRESSURE DURING THE OPERATION, AND THE BLOOD PRESSURE FELL BELOW 50 AFTER REMOVAL. PATIENT¿S HEARTRATE WAS OVER 100. HEMATURIA WAS ALSO CONFIRMED AND PREPARATIONS FOR BLOOD TRANSFUSION WERE MADE. SUSPECTING VENOUS DAMAGE, THE SAME SGC WAS REINSERTED AND THE BLOOD PRESSURE REMAINED IN THE 70S. A BALLOON WAS INSERTED FROM THE LEFT SIDE AND CONTRAST WAS PERFORMED. BLEEDING WAS CONFIRMED FROM THE FEMORAL VEIN NEAR THE PUNCTURE SITE TO THE SITE NEAR THE ILIAC VEIN. A 24FR DRY SEAL(DSF2433 WAS USED AT FEMORAL SITE AND TO STOP THE BLEEDING VIABAHN (REF BXAL087901J 8MM×79MM) WAS USED. AN INCREASE IN ABDOMINAL PRESSURE WAS OBSERVED AND PATIENT BECAME HEMODYNAMICALLY UNSTABLE. VENTRICULAR FIBRILLATION WAS FOLLOWED BY CARDIAC ARREST. PERCUTANEOUS CARDIOPULMONARY SUPPORT(PCPS)WAS PERFORMED TO MAKE PATIENT HEMODYNAMICALLY STABLE. AN INCISION WAS MADE IN ABDOMEN AND DRAINAGE WAS PERFORMED. THERE WAS NO CLINICALLY SIGNIFICANT DELAY. ON (B)(6) 2025, IT WAS REPORTED THAT THE PATIENT DEATH OCCURRED ON (B)(6) 2025 DUE TO BLEEDING FROM PUNCTURE SITE CAUSED BY THE SGC. PER THE PHYSICIAN, THERE WAS SIGNIFICANT BLEEDING, AND PCPS WAS UNABLE TO BE ESTABLISHED. THE PRE- EXISTING CONDITION OF PATIENT MAY HAVE ALSO CONTRIBUTED TO DEATH. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: THE DEVICE IS NOT RETURNING FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22720840&pc=LWP | 2124215-2025-54500 | 2025-07-10 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-08-05 04:00 | LWP | ACUITY? X4 SPIRAL S | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Cardiogenic Shock; Pericardial Effusion | P010012 |  | 1 | Event Description: IT WAS REPORTED THAT THIS PATIENT UNDERWENT AN IMPLANT PROCEDURE FOR A CARDIAC RESYNCHRONIZATION THERAPY PACEMAKER (CRT-P). THE RIGHT VENTRICULAR (RV) LEAD WAS IMPLANTED SUCCESSFULLY AND THEN THE LEFT VENTRICULAR (LV) LEAD WAS ATTEMPTED. THE PATIENT'S BLOOD PRESSURE DROPPED AND THE PATIENT CODED. AN ECHOCARDIOGRAM SHOWED A PERICARDIAL EFFUSION. A PERICARDIOCENTESIS WAS PERFORMED AND 150CC OF FLUID WAS REMOVED. THE PATIENT'S BLOOD PRESSURE IMPROVED; HOWEVER, POST-REMOVAL OF THE EFFUSION THE RV WAS OBSERVED TO BE COMPRESSED WITH A CLOT PRESENT AND THE PATIENT WAS TRANSFERRED TO THE INTENSIVE CARE UNIT (ICU). IT WAS NOTED A SINGLE CHAMBER PACEMAKER HAD BEEN CONNECTED TO THE IMPLANTED RV LEAD TO COMPLETE THE PROCEDURE. WHEN THE LOCAL SALES REPRESENTATIVE WAS AT THE HOSPITAL 11 DAYS LATER, THEY WERE INFORMED THE PATIENT HAD PASSED AWAY THE DAY OF THE IMPLANT PROCEDURE. THE ATTEMPTED LV LEAD WAS NOT INTENDED TO BE RETURNED. Manufacturer Narrative: IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22722961&pc=NAY | 2955842-2025-33066 |  | Death | INTUITIVE SURGICAL, INC | 2025-08-05 04:00 | NAY | DAVINCI XI | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Insufficient Information | K131861 |  | 1 | Event Description: IT WAS REPORTED THAT AFTER A DA VINCI-ASSISTED COLORECTAL PROCEDURE, THE PATIENT EXPIRED; HOWEVER, THE CAUSE OF DEATH WAS NOT DISCLOSED. THE INTUITIVE CLINICAL SALES REPRESENTATIVE WAS INFORMED OF ¿OVERHEARD INFORMATION¿ IN THE HOSPITAL ABOUT THE PATIENT DEATH. ALTHOUGH ADDITIONAL INFORMATION WAS REQUESTED, THE SITE HAS CONFIRMED THAT NO FURTHER DETAILS WILL BE PROVIDED. Manufacturer Narrative: THERE WAS NO STATEMENT OF ANY MALFUNCTION OF AN INTUITIVE SURGICAL INC. (ISI) SYSTEM, INSTRUMENT OR ACCESSORY OCCURRING DURING THE PROCEDURE. NO PRODUCT IS EXPECTED TO BE RETURNED. THE PROCEDURE DATE, TIME AND SURGEON WERE NOT PROVIDED. THEREFORE, NO DA VINCI SYSTEM LOGS ARE AVAILABLE FOR REVIEW. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT DIED FOLLOWING AN UNSPECIFIED COLORECTAL PROCEDURE. NO ADDITIONAL DETAILS ARE PROVIDED. BASED ON THE INFORMATION IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCT OR INSTRUMENT CONTRIBUTED TO THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22722962&pc=FKX | 0008030665-2025-01786 | 2025-07-14 04:00 | Death | ERIKA DE REYNOSA, S.A. DE C.V. | 2025-08-05 04:00 | FKX | LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | K222318 |  | 1 | Event Description: IT WAS REPORTED TO FRESENIUS THAT A PERITONEAL DIALYSIS (PD) PATIENT EXPIRED WHILE UNDERGOING CONTINUOUS CYCLIC PD (CCPD) THERAPY ON THE LIBERTY SELECT CYCLER. FOLLOW-UP WITH THE PATIENTS¿ PD REGISTERED NURSE (PDRN) CONFIRMED THE PATIENT EXPIRED AT HOME DURING CCPD THERAPY IN THE EARLY MORNING HOURS OF (B)(6) 2025. THE PATIENT WAS EXPECTED AT THE CLINIC ON (B)(6) 2025 FOR THEIR REGULARLY SCHEDULED VISIT BUT NEVER SHOWED. ON (B)(6) 2025, THE PDRN PLACED A CALL TO THE POLICE DEPARTMENT FOR A WELLNESS CHECK. AFTER THE PATIENT DID NOT RESPOND TO THE DOORBELL, THE FIRE DEPARTMENT FORCED THE DOOR OPEN AND DISCOVERED THE PATIENT HAD EXPIRED WHILE SLEEPING. THE PDRN STATED THE PATIENT PASSED AWAY PEACEFULLY WITH NO APPARENT SIGNS OF DISTRESS. THE PATIENT¿S TREATMENT DATA FROM (B)(6) 2025 THROUGH (B)(6) 2025 INDICATED THE PATIENT COMPLETED TREATMENT, AND NO ALARMS WERE NOTED. THE PATIENT WAS PICKED UP BY THE CORONER AND BROUGHT TO THE MORGUE UNTIL FAMILY COULD BE CONTACTED. THE PDRN STATED THE PRIMARY CAUSE OF DEATH WAS CARDIAC ARREST DUE TO THEIR SIGNIFICANT CARDIAC HISTORY (CORONARY ARTERY DISEASE, HYPERTENSION, DIABETES MELLITUS TYPE 2, PREVIOUS NSTEMI). THE PDRN STATED THE SERIOUS ADVERSE EVENTS WERE UNRELATED TO ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S) DEFICIENCY OR MALFUNCTION. Manufacturer Narrative: THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL INVESTIGATION: A TEMPORAL RELATIONSHIP EXISTS BETWEEN CCPD THERAPY UTILIZING A LIBERTY SELECT CYCLER AND THE SERIOUS ADVERSE EVENTS OF CARDIAC ARREST AND DEATH, AS THE PATIENT WAS ACTIVELY UNDERGOING CCPD THERAPY WHEN THEY EXPIRED. THE ETIOLOGY OF THE SERIOUS ADVERSE EVENTS IS UNKNOWN; THEREFORE, CAUSALITY CANNOT FIRMLY BE ESTABLISHED. HOWEVER, THE PDRN REPORTED THAT THE PATIENT¿S EXTENSIVE CARDIAC HISTORY WAS THE PRIMARY CAUSE OF THE SERIOUS ADVERSE EVENTS AND WAS UNRELATED TO ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S) MALFUNCTION OR DEFICIENCY. THE END-STAGE RENAL DISEASE (ESRD) POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. CARDIOVASCULAR DISEASE (INCLUDING SUDDEN CARDIAC DEATH) ACCOUNTS FOR THE MOST DEATHS AMONG THE ESRD POPULATION. ADDITIONALLY, ADULTS WITH ESRD HAVE MORTALITY RATES UP TO 30-FOLD HIGHER THAN THE GENERAL POPULATION. BASED ON THE INFORMATION AVAILABLE, THE LIBERTY SELECT CYCLER CAN BE DISASSOCIATED FROM THE SERIOUS ADVERSE EVENTS. THERE IS NO ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) CAUSED OR CONTRIBUTED TO THE SERIOUS ADVERSE EVENTS. FURTHERMORE, THERE WAS NO REPORT A FRESENIUS PRODUCT(S) AND/OR DEVICE(S) FAILED TO MEET THE USERS¿ EXPECTATIONS AND/OR THE MANUFACTURERS¿ SPECIFICATIONS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22723223&pc=NAY | 2955842-2025-33072 | 2025-07-12 04:00 | Death | INTUITIVE SURGICAL, INC | 2025-08-05 04:00 | NAY | DAVINCI X | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Insufficient Information | K171294 |  | 1 | Event Description: IT WAS REPORTED THAT AT AN UNSPECIFIED TIME AFTER A DA VINCI-ASSISTED PARTIAL GASTRECTOMY WITH LYMPHADENECTOMY, THE PATIENT EXPERIENCED A ¿RELAPSE¿, DURING WHICH BLEEDING WAS OBSERVED IN THE SURGICAL DRAIN. THE PATIENT SUBSEQUENTLY EXPIRED. THE INITIAL PROCEDURE WAS REPORTED TO BE SUCCESSFULLY COMPLETED WITHOUT ANY INTRAOPERATIVE COMPLICATIONS, AND THE PLACEMENT OF THE DRAIN WAS FOR ROUTINE CLINICAL MANAGEMENT. TO ADDRESS THE BLEEDING, AN ADDITIONAL PROCEDURE WAS INITIATED LAPAROSCOPICALLY BUT WAS SUBSEQUENTLY CONVERTED TO AN OPEN LAPAROTOMY. ON POST-OPERATIVE DAY THREE, THE PATIENT EXPIRED. ALTHOUGH THE CAUSE OF DEATH REMAINS UNKNOWN, THE SITE REPORTED THAT A DA VINCI SYSTEM, INSTRUMENT, AND/OR ACCESSORY DID NOT CAUSE OR CONTRIBUTE TO THE REPORTED EVENT. ADDITIONAL INFORMATION WAS REQUESTED; HOWEVER, THE SITE HAS INDICATED THAT NO FURTHER DETAILS WILL BE PROVIDED. Manufacturer Narrative: A REVIEW OF THE SYSTEM LOGS FOR THE REPORTED PROCEDURE DATE FOUND THAT NO SYSTEM RELATED ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM SHOWED THE SYSTEM FUNCTIONED NORMALLY, NO INTRAOPERATIVE EVENTS OR ISSUES OCCURRED. A DEVICE HISTORY RECORD (DHR) REVIEW FOUND NO NON-CONFORMANCES DOCUMENTED THAT WOULD BE RELATED TO THIS EVENT. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: 30 DEGREE ENDOSCOPE PLUS, PERMANENT CAUTERY HOOK, LARGE NEEDLE DRIVER, CADIERE FORCEPS, FENESTRATED BIPOLAR FORCEPS. A SITE HISTORY REVIEW FOUND NO COMPLAINTS WERE MADE FOR THE INSTRUMENTS USED DURING THIS PROCEDURE. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT HAD SOME POSTOPERATIVE BLEEDING FOLLOWING A PARTIAL GASTRECTOMY. ALTHOUGH THE PATIENT REQUIRED A RE-EXPLORATION, THE LOCATION, CAUSE, AND AMOUNT OF BLEEDING IS NOT PROVIDED. THE PATIENT SUBSEQUENTLY DIED. THE CAUSE OF DEATH AND EVENTS THAT LED TO THE DEATH ARE NOT PROVIDED. IT IS UNKNOWN IF THE AFOREMENTIONED BLEEDING EVENT CONTRIBUTED TO THE DEATH. NO ALLEGATION WAS MADE AGAINST ANY INTUITIVE SURGICAL PRODUCTS. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22723355&pc=NIP | 2017233-2025-06483 | 2025-07-17 04:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-08-05 04:00 | NIP | GORE® VIABAHN® ENDOPROSTHESIS WITH HEPARIN BIOACTIVE SURFACE | Adverse Event Without Identified Device or Use Problem | Stenosis | P040037 |  | 1 | Event Description: THE FOLLOWING INFORMATION WAS REPORTED TO GORE: ON (B)(6) 2025, THE PATIENT UNDERWENT ENDOVASCULAR TREATMENT FOR A THORACOABDOMINAL AORTIC ANEURYSM USING GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM(CTAGAC) AND GORE® VIABAHN® ENDOPROSTHESIS WITH HEPARIN BIOACTIVE SURFACE(VSX). THE CELIAC ARTERY WAS INTENTIONALLY COVERED BY THE FIRST CTAGAC. ONE VSX WAS PLACED IN THE SUPERIOR MESENTERIC ARTERY AND TWO VSX DEVICES WERE IN THE RIGHT RENAL ARTERY AS ALL CHIMNEY STENT GRAFTS. THE SECOND CTAGAC WAS DEPLOYED DISTALLY, SANDWICHING ALL VSX CHIMNEY STENT GRAFTS BETWEEN THE FIRST AND SECOND CTAGACS. THE PROCEDURE WAS COMPLETED SUCCESSFULLY. APPROXIMATELY SIX HOURS AFTER SURGERY ON THE SAME DAY, THE PATIENT COMPLAINED OF ABDOMINAL PAIN. A CT SCAN REVEALED REDUCED BLOOD FLOW TO THE SUPERIOR MESENTERIC ARTERY, SUGGESTING INTESTINAL ISCHEMIA. THERE WAS NO KINK OBSERVED IN THE VSX PLACED IN THE SUPERIOR MESENTERIC ARTERY. THE DETAILS OF THE TREATMENT ARE UNKNOWN, REPORTEDLY A HEPATIC ARTERY BYPASS WAS PERFORMED. THE PATIENT EXPIRED DUE TO INTESTINAL ISCHEMIA. THE DATE OF DEATH AND THE DETAILS OF THE AUTOPSY ARE UNKNOWN. NO COMMENTS HAVE BEEN RECEIVED FROM THE PHYSICIAN REGARDING THE CAUSE OF INTESTINAL ISCHEMIA. Manufacturer Narrative: B2: THE DATE OF PATIENT DEATH IS UNKNOWN. H6: A REVIEW OF THE MANUFACTURING RECORDS FOR THE DEVICE VERIFIED THAT THE LOT MET ALL PRE-RELEASE SPECIFICATIONS. HEPARIN SURFACE INCORPORATES CARMEDA HEPARIN MANUFACTURED FROM HEPARIN SODIUM API, WHICH IS COVALENTLY BOUND TO THE DEVICE SURFACE AND IS ESSENTIALLY NON-ELUTING. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22723479&pc=MOZ | 3007093114-2025-00009 | 2025-07-01 04:00 | Death | ANIKA | 2025-08-05 04:00 | MOZ | ORTHOVISC | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: THIS REPORT IS BEING SUBMITTED IN PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ANIKA OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE DEVICE, ANIKA, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO A POTENTIAL PATIENT EVENT DOCUMENTED ON THIS REPORT. A MEDWATCH REPORT (MW5172542) WAS RECEIVED FROM THE FDA. A PATIENT OF UNKNOWN AGE AND DEMOGRAPHICS REPORTEDLY EXPIRED. THE CAUSE OF THE PATIENT DEATH HAS NOT BEEN REPORTED. ORTHOVISC (LOT NUMBER WAS NOT PROVIDED) WAS LISTED AS THE DEVICE. THE DATE OF THE ORTHOVISC TREATMENT AND THE PATIENT DEATH WAS NOT REPORTED ON THE MEDWATCH. PATIENT COMORBIDITIES ARE UNKNOWN. THERE WAS NO REPORT OF A DEVICE MALFUNCTION. CONTACT INFORMATION WAS NOT PROVIDED SO A REQUEST FOR ADDITIONAL INFORMATION COULD NOT BE PERFORMED. THE DISTRIBUTOR COPIED THE REPORT AND DID NOT INDICATE THIS CASE WAS REPORTED TO THEM. Manufacturer Narrative: THIS CASE IS UNDER INVESTIGATION. A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22696794&pc=MLV | 2017233-2025-06467 | 2025-07-10 04:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-08-04 04:00 | MLV | GORE® CARDIOFORM ASD OCCLUDER | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P050006 |  | 1 | Event Description: IT WAS REPORTED TO GORE THAT ON (B)(6) 2025 A 44 MM GORE® CARDIOFORM ASD OCCLUDER WAS SELECTED TO TREAT A VENTRICULAR SEPTAL DEFECT POST MYOCARDIAL INFARCTION DURING A PLANNED OFF LABEL PROCEDURE. THE MYCARDIAL INFARCTION WAS REPORTED TO HAVE OCCURRED APROXIMATELY 2.5 WEEKS BEFORE. REPORTEDLY, A GORE® DRY SEAL FLEX INTRODUCER SHEATH WAS ALSO USED IN THE PROCEDURE. THE GORE® CARDIOFORM ASD OCCLUDER WAS SUCESSFULLY DEPLOYED IN THE PATIENT AND SUBSEQUENTLY RELEASED WITH TRANSTHORACIC ECHOCARDIOGRAM ULTRASOUND BEING ON STANDBY. REPORTEDLY, THE PATIENT EXPIRED DURING THE PROCEDURE. THE GORE® CARDIOFORM ASD OCCLUDER REPORTEDLY MIGHT HAVE CAUSED OR CONTRIBUTED TO THE PATIENT'S DEATH. Manufacturer Narrative: A REVIEW OF THE MANUFACTURING RECORDS INDICATED THE LOT MET ALL PRE-RELEASE SPECIFICATIONS. FURTHER INVESTIGATION IS BEING CONDUCTED AND WILL BE INCLUDED IN THE FINAL REPORT. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22696904&pc=MUZ | 1644487-2025-10460 | 2025-07-05 04:00 | Death | LIVANOVA USA, INC. | 2025-08-04 04:00 | MUZ | SYMMETRY GEN MODEL 8103 | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P970003 |  | 1 | Event Description: IT WAS REPORTED PATIENT COMMITTED SUICIDE. THIS EVENT IS REPORTEDLY UNLIKELY RELATED TO STIMULATION, DEVICE AND THE IMPLANT PROCEDURE. NO ADDITIONAL RELEVANT INFORMATION HAS BEEN RECEIVED TO DATE. Manufacturer Narrative: LIVANOVA USA, INC. SUBMITS THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION, BASED ON INFORMATION THAT LIVANOVA HAS OBTAINED, BUT MAY NOT HAVE BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION, OR A CONCLUSION BY FDA OR ANYONE ELSE, THAT THE DEVICE, LIVANOVA, OR LIVANOVA'S EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS¿ OR "MALFUNCTIONS¿. THESE WORDS ARE INCORPORATED INTO THE FDA 3500A MEDWATCH FORM BY THE FDA, AND LIVANOVA OBJECTS TO THEIR USE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22698099&pc=PRL | 2017233-2025-06469 | 2021-01-20 05:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-08-04 04:00 | PRL | GORE® VIABAHN® VBX BALLOON EXPANDABLE ENDOPROSTHESIS | Adverse Event Without Identified Device or Use Problem | Ischemia | P160021 |  | 1 | Event Description: THE FOLLOWING WAS REPORTED TO GORE ON JULY 15, 2025, FROM THE MEDRIO STUDY DATABASE: ON (B)(6) 2021, THIS 68-YEAR-OLD MALE PATIENT UNDERWENT ENDOVASCULAR TREATMENT FOR A JUXTARENAL ANEURYSM. THE PATIENT WAS TREATED WITH A COOK CUSTOM DEVICE, TWO GORE® VIABAHN® VBX BALLOON EXPANDABLE ENDOPROSTHESIS DEVICES (VBX DEVICES) TO THE CELIAC AND SUPERIOR MESENTERIC ARTERIES, AND AN ADVANTA DEVICE TO THE LEFT RENAL ARTERY. THE TWO VBX DEVICES WERE SUCCESSFULLY NAVIGATED TO THEIR INTENDED LOCATION AND SUCCESSFULLY DEPLOYED. DEVICE CATHETERS WERE SUCCESSFULLY REMOVED. ON (B)(6) 2021 THE PATIENT EXPERIENCED A BOWEL ISCHEMIA. FOR THE BOWEL ISCHEMIA AN END ORGAN SURGICAL REINTERVENTION WAS PERFORMED. PART OF THE DESCENDING COLON/RECTOSIGMOID AND 1 METER OF THE JEJUNUM WERE RESECTED. FINALLY, THE PATIENT EXPIRED. Manufacturer Narrative: A1: NO PATIENT SPECIFIC DETAILS HAVE BEEN PROVIDED. THEREFORE, THE PATIENT INITIALS REFLECT THE STUDY NUMBER WITH HOSPITAL AND PATIENT. AS THE DEVICE REMAINS IMPLANTED, NO FURTHER INVESTIGATION OF THE DEVICE CAN BE CONDUCTED. H6 INVESTIGATION FINDINGS C19 REFERS TO THE PRODUCT HISTORY REVIEW: A REVIEW OF THE MANUFACTURING RECORDS INDICATED THE DEVICE MET PRE-RELEASE SPECIFICATIONS. FURTHER INVESTIGATION IS BEING CONDUCTED AND WILL BE INCLUDED IN THE FINAL REPORT. CBAS® HEPARIN SURFACE INCORPORATES CARMEDA-HEPARIN MANUFACTURED FROM HEPARIN SODIUM API, WHICH IS COVALENTLY BOUND TO THE DEVICE SURFACE AND IS ESSENTIALLY NON-ELUTING. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22698133&pc=PRL | 2017233-2025-06470 | 2021-01-20 05:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-08-04 04:00 | PRL | GORE® VIABAHN® VBX BALLOON EXPANDABLE ENDOPROSTHESIS | Adverse Event Without Identified Device or Use Problem | Ischemia | P160021 |  | 1 | Event Description: THE FOLLOWING WAS REPORTED TO GORE ON JULY 15, 2025, FROM THE MEDRIO STUDY DATABASE: ON (B)(6) 2021, THIS 68-YEAR-OLD MALE PATIENT UNDERWENT ENDOVASCULAR TREATMENT FOR A JUXTARENAL ANEURYSM. THE PATIENT WAS TREATED WITH A COOK CUSTOM DEVICE, TWO GORE® VIABAHN® VBX BALLOON EXPANDABLE ENDOPROSTHESIS DEVICES (VBX DEVICES) TO THE CELIAC AND SUPERIOR MESENTERIC ARTERIES, AND AN ADVANTA DEVICE TO THE LEFT RENAL ARTERY. THE TWO VBX DEVICES WERE SUCCESSFULLY NAVIGATED TO THEIR INTENDED LOCATION AND SUCCESSFULLY DEPLOYED. DEVICE CATHETERS WERE SUCCESSFULLY REMOVED. ON (B)(6) 2021 THE PATIENT EXPERIENCED A BOWEL ISCHEMIA. FOR THE BOWEL ISCHEMIA AN END ORGAN SURGICAL REINTERVENTION WAS PERFORMED. PART OF THE DESCENDING COLON/RECTOSIGMOID AND 1 METER OF THE JEJUNUM WERE RESECTED. FINALLY, THE PATIENT EXPIRED. Manufacturer Narrative: A1: NO PATIENT SPECIFIC DETAILS HAVE BEEN PROVIDED. THEREFORE, THE PATIENT INITIALS REFLECT THE STUDY NUMBER WITH HOSPITAL AND PATIENT. AS THE DEVICE REMAINS IMPLANTED, NO FURTHER INVESTIGATION OF THE DEVICE CAN BE CONDUCTED. H6 INVESTIGATION FINDINGS C19 REFERS TO THE PRODUCT HISTORY REVIEW: A REVIEW OF THE MANUFACTURING RECORDS INDICATED THE DEVICE MET PRE-RELEASE SPECIFICATIONS. FURTHER INVESTIGATION IS BEING CONDUCTED AND WILL BE INCLUDED IN THE FINAL REPORT. CBAS® HEPARIN SURFACE INCORPORATES CARMEDA-HEPARIN MANUFACTURED FROM HEPARIN SODIUM API, WHICH IS COVALENTLY BOUND TO THE DEVICE SURFACE AND IS ESSENTIALLY NON-ELUTING. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22698271&pc=CAW | 2518422-2025-108633 | 2025-07-28 04:00 | Death | RESPIRONICS, INC. | 2025-08-04 04:00 | CAW | SIMPLYGO | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K111885 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A SIMPLYGO PORTABLE OXYGEN CONCENTRATOR DEVICE HAS PASSED AWAY. THE PATIENT¿S WIFE REPORTED THE DEATH AND STATED THAT HER DURABLE MEDICAL EQUIPMENT (DME) PROVIDER HAD INDICATED SHE MIGHT RECEIVE PAYMENT UPON RETURNING THE DEVICE. SHE WAS ADVISED THAT THE MANUFACTURER DOES NOT OFFER COMPENSATION FOR RETURNED DEVICES. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT'S DEATH. NO FURTHER DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22699014&pc=MSD | 9680794-2025-00648 | 2025-07-01 04:00 | Death | ARROW INTERNATIONAL LLC | 2025-08-04 04:00 | MSD | ARROW HEMODIALYSIS KIT: 2-L 15 FR X 19 CM | Insufficient Information | Air Embolism; Stroke/CVA; Low Blood Pressure/ Hypotension | K111117 |  | 1 | Event Description: IT WAS REPORTED THAT "INCIDENT INVOLVED A 85-YEAR-OLD PATIENT WHO HAS BEEN ON DIALYSIS SINCE (B)(6) 2024 FOR STAGE V CHRONIC KIDNEY DISEASE. ON (B)(6) 2025, 15 MINUTES AFTER THE START OF THE DIALYSIS SESSION, THE PATIENT SUDDENLY LOST CONSCIOUSNESS WITH A GLASGOW COMA SCALE SCORE OF 8 AND A DROP IN BLOOD PRESSURE TO 79/55 MMHG. THERE WAS NO "AIR BUBBLE" ALARM ON THE GENERATOR, WHICH WOULD NOT HAVE DETECTED THE PRESENCE OF AIR IN THE LINE, DUE TO A GENERATOR MALFUNCTION OR AIR ENTERING BEFORE THE DETECTOR. PATIENT WAS TRANSFERRED TO INTENSIVE CARE FOR SUSPECTED GAS EMBOLISM, HE WAS INTUBATED, VENTILATED. A HYPERBARIC CHAMBER SESSION WAS PERFORMED. THE ELECTROCARDIOGRAM PERFORMED ON (B)(6) REVEALED GENERALIZED CEREBRAL DAMAGE IN THE CONTEXT OF A STROKE DUE TO PROBABLE GAS EMBOLISM. THE PATIENT DIED ON (B)(6) 2025." ASSOCIATE MDR NUMBERS INCLUDE: 3006425876-2025-00715. Manufacturer Narrative: QN #(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22699912&pc=DQX | 3009957947-2025-00003 | 2025-07-12 04:00 | Death | RAPID MEDICAL LTD | 2025-08-04 04:00 | DQX | DRIVEWIRE 24 GUIDEWIRE | Adverse Event Without Identified Device or Use Problem | Ischemia Stroke | K233791 |  | 1 | Event Description: EVENT DESCRIPTION: RAPID MEDICAL RECEIVED A COMPLAINT REGARDING A DRIVEWIRE 24 (DW24) DEVICE USED IN A THROMBECTOMY PROCEDURE. ACCORDING TO THE INFORMATION PROVIDED, THE DW24 WAS INSERTED THROUGH A 3F ESPERANCE CATHETER (PHENOX). DURING ASPIRATION, THE ESPERANCE CATHETER COLLAPSED, RESULTING IN THE TEARING OF ITS DISTAL TIP. THE DW24 REMAINED STRUCTURALLY INTACT, ALTHOUGH MINOR DAMAGE TO ITS DISTAL TIP WAS REPORTED. THE PATIENT, WHO HAD BEEN ADMITTED WITH A SEVERE ISCHEMIC STROKE, LATER PASSED AWAY. BASED ON THE INFORMATION PROVIDED, THE PATIENT'S DEATH APPEARS TO BE RELATED TO THE UNDERLYING MEDICAL CONDITION AND IS NOT ATTRIBUTED TO THE DEVICE OR THE REPORTED EVENT. Manufacturer Narrative: MANUFACTURER NARRATIVE: THE DW24 DEVICE WAS NOT RETURNED FOR INVESTIGATION, SO A DETAILED EXAMINATION COULD NOT BE PERFORMED. WHILE THE LOT NUMBER FOR THIS SPECIFIC DEVICE WAS NOT PROVIDED, ALL DW24 DEVICES UNDERGO BATCH RELEASE TESTING TO ENSURE COMPLIANCE WITH APPLICABLE SPECIFICATIONS. THE DEVICES ARE MANUFACTURED AND RELEASED ACCORDING TO ESTABLISHED PROCEDURES, AND ALL LOTS MEET THE ACCEPTANCE CRITERIA PRIOR TO RELEASE. BASED ON THE COMPLAINT DESCRIPTION AND AVAILABLE INFORMATION, THERE IS NO INDICATION THAT THE DW24 DEVICE MALFUNCTIONED OR CONTRIBUTED TO THE ADVERSE OUTCOME. NO FURTHER ACTION IS REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22700073&pc=DSR | 3007972010-2025-00073 | 2025-07-06 04:00 | Death | CVRX, INC | 2025-08-04 04:00 | DSR | BAROSTIM NEO2 | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Heart Failure/Congestive Heart Failure | P180050 |  | 1 | Event Description: A BAROSTIM SYSTEM WAS IMPLANTED ON (B)(6) 2025. ON (B)(6) 2025, THE PATIENT PASSED AWAY. PER THE OPINION OF THE PHYSICIAN, THE ROOT CAUSE OF THE EVENT WAS CARDIAC ARREST AND CONGESTIVE HEART FAILURE AND THE PATIENT'S DEATH WAS NOT CAUSE OR CONTRIBUTED TO/ BY BAROSTIM. Manufacturer Narrative: WHILE ANALYSIS WAS UNABLE TO BE PERFORMED AS THE DEVICE WAS NOT RETURNED, PER THE OPINION OF THE PHYSICIAN, THE ROOT CAUSE OF THE EVENT WAS CARDIAC ARREST AND CONGESTIVE HEART FAILURE. THE DEVICE HISTORY AND STERILIZATION RECORD FOR THIS DEVICE SERIAL NUMBER HAS BEEN REVIEWED. NO ISSUES OR DISCREPANCIES WERE NOTED DURING THIS REVIEW THAT WOULD HAVE CONTRIBUTED TO THE REPORTED EVENT. THE DEVICE MET MATERIAL, ASSEMBLY, AND QUALITY CONTROL REQUIREMENTS. CVRX ID# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22700523&pc=NIQ | 2024168-2025-07963 | 2016-01-01 05:00 | Death | ABBOTT VASCULAR INC. | 2025-08-04 04:00 | NIQ | XIENCE SKYPOINT¿ | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P110019 |  | 1 | Event Description: THE ARTICLE AIMED TO TEST THE HYPOTHESIS OF NONINFERIORITY OF 1 TO 2 MONTHS OF DUAL ANTIPLATELET THERAPY (DAPT) COMPARED WITH 12 MONTHS OF DAPT FOR A COMPOSITE END POINT OF CARDIOVASCULAR AND BLEEDING EVENTS IN PATIENTS WITH ACUTE CORONARY SYNDROME (ACS). THE STOPDAPT-2 ACS CLINICAL TRIAL ENROLLED 4136 PATIENTS WITH ACS WHO UNDERWENT SUCCESSFUL PCI USING COBALT-CHROMIUM EVEROLIMUS-ELUTING STENTS (XIENCE) AT 96 CENTERS IN JAPAN FROM DECEMBER 2015 THROUGH JUNE 2020. PATIENTS WERE RANDOMIZED EITHER TO 1 TO 2 MONTHS OF DAPT FOLLOWED BY CLOPIDOGREL MONOTHERAPY OR TO 12 MONTHS OF DAPT WITH ASPIRIN AND CLOPIDOGREL. THE PRIMARY END POINT WAS A COMPOSITE OF CARDIOVASCULAR OUTCOMES INCLUDING CARDIOVASCULAR DEATH, MYOCARDIAL INFARCTION, DEFINITE STENT THROMBOSIS, STROKE, BLEEDING, TARGET LESION REVASCULARIZATION, AND HOSPITALIZATION. THE ARTICLE CONCLUDED THAT IN PATIENTS WITH ACS WHO UNDERWENT SUCCESSFUL PCI USING COCR-EES, CLOPIDOGREL MONOTHERAPY AFTER 1 TO 2 MONTHS OF DAPT FAILED TO ATTEST NONINFERIORITY TO 12 MONTHS OF DAPT WITH ASPIRIN AND CLOPIDOGREL FOR THE NET CLINICAL BENEFIT WITH A NUMERICAL INCREASE IN CARDIOVASCULAR EVENTS DESPITE REDUCTION IN MAJOR BLEEDING EVENTS. DETAILS ARE LISTED IN THE ATTACHED ARTICLE, TITLED "COMPARISON OF CLOPIDOGREL MONOTHERAPY AFTER 1 TO 2 MONTHS OF DUAL ANTIPLATELET THERAPY WITH 12 MONTHS OF DUAL ANTIPLATELET THERAPY IN PATIENTS WITH ACUTE CORONARY SYNDROME.¿ NO ADDITIONAL INFORMATION WAS PROVIDED. FOR REPORTING PURPOSES, THE DATE OF PROCEDURE/IMPLANT WILL BE ESTIMATED AS (B)(6) 2015 THIS IS THE FIRST OF THE MONTH THAT THE PATIENTS CONSENTED TO BE IN THE TRIAL. DATE OF EVENT/DEATH WILL BE ESTIMATED AS (B)(6) 2016 AS FOLLOW UP WAS PERFORMED AT 1 MONTH. Manufacturer Narrative: ATTACHMENT ARTICLE TITLED: "COMPARISON OF CLOPIDOGREL MONOTHERAPY AFTER 1 TO 2 MONTHS OF DUAL ANTIPLATELET THERAPY WITH 12 MONTHS OF DUAL ANTIPLATELET THERAPY IN PATIENTS WITH ACUTE CORONARY SYNDROME.¿ THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORDS FOR THIS PRODUCT COULD NOT BE REVIEWED AND A SIMILAR COMPLAINT QUERY COULD NOT BE PERFORMED BECAUSE THE PART NUMBER AND/OR LOT NUMBER WERE NOT REPORTED, AND THE DEVICE WAS NOT RETURNED. THE REPORTED PATIENT EFFECT OF DEATH IS LISTED IN THE XIENCE SKYPOINT, EVEROLIMUS ELUTING CORONARY STENT SYSTEM (EECSS), ELECTRONIC INSTRUCTIONS FOR USE AS A KNOWN PATIENT EFFECT OF CORONARY STENTING PROCEDURES. A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECT OF DEATH AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED. THE INFORMATION RECEIVED IN THE COMPLAINT WAS OBTAINED THROUGH A PROACTIVE LITERATURE SEARCH. SPECIFIC MODEL AND LOT NUMBERS WERE NOT AVAILABLE AND CAUSALITY BETWEEN THE DOCUMENTED DEVICE USAGE AND THE PATIENT EFFECTS COULD NOT BE ESTABLISHED. FACTORS THAT MAY CONTRIBUTE TO THE PERFORMANCE OUTCOMES LISTED IN THE ARTICLE INCLUDE, BUT ARE NOT LIMITED TO, DEVICE TO PATIENT INTERACTION, PATIENT ANATOMICAL CONDITIONS, DEVICE TO USER INTERACTION AND MANUFACTURING OR MATERIAL ISSUES. DUE TO THE LIMITED INFORMATION AVAILABLE, THE EXACT CAUSE CANNOT BE DETERMINED. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN, OR LABELING. B2, B3, D6A: ESTIMATED DATE D4- THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. THE ADDITIONAL PATIENT EFFECTS REPORTED IN THE ARTICLE ARE CAPTURED UNDER A SEPARATE MEDWATCH REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22700606&pc=OZD | 1220648-2025-30962 | 2025-07-22 04:00 | Death | ABIOMED, INC. | 2025-08-04 04:00 | OZD | IMPELLA CP | Pumping Stopped | Arrhythmia | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT IMPLANTED WITH AN IMPELLA CP EXPERIENCED VENTRICULAR TACHYCARDIA THAT WAS MEDICALLY MANAGED. AFTER TRANSFERRING THE PATIENT FROM ONE AUTOMATE IMPELLA CONTROLLER TO ANOTHER, THE PUMP FAILED TO RESTART. THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMARTASSIST SYSTEM LISTS VENTRICULAR ARRYTHMIA AS AN ADVERSE EVENT THAT OCCURRED WHILE IN USE WITH THE IMPELLA DEVICE. IT STATES THAT "ONLY A FRACTION OF THESE RATES WERE ATTRIBUTED TO THE IMPELLA DEVICE AND THE EVENTS RESOLVED WITHOUT RESIDUAL EFFECT IN MOST OF THE CASES, UNLESS THE EVENT OF DEATH OCCURRED. OVERALL, THE RESULTS DID NOT SHOW ANY EVIDENCE OF INCREASED MORBIDITY ASSOCIATED WITH THE IMPELLA SUPPORT IN CARDIOMYOPATHY, MYOCARDITIS, AND PERIPARTUM CARDIOMYOPATHY (PPCM) PATIENTS." SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA CATHETER ¿ACHIEVING AN ACT = 250 SECONDS PRIOR TO REMOVING THE DILATOR WILL HELP PREVENT A THROMBUS FROM ENTERING THE CATHETER AND CAUSING A SUDDEN STOP ON STARTUP.¿ SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA RP WITH SMARTASSIST SYSTEM CATHETER ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA RP WITH SMARTASSIST SYSTEM CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP. MONITOR MOTOR CURRENT AND CONSIDER REPLACING THE IMPELLA RP WITH SMARTASSIST SYSTEM CATHETER WHENEVER A RISE IN MOTOR CURRENT IS SEEN.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22700915&pc=OZD | 1220648-2025-30963 | 2025-07-22 04:00 | Death | ABIOMED, INC. | 2025-08-04 04:00 | OZD | AUTOMATED IMPELLA CONTROLLER | Pumping Stopped | No Clinical Signs, Symptoms or Conditions | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT AFTER TRANSFERRING A PATIENT WHO HAD AN IMPELLA CP IMPLANTED FROM ONE SUPPORTING AUTOMATED IMPELLA CONTROLLER TO A SECOND, THE IMPELLA CP WOULD NOT RESTART. TWO ADDITIONAL AUTOMATED IMPELLA CONTROLLERS WERE TRIED WITHOUT SUCCESS. THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS RETURNED, AND AN EVALUATION IS UNDERWAY. UPON COMPLETION OF OUR ANALYSIS, A SUPPLEMENTAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22700948&pc=DSQ | 3007042319-2025-01874 | 2025-07-31 04:00 | Death | HEARTWARE, INC. | 2025-08-04 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP/DRIVELINE | Pumping Problem | Unspecified Infection; Renal Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE VENTRICULAR ASSIST DEVICE (VAD) IS EXHIBITING LOW PULSATILITY AND LOW FLOW. THE PATIENT HAS UNDERGONE AMPUTATION OF ONE UPPER LIMB, RESULTING IN REDUCED DAILY PHYSICAL ACTIVITY, WHICH HAD AN IMPACT ON FLOW. ADDITIONALLY, THE PATIENT CURRENTLY HAS AN ONGOING INFECTION AND RENAL FAILURE. THESE FACTORS HAVE CONTRIBUTED TO A DECREASE IN BOTH FLOW AND PULSATILITY. HOWEVER, DUE TO OTHER CLINICAL FINDINGS, THE PATIENT¿S CONDITION IS NOT STABLE. THE PATIENT WAS TREATED WITH ANTIBIOTICS AND WILL UNDERGO DIALYSIS. THE VAD IS FUNCTIONING WITHOUT ISSUE AND REMAINS IN USE.Event Description: IT WAS FURTHER REPORTED THAT THE VAD WAS STOPPED AFTER THE PATIENT STOPPED BREATHING, AND HAD NO MORE VISIBLE VITAL SIGNS. THERE ARE NO ALLEGATIONS AGAINST THE VAD. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR ADDITIONAL EVENT DETAILS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22701475&pc=OZD | 1220648-2025-30965 | 2025-07-21 04:00 | Death | ABIOMED, INC. | 2025-08-04 04:00 | OZD | IMPELLA 5.5 | Device Sensing Problem | No Clinical Signs, Symptoms or Conditions | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT IMPLANTED WITH AN IMPELLA 5.5 LOST PLACEMENT SIGNAL. AN IMPELLA RP FLEX WAS ALSO IMPLANTED BUT AFTER 47 HOURS PATIENT CARE WAS WITHDRAWN AND PATIENT EXPIRED. Manufacturer Narrative: A5, ETHNICITY, IS UNKNOWN. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22701648&pc=NPV | 2025587-2025-05458 | 2025-07-11 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-08-04 04:00 | NPV | HARMONY TRANSCATHETER PULMONARY VALVE | Insufficient Information; Patient Device Interaction Problem | Rupture | P200046 |  | 1 | Event Description: IT WAS REPORTED THAT AN UNSPECIFIED AMOUNT OF TIME FOLLOWING THE IMPLANT OF A TRANSCATHETER VALVE, THE PATIENT DIED DUE TO RUPTURE OF AN ALCOHOL-INDUCED ESOPHAGEAL VARICES. Manufacturer Narrative: B2. DATE OF DEATH IS NOT KNOWN. B3. EVENT DATE IS NOT KNOWN. D6A. IMPLANT DATE IS NOT KNOWN. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22701851&pc=NPU | 2015691-2025-06300 | 2024-09-05 04:00 | Death | EDWARDS LIFESCIENCES | 2025-08-04 04:00 | NPU | SAPIEN 3 TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Endocarditis | P140031 |  | 1 | Event Description: AS REPORTED IN THE JOURNAL ARTICLE "ALCOHOL SEPTAL ABLATION FOR LEFT VENTRICLE OUTFLOW TRACT OBSTRUCTION PREVENTION BEFORE TRANSCATHETER MITRAL VALVE REPLACEMENT PROCEDURE: COMPUTED TOMOGRAPHY ANALYSIS SERIES" FROM FRANCE, PATIENTS INDICATED FOR THE TRANSCATHETER MITRAL VALVE REPLACEMENT (TMVR) PROCEDURE USING A SAPIEN 3 PROSTHESIS WITH HIGH LEFT VENTRICULAR OUTFLOW TRACT (LVOT) OBSTRUCTION RISK WERE ENROLLED IN A STUDY. THESE PATIENTS UNDERWENT THE TMVR PROCEDURE BETWEEN MARCH 2021 AND APRIL 2023. THE PATIENTS ENROLLED IN THE STUDY HAD UNDERGONE ALCOHOL SEPTAL ABLATION (ASA) AS THE FIRST CHOICE TECHNIQUE TO REDUCE THE RISK FOR LVOT OBSTRUCTION. THIS RISK WAS BASED ON MULTIPLE AND STAGED CARDIAC COMPUTED TOMOGRAPHY (CT) SCREENING. THIS COMPLAINT INVOLVES A PATIENT WHO UNDERWENT SUCCESSFUL IMPLANTATION OF A SAPIEN 3 VALVE IN THE MITRAL POSITION VIA A FEMORAL VEIN APPROACH. APPROXIMATELY SIX (6) WEEKS POST TMVR PROCEDURE, THE PATIENT PASSED AWAY FOLLOWING A CARDIAC SURGERY FOR EARLY ENDOCARDITIS. THE DATE OF DEATH IS UNKNOWN. Manufacturer Narrative: THIS IS ONE OF TWO MANUFACTURER REPORTS BEING SUBMITTED FOR THIS CASE. PLEASE SEE MANUFACTURER REPORT NUMBER 2015691-2025-06294 FOR DETAILS OF THE OTHER EVENT. PER THE INSTRUCTIONS FOR USE (IFU), INFECTIONS SUCH AS SEPTICEMIA AND ENDOCARDITIS ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. ENDOCARDITIS IS AN INFECTION OF A NATIVE OR PROSTHETIC VALVE, IS TREATED WITH ANTIBIOTICS, AND MAY REQUIRE VALVE REPLACEMENT IF ANTIBIOTIC THERAPY IS NOT EFFECTIVE. CAUSES OF PROSTHETIC VALVE ENDOCARDITIS ARE WELL DOCUMENTED IN THE LITERATURE AND ARE TYPICALLY CLASSIFIED AS EARLY (LESS THAN 60 DAYS) OR LATE (GREATER THAN 60 DAYS). THE TWO CATEGORIES SHOW DEFINITE DIFFERENCES IN CLINICAL FEATURES, MICROBIAL PATTERNS, AND MORTALITY. ACCORDING TO THE LITERATURE REVIEW, AND AS DOCUMENTED IN A TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, EARLY PROSTHETIC VALVE ENDOCARDITIS IS USUALLY CAUSED BY PERIOPERATIVE BACTERIAL CONTAMINATION OF THE VALVE. IN EARLY CASES OF PROSTHETIC ENDOCARDITIS, SUBSEQUENT INFECTIONS MAY BE RELATED TO CONTAMINATION AT THE TIME OF SURGERY (POOR STERILE TECHNIQUE). ADDITIONAL CAUSES OF EARLY PROSTHETIC VALVE ENDOCARDITIS (PVE) CAN OCCUR FROM OTHER SOURCES OF INFECTION AND NOT FROM THE VALVE (E.G., DENTAL TREATMENT OR SURGICAL PROCEDURES INVOLVING THE UPPER RESPIRATORY TRACT, URINARY TRACT INFECTIONS (UTIS), PNEUMONIA). ADDITIONALLY, A NON-CONFORMANCE IN THE STERILITY OR PACKAGING PROCESSES OF THE VALVE MAY ALSO MANIFEST IN THE EARLY POST-OPERATIVE PERIOD CAUSING EARLY PVE. HOWEVER, THERE ARE SEVERAL TYPES OF BACTERIA THAT WOULD NOT SURVIVE THE EDWARDS STERILIZATION PROCESS. EDWARDS LIFESCIENCES PRODUCES AND PROVIDES STERILE TISSUE BIOPROSTHESIS TO ITS CUSTOMERS BY FOLLOWING CAREFULLY DESIGNED ROBUST STERILIZATION PROCESSES. THESE MANUFACTURING PROCESSES HAVE BEEN VALIDATED AND DEMONSTRATED TO CONSISTENTLY PROVIDE A SIGNIFICANT SAFETY FACTOR FROM WHICH MICROORGANISMS COULD NOT SURVIVE. MICROBIOLOGY AND PROCESS MONITORING IS ROUTINELY REVIEWED WITHIN QUALITY SYSTEMS TO MAINTAIN STERILITY CONTROL. VALIDATED TESTING HAS DEMONSTRATED THAT MICROORGANISMS COULD NOT SURVIVE EDWARDS' MULTI-STAGE PROCESSING WITH ENHANCED STERILANT OR HEATED GLUTARALDEHYDE TERMINAL STERILANT SOLUTION. THESE MULTIPLE, REDUNDANT MANUFACTURING CONTROLS ENSURE THE STERILITY OF EDWARDS' VALVES AS PROVIDED TO CUSTOMERS. THEREFORE, THE PROBABILITY OF ENDOCARDITIS RELATED TO EDWARDS' BIOPROSTHESIS IS REMOTE. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. ALTHOUGH A DEFINITIVE ROOT CAUSE WAS UNABLE TO BE DETERMINED, THE EVENT COULD BE RELATED TO THE MECHANISMS DESCRIBED ABOVE. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22701996&pc=FKX | 0008030665-2025-01758 | 2025-07-24 04:00 | Death | ERIKA DE REYNOSA, S.A. DE C.V. | 2025-08-04 04:00 | FKX | LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K181108 |  | 1 | Event Description: ON 22/JUL/2025, FRESENIUS BECAME AWARE THIS PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON CONTINUOUS CYCLIC PERITONEAL DIALYSIS [CC(PD)] UTILIZING A LIBERTY SELECT CYCLER FOR RENAL REPLACEMENT THERAPY (RRT) EXPIRED WHILE UNDERGOING CCPD THERAPY ON (B)(6) 2025.Event Description: ON 25/JUL/2025, FRESENIUS BECAME AWARE THIS PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON CONTINUOUS CYCLIC PERITONEAL DIALYSIS [CC(PD)] UTILIZING A LIBERTY SELECT CYCLER FOR RENAL REPLACEMENT THERAPY (RRT) EXPIRED WHILE UNDERGOING CCPD THERAPY ON (B)(6) 2025. Manufacturer Narrative: PLANT INVESTIGATION: THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL INVESTIGATION: A TEMPORAL RELATIONSHIP EXISTS BETWEEN CCPD THERAPY UTILIZING THE LIBERTY SELECT CYCLER AND THE PATIENT¿S EXPIRATION, AS THE PATIENT WAS ACTIVELY UNDERGOING CCPD THERAPY AT THE TIME OF HER PASSING. THE PDRN REPORTED THE DEFINITIVE CAUSE OF DEATH WAS UNKNOWN; THEREFORE, CAUSALITY COULD NOT BE ESTABLISHED. THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. IT SHOULD BE NOTED THAT WHILE THERE WAS NO REPORT OF A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) MALFUNCTION OR DEFICIENCY, LIMITED CLINICAL FOLLOW-UP INFORMATION PRECLUDED A MORE COMPREHENSIVE INVESTIGATION. BASED ON THE INFORMATION AVAILABLE, THE PATIENT¿S LIBERTY SELECT CYCLER CANNOT BE DISASSOCIATED FROM THE SERIOUS ADVERSE EVENTS. THERE IS NO ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) CAUSED AND/OR CONTRIBUTED TO THE SERIOUS ADVERSE EVENTS. HOWEVER, GIVEN THE LACK OF CLINICAL FOLLOW-UP, TIMELINE OF EVENTS, MEDICAL INTERVENTION, CAUSALITY AND NO MANUFACTURER EVALUATION OF THE DEVICE, THIS CLINICAL INVESTIGATION CANNOT EXCLUDE THE PATIENT¿S LIBERTY SELECT CYCLER FROM PLAYING A POSSIBLE ROLE IN THE SERIOUS ADVERSE EVENTS. Manufacturer Narrative: CORRECTION: B5.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22702813&pc=DSQ | 2916596-2025-04878 | 2025-06-27 04:00 | Death | THORATEC CORPORATION | 2025-08-04 04:00 | DSQ | HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Purulent Discharge; Skin Infection | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS ADMITTED ON (B)(6) 2025 DUE TO LETHARGY AND PURULENT DRAINAGE FROM THE DRIVELINE. CULTURES FROM THE DRIVELINE AND BLOOD WERE POSITIVE FOR PSEUDOMONAS AERUGINOSA, AND LATER POSITIVE VIA STOOL CULTURE FOR CLOSTRIDIOIDES DIFFICILE INFECTION. THE PATIENT RECEIVED INTRAVENOUS (IV) ANTIBIOTICS BUT THE INFECTION WAS NOT RESOLVED. THE PATIENT DECIDED TO GO HOME ON HOSPICE CARE. ON (B)(6) 2025, THE PATIENT PASSED. THE CAUSE OF DEATH WAS DRIVELINE INFECTION, BACTEREMIA. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED. THE DEVICE HAD OPERATED AS EXPECTED. Manufacturer Narrative: SECTION B2 CORRECTION. SECTION H6 CORRECTION: HEALTH EFFECT - IMPACT CODE 4607 - HOSPITALIZATION OR PROLONGED. HOSPITALIZATION ADDED. MANUFACTURER'S INVESTIGATION CONCLUSION: A SPECIFIC CAUSE FOR THE PATIENT¿S DRIVELINE INFECTION COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. FURTHER, A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER: (B)(6) AND THE REPORTED EVENTS COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING INFECTION (LOCAL, DRIVELINE, AND PUMP POCKET), THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, LISTS INFECTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. FURTHERMORE, SEVERAL SECTIONS OF THE HM3 IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22705730&pc=NGV | 2135147-2025-04310 | 2025-07-11 04:00 | Death | ABBOTT MEDICAL | 2025-08-04 04:00 | NGV | AMPLATZER AMULET | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | P200049 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 20MM AMPLATZER AMULET WAS SELECTED FOR IMPLANT USING A 14F AMPLATZER TORQVUE DELIVERY SYSTEM. PRIOR TO PROCEDURE, THE PATIENT WAS IN STABLE CONDITION. BASELINE ECHO SHOWED AN EFFUSION AND FLUID FILLED TRANSVERSE SINUS. NO ANGIOGRAM OF THE LEFT ATRIAL APPENDAGE (LAA) WAS TAKEN. DEVICE SIZING WAS SOLELY BASED ON ECHO. THE LARGEST LANDING ZONE MEASURED 15MM. DURING IMPLANT, THE FIRST DEPLOYMENT "APPEARED TO HAVE A SMALL MITRAL SHOULDER, DISC SHOWED LITTLE SEPARATION." THE AMULET WAS RESHEATHED TO BALL. "SECOND LOBE DEPLOYMENT WAS POSITIONED DIFFERENTLY BUT NOT COAXIAL" AND THE AMULET WAS AGAIN RESHEATHED TO BALL. THE LAST DEPLOYMENT SHOWED ACCEPTABLE POSITION; DISC WAS EXPOSED AND FULL INTEGRATION OF THE DEVICE IN ALL FOUR PLANES WAS PERFORMED. CLOSE CRITERIA WAS ACHIEVED AND THE AMULET WAS RELEASED FOR IMPLANT. THE BASELINE EFFUSION REMAINED UNCHANGED. ON (B)(6) 2025, THE PATIENT BECAME BRADYCARDIC AND HYPOTENSIVE EARLY IN THE MORNING WITHOUT EXPRESSING COMPLAINTS AND WENT INTO PULSELESS ELECTRICAL ACTIVITY (PEA). THE PATIENT PASSED AWAY. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22706224&pc=OZD | 1220648-2025-30981 | 2023-10-15 04:00 | Death | ABIOMED, INC. | 2025-08-04 04:00 | OZD | IMPELLA | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P140003 |  | 1 | Event Description: IT WAS REPORTED VIA A MEDWATCH REPORT (MW5172995) THAT A PATIENT IMPLANTED WITH AN IMPELLA PUMP DECOMPENSATED. THE MEDICAL STAFF AND THE PATIENT'S FAMILY DECIDED TO WITHDRAW CARE AND THE PATIENT EXPIRED. Manufacturer Narrative: A1 IS UNKNOWN. PRODUCT DETAILS ARE UNKNOWN IN SECTION D4. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22707238&pc=NPT | 2025587-2025-05460 | 2025-04-26 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-08-04 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Cardiac Arrest | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT 1 MONTH FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT WAS FOUND COLLAPSED AND EMERGENCY MEDICAL SERVICES WERE CALLED. THE PATIENT WAS TRANSPORTED TO THE HOSPITAL UNDER CARDIOPULMONARY ARREST (CPA) AND A TOTAL OF 7 DOSES OF ADRENALINE WERE ADMINISTERED; HOWEVER, THE CARDIAC ARREST PERSISTED DURING ALL WAVEFORM CHECKS AND THE PATIENT DIED. THE CAUSE OF DEATH WAS UNKNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22708245&pc=NPT | 2015691-2025-06309 | 2025-07-16 04:00 | Death | EDWARDS LIFESCIENCES | 2025-08-04 04:00 | NPT | EDWARDS COMMANDER DELIVERY SYTEM | Burst Container or Vessel; Difficult to Remove | Vascular Dissection | P140031 |  | 1 | Event Description: AS REPORTED BY OUR EDWARDS LIFESCIENCES AFFILIATE IN BRAZIL, REGARDING 29 MM SAPIEN 3 TRANSCATHETER HEART VALVE WAS IMPLANTED IN THE AORTIC POSITION VIA A TRANSFEMORAL APPROACH, DURING THE PROCEDURE, THE BALLOON OF THE COMMANDER DELIVERY SYSTEM RUPTURED AT THE END OF INFLATION, WHICH WAS NOTED DURING DEFLATION DUE TO THE PRESENCE OF BLOOD IN THE SYRINGE. DESPITE THIS, THE VALVE WAS FULLY DEPLOYED. THERE WERE DIFFICULTIES WITHDRAWING THE COMMANDER SYSTEM THROUGH THE ESHEATH AND THE ESHEATH GOT DAMAGED (LINER TORN), NECESSITATING A SURGICAL CUTDOWN TO REMOVE THE DEVICE. A PARTIAL DISSECTION OF THE COMMON ILIAC ARTERY OCCURRED, AND A STENT WAS IMPLANTED. POST-PROCEDURE, THE PATIENT WAS IN STABLE CONDITION BUT PASSED AWAY ON THE SAME DAY AFTER SOME HOURS IN THE ICU. Manufacturer Narrative: THE INVESTIGATION IS ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22708428&pc=FKX | 0008030665-2025-01771 | 2025-07-19 04:00 | Death | ERIKA DE REYNOSA, S.A. DE C.V. | 2025-08-04 04:00 | FKX | LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | Adverse Event Without Identified Device or Use Problem | Sepsis | K222318 |  | 1 | Event Description: A PERITONEAL DIALYSIS REGISTERED NURSE (PDRN) REQUESTED ASSISTANCE IN CANCELING TREATMENT ON THE LIBERTY SELECT CYCLER DURING DWELL 3 OF 4 OF TREATMENT. PATIENT WAS CONNECTED DURING INCIDENT AND NEEDED TO GO TO THE HOSPITAL. TECHNICAL SUPPORT ASSISTED THE PDRN WITH CANCELLING TREATMENT. DURING FOLLOW-UP, THE PATIENT¿S PDRN CONFIRMED THE PATIENT EXPIRED ON (B)(6) 2024, AFTER HAVING BEEN ADMITTED EARLIER IN THE DAY FOR SEPSIS. THE PDRN STATED THE DEFINITIVE CAUSE OF THE PATIENT¿S SEPTICEMIA WAS UNKNOWN. HOWEVER, THE PDRN ALSO REPORTED THAT THE INFECTION WAS NOT FROM A PERITONEAL SOURCE, AS THE PATIENT¿S CELL COUNT RESULTS WERE NEGATIVE FOR ANY FINDINGS (RESULTS UNAVAILABLE, PATIENT¿S ELECTRONIC MEDICAL RECORD CLOSED). THE PATIENT WAS ADMITTED TO THE INTENSIVE CARE UNIT (ICU), WHERE THE PATIENT¿S CONDITION CONTINUED TO RAPIDLY DETERIORATE. THE TIMELINE AND SPECIFIC DETAILS SURROUNDING THE PATIENT¿S FINAL HOURS WERE UNAVAILABLE, HOWEVER IT IS KNOWN THE PATIENT WAS INTUBATED, ON MULTIPLE INTRAVENOUS VASOPRESSORS, AND WAS CONSIDERED IN CRITICAL CONDITION. THE PATIENT¿S FAMILY WAS CONSULTED, AND GIVEN THE PATIENT¿S POOR CHANCE FOR SURVIVAL, THE FAMILY AGREED TO TRANSITION THE PATIENT TO COMFORT MEASURES ONLY. THE PATIENT PASSED QUICKLY ON (B)(6) 2025 WITH FAMILY AT THE BEDSIDE. THE CAUSE OF THE PATIENT¿S DEATH WAS REPORTED TO THE PDRN AS SEPTICEMIA OF UNKNOWN ORIGIN. DESPITE THE CIRCUMSTANCES SURROUNDING THE PATIENTS PASSING, THE FAMILY ELECTED NOT TO HAVE AN AUTOPSY PERFORMED. THE PDRN REPORTED THAT DESPITE THE PATIENT UNDERGOING CCPD WHEN THE INITIAL SYMPTOMS BEGAN, THE PATIENT¿S PASSING WAS UNRELATED TO ANY FRESENIUS PRODUCT(S) AND/OR DEVICE(S) DEFICIENCY OR MALFUNCTION. Manufacturer Narrative: CLINICAL INVESTIGATION: A TEMPORAL RELATIONSHIP DOES NOT EXIST BETWEEN CCPD THERAPY UTILIZING THE LIBERTY SELECT CYCLER AND THE SERIOUS ADVERSE EVENT OF DEATH, AS THE PATIENT WAS NOT UNDERGOING CCPD THERAPY WHEN HE PASSED. THE PDRN REPORTED THE PATIENT BECAME SEPTIC DUE TO AN UNKNOWN INFECTION AND EVENTUALLY WITHDREW FROM CARE ON (B)(6) 2025. THE PATIENT REPORTEDLY PASSED AWAY HOURS LATER DUE TO SEPTICEMIA. PER THE PDRN, THE SERIOUS ADVERSE EVENTS WERE UNRELATED TO THE PATIENT¿S UTILIZATION OF ANY FRESENIUS PRODUCT(S) AND/OR DEVICE(S). THE PATIENT¿S DEATH WAS AN EXPECTED AND INEVITABLE OUTCOME GIVEN HIS WITHDRAWAL FROM CARE. THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. ADDITIONALLY, HOSPICE CARE IS DEFINED AS A MEDICAL INTERVENTION, WITH THE EXPECTED OUTCOME BEING THE PATIENT WILL EXPIRE AS A RESULT. BASED ON THE INFORMATION AVAILABLE, THE PATIENT¿S LIBERTY SELECT CYCLER CAN BE DISASSOCIATED FROM THE SERIOUS ADVERSE EVENTS. THERE IS NO ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) CAUSED OR CONTRIBUTED TO THE SERIOUS ADVERSE EVENTS. FURTHERMORE, THERE IS NO REPORT THAT A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) FAILED TO MEET THE USERS¿ EXPECTATIONS AND/OR THE MANUFACTURERS¿ SPECIFICATIONS, AS THE PATIENT CONTINUES TO UTILIZE THE SAME LIBERTY SELECT CYCLER WITHOUT REPORTED ISSUE. THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22708465&pc=MFE | 2029214-2025-01751 | 2025-04-03 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-08-04 04:00 | MFE | ONYX | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING: "PARTICLES, LIQUIDS, AND ANTIPLATELETS: RESTARTING ANTIPLATELETS ON POSTINTERVENTION DAY 3 OR LAT ER IS ASSOCIATED WITH IMPROVED OUTCOMES AFTER PARTICLE EMBOLIZATION." THE OBJECTIVE WAS TO DETERMINE WHETHER THE TIMING OF ANTIPLATELET RESTART AFTER MIDDLE MENINGEAL ARTERY EMBOLIZATION (MMAE) FOR CHRONIC SUBDURAL HEMATOMA (CSDH) AFFECTS LONG-TERM RADIOGRAPHIC AND CLINICAL OUTCOMES. THE TIME FRAME OF THIS STUDY WAS: FROM 2018 TO 2023, USING DATA FROM 4 ACADEMIC MEDICAL CENTERS THAT PERFORMED MMAE FOR CSDH. MULTIPLE MANUFACTURER¿S DEVICES WERE USED IN THE STUDY POPULATION. 34 OUT OF 128 PATIENTS WERE TREATED WITH ONYX LIQUID EMBOLIC. AMONG PATIENT ADVERSE EVENTS INCLUDED: 90-DAY MORTALITY: 6 PATIENTS (20.7%) IN THE ONYX GROUP. THE CAUSE FOR MORTALITY WAS NOT COLLECTED IN THIS ANALYSIS NEED FOR SURGICAL RETREATMENT OCCURRED IN 12 TOTAL PATIENTS (9.7%). (PAGE 16) FUNCTIONAL INDEPENDENCE AT 90 DAYS: 19 ONYX PATIENTS (65.5%) WERE FUNCTIONALLY INDEPENDENT (MRS 3) NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID UNK-NV-ONYX (UNKNOWN); PRODUCT TYPE: ; IMPLANT DATE N/A; EXPLANT DATE N/A CITATION: PATEL, P. D., SALWI, S., SIOUTAS, G., FILO, J., DOHEIM, M. F., DIAMOND, A., SALEM, M. M., MA, L., HOZ, S. S., SULTANY, A., NOGUEIRA, R. G., AL-BAYATI, A. R., LANG, M. J., GOOCH, M. R., TJOUMAKARIS, S.. PARTICLES, LIQUIDS, AND ANTIPLATELETS: RESTARTING ANTIPLATELETS ON POSTINTERVENTION DAY 3 OR LATER IS ASSOCIATED WITH IMPROVED OUTCOMES AFTER PARTICLE EMBOLIZATION. NEUROSURGERY APRIL 2025. DOI:10.1227/NEU.0000000000003440 EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT. NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22709105&pc=PIF | 1423537-2025-00300 |  | Death | CARDINAL HEALTH, INC. | 2025-08-04 04:00 | PIF | SALEM SUMP | Obstruction of Flow | Aspiration/Inhalation; Vomiting; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K190923 |  | 1 | Event Description: THE CUSTOMER REPORTED AN INSERTION ON (B)(6) 2025 OF A NASOGASTRIC TUBE ON A PATIENT WHO WAS POORLY RELIEVED. THE PATIENT WAS TRANSFERRED TO IMAGING FOR A CT SCAN TO IDENTIFY THE CAUSE OF THE OCCLUSION. PERSISTENCE OF VOMITING ON CT SCAN WITH MASSIVE INHALATION RESPONSIBLE FOR CARDIOPULMONARY ARREST NOT RECOVERED IN RADIOLOGY. IN ALL LIKELIHOOD, THE SUCTION OF THE NASOGASTRIC TUBE WAS UNSATISFACTORY DUE TO THE LACK OF OPENING OF THE AIR INTAKE SYSTEM. INDEED, IN THE ABSENCE OF AN ANTI-BACKFLOW VALVE, THE ENFIT CONNECTOR HAS BEEN CONNECTED TO THE PROBE, WITH THE CAP CLOSED. CURRENT PATIENT CONDITION: ABSENCE OF SUCTION FOR SEVERAL HOURS. CARDIOPULMONARY ARREST FOLLOWED BY THE PATIENT'S DEATH. ON (B)(6) 2025, (B)(6) CARDINAL HEALTH SALES REP, REPORTED SHE HAD A DISCUSSION WITH CUSTOMER, (B)(6), HEAD NURSE AND CONFIRMED THERE WERE NO PROBLEMS WITH THE DEVICE. INSTEAD, THERE WAS A LACK OF KNOWLEDGE ON THE PART OF SOME CAREGIVERS. ADDITIONAL PRODUCT TRAINING WILL BE PROVIDED BY CARDINAL HEALTH SALES TEAM. THE SALES REP ALSO STATED THAT THE CAPPED ENFIT WAS IN THE BLUE PORT. THE FACILITY DID NOT HAVE ACTIVE SUCTION HOOKED UP TO THE CLEAR PART OF THE TUBE. THEY HAD GOTTEN RETURN OF STOMACH CONTENTS FROM THIS TUBE. Manufacturer Narrative: AN INVESTIGATION IS CURRENTLY UNDERWAY. UPON COMPLETION, THE RESULTS WILL BE FORWARDED. SECTION D2A PRODUCT CODE (SECONDARY CODE): FEG, TUBE, DOUBLE LUMEN FOR INTESTINAL DECOMPRESSION AND/OR INTUBATION. SECTION H6 HEALTH EFFECT - CLINICAL CODE, APPROPRIATE CLINICAL SIGNS, SYMPTOMS AND CONDITIONS TERM/CODE NOT AVAILABLE: CARDIOPULMONARY ARREST.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22709318&pc=NPT | 2015691-2025-06322 | 2025-04-18 04:00 | Death | EDWARDS LIFESCIENCES | 2025-08-04 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE SYSTEM | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P140031 |  | 1 | Event Description: AS REPORTED THROUGH EDWARDS LIFESCIENCES IMPLANT PATIENT REGISTRY DEPARTMENT, REGARDING A 23MM SAPIEN 3 ULTRA RESILIA VALVE IN THE AORTIC POSITION, 17 DAYS POST IMPLANT THE PATIENT PASSED AWAY DUE TO A STROKE. NO OTHER INFORMATION WAS ABLE TO BE OBTAINED. Manufacturer Narrative: E1 PHONE NUMBER (B)(6). PER THE INSTRUCTIONS FOR USE (IFU), PERMANENT OR TRANSIENT NEUROLOGICAL EVENTS SUCH AS TRANSIENT ISCHEMIC ATTACK (TIA) AND STROKE ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND THE USE OF THE EDWARDS TRANSCATHETER HEART VALVE (THV) DEVICES. ACCORDING TO THE LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EW, STROKE IS RECOGNIZED IN THE LITERATURE AS A WELL-KNOWN COMPLICATION IN A SMALL NUMBER OF PATIENTS UNDERGOING THV. RISK FACTORS CORRELATING WITH SEVERAL PATIENT CO-MORBIDITIES HAVE BEEN IDENTIFIED. ALTHOUGH IN MANY CASES THE ROOT CAUSE OF THE EVENT IS UNABLE TO BE DETERMINED, STROKES DURING THV ARE UNDOUBTEDLY MULTIFACTORIAL, THE DOMINANT ETIOLOGY LIKELY BEING INTRAPROCEDURE EMBOLIC EVENTS. A TRANSCRANIAL DOPPLER STUDY DURING THV DEMONSTRATED THAT MOST PROCEDURAL EMBOLIC EVENTS OCCURRED DURING BALLOON VALVULOPLASTY, MANIPULATION OF CATHETERS ACROSS THE AORTIC VALVE, AND VALVE IMPLANTATION. AN ANALYSIS IN PATIENTS UNDERGOING VALVE SURGERY REVEALED FOUR BASELINE CHARACTERISTICS AND TWO PROCEDURAL EVENTS THAT WERE ASSOCIATED WITH EARLY POST-PROCEDURE STROKE: FEMALE SEX, EF < 30%, DIABETES, AGE OLDER THAN 70 YEARS, BYPASS PROCEDURE TIME> 120 MIN, AND CALCIFICATION OF THE ASCENDING AORTA. PREDICTORS OF LATE STROKE HAVE INCLUDED FEMALE SEX, AGE OLDER THAN 75 YEARS, ATRIAL FIBRILLATION, AND A HISTORY OF OR CURRENT SMOKING. THERE WERE NO SIGNIFICANT DIFFERENCES IN THE FREQUENCY OF LATE STROKES BETWEEN THV AND AVR PATIENTS. AFTER THV, THERE APPEARS TO BE A MORE CONSIDERABLE PROPORTION OF EARLY STROKES OCCURRING < 24 H POST-PROCEDURE, BUT THV PATIENTS WITH MULTIPLE CO-MORBIDITIES ARE PROBABLY AT HIGHER RISK OF BOTH EARLY AND LATE STROKES. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST/INDICATE PATIENT FACTORS (AGE AND GENDER) CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22710742&pc=NPT | 2135147-2025-04333 | 2025-07-09 04:00 | Death | ABBOTT MEDICAL | 2025-08-04 04:00 | NPT | NAVITOR VISION | Difficult to Fold, Unfold or Collapse; Migration or Expulsion of Device; Perivalvular Leak; Improper or Incorrect Procedure or Method | Perforation; Brain Injury; Vascular Dissection | P190023 |  | 1 | Event Description: CLINICAL INFORMATION: (B)(6) - ENVISION IDE STUDY, PATIENT SITE ID:(B)(6)). IT WAS REPORTED THAT ON (B)(6) 2025, A 27MM NAVITOR VISION VALVE WAS SUCCESSFULLY IMPLANTED IN PATIENT WITH SIGNIFICANT CALCIUM BURDEN, INCLUDING CENTRAL LEAFLET CALCIFICATION AND A LARGE CALCIUM DEPOSIT IN THE LVOT JUST 1MM BELOW THE ANNULUS. THE PATIENT WAS CONSIDERED INTERMEDIATE RISK PRIOR TO THE IMPLANT PROCEDURE. THE AORTIC VALVE AREA (AVA) WAS 0.7 CM², THE INDEXED EFFECTIVE ORIFICE AREA (EOA) WAS 0.3 CM², AND THERE WAS A LARGE CALCIUM DEPOSIT IN THE LEFT VENTRICULAR OUTFLOW TRACT (LVOT) MEASURING 16.9 MM IN CROSS-SECTION, LOCATED 1 MM BELOW THE ANNULUS. THE AORTIC VALVE ANGLE WAS MEASURED AT 53 DEGREES. INITIALLY IT WAS PLANNED TO USE A 21MM NON-ABBOTT FOR PRE-DILATION, CONSIDERING THE HEAVY LVOT CALCIUM BURDEN, AND THIS PLAN WAS AGREED UPON AND PREPARED. THE PATIENT WAS CONSIDERED INTERMEDIATE RISK PRIOR TO THE IMPLANT PROCEDURE. A COMPUTED TOMOGRAPHY SCAN SUMMARY WAS REVIEWED BY A SECOND PHYSICIAN AND OVERRULED THE PRE-DILATION PLAN, INSISTING ON USING AN 18MM NON-ABBOTT BALLOON INSTEAD. DESPITE CONCERNS RAISED REGARDING THE INADEQUACY OF AN 18MM BALLOON FOR A 27MM NAVITOR VISION VALVE WITH AN 81.4 MM PERIMETER, THE PROCEDURE PROCEEDED WITH THE SMALLER BALLOON. DURING INFLATION, IT WAS NOTED THAT THE 18MM BALLOON WAS STRUGGLING TO MOVE THE CALCIUM, AND BOTH PHYSICIANS AGREED IT WAS INSUFFICIENT. NEVERTHELESS, THE VALVE WAS DEPLOYED, THE VALVE WAS DELIVERED AND DEPLOYED AT A DEPTH OF 3¿4 MM ACROSS THE ANNULUS. IT APPEARED CONSTRAINED AND UNDER-EXPANDED, RESULTING IN SEVERE PARAVALVULAR LEAK (PVL) AND CONFIRMED VIA FLUOROSCOPY. THERE WAS SUFFICIENT AORTIC VALVE CALCIUM TO ALLOW ANCHORING. DESPITE MULTIPLE WARNINGS THAT THE VALVE LACKED RADIAL FORCE DUE TO INADEQUATE PRE-DILATION, THE VALVE WAS DEPLOYED PACING AT 130 BPM AND RELEASING THE VALVE SLOWLY WITH FORWARD PRESSURE. IN PREPARATION FOR FINAL DEPLOYMENT, THE GUIDEWIRE WAS PULLED TO A MIDVENTRICULAR POSITION TO DEFLECT THE NOSECONE. DURING FINAL DEPLOYMENT, THE IMPLANTER CONFIRMED THAT ALL THREE TABS WERE SUCCESSFULLY RELEASED USING TWO DIFFERENT VIEWS. UPON RELEASE, THE VALVE IMMEDIATELY EMBOLIZED INTO THE SINUSES. INITIAL ATTEMPTS TO RETRIEVE AND REPOSITION THE VALVE USING A NON-ABBOTT SNARE AND 23MM BALLOON WERE UNSUCCESSFUL, AS THE BALLOON FAILED TO CONTACT THE VALVE FRAME AND BOUNCED WITHIN THE LVOT AND ASCENDING AORTA. A 25MM NON-ABBOTT BALLOON WAS EVENTUALLY USED TO REPOSITION THE VALVE ABOVE THE CORONARY OSTIA. A GOOSENECK SNARE WAS APPLIED TO STABILIZE THE VALVE WHILE A NEW 23MM NON-ABBOTT DEVICE WAS DELIVERED AND DEPLOYED, BUT NOT IN THE 27MM NAVITOR VISION VALVE. POST-DEPLOYMENT, ANGIOGRAPHY REVEALED CONCENTRIC SEVERE PVL, WHICH WAS REDUCED TO MILD-TO-MODERATE AFTER INFLATING THE 25MM NON-ABBOTT BALLOON BY 1 CC. THE TEAM DISCUSSED RETURNING IN A WEEK TO PLUG THE ANNULUS TO ADDRESS THE RESIDUAL PVL. POST-PROCEDURE, IT'S NOTED THAT THE PATIENT'S CONDITION HAD DETERIORATED DUE TO ANNULAR AORTIC RUPTURE AND REQUIRED EMERGENCY SURGERY. FINDINGS INCLUDED ANNULAR RUPTURE, NECESSITATING ROOT REPAIR AND AORTIC VALVE REPLACEMENT (AVR), AS WELL AS REPAIR OF AN AORTIC DISSECTION. THE PATIENT WAS ALSO TRANSFUSED WITH RED BLOOD CELLS. ON (B)(6) 2025, THE PATIENT WAS SUFFERED LEADING UP TO DEATH, THEY EXPERIENCED COMPLETE HEART BLOCK, WERE NEUROLOGICALLY UNRESPONSIVE WITH RIGHT GAZE DEVIATION, AND HAD A PERICARDIAL EFFUSION THAT REQUIRED PERICARDIAL DRAINAGE DURING CPR. ULTIMATELY, HE PATIENT SUFFERED AN ANOXIC BRAIN INJURY AND PASSED AWAY ON (B)(6) 2025. THE ANNULAR RUPTURE AND AORTIC DISSECTION WERE BELIEVED TO BE CAUSED BY CONSTRAINT OF THE VALVE, WHILE THE ANOXIC BRAIN INJURY WAS OF UNKNOWN CAUSE. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22696512&pc=LYJ | 1644487-2025-10454 | 2025-07-12 04:00 | Death | HOUSTON | 2025-08-03 04:00 | LYJ | PULSE GEN MODEL 1000 | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P970003 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN AT TIME OF THE REPORT. NO OTHER RELEVANT INFORMATION HAS BEEN RECEIVED TO DATE. Manufacturer Narrative: LIVANOVA USA, INC. SUBMITS THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION, BASED ON INFORMATION THAT LIVANOVA HAS OBTAINED, BUT MAY NOT HAVE BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION, OR A CONCLUSION BY FDA OR ANYONE ELSE, THAT THE DEVICE, LIVANOVA, OR LIVANOVA'S EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS¿ OR "MALFUNCTIONS¿. THESE WORDS ARE INCORPORATED INTO THE FDA 3500A MEDWATCH FORM BY THE FDA, AND LIVANOVA OBJECTS TO THEIR USE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22696649&pc=BZD | 2518422-2025-108509 | 2025-07-23 04:00 | Death | RESPIRONICS, INC. | 2025-08-03 04:00 | BZD | DREAMSTATION | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO AN ALLEGATION OF A DEATH OF A PATIENT. IT IS NOT YET CLEAR IF THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH OF THE PATIENT. THERE IS NO ALLEGATION OF PRODUCT MALFUNCTION. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A SUPPLEMENTAL FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22684354&pc=OZP | 2032227-2025-223328 | 2025-06-29 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-08-01 04:00 | OZP | MINIMED 780G US SYSTEM BLE CONNECT 3.0 MG/DL | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Hyperglycemia; Hypoglycemia | P160017 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER'S NEXT OF KIN INFORMED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2025, IN THE HOSPITAL. THE CAUSE OF DEATH WAS HEART PROBLEM AND DIABETES COMPLICATIONS. THE LAST KNOWN PRODUCTS FOR THIS CUSTOMER WERE AS FOLLOWS: MMT-1884, UNK_RESERVOIR, MMT-431AJ, MMT-7841ZN. THE TROUBLESHOOTING WAS PERFORMED. THE PUMP AND SENSOR WAS WORN AT THE TIME OF PASSING. NO PRODUCT RETURN IS REQUIRED FOR MMT-1884. NO PRODUCT RETURN IS REQUIRED FOR UNK_RESERVOIR. NO PRODUCT RETURN IS REQUIRED FOR MMT-431AJ. NO PRODUCT RETURN IS REQUIRED FOR MMT-7841ZN. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22684687&pc=PNJ | 9612164-2025-03767 | 2025-07-11 04:00 | Death | MEDTRONIC IRELAND | 2025-08-01 04:00 | PNJ | MICRA | Appropriate Term/Code Not Available | Pseudoaneurysm; Thrombosis/Thrombus | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT POST OPERATIVELY AFTER THE PLACEMENT OF THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG), THE PATIENT PASSED AWAY DUE TO COMPLICATIONS ASSOCIATED WITH THE PROCEDURE. FEMORAL VEIN PSEUDOANEURYSM AND THROMBOSIS WAS NOTED. PATIENT WAS ADMITTED TO THE INTENSIVE CARE UNIT (ICU), DID NOT RECOVER AND LATER PASSED AWAY. Manufacturer Narrative: SECTION D REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [MICRA] PRODUCT ID [MC1VR01-DELSYS] (SERIAL: (B)(6)). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22684697&pc=DYB | 2953200-2025-00010 | 2025-07-11 04:00 | Death | MEDTRONIC VASCULAR | 2025-08-01 04:00 | DYB | MICRA | Appropriate Term/Code Not Available | Pseudoaneurysm; Thrombosis/Thrombus | K132030 |  | 1 | Event Description: IT WAS REPORTED THAT POST OPERATIVELY AFTER THE PLACEMENT OF THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG), THE PATIENT PASSED AWAY DUE TO COMPLICATIONS ASSOCIATED WITH THE PROCEDURE. FEMORAL VEIN PSEUDOANEURYSM AND THROMBOSIS WAS NOTED. PATIENT WAS ADMITTED TO THE INTENSIVE CARE UNIT (ICU), DID NOT RECOVER AND LATER PASSED AWAY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22684834&pc=MKJ | 0003015876-2025-01542 | 2025-07-08 04:00 | Death | PHYSIO-CONTROL, INC. - 3015876 | 2025-08-01 04:00 | MKJ | LIFEPAK(R) CR2 DEFIBRILLATOR | Failure to Power Up | No Clinical Signs, Symptoms or Conditions | P170018 |  | 1 | Event Description: THE CUSTOMER CONTACTED STRYKER TO REPORT THAT THEIR DEVICE DID NOT FUNCTION DURING A RESUSCITATION. THE PATIENT ASSOCIATED WITH THE REPORTED EVENT DID NOT SURVIVE. Manufacturer Narrative: IN COMPLIANCE WITH REGULATION (EU) 2016/679 OF THE EUROPEAN PARLIAMENT AND OF THE COUNCIL, STRYKER COLLECTS ONLY ESSENTIAL AND RELEVANT PATIENT INFORMATION NECESSARY FOR REGULATORY PURPOSES, ENSURING THAT NO DATA CAPABLE OF IDENTIFYING A PERSON, DIRECTLY OR INDIRECTLY, IS GATHERED. PATIENT FIELDS IN WHICH INFORMATION IS NOT REQUESTED OR PROVIDED WERE INTENTIONALLY LEFT BLANK. A STRYKER SERVICE REPRESENTATIVE PERFORMED AN EVALUATION OF THE CUSTOMER¿S DEVICE AND WAS UNABLE TO DUPLICATE THE REPORTED ISSUE. THE ELECTRONIC RECORDS WERE DOWNLOADED FOR REVIEW. THE PRODUCT ASSESSMENT CENTER (PAC) TECHNICIAN EVALUATED THE DEVICE AND THE BATTERY AND COULD NOT DETECT ANY MALFUNCTION. A CAUSE OF THE REPORTED ISSUE COULD NOT BE DETERMINED. A CLINICAL REVIEW OF THIS CASE WAS PERFORMED: INSUFFICIENT DATA WITHIN THE FILE TO DETERMINE THE IMPACT OF THE DEVICE USE. THE DETAILS OF THE MALFUNCTION OF THE DEVICE ARE UNCLEAR. NONE OF THE DEVICE FILES CONTAIN PATIENT DATA. THE CUSTOMER CURRENTLY HAS NO ADDITIONAL INFORMATION REGARDING THE EVENT. OVERALL, THERE IS NOT ENOUGH INFORMATION TO FORM A CONCLUSION REGARDING CONTRIBUTION. DEVICE USE CONTRIBUTION IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22685031&pc=OZP | 2032227-2025-223731 | 2025-06-14 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-08-01 04:00 | OZP | MINIMED 780G US SYSTEM BLE CONNECT 3.0 MG/DL | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Hyperglycemia; Hypoglycemia; Myocardial Infarction; Renal Failure; Heart Failure/Congestive Heart Failure | P160017 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY ON THE (B)(6) 2025 DUE TO NATURAL CAUSES, INCLUDING CHRONIC CONGESTIVE HEART FAILURE, CORONARY ARTERIAL DISEASE, CHRONIC KIDNEY DISEASE AND DIABETES. THE EVENT INVOLVED PRODUCT(S) MMT-1884, MMT-332A, MMT-7841ZN, MMT-7040A & MMT-394A. TROUBLESHOOTING WAS DONE, FOUND THE CUSTOMER PASSED AWAY AT HOME AND WAS WEARING THE INSULIN PUMP AT THE TIME OF PASSING. IT WAS UNKNOWN IF THE AUTOMODE/SMART GUARD FEATURE WAS ACTIVE. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED. THE PRODUCT(S) MMT-1884, MMT-332A, MMT-7841ZN, MMT-7040A & MMT-394A WILL NOT BE RETURNED FOR ANALYSIS. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22685176&pc=OZD | 1220648-2025-30911 | 2025-07-14 04:00 | Death | ABIOMED, INC. | 2025-08-01 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhagic Stroke | P140003 |  | 1 | Event Description: A COMPLAINANT REPORTED THAT A PATIENT WITH PULMONARY CONGESTION PRESENTED FOR PERCUTANEOUS CORONARY INTERVENTION WITH IMPELLA CP SUPPORT. DURING SUPPORT, THE PATIENT DEVELOPED A SUBARACHNOID HEMORRHAGE. THE DECISION WAS MADE TO WITHDRAW THE PROCEDURE DUE TO POOR NEUROLOGICAL PROGNOSIS CAUSED BY THE IMPELLA-INDUCED SUBARACHNOID HEMORRHAGE. THE PATIENT PASSED AWAY TWO DAYS AFTER REMOVAL. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: B.2 THE EXACT DATE OF DEATH IS UNKNOWN. THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22685313&pc=OZD | 1220648-2025-30913 | 2025-07-20 04:00 | Death | ABIOMED, INC. | 2025-08-01 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hematoma | P140003 |  | 1 | Event Description: THE COMPLAINANT HAD IMPELLA CP SUPPORT PLACED FOR A ST-SEGMENT ELEVATION MYOCARDIAL INFARCTION PATIENT WHO HAD SEPSIS AND CHEST PAIN. THE PUMP WAS PLACED BUT WAS EXPLANTED AFTER JUST TWO HOURS DUE TO AN ACCESS SITE HEMATOMA. THE FIELD REPRESENTATIVE LATER UPDATED THE ABIOMED TEAM THAT THE PATIENT DID EXPIRE SHORTLY AFTER THE CASE. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM, SECTION: GENERAL PATIENT CARE CONSIDERATIONS: ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT: USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER, ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22685351&pc=MVK | 3008642652-2025-07389 | 2025-07-01 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-08-01 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Incorrect Interpretation of Signal | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON(B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED THREE INAPPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 09:36:08 ON (B)(6) 2025. AT 20:01:22, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS SINUS TACHYCARDIA @ 120 BPM WITH LOW AMPLITUDE CARDIAC SIGNAL AND MOTION ARTIFACT. AT 20:03:03, THE PATIENT RECEIVED THE FIRST INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL AND MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS SINUS TACHYCARDIA @ 100 BPM WITH PVCS, LOW AMPLITUDE CARDIAC SIGNAL AND MOTION ARTIFACT. POST SHOCK RHYTHM WAS SINUS TACHYCARDIA @ 110 BPM. AT 20:21:45, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS SINUS TACHYCARDIA @ 110 BPM WITH PVCS, LOW AMPLITUDE CARDIAC SIGNAL AND MOTION ARTIFACT. AT 20:22:17, THE PATIENT RECEIVED THE SECOND INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL AND MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS SINUS TACHYCARDIA @ 100 BPM WITH PVCS, LOW AMPLITUDE CARDIAC SIGNAL AND MOTION ARTIFACT. POST SHOCK RHYTHM WAS SINUS TACHYCARDIA @ 110 BPM. AT 20:44:43, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS SINUS RHYTHM @ 90 BPM WITH PVCS, LOW AMPLITUDE CARDIAC SIGNAL AND MOTION ARTIFACT. AT 20:45:38, THE PATIENT RECEIVED THE THIRD INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL AND MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS SINUS TACHYCARDIA @ 110 BPM WITH PVCS, LOW AMPLITUDE CARDIAC SIGNAL AND MOTION ARTIFACT. POST SHOCK RHYTHM WAS SINUS RHYTHM @ 80 BPM WITH PVCS, LOW AMPLITUDE CARDIAC SIGNAL AND MOTION ARTIFACT. THE DEVICE WAS SHUT DOWN AT 23:41:53 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22685359&pc=MVK | 3008642652-2025-07390 | 2025-06-26 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-08-01 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED THREE APPROPRIATE TREATMENTS AND TWO NON-LIFEVEST DEFIBRILLATIONS. THE DEVICE WAS STARTED UP AT 09:43:22 ON (B)(6) 2025. AT 06:43:04 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS SINUS RHYTHM @ 80 BPM WITH PVCS TRANSITIONING TO VT @ 200 BPM. AT 06:43:41, THE PATIENT RECEIVED THE FIRST APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 210 BPM. POST SHOCK RHYTHM WAS VT FROM 115-180 BPM. AT 06:44:07, THE PATIENT RECEIVED THE SECOND APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 210 BPM WITH MOTION ARTIFACT. POST SHOCK RHYTHM WAS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION TRANSITIONING TO SINUS BRADYCARDIA @ 30 BPM WITH CPR/MOTION ARTIFACT. POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY. AT 06:45:47, THE PATIENT RECEIVED THE FIRST NON-LIFEVEST DEFIBRILLATION. RHYTHM AT TIME OF TREATMENT WAS SINUS RHYTHM @ 60 BPM. THE POST SHOCK RHYTHM WAS SINUS BRADYCARDIA @ 40 BPM. AT 07:10:04, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VT @ 190 BPM. AT 07:10:35, THE PATIENT RECEIVED THE THIRD APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS SINUS BRADYCARDIA @ 30 BPM WITH PVCS/PACS AND CPR/MOTION ARTIFACT. AT 07:13:20, THE PATIENT RECEIVED THE SECOND NON-LIFEVEST DEFIBRILLATION. THE RHYTHM AT THE TIME OF THE NON-LIFEVEST DEFIBRILLATIONS WAS VT @ 140 BPM. POST SHOCK RHYTHM WAS SINUS BRADYCARDIA @ 50 BPM WITH CPR/MOTION ARTIFACT. THE PATIENT RECEIVED A NON-LIFEVEST DEFIBRILLATION 46 SECONDS INTO THE DETECTION SEQUENCE. THE ELECTRODE BELT WAS DISCONNECTED AT 07:13:39 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR HAS BEEN COMPLETED. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION. THE ELECTRODE BELT HAS NOT BEEN RECOVERED. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22685363&pc=MVK | 3008642652-2025-07391 | 2025-06-29 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-08-01 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED TWO INAPPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 12:32:29 ON (B)(6) 2025. AT 02:25:02 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH CPR/MOTION ARTIFACT. AT 02:25:39, THE PATIENT RECEIVED THE FIRST INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON-LIFE-SUSTAINING RHYTHM, WITH PACER SPIKES. AT 02:39:53, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH CPR/MOTION ARTIFACT AND INTERMITTENT CARDIAC ACTIVITY. AT 02:41:00, THE PATIENT RECEIVED THE SECOND INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH CPR/MOTION ARTIFACT AND INTERMITTENT CARDIAC ACTIVITY. POST SHOCK RHYTHM WAS PACED RHYTHM @ 70 BPM. THE ELECTRODE BELT WAS DISCONNECTED AT 03:07:02 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22685369&pc=MVK | 3008642652-2025-07392 | 2025-06-06 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-08-01 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED EIGHT INAPPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 13:48:33 ON (B)(6) 2025. AT 20:38:53, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE. AT 20:39:34, THE PATIENT RECEIVED THE FIRST INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON-LIFE-SUSTAINING RHYTHM, WITH CPR/ELECTRODE LEAD FALL OFF. AT 21:15:38, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH CPR/MOTION ARTIFACT. AT 21:17:11, THE PATIENT RECEIVED THE SECOND INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS CPR/MOTION ARTIFACT. AT 21:17:42, THE PATIENT RECEIVED THE THIRD INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS CPR/MOTION ARTIFACT. AT 21:18:05, THE PATIENT RECEIVED THE FOURTH INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS CPR/MOTION ARTIFACT. AT 21:18:37, THE PATIENT RECEIVED THE FIFTH INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS CPR/MOTION ARTIFACT. AT 21:19:13, THE PATIENT RECEIVED THE SIXTH INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS CPR/MOTION ARTIFACT. AT 21:19:55, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH CPR/MOTION ARTIFACT. AT 21:20:40, THE PATIENT RECEIVED THE SEVENTH INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS CPR/MOTION ARTIFACT. AT 21:21:19, THE PATIENT RECEIVED THE EIGHTH INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS CPR/MOTION ARTIFACT. THE ELECTRODE BELT WAS DISCONNECTED AT 21:39:22 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR HAS BEEN COMPLETED. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION. THE ELECTRODE BELT HAS NOT BEEN RECOVERED. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22685440&pc=QZI | MW5173762 | 2025-07-03 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-08-01 04:00 | QZI | FARAPULSE | No Display/Image; Defective Device | Cardiac Perforation |  |  | 1 | Event Description: OUR MOM WENT IN FOR A HEART ABLATION AND THE FARAPULSE SYSTEM FAILED. IMAGE WAS NOT AVAILABLE ON SCREEN AND A HOLE WAS PUNCHED IN OUR MOM'S HEART. SHE DIED DUE TO THIS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22685498&pc=MNS | 2518422-2025-048761 | 2023-03-08 05:00 | Death | RESPIRONICS, INC. | 2025-08-01 04:00 | MNS | DREAMSTATION ST30 | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K102465 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION ST30 DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22685529&pc=PNJ | 9612164-2025-03783 | 2025-07-09 04:00 | Death | MEDTRONIC IRELAND | 2025-08-01 04:00 | PNJ | MICRA¿ VR2 | Appropriate Term/Code Not Available | Heart Failure/Congestive Heart Failure | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IS DECEASED. IT WAS FURTHER REPORTED THAT THE CAUSE OF DEATH WAS DUE TO HEART FAILURE AND WORSENING OF DILATED CARDIOMYOPATHY. THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM REMAINS IN USE. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22685695&pc=IYN | 3019216-2025-000268 | 2025-07-23 04:00 | Death | PHILIPS ULTRASOUND, INC | 2025-08-01 04:00 | IYN | CX50 | Insufficient Information | Insufficient Information | K081802 |  | 1 | Event Description: A CUSTOMER REPORTED A PATIENT HAD EXPIRED THE SAME DAY THEIR CX50 ULTRASOUND SYSTEM WAS USED UPON THEM. NO FURTHER INFORMATION WAS PROVIDED NOR COULD BE OBTAINED AT THIS TIME. THE SERVICE ENGINEER REPORTED THE SYSTEM AND TRANSDUCER TESTING PASSED AND THE SYSTEM WAS FULLY OPERATIONAL AND FUNCTIONAL. PHILIPS HAS STARTED AN INVESTIGATION, AND UPON COMPLETION, A FOLLOW UP REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22686048&pc=OZD | 1220648-2025-30925 | 2025-07-19 04:00 | Death | ABIOMED, INC. | 2025-08-01 04:00 | OZD | IMPELLA CP | Fluid/Blood Leak | Cardiac Arrest | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WITH AN IMPELLA CP FOR SUPPORT DUE TO ACUTE MYOCARDIAL INFARCTION. IT WAS NOTED THAT THE IMPELLA CP SERIAL NUMBER (B)(6) WAS OPENED AND PREPPED. DURING SET UP AND PRIOR TO INSERTION, THE PURGE CASSETTE WAS NOT PROPERLY SEATED AND CRACKED WHEN ATTEMPTING TO CLOSE THE PURGE DOOR. ONCE A NEW PURGE CASSETTE WAS REPLACED, THE PHYSICIAN REALIZED IT WAS A PURGE ISSUE AND NOT A PUMP ISSUE. PUMP SERIAL NUMBER (B)(6) WAS CONNECTED AND IMPLANTED FOR DURATION OF THE CASE WITHOUT ISSUE. PUMP SERIAL NUMBER (B)(6) WAS NEVER CONNECTED AND DISCARDED. DURING SUPPORT IT WAS REPORTED THAT ON BEDSIDE ECHOCARDIOGRAM THE IMPELLA LOOKED TO BE POINTED AT THE MITRAL VALVE. URINE WAS LIGHT PINK. THE PHARMACY WAS GETTING PURGE SOLUTION WITH SODIUM BICARBONATE ADDED AND PEDAL PULSES WERE OPERABLE. THE SITE WAS CLEAN, DRY AND INTACT. ADDITIONALLY, THE PATIENT WAS TAKEN TO THE INTENSIVE CARE UNIT AND WHILE BEING CHECKED IN BEGAN TO CODE. THE PATIENT WAS LOSING VOLUME RAPIDLY OUT OF THE CHEST TUBES. OVER THREE LITERS WERE LOST DURING THE CODE AND THE PATIENT REMAINED IN PULSELESS ELECTRICAL ACTIVITY FOR THE DURATION OF THE CODE. RHYTHM WAS NOT OBTAINED BACK AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22686120&pc=GDW | 1219930-2025-03406 | 2025-05-30 04:00 | Death | COVIDIEN SURGICAL | 2025-08-01 04:00 | GDW | UNKNOWN EGIA TRI STAPLE | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  |  | 1 | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY COMPARED THE OUTCOMES OF PATIENTS WHO UNDERWENT LIVER RESECTION UTILIZING A HYBRID TECHNIQUE VERSUS STAPLER-ONLY TECHNIQUE BETWEEN AUGUST 2014 AND DECEMBER 2021. IT WAS NOTED THAT DURING THE HYBRID TECHNIQUE, THE MAIN BILE DUCTS AND VESSELS IN THE LIVER HILUM AND THE MAIN HEPATIC VEINS WERE DIVIDED USING A MANUAL STAPLER WITH AN ANGULATING 45MM RELOAD. THERE WERE 492 PATIENTS INCLUDED IN THE STUDY, THE TOTAL (N) PATIENTS FOR THE HYBRID PROCEDURE WERE 152 WITH EIGHT (8) DEATHS RESULTING IN THE 5.3% MORTALITY. OUTCOMES AND COST OF MAJOR LIVER RESECTION USING COMBINED LIGASURE AND STAPLER: A PROPENSITY SCORE MATCHING STUDY, SEPEHR ABBASI DEZFOULI, ARASH DOOGHAIE MOGHADAM, NASTARAN SABETKISH, ELIAS KHAJEH, ALI RAMOUZ, ALI MAJLESARA, MARKUS MIETH, DE HUA CHANG, MOHAMMAD GOLRIZ, AND ARIANEB MEHRABI, 2025, JOURNAL OF CLINICAL MEDICINE. Manufacturer Narrative: D10 CONCOMITANT PRODUCTS: UNKNOWN EGIA SU, UNKNOWN ENDO GIA SULU (LOT#:UNKNOWN), UNEGIATRI, UNKNOWN EGIA TRI STAPLE (LOT#:UNKNOWN), UNKNOWN EGIA SU, UNKNOWN ENDO GIA SULU (LOT#:UNKNOWN), UNKNOWN LIGASUR, UNKNOWN LIGASURE INSTRUMENT (LOT#:UNKNOWN), UNKNOWN LIGASUR, UNKNOWN LIGASURE INSTRUMENT (LOT#:UNKNOWN) SEPEHR ABBASI DEZFOULI, ARASH DOOGHAIE MOGHADAM, NASTARAN SABETKISH, ELIAS KHAJEH, ALI RAMOUZ, ALI MAJLESARA, MARKUS MIETH, DE HUA CHANG, MOHAMMAD GOLRIZ, AND ARIANEB MEHRABI, OUTCOMES AND COST OF MAJOR LIVER RESECTION USING COMBINED LIGASURE AND STAPLER: A PROPENSITY SCORE MATCHING STUDY, 2025, JOURNAL OF CLINICAL MEDICINE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22686146&pc=GDW | 1219930-2025-03407 | 2025-05-30 04:00 | Death | COVIDIEN SURGICAL | 2025-08-01 04:00 | GDW | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  |  | 1 | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY COMPARED THE OUTCOMES OF PATIENTS WHO UNDERWENT LIVER RESECTION UTILIZING A HYBRID TECHNIQUE VERSUS STAPLER-ONLY TECHNIQUE BETWEEN AUGUST 2014 AND DECEMBER 2021. IT WAS NOTED THAT DURING THE STAPLER-ONLY METHOD, THE LIVER WAS TRANSECTED, AND THE HEPATIC VEINS AND PORTAL TRIAD WERE DIVIDED USING A STRAIGHT 60MM LOADING UNIT. THERE WERE 492 PATIENTS INCLUDED IN THE STUDY, THE TOTAL (N) PATIENTS FOR THE STAPLER-ONLY PROCEDURE WERE 340 WITH 12 DEATHS RESULTING IN THE 3.9% MORTALITY. OUTCOMES AND COST OF MAJOR LIVER RESECTION USING COMBINED LIGASURE AND STAPLER: A PROPENSITY SCORE MATCHING STUDY, SEPEHR ABBASI DEZFOULI, ARASH DOOGHAIE MOGHADAM, NASTARAN SABETKISH, ELIAS KHAJEH, ALI RAMOUZ, ALI MAJLESARA, MARKUS MIETH, DE HUA CHANG, MOHAMMAD GOLRIZ, AND ARIANEB MEHRABI, 2025, JOURNAL OF CLINICAL MEDICINE. Manufacturer Narrative: D10 CONCOMITANT PRODUCTS: UNEGIATRI, UNKNOWN EGIA TRI STAPLE (LOT#:UNKNOWN), UNEGIATRI, UNKNOWN EGIA TRI STAPLE (LOT#:UNKNOWN), UNKNOWN EGIA SU, UNKNOWN ENDO GIA SULU (LOT#:UNKNOWN), UNKNOWN LIGASUR, UNKNOWN LIGASURE INSTRUMENT (LOT#:UNKNOWN), UNKNOWN LIGASUR, UNKNOWN LIGASURE INSTRUMENT (LOT#:UNKNOWN) SEPEHR ABBASI DEZFOULI, ARASH DOOGHAIE MOGHADAM, NASTARAN SABETKISH, ELIAS KHAJEH, ALI RAMOUZ, ALI MAJLESARA, MARKUS MIETH, DE HUA CHANG, MOHAMMAD GOLRIZ, AND ARIANEB MEHRABI, OUTCOMES AND COST OF MAJOR LIVER RESECTION USING COMBINED LIGASURE AND STAPLER: A PROPENSITY SCORE MATCHING STUDY, 2025, JOURNAL OF CLINICAL MEDICINE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22686149&pc=OZD | 1220648-2025-30927 | 2025-07-19 04:00 | Death | ABIOMED, INC. | 2025-08-01 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Ischemia | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WAS ON AN IMPELLA CP SUPPORT DUE TO METHAMPHETAMINE TOXICITY AND WAS FOUND TO BE IN CARDIOGENIC SHOCK. THE PUMP WAS PLACED AND DISTAL PERFUSION WAS PLACED DUE TO LIMB ISCHEMIA. THE SUPPORT WAS ONGOING WHEN AFTER FIVE AND A HALF HOURS, THE PATIENT HAD CARDIAC ARREST AND CARDIOPULMONARY RESUSCITATION WAS INITIATED. THE FAMILY MADE THE DECISION TO WITHDRAW CARE AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22686271&pc=OZD | 1220648-2025-30928 | 2025-07-20 04:00 | Death | ABIOMED, INC. | 2025-08-01 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Ischemia | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT AFTER IMPELLA CP WAS PLACED THERE WAS OOZING AT THE IMPELLA ACCESS SITE AND DISTAL PERFUSION WAS COMPROMISED. DISTAL REPERFUSION SHEATH WAS PLACED IN THE LSFA AND RIGHT FA SHEATH WAS PLACED FOR EXTERNAL FEM/FEM BYPASS. FIGURE EIGHT SUTURE PLACED ON IMPELLA ACCESS SITE AND PRESSURE DRESSING APPLIED FOR THE BLEEDING. IT WAS FURTHER REPORTED THAT THE FAMILY DECIDED TO WITHDRAW CARE AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE HAS NOT BEEN RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMART ASSIST SYSTEM STATES THE FOLLOWING: SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT: USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22686287&pc=FTL | 9615742-2025-00898 | 2025-05-20 04:00 | Death | SOFRADIM PRODUCTION SAS | 2025-08-01 04:00 | FTL | UNKNOWN PROGRIP MESH PRODUCT | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  |  | 1 | Event Description: A RETROSPECTIVE SINGLE-CENTER STUDY EVALUATED 482 CASES OF OPEN ABDOMINAL WALL HERNIA REPAIR USING THE RETROMUSCULAR SUBLAY TECHNIQUE BETWEEN 2009 AND 2019. OF THESE, 161 PATIENTS RECEIVED THE PROGRIP SELF-GRIPPING MESH, WHILE 321 RECEIVED A COMPARATOR MESH. THE STUDY ASSESSED POSTOPERATIVE INFLAMMATORY RESPONSE AND LONG-TERM OUTCOMES. NO FIXATION WAS USED IN THE PROGRIP GROUP DUE TO ITS SELF-ADHERENT DESIGN. TWO PATIENTS (1.2%) IN THE PROGRIP GROUP DIED WITHIN 30 DAYS OF SURGERY AND CAUSES WERE NOT SPECIFIED. ACUTE INFLAMMATION TRIGGERED BY TWO LIGHTWEIGHT HERNIA MESHES: A COMPARATIVE IN VITRO AND RETROSPECTIVE COHORT STUDY, MARTIN REICHERT, BERNADET MASSAMBO, ANCA-LAURA AMATI, VERONIKA GRAU1, KATRIN RICHTER, ANDREAS HECKER, HERNIA (2025), HTTPS://DOI.ORG/10.1007/S1002 9-025-03391-Y. Manufacturer Narrative: D10 CONCOMITANT PRODUCTS: UNKNOWN PROGRIP, UNKNOWN PROGRIP MESH PRODUCT (LOT#: UNKNOWN) MARTIN REICHERT. "ACUTE INFLAMMATION TRIGGERED BY TWO LIGHTWEIGHT HERNIA MESHES: A COMPARATIVE IN VITRO AND RETROSPECTIVE COHORT STUDY". 2025, HTTPS://DOI.ORG/10.1007/S10029-025-03391-Y. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22686371&pc=FRN | 3012307300-2025-09158 | 2025-02-13 05:00 | Death | ICU MEDICAL, INC. | 2025-08-01 04:00 | FRN | CADD LEGACY PLUS PUMP | Unclear Information | Cancer | K982836 |  | 1 | Event Description: IT WAS REPORTED THAT IT WAS UNCLEAR WHETHER THE PATIENT WAS STILL USING THE PUMP AT THE TIME OF HOSPITAL ADMISSION OR IF IT HAD BEEN DISCONNECTED. THE PATIENT LATER ENTERED HOSPICE CARE ON (B)(6) 2025 AND PASSED AWAY ON (B)(6) 2025, AT 6:01 AM. THE PATIENT RECEIVED ONLY ONE CYCLE OF CHEMOTHERAPY. THE MEDICATION USED WAS 5FU (5150 MG) AT A RATE OF 2.3 ML/HOUR. THE PATIENT WAS BEING TREATED FOR CANCER. Manufacturer Narrative: INVESTIGATION INCLUDING ROOT CAUSE ANALYSIS IS IN PROGRESS. A SUPPLEMENTAL MDR WILL BE FILED AS NECESSARY IN ACCORDANCE WITH 21 CFR 803.56 WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22686479&pc=LJT | 3006260740-2025-05481 | 2025-07-18 04:00 | Death | C.R. BARD, INC. (BASD) -3006260740 | 2025-08-01 04:00 | LJT | POWERPORT IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | Obstruction of Flow; Migration | Obstruction/Occlusion | K060812 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE LITIGATION PROCESS THAT SOMETIME POST A PORT PLACEMENT, THE CATHETER HAD ALLEGEDLY MIGRATED. IT WAS FURTHER REPORTED THAT OCCLUSION WAS OBSERVED. IT WAS ALSO REPORTED THAT THE PATIENT EXPIRED, AND THE CAUSE OF DEATH WAS NOT PROVIDED. Manufacturer Narrative: H11: DATE OF DEATH FOR THIS PATIENT WAS NOT PROVIDED, DATE OF DEATH UPDATED AS (B)(6) 1900 FOR THE MDR SUBMISSION REQUIREMENT. MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO MEDICAL RECORDS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED CATHETER MIGRATION AND OCCLUSION ISSUE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE IS UNKNOWN. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22686575&pc=OYC | 2032227-2025-224099 | 2023-09-07 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-08-01 04:00 | OYC | PARADIGM REAL-TIME INSULIN INFUSION PUMP MMT-754CMS | Adverse Event Without Identified Device or Use Problem | Insufficient Information | EXEMPT |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2023. THE CAUSE OF DEATH WAS HEART DISEASE. THE PUMP WAS WORN AT THE TIME OF PASSING. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: MMT-754CMS. TROUBLESHOOTING WAS PERFORMED FOR THE DECEASED REPORTING. NO ALLEGATIONS AGAINST THE PRODUCT. THE CUSTOMER DISCONTINUED USING THE DEVICE, MMT-754CMS WAS RETURNED FOR ANALYSIS. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. A COMPLETE ANALYSIS AND TESTING OF THE INSULIN PUMP SHOWED THAT IT WAS FUNCTIONING PROPERLY AND PASSED ALL FUNCTIONAL TESTING. AFTER TESTING IT WAS CONCLUDED THAT THE DEVICE OPERATED WITHIN SPECIFICATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. THE CUSTOMER PASSED AWAY ON (B)(6) 2023. THE PUMP PASSED THE FUNCTIONAL TEST, INCLUDING THE STOP (IDLE) CURRENT MEASUREMENT TEST, RUN CURRENT MEASUREMENT TEST, SELF TEST, OFF NO POWER ALARM TEST, A21 ERROR TEST, REWIND TEST, BASIC OCCLUSION TEST, OCCLUSION TEST, PRIME/A33 TEST, EXCESSIVE NO DELIVERY TEST, DISPLACEMENT TEST AND DAT AT 0.0873 INCHES. THE FOLLOWING WERE NOTED DURING VISUAL INSPECTION: MINOR SCRATCHED DISPLAY WINDOW, CRACKED CASE AT DISPLAY WINDOW CORNER, CRACKED BATTERY TUBE THREADS, SCRATCHED RESERVOIR TUBE WINDOW, MISSING END CAP STICKER AND CRACKED RESERVOIR TUBE LIP. THE TEST P-CAP AND RESERVOIR DOES LOCK IN PLACE IN THE RESERVOIR COMPARTMENT. HISTORY DOWNLOAD WAS SUCCESSFUL USING THDS AND CARELINK UPLOAD WAS SUCCESSFUL. THE PUMP WAS RECEIVED WITHOUT A BATTERY. NO DAMAGE NOTED ON THE ORIGINAL BATTERY CAP. UNABLE TO VERIFY DAILY INSULIN TOTAL, BASAL INSULIN DELIVERED TOTAL AND BOLUS INSULIN DELIVERED TOTAL FOR EVENT DATE 07-SEP-2023 DUE TO NO DATA AVAILABLE IN THE PUMP HISTORY FILE. UNABLE TO PERFORM THE HISTORY REVIEW 1 WEEK PRIOR TO THE EVENT DATE 07-SEP-2023 IN THE PUMP HISTORY FILE DUE TO NO DATA AVAILABLE. PUMP WAS CUT OPEN TO PERFORM VISUAL INSPECTION AND FOUND NO EVIDENCE OF PHYSICAL OR MOISTURE DAMAGE ON THE ELECTRONIC ASSEMBLY OR ON THE MOTOR. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22687044&pc=KDJ | 3030306055-2025-00146 | 2025-07-11 04:00 | Death | VANTIVE US HEALTHCARE LLC | 2025-08-01 04:00 | KDJ | NI | Improper or Incorrect Procedure or Method | Abdominal Pain; Fatigue; Sepsis; Peritonitis | NI |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED A BREACH IN ASEPTIC TECHNIQUE WHICH RESULTED IN PERITONITIS. THE BREACH IN ASEPTIC TECHNIQUE WAS NOT FURTHER DESCRIBED. THE PERITONITIS WAS MANIFESTED BY CLOUDY PD EFFLUENT, ABDOMINAL PAIN, AND WEAKNESS IN THE BODY. IT WAS REPORTED THE PATIENT WAS NOT ADMITTED TO A HOSPITAL. IT WAS REPORTED THAT THE PATIENT WAS IN "DAY CARE" AND "LEFT THE DAY CARE AGAINST MEDICAL ADVISE" ON THE SAME DAY AS DATE OF DIAGNOSIS. ON THE SAME DAY OF DIAGNOSIS, THE PATIENT WAS TREATED WITH VANCOMYCIN INJECTION (1GM, EVERY 4TH DAY, INTRAPERITONEAL) AND CEFTAZIDIME INJECTION (500MG, THREE TIMES A DAY, INTRAPERITONEAL) FOR PERITONITIS. PD THERAPY WAS ONGOING. RETRAINING FOR BREAK IN ASEPTIC TECHNIQUE WAS PERFORMED. IT WAS REPORTED THAT THE PATIENT WAS RECOVERING FROM THE EVENTS. FIVE DAYS AFTER THE DATE OF DIAGNOSIS, THE PATIENT SUBSEQUENTLY PASSED. THE CAUSE OF DEATH WAS REPORTED AS SEPSIS AND OTHER INDICATIONS. IT WAS NOT REPORTED IF AN AUTOPSY WAS PERFORMED. NO ADDITIONAL INFORMATION IS AVAILABLE. Manufacturer Narrative: THIS REPORT IS FOR A BREACH IN ASEPTIC TECHNIQUE WHICH RESULTED IN PERITONITIS. PER VANTIVE LABELING, USERS ARE INSTRUCTED TO USE ASEPTIC TECHNIQUE WHEN PERFORMING PERITONEAL DIALYSIS THERAPY. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22687546&pc=NRY | 3014162263-2025-00055 | 2025-06-25 04:00 | Death | BALT USA | 2025-08-01 04:00 | NRY | RAPTOR ASPIRATION CATHETER | Unintended Movement | Rupture | K234083 |  | 1 | Event Description: IT WAS REPORTED THAT: "THE PHYSICIAN WAS CALLED IN FOR AN M1 OCCLUSION; UPON PERFORMING INITIAL RUN, IDENTIFIED THE M 1 WAS OPEN BUT THERE WAS CLOT IN THE M2. SINCE SHE ALREADY OPENED RAPTOR 074 AND LARGE CARRIER, SHE DECIDED TO ADVANCE TO THE M2. SHE WAS USING MEGA BALLAST AS WELL. SHE DECIDED TO PARK THE RAPTOR AT M1 AND ADVANCED CARRIER TO M2 THEN UNSHEATHED A 3MM SOLITAIR. SHE FELT THE M2 WAS FAIRLY LARGE AND BEGAN TO PULL GENTLY BACK ON THE CARRIER WITH THE SOLITAIR AND FELT RESISTANCE. SHE DID A RUN AND FOUND THE ORIGIN OF M2 RUPTURED. SHE THINKS THE RUPTURE OCCURRED FOR 1 OF 2 "POSSIBLE" REASONS: 1) WHEN ALL THE OTHER CATHETERS WOULDN'T HAVE GONE UP SO EASILY DISTALLY, SHE BELIEVES THE CARRIER MAY HAVE STRETCHED THE VESSEL. 2) SHE MAY HAVE STRAIGHTENED OUT THE SYSTEM WHEN TRYING TO RETRIEVE THE STENT TRIEVER/CARRIER AND THE RAPTOR JUMPED FORWARD CAUSING THE RUPTURE. SHE HAD THE MEGA BALLAST AT THE DISTAL PETROUS AND IT WORKED WELL. " ADDITIONAL INFORMATION RECEIVED BY THE ISSUER (B2025: "1. HOW WOULD THE TORTUOSITY OF THE ANATOMY BE DESCRIBED IN THIS PROCEDURE? NORMAL TORTUOSITY 2. WHAT IS THE CURRENT PATIENT STATUS? PATIENT DIED DUE TO RUPTURING THE ARTERY 3. TO CONFIRM, ON THE 1ST POSSIBILITY, THE PHYSICIAN IS POSTULATING THAT SHE TRIED TO ADVANCE THE CARRIER INTO THE M2 AND THIS IS WHAT CAUSED THE RUPTURE? (CARRIER LARGE WAS TOO BIG FOR THE M2) - YES AND THAT IT STRETCHED IT OUT OR POSSIBLE TOOK RAPTOR 074 TOO FAR 4. TO CONFIRM, ON THE 2ND POSSIBILITY, THE PHYSICIAN IS POSTULATING THAT THE RAPTOR WAS TOO LARGE FOR THE VESSEL WHICH WHEN "JUMPING FORWARD" CAUSED THE VESSEL RUPTURE? YES SHE SAID WHEN SHE PULLED THE CARRIER WITH THE SOLITAIRE, SHE FELT A TUG AND IT MAY HAVE STRAIGHTENED THE ARTERY OUT AND CAUSED RAPTOR TO JUMP. 5. WOULD YOU BE ABLE TO PROVIDE THE PART NO. AND LOT NO. FOR THE RAPTOR AND CARRIER UNITS? YES - I'M WORKING ON GETTING THE LOTS FOR BOTH. I AM GOING TO BE UP THERE ON THE (B)(6) AND WILL GET IT." UPDATE 24JUL 2025 - ADDITIONAL INFORMATION RECEIVED FROM THE ISSUER: "THE PATIENT DIDN'T HAVE ANY OTHER CO-MORBIDITIES. THE PHYSICIAN WAS AWARE OF THE WARNING." Manufacturer Narrative: CLINICAL INVESTIGATION: IT IS UNKNOWN IF A TEMPORAL RELATIONSHIP EXISTS BETWEEN CCPD THERAPY UTILIZING A LIBERTY SELECT CYCLER AND THE SERIOUS ADVERSE EVENTS OF HYPOTENSION, CARDIAC ARREST AND DEATH, AS THE TIMELINE OF EVENTS IS UNKNOWN. THE PDRN REPORTED THE PATIENT SUFFERED A CARDIAC EVENT, HOWEVER THE ETIOLOGY OF THE SERIOUS ADVERSE EVENTS IS UNKNOWN; THEREFORE, CAUSALITY CANNOT BE FIRMLY ESTABLISHED. THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. CARDIOVASCULAR DISEASE (INCLUDING SUDDEN CARDIAC DEATH) ACCOUNTS FOR THE MOST DEATHS AMONG THE ESRD POPULATION. ADDITIONALLY, ADULTS WITH ESRD HAVE MORTALITY RATES UP TO 30-FOLD HIGHER THAN THE GENERAL POPULATION. BASED ON THE INFORMATION AVAILABLE, THE LIBERTY SELECT CYCLER CANNOT BE DISASSOCIATED FROM PLAYING A POSSIBLE ROLE IN THE DEVELOPMENT OF THE SERIOUS ADVERSE EVENTS. THERE IS CURRENTLY NO ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) DEFICIENCY OR MALFUNCTION OCCURRED. HOWEVER, GIVEN THE TEMPORAL RELATIONSHIP, UNKNOWN CAUSALITY, NO CAUSE OF DEATH, AND NO MANUFACTURER EVALUATION OF THE PATIENT¿S DEVICE, THIS CLINICAL INVESTIGATION CANNOT EXCLUDE THE PATIENT¿S LIBERTY SELECT CYCLER FROM PLAYING A POSSIBLE ROLE IN THE SERIOUS ADVERSE EVENTS. THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22687686&pc=PYX | 1220648-2025-30933 | 2025-07-04 04:00 | Death | ABIOMED, INC. | 2025-08-01 04:00 | PYX | IMPELLA RP FLEX | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P170011 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS REFERRED TO SURGERY FOR CORONARY ARTERY BYPASS GRAFT. HOWEVER, THE PATIENT WAS SHOWING SIGNS OF BEING SEPTIC WITH DIAGNOSIS OF UROSEPSIS AND BACTEREMIA. THE SURGEON RECOMMENDED ANTIBIOTICS AND FOR THE PATIENT TO FOLLOW UP IN 6 WEEKS. THE PATIENT WAS DISCHARGED TO A REHAB FACILITY. AT THE REHAB FACILITY, THE PATIENT STARTED TO EXPERIENCE CHEST TIGHTNESS AND SHORTNESS OF BREATH. THE PATIENT WAS BROUGHT BACK TO HOSPITAL AND ADMITTED TO THE INTENSIVE CARE UNIT. THE PATIENT'S STATUS CONTINUED TO DECLINE. THE PATIENT REQUIRED INTUBATION AND DEVELOPED BI-VENTRICULAR FAILURE. THE PATIENT WAS NOTED TO BE MAXED OUT ON MULTIPLE PRESSORS. CARE TEAM SPOKE WITH FAMILY ON PLAN OF CARE. THE FAMILY ELECTED TO TAKE TO CARDIAC CATHERIZATION AND PLACE BIPELLA SUPPORT. DURING SUPPORT IT WAS NOTED THAT THE MATTRESS SUTURE BROKE. THE PHYSICIAN ATTEMPTED TO PLACE MULTIPLE NEW MATTRESS SUTURES, BUT WAS UNSUCCESSFUL. LARGE AMOUNTS OF BLEEDING WAS NOTED AT THE SHEATH SITE. MANUAL PRESSURE WAS HELD. VASCULAR SURGEON WAS ABLE TO SUTURE THE SITE. HEMOSTASIS WAS ACHIEVED.Event Description: THE MEDICAL SAFETY OFFICER REVIEWED AND DETERMINED GIVEN THE MASSIVE BLEEDING IN THE GROIN REQUIRING INTERVENTION THERE IS NOT ENOUGH EVIDENCE TO DISSOCIATE THE IMPELLA RP FLEX DEVICE FROM THE PATIENT OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE HAS NOT BEEN RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA IFU: ISECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. THE ROOT CAUSE OF THE ACCESS SITE BLEEDING ISSUE WAS MOST LIKELY USE RELATED, SINCE VASCULAR SURGEON CALLED AND ABLE TO SUTURE SITE TO ACHIEVE HEMOSTASIS. THE MEDICAL SAFETY OFFICER (MSO) REVIEWED AND DETERMINED GIVEN THE MASSIVE BLEEDING IN THE GROIN REQUIRING INTERVENTION THERE IS NOT ENOUGH EVIDENCE TO DISSOCIATE THE IMPELLA RP FLEX DEVICE FROM THE PATIENT OUTCOME. THE FOLLOWING FIELDS WERE CORRECTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-30933 BASED ON MSO REVIEW. B2 DEATH AND DATE OF DEATH. B5 PATIENT OUTCOME. H1 TYPE OF REPORT. H6 HEALTH EFFECT - IMPACT CODE 1802.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22687767&pc=DXO | 2124215-2025-52970 | 2019-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-08-01 04:00 | DXO | COMET | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: BOSTON SCIENTIFIC PERFORMED A RETROSPECTIVE DATA ANALYSIS ON COMET USING THE PINC AI HEALTHCARE DATABASE FROM PREMIER. PATIENTS ARE IDENTIFIED THROUGH USE OF COMET AS IDENTIFIED IN CHARGE MASTER BILLING. ADVERSE EVENTS WERE IDENTIFIED THROUGH THE RESPECTIVE CPT/ICD-10 CODING. THE PROCEDURES WERE PERFORMED FROM JANUARY 2019 THROUGH SEPTEMBER 2024. THESE ANALYSES ARE BEING PERFORMED AS A SPECIFIC ACTIVITY TO ASSESS SAFETY AND PERFORMANCE RATES ASSOCIATED WITH THE SUBJECT DEVICE. COMET OUTCOMES WERE ASSESSED AGAINST OTHER SIMILAR/BENCHMARK DEVICES. RATES OF THE ADVERSE EVENTS INCLUDING DEATH, MYOCARDIAL INFARCTION, PERICARDIAL EFFUSION, CARDIAC TAMPONADE, ACUTE RENAL FAILURE, CONTRAST ALLERGY, BLEEDING, EMBOLISM AND STROKE ARE REPORTED. A TOTAL OF 33342 PATIENTS UNDERWENT A PROCEDURE USING COMET. THE FOLLOWING IS A SUMMARY OF THOSE RESULTS OVER 5 YEARS (JANUARY 2019 - SEPTEMBER 2024): MORTALITY RATES COMET CASES: 112 OUT OF 33342 PATIENTS RESULTING IN A RATE OF 0.34%, COMPARED TO THE COMPETITOR RATE OF 0.29%-0.44%. MORTALITY RATES FOR COMET CLAIMS FALL WITHIN THE RANGE OF THE SIMILAR/BENCHMARK DEVICES. RATES FOR MORTALITY ARE COMPARABLE TO THOSE OF THE SIMILAR/BENCHMARK DEVICES AND ARE THEREFORE ACCEPTABLE. Manufacturer Narrative: B2 DATE OF DEATH AND B3 DATE OF EVENT: DATA WAS PROVIDED FOR EVENTS RECORDED 01JAN2019-30SEP2024. 01JAN2019 SELECTED AS THE EARLIEST POSSIBLE DATE OF EVENT AND DATE OF DEATH. DATA OBTAINED FOR THIS STUDY IS DE-IDENTIFIED BEFORE BEING ACCESSED BY BOSTON SCIENTIFIC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO OUR ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE STUDY DATA DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC AND NO FURTHER INFORMATION IS AVAILABLE. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22687823&pc=QJP | 3014162263-2025-00056 | 2025-06-25 04:00 | Death | BALT USA | 2025-08-01 04:00 | QJP | CARRIER DELIVERY CATHETER | Device-Device Incompatibility; Unintended Movement | Rupture | K230609 |  | 1 | Event Description: IT WAS REPORTED THAT: "THE PHYSICIAN WAS CALLED IN FOR AN M1 OCCLUSION; UPON PERFORMING INITIAL RUN, IDENTIFIED THE M 1 WAS OPEN BUT THERE WAS CLOT IN THE M2. SINCE SHE ALREADY OPENED RAPTOR 074 AND LARGE CARRIER, SHE DECIDED TO ADVANCE TO THE M2. SHE WAS USING MEGA BALLAST AS WELL. SHE DECIDED TO PARK THE RAPTOR AT M1 AND ADVANCED CARRIER TO M2 THEN UNSHEATHED A 3MM SOLITAIR. SHE FELT THE M2 WAS FAIRLY LARGE AND BEGAN TO PULL GENTLY BACK ON THE CARRIER WITH THE SOLITAIR AND FELT RESISTANCE. SHE DID A RUN AND FOUND THE ORIGIN OF M2 RUPTURED. SHE THINKS THE RUPTURE OCCURRED FOR 1 OF 2 POSSSIBLE REASONS: 1) WHEN ALL THE OTHER CATHETERS WOULDN'T HAVE GONE UP SO EASILY DISTALLY, SHE BELIEVES THE CARRIER MAY HAVE STRETCHED THE VESSEL. 2) SHE MAY HAVE STRAIGHTENED OUT THE SYSTEM WHEN TRYING TO RETRIEVE THE STENT TRIEVER/CARRIER AND THE RAPTOR JUMPED FORWARD CAUSING THE RUPTURE. SHE HAD THE MEGA BALLAST AT THE DISTAL PETROUS AND IT WORKED WELL. ADDITIONAL INFORMATION RECEIVED BY THE ISSUER 10JUL2025: 1. HOW WOULD THE TORTUOSITY OF THE ANATOMY BE DESCRIBED IN THIS PROCEDURE? NORMAL TORTUOSITY. 2. WHAT IS THE CURRENT PATIENT STATUS? PATIENT DIED DUE TO RUPTURING THE ARTERY 3. TO CONFIRM, ON THE 1ST POSSIBILITY, THE PHYSICIAN IS POSTULATING THAT SHE TRIED TO ADVANCE THE CARRIER INTO THE M2 AND THIS IS WHAT CAUSED THE RUPTURE? (CARRIER LARGE WAS TOO BIG FOR THE M2) - YES AND THAT IT STRETCHED IT OUT OR POSSIBLE TOOK RAPTOR 074 TOO FAR 4. TO CONFIRM, ON THE 2ND POSSIBILITY, THE PHYSICIAN IS POSTULATING THAT THE RAPTOR WAS TOO LARGE FOR THE VESSEL WHICH WHEN "JUMPING FORWARD" CAUSED THE VESSEL RUPTURE? YES SHE SAID WHEN SHE PULLED THE CARRIER WITH THE SOLITAIRE SHE FELT A TUG AND IT MAY HAVE STRAIGHTENED THE ARTERY OUT AND CAUSED RAPTOR TO JUMP. 5. WOULD YOU BE ABLE TO PROVIDE THE PART NO. AND LOT NO. FOR THE RAPTOR AND CARRIER UNITS? YES - I'M WORKING ON GETTING THE LOTS FOR BOTH. I AM GOING TO BE UP THERE ON THE 17TH AND WILL GET IT." UPDATE 24JUL2025 - ADDITIONAL INFORMATION RECEIVED FROM THE ISSUER: "THE PATIENT DIDN'T HAVE ANY OTHER CO-MORBIDITIES. THE PHYSICIAN WAS AWARE OF THE WARNING." Manufacturer Narrative: BALT USA REFERENCE: (B)(4). AN EVALUATION OF THE ACTUAL COMPLAINT SAMPLE COULD NOT BE PERFORMED AS DEVICE WAS UNAVAILABLE TO BE RETURN. BASED ON THE PROVIDED INFORMATION, ROOT CAUSE COULD NOT BE DEFINITIVELY DETERMINED. IT WAS REPORTED THE PHYSICIAN BEGAN TO PULL GENTLY BACK ON THE CARRIER WITH THE SOLITAIR AND FELT RESISTANCE. PER THE DEVICE'S IFU, THE CARRIER DELIVERY CATHETER IS NOT TO BE USED WITH STENTS, RETRIEVERS, OCCLUSION COILS, GLUE, GLUEMIXTURE OR NON-ADHESIVE LIQUID EMBOLIC AGENTS. LACK OF DEVICE RETURN PREVENTED DEEPER EVALUATION OF THE REPORTED ISSUE. THE LOT NUMBER WAS NOT PROVIDED THEREFORE; A REVIEW OF THE LOT HISTORY RECORDS COULD NOT BE PERFORMED. THE SUBMISSION OF THIS REPORT OR RELATED INFORMATION TO THE FDA, AND ITS RELEASE BY FDA, DOES NOT REFLECT A CONCLUSION BY THE PARTY SUBMITTING THIS REPORT OF THE FDA THAT THE REPORT OR RELATED INFORMATION IS AN ADMISSION THAT THE MANUFACTURER, THEIR EMPLOYEES, OR THE DEVICE CAUSED OR CONTRIBUTED TO THE REPORTABLE EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22688014&pc=LYJ | 1644487-2025-10447 | 2025-06-09 04:00 | Death | HOUSTON | 2025-08-01 04:00 | LYJ | PULSE GEN MODEL 1000 | Adverse Event Without Identified Device or Use Problem | Status Epilepticus | P970003 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ABOUT A MONTH AGO FROM STATUS. NO OTHER RELEVANT INFORMATION HAS BEEN RECEIVED TO DATE. Manufacturer Narrative: LIVANOVA USA, INC. SUBMITS THIS REPORT TO COMPLY WITH 21 CFR PART 803, THE MEDICAL DEVICE REPORTING REGULATION, BASED ON INFORMATION THAT LIVANOVA HAS OBTAINED, BUT MAY NOT HAVE BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION, OR A CONCLUSION BY FDA OR ANYONE ELSE, THAT THE DEVICE, LIVANOVA, OR LIVANOVA'S EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS" OR "MALFUNCTIONS". THESE WORDS ARE INCORPORATED INTO THE FDA 3500A MEDWATCH FORM BY THE FDA, AND LIVANOVA OBJECTS TO THEIR USE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22688017&pc=NPT | 2135147-2025-04240 | 2022-04-01 04:00 | Death | ABBOTT MEDICAL | 2025-08-01 04:00 | NPT | NAVITOR TRANSCATHETER AORTIC VALVE | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P190023 |  | 1 | Event Description: THE ARTICLE, "INTRA- VERSUS SUPRA-ANNULAR SELF-EXPANDING TRANSCATHETER HEART VALVES IN SMALL AORTIC ANNULI", WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE, MULTICENTER STUDY TO COMPARE EARLY CLINICAL OUTCOMES, INCLUDING ECHOCARDIOGRAPHIC PARAMETERS, BETWEEN THE LATEST GENERATION OF SUPRA-ANNULAR SELF-EXPANDING VALVES (SA-SEVS) AND INTRA-ANNULAR (IA)-SEVS AFTER TRANSCATHETER AORTIC VALVE IMPLANTATION (TAVI) IN PATIENTS WITH SMALL AORTIC ANNULUS (SAA). DEVICES INCLUDED IN THE STUDY WERE NAVITOR AND EVOLUT FX. THE ARTICLE CONCLUDED THAT LATEST-GENERATION IA-SEV AND SA-SEV DEMONSTRATED SIMILAR CLINICAL RESULTS EXCEPT FOR A FEW ECHOCARDIOGRAPHIC FINDINGS AFTER TAVI IN PATIENTS WITH SAA. THE PRIMARY AND CORRESPONDING AUTHOR WAS MASANORI YAMAMOTO, DEPARTMENT OF CARDIOLOGY, TOYOHASHI HEART CENTER, 21-1 GOBUTGORI, OYAMACHO, TOYOHASHI, AICHI, 441-8530, JAPAN, WITH CORRESPONDING E-MAIL: MASA-NORI@NMS.AC.JP]. THE TIME FRAME OF THE STUDY WAS FROM APRIL 2022 TO MARCH 2024. A TOTAL OF 919 PATIENTS WERE INCLUDED IN THE STUDY, OF WHICH 518 (56.4%) RECEIVED AN ABBOTT DEVICE. IN THE IA-SEV GROUP, THE AVERAGE AGE WAS 86.0 YEARS AND THE AVERAGE GENDER WAS FEMALE. COMORBIDITIES INCLUDED DIABETES, HYPERTENSION, CHRONIC OBSTRUCTIVE PULMONARY DISEASE, LUNG DISEASE, PRIOR STROKE, PRIOR CORONARY ARTERY BYPASS GRAFT, PRIOR PERCUTANEOUS CORONARY INTERVENTION, PRIOR MYOCARDIAL INFARCTION, ATRIAL FIBRILLATION/FLUTTER, HISTORY OF HEART BLOCK, CORONARY ARTERY DISEASE, PRIOR PACEMAKER, BICUSPID VALVE. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF INTRA- VERSUS SUPRA-ANNULAR SELF-EXPANDING TRANSCATHETER HEART VALVES IN SMALL AORTIC ANNULI WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING DIABETES, HYPERTENSION, CHRONIC OBSTRUCTIVE PULMONARY DISEASE, LUNG DISEASE, PRIOR STROKE, PRIOR CORONARY ARTERY BYPASS GRAFT, PRIOR PERCUTANEOUS CORONARY INTERVENTION, PRIOR MYOCARDIAL INFARCTION, ATRIAL FIBRILLATION/FLUTTER, HISTORY OF HEART , HEART BLOCK, CORONARY ARTERY DISEASE, PRIOR PACEMAKER, BICUSPID VALVE. SOME OF THE COMPLICATIONS REPORTED WERE STROKE, OBSTRUCTION, HEMORRHAGE, RENAL FAILURE, DEATH; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. PLEASE NOTE THAT PER THE INSTRUCTION FOR USE, ARTMT600163776 REVISION D, ¿THE NAVITOR¿ VALVE IS INDICATED FOR PATIENTS WITH SYMPTOMATIC SEVERE NATIVE AORTIC STENOSIS WHO ARE CONSIDERED HIGH OR EXTREME RISK FOR SURGICAL AORTIC VALVE REPLACEMENT (AVR)." SINCE THIS WAS REPORTED THROUGH LITERATURE AND THERE IS NO EVIDENCE THE OFF-LABEL USE CAUSED OR CONTRIBUTED TO THE REPORTED ISSUES, A LETTER WILL NOT BE SENT. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO LABELING DESIGN OR MANUFACTURING OF THE DEVICE. B2: DATE OF DEATH IS ESTIMATED. B3: DATE OF EVENT IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: INTRA- VERSUS SUPRA-ANNULAR SELF-EXPANDING TRANSCATHETER HEART VALVES IN SMALL AORTIC ANNULI. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF INTRA- VERSUS SUPRA-ANNULAR SELF-EXPANDING TRANSCATHETER HEART VALVES IN SMALL AORTIC ANNULI WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING DIABETES, HYPERTENSION, CHRONIC OBSTRUCTIVE PULMONARY DISEASE, LUNG DISEASE, PRIOR STROKE, PRIOR CORONARY ARTERY BYPASS GRAFT, PRIOR PERCUTANEOUS CORONARY INTERVENTION, PRIOR MYOCARDIAL INFARCTION, ATRIAL FIBRILLATION/FLUTTER, HISTORY OF HEART HEART BLOCK, CORONARY ARTERY DISEASE, PRIOR PACEMAKER, BICUSPID VALVE. SOME OF THE COMPLICATIONS REPORTED WERE STROKE, OBSTRUCTION, HEMORRHAGE, RENAL FAILURE, DEATH; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. PLEASE NOTE THAT PER THE INSTRUCTION FOR USE, ARTMT600163776 REVISION D, ¿THE NAVITOR¿ VALVE IS INDICATED FOR PATIENTS WITH SYMPTOMATIC SEVERE NATIVE AORTIC STENOSIS WHO ARE CONSIDERED HIGH OR EXTREME RISK FOR SURGICAL AORTIC VALVE REPLACEMENT (AVR)." SINCE THIS WAS REPORTED THROUGH LITERATURE AND THERE IS NO EVIDENCE THE OFF-LABEL USE CAUSED OR CONTRIBUTED TO THE REPORTED ISSUES, A LETTER WILL NOT BE SENT. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO LABELING DESIGN OR MANUFACTURING OF THE DEVICE. B2: DATE OF DEATH IS ESTIMATED. B3: DATE OF EVENT IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: INTRA- VERSUS SUPRA-ANNULAR SELF-EXPANDING TRANSCATHETER HEART VALVES IN SMALL AORTIC ANNULI.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22688509&pc=DSQ | 2916596-2025-05196 | 2024-02-21 05:00 | Death | THORATEC CORPORATION | 2025-08-01 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22688596&pc=NPT | 2015691-2025-06240 | 2025-07-10 04:00 | Death | EDWARDS LIFESCIENCES | 2025-08-01 04:00 | NPT | COMMANDER DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Discomfort; Cardiac Perforation | P140031 |  | 1 | Event Description: EDWARDS RECEIVED NOTIFICATION FROM OUR AFFILIATE IN MEXICO. AS REPORTED, IT WAS A CASE OF AN IMPLANT OF A 23 MM SAPIEN 3 ULTRA IN AORTIC POSITION BY TRANSFEMORAL APPROACH. DURING PROCEDURE, WHEN THE NATIVE VALVE WAS CROSSED, WITHOUT DIFFICULTIES, THE PATIENT EXPERIENCED SOME DISCOMFORT AND A DROP IN PRESSURE. AN ECHOCARDIOGRAM WAS PERFORMED AND A VENTRICULAR TEAR WAS FOUND. IT WAS PERFORMED A PERICARDIOCENTESIS AND TRANSFERRED TO OPERATING ROOM. UNFORTUNATELY, THE PATIENT PASSED AWAY. THE PATIENT WAS NOTED TO BE FRAGILE. THE DOCTORS ATTRIBUTED THE EVENT TO THE COMMANDER SYSTEM AND THE VENTRICULAR GUIDEWIRE. Manufacturer Narrative: SUPPLEMENTAL REPORT SUBMITTED TO INCLUDE ADDITIONAL INFORMATION RECEIVED THROUGH FOLLOW-UP. ADDED H6: HEALTH EFFECT - IMPACT CODE. PER ADDITIONAL INFORMATION RECEIVED, IT WAS REPORTED THAT THE PATIENT DID UNDERGO A STERNOTOMY FOR OPEN HEART SURGERY. THE DEVICES ARE NOT AVAILABLE FOR EVALUATION. NO ADDITIONAL INFORMATION RECEIVED. THE CONCLUSION REMAINS THE SAME. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. THERE ARE SEVERAL POTENTIAL ETIOLOGIES FOR VENTRICULAR PERFORATION DURING A TVR PROCEDURE, INCLUDING PERFORATION BY THE GUIDEWIRE, THE DELIVERY SYSTEM, OR THE TRANSVENOUS PACER (TVP) LEAD. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVE (THV). TRAINING INCLUDES PROPER GUIDEWIRE POSITIONING, FIXATION OF THE TVP TO PREVENT VENTRICLE PERFORATION, AND CAREFUL MANIPULATION OF DEVICES. PER THE PROCEDURE DIDACTIC, PATIENTS WITH SMALL VENTRICLES ARE AT PARTICULARLY HIGH RISK FOR VENTRICULAR PERFORATION. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST THAT PROCEDURE RELATED FACTORS (COMMANDER SYSTEM AND THE VENTRICULAR GUIDEWIRE) CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22689112&pc=DSQ | 2916596-2025-05194 | 2025-04-10 04:00 | Death | THORATEC CORPORATION | 2025-08-01 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CUSTOMER. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, REV. D IS CURRENTLY AVAILABLE. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22689182&pc=DSQ | 2916596-2025-04870 | 2023-12-21 05:00 | Death | THORATEC CORPORATION | 2025-08-01 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY IN THEIR SLEEP FROM NATURAL CAUSES.Event Description: IT WAS REPORTED THAT THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED. THE DEVICE WOULD NOT BE RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, REV. D IS CURRENTLY AVAILABLE. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22689322&pc=NPT | 9617601-2025-00776 | 2025-04-01 04:00 | Death | MEDTRONIC MEXICO S. DE R.L. DE CV | 2025-08-01 04:00 | NPT | EVOLUT FX PLUS VALVE | Insufficient Information; Patient Device Interaction Problem | Sepsis; Post Traumatic Wound Infection | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WITH A DIAGNOSIS OF AORTIC VALVE STENOSIS AND COMORBIDITIES INCLUDING DIABETES MELLITUS (ON INSULIN), DYSLIPIDEMIA, HYPERTENSION, RENAL FAILURE (NOT ON DIALYSIS), SEVERE CHRONIC OBSTRUCTIVE PULMONARY DISEASE, AND NEUROLOGIC IMPAIRMENT UNDERWENT AN INITIAL IMPLANT PROCEDURE WITH THE EVFXPLUS-26 AORTIC VALVE. FOLLOWING THE OPERATION, THE PATIENT DEVELOPED AN INFECTION AT THE SURGICAL ACCESS SITE, WHICH PROGRESSED TO SEPSIS, PROBABLY CAUSED BY A HOSPITAL-ACQUIRED STRAIN OF ANTIBIOTIC-RESISTANT BACTERIA. THE PATIENT SUBSEQUENTLY DIED DUE TO A NON-CARDIOVASCULAR CAUSE. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22689444&pc=MFE | 2029214-2025-01741 | 2024-11-14 05:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-08-01 04:00 | MFE | ONYX | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage; Paresis; Coma |  |  | 1 | Event Description: YU, J. (2024). DELAYED VENOUS HEMORRHAGE AFTER ENDOVASCULAR TREATMENT FOR A PETROUS RIDGE DURAL ARTERIOVENOUS FISTULA. RADIOLOGY CASE REPORTS, 20(1), 781¿790. HTTPS://DOI.ORG/10.1016/J.RADCR.2024.10.112 MEDTRONIC REVIEW OF THE LITERATURE ARTICLE FOUND A CASE REPORT OF A 59-YEAR-OLD MALE PATIENT WHO UNDERWENT AN ENDOVASCULAR TREATMENT (EVT) PROCEDURE WITH ONYX EMBOLIZATION TO TREAT A PETROUS RIDGE DURAL ARTERIOVENOUS FISTULA (DAVF). IT WAS NOTED IN THE ARTICLE THAT DELAYED HEMORRHAGE IS A KNOWN COMPLICATION FOR THIS THERAPY THAT HAS FATAL OUTCOME, AND THAT WAS THE WHAT OCCURRED IN THE CASE REPORTED IN THE ARTICLE. IT WAS REPORTED THAT THE DAVF A OBLITERATED IN THE EVT PROCEDURE. 2 MARATHON MICROCATHETERS WERE USED IN THE PROCEDURE FOR DELIVERY OF ONYX-18. NO DEVICE MALFUNCTION WAS REPORTED AND THE PATIENT AWOKE NORMALLY AFTER SURGERY WITH NO NEUROLOGICAL DEFICIT. HOWEVER, 20 HOURS AFTER EVT, THE PATIENT EXPERIENCED LEFT SIDE HEMIPLEGIA. COMPUTED TOMOGRAPHY (CT) DID NOT SHOW ANY HEMORRHAGE AND ISCHEMIA WAS INITIALLY CONSIDERED. A NASOGASTRIC FEEDING TUBE WAS PLACED TO MAINTAIN AN ADEQUATE FLUID VOLUME FOR GOOD CIRCULATION. HOWEVER, THE PATIENT¿S CONDITION DID NOT IMPROVE. 30 HOURS POST EVT, THE PATIENT FELL INTO A COMA AND, AT THAT TIME, CT REVEALED HEMORRHAGE OF THE BRAINSTEM AND CEREBELLUM INTO THE VENTRICLE SYSTEM. AFTER 10 HOURS OF CONSERVATIVE TREATMENT, THE PATIENT DIED. THE COMPLICATIONS WERE CONSIDERED KNOWN COMPLICATIONS FOR THE THERAPY/PROCEDURE. Manufacturer Narrative: B2. DATE OF PATIENT DEATH WAS NOT REPORTED IN THE ARTICLE. B3. EVENT DATE IS THE DATE OF ARTICLE ELECTRONIC PUBLICATION. G4. PMA REFERENCE # IS UNKNOWN AS THE EVENT OCCURRED IN ANOTHER COUNTRY AND THE ONYX MODEL NUMBER WAS NOT REPORTED IN THE ARTICLE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22689572&pc=KRA | 2029214-2025-01742 | 2024-11-14 05:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-08-01 04:00 | KRA | MARATHON | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage; Paresis; Coma | K093750 |  | 1 | Event Description: YU, J. (2024). DELAYED VENOUS HEMORRHAGE AFTER ENDOVASCULAR TREATMENT FOR A PETROUS RIDGE DURAL ARTERIOVENOUS FISTULA. RADIOLOGY CASE REPORTS, 20(1), 781¿790. HTTPS://DOI.ORG/10.1016/J.RADCR.2024.10.112 MEDTRONIC REVIEW OF THE LITERATURE ARTICLE FOUND A CASE REPORT OF A 59-YEAR-OLD MALE PATIENT WHO UNDERWENT AN ENDOVASCULAR TREATMENT (EVT) PROCEDURE WITH ONYX EMBOLIZATION TO TREAT A PETROUS RIDGE DURAL ARTERIOVENOUS FISTULA (DAVF). IT WAS NOTED IN THE ARTICLE THAT DELAYED HEMORRHAGE IS A KNOWN COMPLICATION FOR THIS THERAPY THAT HAS FATAL OUTCOME, AND THAT WAS THE WHAT OCCURRED IN THE CASE REPORTED IN THE ARTICLE. IT WAS REPORTED THAT THE DAVF AN OBLITERATED IN THE EVT PROCEDURE. 2 MARATHON MICROCATHETERS WERE USED IN THE PROCEDURE FOR DELIVERY OF ONYX-18. NO DEVICE MALFUNCTION WAS REPORTED AND THE PATIENT AWOKE NORMALLY AFTER SURGERY WITH NO NEUROLOGICAL DEFICIT. HOWEVER, 20 HOURS AFTER EVT, THE PATIENT EXPERIENCED LEFT SIDE HEMIPLEGIA. COMPUTED TOMOGRAPHY (CT) DID NOT SHOW ANY HEMORRHAGE AND ISCHEMIA WAS INITIALLY CONSIDERED. A NASOGASTRIC FEEDING TUBE WAS PLACED TO MAINTAIN AN ADEQUATE FLUID VOLUME FOR GOOD CIRCULATION. HOWEVER, THE PATIENT¿S CONDITION DID NOT IMPROVE. 30 HOURS POST EVT, THE PATIENT FELL INTO A COMA AND, AT THAT TIME, CT REVEALED HEMORRHAGE OF THE BRAINSTEM AND CEREBELLUM INTO THE VENTRICLE SYSTEM. AFTER 10 HOURS OF CONSERVATIVE TREATMENT, THE PATIENT DIED. THE COMPLICATIONS WERE CONSIDERED KNOWN COMPLICATIONS FOR THE THERAPY/PROCEDURE. Manufacturer Narrative: B2. DATE OF PATIENT DEATH WAS NOT REPORTED IN THE ARTICLE. B3. EVENT DATE IS THE DATE OF ARTICLE ELECTRONIC PUBLICATION. G4. PMA REFERENCE # IS UNKNOWN AS THE EVENT OCCURRED IN ANOTHER COUNTRY AND THE ONYX MODEL NUMBER WAS NOT REPORTED IN THE ARTICLE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22689717&pc=OBJ | 3016516752-2025-00007 | 2025-07-09 04:00 | Death | NUVERA MEDICAL | 2025-08-01 04:00 | OBJ | NUVISION ICE CATHETER, 10F | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Cardiac Perforation | K201775 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A WATCHMAN LEFT ATRIAL APPENDAGE CLOSURE WHILE USING A NUVISION ICE CATHETER, THE PATIENT EXPERIENCED PRESSURE DROP, PERICARDIAL EFFUSION, AND CPR WAS PERFORMED--AFTER WHICH, THE PATIENT PASSED AWAY. IT WAS REPORTED THAT DURING A WATCHMAN LEFT ATRIAL APPENDAGE CLOSURE PROCEDURE, THE PHYSICIAN HAD DIFFICULTY CROSSING THE SEPTUM WITH THE NUVISION CATHETER AND AFTER APPROXIMATELY 40 MINUTES, THE PATIENT HAD A PRESSURE DROP. IT WAS OBSERVED BY ULTRASOUND THAT THE PATIENT HAD A LARGE PERICARDIAL EFFUSION AROUND THE RIGHT ATRIUM AND RIGHT VENTRICLE. THEY PERFORMED A PERICARDIAL EFFUSION AND CPR. OVER 2 LITERS WAS REMOVED FROM THE PERICARDIAL SPACE. SURGICAL INTERVENTION WAS CONSIDERED, BUT THE FAMILY OF THE PATIENT DECLINED AS THE PATIENT WAS A DNR (DO NOT RESUSCITATE). THE PATIENT THEN PASSED AWAY. INFORMATION RECEIVED INDICATED THE PHYSICIAN¿S OPINION ON THE CAUSE OF DEATH WAS THAT THE NUVISION CATHETER PERFORATED THE RIGHT ATRIUM. Manufacturer Narrative: D4: UDI: AS THE LOT NUMBER FOR THE DEVICE INVOLVED IN THE EVENT WAS NOT PROVIDED, THE FULL UDI IS CURRENTLY NOT AVAILABLE. AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. AN EVALUATION OF THE MANUFACTURING RECORD COULD NOT BE PERFORMED AS THE REQUIRED PRODUCT IDENTIFICATION NUMBER WAS NOT PROVIDED TO COMPLETE THE EVALUATION. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. BASED ON THE INFORMATION AVAILABLE, THERE IS NO INDICATION THAT A DESIGN OR MANUFACTURING ISSUE HAS CAUSED THE REPORTED COMPLAINT CONDITION. IF ADDITIONAL INFORMATION IS RECEIVED REGARDING THIS EVENT, A SUPPLEMENTAL 3500A REPORT WILL BE SUBMITTED TO THE FDA. MANUFACTURER'S REFERENCE NUMBER: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22689795&pc=NPT | 9617601-2025-00778 | 2025-07-22 04:00 | Death | MEDTRONIC MEXICO S. DE R.L. DE CV | 2025-08-01 04:00 | NPT | EVOLUT FX PLUS VALVE | Patient Device Interaction Problem | Low Blood Pressure/ Hypotension; Perforation; Pericardial Effusion | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A ROUTINE IMPLANTATION OF AN EVOLUT FX+29 MILLIMETER (MM) TRANSCATHETER AORTIC VALVE IN A TRICUSPID MODERATE CALCIFIED VALVE, WITH PRE-DILATATION USING A 23 MM NON-MEDTRONIC BALLOON (VACS3) AND USE OF A NON-MEDTRONIC GUIDEWIRE (SAFARI), THE PATIENT'S BLOOD PRESSURE DROPPED AFTER THE FINAL RELEASE OF THE VALVE. AN ECHOCARDIOGRAM REVEALED A PERICARDIAL EFFUSION. CARDIOPULMONARY RESUSCITATION AND PERICARDIAL DRAINAGE WERE PERFORMED FOR 25 MINUTES. THE PATIENT DIED, WITH THE SUSPECTED CAUSE OF DEATH BEING LEFT VENTRICULAR LEAD PERFORATION ATTRIBUTED TO THE NON-MEDTRONIC GUIDEWIRE (SAFARI). Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22690005&pc=NKM | 2135147-2025-04289 | 2018-04-01 04:00 | Death | ABBOTT MEDICAL | 2025-08-01 04:00 | NKM | MITRACLIP® SYSTEM CLIP DELIVERY SYSTEM | Incomplete Coaptation | No Clinical Signs, Symptoms or Conditions | P100009 |  | 1 | Event Description: THE ARTICLE "TRANSCATHETER EDGE-TO-EDGE REPAIR VS MEDICAL THERAPY IN ATRIAL FUNCTIONAL MITRAL REGURGITATION: A PROPENSITY SCORE-BASED COMPARISON FROM THE OCEAN-MITRAL AND REVEAL-AFMR REGISTRIES" WAS REVIEWED. THE ARTICLE PRESENTED A PROSPECTIVE MULTI CENTER STUDY, TO EVALUATE ATRIAL FUNCTIONAL MITRAL REGURGITATION (AFMR) COMMONLY AFFECTS ELDERLY AND GRAIL INDIVIDUALS. DEVICES MENTIONED INCLUDE MITRACLIP. THE ARTICLE CONCLUDED THAT IN REAL-WORLD DATA, TEER FOR PATIENTS WITH MODERATE OR SEVERE AFMR WAS ASSOCIATED WITH A LOWER INCIDENCE OF ADVERSE EVENTS COMPARED WITH MEDICAL TREATMENT. (B)(6). THE TIME FRAME OF THE STUDY WAS APRIL 2018 TO JUNE 2023. A TOTAL OF 1,081 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH 441 RECEIVED AN ABBOTT DEVICE. AS THIS EVENT IS FROM A LITERATURE REVIEW, THERE IS NO RELEVANT PATIENT INFORMATION (DATE OF BIRTH, AGE, GENDER, WEIGHT, AND MEDICAL HISTORY) TO REPORT. COMORBIDITIES INCLUDED HYPERTENSION, DIABETES, PRIOR SYMPTOMATIC STROKE, PRIOR HEART FAILURE ADMISSION. (B)(4), UNK MITRACLIP PERI- AND POST-PROCEDURAL COMPLICATIONS INCLUDED DEATH, HEART FAILURE, PROLONGED HOSPITALIZATION, TISSUE DAMAGE, SINGLE LEAFLET DEVICE ATTACHMENT (SLDA). Manufacturer Narrative: IT IS UNKNOWN IF THE DEVICE IS RETURNING FOR ANALYSIS. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS, AND A REVIEW OF THE LOT HISTORY RECORD AND SIMILAR COMPLAINT REVIEW COULD NOT BE PERFORMED AS THE PART AND LOT NUMBER REGARDING THE COMPLAINT DEVICE WAS NOT PROVIDED. THE INVESTIGATION WAS UNABLE TO DETERMINE A CAUSE FOR THE REPORTED SLDA, HEART FAILURE, TISSUE INJURY, AND DEATH. HEART FAILURE, TISSUE INJURY, AND DEATH ARE LISTED IN THE INSTRUCTIONS FOR USE AS KNOWN POSSIBLE COMPLICATIONS ASSOCIATED WITH THE MITRACLIP PROCEDURES. THE REPORTED HOSPITALIZATION WAS A RESULT OF CASE SPECIFIC CIRCUMSTANCE. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURE, DESIGN, OR LABELING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22690011&pc=DSQ | 2916596-2025-05110 | 2025-07-05 04:00 | Death | THORATEC CORPORATION | 2025-08-01 04:00 | DSQ | THORATEC® HEARTMATE®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Adverse Event Without Identified Device or Use Problem | Anemia; Gastrointestinal Hemorrhage | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS GASTROINTESTINAL BLEEDING (GIB). THE DEVICE PARAMETERS (FLOW, SPEED, POWER) WITHIN NORMAL LIMITS FOR THIS PATIENT. IT WAS REPORTED THAT THE PATIENT'S OUTCOME WAS NOT DEVICE OR THERAPY RELATED.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT HAD A HISTORY OF ANGI ECTASIS VS DIEULAFOY LESION IN 2ND PORTION OF DUODENUM WITH ATTEMPTS AT CLIPS. THEY HAD 3 ADMISSIONS FOR ACUTE BLEEDING WITH HEMODYNAMICALLY SIGNIFICANT ANEMIA. THEIR INTERNATIONAL NORMALIZED RATION (INR) WAS REDUCED FOR GOAL OF 1.5-2.0. THE PATIENT DID NOT WANT TO TRIAL STOPPING ANTICOAGULATION FOR A PROTRACTED PERIOD FEARFUL OF STROKE RISK. IN PALLIATIVE DISCUSSIONS, THE PATIENT REFERRED TO IF THE ADMISSIONS FOR GIB OR OTHER CLINICAL DIAGNOSIS CONTINUED, THEY WOULD LIKE TO CONSIDER WITHDRAWAL OF DEVICE.Event Description: HISTORY OF GI BLEED WAS PREVIOUSLY REPORTED FOR THIS PATIENT UNDER MFR #: 2916596-2025-02776. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: SECTION B2: OUTCOMES ATTRIBUTED TO ADVERSE CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF GASTROINTESTINAL (GI) BLEEDING COULD NOT BE CONCLUSIVELY DETERMINED. BLEEDING IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. GI BLEEDING HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22690044&pc=FKX | 0008030665-2025-01753 | 2025-07-18 04:00 | Death | ERIKA DE REYNOSA, S.A. DE C.V. | 2025-08-01 04:00 | FKX | LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Low Blood Pressure/ Hypotension | K222318 |  | 1 | Event Description: A PERITONEAL DIALYSIS REGISTERED NURSE (PDRN) REPORTED THAT A PATIENT HAD LOW BLOOD PRESSURE DURING TREATMENT AND THEN AFTER DISCONNECTING HAD A CARDIAC EVENT. FOLLOW-UP WITH THE PATIENT¿S PDRN REVEALED THE PATIENT CONTACTED THE ON-CALL PDRN FOR ASSISTANCE DISCONTINUING CCPD THERAPY DUE TO HYPOTENSION. AFTER ASSISTING THE PATIENT TERMINATE THE TREATMENT, THE PDRN INSTRUCTED HIM TO DRINK EXTRA FLUID AND GO TO THE EMERGENCY ROOM. THE PATIENT¿S SISTER REPORTEDLY STOPPED TO GET GAS WHILE TRANSPORTING THE PATIENT TO THE HOSPITAL, AFTER WHICH THE PATIENT REPORTED HE WAS FEELING BETTER AND DECIDED AGAINST GOING TO THE HOSPITAL. THE PATIENT¿S SISTER DROPPED THE PATIENT BACK OFF AT HIS RESIDENCE AND LEFT THE PATIENT UNTIL THE FOLLOWING MORNING. WHEN THE PATIENT¿S SISTER RETURNED, SHE DISCOVERED THE PATIENT WAS UNRESPONSIVE AND WAS PRONOUNCED DEAD AT THE PATIENT¿S RESIDENCE BY THE COUNTY. THE PDRN REPORTED THE PATIENT WAS NOT TAKEN TO THE HOSPITAL, AND NO RESUSCITATIVE MEASURES WERE UNDERTAKEN. ALTHOUGH THE SPECIFIC TIMELINE OF EVENTS IS UNKNOWN, IT APPEARS THE PATIENT DID NOT RESTART CCPD TREATMENT UPON RETURNING HOME. THE CAUSE OF DEATH WAS NOT PROVIDED, HOWEVER THERE WAS NO ALLEGATION A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) DEFICIENCY OR MALFUNCTION OCCURRED.Event Description: A PERITONEAL DIALYSIS REGISTERED NURSE (PDRN) REPORTED THAT A PATIENT HAD LOW BLOOD PRESSURE DURING TREATMENT AND THEN AFTER DISCONNECTING HAD A CARDIAC EVENT. FOLLOW-UP WITH THE PATIENT¿S PDRN REVEALED THE PATIENT CONTACTED THE ON-CALL PDRN FOR ASSISTANCE DISCONTINUING CCPD THERAPY DUE TO HYPOTENSION. AFTER ASSISTING THE PATIENT TERMINATE THE TREATMENT, THE PDRN INSTRUCTED HIM TO DRINK EXTRA FLUID AND GO TO THE EMERGENCY ROOM. THE PATIENT¿S SISTER REPORTEDLY STOPPED TO GET GAS WHILE TRANSPORTING THE PATIENT TO THE HOSPITAL, AFTER WHICH THE PATIENT REPORTED HE WAS FEELING BETTER AND DECIDED AGAINST GOING TO THE HOSPITAL. THE PATIENT¿S SISTER DROPPED THE PATIENT BACK OFF AT HIS RESIDENCE AND LEFT THE PATIENT UNTIL THE FOLLOWING MORNING. WHEN THE PATIENT¿S SISTER RETURNED, SHE DISCOVERED THE PATIENT WAS UNRESPONSIVE AND WAS PRONOUNCED DEAD AT THE PATIENT¿S RESIDENCE BY THE COUNTY. THE PDRN REPORTED THE PATIENT WAS NOT TAKEN TO THE HOSPITAL, AND NO RESUSCITATIVE MEASURES WERE UNDERTAKEN. ALTHOUGH THE SPECIFIC TIMELINE OF EVENTS IS UNKNOWN, IT APPEARS THE PATIENT DID NOT RESTART CCPD TREATMENT UPON RETURNING HOME. THE CAUSE OF DEATH WAS NOT PROVIDED, HOWEVER THERE WAS NO ALLEGATION A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) DEFICIENCY OR MALFUNCTION OCCURRED. UPON FURTHER FOLLOW UP, THE PDRN STATED THAT THE OFFICER INFORMED HER OVER THE PHONE THAT THE CAUSE OF DEATH WAS DUE TO NATURAL CAUSES. THE CLOSEST ASSUMPTION SHE COULD MAKE IS THAT IT WAS A CARDIAC EVENT, LIKELY CARDIAC ARREST. Manufacturer Narrative: CLINICAL INVESTIGATION: IT IS UNKNOWN IF A TEMPORAL RELATIONSHIP EXISTS BETWEEN CCPD THERAPY UTILIZING A LIBERTY SELECT CYCLER AND THE SERIOUS ADVERSE EVENTS OF HYPOTENSION, CARDIAC ARREST AND DEATH, AS THE TIMELINE OF EVENTS IS UNKNOWN. THE PDRN REPORTED THE PATIENT SUFFERED A CARDIAC EVENT, HOWEVER THE ETIOLOGY OF THE SERIOUS ADVERSE EVENTS IS UNKNOWN; THEREFORE, CAUSALITY CANNOT BE FIRMLY ESTABLISHED. THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. CARDIOVASCULAR DISEASE (INCLUDING SUDDEN CARDIAC DEATH) ACCOUNTS FOR THE MOST DEATHS AMONG THE ESRD POPULATION. ADDITIONALLY, ADULTS WITH ESRD HAVE MORTALITY RATES UP TO 30-FOLD HIGHER THAN THE GENERAL POPULATION. BASED ON THE INFORMATION AVAILABLE, THE LIBERTY SELECT CYCLER CANNOT BE DISASSOCIATED FROM PLAYING A POSSIBLE ROLE IN THE DEVELOPMENT OF THE SERIOUS ADVERSE EVENTS. THERE IS CURRENTLY NO ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) DEFICIENCY OR MALFUNCTION OCCURRED. HOWEVER, GIVEN THE TEMPORAL RELATIONSHIP, UNKNOWN CAUSALITY, NO CAUSE OF DEATH, AND NO MANUFACTURER EVALUATION OF THE PATIENT¿S DEVICE, THIS CLINICAL INVESTIGATION CANNOT EXCLUDE THE PATIENT¿S LIBERTY SELECT CYCLER FROM PLAYING A POSSIBLE ROLE IN THE SERIOUS ADVERSE EVENTS. THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22690098&pc=DSQ | 2916596-2025-05130 | 2025-07-23 04:00 | Death | THORATEC CORPORATION | 2025-08-01 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Dyspnea; Fatigue; Renal Failure; Heart Failure/Congestive Heart Failure; Swelling/ Edema; Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS DUE TO RIGHT HEART FAILURE AND THAT THE PATIENT CHOSE TO HAVE HER VAD DISCONTINUED. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. ADDITIONAL INFORMATION WAS UNKNOWN.Event Description: RIGHT HEART FAILURE DID NOT EXIST PRE-LEFT VENTRICULAR ASSIST DEVICE (LVAD) IMPLANT. A DEVICE RELATED ISSUE DID NOT CAUSE OR CONTRIBUTE TO RIGHT HEART FAILURE. THE PATIENT HAD EXPERIENCED SHORTNESS OF BREATH, SWELLING, KIDNEY FAILURE, AND FATIGUE. THE PATIENT WAS TREATED WITH DOBUTAMINE THEN MILRINONE. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED AND THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RIGHT HEART FAILURE AND RENAL DYSFUNCTION) AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 OF THE IFU, UNDER "RIGHT HEART FAILURE", DISCUSSES THE POTENTIAL DEVELOPMENT OF RIGHT HEART FAILURE FOLLOWING IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS, INCLUDING INOTROPES AND RVAD PLACEMENT. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22691212&pc=DSQ | 2916596-2025-05143 | 2025-07-24 04:00 | Death | THORATEC CORPORATION | 2025-08-01 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Mechanical Problem; Infusion or Flow Problem | Septic Shock; Thrombosis/Thrombus | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD NEW LOW FLOW ALARMS ON THEIR EXCHANGED PUMP. LOG FILES WERE SENT FOR REVIEW. THE EVENT LOG FILE RECORDED (F67) HARMONIC COMPENSATION LEFT VENTRICULAR ASSIST DEVICE (LVAD) AND (F55) MOTOR INSTABILITY FAULT FLAGS SHORTLY AFTER STARTUP. THE CORRESPONDING LVAD LOG FILE DATA RECORDED ELEVATED ROTOR DISPLACEMENT AND MAGLEV VALUES DURING THESE EVENTS. LOW FLOW ALARM FLAGS WERE RECORDED DURING PERIODS OF (F55) MOTOR INSTABILITY EVENTS. THESE EVENTS WERE KNOWN TO OCCUR WHEN THERE WAS THE PRESENCE OF SOMETHING OTHER THAN BLOOD IN THE ROTOR CHAMBER. THE DATA INDICATED THAT THESE EVENTS SUBSIDED AFTER 23:36:57 ON 24JUL2025, WHICH INDICATED THAT THIS OBSTRUCTION PASSED THROUGH THE PUMP. IT WAS LATER NOTED THAT HEMODYNAMICS WAS OFF AND THERE WERE CONCERNS FOR ANOTHER POSSIBLE PUMP THROMBUS. ADDITIONAL LOG FILES WERE SENT FOR REVIEW. THE EVENT LOG CONTAINED CLUSTERS OF PULSATILITY INDEX (PI) EVENTS AS WELL AS ROUTINE POWER SOURCE CHANGES. NO UNUSUAL ROTOR FAULTS, PUMP TEMPERATURE, OR ANY OTHER ABNORMAL PUMP FAULTS WERE SEEN IN THE LOG. BASED ON THE DATA VISIBLE IN THE LOG FILE THE MECHANICAL CIRCULATORY SUPPORT (MCS) EQUIPMENT HAD OPERATED AS INTENDED ELECTRONICALLY AND MECHANICALLY.Event Description: ADDITIONAL INFORMATION RECEIVED STATED THAT THE PATIENT PASSED AWAY ON (B)(6) 2025 DUE TO SEPTIC SHOCK. IT WAS REPORTED THAT THE OUTCOME WAS DEVICE OR THERAPY RELATED AND THAT THERE WAS A PUMP THROMBUS. AN AUTOPSY WAS NOT PERFORMED. THE PUMP WOULD NOT BE EXPLANTED AND WOULD NOT BE RETURNED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22692865&pc=BTR | 3000219639-2025-00053 |  | Death | PARKER MEDICAL | 2025-08-01 04:00 | BTR | PARKER FLEX TIP | Connection Problem; Gas/Air Leak | Cardiac Arrest | UNKNOWN |  | 1 | Event Description: IT WAS REPORTED A LEAK OCCURRED FROM THE CUFF DURING INTUBATION OF THE PATIENT WITH HEART FAILURE. THE DOCTOR PERFORMED BASIC LIFE SUPPORT AND INTUBATED THE TUBE AFTER THE PATIENT HAD GONE INTO CARDIOPULMONARY ARREST. ONCE AN ALARM SOUNDED THAT THE CIRCUIT HAD BEEN DISCONNECTED, BUT THE ALARM STOPPED WHEN AIR WAS INJECTED INTO THE CUFF. AFTER THE PATIENT DIED, A COMPUTED TOMOGRAPHY SCAN (CT) SCAN OF THE INTUBATION REVEALED NO ABNORMALITIES; UPON RETURNING TO THE INTENSIVE CARE UNIT (ICU), IT WAS CONFIRMED THAT AIR WAS LEAKING FROM THE CUFF OF THE ENDOTRACHEAL TUBE. IT WAS ADDITIONALLY REPORTED, THE HOSPITAL BELIEVED THE PATIENT DIED BECAUSE SHE WAS NOT WELL TO BEGIN WITH. Manufacturer Narrative: THE PRODUCT INVOLVED IN THE REPORT WAS RETURNED TO THE VENDOR FOR EVALUATION; ADDITIONALLY, A PHOTOGRAPH WAS PROVIDED. BASED ON THE INVESTIGATION PERFORMED BY THE VENDOR AND REVIEW OF THE PHOTOGRAPH THE CUFF WAS VERIFIED TO BE DAMAGED; THE COMPLAINT IS CONSIDERED CONFIRMED. A REVIEW OF THE DEVICE HISTORY RECORD IS NOT POSSIBLE AS NO LOT NUMBER WAS PROVIDED. THE DOCUMENTED UDI IS BASED ON THE STOCK/PRODUCT CODE PROVIDED BY THE REPORTER; THE UDI-PI IS NOT AVAILABLE AS NO LOT NUMBER WAS PROVIDED ALL INFORMATION REASONABLY KNOWN AS OF 01 AUG 2025 HAS BEEN INCLUDED IN THIS HEALTH AUTHORITY REPORT. SHOULD ADDITIONAL INFORMATION BE OBTAINED, A FOLLOW-UP HEALTH AUTHORITY REPORT WILL BE PROVIDED. THE INFORMATION PROVIDED BY SUN MED HOLDINGS LLC. REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO SUN MED HOLDINGS LLC. SUN MED HOLDINGS LLC. HAS NO INDEPENDENT KNOWLEDGE OF THE EVENT REPORTED BUT IS RELAYING THE INFORMATION THAT WAS PROVIDED BY THE USER FACILITY WHERE THE INCIDENT OCCURRED. THIS PRODUCT INCIDENT IS DOCUMENTED IN THE SUN MED HOLDINGS COMPLAINT DATABASE AND IDENTIFIED AS COMPLAINT: (B)(4). THIS INFORMATION IS SUBMITTED PURSUANT TO 21CFR803, IN COMPLIANCE WITH THE MEDICAL DEVICE REPORTING REQUIREMENT AND SHOULD NOT BE CONSIDERED TO BE AN ADMISSION THAT AN SUN MED HOLDINGS LLC. PRODUCT IS DEFECTIVE OR CAUSED SERIOUS INJURY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22693150&pc=OAE | 2029046-2025-02511 | 2025-07-07 04:00 | Death | BIOSENSE WEBSTER INC | 2025-08-01 04:00 | OAE | QDOT MICRO | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Cardiac Tamponade; Unspecified Nervous System Problem | P210027 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN ATRIAL FIBRILLATION (AF) ABLATION AND THE PATIENT EXPERIENCED A PERICARDIAL EFFUSION THAT REQUIRED PERICARDIOCENTESIS. THEN THE PATIENT EXPERIENCED CARDIAC ARREST REQUIRING CARDIOPULMONARY RESUSCITATION (CPR). THE PATIENT DIED. AN AF PROCEDURE WAS PERFORMED USING THE QDOT CATHETER. BEFORE STARTING THE PROCEDURE, THE DOCTOR WAS INFORMED OF THE RECOMMENDATIONS FOR USING THE CATHETER AND THE QMODE+ MODALITIES FOR THE POSTERIOR WALL AND QMODE FOR THE ANTERIOR WALL. THE PROCEDURE BEGAN WITH LEFT ATRIAL MAPPING WITH OCTARAY. THEN, THE CATHETER WAS EXCHANGED FOR THE QDOT. THE CATHETER WAS INSERTED INTO THE PATIENT, A ZERO WAS PERFORMED TO CALIBRATE THE CATHETER'S FORCE SENSOR, AND THERAPY BEGAN ON THE LEFT ATRIAL POSTERIOR WALL. A QMODE+ (90 W + 4 SEC) WAS USED FOR THESE ABLATIONS. THE IMPEDANCE DROPS WERE CORRECT, AND AN APPROXIMATE WEIGHT OF 6 TO 14 GRAMS WAS USED. THE ABLATIONS WERE COMPLETED, AND A REMAP WAS PERFORMED TO OBSERVE THE EFFECTIVENESS OF THE ABLATIONS. UPON COMPLETION, THE DOCTOR BEGAN TO REVIEW THE INTRACARDIAC ECHOCARDIOGRAM (ICE) TO CHECK FOR PERICARDIAL EFFUSION, AND INDEED, HE OBSERVED AN EFFUSION. THEY SPENT A COUPLE OF HOURS MONITORING THE PATIENT IN CASE OF COMPLICATIONS, BUT NOTHING ABNORMALLY OCCURRED. AT THE TIME OF REPORTING, THE REPORTER STATED THAT THE DOCTOR HAD INFORMED THEM THAT THE PATIENT HAD DEVELOPED COMPLICATIONS IN THE MIDDLE OF THE NIGHT, AND THEY HAD TO PERFORM A PERICARDIOCENTESIS. ADDITIONAL INFORMATION WAS RECEIVED WHICH INDICATED THAT AFTER DOING THE REMAP OF THE PULMONARY VEIN ISOLATION (PVI) AND THE ABLATION WAS SUCCESSFUL, THE DOCTOR CHECKED ON THE ICE AND SAW A SMALL PERICARDIAL EFFUSION, AND WAITED AROUND 2 HOURS WITH THE ICE CATHETER INSIDE TO SEE THE PERICARDIAL EFFUSION. THE PATIENT WAS STABLE DURING THESE TWO HOURS, AND THE DOCTOR DECIDED TO SEND THE PATIENT TO THE INTENSIVE CARE UNIT (ICU). AFTER 5 HOURS, THE PATIENT CONDITION WORSENED AND PERICARDIOCENTESIS WAS PERFORMED. THE PATIENT ENTERED CARDIAC ARREST, AND 30 MINUTES OF CPR WAS DONE. VITAL SIGNS WERE SEEN AFTERWARDS BUT NEUROLOGICAL DAMAGE WAS VISIBLE. DAYS LATER THE PATIENT DIED. THE PROCEDURAL ACTIVATED CLOTTING TIME (ACT) WAS AROUND 260-350 SECONDS. THREE CATHETER EXCHANGES OCCURRED DURING THE PROCEDURE. ONE TRANSSEPTAL PUNCTURE SITE WAS PERFORMED DURING THE PROCEDURE. TRANSSEPTAL PUNCTURE WAS PERFORMED BUT NO NEEDLE WAS USED. NO ABLATIONS WERE PERFORMED WITHIN THE PULMONARY VEINS, AND 68 ABLATIONS WERE PERFORMED OUTSIDE THE PULMONARY VEINS. NO EVIDENCE OF STEAM POP. THE FLOW SETTING FOR IRRIGATION CATHETER WAS QMODE FOR 3 ABLATIONS AND QMODE+ SETTINGS FOR THE REST OF ABLATIONS. THE PATIENT DID NOT REQUIRE CARDIAC SURGERY. THE CORRECT CATHETER SETTINGS WERE SELECTED ON THE GENERATOR. THERE WERE NO ISSUES WITH FLOW RATE CHANGE AT THE START OF ABLATION. NO ERROR MESSAGES OBSERVED ON BIOSENSE WEBSTER EQUIPMENT DURING THE PROCEDURE. THE FORCE VISUALIZATION FEATURES USED WERE GRAPH, DASHBOARD, AND VECTOR. THE VISITAG MODULE PARAMETERS FOR STABILITY USED WERE TAG SIZE 3, TIME 3 SECONDS, AND FOT 25%. NO ADDITIONAL FILTER USED WITH VISITAG. THE COLOR OPTION USED PROSPECTIVELY WAS IMPEDANCE AVERAGE. THE PHYSICIAN OPINION ON THE CAUSE OF DEATH WAS THAT IN THE ICU, THEY GAVE THE PATIENT MEDICATION TO RAISE HIS BLOOD PRESSURE. THIS MAY HAVE CAUSED THE BLOOD CLOT THAT WAS GENERATED TO COVER (BLOCK) THE PERICARDIAL EFFUSION TO POP OUT AND PROVOKED THE PERICARDIAL EFFUSION TO START ONCE AGAIN AND CAUSED THE PATIENT TO ENTER CARDIAC ARREST. THIRTY (30) MIN OF CPR WAS DONE, AND VITAL SIGNS WERE SEEN. NEUROLOGICAL DAMAGE WAS VISIBLE. AFTERWARDS THE PATIENT WAS TRANSFERRED TO A DIFFERENT HOSPITAL WITH A DIFFERENT DOCTOR AND DETAILS AFTER THAT WERE NOT PROVIDED. Manufacturer Narrative: AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT NUMBER: 31469278L AND NO INTERNAL ACTION RELATED TO THE COMPLAINT WAS FOUND DURING THE REVIEW. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. B2. DATE OF DEATH - THE EXACT DATE OF DEATH IS UNKNOWN AT THIS TIME AND SO THIS FIELD WAS POPULATED WITH THE AWARENESS DATE WHICH IS ON (B)(6) 2025 THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. MANUFACTURER'S REFERENCE NUMBER: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22920336&pc=FSA | 3007420694-2025-00221 | 2025-08-08 04:00 | Death | ARJO HOSPITAL EQUIPMENT AB | 2025-08-29 04:00 | FSA | MAXI TWIN | Detachment of Device or Device Component | Fall; Hemorrhage/Bleeding; Laceration(s) |  |  | 1 | Event Description: DURING A PATIENT TRANSFER FROM A BED TO A WHEELCHAIR USING A MAXI TWIN PASSIVE FLOOR LIFT OPERATED BY TWO STAFF MEMBERS, ONE OF THE SLING CLIPS DETACHED. AS A RESULT, THE PATIENT SLIPPED AND STRUCK HER HEAD AGAINST THE FRONT WHEEL OF THE LIFT. THE EVENT LED TO A LACERATION BEHIND THE LEFT SIDE OF THE PATIENT'S NECK AND SUBCUTANEOUS BLEEDING ON HER LEFT ELBOW. IMMEDIATE FIRST AID WAS ADMINISTERED, INCLUDING THE APPLICATION OF ICE TO THE INJURED AREAS IN ACCORDANCE WITH STANDARD PRACTICE. VITAL SIGNS WERE CHECKED AND FOUND TO BE WITHIN NORMAL LIMITS. AN AMBULANCE WAS CALLED, AND THE PATIENT WAS TRANSPORTED TO THE HOSPITAL FOR FURTHER EVALUATION. ON 25 AUG 2025, ARJO WAS INFORMED THAT THE PATIENT DIED IN THE HOSPITAL. AT THIS TIME, THERE IS NO CONFIRMATION THAT THERE IS ANY DIRECT LINK BETWEEN THE FALL AND THE DEATH. THERE WAS NO DEFICIENCY IDENTIFIED ON THE SYSTEM (MAXI TWIN USED WITH A SLING), IT MET THEIR SPECIFICATION. Manufacturer Narrative: THE DEVICE WAS MANUFACTURED BEFORE UDI IMPLEMENTATION AND UDI IS NOT AVAILABLE. PLEASE NOTE THAT PREVIOUS MEDWATCH REPORTS FOR THIS PRODUCT MAY HAVE BEEN SUBMITTED UNDER THE FOLLOWING REGISTRATION NUMBERS: 9611530, 3007420694. CURRENTLY, THESE PRODUCTS ARE TO BE HANDLED BY ARJOHUNTLEIGH AB¿S COMPLAINT HANDLING ESTABLISHMENT AND ANY MEDWATCH REPORTS WILL BE SUBMITTED UNDER REGISTRATION 3007420694. PROCESS OF GATHERING AND ANALYZING INFORMATION IS ONGOING. ADDITIONAL INFORMATION WILL BE PROVIDED UPON CONCLUSION OF THE INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22920337&pc=FSA | 1419652-2025-00051 | 2025-08-08 04:00 | Death | ARJO HOSPITAL EQUIPMENT AB | 2025-08-29 04:00 | FSA | MAXI TWIN | Detachment of Device or Device Component | Fall; Hemorrhage/Bleeding; Laceration(s) |  |  | 1 | Event Description: DURING A PATIENT TRANSFER FROM A BED TO A WHEELCHAIR USING A MAXI TWIN PASSIVE FLOOR LIFT OPERATED BY TWO STAFF MEMBERS, ONE OF THE SLING CLIPS DETACHED. AS A RESULT, THE PATIENT SLIPPED AND STRUCK HER HEAD AGAINST THE FRONT WHEEL OF THE LIFT. THE EVENT LED TO A LACERATION BEHIND THE LEFT SIDE OF THE PATIENT'S NECK AND SUBCUTANEOUS BLEEDING ON HER LEFT ELBOW. IMMEDIATE FIRST AID WAS ADMINISTERED, INCLUDING THE APPLICATION OF ICE TO THE INJURED AREAS IN ACCORDANCE WITH STANDARD PRACTICE. VITAL SIGNS WERE CHECKED AND FOUND TO BE WITHIN NORMAL LIMITS. AN AMBULANCE WAS CALLED, AND THE PATIENT WAS TRANSPORTED TO THE HOSPITAL FOR FURTHER EVALUATION.
ON (B)(6) 2025, ARJO WAS INFORMED THAT THE PATIENT DIED IN THE HOSPITAL. AT THIS TIME, THERE IS NO CONFIRMATION THAT THERE IS ANY DIRECT LINK BETWEEN THE FALL AND THE DEATH.
THERE WAS NO DEFICIENCY IDENTIFIED ON THE SYSTEM (MAXI TWIN USED WITH A SLING), IT MET THEIR SPECIFICATION. Manufacturer Narrative: THE DEVICE WAS MANUFACTURED BEFORE UDI IMPLEMENTATION AND UDI IS NOT AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22920662&pc=MIH | 9612164-2025-04273 |  | Death | MEDTRONIC IRELAND | 2025-08-29 04:00 | MIH | VALIANT CAPTIVIA - CW | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P100040 |  | 1 | Event Description: TWO VALIANT CAPTIVIA STENT GRAFTS WERE IMPLANTED DURING THE ENDOVASCULAR TREATMENT OF A 70MM AAA . IT WAS REPORTED THAT THE PATIENT DEVELOPED A CEREBRAL INFARCTION POST-PROCEDURE AND EXPIRED. ACCORDING TO THE PHYSICIAN, THE CAUSE OF THE CEREBRAL INFARCTION WAS THAT THE CALCIFICATION NEAR THE BASE OF THE BRACHIOCEPHALIC ARTERY WAS REMOVED WITH THE BARE PART OF THE VALIANT TIP, AND THIS WAS IDENTIFIED AS THE CAUSE OF DEATH. Manufacturer Narrative: CONTINUATION OF D10: SECTION D INFORMATION REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER RELEVANT DEVICE(S) ARE: PRODUCT ID: V AMF4642C150TJ, SERIAL/LOT #: (B)(6), UBD: 18-DEC-2025, UDI#:(B)(4), MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22920677&pc=MIH | 9612164-2025-04274 | 2025-07-29 04:00 | Death | MEDTRONIC IRELAND | 2025-08-29 04:00 | MIH | ENDURANT IIS BIFURCATED STENT GRAFT | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Ischemia; Pneumonia; Peritonitis; Multiple Organ Failure | P100021 |  | 1 | Event Description: AN ENDURANT IIS STENT GRAFT SYSTEM WAS IMPLANTED AS PART OF A CLINICAL STUDY. THE MAIN BODY (ESBF2814C103EE) WAS SUCCESSFULLY IMPLANTED, FOLLOWED BY THE PLACEMENT OF AN ILIAC LIMB (ETLW1628C124EE) IN THE LEFT ILIAC ARTERY AND (ETLW1624C93EE) IN THE RIGHT ILIAC ARTERY. PROPER CTA CONTROL WAS PERFORMED CONFIRMING THE CORRECT POSITIONING OF THE IMPLANTS. IT WAS NOTED THAT THE PATIENT DEVELOPED AND ACUTE ABDOMEN AFTER A GASTRECTOMY PERFORMED ONE MONTH PRIOR. IT WAS REPORTED 4 DAYS AFTER THE INDEX PROCEDURE, THE PATIENT DEVELOPED SEPSIS DUE TO PNEUMONIA. THE PATIENT WAS ADMITTED TO THE INTENSIVE CARE UNIT. INTESTINAL ISCHEMIA OCCURRED 2 DAYS LATER. LAPAROTOMY, SPLENECTOMY AND SUBTOTAL COLECTOMY WERE PERFORMED TO TREAT THE ISCHEMIA. ONE WEEK LATER, THE PATIENT DEVELOPED GALLIC PERITONITIS, SEPSIS LEADING TO PROLONGATION OF EXISTING HOSPITALIZATION. THIS EVENT WAS TREATED UNDER LAPAROTOMY BY STITCHING THE BILE DUCT. THE PATIENT DEVELOPED MULTI-ORGAN FAILURE DURING SUBSEQUENT INTERVENTIONS AND EXPIRED. THE SEPSIS DUE TO PNEUMONIA WAS ASSESSED BY THE SITE AS NOT RELATED TO STUDY DEVICE, INDEX PROCEDURE OR AN UNDERLYING CONDITION OR DISEASE. THE SPONSOR ASSESSED THE EVENT AS HAVING A CAUSAL RELATIONSHIP TO INDEX PROCEDURE, POSSIBLE RELATIONSHIP TO STUDY DEVICE AND NOT RELATED TO AN UNDERLYING CONDITION OR DISEASE. THE SITE ASSESSED THE INTESTINAL ISCHEMIA AS NOT RELATED TO THE STUDY DEVICE, POSSIBLY RELATED TO THE INDEX PROCEDURE AND UNDERLYING CONDITION OR DISEASE. THE GALLIC PERITONITIS, SEPSIS EVENT WAS ASSESSED BY THE SPONSOR AS POSSIBLY RELATED TO THE INDEX PROCEDURE AND NOT RELATED TO THE STUDY DEVICE OR AN UNDERLYING CONDITION OR DISEASE. THE SITE ASSESSED THIS EVENT AS NOT RELATED TO THE STUDY DEVICE, PROCEDURE, BUT POSSIBLY RELATED TO AN UNDERLYING CONDITION OR DISEASE. THE INVESTIGATOR REPORTED THE CAUSE OF DEATH AS RELATED TO THE BOWEL ISCHEMIA WITH BOWEL RESECTION AND CREATION OF A STOMA (ANUS PRAETER). NO ADDITIONAL CLINICAL SEQUELAE WERE PROVIDED AND THE PATIENT IS EXPIRED. Manufacturer Narrative: CONTINUATION OF D10: SECTION D INFORMATION REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER RELEVANT DEVICE(S) ARE: PRODUCT ID: ETLW1628C124EE, SERIAL/LOT #: (B)(6), UBD: 26-JUN-2026, UDI#: (B)(4).; PRODUCT ID: ETLW1624C93EE, SERIAL/LOT #: (B)(6), UBD: 15-FEB-2026, UDI#: (B)(4). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22920717&pc=LYJ | 1644487-2025-10563 | 2025-07-23 04:00 | Death | HOUSTON | 2025-08-29 04:00 | LYJ | PULSE GEN MODEL 1000 | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P970003 |  | 1 | Event Description: THE PATIENT HAD A SWALLOWING ASSESSMENT AND MODIFIED BARIUM SWALLOW IN (B)(6) 2025, THE LATTER OF WHICH SHOWED ESOPHAGEAL DYSMOTILITY (WHETHER THE VNS WAS ACTIVATED OR INACTIVATED). IT WAS REPORTED THAT THE PATIENT HAS PASSED AWAY ON (B)(6) FROM ASPIRATION PNEUMONIA IN THE SETTING OF ESOPHAGEAL DYSMOTILITY WITH VNS. NO ADDITIONAL RELEVANT INFORMATION HAS BEEN RECEIVED TO DATE. Manufacturer Narrative: LIVANOVA USA, INC. SUBMITS THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION, BASED ON INFORMATION THAT LIVANOVA HAS OBTAINED, BUT MAY NOT HAVE BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION, OR A CONCLUSION BY FDA OR ANYONE ELSE, THAT THE DEVICE, LIVANOVA, OR LIVANOVA'S EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS¿ OR "MALFUNCTIONS¿. THESE WORDS ARE INCORPORATED INTO THE FDA 3500A MEDWATCH FORM BY THE FDA, AND LIVANOVA OBJECTS TO THEIR USE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22920723&pc=QFG | 3004464228-2025-38742 | 2025-08-08 04:00 | Death | INSULET CORPORATION | 2025-08-29 04:00 | QFG | OMNIPOD 5 AUTOMATED INSULIN DELIVERY SYSTEM | Insufficient Information | Insufficient Information | K203768 |  | 1 | Event Description: ON (B)(6) 2025, A NURSE CALLED TO INFORM INSULET OF A PATIENT DEATH. THE PATIENT WAS USING OMNIPOD 5 (OP5) AND LIBRE 2+ SENSOR. THE PATIENT WAS FOUND DECEASED AT HOME BY PARAMEDICS ON (B)(6) 2025. THE PATIENT HAD STARTED ON OP5 ON (B)(6) 2025. THE REASON FOR DEATH IS CURRENTLY UNCONFIRMED BUT THE NURSE REPORTS ¿THE PUMP HAD NOT RECEIVED A VALUE FROM THE SENSOR SO THEY THINK IT WAS HYPERGLYCEMIA BUT THE PATIENT HAD KETOACIDOSIS BEFORE'. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL REPORT WILL BE SUBMITTED. ABBOTT HAS BEEN NOTIFIED OF THE EVENT. Manufacturer Narrative: THE DEVICE HAS NOT BEEN RETURNED/RECEIVED TO DATE. WE ARE UNABLE TO DETERMINE IF INSULET'S DEVICE CAUSED OR CONTRIBUTED TO THE REPORTED INCIDENT. NO LOT RELEASE RECORDS WERE REVIEWED, AS THE PRODUCT LOT NUMBER WAS NOT PROVIDED. THE INVESTIGATION IS ONGOING AND INSULET IS ATTEMPTING TO RECOVER ADDITIONAL DETAILS. IF ADDITIONAL INFORMATION IS RECEIVED, INSULET WILL SUBMIT A SUPPLEMENTAL REPORT AND CONDUCT ALL POSSIBLE INVESTIGATION. CLOUD - LOCKED DOWN/SMARTPHONE: DATA NOT AVAILABLE. CLOUD - OMNIPOD 5 SOFTWARE APP VERSION: DATA NOT AVAILABLE. CLOUD - SMARTPHONE OPERATING SYSTEM: DATA NOT AVAILABLE. CLOUD - SMARTPHONE HARDWARE: DATA NOT AVAILABLE. CLOUD - CGM SENSOR TYPE: DATA NOT AVAILABLE. PLEASE NOTE, THE DEVICE IDENTIFIERS ARE CAPTURED AS REPORTED BY THE COMPLAINANT AND MAY NOT ALIGN WITH THE DEVICE CONFIGURATION REPORTED IN THIS SECTION AS THIS DATA IS PULLED FROM OUR CLOUD BASED ON THE REPORTED DATE OF EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22921272&pc=LWS | 3004209178-2025-14651 | 2020-01-04 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-08-29 04:00 | LWS | EVERA MRI XT DR SURESCAN | Failure to Deliver Shock/Stimulation | Insufficient Information | P980016 |  | 1 | Event Description: IT WAS REPORTED BY THE PATIENT¿S REPRESENTATIVE THAT THE PATIENT IS DECEASED AND IT WAS NOTED THAT THE PATIENT HAD PASSED AWAY 5 YEARS AGO. THE PATIENT¿S REPRESENTATIVE STATED THAT THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) NEVER WENT OFF WHILE THE PATIENT WAS PASSING. ADDITIONAL INFORMATION RELATED TO THE CAUSE OF DEATH WAS REQUESTED AND NOT RECEIVED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22921565&pc=LJT | 3006260740-2025-06386 | 2025-08-08 04:00 | Death | C.R. BARD, INC. (BASD) -3006260740 | 2025-08-29 04:00 | LJT | UNKNOWN POWER PORT | Device Appears to Trigger Rejection | Insufficient Information; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | UNKNOWN |  | 1 | Event Description: IT WAS REPORTED THROUGH THE LITIGATION PROCESS THAT TWO YEARS AND SIX MONTHS LATER THE POST A PORT PLACEMENT, THE DEVICE ALLEGEDLY REPORTED TO BE NONFUNCTIONAL WITH FIBRIN SHEATH DEVELOPMENT. ON THE SAME DAY, PATIENT UNDERWENT REMOVAL OF PORT. LATER THE PATIENT PASSED AWAY. Manufacturer Narrative: H11: DATE OF DEATH FOR THIS PATIENT WAS NOT PROVIDED, DATE OF DEATH UPDATED AS 01-JAN-1900 FOR THE MDR SUBMISSION REQUIREMENT. AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION / EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22921585&pc=LSZ | 1719232-2025-00003 | 2025-07-28 04:00 | Death | BUNNELL, INC. | 2025-08-29 04:00 | LSZ | BUNNELL LIFEPULSE HIGH FREQUENCY VENTILATOR | Pressure Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P850064 |  | 1 | Event Description: THE USER FACILITY INDICATED THEY WERE TRYING TO ADJUST THE PEEP BUT WERE UNABLE TO. THE USER FACILITY BIOMEDICAL GROUP REPORTED: "A NOTE WITH THE DEVICE INDICATED "PEEP SET AT +6 ALL DAY, WENT TO INCREASE TO +7 AND WOULD NOT INCREASE. BUR WAS SET TO 0. WHEN BUR TURNED ON TO 1, PEEP WAS READING 3. CHECK THE ADAPTER AND COULD NOT FIND A LEAK." HOWEVER, THE WORK ORDER THAT WAS CALLED IN STATED THAT THEY WERE NOT ABLE TO ADJUST THE PEEP, LEAK IN THE ADAPTER. THEY DIDN'T INDICATE IF THERE WAS AN ERROR CODE OR NOT, AND THERE WAS A REPORTED PATIENT DEATH." Manufacturer Narrative: DEVICE RETURNED FOR INVESTIGATION ON 08/12/2025. THIS INVESTIGATION IS ONGOING. THE REPORTED SYMPTOM THAT THE PEEP COULD NOT BE CHANGED HAS NOT BEEN REPRODUCED. PRELIMINARY INVESTIGATION RESULTS: THE HFV PASSED THE SELF TEST AND THE OPERATIONAL VERIFICATION REQUIREMENTS WHERE OPERATION WAS VERY STABLE WITH MINIMUM FLUCTUATIONS OF THE MONITORED DISPLAYED VALUES AND WITH NO SYSTEM ALARMS OF ANY TYPE GENERATED IN THE HFV READY CONDITION. THE PIP ALWAYS ACHIEVED THE SET POINT AND THE RESULTANT MAP AND PEEP REMAINED EXTREMELY STABLE WITH MINIMAL FLUCTUATIONS. AT THIS THIS POINT IN TIME THE INVESTIGATION IS STILL ONGOING. THE USER FACILITY DID NOT RETURN THE ADAPTER FOR INVESTIGATION. THE EVENT DESCRIPTION FROM THE USER FACILITY DOES NOT CLEARLY INDICATE IF THIS IS A BUNNELL LIFEPORT ADAPTER OR OTHER ETT ADAPTER (WHICH IS NOT MANUFACTURED BY BUNNELL). MULTIPLE ATTEMPTS HAVE BEEN MADE, VIA PHONE AND E-MAIL, TO OBTAIN ADDITIONAL INFORMATION FROM THE USER FACILITY. NO ADDITIONAL INFORMATION HAS BEEN MADE AVAILABLE. THE USER FACILITY INDICATED ISSUES SETTING PEEP AND BUR. IT IS NOTED THAT THE LIFEPULSE DOES NOT HAVE SETTINGS FOR PEEP OR BUR. A FOLLOW-UP REPORT WILL BE SUBMITTED FOLLOWING COMPLETION OF THE INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22921761&pc=DQY | 2124215-2025-60759 | 2025-08-11 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-08-29 04:00 | DQY | WATCHMAN FXD CURVE? ACCESS SYSTEM | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Pericardial Effusion; Thrombosis/Thrombus | K212228 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT DEATH OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED, AND A WATCHMAN FXD CURVE ACCESS SYSTEM (WAS) WAS SELECTED TO BE USED. THE PATIENT UNDERWENT GENERAL ANESTHESIA FOR THE IMPLANT PROCEDURE. DURING THE PROCEDURE, PRE TRANSESOPHAGEAL ECHOCARDIOGRAPHY (TEE) IMAGING WAS PERFORMED. THE ACCESS ON THE LEFT AND RIGHT GROIN WAS OBTAINED. A 2-DIMENSIONAL INTRACARDIAC ECHOCARDIOGRAPHY (ICE) WAS INTRODUCED IN THE LEFT GROIN AND THE WAS SHEATH WITH VERSACROSS CONNECT WAS INTRODUCED IN THE RIGHT GROIN. A RF PIGTAIL WIRE WAS ADVANCED INTO THE SUPERIOR VENA CAVA (SVC). THE VERSACROSS CONNECT AND THE WAS SHEATH WERE LOADED OVER THE PIGTAIL WIRE AND DROPPED ONTO THE FOSSA OVALIS. HEPARIN WAS ADMINISTRATED TO THE PATIENT. AFTER CONFIRMING THE WAS SHEATH AND TRANSEPTAL LOCATION ON ICE VIA VISUALIZATION OF TENTING, THE RF PIGTAIL WAS APPLIED WITH BAYLIS GENERATOR AT ONE (1) SECOND AND CONSTANT SETTING. TRANSEPTAL WAS UNSUCCESSFUL AND IT WAS DETERMINED THAT THERE WAS MORE REACH NEEDED, AND BETTER CONTACT REQUIRED. THE PIGTAIL WIRE WAS ADVANCED BACK INTO THE SVC AND THE WAS SHEATH AND VERSACROSS CONNECT DILATOR WERE REMOVED AND RE-SHAPED. THE SYSTEM WAS ADVANCED BACK INTO THE SVC, AND IT WAS DROPPED ONTO THE FOSSA OVALIS. THE SAME RF PIGTAIL WIRE WAS APPLIED TWICE AND UNSUCCESSFULLY CROSSED. THE TRANSEPTAL LOCATION WAS THEN VISUALIZED AND CONFIRMED ON TEE. THE BAYLIS GENERATOR WAS CHANGED TO TWO (2) SECONDS AND PULSE SETTING THE RF PIGTAIL WIRE WAS APPLIED AND UNSUCCESSFUL. THE PHYSICIAN THEN DECIDED TO SWITCH TO A NON-BOSTON SCIENTIFIC (BSC) SHEATH. THE WAS SHEATH AND VERSACROSS CONNECT DILATOR WERE REMOVED AND THE NON-BSC SHEATH WAS INTRODUCED. IT WAS DROPPED ONTO THE FOSSA OVALIS, VISUALIZED ON ICE AND RF PIGTAIL WIRE WAS APPLIED THROUGH VERSACROSS RF WIRE AND BAYLIS GENERATOR AT 2 SECONDS AND PULSE SETTING. THE ATTEMPT WAS UNSUCCESSFUL. A BRK NEEDLE WAS OPENED AND INTRODUCED INTO THE NON-BSC SHEATH. A MANUAL TRANSEPTAL PUNCTURE WAS ATTEMPTED AND UNSUCCESSFUL. THE PHYSICIAN THEN DECIDED TO USE A BOVIE AND WIRE. AFTER INTRODUCING THE WIRE TO THE NON-BSC SHEATH, VISUALIZING THE TRANSEPTAL LOCATION ON ICE, THE PHYSICIAN APPLIED THE BOVIE TO THE WIRE AND IT ADVANCED INTO THE LEFT ATRIUM (LA). THE NON-BSC SHEATH AND DILATOR WERE ADVANCE INTO THE LA OVER THE WIRE. THE WIRE WAS SWITCHED OUT FOR A NON-BSC WIRE AND THE NON-BSC SHEATH AND DILATOR WERE REMOVED. THE WAS SHEATH AND DILATOR WERE INTRODUCED AND ADVANCED INTO THE LA OVER THE NON-BSC WIRE. THE WIRE WAS THEN REMOVED WITH THE DILATOR, AND A PIGTAIL CATHETER WAS INTRODUCED TO THE WAS SHEATH. THE PIGTAIL CATHETER WAS ADVANCED INTO THE LAA UNDER VISUALIZATION ON TEE AND THE WAS SHEATH WAS ADVANCED OVER IT. A 27MM CLOSURE DEVICE WAS OPENED AND PREPARED. THE CLOSURE DEVICE WAS INTRODUCED TO THE WAS SHEATH AFTER PIGTAIL REMOVAL USING A WET-TO-WET CONNECTION. AFTER A FLEX BALL WAS FORMED, THE CLOSURE DEVICE WAS DEPLOYED AND VISUALIZED ON TEE. IT WAS DETERMINED THE CLOSURE DEVICE PLACEMENT WAS TOO PROXIMAL, SO IT WAS RECAPTURED PARTIALLY AND MOVED INTO AN ANTERIOR LOBE AND DEPLOYED. THE DEPLOYMENT WAS TOO PROXIMAL, SO IT WAS PARTIALLY RECAPTURED AGAIN. AFTER THE SECOND PARTIAL RECAPTURE, THE PHYSICIAN NOTICED A DROP IN BLOOD PRESSURE. THE PHYSICIAN SWEPT FOR EFFUSION ON TEE AND VISUALIZED A CONSIDERABLE EFFUSION. THE PHYSICIAN DECIDED TO FULLY RECAPTURE THE CLOSURE DEVICE AND REMOVE THE SYSTEM ENTIRELY AND PERFORMED A PERICARDIOCENTESIS. THEY REVERSED HEPARIN AND CALLED FOR CARDIOTHORACIC SURGEON. BY THE TIME THE PHYSICIANS WERE ABLE TO TAP THE PERICARDIUM, THROMBOSIS HAD FORMED. THE CARDIOTHORACIC SURGEON ATTEMPTED TO RESCUE THE PATIENT, BUT IT WAS UNSUCCESSFUL. THE PATIENT PASSED AWAY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22921765&pc=NIK | 2182208-2025-04321 | 2025-08-15 04:00 | Death | MEDTRONIC, INC. | 2025-08-29 04:00 | NIK | AMPLIA MRI QUAD CRT-D SURESCAN | Appropriate Term/Code Not Available | Insufficient Information | P010031 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT WAS DECEASED AND THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) ALERTED FOUR HOURS POSTMORTEM. THE PATIENT¿S SPOUSE INQUIRED TO THE HCP (HEALTHCARE PROFESSIONAL) ABOUT DEVICE PERFORMANCE AS THE PATIENT HAD PASSED AWAY. ADDITIONAL INFORMATION RELATED TO THE CAUSE OF THE ALERT OR CIRCUMSTANCES SURROUNDING THE DEATH ARE NOT KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22921783&pc=OZD | 1220648-2025-46014 | 2025-07-31 04:00 | Death | ABIOMED, INC. | 2025-08-29 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT IMPLANTED WITH AN IMPELLA 5.5 EXPERIENCED OOZING AT THE AXILLARY SITE. TWO UNITS OF BLOOD PRODUCTS WERE TRANSFUSED. THE PATIENT CONTINUED ON SUPPORT AND EXPERIENCED EPISODES OF LAMINAR FLOW. AN ECHOCARDIOGRAM SHOWED THE INLET WAS DEFLECTED TO THE POSTERIOR WALL. ULTIMATELY, CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22921861&pc=QSY | MW5175399 | 2021-12-21 05:00 | Death | Z-MEDICA, LLC | 2025-08-29 04:00 | QSY | QUICKCLOT | Adverse Event Without Identified Device or Use Problem | Septic Shock; High Oxygen Saturation; Foreign Body In Patient |  |  | 1 | Event Description: (B)(6) 2021 PATIENT WITH EPISTAXIS AND RIGHT NARE PACKED WITH QUICKCLOT X3 SPRAYED WITH AFRIN. (B)(6) RN NOTES INDICATE NASAL PACKING REMOVED. (B)(6) PATIENT DECOMPENSATES AND TRANSFERS TO ICU, UNDERGOES FLEXIBLE FIBEROPTIC BRONCHOSCOPY WITH FINDINGS OF GAUZE-LIKE MATERIAL EXTENDING FROM LEFT BRONCHUS ABOVE LEVEL OF CARINA TO THE DISTAL TRACHEA. CRITICAL CARE NOTES INDICATE MULTIFACTORIAL CIRCULATORY SHOCK WITH HYPERCAPNIA RELATED TO AIRWAY OBSTRUCTION AND SEPTIC PNA. REFERENCE REPORTS: MW5107443, MW5107443-1 AND MW5175400.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22921880&pc=QSY | MW5175400 | 2021-12-21 05:00 | Death | Z-MEDICA, LLC | 2025-08-29 04:00 | QSY | QUICKCLOT | Adverse Event Without Identified Device or Use Problem | Airway Obstruction; Septic Shock; High Oxygen Saturation; Foreign Body In Patient |  |  | 1 | Event Description: (B)(6) 2021 PATIENT WITH EPISTAXIS AND RIGHT NARE PACKED WITH QUICKCLOT X3 SPRAYED WITH AFRIN. (B)(6) RN NOTES INDICATE NASAL PACKING REMOVED. (B)(6) PATIENT DECOMPENSATES AND TRANSFERS TO ICU, UNDERGOES FLEXIBLE FIBEROPTIC BRONCHOSCOPY WITH FINDINGS OF GAUZE-LIKE MATERIAL EXTENDING FROM LEFT BRONCHUS ABOVE LEVEL OF CARINA TO THE DISTAL TRACHEA. CRITICAL CARE NOTES INDICATE MULTIFACTORIAL CIRCULATORY SHOCK WITH HYPERCAPNIA RELATED TO AIRWAY OBSTRUCTION AND SEPTIC PNA. REFERENCE REPORTS: MW5107443, MW5107443-1 AND MW5175399.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22922293&pc=DZE | 3003708589-2025-00212 | 2025-08-05 04:00 | Death | THOMMEN MEDICAL AG | 2025-08-29 04:00 | DZE | ELEMENT MC INICELL | Adverse Event Without Identified Device or Use Problem | Failure of Implant |  |  | 1 | Event Description: SPINNER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22922349&pc=QCA | 2032493-2025-90444 | 2025-03-09 05:00 | Death | MICROVENTION, INC. | 2025-08-29 04:00 | QCA | LVIS HDE | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Coma; Obstruction/Occlusion | P170013 |  | 1 | Event Description: AS REPORTED THROUGH THE LVIS PAS STUDY: EVENT TERM: COMA. EVENT #: 2 EVENT START DATE: (B)(6) 2025. POST OPERATIVE DAYS: 47. EVENT DESCRIPTION: COMA LEADING TO DEATH. SUBJECT WITH NF-1 AND PRIOR INTRACRANIAL VERTEBRAL STENT (LVIS PLACED (B)(6) 2025) PRESENTED ON (B)(6) IN COMA AFTER BEING FOUND UNRESPONSIVE AT HOME (LAST KNOWN NORMAL 0830). HISTORY NOTABLE FOR RECENT MEDICATION NONCOMPLIANCE (STOPPED ANTIPLATELETS AND ANTICOAGULANTS). INTUBATED IN THE FIELD AND BROUGHT TO OSH ED WITH NIHSS 36. ON TRANSFER TO TREATING HOSPITAL, EXAM SHOWED PINPOINT PUPILS, POSITIVE COUGH AND OCRS, AND EXTENSOR POSTURING. CTA SHOWED OCCLUSION OF THE LEFT V4 STENT EXTENDING INTO THE PROXIMAL BASILAR ARTERY, CHRONIC R ICA OCCLUSION, AND R P2 OCCLUSION. NO TNK GIVEN DUE TO EXTENDED LKN; TRANSFERRED FOR EVT. UNDERWENT SUCCESSFUL THROMBECTOMY WITH TICI 2C FLOW. POST-PROCEDURE EXAM IMPROVED SLIGHTLY (REACTIVE R PUPIL, COUGH), BUT MRI SHOWED EXTENSIVE INFARCTS IN POSTERIOR FOSSA AND BILATERAL ANTERIOR AND POSTERIOR CIRCULATION. GIVEN POOR PROGNOSIS, FAMILY OPTED FOR COMFORT CARE. SUBJECT WAS TERMINALLY EXTUBATED AND DIED AT 11:12 ON (B)(6). IMMEDIATE CAUSE OF DEATH: BRAIN COMPRESSION. UNDERLYING CAUSE LEADING TO ABOVE: ACUTE ISCHEMIC STROKE. UNDERLYING CAUSE LEADING TO ABOVE: VERTEBRAL ARTERY STENT OCCLUSION. UNDERLYING CAUSE LEADING TO ABOVE: MEDICATION NON- COMPLIANCE. SITE ASSESSED RELATIONSHIP: UNKNOWN RELATIONSHIP. AE OUTCOME: DEATH. Manufacturer Narrative: MICROVENTION IS SUBMITTING THIS REPORT TO COMPLY WITH FDA REPORTING REGULATIONS UNDER 21 CFR PARTS 4 AND 803. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MICROVENTION, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MICROVENTION HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION AND HAS PROVIDED AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MICROVENTION, OR ITS EMPLOYEES THAT THE DEVICE, MICROVENTION, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. A GOOD FAITH EFFORT WILL BE MADE TO OBTAIN THE APPLICABLE INFORMATION RELEVANT TO THE REPORT. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. A DETAILED MEDICAL REVIEW OF THE MARCH 9, 2025, OPERATIVE NOTE, PERFORMED BY MICROVENTION¿S EXPERT PERSONNEL, WAS PERFORMED ON (B)(6) 2025: ¿AS REPORTED IN THE LVIS PAS STUDY, THE PATIENT WAS TREATED INITIALLY FOR A BASILAR ARTERY OCCLUSION VIA A MECHANICAL THROMBECTOMY WITH THE USE OF AN LVIS STENT WITH INDEX PROCEDURE DATE OF (B)(6) 2025. ON (B)(6) 2025, FORTY-EIGHT DAYS AFTER PERFORMANCE OF THE INDEX SURGICAL PROCEDURE, THE PATIENT WAS DETERMINED TO BE IN COMA STATUS AFTER BEING FOUND UNRESPONSIVE AT HOME WITH NOTABLE HISTORY FOR RECENT MEDICATION NONCOMPLIANCE (STOPPED ANTIPLATELETS AND ANTICOAGULANTS). DATA REVIEW INDICATES THAT THE PATIENT WAS INTUBATED IN THE FIELD AND BROUGHT TO OSH ED WITH NIHSS 36. OPERATIVE NOTE DATA REVIEWED INDICATES THE PERFORMANCE OF CT SCANNING DETERMINED THE PRESENCE OF AN OCCLUSION OF THE BASILAR ARTERY DOWN TO THE LEVEL OF THE STENT IN THE V4 SEGMENT, CHRONIC R ICA OCCLUSION, AND R P2 OCCLUSION. ACCORDING TO THE OPERATIVE REPORT. THE PATIENT UNDERWENT SUCCESSFUL THROMBECTOMY PROCEDURE WHICH DEMONSTRATED SIGNIFICANTLY IMPROVED FLOW WITHIN THE BASILAR SYSTEM AS WELL AS THE RIGHT HEMISPHERE BEING PROFUSED BY A LARGE PCOM DUE TO A RIGHT INTERNAL CAROTID ARTERY OCCLUSION. ACCORDING TO THE OPERATIVE REPORT PATIENT RECEIVED INTEGRILIN TO MAINTAIN PATENCY WITHIN THE STENT AND WAS ALSO GIVE HEPARIN. DATA REVIEWED INDICATES THAT THE PATIENT'S POST-PROCEDURE EXAM IMPROVED SLIGHTLY (REACTIVE R PUPIL, COUGH), BUT MRI SHOWED EXTENSIVE INFARCTS IN POSTERIOR FOSSA AND BILATERAL ANTERIOR AND POSTERIOR CIRCULATION. THE PATIENT'S CONDITION WAS REPORTED AS POOR PROGNOSIS STATUS. BASED ON THE REVIEW OF THE DATA TO INCLUDE PATIENT HISTORY OF MEDICINAL NON-COMPLIANCE (PATIENT STOPPED ANTIPLATELETS AND ANTICOAGULANTS) AND IN MY PROFESSIONAL OPINION, THE RELATIONSHIP OF THE EVENT TO THE DEVICE HAS BEEN DETERMINED TO BE UNKNOWN, AS THE EXACT CIRCUMSTANCES AS TO WHY THE STENT OCCLUSION OCCURRED HAVE NOT BEEN DETERMINED AND PROVIDED.¿ INVESTIGATION FINDINGS: WITHOUT THE RETURN AND PHYSICAL EVALUATION OF THE DEVICE, THE INVESTIGATION CANNOT FURTHER EXAMINE IF A CONDITION EXISTED THAT WOULD HAVE CONTRIBUTED TO THE REPORTED EVENT. WITHOUT IMAGING, THE INVESTIGATION CANNOT FURTHER VERIFY IF THE EVENT OCCURRED AS DESCRIBED, NOR COULD THE INVESTIGATION FURTHER EVALUATE THE CAUSE OF THE REPORTED EVENT. BASED ON A REVIEW OF THE DEVICE¿S RISK DOCUMENTATION, THE REPORTED EVENT DID NOT INDICATE THERE WERE THE PRESENCE OF ANY POTENTIAL OR NEW MANUFACTURING, DESIGN, QUALITY, OR OTHER SYSTEMIC ISSUES, OR NON-CONFORMANCES. THE COMPLAINT CODE IS MONITORED THROUGH THE TRENDING PROCESS; CORRECTIVE ACTION IS DETERMINED, AS NEEDED, THROUGH THIS PROCESS. INVESTIGATIONS OF HISTORIC COMPLAINT FILES WITH SIMILAR COMPLAINT CATEGORY CODING ARE RECORDED IN THE COMPLAINT HANDLING SYSTEM; WITHOUT THE ABILITY TO PERFORM AN ANALYSIS OF THE DEVICE, THIS INVESTIGATION CANNOT IDENTIFY WITH CERTAINTY ANY POTENTIAL ROOT CAUSES. BATCH REVIEW: A SEARCH FOR NON-CONFORMANCES ASSOCIATED WITH THIS PART/LOT NUMBER COMBINATION COULD NOT BE PERFORMED AS THIS INFORMATION WAS NOT AVAILABLE AT THE TIME THIS INVESTIGATION WAS PERFORMED. COMPLAINT SYSTEM REVIEW: A SEARCH OF THE COMPLAINT HANDLING SYSTEM COULD NOT BE PERFORMED TO DETERMINE IF OTHER SIMILAR COMPLAINTS EXIST FOR THIS BATCH NUMBER, BECAUSE THE BATCH NUMBER WAS NOT PROVIDED FOR THE PRODUCT ON THIS COMPLAINT FILE. IFU REVIEW (ADDITIONAL INFORMATION CAN BE FOUND IN THE IFU): ADVERSE EVENTS POSSIBLE ADVERSE EVENTS INCLUDE BUT ARE NOT LIMITED TO THE FOLLOWING: ¿ HEMATOMA AT THE PUNCTURE SITE ¿ PERFORATION OR DISSECTION OF THE VESSEL(S) ¿ INTRAVASCULAR SPASM ¿ HEMORRHAGING ¿ RUPTURE OR PERFORATION OF ANEURYSM ¿ COIL HERNIATION ¿ DEVICE MIGRATION ¿ NEUROLOGIC INSUFFICIENCIES INCLUDING STROKE AND DEATH ¿ ISCHEMIA ¿ VASCULAR OCCLUSION ¿ VESSEL STENOSIS ¿ INCOMPLETE ANEURYSM OCCLUSION ¿ PSEUDOANEURYSM FORMATION ¿ DISTAL EMBOLIZATION ¿ HEADACHE ¿ INFECTION ¿ REACTION TO CONTRAST AGENTS INCLUDING SEVERE ALLERGIC REACTIONS AND RENAL FAILURE WARNINGS SHOULD UNUSUAL RESISTANCE BE FELT AT ANY TIME DURING ACCESS OR REMOVAL, THE INTRODUCER/GUIDE CATHETER/MICROCATHETER AND LVIS DEVICE SHOULD BE REMOVED AS A SINGLE UNIT. APPLYING EXCESSIVE FORCE DURING DELIVERY OR RETRIEVAL OF THE LVIS DEVICE CAN POTENTIALLY RESULT IN LOSS OR DAMAGE TO THE DEVICE AND DELIVERY COMPONENTS. IT IS IMPERATIVE TO USE THE LVIS DEVICE WITH COMPATIBLE MICROCATHETERS. IF REPEATED FRICTION IS ENCOUNTERED DURING LVIS DEVICE DELIVERY, VERIFY MICROCATHETER IS NOT KINKED OR IN EXTREMELY TORTUOUS ANATOMY. CONFIRM THAT THE MICROCATHETER DOES NOT OVALIZE. CONFIRM THAT THERE IS ADEQUATE STERILE FLUSH SOLUTION. DO NOT REPOSITION THE LVIS DEVICE IN THE PARENT VESSEL WITHOUT FULLY RETRIEVING THE DEVICE. THE LVIS DEVICE MUST BE RETRIEVED INTO THE MICROCATHETER AND RE-DEPLOYED AT THE DESIRED TARGET LOCATION OR REMOVED COMPLETELY FROM THE PATIENT. DIRECTIONS FOR USE 15. ADVANCE THE DELIVERY WIRE TO TRANSFER THE LVIS DEVICE FROM WITHIN THE INTRODUCER INTO THE MICROCATHETER. WARNING: DO NOT TORQUE THE DELIVERY WIRE WHILE ADVANCING OR RETRACTING THE LVIS DEVICE. A TORQUE DEVICE SHOULD NOT BE USED. 16. CONTINUE ADVANCING THE DELIVERY WIRE INTO THE MICROCATHETER UNTIL THE PROXIMAL TIP OF THE DELIVERY WIRE ENTERS THE INTRODUCER. LOOSEN THE RHV LOCKING RING, REMOVE THE INTRODUCER, AND SET IT ASIDE. WARNING: DO NOT APPLY UNDUE FORCE. IF RESISTANCE IS ENCOUNTERED AT ANY POINT DURING LVIS DEVICE DELIVERY OR MANIPULATION, WITHDRAW THE UNIT AND SELECT A NEW LVIS DEVICE. 18. POSITION THE LVIS DEVICE FOR DEPLOYMENT BY ALIGNING THE LVIS IMPLANT DISTAL RADIOPAQUE END MARKERS APPROXIMATELY 7 MM PAST THE ANEURYSM NECK. NOTE: A PROPER PUSH/PULL TECHNIQUE, ENCOMPASSING SUFFICIENT DELIVERY WIRE PUSH FORCE, IN ADDITION TO AN OPPOSING MICROCATHETER WITHDRAWAL FORCE, WILL FACILITATE PROPERLY DEPLOYING THE LVIS DEVICE TO ACHIEVE FULL EXPANSION AND GOOD VESSEL APPOSITION. NOTE: SLOWLY ADVANCING THE LVIS DEVICE WHILE ADJUSTING THE MICROCATHETER POSITION WILL ENSURE ACCURATE DEPLOYMENT. MAINTAIN SIMULTANEOUS CONTROL OF THE LVIS DEVICE AND MICROCATHETER IN ORDER TO POSITION AND EXPAND THE DEVICE AT THE PROPER LOCATION. CAUTION: USING A RAPID MICROCATHETER WITHDRAWAL TECHNIQUE TO DEPLOY THE LVIS DEVICE IS NOT RECOMMENDED AND MAY RESULT IN DEVICE ELONGATION. 19. IF LVIS DEVICE POSITIONING IS NOT SATISFACTORY, THE LVIS DEVICE MAY BE RECAPTURED AND REPOSITIONED IF IT IS NOT FULLY DEPLOYED. THE LVIS DEVICE MAY BE RECAPTURED UNTIL THE POINT WHERE THE PROXIMAL END OF THE LVIS DEVICE MARKERS IS ALIGNED 3 MM PROXIMALLY WITH THE MICROCATHETER DISTAL MARKER BAND (APPROXIMATELY 80% DEPLOYED). CAUTION: IF RESISTANCE IS FELT WHILE RECAPTURING THE LVIS DEVICE, DO NOT CONTINUE TO RECAPTURE THE DEVICE. WITHDRAW THE MICROCATHETER SLIGHTLY TO UNSHEATH THE LVIS DEVICE (WITHOUT EXCEEDING THE RECAPTURE LIMIT), AND THEN ATTEMPT TO RECAPTURE THE LVIS DEVICE. CAUTION: THE LVIS DEVICE MUST NOT BE RE-DEPLOYED MORE THAN THREE TIMES. NOTE: THE LVIS DEVICE DELIVERY WIRE SHOULD NOT BE UTILIZED AS A GUIDEWIRE. DO NOT TORQUE THE LVIS DEVICE. A TORQUE DEVICE SHOULD NOT BE USED. 20. IF LVIS DEVICE POSITIONING IS SATISFACTORY, CAREFULLY RETRACT THE MICROCATHETER AND ADVANCE THE DELIVERY WIRE TOGETHER, TO ALLOW THE LVIS DEVICE TO DEPLOY ACROSS THE NECK OF THE ANEURYSM. ENSURE THE DEVICE PROXIMAL RADIOPAQUE END MARKERS ARE APPROXIMATELY 7 MM PROXIMAL TO THE ANEURYSM NECK TO ENSURE AN ADEQUATE LANDING ZONE. THE LVIS DEVICE WILL EXPAND AND TOTAL LENGTH MAY FORESHORTEN UP TO 60% FROM ITS UNDEPLOYED LENGTH AS IT EXITS THE MICROCATHETER. ENSURE MICROCATHETER IS RETRACTED AND CLEAR FROM THE PROXIMAL FLARED ENDS. NOTE: VISUALIZE AND REFER TO THE IMPLANT RADIOPAQUE END MARKERS TO MAINTAIN ADEQUATE IMPLANT LENGTH, APPROXIMATELY 7 MM ON EACH SIDE OF THE ANEURYSM NECK OR TARGET LOCATION TO ENSURE APPROPRIATE NECK COVERAGE. WARNING: DO NOT DETACH THE LVIS DEVICE IF IT IS NOT PROPERLY POSITIONED IN THE PARENT VESSEL. OBSERVE THE DELIVERY WIRE DISTAL TIP TO ASSURE IT REMAINS WITHIN THE DESIRED LOCATION OF THE PARENT VESSEL. 21. PRIOR TO REMOVING THE DELIVERY WIRE AND IF NECESSARY, CAREFULLY POSITION THE MICROCATHETER DISTAL TO THE LVIS DEVICE TO MAINTAIN ACCESS THROUGH THE LVIS DEVICE. REMOVE AND DISCARD THE DELIVERY WIRE. WARNING: THE LVIS DEVICE DELIVERY WIRE SHOULD NOT BE UTILIZED AS A GUIDEWIRE. DO NOT TORQUE THE LVIS DEVICE. A TORQUE DEVICE SHOULD NOT BE USED. 23. USE THE GUIDEWIRE AND MICROCATHETER TO ACCESS THE ANEURYSM THROUGH THE LVIS DEVICE CELLS. WARNING: OBSERVE LVIS DEVICE MARKER POSITION DURING PLACEMENT OF THE MICROCATHETER INTO THE ANEURYSM TO ENSURE THAT THE LVIS DEVICE DOES NOT MIGRATE OR DISLODGE FROM ITS DEPLOYED POSITION. 24. AFTER THE MICROCATHETER IS POSITIONED WITHIN THE ANEURYSM, DETACHABLE COILS MAY BE DELIVERED INTO THE ANEURYSM ACCORDING TO CONVENTIONAL METHODS. WARNING: OBSERVE LVIS DEVICE MARKER POSITION DURING THE COILING PROCEDURE TO ENSURE THAT THE DEVICE DOES NOT MIGRATE FROM ITS DEPLOYED POSITION. 25. AFTER PLACING THE LAST COIL, VERIFY THAT THE LVIS DEVICE HAS REMAINED PATENT AND PROPERLY POSITIONED. ADVANCE A GUIDEWIRE TO THE MICROCATHETER TIP AND CAREFULLY REMOVE THE MICROCATHETER THROUGH THE LVIS DEVICE CELLS. NOTE: A MICROCATHETER MAY BE POSITIONED INTO THE ANEURYSM SAC PRIOR TO DELIVERY OF THE LVIS DEVICE. THE MICROCATHETER WILL BE SUPPORTED BY THE LVIS DEVICE DURING DELIVERY OF EMBOLIC COILING. AFTER COMPLETING THE COILING, THE MICROCATHETER SHOULD BE CAREFULLY REMOVED TO AVOID DISLODGING THE LVIS DEVICE. 27. CAUTION: CAREFULLY WATCH THE LVIS DEVICE DISTAL AND PROXIMAL MARKERS WHEN PASSING THROUGH THE DEPLOYED LVIS DEVICE WITH EMBOLIC COILING MICROCATHETERS TO AVOID DISPLACING THE LVIS DEVICE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22922563&pc=EOQ | 3021325287-2025-00027 | 2025-07-30 04:00 | Death | NOAH MEDICAL CORP. | 2025-08-29 04:00 | EOQ | GALAXY SYSTEM | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Pneumothorax | K223144 |  | 1 | Event Description: A GALAXY-ASSISTED BRONCHOSCOPY PROCEDURE WAS PERFORMED ON A PATIENT WITH NO DEVICE-RELATED MALFUNCTIONS OBSERVED. THE PATIENT WAS DISCHARGED THE SAME DAY; HOWEVER, APPROXIMATELY THREE HOURS LATER, THE PATIENT RETURNED WITH RESPIRATORY DISTRESS, EXPERIENCED CARDIAC ARREST, AND WAS FOUND TO HAVE A PNEUMOTHORAX. THE PHYSICIAN DID NOT ATTRIBUTE THESE EVENTS TO THE USE OF THE GALAXY DEVICE. ON AUGUST 4, 2025, THE COMPANY REPRESENTATIVE WAS INFORMED THAT THE PATIENT HAD UNFORTUNATELY PASSED AWAY. THE PHYSICIAN ATTRIBUTED THE CAUSE OF DEATH TO BILATERAL PNEUMOTHORAX RESULTING FROM THE THYMECTOMY AND CONFIRMED THAT THE GALAXY DEVICE WAS NOT CONSIDERED A CONTRIBUTING FACTOR. Manufacturer Narrative: ON (B)(6) 2025, A BRONCHOSCOPY PROCEDURE WAS PERFORMED USING THE NOAH GALAXY DEVICE, AND NO DEVICE-RELATED ISSUES WERE REPORTED DURING THE CASE. THE PROCEDURE WAS COMPLETED SUCCESSFULLY, AND THE PATIENT WAS DISCHARGED. IMMEDIATE POST-PROCEDURE X-RAY SHOWED NO EVIDENCE OF PNEUMOTHORAX. APPROXIMATELY THREE HOURS LATER, THE PATIENT RETURNED WITH RESPIRATORY DISTRESS AND SUBSEQUENTLY EXPERIENCED CARDIAC ARREST. A PNEUMOTHORAX WAS IDENTIFIED, AND THE PATIENT WAS ADMITTED AND WAS TREATED WITH TARGETED TEMPERATURE MANAGEMENT. THE PHYSICIAN DID NOT ATTRIBUTE THESE SERIOUS ADVERSE EVENTS TO THE GALAXY DEVICE. A CLINICAL ENGINEER CONDUCTED A LOG REVIEW, WHICH DID NOT REVEAL ANY DEVICE-RELATED MALFUNCTIONS OR USER ERRORS. ON AUGUST 4, 2025, THE COMPANY REPRESENTATIVE WAS NOTIFIED THAT THE PATIENT HAD PASSED AWAY. THE PHYSICIAN ATTRIBUTED THE CAUSE OF DEATH TO BILATERAL PNEUMOTHORAX RESULTING FROM THE THYMECTOMY AND CONFIRMED THAT THE GALAXY DEVICE WAS NOT CONSIDERED A CONTRIBUTING FACTOR.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22922579&pc=NGV | 2124215-2025-60733 | 2025-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-08-29 04:00 | NGV | SYSTEM, APPENDAGE CLOSURE, LEFT ATRIAL | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: IT WAS REPORTED VIA JOURNAL ARTICLE THAT DEATHS OCCURRED. THE FOLLOWING EVENT WAS OBTAINED VIA A RETROSPECTIVE ARTICLE FOLLOWING 54,903 PROCEDURES PERFORMED ON PATIENTS WHO UNDERWENT A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WHERE WATCHMAN CLOSURE DEVICES WERE IMPLANTED DURING THE TIME FRAME OF JANUARY 2016 THROUGH DECEMBER 2019. IT WAS REPORTED THE FOLLOWING OCCURRED: PROCEDURE-RELATED DEATHS. Manufacturer Narrative: B2: DEATH DATES WERE NOT PROVIDED IN THE LITERATURE ARTICLE. B3: LITERATURE DATE USED AS EVENT DATE, AS EVENT DATES ARE UNKNOWN. D1: BRAND NAME IS BLANK BECAUSE THE DEVICE IS KNOWN TO BE A WATCHMAN LAA CLOSURE DEVICE, BUT THE EXACT BRAND NAME IS UNKNOWN. D4: DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BASED ON THE NATURE OF THE INFORMATION PROVIDED TO BSC, IT IS NOT POSSIBLE TO PERFORM A GOOD FAITH EFFORT TO OBTAIN ADDITIONAL INFORMATION. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER PRODUCT SPECIFIC INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. LITERATURE CITATION: GAHONA, C.T. ET AL (JAN 1, 2025). PERCUTANEOUS LEFT ATRIAL APPENDAGE OCCLUSION AND PERIPROCEDURAL OUTCOMES TREND IN THE UNITED STATES: AN INSIGHT FROM NATIONWIDE READMISSION DATABASE 2016-2019. DOI: 10.2139/SSRN.5212839.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22922699&pc=OUT | 2029214-2025-01925 | 2023-12-08 05:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-08-29 04:00 | OUT | PIPELINE | Adverse Event Without Identified Device or Use Problem | Aneurysm |  |  | 1 | Event Description: DONG, L., WANG, C., CHEN, X., LI, M., LI, T., LIU, H., ZHAO, Y., DUAN, R., JIN, W., ZHANG, Y., WANG, Y., & LV, M. (2025). PREDICTING PERSISTENT ANEURYSM FILLING AFTER PIPELINE EMBOLIZATION DEVICE TREATMENT IN PATIENTS WITH INTRACRANIAL ANEURYSM: DEVELOPMENT AND EXTERNAL VALIDATION OF A NOMOGRAM MODEL. TRANSLATIONAL STROKE RESEARCH, 16(2), 392¿402. HTTPS://DOI.ORG/10.1007/S12975-023-01222-9 LITERATURE WAS REVIEWED REGARDING A RETROSPECTIVE STUDY THAT INCLUDED PATIENTS WITH INTRACRANIAL ANEURYSM (IA) TREATED WITH A PIPELINE EMBOLIZATION DEVICE (PED), ASSIGNED TO A DERIVATION OR VALIDATION COHORT. IT SPECIFICALLY FOCUSED ON ANALYZING PATIENTS THAT DEVELOPED PERSISTENT ANEURYSM FILLING AFTER PED TREATMENT. MULTIPLE MANUFACTURER¿S DEVICES WERE IMPLANTED IN THE STUDY POPULATION. THE FOLLOWING MEDTRONIC DEVICES WERE USED: PIPELINE, MARKSMAN, PHENOM 27, ECHELON 10. DEATHS OCCURRED IN THE STUDY POPULATION; HOWEVER, THERE WAS NO STATEMENT ESTABLISHING A CAUSAL OR CONTRIBUTORY RELATIONSHIP BETWEEN THE MEDTRONIC PRODUCTS AND THE DEATHS. THE CAUSES OF DEATH WERE NOT SPECIFIED. AMONG ALL MEDTRONIC DEVICES, PATIENTS ADVERSE EVENTS / DEVICE PRODUCT PERFORMANCE ISSUES INCLUDED: POOR WALL APPOSITION REQUIRING BALLOON ANGIOPLASTY, PERSISTENT ANEURYSM FILLING, IN-STENT STENOSIS, TRANSIENT ISCHEMIC ATTACK, IN-STENT THROMBOSIS, CEREBRAL INFARCTION, DISTAL INTRAPARENCHYMAL HEMORRHAGE, DELAYED ANEURYSM RUPTURE, CRANIAL NEUROPATHY, AND BRAINSTEM SYMPTOMS RELATED TO IA COMPRESSION. A TOTAL OF 17 PATIENTS HAD A FOLLOW UP MODIFIED RANKIN SCALE (MRS) SCORE OF 3-5. COMPLETE OCCLUSION AT THE LAST ANGIOGRAPHIC FOLLOW-UP WAS ACHIEVED IN 676 PATIENTS (86.0%) IN THE DERIVATION COHORT AND 188 PATIENTS (85.5%) IN THE VALIDATION COHORT. NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. Manufacturer Narrative: LITERATURE ARTICLE IS INCLUDED IN THE ATTACHMENTS. A.2. THIS VALUE IS THE AVERAGE AGE OF THE PATIENTS REPORTED IN THE ARTICLE AS SPECIFIC PATIENTS COULD NOT BE IDENTIFIED. A.3. THIS VALUE REFLECTS THE GENDER OF THE MAJORITY OF THE PATIENTS REPORTED IN THE ARTICLE AS SPECIFIC PATIENTS COULD NOT BE IDENTIFIED. B.3. PLEASE NOTE THAT THIS DATE IS BASED OFF OF THE DATE OF ONLINE PUBLICATION OF THE ARTICLE. G.4. PMA CODE IS MISSING AS PRODUCT MODEL AND LOT NUMBERS ARE UNKNOWN. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22922796&pc=KRA | 2029214-2025-01926 | 2023-12-08 05:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-08-29 04:00 | KRA | MARKSMAN | Adverse Event Without Identified Device or Use Problem | Aneurysm |  |  | 1 | Event Description: DONG, L., WANG, C., CHEN, X., LI, M., LI, T., LIU, H., ZHAO, Y., DUAN, R., JIN, W., ZHANG, Y., WANG, Y., & LV, M. (2025). PREDICTING PERSISTENT ANEURYSM FILLING AFTER PIPELINE EMBOLIZATION DEVICE TREATMENT IN PATIENTS WITH INTRACRANIAL ANEURYSM: DEVELOPMENT AND EXTERNAL VALIDATION OF A NOMOGRAM MODEL. TRANSLATIONAL STROKE RESEARCH, 16(2), 392¿402. HTTPS://DOI.ORG/10.1007/S12975-023-01222-9 LITERATURE WAS REVIEWED REGARDING A RETROSPECTIVE STUDY THAT INCLUDED PATIENTS WITH INTRACRANIAL ANEURYSM (IA) TREATED WITH A PIPELINE EMBOLIZATION DEVICE (PED), ASSIGNED TO A DERIVATION OR VALIDATION COHORT. IT SPECIFICALLY FOCUSED ON ANALYZING PATIENTS THAT DEVELOPED PERSISTENT ANEURYSM FILLING AFTER PED TREATMENT. MULTIPLE MANUFACTURER¿S DEVICES WERE IMPLANTED IN THE STUDY POPULATION. THE FOLLOWING MEDTRONIC DEVICES WERE USED: PIPELINE, MARKSMAN, PHENOM 27, ECHELON 10. DEATHS OCCURRED IN THE STUDY POPULATION; HOWEVER, THERE WAS NO STATEMENT ESTABLISHING A CAUSAL OR CONTRIBUTORY RELATIONSHIP BETWEEN THE MEDTRONIC PRODUCTS AND THE DEATHS. THE CAUSES OF DEATH WERE NOT SPECIFIED. AMONG ALL MEDTRONIC DEVICES, PATIENTS ADVERSE EVENTS / DEVICE PRODUCT PERFORMANCE ISSUES INCLUDED: POOR WALL APPOSITION REQUIRING BALLOON ANGIOPLASTY, PERSISTENT ANEURYSM FILLING, IN-STENT STENOSIS, TRANSIENT ISCHEMIC ATTACK, IN-STENT THROMBOSIS, CEREBRAL INFARCTION, DISTAL INTRAPARENCHYMAL HEMORRHAGE, DELAYED ANEURYSM RUPTURE, CRANIAL NEUROPATHY, AND BRAINSTEM SYMPTOMS RELATED TO IA COMPRESSION. A TOTAL OF 17 PATIENTS HAD A FOLLOW UP MODIFIED RANKIN SCALE (MRS) SCORE OF 3-5. COMPLETE OCCLUSION AT THE LAST ANGIOGRAPHIC FOLLOW-UP WAS ACHIEVED IN 676 PATIENTS (86.0%) IN THE DERIVATION COHORT AND 188 PATIENTS (85.5%) IN THE VALIDATION COHORT. NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. Manufacturer Narrative: LITERATURE ARTICLE IS INCLUDED IN THE ATTACHMENTS. A.2. THIS VALUE IS THE AVERAGE AGE OF THE PATIENTS REPORTED IN THE ARTICLE AS SPECIFIC PATIENTS COULD NOT BE IDENTIFIED. A.3. THIS VALUE REFLECTS THE GENDER OF THE MAJORITY OF THE PATIENTS REPORTED IN THE ARTICLE AS SPECIFIC PATIENTS COULD NOT BE IDENTIFIED. B.3. PLEASE NOTE THAT THIS DATE IS BASED OFF OF THE DATE OF ONLINE PUBLICATION OF THE ARTICLE. G.4. PMA CODE IS MISSING AS PRODUCT MODEL AND LOT NUMBERS ARE UNKNOWN. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22922814&pc=QJP | 2029214-2025-01927 | 2023-12-08 05:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-08-29 04:00 | QJP | PHENOM CATHETER | Adverse Event Without Identified Device or Use Problem | Aneurysm |  |  | 1 | Event Description: DONG, L., WANG, C., CHEN, X., LI, M., LI, T., LIU, H., ZHAO, Y., DUAN, R., JIN, W., ZHANG, Y., WANG, Y., & LV, M. (2025). PREDICTING PERSISTENT ANEURYSM FILLING AFTER PIPELINE EMBOLIZATION DEVICE TREATMENT IN PATIENTS WITH INTRACRANIAL ANEURYSM: DEVELOPMENT AND EXTERNAL VALIDATION OF A NOMOGRAM MODEL. TRANSLATIONAL STROKE RESEARCH, 16(2), 392¿402. HTTPS://DOI.ORG/10.1007/S12975-023-01222-9 LITERATURE WAS REVIEWED REGARDING A RETROSPECTIVE STUDY THAT INCLUDED PATIENTS WITH INTRACRANIAL ANEURYSM (IA) TREATED WITH A PIPELINE EMBOLIZATION DEVICE (PED), ASSIGNED TO A DERIVATION OR VALIDATION COHORT. IT SPECIFICALLY FOCUSED ON ANALYZING PATIENTS THAT DEVELOPED PERSISTENT ANEURYSM FILLING AFTER PED TREATMENT. MULTIPLE MANUFACTURER¿S DEVICES WERE IMPLANTED IN THE STUDY POPULATION. THE FOLLOWING MEDTRONIC DEVICES WERE USED: PIPELINE, MARKSMAN, PHENOM 27, ECHELON 10. DEATHS OCCURRED IN THE STUDY POPULATION; HOWEVER, THERE WAS NO STATEMENT ESTABLISHING A CAUSAL OR CONTRIBUTORY RELATIONSHIP BETWEEN THE MEDTRONIC PRODUCTS AND THE DEATHS. THE CAUSES OF DEATH WERE NOT SPECIFIED. AMONG ALL MEDTRONIC DEVICES, PATIENTS ADVERSE EVENTS / DEVICE PRODUCT PERFORMANCE ISSUES INCLUDED: POOR WALL APPOSITION REQUIRING BALLOON ANGIOPLASTY, PERSISTENT ANEURYSM FILLING, IN-STENT STENOSIS, TRANSIENT ISCHEMIC ATTACK, IN-STENT THROMBOSIS, CEREBRAL INFARCTION, DISTAL INTRAPARENCHYMAL HEMORRHAGE, DELAYED ANEURYSM RUPTURE, CRANIAL NEUROPATHY, AND BRAINSTEM SYMPTOMS RELATED TO IA COMPRESSION. A TOTAL OF 17 PATIENTS HAD A FOLLOW UP MODIFIED RANKIN SCALE (MRS) SCORE OF 3-5. COMPLETE OCCLUSION AT THE LAST ANGIOGRAPHIC FOLLOW-UP WAS ACHIEVED IN 676 PATIENTS (86.0%) IN THE DERIVATION COHORT AND 188 PATIENTS (85.5%) IN THE VALIDATION COHORT. NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. Manufacturer Narrative: LITERATURE ARTICLE IS INCLUDED IN THE ATTACHMENTS. A.2. THIS VALUE IS THE AVERAGE AGE OF THE PATIENTS REPORTED IN THE ARTICLE AS SPECIFIC PATIENTS COULD NOT BE IDENTIFIED. A.3. THIS VALUE REFLECTS THE GENDER OF THE MAJORITY OF THE PATIENTS REPORTED IN THE ARTICLE AS SPECIFIC PATIENTS COULD NOT BE IDENTIFIED. B.3. PLEASE NOTE THAT THIS DATE IS BASED OFF OF THE DATE OF ONLINE PUBLICATION OF THE ARTICLE. G.4. PMA CODE IS MISSING AS PRODUCT MODEL AND LOT NUMBERS ARE UNKNOWN. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22922861&pc=KRA | 2029214-2025-01928 | 2023-12-08 05:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-08-29 04:00 | KRA | ECHELON | Adverse Event Without Identified Device or Use Problem | Aneurysm |  |  | 1 | Event Description: DONG, L., WANG, C., CHEN, X., LI, M., LI, T., LIU, H., ZHAO, Y., DUAN, R., JIN, W., ZHANG, Y., WANG, Y., & LV, M. (2025). PREDICTING PERSISTENT ANEURYSM FILLING AFTER PIPELINE EMBOLIZATION DEVICE TREATMENT IN PATIENTS WITH INTRACRANIAL ANEURYSM: DEVELOPMENT AND EXTERNAL VALIDATION OF A NOMOGRAM MODEL. TRANSLATIONAL STROKE RESEARCH, 16(2), 392¿402. HTTPS://DOI.ORG/10.1007/S12975-023-01222-9 LITERATURE WAS REVIEWED REGARDING A RETROSPECTIVE STUDY THAT INCLUDED PATIENTS WITH INTRACRANIAL ANEURYSM (IA) TREATED WITH A PIPELINE EMBOLIZATION DEVICE (PED), ASSIGNED TO A DERIVATION OR VALIDATION COHORT. IT SPECIFICALLY FOCUSED ON ANALYZING PATIENTS THAT DEVELOPED PERSISTENT ANEURYSM FILLING AFTER PED TREATMENT. MULTIPLE MANUFACTURER¿S DEVICES WERE IMPLANTED IN THE STUDY POPULATION. THE FOLLOWING MEDTRONIC DEVICES WERE USED: PIPELINE, MARKSMAN, PHENOM 27, ECHELON 10. DEATHS OCCURRED IN THE STUDY POPULATION; HOWEVER, THERE WAS NO STATEMENT ESTABLISHING A CAUSAL OR CONTRIBUTORY RELATIONSHIP BETWEEN THE MEDTRONIC PRODUCTS AND THE DEATHS. THE CAUSES OF DEATH WERE NOT SPECIFIED. AMONG ALL MEDTRONIC DEVICES, PATIENTS ADVERSE EVENTS / DEVICE PRODUCT PERFORMANCE ISSUES INCLUDED: POOR WALL APPOSITION REQUIRING BALLOON ANGIOPLASTY, PERSISTENT ANEURYSM FILLING, IN-STENT STENOSIS, TRANSIENT ISCHEMIC ATTACK, IN-STENT THROMBOSIS, CEREBRAL INFARCTION, DISTAL INTRAPARENCHYMAL HEMORRHAGE, DELAYED ANEURYSM RUPTURE, CRANIAL NEUROPATHY, AND BRAINSTEM SYMPTOMS RELATED TO IA COMPRESSION. A TOTAL OF 17 PATIENTS HAD A FOLLOW UP MODIFIED RANKIN SCALE (MRS) SCORE OF 3-5. COMPLETE OCCLUSION AT THE LAST ANGIOGRAPHIC FOLLOW-UP WAS ACHIEVED IN 676 PATIENTS (86.0%) IN THE DERIVATION COHORT AND 188 PATIENTS (85.5%) IN THE VALIDATION COHORT. NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. Manufacturer Narrative: LITERATURE ARTICLE IS INCLUDED IN THE ATTACHMENTS. A.2. THIS VALUE IS THE AVERAGE AGE OF THE PATIENTS REPORTED IN THE ARTICLE AS SPECIFIC PATIENTS COULD NOT BE IDENTIFIED. A.3. THIS VALUE REFLECTS THE GENDER OF THE MAJORITY OF THE PATIENTS REPORTED IN THE ARTICLE AS SPECIFIC PATIENTS COULD NOT BE IDENTIFIED. B.3. PLEASE NOTE THAT THIS DATE IS BASED OFF OF THE DATE OF ONLINE PUBLICATION OF THE ARTICLE. G.4. PMA CODE IS MISSING AS PRODUCT MODEL AND LOT NUMBERS ARE UNKNOWN. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22923031&pc=NPT | 9617601-2025-01230 | 2025-08-27 04:00 | Death | MEDTRONIC MEXICO S. DE R.L. DE CV | 2025-08-29 04:00 | NPT | EVOLUT FX VALVE | Difficult or Delayed Positioning; Perivalvular Leak; Device Dislodged or Dislocated; Patient Device Interaction Problem | Cardiac Arrest; Valvular Insufficiency/ Regurgitation | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE IN A PATIENT WITH PRE-EXISTING SEVERE AORTIC STENO SIS, AN AORTIC VALVE AREA OF 0.6 CM SQUARED, AND AN EJECTION FRACTION OF 20%, A PRE-IMPLANT BALLOON DILATION WAS PERFORMED WITH A 25 MM X 50 MM BALLOON AT NOMINAL VOLUME. THE FIRST IMPLANTATION ATTEMPT WAS VERY DIFFICULT TO POSITION, BUT THE VALVE WAS RELEASED AT A DEPTH OF 3-5 MM WITH NON-MEDTRONIC GUIDEWIRE RETRACTION, TEMPORARY PACING, AND SLOW MOVEMENTS. AFTER A PROJECTION CHANGE, DISLODGE OF THE VALVE INTO THE LEFT VENTRICLE WAS NOTED. THE VALVE WAS LOCATED AT A HEIGHT OF APPROXIMATELY 20 MM WITHIN THE LEFT VENTRICLE, WITH MODERATE PERIVALVULAR LEAK. A 15 MM SNARE WAS USED ON BOTH PADDLES TO TRACTION THE VALVE. THE VALVE WAS RETURNED TO A 3 MM POSITION AND POST-DILATED. A POST-IMPLANT BALLOON DILATION WITH A 28 MM X 50 MM BALLOON WAS PERFORMED TWICE, BUT THERE WAS NO IMPROVEMENT. AFTER DEFLATING THE BALLOON AND MAINTAINING TRACTION ON THE SNARE, THE VALVE DISLODGED OUT OF THE ASCENDING AORTA AGAIN. THE PATIENT DEVELOPED A FIRST CARDIAC ARREST LASTING 3 MINUTES, WHICH WAS THEN REVERSED. AFTERWARD, AN ATTEMPT WAS MADE TO IMPLANT A SECOND PROSTHESIS, BUT THE DELIVERY DEVICE WAS VERY DIFFICULT TO PASS THROUGH THE FIRST PROSTHESIS. THIS IS DUE TO THE LOCATION THE FIRST VALVE WAS SNARED TO. DURING THE ATTEMPT TO IMPLANT THE SECOND PROSTHESIS, THE PATIENT DEVELOPED ANOTHER CARDIAC ARREST. THE DECISION WAS MADE TO RELEASE THE SECOND VALVE URGENTLY AND ADMINISTER CARDIOPULMONARY RESUSCITATION. AFTER CIRCULATION RETURNED, THE SECOND PROSTHESIS WAS AGAIN MOVED INTO THE LEFT VENTRICLE, WHERE ATTEMPTS WERE MADE TO REPOSITION IT WITH A SNARE, BUT WITHOUT SUCCESS. THE PATIENT WENT INTO CARDIAC ARREST AGAIN AND DIED. OF NOTE, DURING THE PROCEDURE THE PATIENT REQUIRED HIGH DOSE VASOPRESSOR MEDICATION AND MECHANICAL VENTILATION. Manufacturer Narrative: CONTINUATION OF D10: SECTION D INFORMATION REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER RELEVANT DEVICE(S) ARE: PRODUCT ID: D -EVOLUTFX-34, SERIAL/LOT #: (B)(6), UBD: 02-OCT-2026, UDI#: (B)(4) SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22923041&pc=NPT | 9617601-2025-01231 | 2025-08-27 04:00 | Death | MEDTRONIC MEXICO S. DE R.L. DE CV | 2025-08-29 04:00 | NPT | EVOLUT FX VALVE | Difficult or Delayed Positioning; Difficult to Advance; Device Dislodged or Dislocated; Patient Device Interaction Problem | Cardiac Arrest | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE IN A PATIENT WITH PRE-EXISTING SEVERE AORTIC STENO SIS, AN AORTIC VALVE AREA OF 0.6 CM SQUARED, AND AN EJECTION FRACTION OF 20%, A PRE-IMPLANT BALLOON DILATION WAS PERFORMED WITH A 25 MM X 50 MM BALLOON AT NOMINAL VOLUME. THE FIRST IMPLANTATION ATTEMPT WAS VERY DIFFICULT TO POSITION, BUT THE VALVE WAS RELEASED AT A DEPTH OF 3-5 MM WITH NON-MEDTRONIC (SAFARI) GUIDEWIRE RETRACTION, TEMPORARY PACING, AND SLOW MOVEMENTS. AFTER A PROJECTION CHANGE, DISLODGE OF THE VALVE INTO THE LEFT VENTRICLE WAS NOTED. THE VALVE WAS LOCATED AT A HEIGHT OF APPROXIMATELY 20 MM WITHIN THE LEFT VENTRICLE, WITH MODERATE PERIVALVULAR LEAK. A 15 MM SNARE WAS USED ON BOTH PADDLES TO TRACTION THE VALVE. THE VALVE WAS RETURNED TO A 3 MM POSITION AND POST-DILATED. A POST-IMPLANT BALLOON DILATION WITH A 28 MM X 50 MM BALLOON WAS PERFORMED TWICE, BUT THERE WAS NO IMPROVEMENT. AFTER DEFLATING THE BALLOON AND MAINTAINING TRACTION ON THE SNARE, THE VALVE DISLODGED OUT OF THE ASCENDING AORTA AGAIN. THE PATIENT DEVELOPED A FIRST CARDIAC ARREST LASTING 3 MINUTES, WHICH WAS THEN REVERSED. AFTERWARD, AN ATTEMPT WAS MADE TO IMPLANT A SECOND PROSTHESIS, BUT THE DELIVERY DEVICE WAS VERY DIFFICULT TO PASS THROUGH THE FIRST PROSTHESIS. THIS IS DUE TO THE LOCATION THE FIRST VALVE WAS SNARED TO. DURING THE ATTEMPT TO IMPLANT THE SECOND PROSTHESIS, THE PATIENT DEVELOPED ANOTHER CARDIAC ARREST. THE DECISION WAS MADE TO RELEASE THE SECOND VALVE URGENTLY AND ADMINISTER CARDIOPULMONARY RESUSCITATION. AFTER CIRCULATION RETURNED, THE SECOND PROSTHESIS WAS AGAIN MOVED INTO THE LEFT VENTRICLE, WHERE ATTEMPTS WERE MADE TO REPOSITION IT WITH A SNARE, BUT WITHOUT SUCCESS. THE PATIENT WENT INTO CARDIAC ARREST AGAIN AND DIED. OF NOTE, DURING THE PROCEDURE THE PATIENT REQUIRED HIGH DOSE VASOPRESSOR MEDICATION AND MECHANICAL VENTILATION. Manufacturer Narrative: CONTINUATION OF D10: SECTION D INFORMATION REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER RELEVANT DEVICE(S) ARE: UDI#: (B)(4); PRODUCT ID: D-EVOLUTFX-34, SERIAL/LOT #: UNKNOWN, UBD: 02-OCT-2026, SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22923327&pc=NPT | 2015691-2025-07131 | 2025-08-05 04:00 | Death | EDWARDS LIFESCIENCES | 2025-08-29 04:00 | NPT | EDWARDS COMMANDER DELIVERY SYSTEM, 23MM | Difficult to Advance |  | P140031 |  | 1 | Event Description: AS REPORTED BY THE EDWARDS FIELD CLINICAL SPECIALIST, DURING A TRANSFEMORAL TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE WITH A 23 SAPIEN 3 ULTRA RESILIA VALVE, AN ANNULAR RUPTURE AND AORTIC DISSECTION OCCURRED DURING VALVE DEPLOYMENT. THERE WAS DIFFICULTY CROSSING THE NATIVE ANNULUS DUE TO A TIGHT VALVE WITH EXTENSIVE CALCIUM. THE GUIDEWIRE POSITION WAS MAINTAINED, AND THE VALVE WAS SUCCESSFULLY DEPLOYED. FOLLOWING DEPLOYMENT, A CIRCUMFERENTIAL PERICARDIAL EFFUSION WAS OBSERVED WHEN THE PATIENT'S BLOOD PRESSURE DROPPED. THE PATIENT WAS INTUBATED AND PERICARDIOCENTESIS WAS PERFORMED, AND BLEEDING COULD NOT BE CONTROLLED. THE PATIENT WAS NOT CONSIDERED AN OPERABLE CANDIDATE. PER REPORT, THE HEAVY CALCIUM LIKELY CONTRIBUTED TO THE RUPTURE DURING DEPLOYMENT. THE CROSSING DIFFICULTY WITH THE PUSHER BACK AND EXPOSED CAUSED THE DISSECTION IN THE ASCENDING. THE CALCIUM LIKELY CAUSED THE RUPTURE UPON INFLATION. THERE WAS NO DEVICE DAMAGE OBSERVED UPON REMOVAL. THE PATIENT EXPIRED ON THE DAY OF THE PROCEDURE DUE TO ANNULAR RUPTURE AND AORTIC DISSECTION. Manufacturer Narrative: THE INVESTIGATION IS ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22923434&pc=FSA | 2183887-2025-00005 |  | Death | EZ WAY INC | 2025-08-29 04:00 | FSA | LIFT | Insufficient Information | Insufficient Information |  |  | 1 | Event Description: RESIDENT FELL OUT OF SLING. SLING WAS NOT HOOKED UP PROPERLY. RESIDENT SUSTAINED L5 AND L7 FRACTURE ALONG WITH A BAD HEAD BRUISE. Manufacturer Narrative: EZ WAY RECEIVED THIS INFORMATION FROM A LAWYER REPRESENTING THE FAMILY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22923613&pc=NIQ | 2024168-2025-09255 | 2025-05-24 04:00 | Death | ABBOTT VASCULAR INC. | 2025-08-29 04:00 | NIQ | XIENCE XPEDITION¿ | Adverse Event Without Identified Device or Use Problem | Cardiogenic Shock; Pseudoaneurysm | P110019 |  | 1 | Event Description: IT WAS REPORTED THAT THE PROCEDURE WAS TO TREAT THE DISTAL LEFT ANTERIOR DESCENDING (LAD) CORONARY ARTERY WITH MODERATE CALCIFICATION, HEAVY TORTUOSITY AND 90% STENOSIS. THE 3.0X38 MM XIENCE XPEDITION STENT WAS IMPLANTED ON (B)(6) 2025. A PSEUDOANEURYSM OCCURRED AND THE PATIENT EXPERIENCED CARDIOGENIC SHOCK AND WAS ON DUAL INOTROPES. THE PATIENT EXPIRED ON (B)(6) 2025. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22923875&pc=NIQ | 2024168-2025-09258 | 2025-08-06 04:00 | Death | ABBOTT VASCULAR INC. | 2025-08-29 04:00 | NIQ | XIENCE XPEDITION¿ | Adverse Event Without Identified Device or Use Problem | Pseudoaneurysm | P110019 |  | 1 | Event Description: IT WAS REPORTED THAT THE PROCEDURE WAS TO TREAT THE DISTAL LEFT ANTERIOR DESCENDING (LAD) CORONARY ARTERY WITH MODERATE CALCIFICATION, HEAVY TORTUOSITY AND 90% STENOSIS. THE 3.0X38 MM XIENCE XPEDITION STENT WAS IMPLANTED ON (B)(6) 2025. A PSEUDOANEURYSM WAS REPORTED AND IT WAS ATTEMPTED TO INCREASE THE TIMI FLOW. THE PATIENT EXPIRED ON (B)(6) 2025. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22923941&pc=DSQ | 2916596-2025-05595 | 2025-06-18 04:00 | Death | THORATEC CORPORATION | 2025-08-29 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS FAILURE TO THRIVE. THE PATIENT OPTED TO GO ON HOSPICE SECONDARY TO FAILURE TO THRIVE. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED AND THE DEVICE OPERATED AS EXPECTED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22924077&pc=DSQ | 2916596-2025-05641 | 2025-08-15 04:00 | Death | THORATEC CORPORATION | 2025-08-29 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Septic Shock | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED DUE TO SEPTIC SHOCK. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22924099&pc=KDJ | 3030306055-2025-00227 | 2025-07-18 04:00 | Death | VANTIVE US HEALTHCARE LLC | 2025-08-29 04:00 | KDJ | NI | Adverse Event Without Identified Device or Use Problem | Peritonitis | NI |  | 1 | Event Description: A PERITONEAL DIALYSIS (PD) PATIENT EXPERIENCED PERITONITIS. THE CAUSE OF THE PERITONITIS WAS UNKNOWN. THE SAME DAY AS EVENT ONSET, THE PATIENT WAS HOSPITALIZED AND DISCHARGED AFTER 21DAYS. THE PATIENT WAS TREATED WITH UNSPECIFIED ANTIBIOTICS FOR PERITONITIS. IT WAS REPORTED THE PATIENT EXPERIENCED SEPTIC SHOCK AND PASSED AWAY ON AN UNSPECIFIED DATE. THE CAUSE OF DEATH WAS DUE TO "SEPTIC SHOCK OF 24HRS". IT IS NOT REPORTED IF AN AUTOPSY WAS PERFORMED. IT WAS NOT REPORTED IF THE PERITONITIS WAS RESOLVED PRIOR TO DEATH. IT WAS NOT REPORTED IF PD THERAPY WAS ONGOING PRIOR TO DEATH OR AT THE TIME OF DEATH. NO ADDITIONAL INFORMATION WAS AVAILABLE. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED, AND THE LOT NUMBER IS UNKNOWN; THEREFORE, A DEVICE ANALYSIS COULD NOT BE COMPLETED. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22924222&pc=MVK | 3008642652-2025-08378 | 2025-08-05 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-08-29 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED AN INAPPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 13:16:34 ON (B)(6) 2025. AT 15:13:44, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH MOTION/TACTILE ARTIFACT AND ELECTRODE LEAD FALLOFF. AT 15:14:48, THE PATIENT RECEIVED THE INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH MOTION/TACTILE ARTIFACT AND ELECTRODE LEAD FALLOFF. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON-LIFE-SUSTAINING RHYTHM, WITH INTERMITTENT CARDIAC ACTIVITY, MOTION/TACTILE ARTIFACT AND ELECTRODE LEAD FALLOFF. THE DEVICE WAS SHUT DOWN AT 15:19:55 ON (B)(6) 2025. Manufacturer Narrative: THE ELECTRODE BELT AND MONITOR HAVE NOT BEEN RECOVERED. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22924407&pc=MVK | 3008642652-2025-08380 | 2025-04-15 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-08-29 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED THREE INAPPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 17:02:32 ON 4/14/2025. AT 05:22:05 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH CPR/INTERMITTENT CARDIAC ACTIVITY. AT 05:24:30, THE PATIENT RECEIVED THE FIRST INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS ASYSTOLE WITH CPR/MOTION ARTIFACT. AT 05:25:02, THE PATIENT RECEIVED THE SECOND INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. AT 05:25:50, THE PATIENT RECEIVED THE THIRD INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. POST SHOCK RHYTHM ECG IS NOT AVAILABLE. THE ELECTRODE BELT WAS DISCONNECTED AT 05:29:01 ON 4/15/2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22924743&pc=NPT | 2015691-2025-07145 | 2025-08-10 04:00 | Death | EDWARDS LIFESCIENCES | 2025-08-29 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Rupture | P140031 |  | 1 | Event Description: AS REPORTED, IT WAS A CASE OF AN IMPLANT OF A 29MM SAPIEN 3 ULTRA RESILIA VALVE, IN AORTIC POSITION BY TRANSFEMORAL APPROACH. DURING THE PROCEDURE, AN ANNULAR RUPTURE WAS DETECTED AFTER VALVE DEPLOYMENT, AND CARDIOPULMONARY RESUSCITATION AND SURGICAL THORACOTOMY WERE PERFORMED. UNFORTUNATELY, THE PATIENT PASSED AWAY. THE PERCEIVED ROOT CAUSE OF THIS EVENT MIGHT HAVE BEEN AN APPROPRIATE VALVE SIZE. HOWEVER, THE ANNULUS AREA (652MM2) IS IN RANGE OF A 29MM VALVE (540-680MM2). Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. ACCORDING TO THE DEVICE TRAINING MANUALS, RISK FACTORS FOR AORTIC DISSECTION, CARDIAC/AORTIC HEMATOMA, OR ANNULAR RUPTURE DURING THE TVR PROCEDURE INCLUDE SIGNIFICANT VALVE OVER SIZING, SEVERELY OBLITERATED SINUSES OF VALSALVA, PORCELAIN AORTA AND/OR PRESENCE OF BULKY CALCIFICATION, AND NARROW CALCIFIED SINOTUBULAR JUNCTION. IN ADDITION, ADVANCED AGE, FEMALE GENDER, SMALL BODY WEIGHT, AND STEROID DEPENDENCY CAN ALSO BE CONTRIBUTING FACTORS. THE SAPIEN TRANSCATHETER HEART VALVE (ALL MODELS) RELIES CALCIUM TO SECURELY ANCHOR IN THE LANDING ZONE. DESPITE THIS BENEFICIAL ASPECT OF CALCIUM, BULKY CALCIUM CAN INCREASE THE RISK OF CALCIFIC NODULE DISPLACEMENT INTO THE VASCULATURE, WHICH CAN LEAD TO VASCULAR INJURY. AT TIMES, THE EXTENT AND DISTRIBUTION OF CALCIUM CAN IMPAIR EASE OF DELIVERY OF THE VALVE, CORRECT POSITIONING OF THE VALVE, DEPLOYMENT OF THE VALVE, AND PROCEDURAL SUCCESS. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. BASED ON THE AVAILABLE INFORMATION, THE ROOT CAUSE OF THE EVENT REMAINS INDETERMINABLE. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22925624&pc=NVY | 2649622-2025-23798 | 2025-08-01 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-08-29 04:00 | NVY | SPRINT QUATTRO SECURE MRI SURESCAN | Failure to Capture; Under-Sensing | Cardiac Arrest | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED CARDIOPULMONARY ARREST LEADING TO LOSS OF CAPTURE FROM THE RIGHT VENTRICULAR (RV) LEAD DURING PACING. IT WAS NOTED THAT THE LOSS OF CAPTURE WAS LIKELY LARGELY DUE TO THE PATIENT¿S CONDITION FOLLOWING CARDIAC ARREST. IT WAS FURTHER REPORTED THAT LOSS OF CAPTURE THEN LED TO VENTRICULAR FIBRILLATION (VF) WHICH WAS UNDERSENSED BY THE RIGHT VENTRICULAR (RV) LEAD DUE TO VARIATION IN ELECTROGRAMS (EGM) WAVEFORMS AND THE DEVICE WAS UNABLE TO DETECT AND TREAT THE ARRHYTHMIA. EXTERNAL DEFIBRILLATION WAS ATTEMPTED HOWEVER THE PATIENT COULD NOT BE RESCUED AND DIED. Manufacturer Narrative: CONTINUATION OF D10: 5076-52 LEAD, IMPLANTED: (B)(6) 2012. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22925728&pc=QZI | 2124215-2025-60989 | 2025-08-06 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-08-29 04:00 | QZI | FARAWAVE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Inflammation; Perforation; Pericardial Effusion; Swelling/ Edema |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED A PERICARDIAL EFFUSION AND THERE WAS EDEMA, INFLAMMATION AND SWELLING. POST ABLATION PROCEDURE, THE PATIENT WAS LEAVING THE ROOM WHEN THE PATIENT CODED IN RECOVERY. CARDIOPULMONARY RESUSCITATION (CPR) AND PERICARDIOCENTESIS WERE PERFORMED. THE PATIENT WAS ADMITTED INTO THE INTENSIVE CARE UNIT (ICU). IT WAS FURTHER REPORTED THAT THE FARAWAVE CATHETER UPN AND LOT NUMBER IS NOT AVAILABLE. THE PHYSICIAN DID NOT HAVE AN ISSUE WITH THE CATHETER AND DOES NOT THINK IT WAS THE REASON THE PATIENT HAS POST PROCEDURE COMPLICATIONS. THE NUMBER OF LESIONS PERFORMED ON THE LEFT AND RIGHT ATRIUM WAS APPROXIMATELY 77 ABLATIONS. TRANSEPTAL WAS PERFORMED AND OTHER CATHETERS USED WERE NON-BOSTON SCIENTIFIC GRID CATHETER, CS CATHETER, AND ICE. THE NON-BOSTON SCIENTIFIC GRID CATHETER WAS INSERTED VIA THE FARADRIVE. INTRA HEPARIN WAS GIVEN AND THE ACTIVATED CLOT TIME (ACT) WAS GREATER THAN 300 AND WAS MAINTAINED. THE PATIENT WAS ON HOME OXYGEN PRIOR TO THE PROCEDURE, WHICH MAY HAVE CONTRIBUTED TO THE ADVERSE EVENT. THE PATIENT FAMILY TOOK THE PATIENT OFF LIFE SUPPORT AND THE PATIENT PASSED AWAY. Manufacturer Narrative: DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION, BUT IT WAS NOT AVAILABLE BECAUSE THE DEVICE WAS DISCARDED. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION. IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22925757&pc=NKM | 2135147-2025-04879 | 2015-04-01 04:00 | Death | ABBOTT MEDICAL | 2025-08-29 04:00 | NKM | MITRACLIP® SYSTEM CLIP DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P100009 |  | 1 | Event Description: THE ARTICLE "PROGNOSTIC SIGNIFICANCE OF MITRAL ANNULAR CALCIFICATION ASSESSED BY COMPUTED TOMOGRAPHY FOR TRANSCATHETER EDGE-TO-EDGE REPAIR OUTCOMES" WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE SINGLE CENTER STUDY, TO EVALUATE THE ASSOCIATION BETWEEN CT FEATURES OF MITRAL ANNULAR CALCIFICATION (MAC) AND CLINICAL OUTCOMES AFTER TRANSCATHETER EDGE-TO-EDGE REPAIR (TEER) AT 3 YEARS. DEVICES MENTIONED INCLUDE MITRACLIP. THE ARTICLE CONCLUDED THAT ALL-CAUSE MORTALITY DID NOT SIGNIFICANTLY DIFFER BETWEEN PATIENTS WITH VARIED MAC SEVERITY. HOWEVER, GREATER CALCIUM THICKNESS AND LEAFLET INVOLVEMENT WERE ASSOCIATED WITH WORSE CLINICAL OUTCOMES IN PATIENTS UNDERGOING TEER. DETAILED PREOPERATIVE CT EVALUATION CAN FACILITATE THE PREDICTION AND MANAGEMENT OF TEER OUTCOMES. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS RAJ MAKKAR AT CEDARS-SINAI MEDICAL CENTER, LOS ANGELES, CALIFORNIA, USA WITH CORRESPONDING EMAIL: RAJ.MAKKAR@CSHS.ORG. THE TIME FRAME OF THE STUDY WAS APRIL 2015 AND MARCH 2021. A TOTAL OF 391 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH ALL RECEIVED AN ABBOTT DEVICE. AVERAGE AGE WAS 74 AND MAJORITY SEX WAS MALE. COMORBIDITIES INCLUDED HYPERTENSION, PREVIOUS STROKE OR TIA, PORCELAIN AORTA, SMOKER, DIABETES, DIALYSIS, CHRONIC LUNG DISEASE, HEART FAILURE HOSPITALIZATION, PERIPHERAL VASCULAR DISEASE, PERMANENT PACEMAKER, CARDIAC RESYNCHRONIZATION THERAPY, ATRIAL FIBRILLATION, MYOCARDIAL INFARCTION, CABG, AORTIC VALVE REPLACEMENT, PERCUTANEOUS CORONARY INTERVENTION, DEGENERATIVE MITRAL REGURGITATION, FUNCTIONAL MITRAL REGURGITATION, AORTIC REGURGITATION, TRICUSPID REGURGITATION, MITRAL ANNULAR CALCIFICATION. COMPLICATIONS INCLUDED DEATH, HEART FAILURE HOSPITALIZATION, VALVE REINTERVENTION, SURGICAL INTERVENTION, RECURRENT MITRAL REGURGITATION, STROKE, MITRAL STENOSIS. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF MITRACLIP WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING HYPERTENSION, PREVIOUS STROKE OR TIA, PORCELAIN AORTA, SMOKER, DIABETES, DIALYSIS, CHRONIC LUNG DISEASE, HEART FAILURE, HOSPITALIZATION, PERIPHERAL VASCULAR DISEASE, PERMANENT PACEMAKER, CARDIAC RESYNCHRONIZATION THERAPY, ATRIAL FIBRILLATION, MYOCARDIAL INFARCTION, CABG, AORTIC VALVE REPLACEMENT, PERCUTANEOUS CORONARY INTERVENTION, DEGENERATIVE MITRAL REGURGITATION, FUNCTIONAL MITRAL REGURGITATION, AORTIC REGURGITATION, TRICUSPID REGURGITATION, MITRAL ANNULAR CALCIFICATION. COMPLICATIONS REPORTED INCLUDED DEATH, HEART FAILURE HOSPITALIZATION, VALVE REINTERVENTION, SURGICAL INTERVENTION, RECURRENT MITRAL REGURGITATION, STROKE, MITRAL STENOSIS, AND TRICUSPID INTERVENTION; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. B2: DEATH DATE IS ESTIMATED. B3: EVENT DATE IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22925775&pc=NPT | 2015691-2025-07150 | 2025-08-05 04:00 | Death | EDWARDS LIFESCIENCES | 2025-08-29 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA VALVE, 23MM | Adverse Event Without Identified Device or Use Problem | Rupture | P140031 |  | 1 | Event Description: AS REPORTED BY THE EDWARDS FIELD CLINICAL SPECIALIST, DURING A TRANSFEMORAL TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE WITH A 23 SAPIEN 3 ULTRA RESILIA VALVE, AN ANNULAR RUPTURE AND AORTIC DISSECTION OCCURRED DURING VALVE DEPLOYMENT. THERE WAS DIFFICULTY CROSSING THE NATIVE ANNULUS DUE TO A TIGHT VALVE WITH EXTENSIVE CALCIUM. THE GUIDEWIRE POSITION WAS MAINTAINED, AND THE VALVE WAS SUCCESSFULLY DEPLOYED. FOLLOWING DEPLOYMENT, A CIRCUMFERENTIAL PERICARDIAL EFFUSION WAS OBSERVED WHEN THE PATIENT'S BLOOD PRESSURE DROPPED. THE PATIENT WAS INTUBATED AND PERICARDIOCENTESIS WAS PERFORMED, AND BLEEDING COULD NOT BE CONTROLLED. THE PATIENT WAS NOT CONSIDERED AN OPERABLE CANDIDATE. PER REPORT, THE HEAVY CALCIUM LIKELY CONTRIBUTED TO THE RUPTURE DURING DEPLOYMENT. THE CROSSING DIFFICULTY WITH THE PUSHER BACK AND EXPOSED CAUSED THE DISSECTION IN THE ASCENDING BUT THE CALCIUM LIKELY CAUSED THE RUPTURE UPON INFLATION. THERE WAS NO DEVICE DAMAGE OBSERVED UPON REMOVAL, AND NO DEVICE DEFICIENCIES WERE ALLEGED. THE PATIENT EXPIRED ON THE DAY OF THE PROCEDURE DUE TO ANNULAR RUPTURE AND AORTIC DISSECTION. Manufacturer Narrative: THIS REPORT IS 2 OF 2 BEING SUBMITTED FOR THIS COMPLAINT. REFERENCE MFG. NUMBER 2015691-2025-07131. PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. ACCORDING TO THE DEVICE TRAINING MANUALS, RISK FACTORS FOR AORTIC DISSECTION, CARDIAC/AORTIC HEMATOMA, OR ANNULAR RUPTURE DURING THE TVR PROCEDURE INCLUDE SIGNIFICANT VALVE OVER SIZING, SEVERELY OBLITERATED SINUSES OF VALSALVA, PORCELAIN AORTA AND/OR PRESENCE OF BULKY CALCIFICATION, AND NARROW CALCIFIED SINOTUBULAR JUNCTION. IN ADDITION, ADVANCED AGE, FEMALE GENDER, SMALL BODY WEIGHT, AND STEROID DEPENDENCY CAN ALSO BE CONTRIBUTING FACTORS. THE SAPIEN TRANSCATHETER HEART VALVE (ALL MODELS) RELIES ON CALCIUM TO SECURELY ANCHOR IN THE LANDING ZONE. DESPITE THIS BENEFICIAL ASPECT OF CALCIUM, BULKY CALCIUM CAN INCREASE THE RISK OF CALCIFIC NODULE DISPLACEMENT INTO THE VASCULATURE, WHICH CAN LEAD TO VASCULAR INJURY. AT TIMES, THE EXTENT AND DISTRIBUTION OF CALCIUM CAN IMPAIR EASE OF DELIVERY OF THE VALVE, CORRECT POSITIONING OF THE VALVE, DEPLOYMENT OF THE VALVE, AND PROCEDURAL SUCCESS. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION THAT A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST PATIENT (A TIGHT NATIVE VALVE WITH EXTENSIVE CALCIUM; PATIENT AGE [95]) CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED CURRENTLY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22925905&pc=NPT | 2135147-2025-04883 | 2017-10-01 04:00 | Death | ABBOTT MEDICAL | 2025-08-29 04:00 | NPT | PORTICO TRANSCATHETER AORTIC VALVE | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P190023 |  | 1 | Event Description: THE ARTICLE, "HEAD-TO-HEAD COMPARISON OF MERIL MYVAL SERIES BALLOON-EXPANDABLE AND ABBOTT PORTICO SERIES SELF-EXPANDING TRANSCATHETER AORTIC VALVES¿A SINGLE-CENTER EXPERIENCE", WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE, SINGLE CENTER STUDY TO COMPARE THE 30-DAY AND 1-YEAR OUTCOMES FOLLOWING MYVAL BALLOON-EXPANDABLE VALVE (BEV) (MERIL LIFE SCIENCES) AND INTRA-ANNULAR PORTICO SELF-EXPANDING VALVE (SEV) (ABBOTT) IMPLANTATION. DEVICES INCLUDED IN THE STUDY WERE MERIL LIFE SCIENCES MYVAL AND ABBOTT PORTICO. THE ARTICLE CONCLUDED THAT THE MYVAL BEV WAS ASSOCIATED WITH COMPARABLE SHORT- AND MID-TERM OUTCOMES AS THE PORTICO SEV. [(B)(6)]. THE TIME FRAME OF THE STUDY WAS FROM OCTOBER 2017 TO AUGUST 2023. A TOTAL OF 144 PATIENTS WERE INCLUDED IN THE STUDY, OF WHICH 47 (32.6%) RECEIVED AN ABBOTT DEVICE. IN THE MYVAL GROUP, THE AVERAGE AGE WAS 81.0 YEARS AND THE MAJORITY GENDER WAS MALE. IN THE PORTICO GROUP, THE AVERAGE WAS 82.9 YEARS AND THE MAJORITY GENDER WAS FEMALE. COMORBIDITIES INCLUDED AORTIC STENOSIS, AORTIC REGURGITATION, HEART BLOCK, AND ATRIAL FIBRILLATION. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF HEAD-TO-HEAD COMPARISON OF MERIL MYVAL SERIES BALLOON-EXPANDABLE AND ABBOTT PORTICO SERIES SELF-EXPANDING TRANSCATHETER AORTIC VALVES WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING AORTIC STENOSIS, AORTIC REGURGITATION, HEART BLOCK, AND ATRIAL FIBRILLATION. SOME OF THE COMPLICATIONS REPORTED WERE CARDIAC TAMPONADE, ATRIAL FIBRILLATION, HEMORRHAGE, AORTIC VALVE INSUFFICIENCY; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. B2: DATE OF DEATH IS ESTIMATED. B3: DATE OF EVENT IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE REVIEW: HEAD-TO-HEAD COMPARISON OF MERIL MYVAL SERIES BALLOON-EXPANDABLE AND ABBOTT PORTICO SERIES SELF-EXPANDING TRANSCATHETER AORTIC VALVES¿A SINGLE-CENTER EXPERIENCE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22925954&pc=OZD | 1220648-2025-46057 | 2025-07-19 04:00 | Death | ABIOMED, INC. - 1220648 | 2025-08-29 04:00 | OZD | IMPELLA 5.5 | Crack; Fluid/Blood Leak; Decrease in Pressure | Atrial Fibrillation | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT ENDURED ATRIAL FIBRILLATION DURING IMPELLA 5.5 SUPPORT. THE PATIENT INITIALLY CONVERTED BACK TO A NORMAL SINUS RHYTHM, HOWEVER THEY ENDURED ANOTHER EPISODE THE FOLLOWING NIGHT. AN AMIO DRIP WAS STARTED TO COUNTERACT THE CONDITION. ROUGHLY TWELVE DAYS LATER, A CRACK WAS DISCOVERED IN THE PURGE SIDEARM. FLUID WAS SEEN LEAKING FROM THE LOCATION AND A LOW PURGE PRESSURE ALARM APPEARED. NO INTERVENTION WAS TAKEN AND SUPPORT CONTINUED FOR AN ADDITIONAL TWO DAYS BEFORE THE PATIENT PASSED AWAY. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: A5 AND A6 ARE UNKNOWN. THE CAUSE OF THE PURGE LEAK WAS MOST LIKELY A DAMAGED SIDEARM BUT THE LOCATION ON THE SIDEARM IS UNKNOWN SINCE PRODUCT WAS NOT RETURNED FOR REVIEW. THE ROOT CAUSE OF THE ATRIAL FIBRILLATION WAS UNABLE TO BE DETERMINED DUE TO INSUFFICIENT CLINICAL DETAILS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22926006&pc=NPT | 2135147-2025-04890 | 2018-12-01 05:00 | Death | ABBOTT MEDICAL | 2025-08-29 04:00 | NPT | NAVITOR TRANSCATHETER AORTIC VALVE | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P190023 |  | 1 | Event Description: THE ARTICLE, "INTEGRATED SEVERITY STAGING OF PARAVALVULAR REGURGITATION AFTER TRANSCATHETER AORTIC VALVE IMPLANTATION USING ECHOCARDIOGRAPHIC AND HEMODYNAMIC ASSESSMENT", WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE, SINGLE CENTER STUDY ON EFFECTIVENESS OF A COMPREHENSIVE INTEGRATED APPROACH THAT COMBINES BOTH METHODOLOGIES TO ACCURATELY IDENTIFY AND PREDICT THE CLINICAL CONSEQUENCES OF PARAVALVULAR REGURGITATION (PVR). DEVICES INCLUDED IN THE STUDY WERE PORTICO, NAVITOR, AND EVOLUT R. THE ARTICLE CONCLUDED THAT INVASIVE ASSESSMENT OF PVR EFFECTIVELY IDENTIFIES PATIENTS AT RISK FOR EARLY MAJOR ADVERSE CARDIOVASCULAR EVENTS (MACE), PARTICULARLY THOSE WITH MODERATE OR SEVERE PVR. HOWEVER, ECHOCARDIOGRAPHY SIGNIFICANTLY ENHANCES RISK STRATIFICATION IN ALL OTHER CASES. THEREFORE, AN INTEGRATED APPROACH COMBINING BOTH INVASIVE AND NON-INVASIVE MODALITIES MAY REPRESENT THE MOST APPROPRIATE STRATEGY FOR IDENTIFYING PATIENTS AT INCREASED RISK OF ADVERSE OUTCOMES FOLLOWING TRANSCATHETER AORTIC VALVE IMPLANTATION (TAVI). [(B)(6)] THE TIME FRAME OF THE STUDY WAS FROM DECEMBER 2018 TO DECEMBER 2022. A TOTAL OF 126 PATIENTS WERE INCLUDED IN THE STUDY, OF WHICH 114 (90.5%) RECEIVED AN ABBOTT DEVICE. THE AVERAGE AGE WAS 82.1 YEARS AND THE MAJORITY GENDER WAS FEMALE. COMORBIDITIES INCLUDED HYPERTENSION, DYSLIPIDEMIA, DIABETES MELLITUS, PRIOR MYOCARDIAL INFARCTION, PRIOR PERCUTANEOUS CORONARY INTERVENTION, PRIOR BYPASS SURGERY, PRIOR VALVE SURGERY, ATRIAL FIBRILLATION, PRIOR PACEMAKER/INTRACARDIAC DEFIBRILLATOR IMPLANT, PERIPHERAL ARTERY DISEASE, PULMONARY HYPERTENSION, AND CHRONIC KIDNEY DISEASE. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF INTEGRATED SEVERITY STAGING OF PARAVALVULAR REGURGITATION AFTER TRANSCATHETER AORTIC VALVE IMPLANTATION USING ECHOCARDIOGRAPHIC AND HEMODYNAMIC ASSESSMENT WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING HYPERTENSION, DYSLIPIDEMIA, DIABETES MELLITUS, PRIOR MYOCARDIAL INFARCTION, PRIOR PERCUTANEOUS CORONARY INTERVENTION, PRIOR BYPASS SURGERY, PRIOR VALVE SURGERY, ATRIAL FIBRILLATION, PRIOR PACEMAKER/INTRACARDIAC DEFIBRILLATOR IMPLANT, PERIPHERAL ARTERY DISEASE, PULMONARY HYPERTENSION, AND CHRONIC KIDNEY DISEASE. SOME OF THE COMPLICATIONS REPORTED WERE HEART FAILURE, AORTIC VALVE INSUFFICIENCY, STROKE; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. B2: DATE OF DEATH IS ESTIMATED. B3: DATE OF EVENT IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE REVIEW: INTEGRATED SEVERITY STAGING OF PARAVALVULAR REGURGITATION AFTER TRANSCATHETER AORTIC VALVE IMPLANTATION USING ECHOCARDIOGRAPHIC AND HEMODYNAMIC ASSESSMENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22926030&pc=NPT | 2135147-2025-04891 | 2025-08-01 04:00 | Death | ABBOTT MEDICAL | 2025-08-29 04:00 | NPT | FLEXNAV DELIVERY SYSTEM | Difficult to Insert; Difficult to Advance; Positioning Problem | Hematoma; Hemorrhage/Bleeding | P190023 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 25 MM NAVITOR VISION VALVE WAS CHOSEN FOR IMPLANTATION UTILIZING A SMALL FLEXNAV DELIVERY SYSTEM. PRE-PROCEDURE, THE PATIENT PRESENTED IN SEVERE HEMODYNAMIC CONDITION HOWEVER, WAS STABLE WITH NO MAJOR COMPLICATIONS. THERE WAS DIFFICULTY OBTAINING VASCULAR ACCESS BUT WAS EVENTUALLY SUCCESSFUL. AFTER PUNCTURE OF THE LEFT FEMORAL ARTERY, THE PATIENT DEVELOPED A SUPERFICIAL HEMATOMA. PRE-DILATATION WAS PERFORMED WITH A 20 MM CRISTAL BALLOON. DUE TO POOR INITIAL POSITIONING, THE VALVE WAS RECAPTURED THREE TIMES BEFORE BEING IMPLANTED. DURING THIS THE PATIENT DEVELOPED LEFT BUNDLE BRANCH BLOCK (LBBB). THE VALVE WAS IMPLANTED UNDER PACING AT A DEPTH OF 3-4 MM. AFTER DEPLOYMENT, PERIVALVULAR LEAK (PVL) WAS PRESENT. POST IMPLANT DILATATION WAS PERFORMED WITH A 23 MM CRISTAL BALLOON. THE PATIENT LEFT THE OPERATING ROOM WITH A PACEMAKER DUE TO THE LBBB. PATIENT WAS TRANSFERRED TO INTENSIVE CARE UNIT (ICU) PER HOSPITAL PROTOCOL. 24 HOURS POST PROCEDURE, THE PATIENT DIED FROM PERIPHERAL HEMORRHAGE DUE TO PERIPHERAL ACCESS. IN THE PHYSICIAN'S OPINION, THE DEATH WAS UNRELATED TO THE NAVITOR SYSTEM AND WAS ATTRIBUTED TO THE PATIENTS PRE-EXISTING SEVERE HEMODYNAMIC CONDITION. Manufacturer Narrative: N/A. IT IS UNKNOWN IF THE DEVICE IS RETURNING FOR ANALYSIS. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22926036&pc=MIH | 3013164176-2025-02657 | 2025-04-21 04:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-08-29 04:00 | MIH | GORE® EXCLUDER® AAA ENDOPROSTHESIS | Patient-Device Incompatibility | Ischemia; Respiratory Failure; Embolism/Embolus; Renal Impairment | P020004 |  | 1 | Event Description: THE FOLLOWING INFORMATION WAS REPORTED THROUGH THE AAA 24-01 TAMBE POST APPROVAL STUDY: ON (B)(6) 2025, THE PATIENT UNDERWENT ENDOVASCULAR TREATMENT OF A THORACOABDOMINAL AORTIC ANEURYSM USING A GORE® EXCLUDER® THORACOABDOMINAL BRANCH ENDOPROSTHESIS (TAMBE) DEVICE SYSTEM. AN ATAA43720160A TAMBE AORTIC COMPONENT, CEB231410A GORE® EXCLUDER® ILIAC BRANCH ENDOPROSTHESIS ILIAC BRANCH COMPONENT, AND TWO GORE® EXCLUDER® AAA ENDOPROSTHESIS CONTRALATERAL LEG COMPONENTS (PLC141000 AND PLC141200) WERE IMPLANTED. A CSF DRAIN AND INTRAOPERATIVE NEUROPHYSIOLOGICAL MONITORING (MEP/SSEP) WERE UTILIZED PROPHYLACTICALLY DURING THE PROCEDURE. ON THE SAME DAY, CHOLESTEROL EMBOLI SYNDROME, ACUTE KIDNEY INJURY (AKI), AND RIGHT LOWER EXTREMITY ISCHEMIA WERE REPORTED. THE CHOLESTEROL EMBOLI SYNDROME WAS NOTED AS DISEASE RELATED. THE AKI WAS NOTED AS NOT RELATED TO THE DEVICE, PROCEDURE, OR DISEASE. THE PATIENT WAS TAKEN BACK TO THE OPERATING ROOM FOR RIGHT COMMON FEMORAL ARTERY THROMBOENDARTERECTOMY, RIGHT FEMORAL ARTERY INTERPOSITION GRAFT PLACEMENT, AND THROMBECTOMY OF THE RIGHT FEMOROPOPLITEAL ARTERY. THE ISCHEMIA WAS NOTED AS RESOLVED ON THE SAME DAY. ON (B)(6) 2025, THE PATIENT ENTERED RESPIRATORY FAILURE. IT WAS NOTED THAT THE RESPIRATORY FAILURE REQUIRED PROLONGED (>24 HOUR) MECHANICAL VENTILATION OR REINTUBATION. THE RESPIRATORY FAILURE WAS NOTED AS NOT RELATED TO THE DEVICE, PROCEDURE, OR DISEASE. THE PATIENT WAS PLACED ON CONTINUOUS RENAL REPLACEMENT THERAPY (NEW ONSET DIALYSIS) FOR TREATMENT OF THE AKI ON (B)(6) 2025. ON (B)(6) 2025, THE PATIENT EXPIRED. THE CAUSE OF DEATH WAS NOTED AS RELATED TO THE RESPIRATORY FAILURE AND CHOLESTEROL EMBOLI SYNDROME. Manufacturer Narrative: ADDITIONAL DEVICES INCLUDED ON THIS REPORT ARE AS FOLLOWS: CATALOG #PLC141000/ SERIAL #(B)(6)/ UDI #(B)(4) AS A COMPONENT OF THE SAME SYSTEM. CATALOG #ATAA43720160A/ SERIAL #(B)(6)/ UDI #(B)(4) WHICH IS CAPTURED IN MANUFACTURER REPORT #2017233-2025-06209. CATALOG #CEB231410A/ SERIAL #(B)(6)/ UDI #(B)(4) WHICH IS CAPTURED IN MANUFACTURER REPORT #3013164176-2025-02530. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22926044&pc=OZD | 1220648-2025-45974 | 2025-07-29 04:00 | Death | ABIOMED, INC. | 2025-08-29 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT IN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA 5.5 DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. IT WAS REPORTED THAT THE PATIENT EXPERIENCED ECTOPY DURING IMPELLA 5.5 SUPPORT. THE PATIENT WAS INITIALLY IMPLANTED WITH AN IMPELLA CP DEVICE (SERIAL NUMBER (B)(6), HOWEVER WITHIN THE FIRST TWENTY MINUTES OF SUPPORT A HIGH PURGE PRESSURE ALARM OCCURRED WHICH COULD NOT BE RESOLVED. THE FIRST IMPELLA CP PUMP WAS REPLACED WITH A SECOND IMPELLA CP DEVICE (SERIAL NUMBER (B)(6). AFTER 1 DAY OF SUPPORT, THE MEDICAL TEAM DECIDED TO ESCALATE THE PATIENT TO AN IMPELLA 5.5 (SERIAL NUMBER (B)(6). UPON EXPLANATION OF THE SECOND IMPELLA CP, IT WAS REPORTED THAT THE PATIENT EXPERIENCED A VASCULAR INJURY AT THE IMPELLA CP ACCESS SITE AFTER A PERCLOSE DEVICE FAILED. A COVERED STENT WAS PLACED TO RESOLVE THE BLEEDING. IT WAS REPORTED DURING IMPELLA 5.5 SUPPORT, ON (B)(6) 2025, THERE WAS AN IMPELLA POSITIONING ISSUE WHICH CAUSED THE PATIENT TO EXPERIENCE ECTOPY. THE MEDICAL TEAM WAS ADVISED TO REPOSITION THE IMPELLA, WHICH RESOLVED THE ECTOPY. IT WAS ADDITIONALLY NOTED THAT THE PATIENT¿S CARE WAS WITHDRAWN, AND THE PATIENT EXPIRED. THIS COMPLAINT INVOLVES THREE IMPELLA DEVICES: PUMP 1: IMPELLA CP, WITH SERIAL NUMBER (B)(6). PUMP 2: IMPELLA CP, WITH SERIAL NUMBER (B)(6). PUMP 3: IMPELLA 5.5, WITH SERIAL NUMBER (B)(6). THIS MDR SPECIFICALLY ADDRESSES THE IMPELLA 5.5 WITH SERIAL NUMBER (B)(6), WHICH IS THE SUBJECT OF THIS REPORT. SEPARATE MDRS WILL BE SUBMITTED FOR THE OTHER DEVICES ASSOCIATED WITH THIS COMPLAINT. Manufacturer Narrative: THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S DEMISE. PRODUCT STATUS: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22926064&pc=OZD | 1220648-2025-46022 | 2025-08-01 04:00 | Death | ABIOMED, INC. - 1220648 | 2025-08-29 04:00 | OZD | AUTOMATED IMPELLA CONTROLLER | Application Program Problem; Device Sensing Problem | No Clinical Signs, Symptoms or Conditions | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT IMPLANTED WITH AN IMPELLA CP LOST THE OPTICAL SIGNAL. THE SUPPORTING AUTOMATED IMPELLA CONTROLLER (AIC) WAS REPLACED TO RESOLVED THE ISSUE. LATER, THE PATIENT WAS DECLARED DO NOT RESUSCITATE AND EXPIRED. Manufacturer Narrative: A4 IS UNKNOWN. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22926708&pc=OZD | 1220648-2025-46044 | 2025-08-02 04:00 | Death | ABIOMED, INC. - 1220648 | 2025-08-29 04:00 | OZD | IMPELLA CP | Malposition of Device | Tachycardia; Hematuria | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT THE IMPELLA CP WAS PLACED IN THE SETTING OF ACUTE MYOCARDIAL INFARCTION AND CARDIOGENIC SHOCK. IN ICU THE PATIENT DEVELOPED RED TINGED URINE AND LACTATE DEHYDROGENASE WAS HIGH. IN RESPONSE, THE IMPELLA WAS REPOSITIONED. ADDITIONALLY, THE PATIENT HAD AN EPISODE OF VENTRICULAR TACHYCARDIA AND AMIODARONE CONVERSION WAS ADMINISTERED. IT WAS FURTHER REPORTED THAT THE FAMILY DECIDED TO WITHDRAW SUPPORT AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22926801&pc=OZD | 1220648-2025-46049 | 2025-08-04 04:00 | Death | ABIOMED, INC. - 1220648 | 2025-08-29 04:00 | OZD | IMPELLA CP | False Alarm; Restricted Flow rate; Malposition of Device | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT AN IMPLANTED IMPELLA CP DEVICE WAS POSITIONED TOO DEEP AND GENERATING ALARMS. DESPITE REPOSITIONING THE PUMP WOULD RETURN TO A DEEP POSITION. AN ECHOCARDIOGRAM REVEALED THE PATIENT HAD AORTIC STENOSIS CAUSING THE PUMP TO BE MALPOSITIONED. ALBUMIN WAS GIVEN FOR VOLUME. SUCTION ALARMS OCCURRED NECESSITATING A DECREASE IN P LEVEL TO P7. DURING EXPLANTATION OF THE PUMP IT WAS SUSPECTED THAT THE PATIENT HAD A RETROPERITONEAL BLEED AND MANUAL COMPRESSION WAS HELD. ULTIMATELY, THE PATIENT CODED AND EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THIS FILE REPRESENTS THE PUMP THE RELATED MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-46069 REPRESENTS THE INTRODUCER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22926864&pc=NPT | 2135147-2025-04908 | 2022-05-01 04:00 | Death | ABBOTT MEDICAL | 2025-08-29 04:00 | NPT | NAVITOR TRANSCATHETER AORTIC VALVE | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P190023 |  | 1 | Event Description: THE ARTICLE, "DIFFERENCES IN CORONARY ALIGNMENT BETWEEN EVOLUT FX AND NAVITOR VALVES", WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE, MULTICENTER STUDY TO ASSESS AND COMPARE CORONARY ACCESSIBILITY AMONG THE LATEST SELF-EXPANDABLE TRANSCATHETER HEART VALVE (THV) USING POST-TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) MULTIDETECTOR COMPUTED TOMOGRAPHY (MDCT). DEVICES INCLUDED IN THE STUDY WERE EVOLUT FX AND NAVITOR VALVES. THE ARTICLE CONCLUDED THAT CORONARY OVERLAP (CO) AND OSTIAL GEOMETRY RELATIVE TO THE THV FRAME DIFFERED SIGNIFICANTLY BETWEEN THE TWO SELF-EXPANDABLE VALVES (SEV) PLATFORMS. BALANCING THESE ANATOMIC AND DEVICE-SPECIFIC FACTORS MAY BE ESSENTIAL TO ENSURE FUTURE CORONARY ACCESS FOLLOWING TAVR. [THE PRIMARY AND CORRESPONDING AUTHOR WAS HIOKI HIROFUMI, DEPARTMENT OF CARDIOLOGY, IMS TOKYO KATSUSHIKA GENERAL HOSPITAL, TOKYO, JAPAN, WITH CORRESPONDING EMAIL: HIOKI.TEIKYO@GMAIL.COM] THE TIME FRAME OF THE STUDY WAS FROM MAY 2022 TO MAY 2024. A TOTAL OF 402 PATIENTS WERE INCLUDED IN THE STUDY, OF WHICH 152 (37.8%) RECEIVED AN ABBOTT DEVICE. IN THE SUPRA-ANNULAR (SA)-SEV GROUP, THE AVERAGE AGE WAS 85 YEARS AND THE MAJORITY GENDER WAS FEMALE. IN THE INTRA-ANNULAR (IA)-SEV, THE AVERAGE AGE WAS 86 YEARS AND THE MAJORITY GENDER WAS FEMALE. COMORBIDITIES INCLUDED HYPERTENSION, FRAILTY, DIABETES MELLITUS, DYSLIPIDEMIA, CHRONIC KIDNEY DISEASE, PERCUTANEOUS CORONARY INTERVENTION, PULMONARY DISEASE, ATRIAL FIBRILLATION, AND HEART BLOCK. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. B2: DATE OF DEATH IS ESTIMATED B3: DATE OF EVENT IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: DIFFERENCES IN CORONARY ALIGNMENT BETWEEN EVOLUT FX AND NAVITOR VALVES
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22926904&pc=NPT | 2015691-2025-07155 | 2025-08-08 04:00 | Death | EDWARDS LIFESCIENCES | 2025-08-29 04:00 | NPT | EDWARDS LIFESCIENCES SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Migration or Expulsion of Device | Ventricular Fibrillation | P140031 |  | 1 | Event Description: IT WAS REPORTED BY THE IMPLANT PATIENT REGISTRY THAT ON THE SAME DAY AS A TRANSFEMORAL TAVR PROCEDURE WITH A 23MM SAPIEN 3 ULTRA RESILIA VALVE THE PATIENT EXPIRED. AS REPORTED BY THE FIELD CLINICAL SPECIALIST (FCS), A VALVE IN VALVE WAS BEING DONE IN A 29 MM EVOLUT PRO+ TRANSCATHETER VALVE IN AORTIC POSITION TO TREAT STENOSIS. IMPLANTER WAS TRYING TO CANNULATE THE LEFT MAIN. THE PATIENT'S PRESSURE DROPPED AND EKG CHANGED. THE TRANSFEMORAL TAVR PROCEDURE WAS PERFORMED USING A 23 MM SAPIEN 3 ULTRA RESILIA VALVE. THE PATIENT STARTED CODING PRIOR TO THE EDWARDS VALVE INSERTION AND CPR WAS INITIATED. THE COMMANDER DELIVERY SYSTEM (DS) WAS PREPPED AND INSERTED. THE VALVE WAS POSITIONED AND DEPLOYED SUCCESSFULLY. THE PATIENT DID NOT HAVE A PULSE AT THE TIME OF INFLATION AND THE VALVE SHIFTED VENTRICULAR DURING INFLATION AND LANDED TOO LOW AT THE BOTTOM OF THE EVOLUT (THE PLAN WAS TO LAND AT NODE 4). A TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) WAS PERFORMED AND SEVERE PARAVALVULAR LEAK WAS NOTED AND IT WAS DECIDED TO INSERT A SECOND VALVE. THE SECOND VALVE 23 MM SAPIEN 3 ULTRA RESILIA VALVE AND DS WERE INSERTED AND DEPLOYED HIGHER THAN THE FIRST SAPIEN 3 ULTRA RESILIA VALVE. POST DEPLOYMENT OF THE SECOND VALVE, BOTH VALVES EMBOLIZED VENTRICULAR. THE PATIENT CONTINUED TO FIBRILLATE AND LOSE PULSE. CPR WAS PERFORMED THROUGHOUT THE CASE. X-RAY THEN SHOWED THAT BOTH 23MM SAPIEN 3 ULTRA RESILIA VALVES HAD EMBOLIZED IN THE LEFT VENTRICULAR A COUPLE MINUTES AFTER DEPLOYMENT. THE VALVES HAD SHIFTED DURING INFLATION. THE PERCEIVED CAUSE OF DEATH WAS REPORTED AS THE PATIENT CODING PRIOR TO THE EDWARDS VALVE INSERTION AND SEVERE AORTIC INSUFFICIENCY. Manufacturer Narrative: THE INVESTIGATION IS STILL ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22926999&pc=NPT | 2015691-2025-07156 | 2025-08-08 04:00 | Death | EDWARDS LIFESCIENCES | 2025-08-29 04:00 | NPT | EDWARDS LIFESCIENCES SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Migration or Expulsion of Device; Perivalvular Leak; Malposition of Device | Insufficient Information | P140031 |  | 1 | Event Description: IT WAS REPORTED BY THE IMPLANT PATIENT REGISTRY THAT ON THE SAME DAY AS A TRANSFEMORAL TAVR PROCEDURE WITH A 23MM SAPIEN 3 ULTRA RESILIA VALVE THE PATIENT EXPIRED. AS REPORTED BY THE FIELD CLINICAL SPECIALIST (FCS), A VALVE IN VALVE WAS BEING DONE IN A 29 MM EVOLUT PRO+ TRANSCATHETER VALVE IN AORTIC POSITION TO TREAT STENOSIS. IMPLANTER WAS TRYING TO CANNULATE THE LEFT MAIN. THE PATIENT'S PRESSURE DROPPED, AND EKG CHANGED. THE TRANSFEMORAL TAVR PROCEDURE WAS PERFORMED USING A 23 MM SAPIEN 3 ULTRA RESILIA VALVE. THE PATIENT STARTED CODING PRIOR TO THE EDWARDS VALVE INSERTION AND CPR WAS INITIATED. THE COMMANDER DELIVERY SYSTEM (DS) WAS PREPPED AND INSERTED. THE VALVE WAS POSITIONED AND DEPLOYED SUCCESSFULLY. THE PATIENT DID NOT HAVE A PULSE AT THE TIME OF INFLATION AND THE VALVE SHIFTED VENTRICULAR DURING INFLATION AND LANDED TOO LOW AT THE BOTTOM OF THE EVOLUT (THE PLAN WAS TO LAND AT NODE 4). A TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) WAS PERFORMED AND SEVERE PARAVALVULAR LEAK WAS NOTED AND IT WAS DECIDED TO INSERT A SECOND VALVE. THE SECOND VALVE 23 MM SAPIEN 3 ULTRA RESILIA VALVE AND DS WERE INSERTED AND DEPLOYED HIGHER THAN THE FIRST SAPIEN 3 ULTRA RESILIA VALVE. POST DEPLOYMENT OF THE SECOND VALVE, BOTH VALVES EMBOLIZED VENTRICULAR. THE PATIENT CONTINUED TO FIBRILLATE AND LOSE PULSE. CPR WAS PERFORMED THROUGHOUT THE CASE. X-RAY THEN SHOWED THAT BOTH 23MM SAPIEN 3 ULTRA RESILIA VALVES HAD EMBOLIZED IN THE LEFT VENTRICULAR A COUPLE MINUTES AFTER DEPLOYMENT. THE VALVES HAD SHIFTED DURING INFLATION. THE PERCEIVED CAUSE OF DEATH WAS REPORTED AS THE PATIENT CODING PRIOR TO THE EDWARDS VALVE INSERTION AND SEVERE AORTIC INSUFFICIENCY. Manufacturer Narrative: THE INVESTIGATION IS ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22927008&pc=LYJ | 1644487-2025-10566 | 2025-08-03 04:00 | Death | HOUSTON | 2025-08-29 04:00 | LYJ | PULSE GEN MODEL UNKNOWN | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P970003 |  | 1 | Event Description: DESCRIBE EVENT OR PROBLEM IT WAS REPORTED THAT AFTER 4 MONTHS WITH THE VNS, THE PATIENT PASSED AWAY. NO OTHER RELEVANT INFORMATION HAS BEEN RECEIVED TO DATE. Manufacturer Narrative: LIVANOVA USA, INC. SUBMITS THIS REPORT TO COMPLY WITH 21 CFR PART 803, THE MEDICAL DEVICE REPORTING REGULATION, BASED ON INFORMATION THAT LIVANOVA HAS OBTAINED, BUT MAY NOT HAVE BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION, OR A CONCLUSION BY FDA OR ANYONE ELSE, THAT THE DEVICE, LIVANOVA, OR LIVANOVA'S EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS" OR "MALFUNCTIONS". THESE WORDS ARE INCORPORATED INTO THE FDA 3500A MEDWATCH FORM BY THE FDA, AND LIVANOVA OBJECTS TO THEIR USE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22927018&pc=OZD | 1220648-2025-45754 | 2025-07-22 04:00 | Death | ABIOMED, INC. | 2025-08-29 04:00 | OZD | IMPELLA 5.5 | Restricted Flow rate | Cardiac Arrest; Sepsis | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WITH CARDIOMYOPATHY WAS IMPLANTED WITH AN IMPELLA 5.5 FOR MECHANICAL CIRCULATORY SUPPORT. IT WAS REPORTED THAT DURING IMPELLA 5.5 SUPPORT, THE PATIENT BECAME DIZZY AND ALERTED MEDICAL STAFF. THE PATIENT THEN WENT INTO CARDIAC ARREST AND REQUIRED CARDIOPULMONARY RESUSCITATION TO ACHIEVE A RETURN OF SPONTANEOUS CIRCULATION. IT WAS NOTED THAT THE AUTOMATED IMPELLA CONTROLLER DID NOT ALARM FOR SUCTION ALARMS WHICH OCCURRED JUST PRIOR TO THE PATIENT¿S CARDIAC ARREST. TWO DAYS AFTER THIS EVENT, THE PATIENT EXPIRED. IT WAS NOTED THIS WAS LIKELY DUE TO SEPTIC SHOCK. THERE WAS NO INFORMATION AVAILABLE ON THE SOURCE OF SEPSIS. THIS COMPLAINT INVOLVES ONE IMPELLA 5.5 PUMP (SERIAL NUMBER (B)(6)) AND ONE AUTOMATED IMPELLA CONTROLLER (SERIAL NUMBER (B)(6)). THIS MDR SPECIFICALLY ADDRESSES IMPELLA 5.5, WHICH IS THE SUBJECT OF THIS REPORT. A SEPARATE MDR WILL BE SUBMITTED FOR THE AUTOMATED IMPELLA CONTROLLER ASSOCIATED WITH THIS COMPLAINT. Manufacturer Narrative: BASED ON THE INFORMATION AVAILABLE, IT IS UNLIKELY THAT THE SEPSIS WAS RELATED TO THE USE OF THE IMPELLA DEVICE; HOWEVER, A DEVICE-RELATED CONTRIBUTION CANNOT BE DEFINITIVELY RULED OUT. PRODUCT STATUS: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22927033&pc=QBJ | 2954323-2025-34663 | 2025-07-25 04:00 | Death | ABBOTT DIABETES CARE INC | 2025-08-29 04:00 | QBJ | FREESTYLE LIBRE 2 PLUS | Device Displays Incorrect Message | Hypoglycemia; Loss of consciousness | K223435 |  | 1 | Event Description: A SENSOR-RELATED ERROR MESSAGE WAS REPORTED WITH THE ADC DEVICE. A CAREGIVER REPORTED THE CUSTOMER EXPERIENCED A SEVERE HYPOGLYCEMIC EVENT THE NIGHT OF (B)(6) 2025, DUE TO THE ADC SENSOR NO LONGER WORKING. THE CAREGIVER REPORTED THAT THE CUSTOMER LOST CONSCIOUSNESS, WAS HOSPITALIZED, AND RECEIVED INTRAVENOUS GLUCOSE ADMINISTERED BY A HEALTHCARE PROFESSIONAL. THE CAREGIVER FURTHER REPORTED THAT THE CUSTOMER DIED. THE DATE OF DEATH IS UNKNOWN, AND A CORONER IS REPORTEDLY INVOLVED IN THE CASE. ADC CUSTOMER SERVICE ATTEMPTED TO CONTACT THE CAREGIVER THREE TIMES TO GAIN ADDITIONAL DETAILS REGARDING THIS EVENT; HOWEVER, ALL FOLLOW-UP ATTEMPTS WERE UNSUCCESSFUL. NO FURTHER INFORMATION HAS BEEN RECEIVED AT THIS TIME. BASED ON THE INFORMATION CURRENTLY AVAILABLE, IT IS INCONCLUSIVE WHETHER THE ADC DEVICE HAS OR MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED DEATH. Manufacturer Narrative: THE PRODUCT HAS BEEN REQUESTED BACK FOR INVESTIGATION AND THE CUSTOMER AGREED TO RETURN THE DEVICE. A FOLLOW-UP REPORT WILL BE SUBMITTED UPON COMPLETION OF INVESTIGATION ACTIVITIES OR IF ADDITIONAL INFORMATION IS OBTAINED. ALL PERTINENT INFORMATION AVAILABLE TO ABBOTT DIABETES CARE HAS BEEN SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22927046&pc=QNR | 3003306248-2025-00249 | 2025-08-03 04:00 | Death | THORATEC SWITZERLAND GMBH | 2025-08-29 04:00 | QNR | CENTRIMAG BLOOD PUMP | Adverse Event Without Identified Device or Use Problem; Unexpected Shutdown | Low Blood Pressure/ Hypotension | K020271 |  | 1 | Event Description: IT WAS REPORTED THAT THE EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) CEASED OPERATION AT 2015 DUE TO SUSPECTED PUMP FAILURE OR CONSOLE FAILURE RESULTING IN A SIGNIFICANT DROP IN MEAN ARTERIAL PRESSURE (MAP) TO 20MMHG. IMMEDIATE INTERVENTION WAS INITIATED, TITRATING UP ON EPINEPHRINE AND A LEVOPHED (NOREPINEPHRINE) DRIP TO SUPPORT HEMODYNAMIC STABILITY. AFTER ASSESSING THE WHOLE ECMO CIRCUIT, IT APPEARED THAT ALL CABLE CONNECTIONS WERE THOROUGHLY VERIFIED TO BE SECURELY CONNECTED TO THE RED ALTERNATING CURRENT (AC) POWER SOURCE WITH NO LOOSE CONNECTIONS/KINK OBSERVED. EMERGENT MOTOR EXCHANGE WAS PERFORMED. WITHIN TWO MINUTES THE MOTOR AND CONSOLE WERE PROMPTLY SWITCHED AND REPLACED. ROTATIONS PER MINUTE (RPM)/FLOW WAS RESTORED. THE RPM WAS SET BACK TO THE PRESCRIBED PARAMETERS. FLOW AND MAP WAS REESTABLISHED AND STABILIZED AS WELL. THE PATIENT WAS IMMEDIATELY ABLE TO WEAN BACK DOWN ON PRESSORS. PATIENT TOLERATED INTERVENTION WELL AND HEMODYNAMIC STABILITY WAS RECOVERED. THE PATIENT THEN PASSED AWAY. Manufacturer Narrative: NO FURTHER INFO WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE REPORTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22927126&pc=LWS | 3004209178-2025-14697 | 2025-08-04 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-08-29 04:00 | LWS | EVERA MRI XT DR SURESCAN | Pacing Problem; Device Sensing Problem | Insufficient Information | P980016 |  | 1 | Event Description: IT WAS REPORTED THAT THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) DELIVERED ANTI-TACHYCARDIA PACING (ATP) FOR A DETECTED EPISODE OF VENTRICULAR TACHYCARDIA (VT). IT WAS NOTED THAT THE EPISODE MAY HAVE BEEN POSSIBLE ATRIAL TACHYCARDIA/ATRIAL FIBRILLATION (AT/AF) WITH RAPID VENTRICULAR RESPONSE. IT WAS FURTHER REPORTED THAT THREE DAYS POST THE EVENT, THE PATIENT PASSED AWAY. A CAUSE OF DEATH HAS BEEN REQUESTED BUT NOT RECEIVED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22927208&pc=NPT | 2135147-2025-04919 | 2025-08-06 04:00 | Death | ABBOTT MEDICAL | 2025-08-29 04:00 | NPT | NAVITOR VISION | Material Deformation; Patient Device Interaction Problem | Cardiac Arrest; Thrombosis/Thrombus | P190023 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 25MM NAVITOR VISION WAS CHOSEN FOR IMPLANTATION, UTILIZING A SMALL FLEXNAV. THE NAVITOR WAS PERFECTLY DEPLOYED WITH GOOD DEPTH, AND NO CORONARY OBSTRUCTION. POST IMPLANT, THE PATIENT DEVELOPED A THROMBUS. IT WAS NOTED THAT THE THROMBUS WAS NOT DUE TO THE NAVITOR DEVICE. THE PATIENT¿S HEART STOPPED, MANUAL COMPRESSIONS WERE PERFORMED FOR THIRTY MINUTES, THE PATIENT WAS PLACED ON EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) SYSTEM, AND STENTING OF THE INTERVENTRICULAR ARTERY (IVA), CIRCUMFLEX ARTERY (CX), AND MAIN LEFT CORONARY ARTERY WAS PERFORMED. IT WAS NOTED THAT THE STENT IN THE MAIN LEFT WAS CORKED. ON (B)(6) 2025, THE PATIENT DIED. THE CAUSE OF DEATH WAS DUE STOP FLOW IN MAIN LEFT CORONARY. IT WAS NOTED THAT THE PATIENT WAS IN GOOD CONDITION PRIOR TO THE IMPLANT PROCEDURE. IN THE PHYSICIAN'S OPINION, THE CAUSE OF DEATH WAS NOT DUE TO THE NAVITOR VALVE BUT WAS DUE TO THE PATIENT DID NOT TAKE HER MEDICATIONS BEFORE, WHICH CAUSED THE SITUATION OF THROMBUS IN THE ARTERIES. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22928606&pc=FOS | 2245270-2025-00103 | 2025-08-14 04:00 | Death | VYGON | 2025-08-29 04:00 | FOS | UMBILICAL CATHETER | Malposition of Device; Positioning Problem | Insufficient Information | K921374 |  | 1 | Event Description: THE SERIOUS ADVERSE EVENT OCCURRED ON A PREMATURE NEWBORN OF 28 WEEKS + 5, BORN IN ANOTHER HOSPITAL ON (B)(6) 2025 THAT THE HOSPITAL WHERE THE INCIDENT OCCURRED ON (B)(6) 2025. FOLLOWING A DETERIORATION OF HEALTH (CHILD ALREADY HOSPITALIZED IN NEONATAL INTENSIVE CARE AND INTUBATED /VENTILATED), ADDITIONAL EXAMINATIONS (ULTRASOUND AND X-RAY AT BEDSIDE) WERE CARRIED OUT AND A PERFUSOPERITONEUM WAS NOTICED. THIS PREMATURE HAD A CENTRAL VENOUS UMBILICAL CATHETER, INSERTED AT BIRTH AS AN EMERGENCY BY THE INITIAL HOSPITAL. THIS CATHETER HAS BEEN REMOVED BY 1 CM UPON ARRIVAL AT NEONATE'S UNIT OF THIS HOSPITAL. SEVERAL RADIOLOGICAL CONTROLS WERE PERFORMED SINCE HIS ARRIVAL. DESPITE RESUSCITATION WITH FLUID PUNCTURE, EXSUFFLATION AND CHEST DRAIN INSERTION, THIS PREMATURE BABY DIED ON (B)(6) 2025. Manufacturer Narrative: THE INVOLVED UMBILICAL CATHETER IS AVAILABLE; WE ARE WAITING FOR IT FOR THE INVESTIGATION. THE REVIEW OF BOTH POSSIBLE BATCHES DOES NOT SHOW ANY NON-CONFORMITY. THE PRODUCTS ARE IN CONFORMITY WITH THE SPECIFICATIONS. ALL CONTROLS ARE COMPLIANT, INCLUDING THE 100% TESTS CARRIED OUT ON EACH CATHETER DURING THE MANUFACTURING PROCESS. THESE TESTS ARE AS FOLLOWS: THE BLUNTNESS OF THE CATHETER TIP MARKING THE CATHETERS, TIGHTNESS AND NON-OBSTRUCTION. IN ADDITION, 25 UMBILICAL CATHETERS ARE TESTED BY SAMPLING: TO CHECK THE TENSILE STRENGTH ACCORDING TO ACCORDING TO NF EN ISO 10555-1 SPECIFICATION 10N. ALL RESULTS ARE COMPLIANT. TO CHECK THE ABSENCE OF LIQUID LEAKAGE (3BAR/30 SECONDS) AND AIR LEAKAGE WERE DONE, ACCORDING TO NF EN ISO 10555-1. ALL RESULTS COMPLIED (NO AIR OR LIQUID LEAKAGE). THE HISTORICAL DATA ANALYSIS OF COMPLAINTS ON THESE BATCHES DOES NOT SHOW ANY COMPLAINT. PERFUSOPERITONEUM IS A WELL-KNOWN COMPLICATION RELATED TO MALPOSITION OF THE UMBILICAL VENOUS CATHETER. THIS SERIOUS ADVERSE EVENT SEEMS NOT TO BE CAUSED BY A CATHETER MALFUNCTION BUT TRACED TO ITS USE. WE HAVE ASKED FOR ADDITIONAL INFORMATION, AND WE ARE WAITING FOR THEM WITH THE INVOLVED SAMPLE FOR INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22910518&pc=LPB | 2029046-2025-02894 | 2025-08-04 04:00 | Death | BIOSENSE WEBSTER INC | 2025-08-28 04:00 | LPB | EZ STEER NAV | Adverse Event Without Identified Device or Use Problem | Cardiac Perforation |  |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN ATRIAL FIBRILLATION (AFIB) CARDIAC ABLATION/MAPPING PROCEDURE WITH AN OCTARAY MAPPING CATHETER AND AN EZ STEER NAV CATHETER. THE PATIENT EXPERIENCED A PERICARDIAL EFFUSION TREATED WITH PERICARDIOCENTESIS, TRANSFERRED TO THE OPERATING ROOM (OR), AND THE PATIENT PASSED AWAY. ADDITIONALLY, DURING SURGERY A HOLE WAS FOUND IN THE CORONARY SINUS (CS) AND LEFT ATRIAL APPENDAGE (LAA). DURING THE PROCEDURE, A PERICARDIAL EFFUSION OCCURRED. THE PHYSICIAN HAD OBTAINED TRANSSEPTAL ACCESS, AND THE OCTARAY MAPPING CATHETER ADVANCED INTO THE LEFT ATRIUM, AND AFTER ABOUT 5-10 MINUTES GAINING TRANSSEPTAL ACCESS. A PERICARDIAL EFFUSION WAS DISCOVERED AND CONFIRMED ON INTRACARDIAC ECHOCARDIOGRAPHY (ICE) WITH THE SOUNDSTAR CATHETER. THE MEDICAL INTERVENTION PROVIDED WAS PERICARDIOCENTESIS, AND THE AMOUNT OF FLUID REMOVED WAS UNKNOWN. THE STAFF HAD CALLED FOR TWO "UNITS" OF BLOOD, AND THE PATIENT TRANSFERRED TO THE OR. THE PATIENT PASSED AWAY. ADDITIONAL INFORMATION WAS RECEIVED. THE PATIENT¿S DEATH WAS ON (B)(6) 2025. THE PHYSICIAN COMMENTED THAT DURING SURGERY A "HOLE" WAS NOTICED IN THE CS AND LAA THAT WERE NOT PRESENT DURING INITIAL ICE IMAGES AT THE BEGINNING OF THE PROCEDURE. IT WAS A FARAPULSE CASE; THEREFORE, THE NGEN GENERATOR OR FORCE ABLATION CATHETERS WERE NOT BEING USED AT THE TIME. THE ADVERSE EVENT WAS ASSESSED AS MDR REPORTABLE UNDER THE OCTARAY MAPPING CATHETER AND THE EZ STEER NAV CATHETER. Manufacturer Narrative: THE LOT NUMBER PROVIDED OF 2230288 WAS NOT RECOGNIZED IN THE SYSTEM. D4: UDI: AS THE CATALOG/MODEL NUMBER WAS NOT PROVIDED, THE (01)GTIN IS NOT AVAILABLE. AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. AN EVALUATION OF THE MANUFACTURING RECORD COULD NOT BE PERFORMED AS THE REQUIRED PRODUCT IDENTIFICATION NUMBER WAS NOT PROVIDED TO COMPLETE THE EVALUATION. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. HOWEVER, IF THE PRODUCT IS RECEIVED AT A LATER DATE, THE INVESTIGATION WILL BE UPDATED AS APPLICABLE. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. MANUFACTURER'S REFERENCE NUMBER: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22910523&pc=MTD | 2029046-2025-02895 | 2025-08-04 04:00 | Death | BIOSENSE WEBSTER INC | 2025-08-28 04:00 | MTD | OCTARAY MAPPING CATHETER | Adverse Event Without Identified Device or Use Problem | Cardiac Perforation | K193237 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN ATRIAL FIBRILLATION (AFIB) CARDIAC ABLATION/MAPPING PROCEDURE WITH AN OCTARAY MAPPING CATHETER AND AN EZ STEER NAV CATHETER. THE PATIENT EXPERIENCED A PERICARDIAL EFFUSION TREATED WITH PERICARDIOCENTESIS, TRANSFERRED TO THE OPERATING ROOM (OR), AND THE PATIENT PASSED AWAY. ADDITIONALLY, DURING SURGERY A HOLE WAS FOUND IN THE CORONARY SINUS (CS) AND LEFT ATRIAL APPENDAGE (LAA). DURING THE PROCEDURE, A PERICARDIAL EFFUSION OCCURRED. THE PHYSICIAN HAD OBTAINED TRANSSEPTAL ACCESS, AND THE OCTARAY MAPPING CATHETER ADVANCED INTO THE LEFT ATRIUM, AND AFTER ABOUT 5-10 MINUTES GAINING TRANSSEPTAL ACCESS. A PERICARDIAL EFFUSION WAS DISCOVERED AND CONFIRMED ON INTRACARDIAC ECHOCARDIOGRAPHY (ICE) WITH THE SOUNDSTAR CATHETER. THE MEDICAL INTERVENTION PROVIDED WAS PERICARDIOCENTESIS, AND THE AMOUNT OF FLUID REMOVED WAS UNKNOWN. THE STAFF HAD CALLED FOR TWO "UNITS" OF BLOOD, AND THE PATIENT TRANSFERRED TO THE OR. THE PATIENT PASSED AWAY. ADDITIONAL INFORMATION WAS RECEIVED. THE PATIENT¿S DEATH WAS ON (B)(6) 2025. THE PHYSICIAN COMMENTED THAT DURING SURGERY A "HOLE" WAS NOTICED IN THE CS AND LAA THAT WERE NOT PRESENT DURING INITIAL ICE IMAGES AT THE BEGINNING OF THE PROCEDURE. IT WAS A FARAPULSE CASE; THEREFORE, THE NGEN GENERATOR OR FORCE ABLATION CATHETERS WERE NOT BEING USED AT THE TIME. THE ADVERSE EVENT WAS ASSESSED AS MDR REPORTABLE UNDER THE OCTARAY MAPPING CATHETER AND THE EZ STEER NAV CATHETER. Manufacturer Narrative: D4: UDI: AS THE CATALOG/MODEL NUMBER WAS NOT PROVIDED, THE (01)GTIN IS NOT AVAILABLE. AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. AN EVALUATION OF THE MANUFACTURING RECORD COULD NOT BE PERFORMED AS THE REQUIRED PRODUCT IDENTIFICATION NUMBER WAS NOT PROVIDED TO COMPLETE THE EVALUATION. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. HOWEVER, IF THE PRODUCT IS RECEIVED AT A LATER DATE, THE INVESTIGATION WILL BE UPDATED AS APPLICABLE. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. MANUFACTURER'S REFERENCE NUMBER: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22910677&pc=MSD | 3006260740-2025-06333 | 2025-08-01 04:00 | Death | C.R. BARD, INC. (BASD) -3006260740 | 2025-08-28 04:00 | MSD | GLIDEPATH | Disconnection; Fluid/Blood Leak | Hemorrhage/Bleeding | K051748 |  | 1 | Event Description: ON (B)(6) 2025, A PATIENT UNDERWENT A BD GLIDEPATH DIALYSIS CATHETER PLACEMENT PROCEDURE FOR DIALYSIS TREATMENT. ON (B)(6) 2025, IT WAS REPORTED DURING A DIALYSIS SESSION, THE VENOUS-SIDE LUER CONNECTOR ALLEGEDLY DETACHED FROM THE DIALYSIS CIRCUIT WHILE THE CIRCUIT WAS BEING ROTATED, RESULTING IN A SIGNIFICANT BLOOD LOSS. REPORTEDLY, THERE WAS NO VISIBLE DAMAGE TO THE VENOUS SIDE LUER CONNECTOR. THE INCIDENT WAS DISCOVERED DURING A STAFF LUNCH CHANGEOVER, AND THE PATIENT SUBSEQUENTLY PASSED AWAY. THE HOSPITAL BELIEVES THE DISCONNECTION MAY HAVE BEEN DUE TO AN INSECURE CONNECTION, LIKELY A PROCEDURAL ISSUE. THE PATIENT'S CAUSE OF DEATH WAS UNKNOWN. Manufacturer Narrative: H11: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22911035&pc=DYB | 3005168196-2025-00276 | 2025-08-02 04:00 | Death | PENUMBRA, INC. | 2025-08-28 04:00 | DYB | ELEMENT VASCULAR ACCESS SHEATH | Adverse Event Without Identified Device or Use Problem | Cardiac Perforation | K242520 |  | 1 | Event Description: ON (B)(6) 2025, THE PATIENT WAS UNDERGOING A THROMBECTOMY PROCEDURE FOR PULMONARY EMBOLISM IN THE BILATERAL PULMONARY ARTERIES USING A CATHETER (PIGTAIL), GUIDEWIRE, AND AN ELEMENT VASCULAR ACCESS SHEATH (ELEMENT SHEATH). DURING THE PROCEDURE, THE PHYSICIAN ADVANCED THE GUIDEWIRE AND CATHETER BUT WAS UNAWARE THAT THE GUIDEWIRE AND CATHETER WERE POSITIONED IN THE PERICARDIUM. THE PHYSICIAN THEN INSERTED THE ELEMENT SHEATH OVER THE GUIDEWIRE. AT THIS TIME, THE PATIENT DECOMPENSATED, AND A CODE BLUE WAS INITIATED. DURING RESUSCITATION EFFORTS, THE PATIENT WAS ADMINISTERED TISSUE PLASMINOGEN ACTIVATOR (TPA); HOWEVER, THE PATIENT SUBSEQUENTLY EXPIRED. ON 12-AUG-2025, THE PENUMBRA REPRESENTATIVE RECEIVED INFORMATION THAT DURING AN IMAGING REVIEW IN THE DAYS FOLLOWING THE PROCEDURE, IT WAS DETERMINED THAT THE GUIDEWIRE PERFORATED THE PERICARDIUM. Manufacturer Narrative: THE PRODUCT WAS NOT RETURNED FOR EVALUATION. THEREFORE, A PHYSICAL DEVICE INVESTIGATION WAS UNABLE TO BE PERFORMED. FROM THE INFORMATION PROVIDED, THERE IS NO INDICATION THAT THERE WAS ANY DEVICE MALFUNCTION, NONCONFORMANCE, OR MISUSE THAT CONTRIBUTED TO THE REPORTED EVENT. THE MANUFACTURING RECORDS FOR THIS LOT WERE REVIEWED AND DID NOT REVEAL ANY OUTSTANDING DISCREPANCIES, DESIGN, OR QUALITY CONCERNS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22911080&pc=JDI | 0001825034-2025-02703 | 2025-08-04 04:00 | Death | ZIMMER BIOMET, INC. | 2025-08-28 04:00 | JDI | RINGLOC BI-POLAR 28X51MM | Difficult to Insert; Material Twisted/Bent | No Clinical Signs, Symptoms or Conditions | K051569 |  | 1 | Event Description: IT WAS REPORTED DURING AN INITIAL HIP PROCEDURE THAT THERE WAS DIFFICULTY POSITIONING THE RINGLOC BI-POLAR AND HEAD TOGETHER. IT WAS THEN DISCOVERED THAT THE RING IN THE BI-POLAR WAS BENT. AN ADDITIONAL 52 BI-POLAR WAS OPENED, AND THE HEAD WAS REMOVED FROM THE POLYETHYLENE SO THE NEW IMPLANT COULD BE POSITIONED. THERE WAS A 5 (FIVE) MINUTE DELAY DURING THE SURGERY DUE TO THE ISSUE WITH THE BIPOLAR CUP. CEMENT WAS USED FOR THE STEM. SUBSEQUENTLY, AFTER CLOSING THE OPERATION, THE PATIENT DID NOT REGAIN CONSCIOUSNESS AND PASSED AWAY APPROXIMATELY 30 (THIRTY) MINUTES LATER. BASED ON THE INFORMATION CURRENTLY AVAILABLE, THERE IS NO EVIDENCE TO SUGGEST THAT DEVICE CONTRIBUTED TO THE REPORTED DEATH. THE PATIENT WAS 86 YEARS OLD, FRAIL, AND HAD MULTIPLE COMORBIDITIES. AN INVESTIGATION IS IN PROGRESS. NO ADDITIONAL INFORMATION AVAILABLE FOR THE REPORTED EVENT. Manufacturer Narrative: (B)(4). G2: SOUTH AFRICA. CUSTOMER HAS INDICATED THAT THE PRODUCT IS IN PROCESS OF BEING RETURNED TO ZIMMER BIOMET FOR INVESTIGATION. ONCE THE INVESTIGATION HAS BEEN COMPLETED, A SUPPLEMENTAL MDR WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22911212&pc=DSQ | 2916596-2025-05602 | 2024-07-15 04:00 | Death | THORATEC CORPORATION | 2025-08-28 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Hemorrhagic Stroke | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO HEMORRHAGIC STROKE. THE PATIENT PRESENTED TO THE HOSPITAL WITH HEMORRHAGIC STROKE AND CARE WAS WITHDRAWN ON (B)(6) 2025. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED AND THE PUMP WOULD NOT RETURN. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: THE PUMP REMAINS IN USE SUPPORTING THE PATIENT. A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF STROKE COULD NOT BE CONCLUSIVELY DETERMINED. STROKE IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. STROKE HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22911617&pc=CBK | 2518422-2025-109426 | 2025-08-22 04:00 | Death | RESPIRONICS, INC. | 2025-08-28 04:00 | CBK | TRILOGY EVO, JAPAN | Insufficient Information | Apnea | K181166 |  | 1 | Event Description: BASED ON THE INFORMATION CURRENTLY PROVIDED AND AVAILABLE, IT WAS REPORTED THAT ON (B)(6) 2025 AT APPROXIMATELY 2100, A NURSE HAD MADE PATIENT ROUNDS AND CONFIRMED THAT THE DEVICE (TRILOGY EVO) WAS OPERATING WITHOUT ISSUE. AT APPROXIMATELY 2300, THE NURSE HAD CONDUCTED PATIENT ROUNDS AGAIN AND FOUND THE DEVICE WAS IN STANDBY MODE, WHEREBY NO DEVICE THERAPY WAS BEING DELIVERED TO THE PATIENT. IT WAS REPORTED THAT NO ALARM SOUNDED AT THE TIME OF THE EVENT. THE PATIENT WAS FOUND APNEIC, AND A PHYSICIAN WAS CALLED UPON AND MEDICAL INTERVENTION (RESUSCITATION EFFORTS) PROVIDED, HOWEVER THE PATIENT WAS NOTED TO HAVE EXPIRED. IT WAS CONFIRMED THAT THE PATIENT COULD NOT OPERATE THE DEVICE THEMSELVES, AND FURTHER INVESTIGATION IS BEING CONDUCTED BY LOCAL PREFECTURE POLICE. THE 77-YEAR-OLD FEMALE PATIENT HAD A PRIMARY DISEASE OF AMYOTROPHIC LATERAL SCLEROSIS (ALS) AND HAD BEEN USING THE DEVICE FOR 12 MONTHS FOR VENTILATION THERAPY. THE SUBJECT DEVICE HAS BEEN COLLECTED FROM THE HEALTHCARE FACILITY AND SENT TO THE MANUFACTURER'S REPAIR CENTER FOR INVESTIGATION INCLUDING REVIEW OF THE DEVICE EVENT LOGS. FURTHER INFORMATION GATHERING IS REQUIRED TO DETERMINE SPECIFICITIES REGARDING ANY CAUSAL NATURE OR CONTRIBUTORY FACTORS. GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL CLARIFYING INFORMATION ARE CURRENTLY BEING CONDUCTED. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22911685&pc=INO | 2530130-2025-00956 | 2025-05-05 04:00 | Death | PRIDE MOBILITY PRODUCTS CORPORATION | 2025-08-28 04:00 | INO | PRIDE MOBILITY PRODUCTS | Appropriate Term/Code Not Available | No Clinical Signs, Symptoms or Conditions | K091578 |  | 1 | Event Description: ALLEGES THE LITHIUM-ION BATTERY ALLEGEDLY CAUGHT FIRE. Manufacturer Narrative: THE DEVICE HAS NOT YET BEEN MADE AVAILABLE FOR EVALUATION. SHOULD FURTHER INFORMATION OR THE DEVICE BECOME AVAILABLE, A FOLLOW-UP REPORT WILL THEN BE ISSUED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22911807&pc=NPT | 2135147-2025-04829 | 2025-08-04 04:00 | Death | ABBOTT MEDICAL | 2025-08-28 04:00 | NPT | NAVITOR VISION | Incomplete Coaptation; Central Regurgitation | Aortic Valve Insufficiency/ Regurgitation | P190023 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 25MM NAVITOR VISON VALVE WAS CHOSEN FOR A TRANSCATHETER AORTIC VALVE IMPLANTATION (TAVI) PROCEDURE IN A HEMODYNAMICALLY UNSTABLE PATIENT WITH MODERATE CENTRAL LEAK USING A SMALL FLEXNAV DELIVERY SYSTEM. A PRE-DILATATION WAS PERFORMED WITH A 20MM NON-ABBOTT BALLOON. DURING PROCEDURE, THE INFLOW EDGE OF THE VALVE IN CUSP OVERLAP VIEW (COV) WAS POSITIONED USING THE RADIOPAQUE REFERENCE MARKER, AND THE PIGTAIL CATHETER WAS IN THE NON-CORONARY CUSP (NCC). A POST-DILATATION WAS PERFORMED IN AN ATTEMPT TO BETTER ACCOMMODATE THE PROSTHESIS. THE FINAL IMPLANT DEPTH WAS 6MM. AFTER FIVE BALLOONING ATTEMPTS, ONE WAS EFFECTIVE, BUT THE MODERATE CENTRAL LEAK DID NOT IMPROVE. THE ECHOCARDIOGRAPHY SHOWED THAT ONE OF THE LEAFLETS WAS NON-FUNCTIONAL. THE PATIENT LEFT THE PROCEDURE ROOM STABLE AND WAS TRANSFERRED TO THE INTENSIVE CARE UNIT (ICU). A NEW ECHOCARDIOGRAPHIC EVALUATION AND FOLLOW-UP WILL BE CONDUCTED TO ASSESS THE NEED FOR A NEW VALVE. ON (B)(6) 2025, IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS AORTIC INSUFFICIENCY, INFECTION, AND RENAL FAILURE. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22911846&pc=DXY | 2182208-2025-04298 | 2025-07-11 04:00 | Death | MEDTRONIC, INC. | 2025-08-28 04:00 | DXY | MEDTRONIC ICD | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING THE LONG-TERM OUTCOMES AFTER IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) IMPLANTATION IN THE ADULT CONGENITAL HEART DISEASE (ACHD) POPULATION. THE AUTHORS DESCRIBED THIRTY-FIVE PATIENTS STUDIED WITH ACHD AND ICD IMPLANTATION EITHER FOR PRIMARY OR SECONDARY PREVENTION. THERE WERE THREE PATIENT DEATHS; TWO DIED OF HEART FAILURE AND ONE DUE TO BACTEREMIA CAUSED BY A LEAD INFECTION. THERE WERE SIX PATIENTS WHO EXPERIENCED INAPPROPRIATE THERAPY DUE TO ATRIAL ARRHYTHMIAS, T-WAVE OVERSENSING (TWOS), AND NOISE. PATIENTS UNDERWENT ABLATION PROCEDURES FOR ATRIAL ARRHYTHMIAS OR ADJUSTMENTS TO MEDICATION TO PREVENT INAPPROPRIATE THERAPIES. THERE WAS ONE PATIENT WITH AN UNKNOWN LEAD FAILURE WHICH REQUIRED REIMPLANTATION. THE STATUS OF THE DEVICES AND LEADS ISUNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: LONG-TERM OUTCOME OF ADULT CONGENITAL HEART DISEASE PATIENTS WITH IMPLANTABLE CARDIOVERTER-DEFIBRILLATORS. CONGENITAL HEART DISEASE. 2025; 20(3). DOI: 10.32604/CHD.2025.067716 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22911847&pc=DXY | 2182208-2025-04299 | 2025-07-11 04:00 | Death | MEDTRONIC, INC. | 2025-08-28 04:00 | DXY | MEDTRONIC LEAD | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING THE LONG-TERM OUTCOMES AFTER IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) IMPLANTATION IN THE ADULT CONGENITAL HEART DISEASE (ACHD) POPULATION. THE AUTHORS DESCRIBED THIRTY-FIVE PATIENTS STUDIED WITH ACHD AND ICD IMPLANTATION EITHER FOR PRIMARY OR SECONDARY PREVENTION. THERE WERE THREE PATIENT DEATHS; TWO DIED OF HEART FAILURE AND ONE DUE TO BACTEREMIA CAUSED BY A LEAD INFECTION. THERE WERE SIX PATIENTS WHO EXPERIENCED INAPPROPRIATE THERAPY DUE TO ATRIAL ARRHYTHMIAS, T-WAVE OVERSENSING (TWOS), AND NOISE. PATIENTS UNDERWENT ABLATION PROCEDURES FOR ATRIAL ARRHYTHMIAS OR ADJUSTMENTS TO MEDICATION TO PREVENT INAPPROPRIATE THERAPIES. THERE WAS ONE PATIENT WITH AN UNKNOWN LEAD FAILURE WHICH REQUIRED REIMPLANTATION. THE STATUS OF THE DEVICES AND LEADS ISUNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: LONG-TERM OUTCOME OF ADULT CONGENITAL HEART DISEASE PATIENTS WITH IMPLANTABLE CARDIOVERTER-DEFIBRILLATORS. CONGENITAL HEART DISEASE. 2025; 20(3). DOI: 10.32604/CHD.2025.067716 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22912104&pc=QZI | MW5175299 | 2025-04-17 04:00 | Death | MEDTRONIC, INC. | 2025-08-28 04:00 | QZI | PULSESELECT¿ | Insufficient Information | Low Blood Pressure/ Hypotension; Pericardial Effusion |  |  | 1 | Event Description: IT WAS REPORTED THAT THE S70 ULTRASOUND MACHINE DISPLAYED AN ERROR - "SET UP PROBLEM OCCURRED. PLEASE RESELECT PROBE TO CONTINUE" (ERROR CODE UNKNOWN). THE PATIENT INTERFACE UNIT (PIU) WAS REBOOTED WITHOUT RESOLUTION. THEY RESEATED THE CATHETER AND THE PROBE AT THE SAME TIME WITHOUT RESOLUTION. THEY REPLACED THE CATHETER AND THE PROBE AT THE SAME TIME AND THE ISSUE WAS RESOLVED. THE PROCEDURE CONTINUED. IT WAS ALSO REPORTED THAT WHILE ABLATING IN THE LEFT ATRIAL APPENDAGE, THE PHYSICIAN WAS LOOKING AT THE ARTERIAL LINE AND CONFIRMED THERE WAS A SUDDEN DROP IN BLOOD PRESSURE ON THE PATIENT. THE PHYSICIAN CHECKED FOR A PERICARDIAL EFFUSION, BUT A PERICARDIAL EFFUSION WAS NOT PRESENT. THE PHYSICIAN DECIDED TO PERFORM CARDIOPULMONARY RESUSCITATION (CPR) ON THE PATIENT UNTIL THE RESPONSE TEAM ARRIVED. WHEN THE RESPONSE TEAM ARRIVED, THE ELECTRICAL ACTIVITY WAS STILL ACTIVE IN THE HEART AND WAS CONFIRMED BY THE CORONARY SINUS (CS) CATHETER STILL IN THE BODY, BUT THERE WAS NO PULSE. WHILE THE CS CATHETER WAS STILL IN THE BODY, AND THE PATIENT'S CONDITION REMAINED THE SAME WITH LOW BLOOD PRESSURE, ACTIVE ELECTRICAL ACTIVITY IN THE HEART, AND NO PULSE, THE RESPONSE TEAM CONTINUED WITH THE CPR. ANOTHER PHYSICIAN CONFIRMED A SMALL PERICARDIAL EFFUSION IN THE VENTRICULAR WITH INTRACARDIAC ECHOCARDIOGRAPHY (ICE). A PERICARDIOCENTESIS WAS PERFORMED AND THE CPR CONTINUED. AT ONE POINT IN TIME, THE ELECTRICAL ACTIVITY IN THE HEART WAS LOST WHILE THE CS CATHETER WAS STILL IN THE BODY. THE PATIENT WAS DEFIBRILLATED, BUT THERE WAS NO CHANGE IN THE PATIENT SITUATION, AND THE CPR WAS CONTINUED. AT THIS POINT, THE REPORTER STEPPED OUT OF THE ROOM. WHEN THE REPORTER CAME BACK IN THE ROOM, THE PATIENT WAS ANNOUNCED TO BE DECEASED. (B)(6) PULSESELECT¿ PULSED FIELD ABLATION (PFA) SYSTEM WAS IN USE. THE FOLLOWING BWI PRODUCTS WERE IN USE: CS CATHETER (LOT# UNKNOWN; CAT # UNKNOWN); 2 NUVISION ULTRASOUND CATHETER - (LOT - 31424477L; REF - FG10593) (31424483L - FG10593). THIS REPORT REFLECTS INFORMATION RECEIVED BY FDA IN THE FORM OF A NOTIFICATION PER 803.22 (B)(2).Event Description: ADDITIONAL INFORMATION RECEIVED ON 8/29/2025 FOR MW5175299 TO UPDATE BRAND NAME AND MANUFACTURER. THIS REPORT REFLECTS INFORMATION RECEIVED BY FDA IN THE FORM OF A NOTIFICATION PER 803.22 (B)(2).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22912177&pc=NPT | 9612164-2025-04253 | 2025-08-21 04:00 | Death | MEDTRONIC IRELAND | 2025-08-28 04:00 | NPT | EVOLUT FX DCS | Insufficient Information; Patient Device Interaction Problem | Low Blood Pressure/ Hypotension; Cardiac Tamponade | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE IMPLANTATION PROCEDURE; THE VALVE WAS ABOUT TO BE RECAPTURED FOR A SECOND TIME WHEN THE PATIENT'S BLOOD PRESSURE DROPPED. SUBSEQUENTLY, CARDIAC TAMPONADE WAS OBSERVED. THE SYSTEM WAS WITHDRAWN FROM THE PATIENT, AND CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED. THE PATIENT SUBSEQUENTLY DIED. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID EVFXPLUS-29 (SERIAL: (B)(6)); PRODUCT TYPE: 0195-HEART VALVES. SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22912278&pc=PNJ | 9612164-2025-04254 | 2025-08-08 04:00 | Death | MEDTRONIC IRELAND | 2025-08-28 04:00 | PNJ | MICRA¿ AV2 | Appropriate Term/Code Not Available | Cardiac Arrest; Cardiac Tamponade | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. IT WAS FURTHER REPORTED THAT DURING THE IMPLANT OF THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) THE PATIENT EXPERIENCED CARDIAC ARREST AND CARDIAC TAMPONADE. AN ECHOCARDIOGRAM, PERICARDIOCENTESIS AND RESUSCITATIVE MEASURES WERE PERFORMED BUT THE PATIENT PASSED AWAY DUE TO CARDIA TAMPONADE AND COMPLETE HEART BLOCK Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22912339&pc=PNJ | 9612164-2025-04255 | 2025-08-08 04:00 | Death | MEDTRONIC IRELAND | 2025-08-28 04:00 | PNJ | MICRA¿ AV2 | Appropriate Term/Code Not Available | Cardiac Arrest; Cardiac Tamponade | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. IT WAS FURTHER REPORTED THAT DURING THE IMPLANT OF THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) THE PATIENT EXPERIENCED CARDIAC ARREST AND CARDIAC TAMPONADE. AN ECHOCARDIOGRAM, PERICARDIOCENTESIS AND RESUSCITATIVE MEASURES WERE PERFORMED BUT THE PATIENT PASSED AWAY DUE TO CARDIA TAMPONADE AND COMPLETE HEART BLOCK Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22912340&pc=DYB | 9612164-2025-04256 | 2025-08-08 04:00 | Death | MEDTRONIC IRELAND | 2025-08-28 04:00 | DYB | MICRA | Appropriate Term/Code Not Available | Cardiac Arrest; Cardiac Tamponade | K132030 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. IT WAS FURTHER REPORTED THAT DURING THE IMPLANT OF THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) THE PATIENT EXPERIENCED CARDIAC ARREST AND CARDIAC TAMPONADE. AN ECHOCARDIOGRAM, PERICARDIOCENTESIS AND RESUSCITATIVE MEASURES WERE PERFORMED BUT THE PATIENT PASSED AWAY DUE TO CARDIA TAMPONADE AND COMPLETE HEART BLOCK. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22912383&pc=CCW | 3006061749-2025-00042 | 2025-08-23 04:00 | Death | FLEXICARE MEDICAL LTD | 2025-08-28 04:00 | CCW | PROVU | Display or Visual Feedback Problem; Output Problem | Insufficient Information |  |  | 1 | Event Description: "WHEN THE AMR PARAMEDIC WAS PREPARING FOR INTUBATION, IT WAS REQUESTED THAT WE PROVIDE THEM WITH OUR PROVU VIDEO LARYNGOSCOPE FOR INTUBATION. THE DEVICE WAS POWERED ON, CONFIRMED AS BEING 98% BATTERY AND PLACE ONTO THE BLADE WITHOUT ISSUE. WHILE THE AMR PARAMEDIC WAS INSERTING THE BLADE INTO THE PATIENTS MOUTH, THE SCREEN TURNED INTO AN ERROR SCREEN WHICH CONTAINED NUMEROUS PICTOGRAMS/DIAGRAMS AND NOT ALLOWING TO BE CHANGED, TURNED OFF OR RETURN TO THE NORMAL OPERATING SCREEN THAT IT WAS DISPLAYING WHEN IT ENTERED THE PATIENTS MOUTH. THIS WAS THEN REMOVED AND PATIENT CARE WAS CONTINUED BY AMR PARAMEDIC THROUGHOUT. AFTER THE CALL THIS ISSUE BECAME REPRODUCIBLE AT THE STATION WITH MULTIPLE BLADES. OF NOTE, THIS ISSUE WAS NONEXISTENT AT THE START OF THE SHIFT WHEN CONNECTED AND PROPERLY POWERED ON AND OFF. ADDITIONALLY, AMR PARAMEDIC CONFIRMED THEY WERE FAMILIAR AND TRAINED IN THE USE AND FUNCTION OF THE DEVICE PRIOR TO THE EQUIPMENT BEING PROVIDED. THERE WAS NO OBVIOUS ISSUE WITH TECHNIQUE, PLACEMENT OR OPERATION ON BEHALF OF THE PROVIDER AS THIS WAS STRICTLY A DEVICE FAILURE". Manufacturer Narrative: WHEN THE AMR PARAMEDIC WAS PREPARING FOR INTUBATION, IT WAS REQUESTED THAT WE PROVIDE THEM WITH OUR PROVU VIDEO LARYNGOSCOPE FOR INTUBATION. THE DEVICE WAS POWERED ON, CONFIRMED TO BE AT 98% BATTERY AND PLACED ONTO THE BLADE WITHOUT ISSUE. WHILE THE AMR PARAMEDIC WAS INSERTING THE BLADE INTO THE PATIENT'S MOUTH THE SCREEN TURNED INTO AN ERROR SCREEN WHICH CONTAINED NUMEROUS PICTOGRAMS/DIAGRAMS AND NOT ALLOWING TO BE CHANGED, TURNED OFF OR RETURN TO THE NORMAL OPERATING SCREEN THAT IT WAS DISPLAYING WHEN IT ENTERED THE PATIENT'S MOUTH. THIS WAS THEN REMOVED AND PATIENT CARE WAS CONTINUED BY AMR PARAMEDIC THROUGHOUT. AFTER THE CALL THIS ISSUE BECAME REPRODUCIBLE AT THE STATION WITH MULTIPLE BLADES. OF NOTE, THIS ISSUE WAS NONEXISTENT AT THE START OF THE SHIFT WHEN CONNECTED AND PROPERLY POWERED ON AND OFF. ADDITIONALLY, AMR PARAMEDIC CONFIRMED THEY WERE FAMILIAR AND TRAINED IN THE USE AND FUNCTION OF THE DEVICE PRIOR TO THE EQUIPMENT BEING PROVIDED. THERE WAS NO OBVIOUS ISSUE WITH TECHNIQUE, PLACEMENT OR OPERATION ON BEHALF OF THE PROVIDER AS THIS WAS STRICTLY A DEVICE FAILURE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22912702&pc=HHS | 2951250-2025-00545 |  | Death | BAYER PHARMA AG | 2025-08-28 04:00 | HHS | ESSURE | Adverse Event Without Identified Device or Use Problem | Hypersensitivity/Allergic reaction | P020014 |  | 1 | Event Description: BAYER CASE NUMBER: (B)(4). CARDIORESPIRATORY ARREST [CARDIO-RESPIRATORY ARREST]. LIMITED COLLECTION OF FLUID SUGGESTIVE OF AN ABSCESS 40 MM IN DIAMETER IN THE POUCH OF DOUGLAS [DOUGLAS' ABSCESS]. SEPSIS [SEPSIS]. PROCTOSIGMOIDITIS [PROCTOSIGMOIDITIS]. NICKEL ALLERGY WITH THE RIGHT ESSURE IN DISTAL POSITION [NICKEL ALLERGY]. SEVERE ABDOMINAL PAIN [ABDOMINAL PAIN]. DIARRHOEA [DIARRHOEA]. FEVER [FEVER]. ILEITIS REACTIVE ON CONTAC [ILEITIS]. HYPERLEUKOCYTOSIS [HYPERLEUKOCYTOSIS]. FOUND BLOOD PRESSURE AT 99/53 [BLOOD PRESSURE DIASTOLIC LOW]. CASE NARRATIVE: THIS SPONTANEOUS CASE WAS REPORTED BY A PHARMACIST AND DESCRIBES THE OCCURRENCE OF FATAL CARDIO-RESPIRATORY ARREST ("CARDIORESPIRATORY ARREST"), DOUGLAS' ABSCESS ("LIMITED COLLECTION OF FLUID SUGGESTIVE OF AN ABSCESS 40 MM IN DIAMETER IN THE POUCH OF DOUGLAS"), SEPSIS ("SEPSIS"), COLITIS ("PROCTOSIGMOIDITIS") AND ALLERGY TO METALS ("NICKEL ALLERGY WITH THE RIGHT ESSURE IN DISTAL POSITION") IN A 51 YEAR-OLD FEMALE PATIENT WHO HAD ESSURE INSERTED FOR FEMALE STERILISATION. ADDITIONAL NON-SERIOUS EVENTS ARE DETAILED BELOW. THE PATIENT HAD A MEDICAL HISTORY OF INFARCTION (TWO INFARCTS IN 2017 AND 2019, CURRENTLY UNDER TREATMENT), KNEE OPERATION, RECURRENT DEPRESSIVE DISORDER, MEMBRANOUS VENTRICULAR SEPTAL DEFECT, AORTIC STENOSIS, AORTIC INCOMPETENCE, BICUSPID AORTIC VALVE, AORTA COARCTATION REPAIR AND OBESITY. ON (B)(6) 2014, THE PATIENT HAD ESSURE INSERTED. ON (B)(6) 2025, 3982 DAYS AFTER ESSURE INSERTION AND 4 DAYS AFTER ITS REMOVAL, SHE EXPERIENCED ABDOMINAL PAIN ("SEVERE ABDOMINAL PAIN") AND DIARRHOEA ("DIARRHOEA"). ON (B)(6) 2025 SHE EXPERIENCED SEPSIS (SERIOUSNESS CRITERION HOSPITALISATION), PYREXIA ("FEVER") AND HYPERLEUKOCYTOSIS ("HYPERLEUKOCYTOSIS") AND WAS FOUND TO HAVE BLOOD PRESSURE DIASTOLIC DECREASED ("FOUND BLOOD PRESSURE AT 99/53"). ON (B)(6) 2025 SHE EXPERIENCED DOUGLAS' ABSCESS (SERIOUSNESS CRITERIA HOSPITALISATION AND MEDICALLY IMPORTANT), COLITIS (SERIOUSNESS CRITERION HOSPITALISATION) AND ENTERITIS ("ILEITIS REACTIVE ON CONTAC"). ON (B)(6) 2025 SHE EXPERIENCED CARDIO-RESPIRATORY ARREST (SERIOUSNESS CRITERION DEATH). ON UNKNOWN DATE SHE EXPERIENCED ALLERGY TO METALS (SERIOUSNESS CRITERION INTERVENTION REQUIRED). THE PATIENT WAS HOSPITALISED FROM (B)(6) 2025 TO (B)(6) 2025. THE PATIENT WAS TREATED WITH ANTIBIOTIC (CHLORAMPHENICOL), AUGMENTIN (AMOXICILLIN SODIUM, CLAVULANATE POTASSIUM) AND TAZOCILLINE (PIPERACILLIN SODIUM, TAZOBACTAM SODIUM) AS WELL AS SURGERY (TO REMOVE ESSURE, HYSTERECTOMY). THE PATIENT DIED ON (B)(6) 2025 AND THE REPORTED CAUSE OF DEATH WAS CARDIO-RESPIRATORY ARREST. THE REPORTER CONSIDERED ALLERGY TO METALS TO BE RELATED TO ESSURE ADMINISTRATION. NO CAUSALITY ASSESSMENT WAS RECEIVED FOR ESSURE WITH REGARD TO DOUGLAS' ABSCESS, COLITIS, ENTERITIS, CARDIO-RESPIRATORY ARREST, ABDOMINAL PAIN, DIARRHOEA, PYREXIA, HYPERLEUKOCYTOSIS, BLOOD PRESSURE DIASTOLIC DECREASED OR SEPSIS. THE REPORTER COMMENTED: SUDDEN DEATH OF PATIENT: CARDIORESPIRATORY ARREST NOT RECOVERED IN THE HOSPITAL DEPARTMENT. PATIENT PRESENTED WITH SEVERE ABDOMINAL PAIN AND DIARRHEA STARTING ON (B)(6) 2025. SHE PRESENTED WITH A FEVER AND CONSULTED WITH HER TREATING DOCTOR ON (B)(6) 2025. ON THE DOCTOR¿S ADVICE, SHE THEN WENT TO THE GYNECOLOGICAL EMERGENCY DEPARTMENT. UPON THE FEMALE PATIENT¿S ADMISSION, THE PHYSICAL EXAMINATION FOUND BLOOD PRESSURE AT 99/53, HEART RATE AT 102, TEMPERATURE AT 38.4°C. THE LABORATORY BLOOD TEST FOUND A HYPERLEUKOCYTOSIS AT 27,440, A C-REACTIVE PROTEIN (CRP) AT 295, CREATININE CLEARANCE AT 75 ML/MIN, LIVER FUNCTION TEST NORMAL AND COAGULATION NORMAL. THE CYTOBACTERIOLOGICAL URINE TEST DID NOT FIND A URINARY INFECTION. THE BLOOD CULTURES OF (B)(6) 2025 ARE IN PROGRESS. AN ABDOMINAL AND PELVIC COMPUTED TOMOGRAPHY (CT) SCAN FOUND A VERY LIMITED COLLECTION OF FLUID SUGGESTIVE OF AN ABSCESS 40 MM IN DIAMETER IN THE POUCH OF DOUGLAS WITH PROCTOSIGMOIDITIS AND ILEITIS REACTIVE ON CONTACT. THE PATIENT WAS THEREFORE HOSPITALIZED FOR THE MEDICAL MANAGEMENT OF THIS ABSCESS WITH ANTIBIOTIC THERAPY WITH AUGMENTIN SWITCHED ON (B)(6) 2025 TO TAZOCILLINE. ON THE MORNING OF (B)(6) 2025, HEMODYNAMIC PARAMETERS WERE NORMAL. THE PATIENT DID NOT REPORT ANY SPECIFIC COMPLAINT, EXCEPT FOR MODERATE ABDOMINAL PAIN THAT HAD BEEN IMPROVING SINCE ANTIBIOTIC THERAPY WAS INITIATED. AT ABOUT 12.20 P.M., THE PATIENT WAS FOUND IN HER BED IN A STATE OF CARDIORESPIRATORY ARREST, AND SHE DID NOT RECOVER DESPITE RESUSCITATION CONDUCTED FOR 20 MINUTES. THIS CARDIORESPIRATORY ARREST COULD BE EXPLAINED BY HEART FAILURE RELATED TO SEPSIS. CONTEXT: 51-YEAR-OLD PATIENT HOSPITALIZED SINCE (B)(6) 2025 FOR MEDICAL MANAGEMENT OF SEPSIS SECONDARY TO ABDOMINAL ABSCESS OF POUCH OF DOUGLAS. PATIENT WAS ON DAY 9 OF AN INTERVENTION FOR TOTAL HYSTERECTOMY BY LAPAROSCOPY WITHOUT SALPINGO-OOPHORECTOMY. PATIENT HAD ESSURE IN PLACE FOR THE PREVIOUS 11 YEARS WITH GENERALIZED SOMATIC SYMPTOMS AND PROVEN NICKEL ALLERGY WITH THE RIGHT ESSURE IN DISTAL POSITION. DIAGNOSTIC RESULTS (NORMAL RANGES ARE PROVIDED IN PARENTHESIS IF AVAILABLE): BODY MASS INDEX: 34 KG/SQM. [BLOOD PRESSURE MEASUREMENT] IN (B)(6) 2025: 99/53. [BLOOD TEST] IN (B)(6) 2025: A HYPERLEUKOCYTOSIS AT 27,440. [BODY TEMPERATURE] IN (B)(6) 2025: 38.4. [C-REACTIVE PROTEIN] IN (B)(6) 2025: 295. [COAGULATION TEST] IN (B)(6) 2025: NORMAL. [COMPUTERISED TOMOGRAM ABDOMEN] ON (B)(6) 2025: A VERY LIMITED COLLECTION OF FLUID SUGGESTIVE OF AN ABSCESS 40 MM IN DIAMETER IN THE POUCH OF DOUGLAS WITH PROCTOSIGMOIDITIS AND ILEITIS REACTIVE ON CONTACT. [CREATININE RENAL CLEARANCE] IN (B)(6) 2025: 75. [LIVER FUNCTION TEST] IN (B)(6) 2025: NORMAL. [URINE CYTOLOGY] IN (B)(6) 2025: DID NOT FIND A URINARY INFECTION. CASE COMMENTS: BASED ON THE AVAILABLE INFORMATION, A REVIEW OF OUR COMPLAINT RECORDS AND OTHER RELEVANT DATA WILL BE CONDUCTED; ANY NEW AND REPORTABLE INFORMATION THAT BECOMES AVAILABLE FROM OUR INVESTIGATION WILL BE PROVIDED IN A SUPPLEMENTARY REPORT. IN CASE A SAMPLE IS RECEIVED THE INVESTIGATION MIGHT LEAD TO DESTRUCTION OF THE SAMPLE IF REQUIRED TO PERFORM A PROPER ANALYSIS. Manufacturer Narrative: BAYER CASE NUMBER: (B)(4) CARDIORESPIRATORY ARREST [CARDIO-RESPIRATORY ARREST] LIMITED COLLECTION OF FLUID SUGGESTIVE OF AN ABSCESS 40 MM IN DIAMETER IN THE POUCH OF DOUGLAS [DOUGLAS' ABSCESS] SEPSIS [SEPSIS] PROCTOSIGMOIDITIS [PROCTOSIGMOIDITIS] NICKEL ALLERGY WITH THE RIGHT ESSURE IN DISTAL POSITION [NICKEL ALLERGY] SEVERE ABDOMINAL PAIN [ABDOMINAL PAIN] DIARRHOEA [DIARRHOEA] FEVER [FEVER] ILEITIS REACTIVE ON CONTAC [ILEITIS] HYPERLEUKOCYTOSIS [HYPERLEUKOCYTOSIS] FOUND BLOOD PRESSURE AT 99/53 [BLOOD PRESSURE DIASTOLIC LOW]. CASE NARRATIVE: THIS SPONTANEOUS CASE WAS REPORTED BY A PHARMACIST AND DESCRIBES THE OCCURRENCE OF FATAL CARDIO-RESPIRATORY ARREST ("CARDIORESPIRATORY ARREST"), DOUGLAS' ABSCESS ("LIMITED COLLECTION OF FLUID SUGGESTIVE OF AN ABSCESS 40 MM IN DIAMETER IN THE POUCH OF DOUGLAS"), SEPSIS ("SEPSIS"), COLITIS ("PROCTOSIGMOIDITIS") AND ALLERGY TO METALS ("NICKEL ALLERGY WITH THE RIGHT ESSURE IN DISTAL POSITION") IN A 51 YEAR-OLD FEMALE PATIENT WHO HAD ESSURE INSERTED FOR FEMALE STERILISATION. ADDITIONAL NON-SERIOUS EVENTS ARE DETAILED BELOW. THE PATIENT HAD A MEDICAL HISTORY OF INFARCTION (TWO INFARCTS IN 2017 AND 2019, CURRENTLY UNDER TREATMENT), KNEE OPERATION, RECURRENT DEPRESSIVE DISORDER, MEMBRANOUS VENTRICULAR SEPTAL DEFECT, AORTIC STENOSIS, AORTIC INCOMPETENCE, BICUSPID AORTIC VALVE, AORTA COARCTATION REPAIR AND OBESITY. ON (B)(6) 2014, THE PATIENT HAD ESSURE INSERTED. IT WAS REMOVED ON (B)(6) 2025. ON (B)(6) 2025, 3982 DAYS AFTER ESSURE INSERTION AND 4 DAYS AFTER ITS REMOVAL, SHE EXPERIENCED ABDOMINAL PAIN ("SEVERE ABDOMINAL PAIN") AND DIARRHOEA ("DIARRHOEA"). ON (B)(6) 2025 SHE EXPERIENCED SEPSIS (SERIOUSNESS CRITERION HOSPITALISATION), PYREXIA ("FEVER") AND HYPERLEUKOCYTOSIS ("HYPERLEUKOCYTOSIS") AND WAS FOUND TO HAVE BLOOD PRESSURE DIASTOLIC DECREASED ("FOUND BLOOD PRESSURE AT 99/53"). ON (B)(6) 2025 SHE EXPERIENCED DOUGLAS' ABSCESS (SERIOUSNESS CRITERIA HOSPITALISATION AND MEDICALLY IMPORTANT), COLITIS (SERIOUSNESS CRITERION HOSPITALISATION) AND ENTERITIS ("ILEITIS REACTIVE ON CONTAC"). ON (B)(6) 2025 SHE EXPERIENCED CARDIO-RESPIRATORY ARREST (SERIOUSNESS CRITERION DEATH). ON UNKNOWN DATE SHE EXPERIENCED ALLERGY TO METALS (SERIOUSNESS CRITERION INTERVENTION REQUIRED). THE PATIENT WAS HOSPITALISED FROM (B)(6) 2025. THE PATIENT WAS TREATED WITH ANTIBIOTIC (CHLORAMPHENICOL), AUGMENTIN (AMOXICILLIN SODIUM, CLAVULANATE POTASSIUM) AND TAZOCILLINE (PIPERACILLIN SODIUM, TAZOBACTAM SODIUM) AS WELL AS SURGERY (TO REMOVE ESSURE, HYSTERECTOMY). THE PATIENT DIED ON (B)(6) 2025 AND THE REPORTED CAUSE OF DEATH WAS CARDIO-RESPIRATORY ARREST. THE REPORTER CONSIDERED ALLERGY TO METALS TO BE RELATED TO ESSURE ADMINISTRATION. NO CAUSALITY ASSESSMENT WAS RECEIVED FOR ESSURE WITH REGARD TO DOUGLAS' ABSCESS, COLITIS, ENTERITIS, CARDIO-RESPIRATORY ARREST, ABDOMINAL PAIN, DIARRHOEA, PYREXIA, HYPERLEUKOCYTOSIS, BLOOD PRESSURE DIASTOLIC DECREASED OR SEPSIS. THE REPORTER COMMENTED: SUDDEN DEATH OF PATIENT: CARDIORESPIRATORY ARREST NOT RECOVERED IN THE HOSPITAL DEPARTMENT. PATIENT PRESENTED WITH SEVERE ABDOMINAL PAIN AND DIARRHEA STARTING ON (B)(6). SHE PRESENTED WITH A FEVER AND CONSULTED WITH HER TREATING DOCTOR ON (B)(6). ON THE DOCTOR¿S ADVICE, SHE THEN WENT TO THE GYNECOLOGICAL EMERGENCY DEPARTMENT. UPON THE FEMALE PATIENT¿S ADMISSION, THE PHYSICAL EXAMINATION FOUND BLOOD PRESSURE AT 99/53, HEART RATE AT 102, TEMPERATURE AT 38.4°C. THE LABORATORY BLOOD TEST FOUND A HYPERLEUKOCYTOSIS AT 27,440, A C-REACTIVE PROTEIN (CRP) AT 295, CREATININE CLEARANCE AT 75 ML/MIN, LIVER FUNCTION TEST NORMAL AND COAGULATION NORMAL. THE CYTOBACTERIOLOGICAL URINE TEST DID NOT FIND A URINARY INFECTION. THE BLOOD CULTURES OF (B)(6) ARE IN PROGRESS. AN ABDOMINAL AND PELVIC COMPUTED TOMOGRAPHY (CT) SCAN FOUND A VERY LIMITED COLLECTION OF FLUID SUGGESTIVE OF AN ABSCESS 40 MM IN DIAMETER IN THE POUCH OF DOUGLAS WITH PROCTOSIGMOIDITIS AND ILEITIS REACTIVE ON CONTACT. THE PATIENT WAS THEREFORE HOSPITALIZED FOR THE MEDICAL MANAGEMENT OF THIS ABSCESS WITH ANTIBIOTIC THERAPY WITH AUGMENTIN SWITCHED ON (B)(6) TO TAZOCILLINE. ON THE MORNING OF (B)(6), HEMODYNAMIC PARAMETERS WERE NORMAL. THE PATIENT DID NOT REPORT ANY SPECIFIC COMPLAINT, EXCEPT FOR MODERATE ABDOMINAL PAIN THAT HAD BEEN IMPROVING SINCE ANTIBIOTIC THERAPY WAS INITIATED. AT ABOUT 12.20 P.M., THE PATIENT WAS FOUND IN HER BED IN A STATE OF CARDIORESPIRATORY ARREST, AND SHE DID NOT RECOVER DESPITE RESUSCITATION CONDUCTED FOR 20 MINUTES. THIS CARDIORESPIRATORY ARREST COULD BE EXPLAINED BY HEART FAILURE RELATED TO SEPSIS. CONTEXT: 51-YEAR-OLD PATIENT HOSPITALIZED SINCE (B)(6) 2025 FOR MEDICAL MANAGEMENT OF SEPSIS SECONDARY TO ABDOMINAL ABSCESS OF POUCH OF DOUGLAS. PATIENT WAS ON DAY 9 OF AN INTERVENTION FOR TOTAL HYSTERECTOMY BY LAPAROSCOPY WITHOUT SALPINGO-OOPHORECTOMY. PATIENT HAD ESSURE IN PLACE FOR THE PREVIOUS 11 YEARS WITH GENERALIZED SOMATIC SYMPTOMS AND PROVEN NICKEL ALLERGY WITH THE RIGHT ESSURE IN DISTAL POSITION. DIAGNOSTIC RESULTS (NORMAL RANGES ARE PROVIDED IN PARENTHESIS IF AVAILABLE): BODY MASS INDEX: 34 KG/SQM. [BLOOD PRESSURE MEASUREMENT] IN (B)(6) 2025: 99/53 [BLOOD TEST] IN (B)(6) 2025: A HYPERLEUKOCYTOSIS AT 27,440 [BODY TEMPERATURE] IN (B)(6) 2025: 38.4 [C-REACTIVE PROTEIN] IN (B)(6) 2025: 295 [COAGULATION TEST] IN (B)(6) 2025: NORMAL [COMPUTERISED TOMOGRAM ABDOMEN] ON (B)(6) 2025: A VERY LIMITED COLLECTION OF FLUID SUGGESTIVE OF AN ABSCESS 40 MM IN DIAMETER IN THE POUCH OF DOUGLAS WITH PROCTOSIGMOIDITIS AND ILEITIS REACTIVE ON CONTACT [CREATININE RENAL CLEARANCE] IN (B)(6) 2025: 75 [LIVER FUNCTION TEST] IN (B)(6) 2025: NORMAL [URINE CYTOLOGY] IN (B)(6) 2025: DID NOT FIND A URINARY INFECTION. THE FOLLOWING AMENDMENT WAS MADE: COMPLEMENTARY INFORMATION TO COMPANY CAUSALITY COMMENT AND UPDATE OF CORRECTIVE ACTION. NO NEW FOLLOW-UP INFORMATION WAS RECEIVED FROM THE REPORTER. CASE COMMENTS: THIS REPORT DESCRIBES THE OCCURRENCE OF A FATAL CARDIO-RESPIRATORY ARREST, DOUGLAS' ABSCESS SEPSIS, COLITIS AND ALLERGY TO METALS IN A 51-YEAR-OLD WOMAN WHO HAD ESSURE IMPLANTED ON (B)(6) 2014 AND THAT WAS REMOVED ON (B)(6) 2025. WITH THE EXCEPTION OF ALLERGY TO METALS THAT IS LISTED FOR ESSURE, ALL OTHER EVENTS ARE UNLISTED TO THE DEVICE. REGARDING THE ALLERGY TO METALS, COMPANY AGREES WITH THE REPORTER AND CONSIDERS THE EVENT RELATED TO ESSURE. IT WAS REPORTED THAT THE CARDIO-RESPIRATORY ARREST WAS A COMPLICATION OF THE DOUGLAS' ABSCESS SEPSIS. THIS SEPSIS CAN ALSO EXPLAIN THE COLITIS. CONSIDERING THE CLINICAL EVOLUTION OF THE PATIENT WHO PRESENTED THE FIRST SYMPTOMS OF SEPSIS A FEW DAYS AFTER THE REMOVAL PROCEDURE OF ESSURE, THE EVENTS THAT LED TO PATIENT DECEASE ARE PROBABLY A CONSEQUENCE OF AN INFECTION ACQUIRED DURING THE PROCEDURE. THEREFORE, COMPANY ASSESSES CARDIO-RESPIRATORY ARREST, DOUGLAS' ABSCESS SEPSIS, AND COLITIS AS UNRELATED TO ESSURE. SINCE THESE EVENTS ARE CONSIDERED UNRELATED TO THE DEVICE NO CORRECTIVE ACTIONS ARE DEEMED REQUIRED. BASED ON THE AVAILABLE INFORMATION, A REVIEW OF OUR COMPLAINT RECORDS AND OTHER RELEVANT DATA WILL BE CONDUCTED; ANY NEW AND REPORTABLE INFORMATION THAT BECOMES AVAILABLE FROM OUR INVESTIGATION WILL BE PROVIDED IN A SUPPLEMENTARY REPORT. IN CASE A SAMPLE IS RECEIVED THE INVESTIGATION MIGHT LEAD TO DESTRUCTION OF THE SAMPLE IF REQUIRED TO PERFORM A PROPER ANALYSIS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22913090&pc=DSQ | 2916596-2025-05588 | 2023-06-04 04:00 | Death | THORATEC CORPORATION | 2025-08-28 04:00 | DSQ | HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM | Adverse Event Without Identified Device or Use Problem | Renal Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY DUE TO END STAGE RENAL FAILURE. THE ONSET DATE OF THE RENAL DYSFUNCTION WAS ESTIMATED TO BE (B)(6) 2023. THE PATIENT EXPERIENCED SYMPTOMS OF RISING CREATININE LEVELS. THIS EVENT WAS NOT CONSIDERED TO BE DEVICE RELATED. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. D, IS CURRENTLY AVAILABLE. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING RENAL FAILURE AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22913303&pc=NIM | 2124215-2025-59928 | 2025-07-30 04:00 | Death | SILK ROAD MEDICAL | 2025-08-28 04:00 | NIM | ENROUTE TRANSCAROTID STENT | Adverse Event Without Identified Device or Use Problem | Eye Injury; Fatigue; Hemorrhagic Stroke |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED A BRAIN HEMORRHAGE, WHICH ULTIMATELY RESULTED IN DEATH. AN ENROUTE TRANSCAROTID STENT WAS SELECTED FOR USE IN A RIGHT SIDED TRANSCAROTID ARTERY REVASCULARIZATION (TCAR) PROCEDURE. POST TCAR PROCEDURE, THE PATIENT HAD A DILATED RIGHT EYE PUPIL AND WAS LETHARGIC. IMAGING REVEALED A BRAIN HEMORRHAGE. NO ADDITIONAL SURGICAL INTERVENTION WAS PERFORMED. THE PATIENT WAS PLACED ON COMFORT CARE AND ULTIMATELY PASSED AWAY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22913338&pc=DSQ | 2916596-2025-05681 | 2025-08-19 04:00 | Death | THORATEC CORPORATION | 2025-08-28 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN ETIOLOGY. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. IT WAS REPORTED THAT THE PATIENT'S OUTCOME WAS NOT DEVICE OR THERAPY RELATED. THE DEVICE WOULD NOT BE EXPLANTED AND WOULD NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT'S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICLE ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REVISION, REV. B IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22913647&pc=QGM | 3027664504-2025-00018 | 2025-07-29 04:00 | Death | HISTOSONICS, INC | 2025-08-28 04:00 | QGM | EDISON | Adverse Event Without Identified Device or Use Problem | Aspiration/Inhalation; Head Injury; Low Blood Pressure/ Hypotension; Necrosis; Pneumonia; Urinary Retention; Hematuria; Aspiration Pneumonitis | K233466 |  | 1 | Event Description: ON (B)(6) 2025, A 79-YEAR-OLD FEMALE PATIENT WITH METASTATIC GALLBLADDER CANCER UNDERWENT HISTOTRIPSY TO THREE HEPATIC METASTASES ACROSS FOUR TREATMENT TARGETS IN THE RIGHT LIVER FOR A TOTAL PLANNED TREATMENT VOLUME (PTV) OF 96.38 CC. PAST MEDICAL HISTORY INCLUDED CHOLECYSTECTOMY, PULMONARY HYPERTENSION, PREMATURE VENTRICULAR CONTRACTIONS, AND VENOUS THROMBOEMBOLISM (ON LONG-TERM XARELTO). THE ANESTHESIOLOGIST ASSESSED THE PATIENT AS MODERATE RISK. THE PROCEDURE ITSELF WAS UNEVENTFUL APART FROM TRANSIENT HYPOTENSION AND OLIGURIA. BY POST-OPERATIVE DAY (POD) 1, OLIGURIA PERSISTED AND NEPHROLOGY NOTED HEMATURIA AND PROTEINURIA. ON POD2, LEUKOCYTOSIS DEVELOPED, BLOOD CULTURES WERE OBTAINED, AND CREATININE ROSE WHICH INDICATED ACUTE KIDNEY INJURY. DURING THIS PERIOD, THE PATIENT FELL AND WAS FOUND TO BE HYPOTENSION AND MINOR HEAD TRAUMA (HEAD CT NORMAL). A CHEST/ABDOMEN/PELVIS (NON-CONTRAST) CT REVEALED ASPIRATION PNEUMONIA, RIGHT PNEUMOTHORAX, PULMONARY HYPERTENSION, PNEUMOBILIA, INTRAHEPATIC GAS, AND A NECROTIC HEPATIC COLLECTION, ALONG WITH EXTENSIVE HEPATIC AND OSSEOUS METASTASES. BY POD3, BLOOD CULTURES GREW STREPTOCOCCUS AGALACTIAE (GROUP B) AND ESCHERICHIA COLI. ON POD4, THE PATIENT PROGRESSED TO SEPTIC SHOCK WITH RESPIRATORY FAILURE. LEUKOCYTOSIS WORSENED WITH THE SOURCE THOUGHT TO BE THE ASPIRATION PNEUMONIA VERSUS THE NECROTIC LIVER COLLECTION. DRAINAGE OF THE HEPATIC COLLECTION WAS PLANNED, BUT THE PATIENT DETERIORATED FURTHER, AND FOLLOWING FAMILY DISCUSSIONS, CARE WAS TRANSITIONED TO COMFORT MEASURES. THE PATIENT DIED LATER THAT DAY. THE TREATING PHYSICIAN ATTRIBUTED RENAL DYSFUNCTION TO ACUTE KIDNEY INJURY FROM THE PROCEDURE AND/OR SEPSIS-RELATED HYPOTENSION. PNEUMONIA, PNEUMOTHORAX, AND LIVER NECROSIS/ABSCESS WERE CONSIDERED PROCEDURE-RELATED, THOUGH ABSENCE OF RECENT IMAGING LIMITED DETERMINATION OF WHETHER PNEUMONIA OR ABSCESS PREDATED HISTOTRIPSY. FOR THE KIDNEY FAILURE, NEPHROLOGISTS CONCLUDED THAT THERE WAS NON-NEPHROTIC RANGE PROTEINURIA. ADDITIONAL RENAL STUDIES OBTAINED WERE POSITIVE FOR ANTI-NUCLEAR ANTIBODIES, WITH LOW C3 AND C4. WIDE, NON-SPECIFIC DIFFERENTIAL DIAGNOSIS, INCLUDED "BACTERIAL INFECTIONS, VIRAL INFECTIONS, SOME MALIGNANCIES, AUTOIMMUNE CONDITIONS (SYSTEMIC LUPUS ERYTHEMATOSUS, MIXED CONNECTIVE TISSUE DISEASE, SJOGREN'S SYNDROME, AND RHEUMATOID ARTHRITIS), HEREDITARY ANGIOEDEMA, AND CRYOGLOBULINEMIA." Manufacturer Narrative: NO DEVICE MALFUNCTIONS OR OTHER NOTEWORTHY EVENTS OCCURED DURING THE CASE. IN RESPONSE TO REPORTS OF ACUTE KIDNEY INJURY ASSOCIATED WITH HISTOTRIPSY, HISTOSONICS HAS INCORPORATED THE FOLLOWING WARNING INTO ITS LABELING: "AS REPORTED IN OTHER LIVER-DIRECTED THERAPIES, MULTIPLE OR LARGE-VOLUME LIVER TREATMENTS HAVE BEEN ASSOCIATED WITH ACUTE KIDNEY INJURY (AKI). WHEN PLANNING EXTENSIVE OR MULTIPLE TREATMENTS, EVALUATE AND MONITOR THE PATIENT FOR SIGNS OF AKI."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22913668&pc=NPT | 9612164-2025-04259 | 2025-08-11 04:00 | Death | MEDTRONIC IRELAND | 2025-08-28 04:00 | NPT | EVOLUT FX DCS | Patient Device Interaction Problem | Hemorrhage/Bleeding; Low Blood Pressure/ Hypotension; Rupture | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) PROCEDURE, FOLLOWING THE SUCCESSFUL IMPLANT OF THE VALVE, THE DELIVERY CATHETER SYSTEM (DCS) WAS WITHDRAWN AND THE PATIENT'S PRESSURE DROPPED IMMEDIATELY. SUBSEQUENTLY, AT THE ACCESS SITE OF THE ILIAC ARTERY, BLEEDING AND A RUPTURE OCCURRED. DUE TO THE SIGNIFICANT AMOUNT OF BLOOD LOSS, THE PATIENT DIED. IT WAS NOTED THAT CAUSE OF DEATH WAS DUE TO THE VASCULAR COMPLICATIONS. Manufacturer Narrative: CONTINUATION OF D10: SECTION D INFORMATION REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER RELEVANT DEVICE(S) ARE: PRODUCT ID: EVFXPLUS-23, SERIAL/LOT #: (B)(6), UBD: 14-NOV-2026, UDI#: (B)(4). H6. THE CODES PRESENT IN SECTION H6 CORRESPOND TO MULTIPLE COMPONENTS/PRODUCTS THAT COMPRISE THIS REPORTED EVENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22913823&pc=DXF | 2124215-2025-60572 | 2025-08-11 04:00 | Death | BAYLIS MEDICALE CIE INC | 2025-08-28 04:00 | DXF | VERSACROSS CONNECT LAAC ACCESS SOLUTION | Failure to Advance | Low Blood Pressure/ Hypotension; Pericardial Effusion; Thrombosis/Thrombus | K150709 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT DEATH OCCURRED. THE PROCEDURE WAS NOT COMPLETED. DURING A LEFT ATRIAL APPENDAGE CLOSURE (LAAC) PROCEDURE, A VERSACROSS CONNECT LAAC WAS SELECTED FOR USE. THE PATIENT UNDERWENT GENERAL ANESTHESIA FOR WATCHMAN DEVICE PLACEMENT AND PRE-IMAGING WAS OBTAINED ON TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE), WITH NO PERICARDIAL EFFUSION REPORTED. THE GROIN ACCESS WAS OBTAINED ON BOTH RIGHT AND LEFT GROIN. A NON-BOSTON SCIENTIFIC INTRACARDIAC ECHOCARDIOGRAPHY (ICE) CATHETER WAS INTRODUCED IN THE LEFT GROIN AND WATCHMAN SHEATH WITH VERSACROSS CONNECT WAS INTRODUCED IN RIGHT GROIN. THE VERSACROSS RADIO FREQUENCY (RF) WIRE WAS ADVANCED INTO SUPERIOR VENA CAVA (SVC) FOLLOWED BY VERSACROSS CONNECT DILATOR AND WATCHMAN SHEATH LOADED OVER THE WIRE AND DROPPED ONTO THE FOSSA OVALIS. HEPARIN WAS GIVEN. AFTER CONFIRMING THE TRANSSEPTAL LOCATION ON ICE VIA VISUALIZATION OF TENTING, RF WAS APPLIED AT ONE (1) SECOND CONSTANT SETTING. TRANSEPTAL WAS UNSUCCESSFUL AND IT WAS DETERMINED THAT THERE WAS MORE REACH AND BETTER CONTACT REQUIRED. THE PIGTAIL WIRE WAS ADVANCED BACK INTO THE SVC AND THE SHEATH/ DILATOR WERE REMOVED AND RE-SHAPED. THE SYSTEM WAS ADVANCED BACK INTO THE SVC AND WAS DROPPED ONTO THE FOSSA OVALIS. THIS TIME A BETTER TENTING WAS VISUALIZED ON ICE AND RF WAS APPLIED TWICE WITH SAME SETTING AND STILL WAS AN UNSUCCESSFUL CROSS. THE RF GENERATOR SETTING WAS CHANGED TO TWO (2) SECONDS AND PULSE. THE RF WAS APPLIED AND WAS UNSUCCESSFUL. THE PHYSICIAN THEN DECIDED TO SWITCH TO A NON-BOSTON SCIENTIFIC SHEATH TO BE USED WITH VERSACROSS RF WIRE. THE NON-BOSTON SCIENTIFIC SHEATH WAS INTRODUCED, DROPPED ONTO THE FOSSA OVALIS AND VISUALIZED ON ICE. THE RF WAS APPLIED AT TWO (2) SECONDS PULSE SETTING. THE ATTEMPT WAS UNSUCCESSFUL. THEN, A NON-BOSTON SCIENTIFIC TRANSSEPTAL NEEDLE WAS OPENED TO BE USED WITH THAT SAME NON-BOSTON SCIENTIFIC SHEATH. A MANUAL TRANSEPTAL PUNCTURE WAS ATTEMPTED AND UNSUCCESSFUL. THE PHYSICIAN THEN DECIDED TO USE A NON-BOSTON SCIENTIFIC RF DEVICE. AFTER INTRODUCING THE WIRE TO THE NON-BOSTON SCIENTIFIC SHEATH AGAIN, THE PHYSICIAN APPLIED RF TO THE WIRE AND IT ADVANCED INTO THE LEFT ATRIUM (LA). THE NON-BOSTON SCIENTIFIC SHEATH/DILATOR WERE ADVANCE INTO THE LA OVER THE WIRE. THE WIRE WAS SWITCHED OUT FOR A TORAY WIRE AND THE SHEATH/DILATOR WERE REMOVED AND THE WATCHMAN SHEATH AND DILATOR WERE INTRODUCED AND ADVANCED INTO THE LA OVER THE TORAY WIRE. A NON-BOSTON SCIENTIFIC PIGTAIL CATHETER WAS INTRODUCED THROUGH WATCHMAN SHEATH AND WAS ADVANCED INTO THE LEFT ATRIAL APPENDAGE UNDER VISUALIZATION ON TEE. A CONTRAST SHOT WAS PERFORMED. A 27 MM WATCHMAN DEVICE WAS INTRODUCED TO THE WATCHMAN SHEATH AFTER PIGTAIL REMOVAL USING A WET TO WET CONNECTION. AFTER FORMED A FLEX BALL, THE DEVICE WAS DEPLOYED AND VISUALIZED ON TEE. IT WAS DETERMINED DEVICE PLACEMENT WAS TOO PROXIMAL SO IT WAS RECAPTURED PARTIALLY AND MOVED INTO AN ANTERIOR LOBE AND DEPLOYED. AGAIN, THE DEPLOYMENT WAS TOO PROXIMAL SO IT WAS PARTIALLY RECAPTURED AGAIN. AFTER THE SECOND PARTIAL RECAPTURE, THERE WAS A NOTICE OF DROP IN BLOOD PRESSURE. THEY SWEPT FOR EFFUSION ON TEE AND VISUALIZED A CONSIDERABLE EFFUSION AT THIS POINT. IT WAS DECIDED TO FULLY RECAPTURE THE WATCHMAN DEVICE AND REMOVE THE SYSTEM ENTIRELY AND TAP THE PATIENT. HEPARIN WAS REVERSED AND A CARDIAC THORACIC (CT) SURGEON WAS CALLED. BY THE TIME THEY WERE ABLE TO TAP THE PERICARDIUM, THROMBOSIS HAD FORMED. CT SURGEON CAME IN AND ATTEMPTED TO RESCUE THE PATIENT BY PERFORMING CARDIAC SURGERY, BLOOD TRANSFUSION, CHEST COMPRESSIONS BUT WERE UNSUCCESSFUL. THE PATIENT EXPIRED. IT WAS FURTHER CONFIRMED THAT THE PATIENT HAD A SMALL FOSSA OVALIS, AND A POOR ICE IMAGINE FROM ICE OPERATOR WAS NOTED. THE PHYSICIAN DID NOT BELIEVE VERSACROSS SYSTEM DIRECTLY CAUSED EFFUSION, IT WAS SUSPECTED DELIVERING A NON-BOSTON SCIENTIFIC ELECTROSURGERY DEVICE THROUGH WIRE IN A THICK PART OF SEPTUM MAY HAVE CAUSED IT. THE DEVICE IS NOT EXPECTED TO BE RETURNED FOR ANALYSIS (DISPOSED). Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION, BUT IT WAS NOT AVAILABLE. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION. GOOD FAITH EFFORT ATTEMPTS WERE MADE TO TRY AND RETRIEVE ADDITIONAL DETAILS REGARDING THE REPORTED EVENT, BUT FURTHER INFORMATION WAS UNABLE TO BE OBTAINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22913831&pc=NPT | 2025587-2025-05562 | 2024-05-18 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-08-28 04:00 | NPT | COREVALVE TRANSCATHETER AORTIC VALVE | Insufficient Information | Insufficient Information | P130021 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING LONG-TERM STROKE AND MORTALITY RISK IN NONAGENARIANS AFTER TRANSCATHETER AORTIC VALVE IMPLANTATION (TAVI). THE STUDY POPULATION INCLUDED 187 PATIENTS WHO WERE PREDOMINANTLY MALE WITH A MEAN AGE OF (B)(6) YEARS OLD. MULTIPLE MANUFACTURERS¿ DEVICES WERE IMPLANTED IN THE STUDY POPULATION; 25 PATIENTS WERE IMPLANTED WITH A MEDTRONIC COREVALVE BIOPROSTHETIC VALVE. THREE PROCEDURE-RELATED DEATHS AND 150 DEATHS DURING FOLLOW-UP OCCURRED IN THE STUDY POPULATION. NO FURTHER DETAILS WERE PROVIDED REGARDING THESE DEATHS. AMONG ALL PATIENTS, OTHER CLINICAL OBSERVATIONS INCLUDED: ARRHYTHMIA REQUIRING PERMANENT PACEMAKER IMPLANT, BLEEDING/VASCULAR COMPLICATION, STROKE, MODERATE OR GREATER PARAVALVULAR REGURGITATION, AND REPEAT HOSPITALIZATION. NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. Manufacturer Narrative: CITATION: JUAREZ-CASSO ET AL. LONG-TERM STROKE AND MORTALITY RISK IN NONAGENARIANS AFTER TRANSCATHETER AORTIC VALVE INSERTION. ANN THORAC SURG. 2024 NOV;118(5):1035-1042. DOI: 10.1016/J.ATHORACSUR.2024.04.030. EPUB 2024 MAY 18. EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT. NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22913984&pc=NPT | 2135147-2025-04840 | 2025-08-05 04:00 | Death | ABBOTT MEDICAL | 2025-08-28 04:00 | NPT | NAVITOR VISION | Migration or Expulsion of Device; Improper or Incorrect Procedure or Method; Patient-Device Incompatibility | Vascular Dissection; Pericardial Effusion | P190023 |  | 1 | Event Description: IT WAS REPORTED ON 05 AUGUST 2025 A 27MM NAVITOR VISION VALVE WAS SELECTED FOR IMPLANT USING A LARGE FLEXNAV DELIVERY SYSTEM. THE PATIENT HAD HORIZONTAL ANATOMY. THE PATIENT'S ANNULUS WAS MEASURED WITH AN AVERAGE DIAMETER OF 24MM, AN AREA OF 442.8MM^2, AND A PERIMETER OF 75.1MM. THE LEFT VENTRICULAR OUTFLOW TRACT (LVOT) AVERAGE DIAMETER WAS 23.3MM. THE PATIENT'S ACTIVATED CLOTTING TIME (ACT) LEVEL WAS GREATER THAN 250 SECONDS. HEPARIN WAS ADMINISTERED. A PRE-BALLOON AORTIC VALVULOPLASTY (BAV) WAS PERFORMED WITH A 23MM NON-ABBOTT BALLOON. DURING THE PROCEDURE THE VALVE WAS RE-SHEATHED TWICE. THE PATIENT'S AORTIC VALVE ANGLE WAS 56 DEGREES. THE VALVE WAS IMPLANTED. THE FINAL IMPLANT DEPTH WAS 3-4MM FROM THE NON-CORONARY CUSP (NCC) AND BELOW THE ANNULUS ON THE LEFT CORONARY CUSP (LCC). ECHOCARDIOGRAM SHOWED A GRADIENT OF 12MMHG. AN AORTIC FLUOROSCOPY SHOWED THE VALVE MIGRATED ABOVE THE ANNULUS. THE DECISION WAS MADE TO SNARE THE VALVE ABOVE SINOTUBULAR JUNCTION (STJ) AND IMPLANT A SECOND VALVE. THE VALVE WAS SNARED BY THE INNER CURVATURE TAB WHICH THEN DUG INTO THE ASCENDING AORTA ANATOMY. A PERICARDIAL EFFUSION WAS THEN OBSERVED ON ECHOCARDIOGRAM. THE PERICARDIAL EFFUSION LED TO A CARDIAC TAMPONADE. THE PATIENT WAS INTUBATED AND A TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) AND ANGIOGRAPHY SHOWED A DISSECTION ABOVE THE VALVE. A PERICARDIOCENTESIS WAS PERFORMED, HOWEVER THE PATIENT DID NOT SHOW SIGNS OF IMPROVEMENT. AT THAT MOMENT 30-45 MINUTES HAD PASSED. THE PATIENT'S FAMILY WAS CONSULTED AND THE DECISION WAS MADE TO STOP TREATMENT. THE PATIENT PASSED AWAY. THE OFFICIAL CAUSE OF DEATH WAS THE AORTIC DISSECTION. Manufacturer Narrative: IT IS UNKNOWN IF THE DEVICE IS RETURNING FOR ANALYSIS. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22914071&pc=QZI | MW5175336 | 2025-08-13 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-08-28 04:00 | QZI | FARAWAVE CATHETER - FARAWAVE PULSED FIELD ABLATION (PFA) | Unintended Movement | Cardiac Arrest; Tachycardia |  |  | 1 | Event Description: IT WAS REPORTED BY THE CALLER THAT THE PATIENT HAD AN ADVERSE EVENT. THE CALLER REPORTED THAT THE PATIENT HAD A CARDIAC ARREST. THE CALLER STATED THAT THE PROCEDURE WAS PFA WITH THE BOSTON SCIENTIFIC FARAPULSE SYSTEM. THE CALLER REPORTED THE LEFT ATRIAL ANATOMY WAS MAPPED AFTER A DIFFICULT TRANSEPTAL PUNCTURE. A MAP SHIFT WAS NOTED WHEN THE FARAWAVE CATHETER WAS ADVANCED INTO THE LEFT ATRIUM. THE CATHETER APPEARED TO BE POSTERIOR AND INFERIOR TO THE CARTO 3 MAP. THE FARAPULSE CATHETER WAS REMOVED AND A NEW MAP WAS CREATED USING THE OCTARAY CATHETER. THE CALLER STATED THAT TWO CARTO MAPS ALIGNED. THE PROCEDURE CONTINUED WITH 36 PFA APPLICATIONS FROM THE FARAPULSE CATHETER. THE CALLER DID NOT STATE IF THE MAP SHIFT HAD RESOLVED. THE CALLER STATED THAT THERE WAS SIGNIFICANT METAL IN THE PATIENTS BACK THAT COULD HAVE CONTRIBUTED TO THE MAP SHIFT. INTRACARDIAC ECHO WAS USED TO CONFIRM THAT THE PATIENT DID NOT HAVE AN EFFUSION POST PROCEDURE. THE PATIENT WAS STABLE ON TRANSFER TO RECOVERY. IT WAS REPORTED BY THE CALLER THAT APPROXIMATELY FIVE MINUTES AFTER THE PATIENT WAS TRANSFERRED FROM THE ROOM POST PROCEDURE THE STAFF AND ANESTHESIA BROUGHT THE PATIENT BACK IN FULL CARDIAC ARREST. CPR WAS INITIATED, THE PATIENT WAS INTUBATED AND MEDICATIONS WERE GIVEN TO RESUSCITATE THE PATIENT. THE PATIENT WAS FULLY RESUSCITATED. THE CALLER REPORTS A TEE WAS COMPLETED TO CONFIRM THAT THERE WAS NO CARDIAC EFFUSION. THE CALLER REPORTS THAT THE PATIENT WENT INTO VENTRICULAR TACHYCARDIA WHILE BEING PREPPED FOR TRANSFER. THE PATIENT WAS DEFIBRILLATED AND FULLY RECOVERED. THE PATIENT WAS STABLE ON TRANSFER TO ICU FOR FURTHER MONITORING. THE CALLER STATED THAT THE PHYSICIAN FELT THAT THE CARDIAC ARREST/VENTRICULAR ARRHYTHMIA WAS TIED TO PREVIOUS SHORTER EPISODES OF VENTRICULAR TACHYCARDIA. THE PHYSICIAN STATED THAT THERE WERE PREVIOUS SHORT VENTRICULAR EVENTS RECORDED ON THE PATIENT'S BIOTRONIC PACEMAKER. BWI CATHETERS USED DURING THE PROCEDURE: OCTARAY CATHETER F CURVE 2-2-2-2-2 SPACING (D160904 LOT UNKNOWN), CS D/F CURVE (BD710DF282CT LOT UNKNOWN), SIEMANS 8FR SOUNDSTAR (10439011 LOT UNKNOWN). NON-BWI CATHETERS USED: (B)(6) FARADRIVE SHEATH, (B)(6) FARAPULSE CATHETER, (B)(6) VERSACROSS TRANSEPTAL NEEDLE. AFTERWARD, IT WAS REPORTED BY THE CALLER THAT THE PATIENT PASSED AWAY ON (B)(6) 2025. THE CALLER DOES NOT KNOW THE CAUSE OF DEATH. THE CALLER STATED THAT HE WAS INFORMED THAT THE PATIENT HAD "BLOOD IN HER LUNGS" BUT HAD NO OTHER INFORMATION. THIS REPORT REFLECTS INFORMATION RECEIVED BY FDA IN THE FORM OF A NOTIFICATION PER 803.22 (B)(2).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22914083&pc=QZI | MW5175338 | 2025-08-08 04:00 | Death | MEDTRONIC, INC. | 2025-08-28 04:00 | QZI | PULSESELECT PULSED FIELD ABLATION (PFA) SYSTEM | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Foreign Body In Patient |  |  | 1 | Event Description: DURING AN ATRIAL FIBRILLATION PROCEDURE, THE PATIENT'S BLOOD PRESSURE RAPIDLY DROPPED DURING ABLATION. (B)(6) ICE AND MEDTRONIC PULSE SELECT CATHETERS WERE IN THE BODY AT THIS POINT IN THE CASE AND CODE WAS CALLED. AFTER AN HOUR AND A HALF OF ATTEMPTED RESUSCITATION THE PATIENT WAS PRONOUNCED DEAD. THIS REPORT REFLECTS INFORMATION RECEIVED BY FDA IN THE FORM OF A NOTIFICATION PER 803.22 (B)(2).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22914202&pc=CBK | 3013095415-2025-00767 | 2025-08-04 04:00 | Death | VENTEC LIFE SYSTEMS, INC | 2025-08-28 04:00 | CBK | VOCSN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K162877 |  | 1 | Event Description: IT WAS REPORTED TO VENTEC THAT A PATIENT PASSED AWAY WHILE ON THE DEVICE, AND THAT NO ALARMS WENT OFF. THE PATIENT'S CAREGIVER ASKED "IF IT [ALARMS] COULD HAVE BEEN PUT ON SILENCE OR NOT." THE CARETAKER DID NOT ALLEGE A SPECIFIC DEVICE OR USE PROBLEM, NOR DID THE THEY ADVISE WHETHER OR NOT THEY HAD EXPECTED AN ALARM. NO FURTHER DETAILS WERE PROVIDED. VENTEC HAS MADE MULTIPLE ATTEMPTS TO CONTACT THE DME, BOTH IN WRITING AND VIA PHONE, IN ORDER TO OBTAIN ADDITIONAL INFORMATION ABOUT THE PATIENT AND THE EVENT, BUT NO RESPONSE HAS BEEN RECEIVED. Manufacturer Narrative: H6: VENTEC HAS MADE MULTIPLE ATTEMPTS TO CONTACT THE DME FOR ADDITIONAL INFORMATION, HOWEVER, NO RESPONSE HAS BEEN RECEIVED. THE DEVICE'S CURRENT DISPOSITION IS UNKNOWN. A REVIEW OF THE MANUFACTURING RECORDS SHOWED ALL REQUIREMENTS WERE MET. FINAL TEST VERIFICATION SPECIFICATIONS ARE ACCEPTABLE. NO NON-CONFORMITIES OR ANOMALIES FOUND RELATED TO THIS COMPLAINT WHEN REVIEWING THE DEVICE HISTORY RECORD. THE DEVICE HAS NOT BEEN RETURNED TO VENTEC FOR EVALUATION. WITH THE INFORMATION AVAILABLE, VENTEC IS UNABLE TO DETERMINE IF THE DEVICE USE MAY HAVE CAUSED OR CONTRIBUTED TO THE PATIENT'S OUTCOME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22914369&pc=PNJ | 9612164-2025-04261 | 2025-08-11 04:00 | Death | MEDTRONIC IRELAND | 2025-08-28 04:00 | PNJ | MICRA¿ AV2 | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Cardiac Perforation; Pericardial Effusion | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT INTRA OPERATIVELY DURING THE PLACEMENT OF THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG), DUE TO COMPLEX ANATOMY, WHENEVER GOOSENECK WAS APPLIED THE DEVICE SLIPPED FOR MULTIPLE ATTEMPTS. PATIENT EXPERIENCED CARDIAC PERFORATION AND PERICARDIAL EFFUSION. BLOOD PRESSURE WAS LOW. MEDICATION WAS PROVIDED. CODE BLUE WAS CALLED. CARDIOPULMONARY RESUSCITATION (CRP) AND INTUBATION WAS PERFORMED. LATER, IT WAS NOTED THAT THE PATIENT PASSED AWAY. Manufacturer Narrative: THIS IS A SYSTEM REPORT. SECTION D REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: BRAND NAME [MICRA] PRODUCT ID [MC2AVR1-DELSYS] (SERIAL: [MVD663855E]). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22914450&pc=NGV | 2124215-2025-60264 | 2025-08-19 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-08-28 04:00 | NGV | SYSTEM, APPENDAGE CLOSURE, LEFT ATRIAL | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P130013 |  | 1 | Event Description: THE FOLLOWING EVENT ORIGINATED FROM SOCIAL MEDIA, AND THEREFORE DESPITE GOOD FAITH EFFORTS THE INFORMATION IS OFTEN INCOMPLETE. ADDITIONALLY, DUE TO THE LIMITED INFORMATION RECEIVED THROUGH GOOD FAITH EFFORTS, THIS EVENT MAY REPRESENT A DUPLICATE OF AN EVENT WHICH WAS ALREADY KNOWN TO BOSTON SCIENTIFIC. IT WAS REPORTED THAT PATIENT DEATH OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS PERFORMED, AND A WATCHMAN LAA CLOSURE DEVICE WAS IMPLANTED. THE PATIENT WAS DISCHARGED. PER REPORT, THE PATIENT PASSED AWAY WITHIN ONE MONTH AFTER THE CLOSURE DEVICE WAS IMPLANTED. NO FURTHER INFORMATION IS AVAILABLE AT THE TIME OF THIS REPORT. Manufacturer Narrative: B3: AWARE DATE WAS USED AS EVENT DATE AS ACTUAL EVENT DATE WAS NOT KNOWN. D1: BRAND NAME IS BLANK BECAUSE THE DEVICE IS KNOWN TO BE A WATCHMAN LAA CLOSURE DEVICE, BUT THE EXACT BRAND NAME IS UNKNOWN. D4: DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UDI AND OTHER PRODUCT SPECIFIC INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE A SUPPLEMENTAL REPORT WILL BE SUBMITTED
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22914625&pc=LWS | 2649622-2025-23619 | 2025-08-08 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-08-28 04:00 | LWS | SPRINT QUATTRO SECURE S MRI SURESCAN | Under-Sensing | Syncope/Fainting | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED TO THE EMERGENCY DEPARTMENT DUE TO SYNCOPE. THE RIGHT ATRIAL (RA) LEAD EXHIBITED UNDERSENSING ON STORED EPISODES. THE RIGHT VENTRICULAR (RV) LEAD EXHIBITED UNDERSENSING THAT WAS NOTED AS ONE BEAT BEFORE DETECTION. THE PATIENT PASSED AWAY THE FOLLOWING DAY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22915804&pc=BZD | 2518422-2025-109435 | 2025-08-18 04:00 | Death | RESPIRONICS, INC. | 2025-08-28 04:00 | BZD | PHILIPS CPAP DEVICE | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: THIS NOTIFICATION IS BEING REVIEWED DUE TO A USER OF PHILIPS CPAP DEVICE STATING THAT HIS WIFE PASSED AND HE WOULD WANT TO HAVE THE DEVICE REPAIRED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22916144&pc=DSQ | 2916596-2025-05707 | 2025-07-27 04:00 | Death | THORATEC CORPORATION | 2025-08-28 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WENT TO THE BATHROOM. THEIR SPOUSE HEARD VENTRICULAR ASSIST DEVICE ALARMS AND WENT TO CHECK ON THE PATIENT. THERE THEY FOUND THE PATIENT SLUMPED OVER UNCONSCIOUS IN THE BATHROOM WITH THEIR VENTRICULAR DEVICE ALARMING. EMERGENCY MEDICAL SERVICES WERE CALLED, AND THE PATIENT WAS TAKEN TO THE HOSPITAL WHERE THEY WERE THEN PRONOUNCED TO HAVE PASSED. THE COORDINATOR REPORTED THAT THE DEVICE WAS OPERATING AS EXPECTED AND WOULD NOT BE RETURNED. THE DETAILS ON THE ALARMS WERE NOT ABLE TO BE PROVIDED BY THE FAMILY. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22916154&pc=DSQ | 2916596-2025-05506 | 2025-08-06 04:00 | Death | THORATEC CORPORATION | 2025-08-28 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO SUDDEN INTRACRANIAL BLEEDING. NO HEAD TRAUMA WAS REPORTED. THE PUMP WORKED AS INTENDED DURING THE TIME OF SUPPORT. THE DEATH WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED, AND THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22916164&pc=LWS | 9614453-2025-03430 | 2025-08-20 04:00 | Death | MEDTRONIC EUROPE SARL | 2025-08-28 04:00 | LWS | COBALT¿ DR MRI SURESCAN¿ | Unable to Obtain Readings | Ventricular Fibrillation | P980016 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS DECEASED. IT WAS REPORTED THAT THE PATIENT SUFFERED A SPONTANEOUS VENTRICULAR FIBRILLATION (VF) EPISODE. WHICH WAS DETECTED BY THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) AND THERAPY WAS DELIVERED FOR A TOTAL OF TWELVE SHOCKS. IT WAS NOTED THAT SOME OF THE SHOCKS WERE EFFECTIVE, HOWEVER THE VF RECURRED EACH TIME. THE PHYSICIAN REPORTED THEY WERE UNABLE TO SEE THE LAST SHOCK AND INSTEAD OF THE ENERGY QUANTITY IN JOULES, THERE WERE QUESTION MARKS ON THE ELECTROGRAMS (EGM) RECORDING. THE DISPOSITION OF THE ICD IS UNKNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22916461&pc=NPT | 2015691-2025-07098 | 2025-08-07 04:00 | Death | EDWARDS LIFESCIENCES | 2025-08-28 04:00 | NPT | SAPIEN 3 TRANSCATHETER HEART VALVE | Device Stenosis | Cardiogenic Shock | P140031 |  | 1 | Event Description: AS REPORTED BY THE FIELD SPECIALIST, APPROXIMATELY 7 YEARS, 1 MONTH, AND 24 DAYS POST TRANSFEMORAL TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) PROCEDURE WITH A 20 MM SAPIEN 3 VALVE, THE PATIENT WAS IN THE INTENSIVE CARE UNIT (ICU) WITH CLASS IV HEART FAILURE DUE TO STENOSIS. ADDITIONAL INFORMATION WAS PROVIDED INDICATING THE PATIENT WAS IN CARDIOGENIC SHOCK WITH BILATERAL PULMONARY EMBOLISMS. A DECISION WAS MADE TO PLACE THE PATIENT ON PALLIATIVE CARE. SUBSEQUENTLY, THE PATIENT PASSED AWAY. Manufacturer Narrative: PER THE INSTRUCTION FOR USE (IFU), VALVE STENOSIS IS A POTENTIAL RISK ASSOCIATED WITH BIOPROSTHETIC HEART VALVES. PER THE VALVE ACADEMIC RESEARCH CONSORTIUM (VARC), VALVE STENOSIS CAN RESULT FROM A NUMBER OF FACTORS, INCLUDING PANNUS, CALCIFICATION, SUPPORT STRUCTURE DEFORMATION (OUT-OF-ROUND CONFIGURATION), TRAUMA (CARDIA-PULMONARY RESUSCITATION, BLUNT CHEST TRAUMA), ENDOCARDITIS, PROSTHETIC VALVE THROMBOSIS, AND NATIVE LEAFLET PROLAPSE IMPEDING PROSTHETIC LEAFLET MOTION. VALVE STENOSIS MAY RESULT IN SYMPTOMS SUCH AS SHORTNESS OF BREATH AND DECREASED EXERCISE TOLERANCE, WHICH MAY BE ACCOMPANIED BY AN INCREASED GRADIENT ACROSS THE VALVE. PER THE INSTRUCTIONS FOR USE (IFU), STRUCTURAL VALVE DETERIORATION (WEAR, FRACTURE, CALCIFICATION, LEAFLET TEAR/TEARING FROM THE STENT POSTS, LEAFLET RETRACTION, SUTURE LINE DISRUPTION OF COMPONENTS OF A PROSTHETIC VALVE, THICKENING, STENOSIS), AND DEVICE DEGENERATION ARE KNOWN POTENTIAL RISKS ASSOCIATED WITH BIOPROSTHETIC HEART VALVES. THE IFU CAUTIONS THAT ACCELERATED DETERIORATION OF THE VALVE MAY OCCUR IN PATIENTS WITH AN ALTERED CALCIUM METABOLISM. LONG-TERM DURABILITY HAS NOT BEEN ESTABLISHED FOR THE VALVE. REGULAR MEDICAL FOLLOW-UP IS ADVISED TO EVALUATE VALVE PERFORMANCE. STRUCTURAL VALVE DETERIORATION (SVD) MAY BE MANIFESTED AS STENOSIS WITH THICKENED LEAFLETS. SVD MAY BE MILD AND NOT REQUIRE ANY INTERVENTION OR IT MAY BE MODERATE TO SEVERE. IT CAN CAUSE THE HEART TO WORK HARDER TO EJECT BLOOD FROM THE VENTRICLE. DEPENDING ON THE SEVERITY IT COULD BE AN INDICATION FOR VALVE REPLACEMENT OR MEDICAL INTERVENTION. TISSUE CALCIFICATION IS A COMMON FAILURE MODE. THE MECHANISMS FOR BIOPROSTHETIC HEART VALVE TISSUE CALCIFICATION ARE NOT YET FULLY UNDERSTOOD. MANY FACTORS CAN CONTRIBUTE TO THE ONSET AND PROPAGATION OF CALCIFICATION INCLUDING PATIENT-RELATED (E.G., PATIENT AGE, DISEASE STATE, IMMUNE STATUS, AND OTHER CO-MORBIDITIES), PHARMACOLOGICAL, AND INTRINSIC PROPERTIES OF THE VALVE ITSELF. IT IS WIDELY UNDERSTOOD THAT PATIENTS WITH CHRONIC RENAL DISEASE AND PRIOR HISTORY OF CALCIFIC STENOSIS OF THE NATIVE VALVE MAY BE PREDISPOSED TO BIOPROSTHETIC CALCIFICATION. DURING THE MANUFACTURING PROCESS, THE SAPIEN TRANSCATHETER HEART VALVES (THV) ARE 100% VISUALLY INSPECTED FOR DEFECTS AND 100% FUNCTIONALLY TESTED FOR PROPER COAPTATION UNDER PHYSIOLOGICAL BACKPRESSURE CONDITIONS PRIOR TO RELEASE FOR DISTRIBUTION. THEREFORE, IT IS HIGHLY UNLIKELY THAT A MANUFACTURING DEFECT OR DEVICE MALFUNCTION CONTRIBUTED TO THE EVENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A MANUFACTURING NON-CONFORMANCE CONTRIBUTED TO THIS PRODUCT PROBLEM. ALTHOUGH A DEFINITIVE ROOT CAUSE WAS UNABLE TO BE DETERMINED, THE EVENT COULD BE RELATED TO THE MECHANISMS DESCRIBED ABOVE. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22917213&pc=LPB | 2029046-2025-02911 | 2025-04-24 04:00 | Death | BIOSENSE WEBSTER INC | 2025-08-28 04:00 | LPB | THERMOCOOL SMARTTOUCH SF | Adverse Event Without Identified Device or Use Problem | Cardiac Perforation | P030031/S078 |  | 1 | Event Description: THIS COMPLAINT IS FROM A LITERATURE SOURCE. THE FOLLOWING LITERATURE CITE HAS BEEN REVIEWED: KASAI Y, KITAI T, MORITA J, HORITA R, KASAI J, FUJITA T. DELAYED LEFT VENTRICULAR RUPTURE AFTER RADIOFREQUENCY CATHETER ABLATION OF PREMATURE VENTRICULAR CONTRACTIONS TRIGGERING VENTRICULAR FIBRILLATION. CJC OPEN. 2025 APR 24;7(7):952-954. DOI: 10.1016/J.CJCO.2025.04.013. PMID: 40698295; PMCID: PMC12277838. BACKGROUND: DELAYED LEFT VENTRICULAR RUPTURE CAN OCCUR AFTER RADIOFREQUENCY CATHETER ABLATION OF PREMATURE VENTRICULAR CONTRACTIONS TRIGGERING VENTRICULAR FIBRILLATION, EVEN IN THE ABSENCE OF PROCEDURAL COMPLICATIONS SUCH AS STEAM POPS, HIGHLIGHTING A RARE BUT LIFE-THREATENING LATE COMPLICATION. OBJECTIVE: THIS CASE HIGHLIGHTS A RARE BUT CATASTROPHIC COMPLICATION OF RFCA, EMPHASIZING THE NEED FOR VIGILANT POSTPROCEDURAL MONITORING TO DETECT DELAYED STRUCTURAL DAMAGE. METHODS: AN 84-YEAR-OLD MAN WITHOUT ORGANIC HEART DISEASE (MAGNETIC RESONANCE IMAGING WAS NOT PERFORMED BECAUSE OF CLAUSTROPHOBIA EXPERIENCED FREQUENT EPISODES OF VENTRICULAR FIBRILLATION (VF), RESULTING IN AN ELECTRICAL STORM. DESPITE ADMINISTRATION OF ANTIARRHYTHMIC DRUGS, THE VF EPISODES COULD NOT BE CONTROLLED. A SINGLE MORPHOLOGY OF PREMATURE VENTRICULAR CONTRACTIONS WAS CONSISTENTLY IDENTIFIED AS THE TRIGGER FOR VF, AND RADIOFREQUENCY CATHETER ABLATION (RFCA) WAS PERFORMED (FIG. 1, A AND B). DURING THE PROCEDURE, VF REPEATEDLY OCCURRED AGAIN (ELECTRICAL STORM). IMMEDIATE EXTRACORPOREAL MEMBRANE OXYGENATION WAS INITIATED FOR CIRCULATORY SUPPORT. RADIOFREQUENCY ENERGY WAS APPLIED TO THE EARLIEST ACTIVATION SITE OF THE PREMATURE VENTRICULAR CONTRACTIONS, LOCATED AT THE INFERIOR WALL OF THE LEFT VENTRICLE (LV), USING A 3.5-MM OPEN IRRIGATED CATHETER (THERMOCOOL SMARTTOUCH SURROUND FLOW; BIOSENSE WEBSTER, IRVINE, CA). CONCLUSIONS: MOREOVER, IT RAISES CONCERNS ABOUT THE INHERENT VULNERABILITY OF THE INFERIOR LV WALL5 TO RADIOFREQUENCY-INDUCED INJURY AND THE POTENTIAL EXACERBATION OF MYOCARDIAL STRESS DUE TO INCREASED INTRATHORACIC OR INTRACARDIAC PRESSURE DURING THE POSTPROCEDURAL PERIOD. CARVEDILOL HAD ALREADY BEEN INITIATED, BUT THE DOSE WAS RELATIVELY LOW (5 MG/D). MORE AGGRESSIVE BLOOD PRESSURE TITRATION MIGHT REDUCE RUPTURE RISK. LOT, MODEL, AND CATALOG NUMBER ARE NOT AVAILABLE, BUT THE SUSPECTED BWI DEVICE IS POSSIBLY ASSOCIATED WITH REPORTED ADVERSE EVENTS: A 3.5-MM OPEN IRRIGATED CATHETER (THERMOCOOL SMARTTOUCH SURROUND FLOW; BIOSENSE WEBSTER, IRVINE, CA). OTHER BWI PRODUCTS: NONE. NON-BWI DEVICES: NONE. ADVERSE EVENT(S) AND PROVIDED INTERVENTIONS FOR BWI DEVICES: QTY 1: 1 PATIENT OF AN 84-YEAR-OLD MALE, DEVELOPED SUDDEN-ONSET HYPOTENSION FOLLOWING DEFECATION. URGENT LEFT VENTRICULOGRAPHY REVEALED A RUPTURE OF THE LV INFERIOR WALL AT THE ABLATION SITE (FIGURE 2B; VIDEO 1 , VIEW VIDEO ONLINE). DESPITE EMERGENCY SURGICAL REPAIR, THE PATIENT DIED WITHIN 2 HOURS POST-PROCEDURE. ALTHOUGH LV RUPTURE FOLLOWING RFCA HAS BEEN PREVIOUSLY REPORTED TO OUR KNOWLEDGE, THIS IS THE FIRST DOCUMENTED CASE OF ¿DELAYED¿ LV RUPTURE FOLLOWING RFCA. THE RUPTURE WAS LIKELY CAUSED BY LOCALIZED MYOCARDIAL INJURY FROM RADIOFREQUENCY ENERGY, WHICH WEAKENED THE TISSUE AT THE ABLATION SITE. ADDITIONALLY, INCREASED INTRACARDIAC PRESSURE DUE TO STRAINING DURING DEFECATION, COMBINED WITH PRESERVED LV FUNCTION (EJECTION FRACTION: 64%), MAY HAVE CONTRIBUTED TO THIS FATAL EVENT. LV RUPTURE AFTER RFCA HAS BEEN PRIMARILY ASSOCIATED WITH CASES WHERE STEAM POPS OCCUR, WHICH ARE OFTEN ACCOMPANIED BY ABRUPT IMPEDANCE RISES. IN THIS CASE, ALTHOUGH IMPEDANCE DID NOT EXCEED BASELINE VALUES, TRANSIENT RISES WERE OBSERVED DURING APPLICATIONS. Manufacturer Narrative: B2. DATE OF DEATH - THE EXACT DATE OF DEATH WAS NOT SPECIFIED. HOWEVER, SINCE IT WAS STATED THAT THE PATIENT DIED WITHIN 2 HOURS POST-PROCEDURE, THE DATE OF DEATH WAS POPULATED WITH THE EVENT DATE. D4. CATALOG: UNK_SMART TOUCH UNIDIRECTIONAL SF. D4: UDI: AS THE CATALOG/MODEL NUMBER WAS NOT PROVIDED, THE (01)GTIN IS NOT AVAILABLE. THIS COMPLAINT IS FROM A LITERATURE SOURCE. THE FOLLOWING LITERATURE CITE HAS BEEN REVIEWED: KASAI Y, KITAI T, MORITA J, HORITA R, KASAI J, FUJITA T. DELAYED LEFT VENTRICULAR RUPTURE AFTER RADIOFREQUENCY CATHETER ABLATION OF PREMATURE VENTRICULAR CONTRACTIONS TRIGGERING VENTRICULAR FIBRILLATION. CJC OPEN. 2025 APR 24;7(7):952-954. DOI: 10.1016/J.CJCO.2025.04.013. PMID: 40698295; PMCID: PMC12277838. AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. AN EVALUATION OF THE MANUFACTURING RECORD COULD NOT BE PERFORMED AS THE REQUIRED PRODUCT IDENTIFICATION NUMBER WAS NOT PROVIDED TO COMPLETE THE EVALUATION. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. MANUFACTURER'S REFERENCE NUMBER: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22917263&pc=BZD | 2518422-2025-109443 | 2025-08-08 04:00 | Death | RESPIRONICS, INC. | 2025-08-28 04:00 | BZD | NO EQUIPMENT NUMBER | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: MANUFACTURER RECEIVED A CALL FROM THE PATIENT'S DAUGHTER STATING THAT THEY RECEIVED A REPLACEMENT FOR HER FATHER'S MACHINE, BUT HE HAS PASSED AWAY AND THEY DON'T NEED IT. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22917333&pc=NIK | 9614453-2025-03435 | 2025-08-21 04:00 | Death | MEDTRONIC EUROPE SARL | 2025-08-28 04:00 | NIK | VIVA¿ S CRT-D | Appropriate Term/Code Not Available | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P010031 |  | 1 | Event Description: IT WAS REPORTED THAT THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) MALFUNCTIONED, AND THE PATIENT PASSED AWAY. THE PATIENT'S CAUSE OF DEATH WAS REQUESTED, BUT HAS NOT BEEN RECEIVED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22917357&pc=QBJ | 3004753838-2025-240049 |  | Death | DEXCOM, INC. | 2025-08-28 04:00 | QBJ | DEXCOM G7 CONTINUOUS GLUCOSE MONITORING SYSTEM | No Audible Alarm | Hypoglycemic Shock | K213919 |  | 1 | Event Description: IT WAS REPORTED VIA EMAIL THAT THE PATIENT PASSED AWAY. ACCORDING TO THE EMAIL, THE PATIENT WITH TYPE 1 DIABETES PASSED AWAY AFTER BEING IN A COMA ON (B)(6) 2025. ON(B)(6) 2025, SHE WAS STILL WEARING HER SENSOR. WHEN THE DEXCOM MOBILE APPLICATION DATA WAS REVIEWED BY POLICE INVESTIGATOR, DATA WAS NO LONGER SEEN RECORDED AFTER (B)(6) 2025 AT 1800. IN ADDITION, IT WAS REPORTED THAT THERE DID NOT SEEM TO HAVE BEEN AN ALERT AROUND 1800 WHEN THE SENSOR STOPPED SENDING DATA AND THE PATIENT WAS VERY LOW IN BLOOD SUGAR. NO FURTHER EVENT INFORMATION WAS PROVIDED. DATA HAS BEEN REQUESTED BUT IS PENDING EVALUATION. A FOLLOW UP REPORT WILL BE SUBMITTED UPON COMPLETION. NO ADDITIONAL PATIENT OR EVENT INFORMATION IS AVAILABLE. Manufacturer Narrative: (B)(4). DIABETES MELLITUS IS A KNOWN CAUSE OF HYPOGLYCEMIA.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22917447&pc=NPT | 9617601-2025-01219 | 2025-08-06 04:00 | Death | MEDTRONIC MEXICO S. DE R.L. DE CV | 2025-08-28 04:00 | NPT | EVOLUT FX VALVE | Material Invagination; Device Difficult to Setup or Prepare; Patient Device Interaction Problem | Stroke/CVA; Paresis | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE IMPLANT PROCEDURE, THE VALVE WAS MISLOADED BY THE MEDTRONIC REPRESENTATIVE DUE TO FOLDING/OVERLAP. A NEW VALVE WAS LOADED ONTO A DIFFERENT DELIVERY CATHETER SYSTEM (DCS) AND DEPLOYED WITH NO ISSUES. FOLLOWING THE PROCEDURE, THE PATIENT SUFFERED A STROKE IN RECOVERY AND SUBSEQUENTLY DIED OVERNIGHT IN THE INTENSIVE CARE UNIT (ICU). ADDITIONAL INFORMATION WAS RECEIVED THAT THE STROKE WAS CONFIRMED WITH A COMPUTED TOMOGRAPHY (CT) SCAN. AS A RESULT OF THE STROKE, THE PATIENT WAS STRUGGLING TO WAKE UP IN RECOVERY, WAS NOT ABLE TO MOVE THEIR LEFT ARM AND HAD LEFT-SIDED FACIAL DROOPING. UNSPECIFIED TREATMENT WAS PROVIDED IN THE ICU. PER THE PHYSICIAN, THE CAUSE OF THE STROKE WAS UNKNOWN AND DIFFICULT TO DETERMINE, BUT IT IS A KNOWN RISK FOR ANY INTERVENTION AND ANY EQUIPMENT USED COULD HAVE CONTRIBUTED TO THE STROKE. THE PHYSICIAN DID NOT ATTRIBUTE THE STROKE TO THE DCS. THE MISLOAD OR FOLDING WAS NOT DETECTED ON THE SCREENING, BUT WHEN THE VALVE WAS DEPLOYED UP UNTIL 80% MARK, THE IN FOLD WAS OBSERVED EXTENDING ALL THE WAY UP THE VALVE. THE VALVE WAS QUICKLY RECAPTURED AND THE NEW VALVE LOADED. Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME: EVOLUTFX DCS, PRODUCT ID: D-EVOLUTFX-34, LOT: 0012425596, USE BY DATE: 2026-09-03, UDI: (B)(4). CONTINUATION OF D10: PRODUCT ID D-EVOLUTFX-34 (LOT: 0012425596); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE N/A; EXPLANT DATE N/A PRODUCT ID EVOLUTFX-34 (SERIAL: (B)(6); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE (B)(6) 2025; EXPLANT DATE N/A PRODUCT ID L-EVOLUTFX-34 (LOT: 0012210393); PRODUCT TYPE: 0195-HEART VALVES; MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22918901&pc=LPB | 2029046-2025-02913 | 2025-04-06 04:00 | Death | BIOSENSE WEBSTER INC | 2025-08-28 04:00 | LPB | UNK_CELSIUS THERMOCOOL | Adverse Event Without Identified Device or Use Problem | Pulmonary Embolism; Pneumonia; Cardiac Tamponade; Cardiogenic Shock; Heart Failure/Congestive Heart Failure |  |  | 1 | Event Description: THIS COMPLAINT IS FROM A LITERATURE SOURCE AND THE FOLLOWING CITATION WAS REVIEWED: SCIACCA V, FELDT J, ROTTNER L, HEEGER CH, SOHNS C, REISSMANN B, SAGUNER AM, SANTORO F, TILZ RR, MAURER T, RILLIG A, OUYANG F, LINZ D, VERNOOY K, SOMMER P, SULTAN A, WILLEMS S, KUCK KH, METZNER A, FINK T. PERIPROCEDURAL SAFETY OF INTERVENTIONAL ELECTROPHYSIOLOGICAL PROCEDURES IN OCTOGENARIANS AND NONAGENARIANS. J CARDIOVASC ELECTROPHYSIOL. 2025 MAY 26. DOI: 10.1111/JCE.16689. EPUB AHEAD OF PRINT. PMID: 40420477. BACKGROUND AND PURPOSE: CATHETER ABLATION IS AN ESTABLISHED TREATMENT FOR CARDIAC ARRHYTHMIA. THERE IS A LACK OF DATA ON INVASIVE ELECTROPHYSIOLOGICAL (EP) PROCEDURES IN AGED PATIENTS. METHODS: CONSECUTIVE PATIENTS = 80 YEARS WHO UNDERWENT CATHETER ABLATION OR LEFT ATRIAL APPENDAGE CLOSURE PROCEDURES BETWEEN JANUARY 2005 AND DECEMBER 2017 IN A HIGH-VOLUME CENTER WERE RETROSPECTIVELY STUDIED AND COMPARED TO A MATCHED CONTROL GROUP OF INDIVIDUALS < 80 YEARS OF AGE. BWI DEVICES THAT WERE USED IN THIS STUDY: 7 F 3.5 MM TIP ABLATION CATHETERS (CELSIUS OR THERMOCOOL, BIOSENSE WEBSTER INC., (B)(6), USA) NON-BWI DEVICES: ENDOCARDIAL DEVICE (WATCHMAN, BOSTON SCIENTIFIC, (B)(6), USA) NOTE: THE VASCULAR COMPLICATION NOTED IN THE ARTICLE WAS NOT IDENTIFIED WHETHER IT IS BWI DEVICE OR NON-BWI DEVICE. ADVERSE EVENT(S) AND PROVIDED INTERVENTIONS POSSIBLY ASSOCIATED WITH UNK_CELSIUS THERMOCOOL (QTY 51) - 17 CASES HAD BLEEDING AT VASCULAR ACCESS SITE, NO INTERVENTION NOTED. - 13 CASES HAD CARDIAC TAMPONADE, NO INTERVENTION NOTED. - 2 CASES HAD PULMONARY EMOBOLISM, NO INTERVENTION NOTED. - 3 CASES HAD STROKE, NO INTERVENTION NOTED. - 2 CASES HAD RESPIRATORY FAILURE WITH NEED FOR INVASIVE VENTILLATION. - 2 CASES UNDER AGED GROUP HAD AV BLOCK WITH SUBSEQUENT PERMANENT PEACEMAKER IMPLANTATION. - 2 CASES HAD PNEUMOTHORAX WITH DRAINAGE. - 2 CASES HAD PHERENIC NERVE PALSY. - 1 CASE HAD GASTROPARESIS. - 7 CASES OF DEATH (2 CARDIAC TAMPONADE, 2 PULMONARY ARTERY EMBOLISM, 2 DECOMPENSATED LEFT HEART FAILURE DURING VT STORY, 1 PNEUMONIA AFTER ASPIRATION DURING THE ABLATION PROCEDURE). Manufacturer Narrative: D4: UDI: AS THE CATALOG/MODEL NUMBER WAS NOT PROVIDED, THE (01) GTIN IS NOT AVAILABLE. THIS COMPLAINT IS FROM A LITERATURE SOURCE AND THE FOLLOWING CITATION WAS REVIEWED: SCIACCA V, FELDT J, ROTTNER L, HEEGER CH, SOHNS C, REISSMANN B, SAGUNER AM, SANTORO F, TILZ RR, MAURER T, RILLIG A, OUYANG F, LINZ D, VERNOOY K, SOMMER P, SULTAN A, WILLEMS S, KUCK KH, METZNER A, FINK T. PERIPROCEDURAL SAFETY OF INTERVENTIONAL ELECTROPHYSIOLOGICAL PROCEDURES IN OCTOGENARIANS AND NONAGENARIANS. J CARDIOVASC ELECTROPHYSIOL. 2025 MAY 26. DOI: 10.1111/JCE.16689. EPUB AHEAD OF PRINT. PMID: 40420477. NO DEVICE WAS RECEIVED FOR ANALYSIS AT THE TIME OF SUBMISSION OF THE INITIAL 3500A. SINCE THE PRODUCT WAS NOT RETURNED FOR ANALYSIS, NO PRODUCT FAILURE ANALYSIS CAN BE CONDUCTED, AND NO DETERMINATION OF POSSIBLE CONTRIBUTING FACTORS COULD BE MADE. DEVICE HISTORY RECORD (DHR) REVIEW CANNOT BE CONDUCTED BECAUSE THE LOT NUMBER WAS PROVIDED BY THE CUSTOMER. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 4. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY BIOSENSE WEBSTER INC, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, BIOSENSE WEBSTER INC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. MANUFACTURER¿S REFERENCE NUMBER: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22918938&pc=OZD | 1220648-2025-45989 | 2025-07-04 04:00 | Death | ABIOMED, INC. - 1220648 | 2025-08-28 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Thrombosis/Thrombus | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT IMPLANTED WITH AN IMPELLA CP DEVELOPED A THROMBUS IN THE LEFT VENTRICLE. THE IMPELLA WAS EXPLANTED. LATER, CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22918939&pc=OZD | 1220648-2025-45981 | 2025-07-30 04:00 | Death | ABIOMED, INC. - 1220648 | 2025-08-28 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT IMPLANTED WITH AN IMPELLA CP EXPERIENCED LIMB ISCHEMIA. A FEM FEM BYPASS WAS DONE HOWEVER, THE PATIENT CODED AND EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22919289&pc=CAW | 3013095415-2025-00772 | 2025-07-13 04:00 | Death | INVACARE FLORIDA | 2025-08-28 04:00 | CAW | INVACARE® PERFECTO2¿ OXYGEN CONCENTRATOR | Insufficient Information | Insufficient Information | K200890 |  | 1 | Event Description: THE DEVICE¿S PREVIOUS MANUFACTURER, INVACARE, CONTACTED VENTEC VIA EMAIL TO REPORT THAT THEIR LEGAL DEPARTMENT HAD RECEIVED A LETTER FROM A LAW FIRM WHICH STATED THE FOLLOWING: ¿TO WHOM IT MAY CONCERN, THIS LAW FIRM HAS BEEN RETAINED BY (B)(6) TO REPRESENT ITS SUBROGATION INTERESTS ARISING FROM A FIRE ON (B)(6) 2025, AT (B)(6) (THE ¿PROPERTY¿). (B)(6) PROVIDES PROPERTY INSURANCE TO (B)(6). (B)(6) FOR ITS INTERESTS AT THE PROPERTY. OUR INVESTIGATION INTO THE FIRE IS UNDERWAY. WITHIN THE ROOM OF THE FIRE ORIGIN, WE LOCATED THE FOLLOWING PRODUCTS MANUFACTURED AND/OR DISTRIBUTED BY YOUR COMPANY: - INVACARE O2 HOMEFILL II COMPRESSOR (MODEL: IOH200). - INVACARE PERFECTO 2 V OXYGEN CONCENTRATOR (MODEL: IRC5PO2V / SERIAL: (B)(6)).¿ [THE SERIAL NUMBER FOR THE INVACARE O2 HOMEFILL II COMPRESSOR IS CURRENTLY UNKNOWN]. THE INVACARE LEGAL DEPARTMENT THEN COMPLETED SOME RESEARCH REGARDING THE "(B)(6)" FIRE. THROUGH THIS RESEARCH, INVACARE BECAME AWARE THAT 10 PEOPLE HAD DIED AS A RESULT OF THE FIRE. THIS MEDWATCH REPORT IS TO DOCUMENT THE DEATH OF PATIENT ¿(B)(6)¿ AS IT PERTAINS TO THE INVACARE® PERFECTO2¿ OXYGEN CONCENTRATOR. NO FURTHER DETAILS ABOUT THE PATIENT OR THEIR CAUSE OF DEATH WAS PROVIDED. Manufacturer Narrative: H6: VENTEC HAS DETERMINED THAT THIS DEVICE WAS PLACED INTO COMMERCIAL DISTRIBUTION BACK IN FEBRUARY OF 2015, WHICH PREDATES THE REQUIRED UDI COMPLIANCE DATE FOR CLASS II MEDICAL DEVICES. AS A RESULT, NO UDI# EXISTS FOR THIS DEVICE AND SECTION D4, UDI#, SHALL STATE ¿NONE.¿ THE DEVICE HAS NOT BEEN RETURNED TO INVACARE OR VENTEC (REACT HEALTH) FOR AN EVALUATION. INVACARE CONTINUES TO INVESTIGATE THE REPORTED ISSUE. ONCE COMPLETE, VENTEC WILL BE PROVIDED WITH THEIR INVESTIGATION FINDINGS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE AS DEFINED BY 21 CFR 803.56.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22919349&pc=CAW | 3013095415-2025-00773 | 2025-07-13 04:00 | Death | INVACARE FLORIDA | 2025-08-28 04:00 | CAW | INVACARE® PERFECTO2¿ OXYGEN CONCENTRATOR | Insufficient Information | Insufficient Information | K200890 |  | 1 | Event Description: THE DEVICE¿S PREVIOUS MANUFACTURER, INVACARE, CONTACTED VENTEC VIA EMAIL TO REPORT THAT THEIR LEGAL DEPARTMENT HAD RECEIVED A LETTER FROM A LAW FIRM WHICH STATED THE FOLLOWING: ¿TO WHOM IT MAY CONCERN, THIS LAW FIRM HAS BEEN RETAINED BY (B)(6) TO REPRESENT ITS SUBROGATION INTERESTS ARISING FROM A FIRE ON (B)(6) 2025, AT (B)(6) (THE ¿PROPERTY¿). (B)(6) PROVIDES PROPERTY INSURANCE TO (B)(6) FOR ITS INTERESTS AT THE PROPERTY. OUR INVESTIGATION INTO THE FIRE IS UNDERWAY. WITHIN THE ROOM OF THE FIRE ORIGIN, WE LOCATED THE FOLLOWING PRODUCTS MANUFACTURED AND/OR DISTRIBUTED BY YOUR COMPANY: INVACARE O2 HOMEFILL II COMPRESSOR (MODEL: IOH200). INVACARE PERFECTO 2 V OXYGEN CONCENTRATOR (MODEL: IRC5PO2V / SERIAL: (B)(6). [THE SERIAL NUMBER FOR THE INVACARE O2 HOMEFILL II COMPRESSOR IS CURRENTLY UNKNOWN]. THE INVACARE LEGAL DEPARTMENT THEN COMPLETED SOME RESEARCH REGARDING THE ¿(B)(6)¿ FIRE. THROUGH THIS RESEARCH, INVACARE BECAME AWARE THAT 10 PEOPLE HAD DIED AS A RESULT OF THE FIRE. THIS MEDWATCH REPORT IS TO DOCUMENT THE DEATH OF PATIENT ¿(B)(6)¿ AS IT PERTAINS TO THE INVACARE® PERFECTO2¿ OXYGEN CONCENTRATOR. NO FURTHER DETAILS ABOUT THE PATIENT OR THEIR CAUSE OF DEATH WAS PROVIDED. Manufacturer Narrative: H6: VENTEC HAS DETERMINED THAT THIS DEVICE WAS PLACED INTO COMMERCIAL DISTRIBUTION BACK IN FEBRUARY OF 2015, WHICH PREDATES THE REQUIRED UDI COMPLIANCE DATE FOR CLASS II MEDICAL DEVICES. AS A RESULT, NO UDI# EXISTS FOR THIS DEVICE AND SECTION D4, UDI#, SHALL STATE ¿NONE.¿ THE DEVICE HAS NOT BEEN RETURNED TO INVACARE OR VENTEC (REACT HEALTH) FOR AN EVALUATION. INVACARE CONTINUES TO INVESTIGATE THE REPORTED ISSUE. ONCE COMPLETE, VENTEC WILL BE PROVIDED WITH THEIR INVESTIGATION FINDINGS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE AS DEFINED BY 21 CFR 803.56.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22919363&pc=NAY | 2955842-2025-36085 | 2025-07-30 04:00 | Death | INTUITIVE SURGICAL, INC | 2025-08-28 04:00 | NAY | DAVINCI XI | Insufficient Information | Hemorrhage/Bleeding; Great Vessel Perforation | K131861 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED PULMONARY LOBECTOMY, THE PATIENT EXPERIENCED BLEEDING AND DIED. THE SITE THORACIC LEAD REPORTED THAT IT WAS A LONG LOBECTOMY SURGERY, AND THE BLEEDING OCCURRED TOWARDS THE END OF THE PROCEDURE. THE SURGEON WAS WORKING NEAR THE PULMONARY ARTERY (PA) WHEN THIS EVENT OCCURRED. THE THORACIC LEAD THOUGHT THE SURGEON WAS USING A DA VINCI STAPLER ON THE BLOOD VESSELS NEAR THE PA, OR MAYBE ON THE PA DIRECTLY WHEN THE BLEEDING BEGAN; HOWEVER, IT WAS UNKNOWN IF THE BLEEDING RESULTED FROM A DA VINCI STAPLER AND STAPLE LINE, BUT KNOWS THE SURGEON WAS WORKING ON VESSELS NEAR THE PA AND USED A DA VINCI STAPLER DURING THE PROCEDURE. THE OR STAFF REPORTED THAT THE PATIENT'S TISSUE WAS IN POOR CONDITION WHICH CONTRIBUTED TO THE BLEEDING COMPLICATION. ATTEMPTS TO CONTACT THE SURGEON FOR ADDITIONAL INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: REVIEW OF THE SYSTEM LOGS FOUND NO RELEVANT ERRORS OCCURRED DURING THE PROCEDURE. LOGS FOR THE FOUR SUBSEQUENT PROCEDURES FOUND THE SYSTEM FUNCTIONED NORMALLY, WITH NO INTRAOPERATIVE EVENTS OR ISSUES. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: 30 DEGREE ENDOSCOPE PLUS, TIP-UP FENESTRATED GRASPER, FENESTRATED BIPOLAR FORCEPS, LONG BIPOLAR GRASPER, TWO SUREFORM 45 STAPLER INSTRUMENTS. A SITE HISTORY REVIEW FOUND NO COMPLAINTS WERE MADE FOR THE INSTRUMENTS USED DURING THIS PROCEDURE. REVIEW OF THE STAPLER LOGS FOUND THAT THE FIRST SUREFORM 45 CURVED TIP STAPLER (PART NUMBER 480545-06, LOT NUMBER K10250303-0053) WAS INSTALLED 16 TIMES AND FIRED 12 RELOADS (1 BLUE, 4 GREEN 1 BLUE, 4 GREEN, 2 WHITE RESPECTIVELY). ALL FIRINGS WERE COMPLETED WITH NO PAUSES FOR COMPRESSION. THERE WERE NO INCOMPLETE CLAMPS OR UNCLAMP FAILURES BY THIS INSTRUMENT. THERE WERE FOUR CONSECUTIVE INITIALIZATION FAILURES DUE TO HIGH DRIVETRAIN FRICTION AFTER THE 3RD FIRING, HOWEVER, THE INITIALIZATION FAILURES RESOLVED AND THE INSTRUMENT SUCCESSFULLY COMPLETED 9 SUBSEQUENT FIRINGS. THE SECOND STAPLER USED WAS CURVED-TIP SUREFORM 45 (PART NUMBER 480545-06, LOT NUMBER K10250303-0050). THIS INSTRUMENT WAS INSTALLED 2 TIMES AND FIRED TWO GREEN RELOADS. BOTH FIRINGS WERE COMPLETED WITH NO PAUSES FOR COMPRESSION. THERE WERE NO INCOMPLETE CLAMPS OR UNCLAMP FAILURES. THE SYSTEM LOGS DID NOT SHOW ANY STAPLER RELATED ERRORS. A DEVICE HISTORY RECORD (DHR) REVIEW FOUND NO NON-CONFORMANCES RELATED TO THE DEVICES USED DURING THIS PROCEDURE. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT DIED FOLLOWING BLEEDING IN A PULMONARY LOBECTOMY PROCEDURE. THE EXACT CAUSE OF THE BLEEDING AND THE EVENTS THAT LED TO THE BLEEDING ARE NOT REPORTED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT. SECTION A2: PATIENT AGE WAS REPORTED TO BE "AROUND 60 YEARS OLD". .
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22919368&pc=CAW | 3013095415-2025-00774 | 2025-07-13 04:00 | Death | INVACARE FLORIDA | 2025-08-28 04:00 | CAW | INVACARE® PERFECTO2¿ OXYGEN CONCENTRATOR | Insufficient Information | Insufficient Information | K200890 |  | 1 | Event Description: THE DEVICE¿S PREVIOUS MANUFACTURER, INVACARE, CONTACTED VENTEC VIA EMAIL TO REPORT THAT THEIR LEGAL DEPARTMENT HAD RECEIVED A LETTER FROM A LAW FIRM WHICH STATED THE FOLLOWING: ¿TO WHOM IT MAY CONCERN, THIS LAW FIRM HAS BEEN RETAINED BY (B)(6) COMPANY (B)(6) TO REPRESENT ITS SUBROGATION INTERESTS ARISING FROM A FIRE ON (B)(6) 2025, (B)(6) (THE ¿PROPERTY¿). (B)(6) PROVIDES PROPERTY INSURANCE TO (B)(6) FOR ITS INTERESTS AT THE PROPERTY. OUR INVESTIGATION INTO THE FIRE IS UNDERWAY. WITHIN THE ROOM OF THE FIRE ORIGIN, WE LOCATED THE FOLLOWING PRODUCTS MANUFACTURED AND/OR DISTRIBUTED BY YOUR COMPANY: INVACARE O2 HOMEFILL II COMPRESSOR (MODEL: IOH200) INVACARE PERFECTO 2 V OXYGEN CONCENTRATOR (MODEL: IRC5PO2V / SERIAL: (B)(6). [THE SERIAL NUMBER FOR THE INVACARE O2 HOMEFILL II COMPRESSOR IS CURRENTLY UNKNOWN] THE INVACARE LEGAL DEPARTMENT THEN COMPLETED SOME RESEARCH REGARDING THE (B)(6) FIRE. THROUGH THIS RESEARCH, INVACARE BECAME AWARE THAT 10 PEOPLE HAD DIED AS A RESULT OF THE FIRE. THIS MEDWATCH REPORT IS TO DOCUMENT THE DEATH OF PATIENT (B)(6) AS IT PERTAINS TO THE INVACARE® PERFECTO2¿ OXYGEN CONCENTRATOR. NO FURTHER DETAILS ABOUT THE PATIENT OR THEIR CAUSE OF DEATH WAS PROVIDED. Manufacturer Narrative: H6: VENTEC HAS DETERMINED THAT THIS DEVICE WAS PLACED INTO COMMERCIAL DISTRIBUTION BACK IN FEBRUARY OF 2015, WHICH PREDATES THE REQUIRED UDI COMPLIANCE DATE FOR CLASS II MEDICAL DEVICES. AS A RESULT, NO UDI# EXISTS FOR THIS DEVICE AND SECTION D4, UDI#, SHALL STATE ¿NONE.¿ THE DEVICE HAS NOT BEEN RETURNED TO INVACARE OR VENTEC (REACT HEALTH) FOR AN EVALUATION. INVACARE CONTINUES TO INVESTIGATE THE REPORTED ISSUE. ONCE COMPLETE, VENTEC WILL BE PROVIDED WITH THEIR INVESTIGATION FINDINGS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE AS DEFINED BY 21 CFR 803.56.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22919389&pc=CAW | 3013095415-2025-00775 | 2025-07-13 04:00 | Death | INVACARE FLORIDA | 2025-08-28 04:00 | CAW | INVACARE® PERFECTO2¿ OXYGEN CONCENTRATOR | Insufficient Information | Insufficient Information | K200890 |  | 1 | Event Description: THE DEVICE¿S PREVIOUS MANUFACTURER, INVACARE, CONTACTED VENTEC VIA EMAIL TO REPORT THAT THEIR LEGAL DEPARTMENT HAD RECEIVED A LETTER FROM A LAW FIRM WHICH STATED THE FOLLOWING: ¿TO WHOM IT MAY CONCERN, THIS LAW FIRM HAS BEEN RETAINED BY (B)(6) TO REPRESENT ITS SUBROGATION INTERESTS ARISING FROM A FIRE ON (B)(6) 2025, AT (B)(6) (THE ¿PROPERTY¿). (B)(6) PROVIDES PROPERTY INSURANCE TO (B)(6). (¿(B)(6)¿) FOR ITS INTERESTS AT THE PROPERTY. OUR INVESTIGATION INTO THE FIRE IS UNDERWAY. WITHIN THE ROOM OF THE FIRE ORIGIN, WE LOCATED THE FOLLOWING PRODUCTS MANUFACTURED AND/OR DISTRIBUTED BY YOUR COMPANY: - INVACARE O2 HOMEFILL II COMPRESSOR (MODEL: IOH200). - INVACARE PERFECTO 2 V OXYGEN CONCENTRATOR (MODEL: IRC5PO2V / SERIAL: (B)(6)).¿ [THE SERIAL NUMBER FOR THE INVACARE O2 HOMEFILL II COMPRESSOR IS CURRENTLY UNKNOWN]. THE INVACARE LEGAL DEPARTMENT THEN COMPLETED SOME RESEARCH REGARDING THE ¿(B)(6)¿ FIRE. THROUGH THIS RESEARCH, INVACARE BECAME AWARE THAT 10 PEOPLE HAD DIED AS A RESULT OF THE FIRE. THIS MEDWATCH REPORT IS TO DOCUMENT THE DEATH OF PATIENT ¿(B)(6)¿ AS IT PERTAINS TO THE INVACARE® PERFECTO2¿ OXYGEN CONCENTRATOR. NO FURTHER DETAILS ABOUT THE PATIENT OR THEIR CAUSE OF DEATH WAS PROVIDED. Manufacturer Narrative: H6: VENTEC HAS DETERMINED THAT THIS DEVICE WAS PLACED INTO COMMERCIAL DISTRIBUTION BACK IN FEBRUARY OF 2015, WHICH PREDATES THE REQUIRED UDI COMPLIANCE DATE FOR CLASS II MEDICAL DEVICES. AS A RESULT, NO UDI# EXISTS FOR THIS DEVICE AND SECTION D4, UDI#, SHALL STATE ¿NONE.¿ THE DEVICE HAS NOT BEEN RETURNED TO INVACARE OR VENTEC (REACT HEALTH) FOR AN EVALUATION. INVACARE CONTINUES TO INVESTIGATE THE REPORTED ISSUE. ONCE COMPLETE, VENTEC WILL BE PROVIDED WITH THEIR INVESTIGATION FINDINGS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE AS DEFINED BY 21 CFR 803.56.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22919471&pc=CAW | 3013095415-2025-00776 | 2025-07-13 04:00 | Death | INVACARE FLORIDA | 2025-08-28 04:00 | CAW | INVACARE® PERFECTO2¿ OXYGEN CONCENTRATOR | Insufficient Information | Insufficient Information | K200890 |  | 1 | Event Description: THE DEVICE¿S PREVIOUS MANUFACTURER, INVACARE, CONTACTED VENTEC VIA EMAIL TO REPORT THAT THEIR LEGAL DEPARTMENT HAD RECEIVED A LETTER FROM A LAW FIRM WHICH STATED THE FOLLOWING: ¿TO WHOM IT MAY CONCERN, THIS LAW FIRM HAS BEEN RETAINED BY (B)(6) TO REPRESENT ITS SUBROGATION INTERESTS ARISING FROM A FIRE ON (B)(6) 2025, AT (B)(6) (THE ¿PROPERTY¿). (B)(6) PROVIDES PROPERTY INSURANCE TO (B)(6). (¿(B)(6)¿) FOR ITS INTERESTS AT THE PROPERTY. OUR INVESTIGATION INTO THE FIRE IS UNDERWAY. WITHIN THE ROOM OF THE FIRE ORIGIN, WE LOCATED THE FOLLOWING PRODUCTS MANUFACTURED AND/OR DISTRIBUTED BY YOUR COMPANY: - INVACARE O2 HOMEFILL II COMPRESSOR (MODEL: IOH200). - INVACARE PERFECTO 2 V OXYGEN CONCENTRATOR (MODEL: IRC5PO2V / SERIAL: (B)(6)).¿ [THE SERIAL NUMBER FOR THE INVACARE O2 HOMEFILL II COMPRESSOR IS CURRENTLY UNKNOWN]. THE INVACARE LEGAL DEPARTMENT THEN COMPLETED SOME RESEARCH REGARDING THE ¿(B)(6)¿ FIRE. THROUGH THIS RESEARCH, INVACARE BECAME AWARE THAT 10 PEOPLE HAD DIED AS A RESULT OF THE FIRE. THIS MEDWATCH REPORT IS TO DOCUMENT THE DEATH OF PATIENT ¿(B)(6)¿ AS IT PERTAINS TO THE INVACARE® PERFECTO2¿ OXYGEN CONCENTRATOR. NO FURTHER DETAILS ABOUT THE PATIENT OR THEIR CAUSE OF DEATH WAS PROVIDED. Manufacturer Narrative: H6: VENTEC HAS DETERMINED THAT THIS DEVICE WAS PLACED INTO COMMERCIAL DISTRIBUTION BACK IN FEBRUARY OF 2015, WHICH PREDATES THE REQUIRED UDI COMPLIANCE DATE FOR CLASS II MEDICAL DEVICES. AS A RESULT, NO UDI# EXISTS FOR THIS DEVICE AND SECTION D4, UDI#, SHALL STATE ¿NONE.¿ THE DEVICE HAS NOT BEEN RETURNED TO INVACARE OR VENTEC (REACT HEALTH) FOR AN EVALUATION. INVACARE CONTINUES TO INVESTIGATE THE REPORTED ISSUE. ONCE COMPLETE, VENTEC WILL BE PROVIDED WITH THEIR INVESTIGATION FINDINGS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE AS DEFINED BY 21 CFR 803.56.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22919524&pc=CAW | 3013095415-2025-00777 | 2025-07-13 04:00 | Death | INVACARE FLORIDA | 2025-08-28 04:00 | CAW | INVACARE® PERFECTO2¿ OXYGEN CONCENTRATOR | Insufficient Information | Insufficient Information | K200890 |  | 1 | Event Description: THE DEVICE¿S PREVIOUS MANUFACTURER, INVACARE, CONTACTED VENTEC VIA EMAIL TO REPORT THAT THEIR LEGAL DEPARTMENT HAD RECEIVED A LETTER FROM A LAW FIRM WHICH STATED THE FOLLOWING: ¿TO WHOM IT MAY CONCERN, THIS LAW FIRM HAS BEEN RETAINED BY (B)(6) COMPANY (B)(6) TO REPRESENT ITS SUBROGATION INTERESTS ARISING FROM A FIRE ON (B)(6) 2025, (B)(6) (THE ¿PROPERTY¿). (B)(6) PROVIDES PROPERTY INSURANCE TO (B)(6). FOR ITS INTERESTS AT THE PROPERTY. OUR INVESTIGATION INTO THE FIRE IS UNDERWAY. WITHIN THE ROOM OF THE FIRE ORIGIN, WE LOCATED THE FOLLOWING PRODUCTS MANUFACTURED AND/OR DISTRIBUTED BY YOUR COMPANY: INVACARE O2 HOMEFILL II COMPRESSOR (MODEL: IOH200) INVACARE PERFECTO 2 V OXYGEN CONCENTRATOR (MODEL: IRC5PO2V / SERIAL: (B)(6).¿ [THE SERIAL NUMBER FOR THE INVACARE O2 HOMEFILL II COMPRESSOR IS CURRENTLY UNKNOWN] THE INVACARE LEGAL DEPARTMENT THEN COMPLETED SOME RESEARCH REGARDING THE ¿(B)(6)¿ FIRE. THROUGH THIS RESEARCH, INVACARE BECAME AWARE THAT 10 PEOPLE HAD DIED AS A RESULT OF THE FIRE. THIS MEDWATCH REPORT IS TO DOCUMENT THE DEATH OF PATIENT (B)(6) AS IT PERTAINS TO THE INVACARE® PERFECTO2¿ OXYGEN CONCENTRATOR. NO FURTHER DETAILS ABOUT THE PATIENT OR THEIR CAUSE OF DEATH WAS PROVIDED. Manufacturer Narrative: H6: VENTEC HAS DETERMINED THAT THIS DEVICE WAS PLACED INTO COMMERCIAL DISTRIBUTION BACK IN FEBRUARY OF 2015, WHICH PREDATES THE REQUIRED UDI COMPLIANCE DATE FOR CLASS II MEDICAL DEVICES. AS A RESULT, NO UDI# EXISTS FOR THIS DEVICE AND SECTION D4, UDI#, SHALL STATE ¿NONE.¿ THE DEVICE HAS NOT BEEN RETURNED TO INVACARE OR VENTEC (REACT HEALTH) FOR AN EVALUATION. INVACARE CONTINUES TO INVESTIGATE THE REPORTED ISSUE. ONCE COMPLETE, VENTEC WILL BE PROVIDED WITH THEIR INVESTIGATION FINDINGS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE AS DEFINED BY 21 CFR 803.56.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22919578&pc=CAW | 3013095415-2025-00778 | 2025-07-13 04:00 | Death | INVACARE FLORIDA | 2025-08-28 04:00 | CAW | INVACARE® PERFECTO2¿ OXYGEN CONCENTRATOR | Insufficient Information | Insufficient Information | K200890 |  | 1 | Event Description: THE DEVICE¿S PREVIOUS MANUFACTURER, INVACARE, CONTACTED VENTEC VIA EMAIL TO REPORT THAT THEIR LEGAL DEPARTMENT HAD RECEIVED A LETTER FROM A LAW FIRM WHICH STATED THE FOLLOWING: ¿TO WHOM IT MAY CONCERN, THIS LAW FIRM HAS BEEN RETAINED BY (B)(6) TO REPRESENT ITS SUBROGATION INTERESTS ARISING FROM A FIRE ON (B)(6) 2025, AT (B)(6) (THE ¿PROPERTY¿). (B)(6) PROVIDES PROPERTY INSURANCE TO (B)(6). (¿(B)(6)¿) FOR ITS INTERESTS AT THE PROPERTY. OUR INVESTIGATION INTO THE FIRE IS UNDERWAY. WITHIN THE ROOM OF THE FIRE ORIGIN, WE LOCATED THE FOLLOWING PRODUCTS MANUFACTURED AND/OR DISTRIBUTED BY YOUR COMPANY: - INVACARE O2 HOMEFILL II COMPRESSOR (MODEL: IOH200). - INVACARE PERFECTO 2 V OXYGEN CONCENTRATOR (MODEL: IRC5PO2V / SERIAL: (B)(6)).¿ [THE SERIAL NUMBER FOR THE INVACARE O2 HOMEFILL II COMPRESSOR IS CURRENTLY UNKNOWN]. THE INVACARE LEGAL DEPARTMENT THEN COMPLETED SOME RESEARCH REGARDING THE ¿(B)(6)¿ FIRE. THROUGH THIS RESEARCH, INVACARE BECAME AWARE THAT 10 PEOPLE HAD DIED AS A RESULT OF THE FIRE. THIS MEDWATCH REPORT IS TO DOCUMENT THE DEATH OF PATIENT ¿(B)(6)¿ AS IT PERTAINS TO THE INVACARE® PERFECTO2¿ OXYGEN CONCENTRATOR. NO FURTHER DETAILS ABOUT THE PATIENT OR THEIR CAUSE OF DEATH WAS PROVIDED. Manufacturer Narrative: H6: VENTEC HAS DETERMINED THAT THIS DEVICE WAS PLACED INTO COMMERCIAL DISTRIBUTION BACK IN FEBRUARY OF 2015, WHICH PREDATES THE REQUIRED UDI COMPLIANCE DATE FOR CLASS II MEDICAL DEVICES. AS A RESULT, NO UDI# EXISTS FOR THIS DEVICE AND SECTION D4, UDI#, SHALL STATE ¿NONE.¿ THE DEVICE HAS NOT BEEN RETURNED TO INVACARE OR VENTEC (REACT HEALTH) FOR AN EVALUATION. INVACARE CONTINUES TO INVESTIGATE THE REPORTED ISSUE. ONCE COMPLETE, VENTEC WILL BE PROVIDED WITH THEIR INVESTIGATION FINDINGS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE AS DEFINED BY 21 CFR 803.56.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22919585&pc=CAW | 3013095415-2025-00779 | 2025-07-13 04:00 | Death | INVACARE FLORIDA | 2025-08-28 04:00 | CAW | INVACARE® PERFECTO2¿ OXYGEN CONCENTRATOR | Insufficient Information | Insufficient Information | K200890 |  | 1 | Event Description: THE DEVICE¿S PREVIOUS MANUFACTURER, INVACARE, CONTACTED VENTEC VIA EMAIL TO REPORT THAT THEIR LEGAL DEPARTMENT HAD RECEIVED A LETTER FROM A LAW FIRM WHICH STATED THE FOLLOWING: ¿TO WHOM IT MAY CONCERN, THIS LAW FIRM HAS BEEN RETAINED BY (B)(6) TO REPRESENT ITS SUBROGATION INTERESTS ARISING FROM A FIRE ON (B)(6) 2025, AT (B)(6) (THE ¿PROPERTY¿). (B)(6) PROVIDES PROPERTY INSURANCE TO (B)(6) FOR ITS INTERESTS AT THE PROPERTY. OUR INVESTIGATION INTO THE FIRE IS UNDERWAY. WITHIN THE ROOM OF THE FIRE ORIGIN, WE LOCATED THE FOLLOWING PRODUCTS MANUFACTURED AND/OR DISTRIBUTED BY YOUR COMPANY: INVACARE O2 HOMEFILL II COMPRESSOR (MODEL: IOH200). INVACARE PERFECTO 2 V OXYGEN CONCENTRATOR (MODEL: IRC5PO2V / SERIAL: (B)(6). [THE SERIAL NUMBER FOR THE INVACARE O2 HOMEFILL II COMPRESSOR IS CURRENTLY UNKNOWN]. THE INVACARE LEGAL DEPARTMENT THEN COMPLETED SOME RESEARCH REGARDING THE ¿(B)(6)¿ FIRE. THROUGH THIS RESEARCH, INVACARE BECAME AWARE THAT 10 PEOPLE HAD DIED AS A RESULT OF THE FIRE. THIS MEDWATCH REPORT IS TO DOCUMENT THE DEATH OF PATIENT ¿(B)(6)¿ AS IT PERTAINS TO THE INVACARE® PERFECTO2¿ OXYGEN CONCENTRATOR. NO FURTHER DETAILS ABOUT THE PATIENT OR THEIR CAUSE OF DEATH WAS PROVIDED. Manufacturer Narrative: H6: VENTEC HAS DETERMINED THAT THIS DEVICE WAS PLACED INTO COMMERCIAL DISTRIBUTION BACK IN FEBRUARY OF 2015, WHICH PREDATES THE REQUIRED UDI COMPLIANCE DATE FOR CLASS II MEDICAL DEVICES. AS A RESULT, NO UDI# EXISTS FOR THIS DEVICE AND SECTION D4, UDI#, SHALL STATE ¿NONE.¿ THE DEVICE HAS NOT BEEN RETURNED TO INVACARE OR VENTEC (REACT HEALTH) FOR AN EVALUATION. INVACARE CONTINUES TO INVESTIGATE THE REPORTED ISSUE. ONCE COMPLETE, VENTEC WILL BE PROVIDED WITH THEIR INVESTIGATION FINDINGS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE AS DEFINED BY 21 CFR 803.56.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22919662&pc=CAW | 3013095415-2025-00780 | 2025-07-13 04:00 | Death | INVACARE FLORIDA | 2025-08-28 04:00 | CAW | INVACARE® PERFECTO2¿ OXYGEN CONCENTRATOR | Insufficient Information | Insufficient Information | K200890 |  | 1 | Event Description: THE DEVICE¿S PREVIOUS MANUFACTURER, INVACARE, CONTACTED VENTEC VIA EMAIL TO REPORT THAT THEIR LEGAL DEPARTMENT HAD RECEIVED A LETTER FROM A LAW FIRM WHICH STATED THE FOLLOWING: ¿TO WHOM IT MAY CONCERN, THIS LAW FIRM HAS BEEN RETAINED BY (B)(6) COMPANY (B)(6)¿TO REPRESENT ITS SUBROGATION INTERESTS ARISING FROM A FIRE ON (B)(6) 2025, AT (B)(6) (THE ¿PROPERTY¿). (B)(6) PROVIDES PROPERTY INSURANCE TO (B)(6) FOR ITS INTERESTS AT THE PROPERTY. OUR INVESTIGATION INTO THE FIRE IS UNDERWAY. WITHIN THE ROOM OF THE FIRE ORIGIN, WE LOCATED THE FOLLOWING PRODUCTS MANUFACTURED AND/OR DISTRIBUTED BY YOUR COMPANY: INVACARE O2 HOMEFILL II COMPRESSOR (MODEL: IOH200) INVACARE PERFECTO 2 V OXYGEN CONCENTRATOR (MODEL: IRC5PO2V / SERIAL: (B)(6).¿ [THE SERIAL NUMBER FOR THE INVACARE O2 HOMEFILL II COMPRESSOR IS CURRENTLY UNKNOWN] THE INVACARE LEGAL DEPARTMENT THEN COMPLETED SOME RESEARCH REGARDING THE ¿(B)(6)¿ FIRE. THROUGH THIS RESEARCH, INVACARE BECAME AWARE THAT 10 PEOPLE HAD DIED AS A RESULT OF THE FIRE. THIS MEDWATCH REPORT IS TO DOCUMENT THE DEATH OF PATIENT ¿(B)(6)¿ AS IT PERTAINS TO THE INVACARE® PERFECTO2¿ OXYGEN CONCENTRATOR. NO FURTHER DETAILS ABOUT THE PATIENT OR THEIR CAUSE OF DEATH WAS PROVIDED. Manufacturer Narrative: H6: VENTEC HAS DETERMINED THAT THIS DEVICE WAS PLACED INTO COMMERCIAL DISTRIBUTION BACK IN FEBRUARY OF 2015, WHICH PREDATES THE REQUIRED UDI COMPLIANCE DATE FOR CLASS II MEDICAL DEVICES. AS A RESULT, NO UDI# EXISTS FOR THIS DEVICE AND SECTION D4, UDI#, SHALL STATE ¿NONE.¿ THE DEVICE HAS NOT BEEN RETURNED TO INVACARE OR VENTEC (REACT HEALTH) FOR AN EVALUATION. INVACARE CONTINUES TO INVESTIGATE THE REPORTED ISSUE. ONCE COMPLETE, VENTEC WILL BE PROVIDED WITH THEIR INVESTIGATION FINDINGS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE AS DEFINED BY 21 CFR 803.56.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22919759&pc=CAW | 3013095415-2025-00781 | 2025-07-13 04:00 | Death | INVACARE FLORIDA | 2025-08-28 04:00 | CAW | INVACARE® PERFECTO2¿ OXYGEN CONCENTRATOR | Insufficient Information | Insufficient Information | K200890 |  | 1 | Event Description: THE DEVICE¿S PREVIOUS MANUFACTURER, INVACARE, CONTACTED VENTEC VIA EMAIL TO REPORT THAT THEIR LEGAL DEPARTMENT HAD RECEIVED A LETTER FROM A LAW FIRM WHICH STATED THE FOLLOWING: ¿TO WHOM IT MAY CONCERN, THIS LAW FIRM HAS BEEN RETAINED BY (B)(6) TO REPRESENT ITS SUBROGATION INTERESTS ARISING FROM A FIRE ON (B)(6) 2025, AT (B)(6) (THE ¿PROPERTY¿). (B)(6) PROVIDES PROPERTY INSURANCE TO (B)(6). (¿(B)(6)¿) FOR ITS INTERESTS AT THE PROPERTY. OUR INVESTIGATION INTO THE FIRE IS UNDERWAY. WITHIN THE ROOM OF THE FIRE ORIGIN, WE LOCATED THE FOLLOWING PRODUCTS MANUFACTURED AND/OR DISTRIBUTED BY YOUR COMPANY: - INVACARE O2 HOMEFILL II COMPRESSOR (MODEL: IOH200). - INVACARE PERFECTO 2 V OXYGEN CONCENTRATOR (MODEL: IRC5PO2V / SERIAL: (B)(6))¿ [THE SERIAL NUMBER FOR THE INVACARE O2 HOMEFILL II COMPRESSOR IS CURRENTLY UNKNOWN]. THE INVACARE LEGAL DEPARTMENT THEN COMPLETED SOME RESEARCH REGARDING THE ¿(B)(6)¿ FIRE. THROUGH THIS RESEARCH, INVACARE BECAME AWARE THAT 10 PEOPLE HAD DIED AS A RESULT OF THE FIRE. THIS MEDWATCH REPORT IS TO DOCUMENT THE DEATH OF PATIENT ¿(B)(6)¿ AS IT PERTAINS TO THE INVACARE® PERFECTO2¿ OXYGEN CONCENTRATOR. (B)(6) WAS INITIALLY TRANSPORTED TO A HOSPITAL, WHERE SHE LATER PASSED AWAY FROM THE INJURIES SHE HAD SUSTAINED IN THE FIRE. NO FURTHER DETAILS ABOUT THE PATIENT OR HER CAUSE OF DEATH WAS PROVIDED. Manufacturer Narrative: H6: VENTEC HAS DETERMINED THAT THIS DEVICE WAS PLACED INTO COMMERCIAL DISTRIBUTION BACK IN FEBRUARY OF 2015, WHICH PREDATES THE REQUIRED UDI COMPLIANCE DATE FOR CLASS II MEDICAL DEVICES. AS A RESULT, NO UDI# EXISTS FOR THIS DEVICE AND SECTION D4, UDI#, SHALL STATE ¿NONE.¿ THE DEVICE HAS NOT BEEN RETURNED TO INVACARE OR VENTEC (REACT HEALTH) FOR AN EVALUATION. INVACARE CONTINUES TO INVESTIGATE THE REPORTED ISSUE. ONCE COMPLETE, VENTEC WILL BE PROVIDED WITH THEIR INVESTIGATION FINDINGS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE AS DEFINED BY 21 CFR 803.56.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22919791&pc=CAW | 3013095415-2025-00782 | 2025-07-13 04:00 | Death | INVACARE TAYLOR STREET | 2025-08-28 04:00 | CAW | INVACARE® HOMEFILL® OXYGEN SYSTEM | Insufficient Information | Insufficient Information | K003939 |  | 1 | Event Description: THE DEVICE¿S PREVIOUS MANUFACTURER, INVACARE, CONTACTED VENTEC VIA EMAIL TO REPORT THAT THEIR LEGAL DEPARTMENT HAD RECEIVED A LETTER FROM A LAW FIRM WHICH STATED THE FOLLOWING: ¿TO WHOM IT MAY CONCERN, THIS LAW FIRM HAS BEEN RETAINED BY (B)(6) TO REPRESENT ITS SUBROGATION INTERESTS ARISING FROM A FIRE ON (B)(6) 2025, AT (B)(6) (THE ¿PROPERTY¿). (B)(6) PROVIDES PROPERTY INSURANCE TO (B)(6) FOR ITS INTERESTS AT THE PROPERTY. OUR INVESTIGATION INTO THE FIRE IS UNDERWAY. WITHIN THE ROOM OF THE FIRE ORIGIN, WE LOCATED THE FOLLOWING PRODUCTS MANUFACTURED AND/OR DISTRIBUTED BY YOUR COMPANY: INVACARE O2 HOMEFILL II COMPRESSOR (MODEL: IOH200). INVACARE PERFECTO 2 V OXYGEN CONCENTRATOR (MODEL: IRC5PO2V / SERIAL: (B)(6). [THE SERIAL NUMBER FOR THE INVACARE O2 HOMEFILL II COMPRESSOR IS CURRENTLY UNKNOWN]. THE INVACARE LEGAL DEPARTMENT THEN COMPLETED SOME RESEARCH REGARDING THE (B)(6) FIRE. THROUGH THIS RESEARCH, INVACARE BECAME AWARE THAT 10 PEOPLE HAD DIED AS A RESULT OF THE FIRE. THIS MEDWATCH REPORT IS TO DOCUMENT THE DEATH OF PATIENT ¿(B)(6)¿ AS IT PERTAINS TO THE INVACARE® HOMEFILL® OXYGEN SYSTEM. NO FURTHER DETAILS ABOUT THE PATIENT OR THEIR CAUSE OF DEATH WAS PROVIDED. Manufacturer Narrative: H6: THE INITIAL REPORTER DID NOT PROVIDE A DEVICE SERIAL NUMBER FOR THE INVACARE® HOMEFILL® OXYGEN SYSTEM. AS A RESULT, SECTION D4, SERIAL #, AND PRIMARY UDI NUMBER ARE CURRENTLY "UNKNOWN," AND SECTION H4, DEVICE MANUFACTURER DATE, SHALL BE LEFT BLANK. IF THE SERIAL NUMBER IS RECEIVED, THESE FIELDS SHALL BE UPDATED ACCORDINGLY. THE DEVICE HAS NOT BEEN RETURNED TO INVACARE OR VENTEC (REACT HEALTH) FOR AN EVALUATION. INVACARE CONTINUES TO INVESTIGATE THE REPORTED ISSUE. ONCE COMPLETE, VENTEC WILL BE PROVIDED WITH THEIR INVESTIGATION FINDINGS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE AS DEFINED BY 21 CFR 803.56.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22919792&pc=CAW | 3013095415-2025-00783 | 2025-07-13 04:00 | Death | INVACARE TAYLOR STREET | 2025-08-28 04:00 | CAW | INVACARE® HOMEFILL® OXYGEN SYSTEM | Insufficient Information | Insufficient Information | K003939 |  | 1 | Event Description: THE DEVICE¿S PREVIOUS MANUFACTURER, INVACARE, CONTACTED VENTEC VIA EMAIL TO REPORT THAT THEIR LEGAL DEPARTMENT HAD RECEIVED A LETTER FROM A LAW FIRM WHICH STATED THE FOLLOWING: ¿TO WHOM IT MAY CONCERN, THIS LAW FIRM HAS BEEN RETAINED BY (B)(6) COMPANY (B)(6) TO REPRESENT ITS SUBROGATION INTERESTS ARISING FROM A FIRE ON (B)(6) 2025, (B)(6) (THE ¿PROPERTY¿). (B)(6) PROVIDES PROPERTY INSURANCE TO (B)(6) FOR ITS INTERESTS AT THE PROPERTY. OUR INVESTIGATION INTO THE FIRE IS UNDERWAY. WITHIN THE ROOM OF THE FIRE ORIGIN, WE LOCATED THE FOLLOWING PRODUCTS MANUFACTURED AND/OR DISTRIBUTED BY YOUR COMPANY: INVACARE O2 HOMEFILL II COMPRESSOR (MODEL: IOH200) INVACARE PERFECTO 2 V OXYGEN CONCENTRATOR (MODEL: IRC5PO2V / SERIAL: (B)(6).¿ [THE SERIAL NUMBER FOR THE INVACARE O2 HOMEFILL II COMPRESSOR IS CURRENTLY UNKNOWN] THE INVACARE LEGAL DEPARTMENT THEN COMPLETED SOME RESEARCH REGARDING THE ¿(B)(6)¿ FIRE. THROUGH THIS RESEARCH, INVACARE BECAME AWARE THAT 10 PEOPLE HAD DIED AS A RESULT OF THE FIRE. THIS MEDWATCH REPORT IS TO DOCUMENT THE DEATH OF PATIENT (B)(6) AS IT PERTAINS TO THE INVACARE® HOMEFILL® OXYGEN SYSTEM. NO FURTHER DETAILS ABOUT THE PATIENT OR THEIR CAUSE OF DEATH WAS PROVIDED. Manufacturer Narrative: H6: THE INITIAL REPORTER DID NOT PROVIDE A DEVICE SERIAL NUMBER FOR THE INVACARE® HOMEFILL® OXYGEN SYSTEM. AS A RESULT, SECTION D4, SERIAL #, AND PRIMARY UDI NUMBER ARE CURRENTLY "UNKNOWN," AND SECTION H4, DEVICE MANUFACTURER DATE, SHALL BE LEFT BLANK. IF THE SERIAL NUMBER IS RECEIVED, THESE FIELDS SHALL BE UPDATED ACCORDINGLY. THE DEVICE HAS NOT BEEN RETURNED TO INVACARE OR VENTEC (REACT HEALTH) FOR AN EVALUATION. INVACARE CONTINUES TO INVESTIGATE THE REPORTED ISSUE. ONCE COMPLETE, VENTEC WILL BE PROVIDED WITH THEIR INVESTIGATION FINDINGS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE AS DEFINED BY 21 CFR 803.56.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22919794&pc=CAW | 3013095415-2025-00784 | 2025-07-13 04:00 | Death | INVACARE TAYLOR STREET | 2025-08-28 04:00 | CAW | INVACARE® HOMEFILL® OXYGEN SYSTEM | Insufficient Information | Insufficient Information | K003939 |  | 1 | Event Description: THE DEVICE¿S PREVIOUS MANUFACTURER, INVACARE, CONTACTED VENTEC VIA EMAIL TO REPORT THAT THEIR LEGAL DEPARTMENT HAD RECEIVED A LETTER FROM A LAW FIRM WHICH STATED THE FOLLOWING: ¿TO WHOM IT MAY CONCERN, THIS LAW FIRM HAS BEEN RETAINED BY (B)(6) TO REPRESENT ITS SUBROGATION INTERESTS ARISING FROM A FIRE ON (B)(6) 2025, AT (B)(6) (THE ¿PROPERTY¿). (B)(6) PROVIDES PROPERTY INSURANCE TO (B)(6) FOR ITS INTERESTS AT THE PROPERTY. OUR INVESTIGATION INTO THE FIRE IS UNDERWAY. WITHIN THE ROOM OF THE FIRE ORIGIN, WE LOCATED THE FOLLOWING PRODUCTS MANUFACTURED AND/OR DISTRIBUTED BY YOUR COMPANY: INVACARE O2 HOMEFILL II COMPRESSOR (MODEL: IOH200). INVACARE PERFECTO 2 V OXYGEN CONCENTRATOR (MODEL: IRC5PO2V / SERIAL: (B)(6). [THE SERIAL NUMBER FOR THE INVACARE O2 HOMEFILL II COMPRESSOR IS CURRENTLY UNKNOWN]. THE INVACARE LEGAL DEPARTMENT THEN COMPLETED SOME RESEARCH REGARDING THE ¿(B)(6)" FIRE. THROUGH THIS RESEARCH, INVACARE BECAME AWARE THAT 10 PEOPLE HAD DIED AS A RESULT OF THE FIRE. THIS MEDWATCH REPORT IS TO DOCUMENT THE DEATH OF PATIENT ¿(B)(6)¿ AS IT PERTAINS TO THE INVACARE® HOMEFILL® OXYGEN SYSTEM. NO FURTHER DETAILS ABOUT THE PATIENT OR THEIR CAUSE OF DEATH WAS PROVIDED. Manufacturer Narrative: H6: THE INITIAL REPORTER DID NOT PROVIDE A DEVICE SERIAL NUMBER FOR THE INVACARE® HOMEFILL® OXYGEN SYSTEM. AS A RESULT, SECTION D4, SERIAL #, AND PRIMARY UDI NUMBER ARE CURRENTLY "UNKNOWN," AND SECTION H4, DEVICE MANUFACTURER DATE, SHALL BE LEFT BLANK. IF THE SERIAL NUMBER IS RECEIVED, THESE FIELDS SHALL BE UPDATED ACCORDINGLY. THE DEVICE HAS NOT BEEN RETURNED TO INVACARE OR VENTEC (REACT HEALTH) FOR AN EVALUATION. INVACARE CONTINUES TO INVESTIGATE THE REPORTED ISSUE. ONCE COMPLETE, VENTEC WILL BE PROVIDED WITH THEIR INVESTIGATION FINDINGS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE AS DEFINED BY 21 CFR 803.56.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22919799&pc=CAW | 3013095415-2025-00785 | 2025-07-13 04:00 | Death | INVACARE TAYLOR STREET | 2025-08-28 04:00 | CAW | INVACARE® HOMEFILL® OXYGEN SYSTEM | Insufficient Information | Insufficient Information | K003939 |  | 1 | Event Description: THE DEVICE¿S PREVIOUS MANUFACTURER, INVACARE, CONTACTED VENTEC VIA EMAIL TO REPORT THAT THEIR LEGAL DEPARTMENT HAD RECEIVED A LETTER FROM A LAW FIRM WHICH STATED THE FOLLOWING: ¿TO WHOM IT MAY CONCERN, THIS LAW FIRM HAS BEEN RETAINED BY (B)(6) TO REPRESENT ITS SUBROGATION INTERESTS ARISING FROM A FIRE ON (B)(6) 2025, AT (B)(6) (THE ¿PROPERTY¿). (B)(6) PROVIDES PROPERTY INSURANCE TO (B)(6). (¿(B)(6)¿) FOR ITS INTERESTS AT THE PROPERTY. OUR INVESTIGATION INTO THE FIRE IS UNDERWAY. WITHIN THE ROOM OF THE FIRE ORIGIN, WE LOCATED THE FOLLOWING PRODUCTS MANUFACTURED AND/OR DISTRIBUTED BY YOUR COMPANY: - INVACARE O2 HOMEFILL II COMPRESSOR (MODEL: IOH200). - INVACARE PERFECTO 2 V OXYGEN CONCENTRATOR (MODEL: IRC5PO2V / SERIAL: (B)(6)).¿ [THE SERIAL NUMBER FOR THE INVACARE O2 HOMEFILL II COMPRESSOR IS CURRENTLY UNKNOWN]. THE INVACARE LEGAL DEPARTMENT THEN COMPLETED SOME RESEARCH REGARDING THE ¿(B)(6)¿ FIRE. THROUGH THIS RESEARCH, INVACARE BECAME AWARE THAT 10 PEOPLE HAD DIED AS A RESULT OF THE FIRE. THIS MEDWATCH REPORT IS TO DOCUMENT THE DEATH OF PATIENT ¿(B)(6)¿ AS IT PERTAINS TO THE INVACARE® HOMEFILL® OXYGEN SYSTEM. NO FURTHER DETAILS ABOUT THE PATIENT OR THEIR CAUSE OF DEATH WAS PROVIDED. Manufacturer Narrative: H6: THE INITIAL REPORTER DID NOT PROVIDE A DEVICE SERIAL NUMBER FOR THE INVACARE® HOMEFILL® OXYGEN SYSTEM. AS A RESULT, SECTION D4, SERIAL #, AND PRIMARY UDI NUMBER ARE CURRENTLY "UNKNOWN," AND SECTION H4, DEVICE MANUFACTURER DATE, SHALL BE LEFT BLANK. IF THE SERIAL NUMBER IS RECEIVED, THESE FIELDS SHALL BE UPDATED ACCORDINGLY. THE DEVICE HAS NOT BEEN RETURNED TO INVACARE OR VENTEC (REACT HEALTH) FOR AN EVALUATION. INVACARE CONTINUES TO INVESTIGATE THE REPORTED ISSUE. ONCE COMPLETE, VENTEC WILL BE PROVIDED WITH THEIR INVESTIGATION FINDINGS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE AS DEFINED BY 21 CFR 803.56.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22919822&pc=CAW | 3013095415-2025-00786 | 2025-07-13 04:00 | Death | INVACARE TAYLOR STREET | 2025-08-28 04:00 | CAW | INVACARE® HOMEFILL® OXYGEN SYSTEM | Insufficient Information | Insufficient Information | K003939 |  | 1 | Event Description: THE DEVICE¿S PREVIOUS MANUFACTURER, INVACARE, CONTACTED VENTEC VIA EMAIL TO REPORT THAT THEIR LEGAL DEPARTMENT HAD RECEIVED A LETTER FROM A LAW FIRM WHICH STATED THE FOLLOWING: ¿TO WHOM IT MAY CONCERN, THIS LAW FIRM HAS BEEN RETAINED BY (B)(6) TO REPRESENT ITS SUBROGATION INTERESTS ARISING FROM A FIRE ON (B)(6) 2025, AT (B)(6) (THE ¿PROPERTY¿). (B)(6) PROVIDES PROPERTY INSURANCE TO (B)(6). (¿(B)(6)¿) FOR ITS INTERESTS AT THE PROPERTY. OUR INVESTIGATION INTO THE FIRE IS UNDERWAY. WITHIN THE ROOM OF THE FIRE ORIGIN, WE LOCATED THE FOLLOWING PRODUCTS MANUFACTURED AND/OR DISTRIBUTED BY YOUR COMPANY: - INVACARE O2 HOMEFILL II COMPRESSOR (MODEL: IOH200). - INVACARE PERFECTO 2 V OXYGEN CONCENTRATOR (MODEL: IRC5PO2V / SERIAL: (B)(6)).¿ [THE SERIAL NUMBER FOR THE INVACARE O2 HOMEFILL II COMPRESSOR IS CURRENTLY UNKNOWN]. THE INVACARE LEGAL DEPARTMENT THEN COMPLETED SOME RESEARCH REGARDING THE ¿(B)(6)¿ FIRE. THROUGH THIS RESEARCH, INVACARE BECAME AWARE THAT 10 PEOPLE HAD DIED AS A RESULT OF THE FIRE. THIS MEDWATCH REPORT IS TO DOCUMENT THE DEATH OF PATIENT ¿(B)(6)¿ AS IT PERTAINS TO THE INVACARE® HOMEFILL® OXYGEN SYSTEM. NO FURTHER DETAILS ABOUT THE PATIENT OR THEIR CAUSE OF DEATH WAS PROVIDED. Manufacturer Narrative: H6: THE INITIAL REPORTER DID NOT PROVIDE A DEVICE SERIAL NUMBER FOR THE INVACARE® HOMEFILL® OXYGEN SYSTEM. AS A RESULT, SECTION D4, SERIAL #, AND PRIMARY UDI NUMBER ARE CURRENTLY "UNKNOWN," AND SECTION H4, DEVICE MANUFACTURER DATE, SHALL BE LEFT BLANK. IF THE SERIAL NUMBER IS RECEIVED, THESE FIELDS SHALL BE UPDATED ACCORDINGLY. THE DEVICE HAS NOT BEEN RETURNED TO INVACARE OR VENTEC (REACT HEALTH) FOR AN EVALUATION. INVACARE CONTINUES TO INVESTIGATE THE REPORTED ISSUE. ONCE COMPLETE, VENTEC WILL BE PROVIDED WITH THEIR INVESTIGATION FINDINGS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE AS DEFINED BY 21 CFR 803.56.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22919829&pc=CAW | 3013095415-2025-00787 | 2025-07-13 04:00 | Death | INVACARE TAYLOR STREET | 2025-08-28 04:00 | CAW | INVACARE® HOMEFILL® OXYGEN SYSTEM | Insufficient Information | Insufficient Information | K003939 |  | 1 | Event Description: THE DEVICE¿S PREVIOUS MANUFACTURER, INVACARE, CONTACTED VENTEC VIA EMAIL TO REPORT THAT THEIR LEGAL DEPARTMENT HAD RECEIVED A LETTER FROM A LAW FIRM WHICH STATED THE FOLLOWING: ¿TO WHOM IT MAY CONCERN, THIS LAW FIRM HAS BEEN RETAINED BY (B)(6) TO REPRESENT ITS SUBROGATION INTERESTS ARISING FROM A FIRE ON (B)(6) 2025, AT (B)(6) (THE ¿PROPERTY¿). (B)(6) PROVIDES PROPERTY INSURANCE TO (B)(6). (¿(B)(6)¿) FOR ITS INTERESTS AT THE PROPERTY. OUR INVESTIGATION INTO THE FIRE IS UNDERWAY. WITHIN THE ROOM OF THE FIRE ORIGIN, WE LOCATED THE FOLLOWING PRODUCTS MANUFACTURED AND/OR DISTRIBUTED BY YOUR COMPANY: - INVACARE O2 HOMEFILL II COMPRESSOR (MODEL: IOH200) - INVACARE PERFECTO 2 V OXYGEN CONCENTRATOR (MODEL: IRC5PO2V / SERIAL: (B)(6))¿ [THE SERIAL NUMBER FOR THE INVACARE O2 HOMEFILL II COMPRESSOR IS CURRENTLY UNKNOWN] THE INVACARE LEGAL DEPARTMENT THEN COMPLETED SOME RESEARCH REGARDING THE ¿(B)(6)¿ FIRE. THROUGH THIS RESEARCH, INVACARE BECAME AWARE THAT 10 PEOPLE HAD DIED AS A RESULT OF THE FIRE. THIS MEDWATCH REPORT IS TO DOCUMENT THE DEATH OF PATIENT ¿(B)(6)¿ AS IT PERTAINS TO THE INVACARE® HOMEFILL® OXYGEN SYSTEM. NO FURTHER DETAILS ABOUT THE PATIENT OR THEIR CAUSE OF DEATH WAS PROVIDED. Manufacturer Narrative: H6: THE INITIAL REPORTER DID NOT PROVIDE A DEVICE SERIAL NUMBER FOR THE INVACARE® HOMEFILL® OXYGEN SYSTEM. AS A RESULT, SECTION D4, SERIAL #, AND PRIMARY UDI NUMBER ARE CURRENTLY "UNKNOWN," AND SECTION H4, DEVICE MANUFACTURER DATE, SHALL BE LEFT BLANK. IF THE SERIAL NUMBER IS RECEIVED, THESE FIELDS SHALL BE UPDATED ACCORDINGLY. THE DEVICE HAS NOT BEEN RETURNED TO INVACARE OR VENTEC (REACT HEALTH) FOR AN EVALUATION. INVACARE CONTINUES TO INVESTIGATE THE REPORTED ISSUE. ONCE COMPLETE, VENTEC WILL BE PROVIDED WITH THEIR INVESTIGATION FINDINGS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE AS DEFINED BY 21 CFR 803.56.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22919833&pc=CAW | 3013095415-2025-00788 | 2025-07-13 04:00 | Death | INVACARE TAYLOR STREET | 2025-08-28 04:00 | CAW | INVACARE® HOMEFILL® OXYGEN SYSTEM | Insufficient Information | Insufficient Information | K003939 |  | 1 | Event Description: THE DEVICE¿S PREVIOUS MANUFACTURER, INVACARE, CONTACTED VENTEC VIA EMAIL TO REPORT THAT THEIR LEGAL DEPARTMENT HAD RECEIVED A LETTER FROM A LAW FIRM WHICH STATED THE FOLLOWING: ¿TO WHOM IT MAY CONCERN, THIS LAW FIRM HAS BEEN RETAINED BY (B)(6) TO REPRESENT ITS SUBROGATION INTERESTS ARISING FROM A FIRE ON (B)(6) 2025, AT (B)(6) (THE ¿PROPERTY¿). (B)(6) PROVIDES PROPERTY INSURANCE TO (B)(6) FOR ITS INTERESTS AT THE PROPERTY. OUR INVESTIGATION INTO THE FIRE IS UNDERWAY. WITHIN THE ROOM OF THE FIRE ORIGIN, WE LOCATED THE FOLLOWING PRODUCTS MANUFACTURED AND/OR DISTRIBUTED BY YOUR COMPANY: INVACARE O2 HOMEFILL II COMPRESSOR (MODEL: IOH200). INVACARE PERFECTO 2 V OXYGEN CONCENTRATOR (MODEL: IRC5PO2V / SERIAL: (B)(6)¿. [THE SERIAL NUMBER FOR THE INVACARE O2 HOMEFILL II COMPRESSOR IS CURRENTLY UNKNOWN]. THE INVACARE LEGAL DEPARTMENT THEN COMPLETED SOME RESEARCH REGARDING THE ¿(B)(6)¿ FIRE. THROUGH THIS RESEARCH, INVACARE BECAME AWARE THAT 10 PEOPLE HAD DIED AS A RESULT OF THE FIRE. THIS MEDWATCH REPORT IS TO DOCUMENT THE DEATH OF PATIENT ¿(B)(6)¿ AS IT PERTAINS TO THE INVACARE® HOMEFILL® OXYGEN SYSTEM. NO FURTHER DETAILS ABOUT THE PATIENT OR THEIR CAUSE OF DEATH WAS PROVIDED. Manufacturer Narrative: H6: THE INITIAL REPORTER DID NOT PROVIDE A DEVICE SERIAL NUMBER FOR THE INVACARE® HOMEFILL® OXYGEN SYSTEM. AS A RESULT, SECTION D4, SERIAL#, AND PRIMARY UDI NUMBER ARE CURRENTLY "UNKNOWN," AND SECTION H4, DEVICE MANUFACTURER DATE, SHALL BE LEFT BLANK. IF THE SERIAL NUMBER IS RECEIVED, THESE FIELDS SHALL BE UPDATED ACCORDINGLY. THE DEVICE HAS NOT BEEN RETURNED TO INVACARE OR VENTEC (REACT HEALTH) FOR AN EVALUATION. INVACARE CONTINUES TO INVESTIGATE THE REPORTED ISSUE. ONCE COMPLETE, VENTEC WILL BE PROVIDED WITH THEIR INVESTIGATION FINDINGS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE AS DEFINED BY 21 CFR 803.56.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22919843&pc=CAW | 3013095415-2025-00789 | 2025-07-13 04:00 | Death | INVACARE TAYLOR STREET | 2025-08-28 04:00 | CAW | INVACARE® HOMEFILL® OXYGEN SYSTEM | Insufficient Information | Insufficient Information | K003939 |  | 1 | Event Description: THE DEVICE¿S PREVIOUS MANUFACTURER, INVACARE, CONTACTED VENTEC VIA EMAIL TO REPORT THAT THEIR LEGAL DEPARTMENT HAD RECEIVED A LETTER FROM A LAW FIRM WHICH STATED THE FOLLOWING: ¿TO WHOM IT MAY CONCERN, THIS LAW FIRM HAS BEEN RETAINED BY (B)(6) TO REPRESENT ITS SUBROGATION INTERESTS ARISING FROM A FIRE ON (B)(6) 2025, AT (B)(6)(THE ¿PROPERTY¿). (B)(6) PROVIDES PROPERTY INSURANCE TO (B)(6). (¿(B)(6)¿) FOR ITS INTERESTS AT THE PROPERTY. OUR INVESTIGATION INTO THE FIRE IS UNDERWAY. WITHIN THE ROOM OF THE FIRE ORIGIN, WE LOCATED THE FOLLOWING PRODUCTS MANUFACTURED AND/OR DISTRIBUTED BY YOUR COMPANY: - INVACARE O2 HOMEFILL II COMPRESSOR (MODEL: IOH200) - INVACARE PERFECTO 2 V OXYGEN CONCENTRATOR (MODEL: IRC5PO2V / SERIAL: (B)(6)¿. [THE SERIAL NUMBER FOR THE INVACARE O2 HOMEFILL II COMPRESSOR IS CURRENTLY UNKNOWN] THE INVACARE LEGAL DEPARTMENT THEN COMPLETED SOME RESEARCH REGARDING THE ¿(B)(6)¿ FIRE. THROUGH THIS RESEARCH, INVACARE BECAME AWARE THAT 10 PEOPLE HAD DIED AS A RESULT OF THE FIRE. THIS MEDWATCH REPORT IS TO DOCUMENT THE DEATH OF PATIENT ¿(B)(6)¿ AS IT PERTAINS TO THE INVACARE® HOMEFILL® OXYGEN SYSTEM. NO FURTHER DETAILS ABOUT THE PATIENT OR THEIR CAUSE OF DEATH WAS PROVIDED. Manufacturer Narrative: H6: THE INITIAL REPORTER DID NOT PROVIDE A DEVICE SERIAL NUMBER FOR THE INVACARE® HOMEFILL® OXYGEN SYSTEM. AS A RESULT, SECTION D4, SERIAL #, AND PRIMARY UDI NUMBER ARE CURRENTLY "UNKNOWN," AND SECTION H4, DEVICE MANUFACTURER DATE, SHALL BE LEFT BLANK. IF THE SERIAL NUMBER IS RECEIVED, THESE FIELDS SHALL BE UPDATED ACCORDINGLY. THE DEVICE HAS NOT BEEN RETURNED TO INVACARE OR VENTEC (REACT HEALTH) FOR AN EVALUATION. INVACARE CONTINUES TO INVESTIGATE THE REPORTED ISSUE. ONCE COMPLETE, VENTEC WILL BE PROVIDED WITH THEIR INVESTIGATION FINDINGS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE AS DEFINED BY 21 CFR 803.56.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22919851&pc=NFA | 2124215-2025-60503 | 2017-09-01 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-08-28 04:00 | NFA | FILTERWIRE EZ EMBOLIC PROTECTION SYSTEM | Adverse Event Without Identified Device or Use Problem | Myocardial Infarction |  |  | 1 | Event Description: IT WAS REPORTED VIA LITERATURE THAT A PATIENT WHO UNDERWENT A CAROTID ARTERY STENTING (CAS) PROCEDURE USING A FILTER WIRE EZ DIED EVEN AFTER UNDERGOING EMERGENT PERCUTANEOUS CORONARY INTERVENTION (PCI) DUE TO MYOCARDIAL INFARCTION (MI). THE AIM OF THE STUDY WAS TO INVESTIGATE THE PERI-PROCEDURAL RISK FACTORS AND CLINICAL OUTCOMES ASSOCIATED WITH LONG-PERIOD HEMODYNAMIC INSTABILITY (HI) FOLLOWING CAROTID ARTERY STENTING (CAS). HEMODYNAMIC INSTABILITY, MANIFESTING AS HYPERTENSION, HYPOTENSION, OR BRADYCARDIA, IS A KNOWN COMPLICATION AFTER CAS; HOWEVER, ITS PROLONGED FORM-LASTING MORE THAN 24 HOURS-HAS NOT BEEN THOROUGHLY CHARACTERIZED IN TERMS OF PREDICTORS AND CONSEQUENCES. THE RESEARCHERS SOUGHT TO IDENTIFY WHICH PATIENT CHARACTERISTICS, PROCEDURAL TECHNIQUES, AND INTRAOPERATIVE VARIABLES WERE INDEPENDENTLY ASSOCIATED WITH THE DEVELOPMENT OF LONG-PERIOD HI. ADDITIONALLY, THEY EVALUATED SHORT-TERM CLINICAL OUTCOMES SUCH AS TRANSIENT ISCHEMIC ATTACK (TIA), STROKE, MYOCARDIAL INFARCTION, AND MORTALITY LINKED TO THIS CONDITION. THE STUDY INVOLVED 168 PATIENTS WHO UNDERWENT CAS BETWEEN (B)(6) 2017 AND (B)(6) 2020 AT A SINGLE ACADEMIC MEDICAL CENTER IN CHINA. THE AVERAGE AGE OF THE PARTICIPANTS WAS 68.2 YEARS, AND THE MAJORITY WERE MALE (81.5%). IN-HOSPITAL MORTALITY OCCURRED IN 1 PATIENT (0.8%) DESPITE UNDERGOING EMERGENT PERCUTANEOUS CORONARY INTERVENTION (PCI) DUE TO MYOCARDIAL INFARCTION (MI). Manufacturer Narrative: B2: THE EXACT DATE OF DEATH WAS NOT PROVIDED AND WAS ESTIMATED BASED ON THE EARLIEST KNOWN CAS PROCEDURE IN THE COHORT. B3: THE EVENT DATE WAS ESTIMATED TO THE EARLIEST KNOWN PROCEDURE INCLUDED IN THE COHORT. E1: (B)(6). DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER PRODUCT SPECIFIC INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. LI, B., FAN, W., YANG, Y., QU, X., TONG, J., LIU, Y., TAN, J., JIANG, W., & YU, B. (2022). PERI-PROCEDURAL VARIABLES AND OUTCOMES OF LONG-PERIOD HEMODYNAMIC INSTABILITY AFTER CAROTID ARTERY ANGIOPLASTY AND STENTING. VASCULAR, 31(5), 892-901. HTTPS://DOI.ORG/10.1177/17085381221091369.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22919881&pc=CAW | 3013095415-2025-00790 | 2025-07-13 04:00 | Death | INVACARE TAYLOR STREET | 2025-08-28 04:00 | CAW | INVACARE® HOMEFILL® OXYGEN SYSTEM | Insufficient Information | Insufficient Information | K003939 |  | 1 | Event Description: THE DEVICE¿S PREVIOUS MANUFACTURER, INVACARE, CONTACTED VENTEC VIA EMAIL TO REPORT THAT THEIR LEGAL DEPARTMENT HAD RECEIVED A LETTER FROM A LAW FIRM WHICH STATED THE FOLLOWING: ¿TO WHOM IT MAY CONCERN, THIS LAW FIRM HAS BEEN RETAINED BY (B)(6) TO REPRESENT ITS SUBROGATION INTERESTS ARISING FROM A FIRE ON (B)(6) 2025, AT (B)(6) (THE ¿PROPERTY¿). (B)(6) PROVIDES PROPERTY INSURANCE TO (B)(6) FOR ITS INTERESTS AT THE PROPERTY. OUR INVESTIGATION INTO THE FIRE IS UNDERWAY. WITHIN THE ROOM OF THE FIRE ORIGIN, WE LOCATED THE FOLLOWING PRODUCTS MANUFACTURED AND/OR DISTRIBUTED BY YOUR COMPANY: INVACARE O2 HOMEFILL II COMPRESSOR (MODEL: IOH200). INVACARE PERFECTO 2 V OXYGEN CONCENTRATOR (MODEL: IRC5PO2V / SERIAL: (B)(6). [THE SERIAL NUMBER FOR THE INVACARE O2 HOMEFILL II COMPRESSOR IS CURRENTLY UNKNOWN]. THE INVACARE LEGAL DEPARTMENT THEN COMPLETED SOME RESEARCH REGARDING THE ¿(B)(6)¿ FIRE. THROUGH THIS RESEARCH, INVACARE BECAME AWARE THAT 10 PEOPLE HAD DIED AS A RESULT OF THE FIRE. THIS MEDWATCH REPORT IS TO DOCUMENT THE DEATH OF PATIENT ¿(B)(6)¿ AS IT PERTAINS TO THE INVACARE® HOMEFILL® OXYGEN SYSTEM. (B)(6) WAS INITIALLY TRANSPORTED TO A HOSPITAL, WHERE SHE LATER PASSED AWAY FROM THE INJURIES SHE HAD SUSTAINED IN THE FIRE. NO FURTHER DETAILS ABOUT THE PATIENT OR HER CAUSE OF DEATH WAS PROVIDED. Manufacturer Narrative: H6: THE INITIAL REPORTER DID NOT PROVIDE A DEVICE SERIAL NUMBER FOR THE INVACARE® HOMEFILL® OXYGEN SYSTEM. AS A RESULT, SECTION D4, SERIAL #, AND PRIMARY UDI NUMBER ARE CURRENTLY "UNKNOWN," AND SECTION H4, DEVICE MANUFACTURER DATE, SHALL BE LEFT BLANK. IF THE SERIAL NUMBER IS RECEIVED, THESE FIELDS SHALL BE UPDATED ACCORDINGLY. THE DEVICE HAS NOT BEEN RETURNED TO INVACARE OR VENTEC (REACT HEALTH) FOR AN EVALUATION. INVACARE CONTINUES TO INVESTIGATE THE REPORTED ISSUE. ONCE COMPLETE, VENTEC WILL BE PROVIDED WITH THEIR INVESTIGATION FINDINGS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE AS DEFINED BY 21 CFR 803.56.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22919886&pc=CAW | 3013095415-2025-00791 | 2025-07-13 04:00 | Death | INVACARE TAYLOR STREET | 2025-08-28 04:00 | CAW | INVACARE® HOMEFILL® OXYGEN SYSTEM | Insufficient Information | Insufficient Information | K003939 |  | 1 | Event Description: THE DEVICE¿S PREVIOUS MANUFACTURER, INVACARE, CONTACTED VENTEC VIA EMAIL TO REPORT THAT THEIR LEGAL DEPARTMENT HAD RECEIVED A LETTER FROM A LAW FIRM WHICH STATED THE FOLLOWING: ¿TO WHOM IT MAY CONCERN, THIS LAW FIRM HAS BEEN RETAINED BY (B)(6) TO REPRESENT ITS SUBROGATION INTERESTS ARISING FROM A FIRE ON (B)(6) 2025, AT (B)(6) (THE ¿PROPERTY¿). (B)(6) PROVIDES PROPERTY INSURANCE TO (B)(6). (¿(B)(6)¿) FOR ITS INTERESTS AT THE PROPERTY. OUR INVESTIGATION INTO THE FIRE IS UNDERWAY. WITHIN THE ROOM OF THE FIRE ORIGIN, WE LOCATED THE FOLLOWING PRODUCTS MANUFACTURED AND/OR DISTRIBUTED BY YOUR COMPANY: - INVACARE O2 HOMEFILL II COMPRESSOR (MODEL: IOH200). - INVACARE PERFECTO 2 V OXYGEN CONCENTRATOR (MODEL: IRC5PO2V / SERIAL: (B)(6)).¿ [THE SERIAL NUMBER FOR THE INVACARE O2 HOMEFILL II COMPRESSOR IS CURRENTLY UNKNOWN]. THE INVACARE LEGAL DEPARTMENT THEN COMPLETED SOME RESEARCH REGARDING THE ¿(B)(6)¿ FIRE. THROUGH THIS RESEARCH, INVACARE BECAME AWARE THAT 10 PEOPLE HAD DIED AS A RESULT OF THE FIRE. THIS MEDWATCH REPORT IS TO DOCUMENT THE DEATH OF PATIENT ¿(B)(6)¿ AS IT PERTAINS TO THE INVACARE® HOMEFILL® OXYGEN SYSTEM. NO FURTHER DETAILS ABOUT THE PATIENT OR THEIR CAUSE OF DEATH WAS PROVIDED. Manufacturer Narrative: H6: THE INITIAL REPORTER DID NOT PROVIDE A DEVICE SERIAL NUMBER FOR THE INVACARE® HOMEFILL® OXYGEN SYSTEM. AS A RESULT, SECTION D4, SERIAL #, AND PRIMARY UDI NUMBER ARE CURRENTLY "UNKNOWN," AND SECTION H4, DEVICE MANUFACTURER DATE, SHALL BE LEFT BLANK. IF THE SERIAL NUMBER IS RECEIVED, THESE FIELDS SHALL BE UPDATED ACCORDINGLY. THE DEVICE HAS NOT BEEN RETURNED TO INVACARE OR VENTEC (REACT HEALTH) FOR AN EVALUATION. INVACARE CONTINUES TO INVESTIGATE THE REPORTED ISSUE. ONCE COMPLETE, VENTEC WILL BE PROVIDED WITH THEIR INVESTIGATION FINDINGS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE AS DEFINED BY 21 CFR 803.56.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22919965&pc=DSQ | 2916596-2025-05419 | 2025-08-08 04:00 | Death | THORATEC CORPORATION | 2025-08-28 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22901485&pc=EOQ | 2955842-2025-35791 | 2025-07-23 04:00 | Death | INTUITIVE SURGICAL, INC | 2025-08-27 04:00 | EOQ | ION | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Insufficient Information | K182188 |  | 1 | Event Description: IT WAS REPORTED THAT AFTER AN ION ENDOLUMINAL LUNG BIOPSY PROCEDURE, THE PATIENT DID NOT WAKE UP AFTER ANESTHESIA. THE INTUITIVE ENDOLUMINAL TERRITORY ASSOCIATE WAS INFORMED BY AN UNSPECIFIED SITE EMPLOYEE THAT THE EVENT WAS LIKELY DUE TO A STROKE. NO ADDITIONAL INFORMATION WAS PROVIDED. THERE WAS NO REPORT OF ANY ION SYSTEM ISSUES. MULTIPLE ATTEMPTS TO CONTACT THE PHYSICIAN FOR ADDITIONAL INFORMATION WERE MADE; NO RESPONSE HAS BEEN RECEIVED. Manufacturer Narrative: A REVIEW OF THE SYSTEM LOGS FOR THE REPORTED PROCEDURE DATE SHOWED THERE WERE NO RELATED SYSTEM ERRORS DURING THE PROCEDURE THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT A PATIENT OF UNKNOWN AGE AND UNKNOWN MEDICAL HISTORY UNDERWENT AN ION LUNG BIOPSY. POST PROCEDURE THE PATIENT DID NOT WAKE UP AND DIED. THE EVENT WAS ATTRIBUTED TO A STROKE. NO FURTHER DATA IS AVAILABLE FOR REVIEW INCLUDING WHETHER THE STROKE WAS CONFIRMED ON DIAGNOSTIC WORK UP OR SUSPECTED. ATTEMPTS TO OBTAIN FURTHER INFORMATION HAVE BEEN UNSUCCESSFUL TO DATE. THERE WAS NO ALLEGATION OF MALFUNCTION OF THE ION SYSTEM, INSTRUMENTS, OR ACCESSORIES. BASED ON THE AVAILABLE DATA THE EVENT WAS LIKELY PROCEDURE RELATED AND NOT DEVICE RELATED. BRONCHOSCOPY AND BIOPSY IS A MINIMALLY INVASIVE PROCEDURE WITH A LOW COMPLICATION RATE AND RARELY ASSOCIATED WITH FATAL COMPLICATIONS. IN A MULTICENTER PROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES THE TOTAL NUMBER OF ANY COMPLICATIONS WAS REPORTED TO BE 227 (1.08%) WITH 5 (0.02%) TRANSIENT ISCHEMIC ATTACKS AND 4 (0.02%) TOTAL DEATHS BUT NO STROKES. ANOTHER MULTICENTER STUDY REPORTED 13 (2.2%) COMPLICATIONS WITH NO STROKES AND NO DEATHS ASSOCIATED WITH 581 CASES. A RECENT PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 BRONCHOSCOPIES REPORTED 1 ASSOCIATED DEATH (0.08%) AND NO STROKES. ANOTHER PROSPECTIVE SERIES OF 415 ION PROCEDURES REPORTED 1 CVA (0.2%). THE RISK OF STROKE ASSOCIATED WITH GENERAL ANESTHESIA HAS BEEN REPORTED TO VARY FROM 0.1-1.0% IN NON-CARDIAC, NON-NEUROLOGIC, AND NON-MAJOR SURGERY AND AS HIGH AS 10% IN THE SETTING OF BRAIN OR HIGH-RISK CARDIOVASCULAR SURGERY. FACCIOLONGO N, PATELLI M, GASPARINI S, ET AL. INCIDENCE OF COMPLICATIONS IN BRONCHOSCOPY. MULTICENTRE PROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES. MONALDI ARCHIVES FOR CHEST DISEASE. 2009;71(1). OST DE, ERNST A, LEI X, ET AL. DIAGNOSTIC YIELD AND COMPLICATIONS OF BRONCHOSCOPY FOR PERIPHERAL LUNG LESIONS. RESULTS OF THE AQUIRE REGISTRY. AM J RESPIR CRIT CARE MED. 2016;193(1):68-77. FOLCH EE, PRITCHETT MA, NEAD MA, ET AL. ELECTROMAGNETIC NAVIGATION BRONCHOSCOPY FOR PERIPHERAL PULMONARY LESIONS: ONE-YEAR RESULTS OF THE PROSPECTIVE, MULTICENTER NAVIGATE STUDY. JOURNAL OF THORACIC ONCOLOGY. 2019. BROWNLEE AR, WATSON JJJ, AKHMEROV A, ET AL. ROBOTIC NAVIGATIONAL BRONCHOSCOPY IN A THORACIC SURGICAL PRACTICE: LEVERAGING TECHNOLOGY IN THE MANAGEMENT OF PULMONARY NODULES. JTCVS. 2023. BATEMAN BT, SCHUMACHER HC, WANG S, SHAEFI S, BERMAN MF. PERIOPERATIVE ACUTE ISCHEMIC STROKE IN NONCARDIAC AND NONVASCULAR SURGERY: INCIDENCE, RISK FACTORS, AND OUTCOMES. ANESTHESIOLOGY. 2009. SELIM M. PERIOPERATIVE STROKE. NEW ENGLAND JOURNAL OF MEDICINE. 2007.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22901512&pc=DRA | 2124215-2025-60104 | 2025-08-04 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-08-27 04:00 | DRA | FARADRIVE STEERABLE SHEATH CLEAR | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Tachycardia; Cardiac Tamponade; Cardiac Perforation; Pericardial Effusion |  |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT DEATH OCCURRED. DURING A ABLATION PROCEDURE TO TREAT PERSISTENT ATRIAL FIBRILLATION, FEMORAL ACCESS WAS OBTAINED VIA THE LEFT FEMORAL VEIN IN TWO LOCATIONS AND THE RIGHT FEMORAL VEIN IN ONE LOCATION. A NON-BOSTON SCIENTIFIC BIOSENSE WEBSTER CATHETER WAS PLACED, AND A NON-BOSTON SCIENTIFIC CARTO SOUNDSTAR INTRACARDIAC ECHOCARDIOGRAPHY CATHETER (8 FRENCH) WAS INSERTED THROUGH THE LEFT-SIDED ACCESS. TRANSSEPTAL PREPARATION WAS PERFORMED USING THE VERSACROSS CONNECT SYSTEM FOR THE FARADRIVE SHEATH. A NON-BOSTON SCIENTIFIC BIOSENSE RF WIRE WAS FOLLOWED BY THE FARADRIVE SHEATH, WHICH WAS LOADED WITH THE VERSACROSS CONNECT DILATOR AND ADVANCED THROUGH THE RIGHT FEMORAL VEIN. A MID-MID TRANSSEPTAL PUNCTURE WAS SUCCESSFULLY ACHIEVED AND CONFIRMED USING INTRACARDIAC ECHOCARDIOGRAPHY. THE PHYSICIAN ADVANCED THE VERSACROSS WIRE, DILATOR, AND FARADRIVE SHEATH INTO THE LEFT ATRIUM, CONFIRMED TO BE FLOATING IN THE BLOOD POOL VIA ECHOCARDIOGRAPHIC IMAGING. THE DILATOR AND WIRE WERE THEN REMOVED FROM THE FARADRIVE SHEATH. A NON-BOSTON SCIENTIFIC OCTARAY MAPPING CATHETER WITH A 2-2-2 F CURVE WAS ADVANCED INTO THE LEFT ATRIUM, AND MAPPING OF THE LEFT PULMONARY VEINS WAS INITIATED. DURING MAPPING, THE PATIENT EXPERIENCED A SUDDEN AND SEVERE DROP IN BLOOD PRESSURE ALONG WITH AN ELEVATED HEART RATE. INTRACARDIAC ECHOCARDIOGRAPHY REVEALED A PERICARDIAL EFFUSION THAT HAD NOT BEEN PRESENT AT BASELINE. THE ANESTHESIA TEAM PROVIDED HEMODYNAMIC SUPPORT AND INITIATED AUTOTRANSFUSION TO STABILIZE THE PATIENT, WHO WAS ALREADY UNDER GENERAL ANESTHESIA AND INTUBATED. THE PATIENT WAS TRANSFERRED FROM THE ELECTROPHYSIOLOGY PROCEDURAL LABORATORY TO THE CARDIOTHORACIC SURGERY DEPARTMENT TO HAVE A PERICARDIOCENTESIS AND PERICARDIAL DRAIN PLACEMENT. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. IT WAS FURTHER REPORTED THAT THIS PATIENT DID NOT SURVIVE OPEN HEART SURGERY. THE CAUSE OF TAMPONADE WAS IDENTIFIED TO BE PERFORATION OF THE LEFT ATRIAL APPENDAGE. IT IS UNKNOWN TO WHETHER THE DILATOR, SHEATH, OR MAPPING CATHETER HAD CAUSED THE PERFORATION. AS STATED PREVIOUSLY, TRANSSEPTAL PUNCTURE HAD BEEN PERFORMED AND WIRE AND DILATOR EXCHANGED FOR MAPPING CATHETER. THE PATIENT BECAME UNSTABLE, AND EFFUSION WAS IDENTIFIED WHILE MAPPING THE LEFT SIDED STRUCTURES. Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22901539&pc=DRA | 2124215-2025-60081 | 2025-07-29 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-08-27 04:00 | DRA | FARADRIVE STEERABLE SHEATH CLEAR | Difficult to Advance | Pericardial Effusion |  |  | 1 | Event Description: IT WAS REPORTED THAT DURING A CONCOMITANT PULSE FIELD ABLATION (PFA) AND WATCHMAN LEFT ATRIAL APPENDAGE CLOSURE PROCEDURE USING A FARADRIVE SHEATH THE PATIENT EXPERIENCED PERICARDIAL EFFUSION AND PASSED AWAY. PRIOR TO THE PROCEDURE A BASELINE EFFUSION WAS NOTED. DURING THE PROCEDURE IT WAS DIFFICULT TO ADVANCE THE FARADRIVE SHEATH THROUGH THE FEMORAL VEIN. DURING TRANSSEPTAL IT WAS NOTED TO BE DIFFICULT, AND THE PHYSICIAN FLOSSED THE SEPTUM WITH THE VERSACROSS LARGE ACCESS DILATOR, BUT THEY WERE STILL UNABLE TO CROSS WITH THE FARADRIVE SHEATH. THE CROSSING WAS THEN ATTEMPTED WITH THE FARADRIVE'S NATIVE DILATOR, WHICH STILL DID NOT CROSS. FINALLY, THE FARADRIVE WAS ABLE TO CROSS WHEN USING A THIRD DIFFERENT TYPE OF DILATOR. IT WAS NOTED THAT THE PATIENT'S PRESSURE DROPPED OCCASIONALLY DURING THE CROSSING AND THE "PRESSURE WAS TREATED". DURING THE PROCEDURE IT WAS OBSERVED THAT THE EXISTING EFFUSION STARTED TO GROW AND PERICARDIOCENTESIS WAS PERFORMED TO DRAIN 1 LITER FROM THE PATIENT. THE PATIENT WENT TO THE ICU WHERE THEY LATER PASSED AWAY. THE SHEATH IS NOT AVAILABLE FOR RETURN AS IT WAS DISCARDED BY THE FACILITY. Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22901785&pc=LJT | 3006260740-2025-06284 | 2024-11-22 05:00 | Death | C.R. BARD, INC. (BASD) -3006260740 | 2025-08-27 04:00 | LJT | UNKNOWN POWER PORT | Fluid/Blood Leak; Difficult to Flush; Fracture | Extravasation; Unspecified Infection; Necrosis; Discomfort; Swelling/ Edema | UNKNOWN |  | 1 | Event Description: IN (B)(6) 2023, A PATIENT UNDERWENT A BD PORT PLACEMENT FOR CHEMOTHERAPY ADMINISTRATION. ON (B)(6) 2024, DURING A ROUTINE FLUSH FOLLOWING CHEMOTHERAPY PUMP REMOVAL, IT WAS REPORTED THAT THE DEVICE WAS ALLEGEDLY FLUSHED RAPIDLY BY NURSING STAFF. THE PATIENT IMMEDIATELY EXPRESSED DISCOMFORT, AND SWELLING WAS NOTED. THE NURSE REPORTEDLY ADVISED THE PATIENT TO APPLY AN ICE PACK. IT WAS FURTHER REPORTED THAT WITHIN DAYS, THE PATIENT WAS HOSPITALIZED WITH A SEVERE INFECTION AND TISSUE NECROSIS. CLINICAL FINDINGS ALLEGEDLY SUGGESTED EXTRAVASATION OF CHEMOTHERAPY AND FLUSH SOLUTION INTO THE CHEST CAVITY, POTENTIALLY DUE TO A RUPTURE OF THE DEVICE. THE PORT WAS SURGICALLY REMOVED AND DISCARDED. THE PATIENT CONTINUED TREATMENT UNTIL HIS DEATH ON (B)(6) 2025. THE REPORTER ATTRIBUTES THE OUTCOME TO IMPROPER USE BY HOSPITAL STAFF AND DOES NOT ALLEGE A PRODUCT DEFECT. HOWEVER, FOLLOW-UP IS BEING CONDUCTED TO ASSESS ANY POTENTIAL RELATIONSHIP BETWEEN THE DEVICE AND THE REPORTED DEATH. Manufacturer Narrative: H11: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, PHOTOS WERE PROVIDED FOR REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. SECTION A THROUGH F: THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22903537&pc=BZD | 2518422-2025-109392 | 2025-05-15 04:00 | Death | RESPIRONICS, INC. | 2025-08-27 04:00 | BZD | DREAMMAPPER WEB (DM) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K152356 |  | 1 | Event Description: MANUFACTURER RECEIVED A CALL FROM THE PATIENT'S RELATIVE ASKING WHAT TO DO WITH THE MACHINE, BECAUSE THE PATIENT HAS PASSED AWAY. IT IS UNCLEAR IF AN EVENT MAY HAVE CAUSED OR CONTRIBUTED TO THE PATIENT'S DEATH. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22903650&pc=NIM | 2124215-2025-59192 | 2015-01-01 05:00 | Death | SILK ROAD MEDICAL | 2025-08-27 04:00 | NIM | ENROUTE TRANSCAROTID STENT | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: IT WAS REPORTED VIA LITERATURE THAT SOME PATIENTS WHO UNDERWENT A TRANSCAROTID ARTERY REVASCULARIZATION (TCAR) PROCEDURE EXPERIENCED DEATH. THIS STUDY WAS A RETROSPECTIVE DATA REVIEW OF 321 PATIENTS BETWEEN 2015 THROUGH 2024 WHO UNDERWENT TCAR. THE DATA WAS ANALYZED AS THE POSTOPERATIVE (TIMELINE NOT DEFINED) OUTCOMES BY ANESTHESIA TYPE. THE PERIOPERATIVE ALL-CAUSE MORTALITY IDENTIFIED FOR PATIENTS THAT UNDERWENT GENERAL ANESTHESIA WAS ABOUT 2%. THE PERIOPERATIVE ALL-CAUSE MORTALITY IDENTIFIED FOR PATIENTS THAT UNDERWENT LOCAL ANESTHESIA WAS ABOUT 1%. NO FURTHER INFORMATION WAS PROVIDED TO BOSTON SCIENTIFIC. Manufacturer Narrative: B2. THE DATE OF DEATH WAS NOT PROVIDED FOR ANY OF THE REPORTED DEATHS. B3. EVENT DATE WAS ESTIMATED BASED ON THE EARLIEST PROCEDURE DATE NOTED IN THE LITERATURE ARTICLE. D6A. IMPLANT DATE WAS ESTIMATED BASED ON THE EARLIEST PROCEDURE DATE NOTED IN THE LITERATURE ARTICLE. H6. E2401 WAS USED, AS THE CAUSE OF DEATH WAS NOT REPORTED IN ANY CASE. DETAILS OF THE EVENT, INCLUDING PRODUCT INFORMATION, WERE NOT PROVIDED TO BSC. BECAUSE THE PRODUCT LOT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) NUMBER AND OTHER SPECIFIC PRODUCT INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. KEDWAI, B. J., ESPER, B., LYONS, D. C., & STONER, M. C. (2025). LOCAL ANESTHESIA AND ENHANCED RECOVERY AFTER TRANSCAROTID ARTERY REVASCULARIZATION. ANNALS OF VASCULAR SURGERY, 113, 363-369. HTTPS://DOI.ORG/10.1016/J.AVSG.2024.08.012.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22904100&pc=NIK | 9614453-2025-03404 | 2025-08-01 04:00 | Death | MEDTRONIC EUROPE SARL | 2025-08-27 04:00 | NIK | COBALT¿ HF QUAD CRT-D MRI SURESCAN¿ | Inappropriate/Inadequate Shock/Stimulation; Device Sensing Problem | Shock from Patient Lead(s) | P010031 |  | 1 | Event Description: IT WAS REPORTED THAT THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) DELIVERED POSSIBLE INAPPROPRIATE THERAPY FOR ATRIAL FIBRILLATION WITH RAPID VENTRICULAR RESPONSE THAT WAS DETECTED AS VENTRICULAR FIBRILLATION. IT WAS FURTHER NOTED THAT THE PATIENT PASSED AWAY NINE DAYS LATER. Manufacturer Narrative: CONTINUATION OF D10: 429888 LEAD, IMPLANTED: (B)(6) 2018. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22904359&pc=OZD | 1220648-2025-46012 | 2025-07-29 04:00 | Death | ABIOMED, INC. | 2025-08-27 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT IMPLANTED WITH AN IMPELLA CP LOST DISTAL PULSES. THE IMPELLA CP WAS EXPLANTED AND REPLACED WITH AN IMPELLA 5.5. DISTAL PULSES WERE RESTORED. THE COMPLAINANT REPORTED THAT A PATIENT UNDERWENT SINGLE VESSEL CORONARY ARTERY BYPASS GRAFT SURGERY AND RIGHT CORONARY ARTERY FEEDING THE LEFT ANTERIOR DESCENDING ARTERY. THE PLAN WAS TO REST AND RECOVER THE HEART AS THE PATIENT TOLERATED SURGERY WELL. UPON EXITING THE OPERATING ROOM THE PATIENT HAD CARDIAC ARREST. CARDIOPULMONARY RESUSCITATION WAS PERFORMED AND THE PATIENT WAS DEFIBRILLATED MULTIPLE TIMES. THE DECISION WAS MADE TO PLACE AN IMPELLA CP. DUE TO SIGNIFICANT CALCIFIC DISEASE THERE WAS NO FLOW FOUND DOWN THE IMPELLA ACCESS LEG. THE PATIENT WAS TAKEN TO THE CATHETERIZATION LAB FOR PLACEMENT OF A PULMONARY ARTERY CATHETER AND EXTERNAL FEMORAL TO FEMORAL BYPASS DUE TO THE LIMB ISCHEMIA. ON DAY TWO OF IMPELLA CP SUPPORT DISTAL PULSES WERE NOT DETECTED AND A DECISION WAS MADE TO ESCALATE THE PATIENT TO AN IMPELLA 5.5. THE IMPELLA CP WAS SUCCESSFULLY REMOVED SURGICALLY AND DISTAL PULSES RETURNED. UPON IMPELLA CP EXPLANT, THE PATIENT BEGAN DECOMPENSATING. THE PATIENT CONTINUED TO DECOMPENSATE AFTER THE IMPELLA 5.5 WAS IMPLANTED AND THE PATIENT WAS PLACED ON SUPPORT. THE PATIENT WAS THEN ADDITIONALLY CANNULATED FOR VENO-ARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION. IT WAS FURTHER REPORTED THAT THE PATIENT EXPIRED WHILE ON IMPELLA 5.5 SUPPORT, 8 DAYS AFTER THE IMPELLA CP WAS EXPLANTED. Manufacturer Narrative: THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN. BASED ON THE INFORMATION AVAILABLE THE IMPELLA CP CANNOT BE EXCLUDED AS A POSSIBLE CONTRIBUTING FACTOR TO THE PATIENT¿S OUTCOME. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22904562&pc=PCU | 3005099803-2025-04348 | 2025-07-16 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-08-27 04:00 | PCU | AXIOS STENT AND ELECTROCAUTERY ENHANCED DELIVERY SYSTEM | Migration | Abdominal Pain; Abscess; Fever; Hematoma; Hemorrhage/Bleeding; Unspecified Infection; Perforation; Insufficient Information |  |  | 1 | Event Description: THIS REPORT PERTAINS TO ONE OF THREE AXIOS DEVICES USED IN THE SAME PROCEDURE. REFER TO MANUFACTURER REPORT# 3005099803-2025-03893, 3005099803-2025-04348 AND 3005099803-2025-04349 FOR THE ASSOCIATED DEVICE INFORMATION. IT WAS REPORTED TO BOSTON SCIENTIFIC CORPORATION THAT THREE AXIOS STENT AND ELECTROCAUTERY-ENHANCED DELIVERY SYSTEMS WERE IMPLANTED O TREAT A GASTRIC BILIARY OBSTRUCTION DURING A PROCEDURE PERFORMED ON (B)(6) 2025. DURING THE PROCEDURE, IT WAS DIFFICULT TO PASS THE GUIDEWIRE THROUGH THE DUODENAL STENT TO THE JEJUNUM. AFTER LOCATING THE JEJUNAL LOOP, THE FIRST AXIOS STENT (SUBJECT OF MANUFACTURER REPORT # 3005099803-2025-03893) WAS IMPLANTED, BUT NO DRAINAGE OF METHYLENE BLUE WAS NOTED. THE AXIOS STENT WAS PLACED SMOOTHLY, BUT ON DEPLOYMENT OF THE PROXIMAL FLANGE, THE STENT WAS MISPLACED WITH A TYPE 3 COMPLICATION. THE FIRST AXIOS WAS DEPLOYED INTO A FLUID FILLED LOOP OF BOWEL WHICH APPEARED TO BE SMALL BOWEL BUT ULTIMATELY WAS FOUND TO BE THE COLON. ONCE THIS WAS DISCOVERED, THE STENT WAS LEFT IN SITU WITH A PLAN TO REMOVE IT A FEW WEEKS LATER. A CT SCAN (COMPUTED TOMOGRAPHY) WAS PERFORMED, AND IT SHOWED THE SMALL BOWEL WAS DISPLACED TO THE LEFT OF MIDLINE AND TO THE LEFT OF THE COLON, AND A SMALL BOWEL LOOP WAS IDENTIFIED ENDOSONOGRAPHICALLY. A SECOND AXIOS STENT (SUBJECT OF THIS REPORT) WAS DEPLOYED. WHEN THE PROXIMAL FLANGE WAS DEPLOYED, THE PHYSICIAN OBSERVED THE STENT RETRACT OUT OF THE STOMACH. WIRE ACCESS WAS MAINTAINED THROUGH THE MISDEPLOYED STENT INTO THE SMALL BOWEL, AND THIS WAS USED TO SUCCESSFULLY PLACE THE THIRD AXIOS STENT (SUBJECT OF MANUFACTURER REPORT # 3005099803-2025-04349) COAXIALLY. TWO HOURS POST STENT-IN STENT PROCEDURE, THE PATIENT EXPERIENCED ABDOMINAL PAIN. DURING THE NIGHT, THE STENT-IN-STENT WAS DISLODGED. THE FIRST STENT WAS SEEN IN SITU. THE OUTER STENT, REFERRING TO THE SECOND STENT, WAS COMPLETELY DETACHED, AND THE INNER STENT, REFERRING TO THE THIRD STENT, WAS IN PROPER POSITION. THE PATIENT WAS SENT TO SURGERY AND SURGICAL GASTROENTEROSTOMY WAS PERFORMED, AND A PERFORATION WAS NOTED. POST SURGICAL PROCEDURE, THE PATIENT DEVELOPED AN INFECTED HEMATOMA (PERITONEAL) AND WAS DRAINED PERCUTANEOUSLY. THE HEMATOMA WAS REPORTED AS A CLINICALLY SIGNIFICANT BLEEDING EVENT WHICH RESULTED IN A TRANSFUSION. THE STENTS WERE THEN REMOVED, AND PROPHYLACTIC ANTIBIOTICS WERE ADMINISTERED, AND THE PATIENT WAS HOSPITALIZED. ON (B)(6) 2025, THE PATIENT EXPERIENCED FEVER UNDER ANTIBIOTICS. A CT ABDOMEN SCAN (COMPUTED TOMOGRAPHY OF THE ABDOMEN) WAS PERFORMED, AND REVEALED A DEVELOPMENT OF A DISTINCT, BRANCHING, HETEROGENOUS ABSCESS, LARGELY IN THE LEFT HEMI-ABDOMEN. ON (B)(6) 2025, A DIGESTIVE TRACT LEAK/PERFORATION WAS NOTED. ON (B)(6) 2025, THE PATIENT PASSED AWAY. Manufacturer Narrative: BLOCK E1: INITIAL REPORTER'S SECOND PHONE NUMBER IS (B)(6); REPORTED HERE AS PHONE NUMBER EXCEEDED CHARACTER LIMIT FOR DESIGNATED FIELD. BLOCK H6: IMDRF DEVICE CODE A010402 CAPTURES THE REPORTABLE EVENT OF STENT MIGRATION. IMDRF PATIENT CODE E2114 CAPTURES THE REPORTABLE EVENT OF PERFORATION. IMDRF PATIENT CODE E1002 CAPTURES THE REPORTABLE EVENT OF ABDOMINAL PAIN. IMDRF PATIENT CODE E0505 CAPTURES THE REPORTABLE EVENT OF HEMATOMA. IMDRF PATIENT CODE E1906 CAPTURES THE REPORTABLE EVENT OF INFECTION. IMDRF PATIENT CODE E0506 CAPTURES THE REPORTABLE EVENT OF MAJOR HEMORRHAGE. IMDRF PATIENT CODE E2401 CAPTURES THE REPORTABLE EVENT OF INJURY, NOS (NOT OTHERWISE SPECIFIED). IMDRF PATIENT CODE E172001 CAPTURES THE REPORTABLE EVENT OF ABSCESS. IMDRF PATIENT CODE E230101 CAPTURES THE REPORTABLE EVENT OF FEVER. IMDRF IMPACT CODE F08 CAPTURES THE REPORTABLE EVENT OF PATIENT WAS HOSPITALIZED. IMDRF IMPACT CODE F19 CAPTURES THE REPORTABLE EVENT OF SURGICAL GASTROENTEROSTOMY WAS PERFORMED. IMDRF IMPACT CODE F2301 CAPTURES THE REPORTABLE EVENT OF STENT-IN-STENT PROCEDURE. IMDRF IMPACT CODE F2303 CAPTURES THE REPORTABLE EVENT OF THE PROPHYLACTIC ANTIBIOTICS WERE ADMINISTERED. IMDRF IMPACT CODE F02 CAPTURES THE REPORTABLE EVENT OF THE PATIENT PASSED AWAY. IMDRF IMPACT CODE F2302 CAPTURES THE REPORTABLE EVENT OF THE TRANSFUSION. IMDRF IMPACT CODES F23 AND F22 CAPTURES THE REPORTABLE EVENT OF A CT ABDOMEN SCAN (COMPUTED TOMOGRAPHY OF THE ABDOMEN) WAS PERFORMED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22905011&pc=PNJ | 2017865-2025-98757 | 2025-08-08 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-08-27 04:00 | PNJ | AVEIR¿ DELIVERY CATHETER | Activation, Positioning or Separation Problem | Cardiac Tamponade; Cardiac Perforation; Pericardial Effusion | P150035 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED IN CLINIC FOR LEADLESS PACEMAKER (LP) IMPLANT PROCEDURE. DURING THE PROCEDURE, THE LP EXHIBITED AN LARGE INCREASE IN PACING IMPEDANCE UPON FIXATION. IT WAS NOTED THAT DURING REPOSITIONING, THE CATHETER SLIPPED FORWARD, TRAVELING BEYOND THE DEVICE. CARDIAC PERFORATION WAS NOTED AFTER INJECTION OF CONTRAST. ECHOCARDIOGRAM CONFIRMED PERCIARDIAL EFFUSION. DURING TRANSPORT, THE PATIENT DEVELOPED CARDIAC TAMPONADE AND DUE TO MEDICAL DIRECTIVE OF DNR/DNI, THE PATIENT HAD DECEASED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22905122&pc=SCZ | 3030471525-2025-00004 | 2025-07-29 04:00 | Death | PI-CARDIA LTD. | 2025-08-27 04:00 | SCZ | SHORTCUT | Difficult or Delayed Positioning; Inadequate or Insufficient Training | Pulmonary Embolism; Cardiac Tamponade; Cardiac Perforation | DEN240017 |  | 1 | Event Description: DURING THE PLANNED VALVE-IN-VALVE (VIV) PROCEDURE IN A HIGH SURGICAL RISK PATIENT, A SHORTCUT DEVICE WAS USED FOR A SINGLE RC LEAFLET SPLIT DUE TO RISK OF CORONARY OBSTRUCTION. THE PATIENT, A (B)(6) YEAR-OLD FRAIL FEMALE WITH STRUCTURAL VALVE DETERIORATION, MANIFESTING AS BOTH AORTIC STENOSIS (AS) AND AORTIC REGURGITATION (AR) IN A PREVIOUSLY IMPLANTED 23 MM SAPIEN 3 VALVE, ALSO PRESENTED WITH A PRE-EXISTING HYPERTROPHIC HEART AND A TRACE-TO-MILD PERICARDIAL EFFUSION. THE PROCEDURE WAS INITIATED AS PLANNED. WHILE ADVANCING THROUGH THE DESCENDING AORTA AND INTO THE ARCH, THE GW INITIALLY APPEARED WELL POSITIONED IN THE LV APEX BUT SUBSEQUENTLY PROLAPSED, REQUIRING REPOSITIONING WITH TRACTION APPLIED BY THE SECOND OPERATOR UNDER PHYSICIAN INSTRUCTION TO FACILITATE TRAVERSAL OF THE ARCH. THE SHORTCUT WAS ADVANCED OVER A SAFARI SMALL GUIDEWIRE. AFTER CROSSING THE ARCH, THE DEVICE WAS UNSHEATHED AND MANIPULATED TOWARD THE RC LEAFLET. AT THAT POINT, THE PATIENT BECAME HEMODYNAMICALLY UNSTABLE, AND ECHOCARDIOGRAPHY SHOWED PROGRESSION OF A PRE-EXISTING PERICARDIAL EFFUSION. THE PA APPEARED TO REORIENT MORE ANTERIORLY AND BECAME ENGAGED AT THE RC/LC COMMISSURE OF THE S3 FRAME. ADDITIONAL FLEXION WAS APPLIED TO DIRECT THE DEVICE AWAY FROM THE COMMISSURE, ALLOWING THE PA TO DESCEND INTO APPROPRIATE POSITION ON THE RC LEAFLET. THE ANESTHESIA TEAM REPORTED THAT LV FUNCTION APPEARED ABNORMAL, WITH A REDUCTION IN CONTRACTILITY OBSERVED. TWO ACTIVATION ATTEMPTS WERE MADE, BUT CORRECT POSITIONING OF THE SPLITTING ELEMENT COULD NOT BE CONFIRMED. THE PATIENT DEVELOPED TAMPONADE REQUIRING CPR AND PERICARDIOCENTESIS WITH EVACUATION OF A LARGE BLOODY EFFUSION. THE PROCEDURE WAS ABORTED, AND THE DEVICE WAS RE-SHEATHED AND WITHDRAWN. SUSPECTED LV PERFORATION WAS NOTED, AND CPR AND PERICARDIOCENTESIS WAS PERFORMED TO REMOVE THE BLOOD. AFTER THE PATIENT WAS STABILIZED, CLOTS WERE REMOVED FROM THE RIGHT ARTIUM, AND THE PATIENT WAS MOVED TO THE ICU WITH A DRAIN. PI-CARDIA WAS INFORMED BY THE CLINICAL SITE ON THE DAY FOLLOWING THE PROCEDURE THAT THE PATIENT HAD PASSED AWAY. NO ADDITIONAL DETAILS REGARDING THE EXACT TIMING OR SPECIFIC CIRCUMSTANCES OF DEATH WERE PROVIDED. Manufacturer Narrative: THE PROCEDURE WAS INITIATED AS PLANNED. WHILE ADVANCING THROUGH THE DESCENDING AORTA AND INTO THE ARCH, THE GW INITIALLY APPEARED WELL POSITIONED IN THE LV APEX BUT SUBSEQUENTLY PROLAPSED, REQUIRING REPOSITIONING WITH TRACTION APPLIED BY THE SECOND OPERATOR UNDER PHYSICIAN INSTRUCTION TO FACILITATE TRAVERSAL OF THE ARCH. THE SHORTCUT WAS ADVANCED OVER A SAFARI SMALL GUIDEWIRE. AFTER CROSSING THE ARCH, THE DEVICE WAS UNSHEATHED AND MANIPULATED TOWARD THE RC LEAFLET. AT THAT POINT, THE PATIENT BECAME HEMODYNAMICALLY UNSTABLE, AND ECHOCARDIOGRAPHY SHOWED PROGRESSION OF A PRE-EXISTING PERICARDIAL EFFUSION. THE PA APPEARED TO REORIENT MORE ANTERIORLY AND BECAME ENGAGED AT THE RC/LC COMMISSURE OF THE S3 FRAME. ADDITIONAL FLEXION WAS APPLIED TO DIRECT THE DEVICE AWAY FROM THE COMMISSURE, ALLOWING THE PA TO DESCEND INTO APPROPRIATE POSITION ON THE RC LEAFLET. THE ANAESTHESIA TEAM REPORTED THAT LV FUNCTION APPEARED ABNORMAL, WITH A REDUCTION IN CONTRACTILITY OBSERVED. TWO ACTIVATION ATTEMPTS WERE MADE, BUT CORRECT POSITIONING OF THE SPLITTING ELEMENT COULD NOT BE CONFIRMED. THE PATIENT DEVELOPED TAMPONADE REQUIRING CPR AND PERICARDIOCENTESIS WITH EVACUATION OF A LARGE BLOODY EFFUSION. THE PROCEDURE WAS ABORTED, AND THE DEVICE WAS RE-SHEATHED AND WITHDRAWN. SUSPECTED LV PERFORATION WAS NOTED, AND CPR AND PERICARDIOCENTESIS WAS PERFORMED TO REMOVE THE BLOOD. AFTER THE PATIENT WAS STABILIZED, CLOTS WERE REMOVED FROM THE RIGHT ARTIUM, AND THE PATIENT WAS MOVED TO THE ICU WITH A DRAIN. PI-CARDIA WAS INFORMED BY THE CLINICAL SITE ON THE DAY FOLLOWING THE PROCEDURE THAT THE PATIENT HAD PASSED AWAY. NO ADDITIONAL DETAILS REGARDING THE EXACT TIMING OR SPECIFIC CIRCUMSTANCES OF DEATH WERE PROVIDED. THE SHORTCUT DEVICE USED DURING THE PROCEDURE WAS NOT RETURNED TO PI-CARDIA FOR INVESTIGATION DUE TO INTERNAL HOSPITAL POLICY. THEREFORE, A DIRECT DEVICE INSPECTION COULD NOT BE PERFORMED. THE INVESTIGATION RELIED ON INFORMATION PROVIDED BY THE INFORMATION FROM HOSPITAL PROCEDURAL SUMMARY AND PI-CARDIA'S TD TEAM MEMBERS WHO WERE PRESENT DURING THE PROCEDURE. A LIMITED INVESTIGATION WAS UNDERTAKEN, WHICH INCLUDED REVIEW OF TRAINING RECORDS, PROCEDURE SUMMARY, AND THE DEVICE HISTORY RECORD (DHR). PI-CARDIA REP CONFIRMED THAT DURING THE PREPARATION OF THE DEVICE IT WAS STRUCTURALLY INTACT AND FUNCTIONING APPROPRIATELY WITH NO EVIDENCE OF MANUFACTURING OR RELEASE-RELATED ISSUES. PER THE CASE RECORDS AND INVESTIGATION FINDINGS, THE PERFORATION WAS POSSIBLY RELATED TO GUIDEWIRE MANIPULATIONS DURING ADVANCEMENT THROUGH THE DESCENDING AORTA AND ARCH, WHEN THE WIRE PROLAPSED FROM THE LV APEX AND REQUIRED REPOSITIONING WITH TRACTION. THE PROCEDURE WAS PERFORMED BY A TRAINED PHYSICIAN ASSISTED BY UNTRAINED LAB TECHNICIAN. THE WIRE, WHICH HAD INITIALLY APPEARED WELL POSITIONED IN THE LV APEX, SUBSEQUENTLY PROLAPSED AND REQUIRED TRACTION AND REPOSITIONING. THESE WIRE MANIPULATIONS WERE PERFORMED BY THE UNTRAINED LAB TECHNICIAN, MAY HAVE RESULTED IN EXCESSIVE WIRE PRESSURE WITHIN THE VENTRICLE WHICH COULD INCREASE THE RISK OF PERFORATION, PARTICULARLY IN THE CONTEXT OF THE PATIENT'S STRUCTURALLY VULNERABLE HEART. THE INSTRUCTIONS FOR USE SPECIFY THAT THE SHORTCUT IS INTENDED FOR USE BY TRAINED INTERVENTIONAL CARDIOLOGISTS OR CARDIAC SURGEONS. IN THIS CASE, GUIDEWIRE MANIPULATIONS WERE PERFORMED BY A LAB TECHNICIAN WHO HAD NOT RECEIVED SHORTCUT TRAINING. THE QUALIFIED PHYSICIAN OPERATOR WAS APPROPRIATELY TRAINED, BUT DUE TO THE ABSENCE OF TRAINED ASSISTANTS, THE TECHNICIAN WAS ENGAGED IN A SUPPORTING ROLE. LV PERFORATION IS A KNOWN PROCEDURAL RISK IN STRUCTURAL HEART INTERVENTIONS ASSOCIATED WITH GUIDE WIRE MANIPULATION, INCLUDING TAVR. THE DEVICE'S DESIGN, INSTRUCTIONS FOR USE, AND PHYSICIAN TRAINING ADDRESS SUCH RISKS. AS IN TAVR PROCEDURES, CAREFUL ATTENTION TO WIRE MANIPULATION IS CRITICAL TO PREVENT COMPLICATIONS LIKE LEFT VENTRICULAR (LV) PERFORATION. BASED ON THE INVESTIGATION, THE RESIDUAL RISK REMAINS ACCEPTABLE, AND NO ADDITIONAL RISK CONTROLS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22905227&pc=PCU | 3005099803-2025-04349 | 2025-07-16 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-08-27 04:00 | PCU | AXIOS STENT AND ELECTROCAUTERY ENHANCED DELIVERY SYSTEM | Migration | Abdominal Pain; Abscess; Fever; Hematoma; Hemorrhage/Bleeding; Unspecified Infection; Perforation; Insufficient Information |  |  | 1 | Event Description: THIS REPORT PERTAINS TO ONE OF THREE AXIOS DEVICES USED IN THE SAME PROCEDURE. REFER TO MANUFACTURER REPORT#: 3005099803-2025-03893, 3005099803-2025-04348. FOR THE ASSOCIATED DEVICE INFORMATION. IT WAS REPORTED TO BOSTON SCIENTIFIC CORPORATION THAT THREE AXIOS STENT AND ELECTROCAUTERY-ENHANCED DELIVERY SYSTEMS WERE IMPLANTED O TREAT A GASTRIC BILIARY OBSTRUCTION DURING A PROCEDURE PERFORMED ON (B)(6) 2025. DURING THE PROCEDURE, IT WAS DIFFICULT TO PASS THE GUIDEWIRE THROUGH THE DUODENAL STENT TO THE JEJUNUM. AFTER LOCATING THE JEJUNAL LOOP, THE FIRST AXIOS STENT (SUBJECT OF MANUFACTURER REPORT#: 3005099803-2025-03893) WAS IMPLANTED, BUT NO DRAINAGE OF METHYLENE BLUE WAS NOTED. THE AXIOS STENT WAS PLACED SMOOTHLY, BUT ON DEPLOYMENT OF THE PROXIMAL FLANGE, THE STENT WAS MISPLACED WITH A TYPE 3 COMPLICATION. THE FIRST AXIOS WAS DEPLOYED INTO A FLUID FILLED LOOP OF BOWEL WHICH APPEARED TO BE SMALL BOWEL BUT ULTIMATELY WAS FOUND TO BE THE COLON. ONCE THIS WAS DISCOVERED, THE STENT WAS LEFT IN SITU WITH A PLAN TO REMOVE IT A FEW WEEKS LATER. A CT SCAN (COMPUTED TOMOGRAPHY) WAS PERFORMED, AND IT SHOWED THE SMALL BOWEL WAS DISPLACED TO THE LEFT OF MIDLINE AND TO THE LEFT OF THE COLON, AND A SMALL BOWEL LOOP WAS IDENTIFIED ENDOSONOGRAPHICALLY. A SECOND AXIOS STENT (SUBJECT OF MANUFACTURER REPORT#: 3005099803-2025-04348) WAS DEPLOYED. WHEN THE PROXIMAL FLANGE WAS DEPLOYED, THE PHYSICIAN OBSERVED THE STENT RETRACT OUT OF THE STOMACH. WIRE ACCESS WAS MAINTAINED THROUGH THE MISDEPLOYED STENT INTO THE SMALL BOWEL, AND THIS WAS USED TO SUCCESSFULLY PLACE THE THIRD AXIOS STENT (SUBJECT OF THIS REPORT) COAXIALLY. TWO HOURS POST STENT-IN STENT PROCEDURE, THE PATIENT EXPERIENCED ABDOMINAL PAIN. DURING THE NIGHT, THE STENT-IN-STENT WAS DISLODGED. THE FIRST STENT WAS SEEN IN SITU. THE OUTER STENT, REFERRING TO THE SECOND STENT, WAS COMPLETELY DETACHED, AND THE INNER STENT, REFERRING TO THE THIRD STENT, WAS IN PROPER POSITION. THE PATIENT WAS SENT TO SURGERY AND SURGICAL GASTROENTEROSTOMY WAS PERFORMED, AND A PERFORATION WAS NOTED. POST SURGICAL PROCEDURE, THE PATIENT DEVELOPED AN INFECTED HEMATOMA (PERITONEAL) AND WAS DRAINED PERCUTANEOUSLY. THE HEMATOMA WAS REPORTED AS A CLINICALLY SIGNIFICANT BLEEDING EVENT WHICH RESULTED IN A TRANSFUSION. THE STENTS WERE THEN REMOVED, AND PROPHYLACTIC ANTIBIOTICS WERE ADMINISTERED, AND THE PATIENT WAS HOSPITALIZED. ON (B)(6) 2025, THE PATIENT EXPERIENCED FEVER UNDER ANTIBIOTICS. A CT ABDOMEN SCAN (COMPUTED TOMOGRAPHY OF THE ABDOMEN) WAS PERFORMED, AND REVEALED A DEVELOPMENT OF A DISTINCT, BRANCHING, HETEROGENOUS ABSCESS, LARGELY IN THE LEFT HEMI-ABDOMEN. ON (B)(6) 2025, A DIGESTIVE TRACT LEAK/PERFORATION WAS NOTED. ON (B)(6) 2025, THE PATIENT PASSED AWAY. Manufacturer Narrative: BLOCK E1: INITIAL REPORTER'S SECOND PHONE NUMBER IS (B)(6); REPORTED HERE AS PHONE NUMBER EXCEEDED CHARACTER LIMIT FOR DESIGNATED FIELD. BLOCK H6: IMDRF DEVICE CODE A010402 CAPTURES THE REPORTABLE EVENT OF STENT MIGRATION. IMDRF PATIENT CODE E2114 CAPTURES THE REPORTABLE EVENT OF PERFORATION. IMDRF PATIENT CODE E1002 CAPTURES THE REPORTABLE EVENT OF ABDOMINAL PAIN. IMDRF PATIENT CODE E0505 CAPTURES THE REPORTABLE EVENT OF HEMATOMA. IMDRF PATIENT CODE E1906 CAPTURES THE REPORTABLE EVENT OF INFECTION. IMDRF PATIENT CODE E0506 CAPTURES THE REPORTABLE EVENT OF MAJOR HEMORRHAGE. IMDRF PATIENT CODE E2401 CAPTURES THE REPORTABLE EVENT OF INJURY, NOS (NOT OTHERWISE SPECIFIED). IMDRF PATIENT CODE E172001 CAPTURES THE REPORTABLE EVENT OF ABSCESS. IMDRF PATIENT CODE E230101 CAPTURES THE REPORTABLE EVENT OF FEVER. IMDRF IMPACT CODE F08 CAPTURES THE REPORTABLE EVENT OF PATIENT WAS HOSPITALIZED. IMDRF IMPACT CODE F19 CAPTURES THE REPORTABLE EVENT OF SURGICAL GASTROENTEROSTOMY WAS PERFORMED. IMDRF IMPACT CODE F2301 CAPTURES THE REPORTABLE EVENT OF STENT-IN-STENT PROCEDURE. IMDRF IMPACT CODE F2303 CAPTURES THE REPORTABLE EVENT OF THE PROPHYLACTIC ANTIBIOTICS WERE ADMINISTERED. IMDRF IMPACT CODE F02 CAPTURES THE REPORTABLE EVENT OF THE PATIENT PASSED AWAY. IMDRF IMPACT CODE F2302 CAPTURES THE REPORTABLE EVENT OF THE TRANSFUSION. IMDRF IMPACT CODES F23 AND F22 CAPTURES THE REPORTABLE EVENT OF A CT ABDOMEN SCAN (COMPUTED TOMOGRAPHY OF THE ABDOMEN) WAS PERFORMED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22905419&pc=PNJ | 2017865-2025-98756 | 2025-08-08 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-08-27 04:00 | PNJ | AVEIR¿ LEADLESS PACEMAKER, RIGHT VENTRICLE | High impedance | Cardiac Tamponade; Cardiac Perforation; Pericardial Effusion | P150035 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED IN CLINIC FOR LEADLESS PACEMAKER (LP) IMPLANT PROCEDURE. DURING THE PROCEDURE, THE LP EXHIBITED AN LARGE INCREASE IN PACING IMPEDANCE UPON FIXATION. IT WAS NOTED THAT DURING REPOSITIONING, THE CATHETER SLIPPED FORWARD, TRAVELING BEYOND THE DEVICE. CARDIAC PERFORATION WAS NOTED AFTER INJECTION OF CONTRAST. ECHOCARDIOGRAM CONFIRMED PERCIARDIAL EFFUSION. DURING TRANSPORT, THE PATIENT DEVELOPED CARDIAC TAMPONADE AND DUE TO MEDICAL DIRECTIVE OF DNR/DNI, THE PATIENT HAD DECEASED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22905444&pc=MCX | 3004742232-2025-00252 | 2025-08-06 04:00 | Death | CARDIOVASCULAR SYSTEMS INCORPORATED | 2025-08-27 04:00 | MCX | DIAMONDBACK CORONARY | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Perforation of Vessels | P130005 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DECLINED BYPASS SURGERY AND IT WAS ACKNOWLEDGED TO BE HIGH-RISK PROCEDURE. A DIAMONDBACK 360 CORONARY ORBITAL ATHERECTOMY DEVICE (OAD) WAS USED TO TREAT A HEAVILY CALCIFIED AND HEAVILY TORTUOUS MID LATERAL ANTERIOR DESCENDING ARTERY (LAD). AN UNSPECIFIED CATHETER WAS INSERTED AT 90 DEGREES AND THE WIRE WAS EXCHANGED FOR A VIPERWIRE ADVANCE CORONARY GUIDE WIRE. THREE TREATMENTS WERE PERFORMED AND AT LEAST ONE WAS AT HIGH-SPEED. FOLLOWING REMOVAL OF THE SYSTEM, A PERFORATION WAS OBSERVED IN THE MID LAD. BALLOON TAMPONADE WAS PERFORMED, BUT PERICARDIOCENTESIS WAS OPTED NOT TO BE PERFORMED DUE TO PATIENT CONDITION. AN ATTEMPT WAS MADE TO PLACE A STENT; HOWEVER, THE STENT COULD NOT ADVANCE TO THE LESION. ANOTHER UNSPECIFIED BALLOON WAS USED TO TREAT THE AFFECTED AREA. THE PATIENT CODED, BALLOON TAMPONADE WAS ATTEMPTED AND CARDIOPULMONARY RESUSCITATION WAS PERFORMED, HOWEVER, THE PATIENT EXPIRED. IN THE PHYSICIAN'S OPINION, THE OAD CONTRIBUTED TO THE PATIENT ADVERSE EVENTS AND PATIENT DEATH. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. PRODUCTION RECORD AND CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A QUERY OF THE COMPLAINT HANDLING DATABASE FOR THE REPORTED LOT REVEALED THERE IS NO INDICATION OF A LOT SPECIFIC ISSUE. BASED ON THE INFORMATION RECEIVED, THE INVESTIGATION DETERMINED THAT THE REPORTED PATIENT PERFORATION OF VESSELS, CARDIAC ARREST AND DEATH APPEARS TO BE RELATED TO OPERATIONAL CONTEXT. IN THIS CASE IT IS POSSIBLE THAT THE REPORTED PATIENT PERFORATION OF VESSELS, CARDIAC ARREST AND DEATH ARE A RESULT OF THE PATIENT¿S DISEASE SEVERITY AND/OR USE TECHNIQUES EMPLOYED; HOWEVER, SINCE THE DEVICE WAS NOT RETURNED THIS COULD NOT BE CONFIRMED. BASED ON THE RESULTS OF THE COMPLAINT INVESTIGATION, THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO THE DESIGN, MANUFACTURE, OR LABELING OF THE DEVICE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22905554&pc=QBJ | 3004753838-2025-238390 | 2025-06-17 04:00 | Death | DEXCOM, INC. | 2025-08-27 04:00 | QBJ | DEXCOM G7 CONTINUOUS GLUCOSE MONITORING SYSTEM | Appropriate Term/Code Not Available | Insufficient Information |  |  | 1 | Event Description: IT WAS REPORTED BY THE PATIENT'S SISTER THAT THE PATIENT WITH TYPE 1 DIABETES PASSED AWAY ON (B)(6) 2025. SHE DECLINED TO DISCLOSE THE CAUSE OF DEATH. SHE REPORTED CONCERNS ABOUT CGM ACCURACY ON SOCIAL MEDIA AND REITERATED THIS CONCERN ON THE PHONE CALL WITH DEXCOM ON 08/20/2025. SHE FELT THAT THE CGM MAY HAVE CONTRIBUTED TO HIS DEATH SINCE HE WAS WEARING THE G7 IN HIS ARM AND HAD THE MOBILE DEVICE NEXT TO HIM WHEN FOUND DEAD IN BED. SHE NOTED THAT THE LAST CGM READING WAS 155 MG/DL ON (B)(6) 2025, 3 DAYS PRIOR TO THE DEATH. SHE WANTED TO KNOW WHAT HAPPENED WITH THE CGM IN THE 3 DAYS BETWEEN. SHE DECLINED TO PROVIDE ANY ADDITIONAL DETAILS. NO PRODUCT HAS BEEN PROVIDED FOR INVESTIGATION. DATA HAS BEEN REQUESTED BUT IS PENDING EVALUATION. A FOLLOW UP REPORT WILL BE SUBMITTED UPON COMPLETION. NO ADDITIONAL PATIENT OR EVENT INFORMATION IS AVAILABLE. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22905612&pc=NIQ | 2024168-2025-09140 | 2024-01-01 05:00 | Death | ABBOTT VASCULAR | 2025-08-27 04:00 | NIQ | XIENCE PRIME EVEROLIMUS ELUTING CORONARY STENT SYSTEM | Adverse Event Without Identified Device or Use Problem | Unspecified Vascular Problem | P110019 |  | 1 | Event Description: IT WAS REPORTED IN A RESEARCH ARTICLE REGARDING INTRAVASCULAR ULTRASOUND, OPTICAL COHERENCE TOMOGRAPHY AND THE EVOLUTION IN THE FIELD OF HYBRID INTRACORONARY IMAGING IT WAS NOTED THAT RESEARCH SUPPORTING THE ROLE OF STANDALONE INTRACORONARY IMAGING IN THE CLINICAL PRACTICE WITH VARIOUS TRIALS INCLUDING THE XIENCE PRIME STENT OBSERVED THE RATE OF MAJOR ADVERSE CARDIOVASCULAR EVENTS (MACES) WHICH INCLUDE (CARDIAC DEATH, STENT THROMBOSIS, TARGET LESION REVASCULARIZATION AND TARGET VESSEL REVASCULARIZATION) RATE WAS SIGNIFICANTLY LOWER IN THE IVUS-GUIDED REVASCULARIZATION GROUP COMPARED TO ANGIOGRAPHY. POSTMORTEM HISTOLOGY STUDIES HAVE DOCUMENTED THAT HYBRID IMAGING CATHETERS CAN ENABLE MORE ACCURATE CHARACTERIZATION OF PLAQUE MORPHOLOGY THAN STANDALONE IMAGING. THE USE OF MULTIMODALITY IMAGING MAY IMPROVE OUTCOMES FOLLOWING INTERVENTIONAL TREATMENT IN CHALLENGING CLINICAL SCENARIOS. TO DETERMINE THE ROLE OF HYBRID INTRACORONARY IMAGING IN GUIDING REVASCULARIZATION, CARDIOVASCULAR RISK STRATIFICATION AND PLAQUE DETECTION A LARGER SCALE OF STUDIES ARE NEEDED. SPECIFIC PATIENT INFORMATION REMAINS UNKNOWN. ADDITIONAL INFORMATION CAN BE FOUND IN THE ATTACHED ARTICLE "EMERGING HYBRID INTRACORONARY IMAGING TECHNOLOGIES AND THEIR APPLICATIONS IN CLINICAL PRACTICE AND RESEARCH." NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: THE DEVICES WERE NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORDS FOR THIS PRODUCT COULD NOT BE REVIEWED AND A SIMILAR COMPLAINT QUERY COULD NOT BE PERFORMED BECAUSE THE PART NUMBERS AND/OR LOT NUMBERS WERE NOT REPORTED, AND THE DEVICE WAS NOT RETURNED. A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECT, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED; HOWEVER, THE SUBSEQUENT TREATMENT APPEARS TO BE RELATED TO THE OPERATIONAL CONTEXT OF THE PROCEDURE. THE REPORTED PATIENT EFFECT DEATH IS LISTED IN THE XIENCE PRIME, XIENCE PRIME SMALL VESSEL (SV), AND XIENCE PRIME LONG LENGTH (LL), EVEROLIMUS ELUTING CORONARY STENT SYSTEMS INSTRUCTIONS FOR USE AS A KNOWN PATIENT EFFECT OF CORONARY STENTING PROCEDURES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING. B2, B3, D6: DATES ESTIMATED. D4: THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22905635&pc=DSQ | 2916596-2025-05585 | 2023-07-01 04:00 | Death | THORATEC CORPORATION | 2025-08-27 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. AN AUTOPSY WAS NOT PERFORMED. IT WAS UNKNOWN IF THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER: (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22906285&pc=KNS | 3005099803-2025-04353 | 2023-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-08-27 04:00 | KNS | AXIOS STENT AND ELECTROCAUTERY ENHANCED DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K233318 |  | 1 | Event Description: BOSTON SCIENTIFIC BECAME AWARE OF EVENTS INVOLVING AXIOS STENT AND ELECTROCAUTERY ENHANCED DELIVERY SYSTEMS THROUGH THE ARTICLE: "EUS-GUIDED DRAINAGE OF LARGE WALLED-OFF PANCREATIC NECROSES USING PLASTIC VERSUS LUMEN-APPOSING METAL STENTS: A SINGLE-CENTRE RANDOMISED CONTROLLED TRIAL," BY JOHN GASDAL KARSTENSEN, ET AL. PER THE ARTICLE, A SINGLE-CENTRE, OPEN-LABEL, RANDOMIZED CONTROLLED SUPERIORITY TRIAL WAS CONDUCTED USING AN ENDOSCOPIC STEP-UP APPROACH IN PATIENTS WITH PANCREATIC WALLED-OFF NECROSIS (WON) LARGER THAN 15 CM. THE STUDY INCLUDED 22 PATIENTS TREATED WITH THE TRADITIONAL DOUBLE PIGTAIL TECHNIQUE (DPT) AND 20 PATIENTS TREATED WITH LUMEN-APPOSING METAL STENTS (LAMS). THE OVERALL MORTALITY RATE AMONG PARTICIPANTS WAS 4.8%. PLEASE SEE THE REFERENCED ARTICLE FOR FULL DETAILS AND FULL LISTING OF PHYSICIANS AND HEALTHCARE FACILITIES. Manufacturer Narrative: BLOCK D4, H4: THE COMPLAINANT WAS UNABLE TO PROVIDE THE COMPLAINT DEVICE UPN AND/OR LOT NUMBER. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER PRODUCT SPECIFIC INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. LITERATURE SOURCE: KARSTENSEN, J. G.; NOVOVIC, S.; HANSEN, E. F.; JENSEN, A. B.; JOERGENSEN, H. L.; LAURITSEN, M. L.; WERGE, M. P.; SCHMIDT, P. N. "EUS-GUIDED DRAINAGE OF LARGE WALLED-OFF PANCREATIC NECROSES USING PLASTIC VERSUS LUMEN-APPOSING METAL STENTS: A SINGLE-CENTRE RANDOMISED CONTROLLED TRIAL." GUT, 72(6), 1167-1173. ARTICLE 328225. HTTPS://DOI.ORG/10.1136/GUTJNL-2022-328225. BLOCK H6: IMDRF IMPACT CODE F02 CAPTURES THE REPORTABLE EVENT OF DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22906385&pc=LWS | 2649622-2025-23466 | 2023-07-05 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-08-27 04:00 | LWS | SPRINT QUATTRO SECURE MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Fall; Sepsis; Fluid Discharge; Blister | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IS DECEASED. THE PATIENT'S SPOUSE INDICATED THAT DURING USE OF THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) SYSTEM THE PATIENT¿S LEGS BEGAN TO SEEP FLUID AND FORM BLISTERS AND WAS INSTRUCTED BY A PHYSICIAN TO ADMINISTER DIURETIC TWO TIMES A DAY. ADDITIONALLY, IT WAS STATED THAT LATER THE PATIENT EXPERIENCED A FALL AND COULD NO LONGER STAND. THE PATIENT PRESENTED TO THE HOSPITAL, PROCEDURES AND TEST WERE PERFORMED AND THE PATIENT WAS DIAGNOSED WITH AN INFECTED DEVICE, AND THE INFECTION OF SEPSIS AND STAPH HAD SPREAD THROUGHOUT THE PATIENT¿S SYSTEM. IT WAS ALSO REPORTED THAT MANY PROCEDURES INCLUDING DIALYSIS WERE PERFORMED, HOWEVER, THE PATIENT SUBSEQUENTLY EXPIRED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22907469&pc=FRN | 3000204189-2025-00001 | 2025-07-03 04:00 | Death | WALKMED TECHNOLOGIES, LLC | 2025-08-27 04:00 | FRN | WALKMED® 350VL AMBULATORY INFUSION PUMP | Pressure Problem | Pulmonary Embolism | K991275 |  | 1 | Event Description: A PATIENT UNDERGOING HOME INFUSION THERAPY USING A WALKMED 350VL AMBULATORY INFUSION PUMP (SERIAL NUMBER (B)(6)) EXPERIENCED CARDIAC ARREST ON (B)(6) 2025, WHILE RECEIVING 5-FU VIA CONTINUOUS INFUSION. THE PATIENT'S SPOUSE INITIATED CPR, AND EMERGENCY RESPONDERS PROVIDED FURTHER RESUSCITATION, BUT THE PATIENT WAS LATER PRONOUNCED DECEASED ON (B)(6) 2025 11:15 PM. THE INFUSION PUMP WAS STILL CONNECTED AT THE TIME OF THE EVENT. TIME OF DEATH DECLARED 22:15 PM (EXTENSIVE DOWNTIME) THE REPORTING FACILITY, VIA DISTRIBUTOR INFUSYSTEM, LATER INDICATED THAT THE CAUSE OF DEATH WAS DETERMINED TO BE A PULMONARY EMBOLISM AND DID NOT ATTRIBUTE THE OUTCOME TO A DEVICE MALFUNCTION. HOWEVER, DUE TO THE CIRCUMSTANCES, THE DEVICE WAS RETURNED TO WALKMED UNDER RGA #(B)(4) FOR EVALUATION. WALKMED CONDUCTED A VISUAL AND FUNCTIONAL EVALUATION PER SOP 100-601 USING FORM 100-601F6. THE PUMP FAILED PRESSURE PERFORMANCE TESTING AND EXHIBITED CLAMP BAR MALALIGNMENT, RESULTING IN IMPAIRED OCCLUSION DETECTION AND AN OUT-OF-SPEC PRESSURE PROFILE. THESE FINDINGS CONFIRM THAT THE DEVICE WAS NOT FUNCTIONING WITHIN DESIGN SPECIFICATIONS AT THE TIME OF EVALUATION. ALTHOUGH A DEFINITIVE CAUSAL LINK TO THE PATIENT'S DEATH HAS NOT BEEN ESTABLISHED, THE CONFIRMED FAILURE MODE PRESENTS A SIGNIFICANT RISK OF SERIOUS INJURY OR DELAY OF THERAPY IF THE MALFUNCTION WERE TO RECUR. IN ACCORDANCE WITH 21 CFR 803.50, THIS MEDICAL DEVICE REPORT IS SUBMITTED OUT OF AN ABUNDANCE OF CAUTION. NO AUTOPSY WAS PERFORMED. WALKMED HAS RECEIVED CONFIRMATION FROM THE FACILITY THAT THE ONLY AVAILABLE DOCUMENTATION IS THE HOSPITAL REPORT PREVIOUSLY PROVIDED. AVAILABLE CLINICAL DOCUMENTATION INDICATES THAT THE PATIENT HAD EXTENSIVE COMORBIDITIES, INCLUDING CHRONIC PULMONARY EMBOLISM, CONGESTIVE HEART FAILURE, AND PANCREATIC CANCER. WALKMED HAS REQUESTED ADDITIONAL CLINICAL DOCUMENTATION FOR ONGOING REVIEW. Manufacturer Narrative: WALKMED EVALUATED THE RETURNED 350VL INFUSION PUMP (SN (B)(6)) UNDER RGA #(B)(4) IN RESPONSE TO A REPORTED PATIENT DEATH. THE DEVICE UNDERWENT A FULL VISUAL AND FUNCTIONAL EVALUATION IN ACCORDANCE WITH SOP 100-700 AND FORM 100-601F6. THE PUMP FAILED PRESSURE PROFILE TESTING DUE TO CLAMP BAR MISALIGNMENT AND OUT-OF-SPEC PRESSURE PERFORMANCE. THIS CONDITION COULD RESULT IN IMPROPER OCCLUSION DETECTION OR ALTERED INFUSION ACCURACY, POTENTIALLY DELAYING THERAPY. NO OBSERVABLE PHYSICAL DAMAGE OR CONTAMINATION WAS FOUND. ALL OTHER FUNCTIONAL CHECKS PASSED. NOTABLY, THE DEVICE WAS ORIGINALLY MANUFACTURED AND RELEASED IN 2015, AND REVIEW OF THE DEVICE HISTORY RECORD CONFIRMED FULL CONFORMANCE TO MANUFACTURING AND QUALITY RELEASE REQUIREMENTS AT THAT TIME. NO PRODUCTION DEVIATIONS OR NON-CONFORMANCES WERE NOTED. BASED ON THE SERVICE TECHNICIAN'S ASSESSMENT AND THE AGE OF THE DEVICE, THE CLAMP BAR MISALIGNMENT AND PRESSURE PROFILE DEVIATION ARE MORE LIKELY ATTRIBUTABLE TO WEAR OVER TIME RATHER THAN A LATENT MANUFACTURING DEFECT. THE PUMP HAD NO RECORD OF PRIOR WALKMED SERVICE. WHILE THERE IS NO DIRECT EVIDENCE LINKING THE MALFUNCTION TO THE PATIENT'S DEATH, THE CONFIRMED OUT-OF-SPEC PERFORMANCE CONSTITUTES A POTENTIAL RISK OF SERIOUS INJURY OR DELAY OF THERAPY. THIS MDR IS SUBMITTED IN ACCORDANCE WITH 21 CFR 803.50.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22907642&pc=NTE | 2124215-2025-58136 | 2017-09-01 04:00 | Death | SILK ROAD MEDICAL | 2025-08-27 04:00 | NTE | ENROUTE TRANSCAROTID NEUROPROTECTION SYSTEM | Adverse Event Without Identified Device or Use Problem | Pulmonary Embolism; Stroke/CVA |  |  | 1 | Event Description: IT WAS REPORTED VIA LITERATURE THAT PATIENTS WHO UNDERWENT TRANSCAROTID ARTERY REVASCULARIZATION (TCAR) DIED DURING THE 30-DAY PERIOPERATIVE PERIOD. THE AIM OF THE STUDY WAS TO EVALUATE THE OUTCOMES OF TCAR. (B)(6) 2017 TO (B)(6) 2021, 429 TCAR PROCEDURES WERE PERFORMED. DURING THE 30-DAY PERIOPERATIVE PERIOD, THREE PATIENTS DIED FROM STROKE-RELATED CAUSES. TWO ADDITIONAL DEATHS WERE ATTRIBUTED TO CARDIOPULMONARY EVENTS SECONDARY TO ASPIRATION AND LIKELY PULMONARY EMBOLUS, ALSO WITHIN THE 30-DAY PERIOPERATIVE PERIOD. NO FURTHER INFORMATION WAS PROVIDED TO BOSTON SCIENTIFIC. Manufacturer Narrative: A2: THE MEAN AGE OF PATIENTS WAS 74.3 YEARS, WITH A STANDARD DEVIATION OF 8.9 YEARS. A3: WOMEN COMPRISED 45.5% OF THE STUDY POPULATION. B2: THE EXACT DATE OF DEATH WAS NOT PROVIDED FOR ANY OF THE REPORTED CASES AND WAS ESTIMATED BASED ON THE EARLIEST KNOWN TCAR PROCEDURE IN THE COHORT. B3: THE EVENT DATE WAS ESTIMATED TO THE EARLIEST KNOWN TCAR PROCEDURE INCLUDED IN THE COHORT. DETAILS OF THE EVENT, INCLUDING PRODUCT INFORMATION, WERE NOT PROVIDED TO BSC. BECAUSE THE PRODUCT LOT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) NUMBER AND OTHER SPECIFIC PRODUCT INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. HUSMAN, R., TANAKA, A., HARLIN, S. A., MARTIN, G. H., SAQIB, N. U., KEYHANI, A., KEYHANI, K., & WANG, S. K. (2022). RESULTS ASSOCIATED WITH THE HEALTH SYSTEM-WIDE ADOPTION OF TRANSCAROTID REVASCULARIZATION. JOURNAL OF VASCULAR SURGERY, 76(4), 967-972. HTTPS://DOI.ORG/10.1016/J.JVS.2022.04.028.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22908296&pc=DQY | 2032493-2025-90435 | 2025-07-30 04:00 | Death | MICROVENTION, INC. | 2025-08-27 04:00 | DQY | VIA MICROCATHETER | Adverse Event Without Identified Device or Use Problem | Ruptured Aneurysm | K192135 |  | 1 | Event Description: IT WAS REPORTED TO MICROVENTION THAT THE PATIENT WAS INTUBATED AND UNDER GENERAL ANESTHESIA, NOT UNDER CONSCIOUS SEDATION, FOR AN OFF-LABEL USE OF THE WEB IN A PARAOPHTHALMIC ANEURYSM. THE PATIENT WAS REPORTED TO HAVE COUGHED OVER THEIR ET TUBE. THIS REPORTEDLY CAUSED THE SYSTEM OF CATHETERS USED FOR TREATMENT, INCLUDING THE SOFIA, VIA, AND ¿OTHER DEVICES¿, TO MOVE UNEXPECTEDLY. THESE DEVICES WERE IN THE PATIENT WHEN IT WAS THEN NOTICED THAT THE VESSEL WAS PERFORATED. A WEB DEVICE WAS THEN INSERTED AFTER THE PERFORATION AND REPORTEDLY DID NOT INITIALLY DETACH. THE 90 DEGREE VIA WAS BROUGHT UP TO BRACE AGAINST THE WEB TO DETACH AND THE WEB CAME TO REST IN THE PARENT VESSEL. INTERVENTION WITH SNARE AND TWO OTHER DEVICES WERE USED IN AN ATTEMPT TO RETRIEVE THE WEB, WHICH PROVED UNSUCCESSFUL. THE CASE WAS TERMINATED DUE TO THE VESSEL WALL BLEEDING, AND THE WEB REMAINED IN THE PARENT VESSEL. THE PATIENT PASSED AWAY THE NEXT DAY. THIS WEB DEVICE USED IN THIS CASE WAS SUBMITTED UNDER MDR 2032493-2025-90428. Manufacturer Narrative: BASED ON A REVIEW OF THE DEVICE¿S RISK DOCUMENTATION, THE REPORTED EVENT DID NOT INDICATE THERE WERE THE PRESENCE OF ANY POTENTIAL OR NEW MANUFACTURING, DESIGN, QUALITY, OR OTHER SYSTEMIC ISSUES, OR NON-CONFORMANCES. THE COMPLAINT CODE IS MONITORED THROUGH THE TRENDING PROCESS; CORRECTIVE ACTION IS DETERMINED, AS NEEDED, THROUGH THIS PROCESS. INVESTIGATIONS OF HISTORIC COMPLAINT FILES WITH SIMILAR COMPLAINT CATEGORY CODING ARE RECORDED IN THE COMPLAINT HANDLING SYSTEM; WITHOUT THE ABILITY TO PERFORM AN ANALYSIS OF THE DEVICE, THIS INVESTIGATION CANNOT IDENTIFY WITH CERTAINTY ANY POTENTIAL ROOT CAUSES. BATCH REVIEW: A SEARCH FOR NON-CONFORMANCES ASSOCIATED WITH THIS PART/LOT NUMBER COMBINATION DID NOT REVEAL ANY PRODUCTION-RELATED ISSUES RELEVANT TO THE COMPLAINT THAT OCCURRED DURING MANUFACTURING OF THE DEVICE. COMPLAINT SYSTEM REVIEW: THERE ARE NO SIMILAR COMPLAINTS BASED ON THE COMPLAINT CATEGORY REGARDING THIS BATCH NUMBER FROM THE LAST TWO YEARS RECORDED IN THE COMPLAINT SYSTEM AT THE TIME OF THIS INVESTIGATION. IFU REVIEW (ADDITIONAL INFORMATION CAN BE FOUND IN THE IFU): POTENTIAL COMPLICATIONS POTENTIAL COMPLICATIONS INCLUDE BUT ARE NOT LIMITED TO THE FOLLOWING: HEMATOMA AT THE SITE OF ENTRY, ANEURYSM RUPTURE, EMBOLI, VESSEL PERFORATION, PARENT ARTERY OCCLUSION, HEMORRHAGE, ISCHEMIA, VASOSPASM, VASCULAR THROMBOSIS AND NEUROLOGICAL DEFICITS INCLUDING STROKE AND DEATH. POTENTIAL X-RAY RADIATION EXPOSURE RELATED ADVERSE EVENTS INCLUDE BUT ARE NOT LIMITED TO: ALOPECIA, BURNS RANGING IN SEVERITY FROM SKIN REDDENING TO ULCERS, CATARACTS, DELAYED NEOPLASIA, TISSUE NECROSIS, AND RISKS ASSOCIATED WITH CONTRAST DYE. WARNINGS ¿ NEVER ADVANCE OR WITHDRAW A DEVICE AGAINST RESISTANCE UNTIL THE CAUSE OF THE RESISTANCE IS DETERMINED BY FLUOROSCOPY. EXCESSIVE FORCE AGAINST RESISTANCE MAY RESULT IN DAMAGE TO THE DEVICE OR VESSEL PERFORATION. ¿ WHEN INJECTING CONTRAST FOR ANGIOGRAPHY, ENSURE THE CATHETER IS NOT KINKED OR OCCLUDED. ¿ DO NOT EXCEED 300 PSI MAXIMUM RECOMMENDED INFUSION PRESSURE. EXCESS PRESSURE MAY RESULT IN CATHETER DAMAGE OR PATIENT INJURY. ¿ STEAM SHAPING MAY RESULT IN IMPROPER DEVICE DELIVERY AND DEPLOYMENT, DEPENDING ON THE DEGREE OF SHAPING AND CATHETER DEFLECTION DURING DEVICE DELIVERY. ¿ USING THE VIA MICROCATHETER WITH GUIDE CATHETERS SMALLER THAN WHAT IS RECOMMENDED MAY RESULT IN DAMAGING THE HYDROPHILIC COATING. PRECAUTIONS ¿ PRIOR TO USE, EXAMINE THE VIA MICROCATHETER TO VERIFY THAT IT HAS NOT BEEN DAMAGED DURING SHIPMENT. ¿ THE VIA MICROCATHETER HAS A LUBRICIOUS HYDROPHILIC COATING ON THE OUTSIDE OF THE CATHETER. IT MUST BE KEPT HYDRATED IN ORDER TO BE LUBRICIOUS. THIS CAN BE ACCOMPLISHED BY ATTACHING A Y-CONNECTOR TO A CONTINUOUS SALINE DRIP. ¿ HIGH QUALITY, DIGITAL SUBTRACTION FLUOROSCOPIC ROAD MAPPING, WITH ORTHOGONAL VIEWS, IS MANDATORY TO ACHIEVE CORRECT PLACEMENT OF THE MICROCATHETER AND EMBOLIZATION DEVICE. ¿ IF REPOSITIONING IS REQUIRED, TAKE SPECIAL CARE TO RETRACT OR TO ADVANCE THE DEVICE UNDER FLUOROSCOPY ¿ THE OPERATOR SHOULD BE AWARE THAT MICROCATHETERS, IN DISTAL BLOOD VESSELS, MAY INCREASE THE RISK OF THROMBOEMBOLI. ¿ IF REMOVED FROM THE PATIENT, THE HYDROPHILIC COATING ON THE VIA MICROCATHETER SHOULD BE HYDRATED WITH HEPARINIZED SALINE. DO NOT ALLOW THE COATING TO DRY AS THIS MAY IMPACT THE COATING SAFETY AND PERFORMANCE. ¿ AVOID PRE-SOAKING DEVICES FOR LONG DURATIONS WHEN THE DEVICE IS NOT IN USE AS THIS MAY IMPACT THE COATING SAFETY AND PERFORMANCE. ¿ AVOID WIPING THE DEVICE WITH DRY GAUZE AS THIS MAY DAMAGE THE DEVICE COATING. AVOID EXCESSIVE WIPING OF THE COATED DEVICE. PROCEDURE CATHETERIZATION OF THE LESION 5. IF DESIRED, CAREFULLY INTRODUCE THE MICROCATHETER THROUGH THE INTRODUCER SHEATH. 6. PREPARE AN APPROPRIATELY SIZED GUIDEWIRE AND INSERT INTO THE MICROCATHETER PER THE MANUFACTURER¿S INSTRUCTIONS. 7. INTRODUCE THE GUIDEWIRE AND MICROCATHETER AS A UNIT INTO THE GUIDING CATHETER UNTIL THE DISTAL TIP OF THE GUIDE CATHETER HAS BEEN REACHED. ALTERNATIVELY ADVANCE THE GUIDEWIRE AND MICROCATHETER UNTIL THE DESIRED SITE HAS BEEN REACHED. VERIFY CATHETER POSITION USING FLUOROSCOPY. GENTLY TIGHTEN RHV, AS NEEDED, WITHOUT CRUSHING THE MICROCATHETER. 8. PEEL AWAY INTRODUCER SHEATH BY PULLING ON THE TAB, IF USED. 9. AFTER THE MICROCATHETER HAS BEEN POSITIONED INSIDE THE LESION, REMOVE THE GUIDEWIRE. 10. FLUSH THE ID OF THE MICROCATHETER WITH STERILE FLUSH SOLUTION BY ATTACHING A SYRINGE TO THE CATHETER HUB 11. ATTACH A SECOND RHV TO THE HUB OF THE MICROCATHETER. ATTACH A ONE-WAY STOPCOCK TO THE SIDEARM OF THE SECOND RHV AND CONNECT THE FLUSH SOLUTION LINE TO THE STOPCOCK. 12. OPEN THE STOPCOCK TO ALLOW FLUSH THROUGH THE MICROCATHETER WITH STERILE FLUSH SOLUTION. REMOVAL OF THE VIA MICROCATHETER 13. UNDER FLUOROSCOPIC GUIDANCE, WITHDRAW THE VIA MICROCATHETER UNTIL THE ENTIRE DEVICE HAS BEEN REMOVED FROM THE PATIENT. THE PHYSICAL DEVICE WAS NOT AVAILABLE FOR EVALUATION TO INVESTIGATE IF A CONDITION EXISTED THAT WOULD HAVE CONTRIBUTED TO EVENT. SUPPLEMENTAL IMAGING WAS ALSO NOT PROVIDED FOR REVIEW; WITHOUT IMAGING, THE INVESTIGATION CANNOT VERIFY THE EVENT OCCURRED AS DESCRIBED, NOR COULD THE INVESTIGATION FURTHER EXAMINE THE CAUSE OF THE REPORTED EVENT. MICROVENTION IS SUBMITTING THIS REPORT TO COMPLY WITH FDA REPORTING REGULATIONS UNDER 21 CFR PARTS 4 AND 803. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MICROVENTION, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MICROVENTION HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION AND HAS PROVIDED AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MICROVENTION, OR ITS EMPLOYEES THAT THE DEVICE, MICROVENTION, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. A GOOD FAITH EFFORT WILL BE MADE TO OBTAIN THE APPLICABLE INFORMATION RELEVANT TO THE REPORT. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22908891&pc=LYJ | 1644487-2025-10550 | 2025-06-24 04:00 | Death | LIVANOVA USA, INC. | 2025-08-27 04:00 | LYJ | PULSE GEN MODEL 1000 | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P970003 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY IN THEIR SLEEP. NO OTHER RELEVANT INFORMATION HAS BEEN RECEIVED TO DATE. Manufacturer Narrative: LIVANOVA USA, INC. SUBMITS THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION, BASED ON INFORMATION THAT LIVANOVA HAS OBTAINED, BUT MAY NOT HAVE BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION, OR A CONCLUSION BY FDA OR ANYONE ELSE, THAT THE DEVICE, LIVANOVA, OR LIVANOVA'S EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS¿ OR "MALFUNCTIONS¿. THESE WORDS ARE INCORPORATED INTO THE FDA 3500A MEDWATCH FORM BY THE FDA, AND LIVANOVA OBJECTS TO THEIR USE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22909046&pc=NIQ | 2024168-2025-09152 | 2020-12-17 05:00 | Death | ABBOTT VASCULAR | 2025-08-27 04:00 | NIQ | XIENCE EVEROLIMUS ELUTING CORONARY STENT SYSTEM | Adverse Event Without Identified Device or Use Problem | Unspecified Vascular Problem | P070015 |  | 1 | Event Description: IT WAS REPORTED THROUGH THIS RESEARCH ARTICLE THAT THE LONG-TERM SAFETY AND EFFICACY IN THE SETTING OF ACUTE CORONARY SYNDROMES (ACS) ARE NOT KNOWN THEREFORE, THIS STUDY COMPARES THE LONG-ER OUTCOME IN ACS VERSUS CHRONIC CORONARY SYNDROMES (CCS) WITH BIODEGRADABLE POLYMER SIROLIMUS-ELUTING NON-ABBOTT STENT AND THE DURABLE POLYMER (DP) EVEROLIMUS-ELUTING XIENCE STENT. THE TARGET AC STUDY CONTAINED OF 1,653 PATIENTS (728 WITH ACS AND 922 WITH CCS) WITH EITHER THE XIENCE OR NON-ABBOTT STENT RANDOMLY CHOSEN TO AS THE STENT IMPLANT. HOWEVER, BOTH GROUPS ACS AND CCS CLINICAL OUTCOMES NOTED PATIENTS SUFFERED (CARDIAC DEATH, TARGET VESSEL MYOCARDIAL INFARCTION, ISCHEMIA-DRIVEN TARGET LESION REVASCULARIZATION, STENT THROMBOSIS (ACUTE, LATE, VERY LATE). IN CONCLUSION, IN THE TARGET AC TRIAL,1 IN 3 PATIENT HAD A MAJOR ADVERSE EVENT AT 5 YEARS, IRRESPECTIVE OF CCS OR ACS PRESENTATION. LONG-TERM, THE BP SIROLIMUS-ELUTING NON-ABBOTT STENT WAS AS SAFE AND EFFECTIVE AS THE DP EVEROLIMUS-ELUTING XIENCE STENT ACROSS THE SPECTRUM OF CLINICAL PRESENTATIONS. SPECIFIC PATIENT INFORMATION IS DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE, TITLED "LONG-TERM PERCUTANEOUS CORONARY INTERVENTION OUTCOMES IN CHRONIC VERSUS ACUTE CORONARY SYNDROMES (TARGET ALL COMERS TRIAL)". NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: THE DEVICES WERE NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORDS FOR THIS PRODUCT COULD NOT BE REVIEWED AND A SIMILAR COMPLAINT QUERY COULD NOT BE PERFORMED BECAUSE THE PART NUMBERS AND/OR LOT NUMBERS WERE NOT REPORTED, AND THE DEVICE WAS NOT RETURNED. A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECT, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED; HOWEVER, THE SUBSEQUENT TREATMENT APPEARS TO BE RELATED TO THE OPERATIONAL CONTEXT OF THE PROCEDURE. THE REPORTED PATIENT EFFECT OF DEATH ARE LISTED IN THE XIENCE V EVEROLIMUS ELUTING CORONARY STENT SYSTEMS INSTRUCTIONS FOR USE AS A KNOWN PATIENT EFFECT OF CORONARY STENTING PROCEDURES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING. B2, B3, D6A: DATES ESTIMATED .. ATTACHMENT: ARTICLE TITLED: "LONG-TERM PERCUTANEOUS CORONARY INTERVENTION OUTCOMES IN CHRONIC VERSUS ACUTE CORONARY SYNDROMES (TARGET ALL COMERS TRIAL)".
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22910006&pc=OUT | 2029214-2025-01912 | 2025-04-22 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-08-27 04:00 | OUT | PIPELINE | Adverse Event Without Identified Device or Use Problem | Aneurysm |  |  | 1 | Event Description: WANG, C., ZHOU, Y., ZHANG, Y., WEI, D., TURHON, M., LIU, J., ZHANG, Y., WANG, K., ZHANG, H., LI, T., MAIMAITILI, A., MAO, G., SONG, D., WANG, Y., FENG, W., WANG, Y., SHI, H., WAN, J., LIU, J; FRONTIERS IN NEUROLOGY; 2025; 16, 1561965; DELAYED ISCHEMIC STROKE AFTER PED PLACEMENT FOR ANEURYSMS: OPTIMAL DURATION OF DUAL ANTIPLATELET THERAPY AND RISK FACTORS; DOI.ORG/10.3389/FNEUR.2025.1561965 LITERATURE WAS REVIEWED REGARDING A STUDY AIMING TO INVESTIGATE THE INCIDENCE AND POTENTIAL RISK FACTORS OF DIS IN A LARGE COHORT OF PATIENTS IN CHINA TREATED WITH PED FOR INTRACRANIAL ANEURYSMS. ADDITIONALLY, THE STUDY EXPLORES THE ASSOCIATION BETWEEN THE DURATION OF DUAL ANTIPLATELET THERAPY (DAPT) AND THE OCCURRENCE OF DELAYED ISCHEMIC STROKE (DIS). THE FOLLOWING MEDTRONIC DEVICES WERE USED: PIPELINE EMBOLIZATION DEVICE (EITHER CLASSIC OR FLEX), MARKSMAN CATHETER, AND PHENOM 27 CATHETER. THE TIMEFRAME OF THE STUDY WAS BETWEEN (B)(6) 2014 AND (B)(6) 2019. 15 DEATHS OCCURRED IN THE STUDY POPULATION WITHIN ON MONTH; HOWEVER, THERE WAS NO STATEMENT ESTABLISHING A CAUSAL OR CONTRIBUTORY RELATIONSHIP BETWEEN MEDTRONIC PRODUCT AND THE DEATHS. THE CAUSES OF DEATH WERE NOT REPORTED AMONG ALL MEDTRONIC DEVICES PATIENTS¿ ADVERSE EVENTS / DEVICE PRODUCT PERFORMANCE ISSUES INCLUDED: 12 PATIENTS EXPERIENCED DELAYED ISCHEMIC STROKE (DIS). OF WHICH FIVE PATIENTS SUFFERED TRANSIENT ISCHEMIC ATTACKS, AND SEVEN EXPERIENCED INFARCTIONS. IMAGING REVEALED THAT THE EMBOLIZED AREA CORRESPONDED TO THE LOCATION WHERE THE PED HAD BEEN DEPLOYED, WITH NO ADDITIONAL INFARCTS DETECTED. TWO PATIENTS EXPERIENCED STROKES DUE TO STENOSIS OR OCCLUSION OF THE PARENT ARTERY COVERED BY THE PED. ONE PATIENT UNDERWENT ASPIRATION THROMBECTOMY AND BYPASS SURGERY. THE OTHER PATIENTS RECEIVED MEDICAL TREATMENT AND REHABILITATION, WITH ONLY ONE PATIENT EXHIBITING POOR NEUROLOGICAL FUNCTION (MODIFIED RANKIN SCALE SCORE = 3) AT THE LAST FOLLOW-UP. ONE PATIENT EXPERIENCED SUDDEN HEADACHES ACCOMPANIED BY PROJECTILE VOMITING, FACIAL ASYMMETRY, AND RIGHT-SIDED HEMIPLEGIA, WITH RIGHT CEREBRAL INFARCTION BY CT 10 MONTHS AFTER PROCEDURE. NO MAJOR BLEEDING EVENTS, INCLUDING INTRACRANIAL OR VISCERAL HEMORRHAGES, WERE OBSERVED DURING THE FOLLOW-UP PERIOD. NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. Manufacturer Narrative: CITATION: WANG, C., ZHOU, Y., ZHANG, Y., WEI, D., TURHON, M., LIU, J., ZHANG, Y., WANG, K., ZHANG, H., LI, T., MAIMAITILI, A., MAO, G., SONG, D., WANG, Y., FENG, W., WANG, Y., SHI, H., WAN, J., LIU, J. DELAYED ISCHEMIC STROKE AFTER PED PLACEMENT FOR ANEURYSMS: OPTIMAL DURATION OF DUAL ANTIPLATELET THERAPY AND RISK FACTORS. FRONTIERS IN NEUROLOGY 16, 1561965 2025. DOI.ORG/10.3389/FNEUR.2025.1561965 EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT. NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. A2 AGE AND A3A SEX ARE AVERAGE OF COHORT. SEE ATTACHMENTS FOR LITERATURE ARTICLE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22910014&pc=KRA | 2029214-2025-01914 | 2025-04-22 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-08-27 04:00 | KRA | MARKSMAN | Adverse Event Without Identified Device or Use Problem | Aneurysm |  |  | 1 | Event Description: WANG, C., ZHOU, Y., ZHANG, Y., WEI, D., TURHON, M., LIU, J., ZHANG, Y., WANG, K., ZHANG, H., LI, T., MAIMAITILI, A., MAO, G., SONG, D., WANG, Y., FENG, W., WANG, Y., SHI, H., WAN, J., LIU, J; FRONTIERS IN NEUROLOGY; 2025; 16, 1561965; DELAYED ISCHEMIC STROKE AFTER PED PLACEMENT FOR ANEURYSMS: OPTIMAL DURATION OF DUAL ANTIPLATELET THERAPY AND RISK FACTORS; DOI.ORG/10.3389/FNEUR.2025.1561965 LITERATURE WAS REVIEWED REGARDING A STUDY AIMING TO INVESTIGATE THE INCIDENCE AND POTENTIAL RISK FACTORS OF DIS IN A LARGE COHORT OF PATIENTS IN CHINA TREATED WITH PED FOR INTRACRANIAL ANEURYSMS. ADDITIONALLY, THE STUDY EXPLORES THE ASSOCIATION BETWEEN THE DURATION OF DUAL ANTIPLATELET THERAPY (DAPT) AND THE OCCURRENCE OF DELAYED ISCHEMIC STROKE (DIS). THE FOLLOWING MEDTRONIC DEVICES WERE USED: PIPELINE EMBOLIZATION DEVICE (EITHER CLASSIC OR FLEX), MARKSMAN CATHETER, AND PHENOM 27 CATHETER. THE TIMEFRAME OF THE STUDY WAS BETWEEN NOVEMBER 2014 AND OCTOBER 2019. 15 DEATHS OCCURRED IN THE STUDY POPULATION WITHIN ON MONTH; HOWEVER, THERE WAS NO STATEMENT ESTABLISHING A CAUSAL OR CONTRIBUTORY RE LATIONSHIP BETWEEN MEDTRONIC PRODUCT AND THE DEATHS. THE CAUSES OF DEATH WERE NOT REPORTED AMONG ALL MEDTRONIC DEVICES PATIENTS¿ ADVERSE EVENTS / DEVICE PRODUCT PERFORMANCE ISSUES INCLUDED: 12 PATIENTS EXPERIENCED DELAYED ISCHEMIC STROKE (DIS). OF WHICH FIVE PATIENTS SUFFERED TRANSIENT ISCHEMIC ATTACKS, AND SEVEN EXPERIENCED INFARCTIONS. IMAGING REVEALED THAT THE EMBOLIZED AREA CORRESPONDED TO THE LOCATION WHERE THE PED HAD BEEN DEPLOYED, WITH NO ADDITIONAL INFARCTS DETECTED. TWO PATIENTS EXPERIENCED STROKES DUE TO STENOSIS OR OCCLUSION OF THE PARENT ARTERY COVERED BY THE PED. ONE PATIENT UNDERWENT ASPIRATION THROMBECTOMY AND BYPASS SURGERY. THE OTHER PATIENTS RECEIVED MEDICAL TREATMENT AND REHABILITATION, WITH ONLY ONE PATIENT EXHIBITING POOR NEUROLOGICAL FUNCTION (MODIFIED RANKIN SCALE SCORE = 3) AT THE LAST FOLLOW-UP. ONE PATIENT EXPERIENCED SUDDEN HEADACHES ACCOMPANIED BY PROJECTILE VOMITING, FACIAL ASYMMETRY, AND RIGHT-SIDED HEMIPLEGIA, WITH RIGHT CEREBRAL INFARCTION BY CT 10 MONTHS AFTER PROCEDURE. NO MAJOR BLEEDING EVENTS, INCLUDING INTRACRANIAL OR VISCERAL HEMORRHAGES, WERE OBSERVED DURING THE FOLLOW-UP PERIOD. NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. Manufacturer Narrative: CITATION: WANG, C., ZHOU, Y., ZHANG, Y., WEI, D., TURHON, M., LIU, J., ZHANG, Y., WANG, K., ZHANG, H., LI, T., MAIMAITILI, A., MAO, G., SONG, D., WANG, Y., FENG, W., WANG, Y., SHI, H., WAN, J., LIU, J. DELAYED ISCHEMIC STROKE AFTER PED PLACEMENT FOR ANEURYSMS: OPTIMAL DURATION OF DUAL ANTIPLATELET THERAPY AND RISK FACTORS. FRONTIERS IN NEUROLOGY 16, 1561965 2025. DOI.ORG/10.3389/FNEUR.2025.1561965 EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT. NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. A2 AGE AND A3A SEX ARE AVERAGE OF COHORT. SEE ATTACHMENTS FOR LITERATURE ARTICLE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22910036&pc=QJP | 2029214-2025-01916 | 2025-04-22 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-08-27 04:00 | QJP | PHENOM CATHETER | Adverse Event Without Identified Device or Use Problem | Aneurysm |  |  | 1 | Event Description: WANG, C., ZHOU, Y., ZHANG, Y., WEI, D., TURHON, M., LIU, J., ZHANG, Y., WANG, K., ZHANG, H., LI, T., MAIMAITILI, A., MAO, G., SONG, D., WANG, Y., FENG, W., WANG, Y., SHI, H., WAN, J., LIU, J; FRONTIERS IN NEUROLOGY; 2025; 16, 1561965; DELAYED ISCHEMIC STROKE AFTER PED PLACEMENT FOR ANEURYSMS: OPTIMAL DURATION OF DUAL ANTIPLATELET THERAPY AND RISK FACTORS; DOI.ORG/10.3389/FNEUR.2025.1561965 LITERATURE WAS REVIEWED REGARDING A STUDY AIMING TO INVESTIGATE THE INCIDENCE AND POTENTIAL RISK FACTORS OF DIS IN A LARGE COHORT OF PATIENTS IN CHINA TREATED WITH PED FOR INTRACRANIAL ANEURYSMS. ADDITIONALLY, THE STUDY EXPLORES THE ASSOCIATION BETWEEN THE DURATION OF DUAL ANTIPLATELET THERAPY (DAPT) AND THE OCCURRENCE OF DELAYED ISCHEMIC STROKE (DIS). THE FOLLOWING MEDTRONIC DEVICES WERE USED: PIPELINE EMBOLIZATION DEVICE (EITHER CLASSIC OR FLEX), MARKSMAN CATHETER, AND PHENOM 27 CATHETER. THE TIMEFRAME OF THE STUDY WAS BETWEEN NOVEMBER 2014 AND OCTOBER 2019. 15 DEATHS OCCURRED IN THE STUDY POPULATION WITHIN ON MONTH; HOWEVER, THERE WAS NO STATEMENT ESTABLISHING A CAUSAL OR CONTRIBUTORY RE LATIONSHIP BETWEEN MEDTRONIC PRODUCT AND THE DEATHS. THE CAUSES OF DEATH WERE NOT REPORTED AMONG ALL MEDTRONIC DEVICES PATIENTS¿ ADVERSE EVENTS / DEVICE PRODUCT PERFORMANCE ISSUES INCLUDED: 12 PATIENTS EXPERIENCED DELAYED ISCHEMIC STROKE (DIS). OF WHICH FIVE PATIENTS SUFFERED TRANSIENT ISCHEMIC ATTACKS, AND SEVEN EXPERIENCED INFARCTIONS. IMAGING REVEALED THAT THE EMBOLIZED AREA CORRESPONDED TO THE LOCATION WHERE THE PED HAD BEEN DEPLOYED, WITH NO ADDITIONAL INFARCTS DETECTED. TWO PATIENTS EXPERIENCED STROKES DUE TO STENOSIS OR OCCLUSION OF THE PARENT ARTERY COVERED BY THE PED. ONE PATIENT UNDERWENT ASPIRATION THROMBECTOMY AND BYPASS SURGERY. THE OTHER PATIENTS RECEIVED MEDICAL TREATMENT AND REHABILITATION, WITH ONLY ONE PATIENT EXHIBITING POOR NEUROLOGICAL FUNCTION (MODIFIED RANKIN SCALE SCORE = 3) AT THE LAST FOLLOW-UP. ONE PATIENT EXPERIENCED SUDDEN HEADACHES ACCOMPANIED BY PROJECTILE VOMITING, FACIAL ASYMMETRY, AND RIGHT-SIDED HEMIPLEGIA, WITH RIGHT CEREBRAL INFARCTION BY CT 10 MONTHS AFTER PROCEDURE. NO MAJOR BLEEDING EVENTS, INCLUDING INTRACRANIAL OR VISCERAL HEMORRHAGES, WERE OBSERVED DURING THE FOLLOW-UP PERIOD. NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. Manufacturer Narrative: CITATION: WANG, C., ZHOU, Y., ZHANG, Y., WEI, D., TURHON, M., LIU, J., ZHANG, Y., WANG, K., ZHANG, H., LI, T., MAIMAITILI, A., MAO, G., SONG, D., WANG, Y., FENG, W., WANG, Y., SHI, H., WAN, J., LIU, J. DELAYED ISCHEMIC STROKE AFTER PED PLACEMENT FOR ANEURYSMS: OPTIMAL DURATION OF DUAL ANTIPLATELET THERAPY AND RISK FACTORS. FRONTIERS IN NEUROLOGY 16, 1561965 2025. DOI.ORG/10.3389/FNEUR.2025.1561965 EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT. NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. A2 AGE AND A3A SEX ARE AVERAGE OF COHORT. SEE ATTACHMENTS FOR LITERATURE ARTICLE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22893301&pc=OZD | 1220648-2025-45925 | 2025-07-24 04:00 | Death | ABIOMED, INC. | 2025-08-26 04:00 | OZD | IMPELLA 5.5 | Failure to Advance | Low Blood Pressure/ Hypotension; Ventricular Fibrillation | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT IN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS BEING IMPLANTED WITH AN IMPELLA 5.5 DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. THE IMPELLA 5.5 WAS PREPPED IN THE OPERATING ROOM (OR) BUT WAS UNABLE TO ADVANCE PAST BRACHIOCEPHALIC TRUNK DUE TO CALCIFIED STENOSIS. THE PATIENT CONTINUED TO DECOMPENSATE AND AN IMPELLA CP WAS PREPPED AND URGENTLY IMPLANTED INSTEAD OF THE IMPELLA 5.5. THE IMPELLA 5.5 WAS SAVED IN A STERILE ENVIRONMENT FOR POSSIBLE ESCALATION AND THE PATIENT WAS SENT TO THE CARDIAC CATHETERIZATION LABORATORY (CCL) FOR A POSSIBLE PERCUTANEOUS CORONARY INTERVENTION (PCI). IN THE CCL, THE PATIENT WAS FOUND TO HAVE TRIPLE VESSEL DISEASE, AND THE MEDICAL TEAM WAS UNABLE TO PERFORM THE PCI. THE PATIENT STABILIZED WITH THE IMPELLA CP BUT REMAINED IN CARDIOGENIC SHOCK AND THE PATIENT¿S LACTATE REMAINED ELEVATED. THE MEDICAL TEAM DECIDED TO ESCALATE THE PATIENT TO THE IMPELLA 5.5 FOR INCREASED SUPPORT AND THE PATIENT WAS TAKEN BACK TO THE OR FOR IMPELLA 5.5 INSERTION THROUGH THE AXILLARY ARTERY. THE IMPELLA 5.5 WAS SUCCESSFULLY PLACED AND SUPPORT INITIATED. THE PATIENT WAS TAKEN BACK TO THE ICU FOR RECOVERY AND MONITORING. THE PATIENT WAS NOTED TO HAVE BEEN TRENDING POSITIVELY WITH LACTATE RESULTS DOWN TRENDING, HOWEVER THE PATIENT ACUTELY DECOMPENSATED WITH NO PULSATILITY. THE PATIENT CONTINUED TO DECOMPENSATE WITH PROFOUND HYPOTENSION AND EVENTUALLY WENT INTO VENTRICULAR FIBRILLATION ARREST AND ADVANCED CARDIAC LIFE SUPPORT WAS INITIATED AS A RESULT. THE MEDICAL TEAM WAS NOT CERTAIN OF THE CAUSE OF THE ACUTE DECOMPENSATION, BUT SUSPECTED THE CAUSE WAS DUE TO QUICKLY DECLINING RIGHT VENTRICULAR FAILURE. Manufacturer Narrative: ALTHOUGH THE PATIENT'S PRESENTING CONDITION MAY BE MORE LIKELY HAVE IMPACTED THE EVENT THE IMPELLA CANNOT BE EXCLUDED AS A POSSIBLE CONTRIBUTING FACTOR IN THE PATIENT'S DEMISE. DEVICE STATUS: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22893655&pc=OZD | 1220648-2025-45926 | 2025-07-29 04:00 | Death | ABIOMED, INC. | 2025-08-26 04:00 | OZD | IMPELLA CP | Restricted Flow rate | Cardiac Arrest; Pericardial Effusion | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT IN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT DURING A PERCUTANEOUS CARDIAC INTERVENTION (PCI). IT WAS REPORTED THAT DURING IMPELLA CP SUPPORT, THE PUMP EXPERIENCED SUCTION EVENTS THAT WERE RESOLVED BY GIVING THE PATIENT INTRAVENOUS FLUIDS. AFTER THE PCI, THE PATIENT RAPIDLY DECOMPENSATED, REQUIRING INCREASED VASOPRESSORS, THE PATIENT WAS NOTED TO BE HYPOXIC ON ARTERIAL BLOOD GAS RESULTS, AND THERE WERE RECURRENT SUCTION ALARMS ON THE IMPELLA. THE PATIENT PROCEEDED TO HAVE A PULSELESS ELECTRICAL ACTIVITY ARREST. AN ECHOCARDIOGRAM WAS PERFORMED AND REVEALED MASSIVE BILATERAL PLEURAL EFFUSIONS AND PERICARDIAL EFFUSION. THE PHYSICIAN BELIEVED THESE TO BE INJURIES FROM A PREVIOUS CARDIOPULMONARY RESUSCITATION AND THEN WAS HEPARINIZED FOR THE PCI AND IMPELLA IMPLANTATION WHICH POSSIBLY LED TO A SECOND ARREST. THE FAMILY DECIDED TO WITHDRAW CARE, AND THE PATIENT EXPIRED. Manufacturer Narrative: ALTHOUGH THE PATIENT'S PRESENTING CONDITION MAY BE MORE LIKELY HAVE IMPACTED THE EVENT THE IMPELLA CANNOT BE EXCLUDED AS A POSSIBLE CONTRIBUTING FACTOR IN THE PATIENT'S DEMISE. PRODUCT STATUS: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22893756&pc=MAF | 2024168-2025-09055 | 2025-06-30 04:00 | Death | ABBOTT VASCULAR | 2025-08-26 04:00 | MAF | MULTI-LINK 8 | Adverse Event Without Identified Device or Use Problem | Angina; Dyspnea; Myocardial Infarction; Obstruction/Occlusion; Heart Failure/Congestive Heart Failure | P020047 |  | 1 | Event Description: PATIENT ID: (B)(6). IT WAS REPORTED THAT ON (B)(6) 2025, ONE 3.5 X 15 MM MULTI-LINK 8 STENT WAS IMPLANTED IN THE HEAVILY CALCIFIED, PROXIMAL LEFT MAIN CORONARY ARTERY (LMCA) AND TWO 3.5 X 18 MM MULTI-LINK 8 STENTS WERE IMPLANTED IN THE MID AND PROXIMAL RIGHT CORONARY ARTERY. ON (B)(6) 2025, THE PATIENT PRESENTED WITH DYSPNEA ON EXERTION FOR DAYS WITH CHEST DISCOMFORT. MEDICATIONS (ISOSORBIDE DINITRATE, HEPARIN AND ANGIDIL) WERE GIVEN IN THE EMERGENCY ROOM. THE ELECTROCARDIOGRAM SHOWED AN ACUTE ANTERIOR INFARCTION IN THE LEFT ANTERIOR DESCENDING (LAD) CORONARY ARTERY. UNDER THE TENTATIVE DIAGNOSIS OF NON-ST ELEVATION ACUTE CORONARY SYNDROME WITH ACUTE DECOMPENSATED HEART FAILURE, THE PATIENT WAS ADMITTED FOR FURTHER CARE. RESTENOSIS WAS NOTED IN THE MID RCA AND THE LEFT MAIN STENTS ONLY. REVASCULARIZATION WAS PERFORMED ON (B)(6) 2025 IN THE STUDY TREATED LMCA AND THE MID RIGHT CORONARY ARTERY AND THE PROXIMAL LAD. THE PATIENT EXPIRED IN THE HOSPITAL ON (B)(6) 2025 FROM ISCHEMIC HEART FAILURE. IN THE OPINION OF THE PHYSICIAN, THE PATIENT'S DEATH WAS INDIRECTLY RELATED TO THE STENT. THE PATIENT WAS ADMITTED TO THE HOSPITAL ON (B)(6) 2025 UNTIL THE TIME OF HIS DEATH. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. PRODUCTION RECORD AND CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A QUERY OF THE COMPLAINT HANDLING DATABASE FOR THE REPORTED LOT REVEALED THERE IS NO INDICATION OF A LOT SPECIFIC ISSUE. THE REPORTED PATIENT EFFECTS OF DEATH, ANGINA, OCCLUSION AND MYOCARDIAL INFARCTION ARE LISTED IN THE MULTI-LINK 8, MULTI-LINK 8 SMALL VESSEL (SV) AND MULTI-LINK LONG LENGTH (LL) CORONARY STENT SYSTEMS (CSS) INSTRUCTIONS FOR USE AND ARE KNOWN PATIENT EFFECTS THAT MAY BE ASSOCIATED WITH USE OF A CORONARY STENT IN NATIVE CORONARY ARTERIES. A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECTS, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED; HOWEVER, THE TREATMENT APPEARS TO BE RELATED TO THE OPERATIONAL CONTEXT OF THE PROCEDURE. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN, OR LABELING; THEREFORE, NO PRODUCT-RELATED CORRECTIVE ACTION WILL BE IMPLEMENTED IN THIS CASE. THE ADDITIONAL DEVICE REFERENCED IN B5 IS FILED UNDER A SEPARATE MEDWATCH REPORT NUMBER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22893781&pc=NIM | 2124215-2025-58835 | 2020-05-16 04:00 | Death | SILK ROAD MEDICAL | 2025-08-26 04:00 | NIM | ENROUTE TRANSCAROTID STENT | Adverse Event Without Identified Device or Use Problem | Stroke/CVA |  |  | 1 | Event Description: IT WAS REPORTED VIA LITERATURE THAT ONE PATIENT WHO UNDERWENT A TRANSCAROTID ARTERY REVASCULARIZATION (TCAR) PROCEDURE EXPERIENCED STROKE/DEATH. THE AIM OF THIS ARTICLE IS TO DESCRIBE THE TECHNIQUE AND DISCUSS POSSIBLE COMPLICATIONS WHEN PERFORMING TCAR, WHICH IS POISED TO BECOME THE GOLD STANDARD CAROTID ARTERY MINIMAL ACCESS REVASCULARIZATION. THIS STUDY WAS A RETROSPECTIVE REVIEW OF 70 PATIENTS WHO UNDERWENT 112 TCAR PROCEDURES. IN THE PERIOPERATIVE PERIOD, DEFINED AS 30 DAYS POST TCAR, ONE PATIENT EXPERIENCED STROKE/DEATH. NO FURTHER INFORMATION WAS PROVIDED TO BOSTON SCIENTIFIC. Manufacturer Narrative: B2. THE DATE OF DEATH WAS NOT PROVIDED. B3 AND D6A. EVENT DATE/IMPLANT DATE WAS NOT PROVIDED; THEREFORE, DATE OF PUBLICATION WAS USED. DETAILS OF THE EVENT, INCLUDING PATIENT STATUS AND PRODUCT INFORMATION, WERE NOT PROVIDED TO BSC. BECAUSE THE PRODUCT LOT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) NUMBER AND OTHER SPECIFIC PRODUCT INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. SCHROEDER, A. C., BALCENIUK, M. D., SEBASTIAN, A., & STONER, M. C. (2020). TECHNICAL TIPS FOR SUCCESS IN TRANSCAROTID ARTERY REVASCULARIZATION. SEMINARS IN VASCULAR SURGERY, 33 (1 TO 2), 4 TO 9. HTTPS://DOI.ORG/10.1053/J.SEMVASCSURG.2020.05.001.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22893856&pc=DSQ | 2916596-2025-05600 | 2025-08-12 04:00 | Death | THORATEC CORPORATION | 2025-08-26 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE DEVICE WAS NOT RETURNED FOR EVALUATION. DUE TO PATIENT PRIVACY LAWS, THE MANAGING HOSPITAL WOULD NOT COMMUNICATE ANY ADDITIONAL INFORMATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. D IS CURRENTLY AVAILABLE. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22893937&pc=NVN | 2649622-2025-23251 | 2025-08-23 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-08-26 04:00 | NVN | CAPSURE FIX MRI | Impedance Problem | Insufficient Information | P090013 |  | 1 | Event Description: IT WAS REPORTED THAT THE RIGHT VENTRICULAR (RV) LEAD IMPEDANCE WAS OUT OF RANGE AND THE PATIENT IS DECEASED. Manufacturer Narrative: CONTINUATION OF D10: 5086-52 LEAD IMPLANTED (B)(6) 2012. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22893989&pc=NPT | 9617601-2025-01167 | 2024-11-22 05:00 | Death | MEDTRONIC MEXICO S. DE R.L. DE CV | 2025-08-26 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Insufficient Information | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 1 YEAR AND 4 MONTHS FOLLOWING THE IMPLANT OF A TRANSCATHETER VALVE, THE PATIENT DIED DUE TO AN UNKNOWN REASON. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22894434&pc=FRN | 2016493-2025-108143 | 2025-05-27 04:00 | Death | CAREFUSION SD | 2025-08-26 04:00 | FRN | ALARIS SYSTEM | Excess Flow or Over-Infusion | Insufficient Information | K133532 |  | 1 | Event Description: IT WAS INITIALLY REPORTED AN OVER INFUSION EVENT THAT OCCURRED ON (B)(6) 2025. PER REPORT, THE PATIENT "DIED." THE CUSTOMER IS REQUESTING THE MANUFACTURER TO REVIEW THE DEVICE LOGS TO IDENTIFY WHICH MEDICATIONS WERE INFUSED FROM 3:00 PM ON (B)(6) 2025 TO 1:30 AM ON (B)(6) 2025. NO FURTHER INFORMATION WAS PROVIDED. BD RECEIVED ADDITIONAL INFORMATION. IT WAS REPORTED THAT "THE REASON TO DO THIS (REVIEW OF DEVICE LOGS) ONLY TO CHECK THE MEDICATIONS AND THE RATE OF IT GIVEN BY THE MACHINE. I¿M NOT BELIEVE THAT THE DEVICE CAUSED THE DEATH BUT MAYBE HUMAN MISTAKE." PER CUSTOMER (UNIT MANAGER OF (B)(6)), "PATIENT CAME TO OUR UNIT FOR CHEMOTHERAPY, STARTED HIS MEDICATION AT 16:30 (RITUXIMAB) 660 ML OVER 4 HOURS, THEN (GEMCITABINE) 1000ML OVER 1 HOUR, THEN CISPLATIN 1000 ML OVER 1 HOUR." THE CUSTOMER STATED THEY DO NOT BELIEVE THAT THE DEVICE CAUSED THE DEATH. THE PRIMARY CAUSE OF DEATH WAS REPORTED AS "CARDIOPULMONARY ARREST." THE PATIENT DIED ON (B)(6) 2025 AT 01:48 AM. NO AUTOPSY WAS PERFORMED. THE DIAGNOSIS WAS "DIFFUSE LARGE B CELL LYMPHOMA." Manufacturer Narrative: THE EVENT LOGS HAVE BEEN RECEIVED AND THE EVALUATION IS PENDING. A FOLLOW UP REPORT WILL BE SUBMITTED WITH INVESTIGATION RESULTS ONCE THE EVALUATION HAS BEEN COMPLETED. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22894616&pc=PNJ | 9612164-2025-04205 | 2025-08-03 04:00 | Death | MEDTRONIC IRELAND | 2025-08-26 04:00 | PNJ | MICRA¿ VR2 | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT TWO WEEKS POST LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) IMPLANT THE PATIENT PASSED AWAY. CONGESTIVE HEART FAILURE WAS SUSPECTED TO BE THE POSSIBLE CAUSE OF DEATH. Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [MICRA] PRODUCT ID [MC2VR01-DELSYS] (SERIAL: (B)(6)). D9: NO. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22894663&pc=KDI | 9616026-2025-00063 | 2025-08-02 04:00 | Death | VANTIVE US HEALTHCARE LLC | 2025-08-26 04:00 | KDI | AK 98 MACHINES | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NI |  | 1 | Event Description: IT WAS REPORTED THAT DURING HEMODIALYSIS WITH AN AK 98 MACHINE, THE PATIENT PASSED AWAY. THE CAUSE OF THE DEATH WAS NOT REPORTED. IT WAS REPORTED THAT THE PATIENT WAS "RECEIVING DIALYSIS" AT THE TIME OF DEATH IT WAS NOT REPORTED IF AN AUTOPSY WAS PERFORMED. THE REPORTER STATED THAT "THIS WAS NOT AN UNEXPECTED DEATH AND MADE NO CLAIM AGAINST THE PRODUCT". THE REPORTER DECLINED TO PROVIDE ADDITIONAL INFORMATION. Manufacturer Narrative: SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22894687&pc=DYB | 2182208-2025-04251 | 2025-08-03 04:00 | Death | MEDTRONIC, INC. | 2025-08-26 04:00 | DYB | MICRA | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | K132030 |  | 1 | Event Description: IT WAS REPORTED THAT TWO WEEKS POST LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) IMPLANT THE PATIENT PASSED AWAY. CONGESTIVE HEART FAILURE WAS SUSPECTED TO BE THE POSSIBLE CAUSE OF DEATH. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22895222&pc=NPW | 2015691-2025-06996 | 2025-08-11 04:00 | Death | EDWARDS LIFESCIENCES | 2025-08-26 04:00 | NPW | EVOQUE TRICUSPID VALVE REPLACEMENT SYSTEM | Difficult to Insert | Great Vessel Perforation | P230013 |  | 1 | Event Description: AS REPORTED BY THE EDWARDS LIFESCIENCES AFFILIATE IN GERMANY, EDWARDS RECEIVED NOTIFICATION OF AN EVOQUE PROCEDURE IN TRICUSPID POSITION. DURING THE PROCEDURE, THE PHYSICIAN NOTED DIFFICULTY ADVANCING THE INITIAL 24 FR. DILATOR AND SWAPPED IT OUT FOR A 33 FR. DILATOR WITHOUT FURTHER ISSUES. HOWEVER, ON POST-OPERATIVE DAY ONE (1), THE PATIENT DEVELOPED A HARD STOMACH DUE TO INTERNAL BLEEDING. A COVERED STENT WAS ATTEMPTED TO CONTROL THE HEMORRHAGE, BUT THE INTERVENTION WAS UNSUCCESSFUL. IT WAS FURTHER REPORTED THAT THE PATIENT EXPIRED THEREAFTER. Manufacturer Narrative: THE MANUFACTURER'S INVESTIGATION IS ONGOING AND A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22895453&pc=DSQ | 2916596-2025-05378 | 2024-08-25 04:00 | Death | THORATEC CORPORATION | 2025-08-26 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED. THE MEDICAL CENTER MENTIONED THE PATIENT HAD SEVERAL STROKES. A CAUSE OF DEATH WAS NOT PROVIDED. THE MEDICAL CENTER DECOMMISSIONED THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) AFTER THE PATIENT EXPIRED. NO ALARMS WERE REPORTED. LOG FILES WERE NOT PROVIDED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22895644&pc=MNT | 2518422-2025-109356 | 2025-06-17 04:00 | Death | RESPIRONICS, INC. | 2025-08-26 04:00 | MNT | BIPAP A40 CANADA, HEATED TUBE CORE PKG | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K121623 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH A BIPAP A40 HAS PASSED AWAY. THE REPORTED EVENT MAKES NO ALLEGATION OF ANY SPECIFIC DEVICE MALFUNCTION, USE ERROR, FAILURE, LABELING, OR OTHER DEFICIENCY THAT IS RELEVANT TO THE HARMS REPORTED; THEREFORE, BASED ON INFORMATION AVAILABLE AT THIS TIME, THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO A SERIOUS INJURY OR DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22896289&pc=NVN | 2124215-2025-59839 | 2025-08-05 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-08-26 04:00 | NVN | INGEVITY? MRI | High impedance; High Capture Threshold | No Clinical Signs, Symptoms or Conditions | P150012 |  | 1 | Event Description: IT WAS REPORTED THAT THIS RIGHT VENTRICULAR (RV) LEAD EXHIBITED ELEVATED PACING IMPEDANCES AT 1000 OHMS AND CAPTURE THRESHOLD. DURING THE EXPLANT PROCEDURE, THE PATIENT EXPIRED DUE TO EXTRACTION COMPLICATIONS. THIS LEAD REMAINS IN SITU. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22897160&pc=OUT | 2029214-2025-01901 | 2025-07-29 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-08-26 04:00 | OUT | PIPELINE FLEX W/SHIELD TECHNOLOGY | Adverse Event Without Identified Device or Use Problem | Aneurysm | P100018 |  | 1 | Event Description: MEDTRONIC RECEIVED A REPORT THAT THE PATIENT DIED OF AN UNKNOWN CAUSE. THE PATIENT RECEIVED TREATMENT FOR A SACCULAR, UNRUPTURED ANEURYSM IN THE LEFT INTERNAL CAROTID ARTERY IN THE C7 SEGMENT IN THE SIDEWALL WITH A MAX DIAMETER OF 13.1 MM AND A NECK SIZE OF 4 MM. POST IMPLANT THERE WAS COMPLETE STASIS. THERE WERE NO SIDE BRANCHES COVERED BY THE PIPELINE, THERE WAS NO FLATTENING OR TWISTING. ANCILLARY DEVICES INCLUDE A BENCHMARK, PHENOM PLUS, AND A PHENOM 27 CATHETERS. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22897672&pc=NVN | 2124215-2025-59914 | 2025-08-11 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-08-26 04:00 | NVN | INGEVITY+ | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P150012 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A GENERAL CHANGEOUT PROCEDURE, THIS RIGHT VENTRICULAR (RV) LEAD WAS OPENED AND WHEN THE PHYSICIAN ATTEMPTED TO GAIN VENOUS ACCESS, THE PATIENT ENDED UP DYING SHORTLY AFTERWARDS ON THE OPERATING TABLE. THE RV LEAD WAS NEVER IMPLANTED IN THE PATIENT, AND NO ALLEGATION WAS MADE AGAINST IT IN RELATION TO THE PATIENT'S PASSING. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22897933&pc=OZD | 1220648-2025-45940 | 2025-07-28 04:00 | Death | ABIOMED, INC. | 2025-08-26 04:00 | OZD | IMPELLA CP | Difficult or Delayed Positioning; Fluid/Blood Leak; Deformation Due to Compressive Stress; Device Sensing Problem | No Clinical Signs, Symptoms or Conditions | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT IN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT AND IT WAS REPORTED THERE WERE ISSUES WITH THE PUMP POSITION, ALARMS FOR IMPELLA IN AORTA, BLOOD SEEN ON THE ANTI-CONTAMINATION SLEEVE, AND A CATHETER KINK. THE PATIENT PRESENTED WITH TO THE HOSPITAL ON (B)(6) 2025 IN ST ELEVATION MYOCARDIAL INFARCTION WITH AN OCCULTED RAMUS WHICH WAS PERFORATED DURING A PERCUTANEOUS CORONARY INTERVENTION. THE SURGEON WAS UNABLE TO RESOLVE THE PERFORATION WITH BALLOON INFLATIONS, AND CARDIOVASCULAR SURGERY WAS CONSULTED. AN INTRA-AORTIC BALLOON PUMP (IABP) WAS PLACED, AND THE PATIENT WAS TAKEN FOR 2 CORONARY ARTERY BYPASS GRAFTS. THE PATIENT REMAINED ON THE IABP FOR FOUR DAYS AND THEN REMOVED FROM SUPPORT. AFTER THE REMOVAL OF THE IABP, THE PATIENT BEGAN SHOWING SIGNS OF SHOCK AND WENT INTO CARDIAC ARREST REQUIRING ONE ROUND OF CARDIOVASCULAR RESUSCITATION, DEFIBRILLATION, AND WAS RE-INTUBATED. THE PATIENT WAS THEN EMERGENTLY TAKEN TO THE CATHETERIZATION LABORATORY FOR AN IMPELLA CP IMPLANT. IT WAS REPORTED THAT DURING THE FIRST DAY OF IMPELLA CP SUPPORT THERE WAS A PERSISTENT IMPELLA IN AORTA ALARM. REPOSITIONING WAS PERFORMED AND IT WAS NOTED THAT THE DEVICE WAS CAUGHT IN THE PAPILLARY MUSCLE. THE PUMP WAS FREED FROM THE MUSCLE, BUT THE ALARMS CONTINUED. ON THE SECOND DAY OF SUPPORT, BLOOD WAS FOUND IN THE ANTI-CONTAMINATION SLEEVE AND A CATHETER KINK WAS NOTED, WITH NO NOTED INTERVENTION REQUIRED. ON THE FOURTH DAY OF SUPPORT, CARE WAS WITHDRAWN AT THE REQUEST OF THE FAMILY AND THE PATIENT EXPIRED. Manufacturer Narrative: ALTHOUGH THE PATIENT'S PRESENTING CONDITION MAY BE MORE LIKELY HAVE IMPACTED THE EVENT THE IMPELLA CANNOT BE EXCLUDED AS A POSSIBLE CONTRIBUTING FACTOR IN THE PATIENT'S DEMISE. PRODUCT STATUS: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22898116&pc= | MW5175265 | 2025-06-26 04:00 | Death | MEDTRONIC, INC. | 2025-08-26 04:00 | --- | CAPSUREFIX NOVUS MRI SURESCAN PACING LEAD | Adverse Event Without Identified Device or Use Problem | Septic Shock; Cardiogenic Shock |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO CARDIOGENIC SHOCK AND SEPTIC SHOCK. THERE IS NO INFORMATION CONCERNING THE RELATIONSHIP BETWEEN CAUSE OF DEATH AND THE RIGHT ATRIAL (RA) AND RIGHT VENTRICULAR (RV) LEADS. REFERENCE REPORT #: MW5175303. THIS REPORT REFLECTS INFORMATION RECEIVED BY FDA IN THE FORM OF A NOTIFICATION PER 803.22 (B)(2).Event Description: ADDITIONAL INFORMATION RECEIVED FOR ON 8/29/2025 FOR REPORT MW5175265 TO UPDATE STATUS OF THE REPORT TO DEATH. REFERENCE REPORT MW5175303. THIS REPORT REFLECTS INFORMATION RECEIVED BY FDA IN THE FORM OF A NOTIFICATION PER 803.22 (B)(2).Event Description: ADDITIONAL INFORMATION RECEIVED FOR ON 9/1/2025 FOR REPORT MW5175265 TO UPDATE PROCODE. REFERENCE RERPORT MW5175303. THIS REPORT REFLECTS INFORMATION RECEIVED BY FDA IN THE FORM OF A NOTIFICATION PER 803.22 (B)(2).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22898156&pc=MCX | 2124215-2025-59032 | 2025-07-30 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-08-26 04:00 | MCX | ROTAPRO? | Failure to Advance; Output Problem | Myocardial Infarction | P900056 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT DEATH OCCURRED. A ROTAPRO 1.50MM WAS SELECTED FOR TREATMENT OF THE TARGET LESION. AS THE PROCEDURE BEGAN, THE PHYSICIAN WAS ABLE TO BEGIN ROTATIONAL ATHERECTOMY AT 140,000 RPMS. IT WAS NOTED THAT THE ROTATIONAL SPEED DECREASED TO 120,000 RPMS, AND A FURTHER DROP IN PRESSURE WAS OBSERVED. AFTER FIVE MINUTES THERE WAS STILL NO SUCCESS IN ADVANCING TO THE TARGET LESION. BECAUSE OF THIS, THE PHYSICIAN CONCLUDED THAT THE DEVICE WAS NOT ABLE TO ADVANCE DUE TO THE DROP IN RPMS. THE ROTAPRO WAS SWAPPED OUT FOR ANOTHER OF THE SAME DEVICE AND WAS ABLE TO BE ADVANCED. IT WAS AT THIS POINT THAT THE PHYSICIAN NOTICED SLOW REFLOW. THE PATIENT THEN PROGRESSED TO NO REFLOW AND CARDIAC ARREST. AN ATTEMPT TO ADMINISTER EPINEPHRINE AND RESUSCITATION MANEUVERS WERE PERFORMED, BUT THE PATIENT DIED. Manufacturer Narrative: E1 INITIAL REPORTER ADDRESS: (B)(6).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22898307&pc=QZI | MW5175275 | 2025-07-22 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-08-26 04:00 | QZI | FARWAVE CATHETER-FARAWAVE PULSED FIELD ABLATION. | Adverse Event Without Identified Device or Use Problem | Ventricular Fibrillation |  |  | 1 | Event Description: IT WAS REPORTED THAT DURING A PFA ATRIAL FIBRILLATION CASE, THE PATIENT HAD PASSED AWAY. THEY HAD COMPLETED PFA ABLATION ON THE LEFT VEINS AND THE POSTERIOR WALL AND THEN MOVED ONTO ABLATE ON THE RIGHT SUPERIOR VEIN. AFTER APPLYING THE SECOND PFA LESION ON THE MIDDLE OF THE RIGHT SUPERIOR VEIN, IT WAS NOTICED ON THE CARTO¿ 3 AND RECORDING SYSTEM, THAT THE PATIENT HAD GONE INTO "VF" OR VENTRICULAR FIBRILLATION. THEY BEGAN TO ADMINISTER CHEST COMPRESSIONS TO THE PATIENT. IT WAS CONFIRMED THAT THERE WAS NO EFFUSION PRESENT IN THE PATIENT. THEY HAD CALLED A "CODE" FOR THE PATIENT. THE CALLER THEN REPORTED THAT THEY HAD TO LEAVE THE ROOM AT THIS POINT AND THE CALLER DID NOT HAVE ANY FURTHER EVENT DETAILS AVAILABLE. CONFIRMED THAT THE PATIENT HAD PASSED AWAY ON (B)(6) 2025. THE CALLER STATED THAT THE PHYSICIAN SUSPECTS THE PATIENT MAY HAVE HAD AN "MI" OR MYOCARDIAL INFARCTION, BUT THE CALLER STATED THIS INFORMATION HAS NOT YET BEEN CONFIRMED. THE ACUSON ACUNAV¿ ULTRASOUND CATHETER WAS STILL IN THE BODY, WHEN THE PATIENT'S VF WAS NOTICED, BUT THE PENTARAY® NAV ECO CATHETER WAS NO LONGER IN THE BODY. THEY DID MAP WITH THE PENTARAY® NAV ECO CATHETER EARLIER IN THE PROCEDURE. THE FOLLOWING BWI DEVICES WERE ALSO IN USE: ACUSON ACUNAV¿ ULTRASOUND CATHETER THE CALLER STATED THAT THE ACUSON ACUNAV¿ ULTRASOUND CATHETER AND THE PENTARAY® NAV ECO CATHETER ARE NOT AVAILABLE FOR RETURN, AS THE ACCOUNT HAS REQUESTED TO RETAIN THEM FOR THE PATIENT'S AUTOPSY. FARAPULSE¿ SYSTEM WAS IN USE. XML SOFTWARE WAS IN USE. THIS REPORT REFLECTS INFORMATION RECEIVED BY FDA IN THE FORM OF A NOTIFICATION PER 803.22 (B)(2).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22898566&pc=DSQ | 3007042319-2025-01966 | 2025-07-18 04:00 | Death | HEARTWARE, INC. | 2025-08-26 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE VENTRICULAR ASSIST DEVICE (VAD) PATIENT SUBSEQUENTLY PASSED AWAY PEACEFULLY DUE TO CLINICAL ISSUES UNRELATED TO THE HVAD THERAPY. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22898573&pc=NIQ | 2024168-2025-09094 | 2017-01-01 05:00 | Death | ABBOTT VASCULAR | 2025-08-26 04:00 | NIQ | XIENCE PRIME EVEROLIMUS ELUTING CORONARY STENT SYSTEM | Adverse Event Without Identified Device or Use Problem | Unspecified Heart Problem | P110019 |  | 1 | Event Description: IT WAS REPORTED IN A SUMMARY RESEARCH ARTICLE A COMPARISON OF THE LONG TERM CLINICAL OUTCOMES OF BIOABSORBABLE POLYMER DRUG ELUTING STENTS (DES) VERSUS DURABLE POLYMER DES. THE ARTICLE ANALYZED A TOTAL OF 18 CLINICAL STUDIES THAT INVOLVED 28,874 PATIENTS. THE PROCEDURE INCLUDED THE IMPLANTATION OF AN UNSPECIFIED XIENCE STENT, THE XIENCE PRIME STENT, THE XIENCE EXPEDITION STENT, THE XIENCE V STENT AND THE PROMUS STENT. IN THE BEFORE AND AFTER TWO YEAR FOLLOW-UP SOME OF THE LONG TERM CLINICAL OUTCOMES INCLUDED CARDIAC DEATH, TARGET LESION REVASCULARIZATION (TLR), LATE STENT THROMBOSIS, TARGET LESION FAILURE (TLF), MYOCARDIAL INFARCTION (MI) AND TARGET VESSEL REVASCULARIZATION (TVR). THE STUDY REVEALED NO CLINICALLY SIGNIFICANT DIFFERENCES IN THE CLINICAL OUTCOMES OF BIOABSORBABLE POLYMER DES VERSUS DURABLE POLYMER DES. ADDITIONAL INFORMATION CAN BE FOUND IN THE LITERATURE REVIEW ARTICLE, " COMPARISON OF LONG-TERM CLINICAL OUTCOMES OF BIOABSORBABLE POLYMER VERSUS DURABLE POLYMER DRUG-ELUTING STENTS: A SYSTEMATIC REVIEW AND META-ANALYSIS". Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORDS FOR THIS PRODUCT COULD NOT BE REVIEWED AND A SIMILAR COMPLAINT QUERY COULD NOT BE PERFORMED BECAUSE THE PART NUMBER AND/OR LOT NUMBER WAS NOT REPORTED, AND THE DEVICE WAS NOT RETURNED. A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT DEATH AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED. THE REPORTED PATIENT DEATH IS LISTED IN THE XIENCE PRIME EVEROLIMUS ELUTING CORONARY STENT SYSTEMS INSTRUCTIONS FOR USE AS A KNOWN PATIENT EFFECT(S) OF CORONARY STENTING PROCEDURES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING; THEREFORE, NO PRODUCT-RELATED CORRECTIVE ACTION WILL BE IMPLEMENTED IN THIS CASE. B2, B3, D6A: ESTIMATED DATE. . THE ADDITIONAL PATIENT EFFECTS REPORTED IN B5 ARE CAPTURED UNDER A SEPARATE MEDWATCH REPORT. THE ADDITIONAL DEVICES REPORTED IN B5 ARE CAPTURED UNDER A SEPARATE MEDWATCH REPORT. ATTACHMENT ARTICLE TITLED: "COMPARISON OF LONG-TERM CLINICAL OUTCOMES OF BIOABSORBABLE POLYMER VERSUS DURABLE POLYMER DRUG-ELUTING STENTS: A SYSTEMATIC REVIEW AND META-ANALYSIS".
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22898901&pc=FRN | 3012307300-2025-10054 | 2025-08-04 04:00 | Death | ICU MEDICAL, INC. | 2025-08-26 04:00 | FRN | CADD-SOLIS VIP AMBULATORY INFUSION PUMP | Adverse Event Without Identified Device or Use Problem | Convulsion/Seizure | K111275 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WITH B-CELL ACUTE LYMPHOBLASTIC LEUKEMIA PASSED AWAY. NO INFORMATION HAS BEEN PROVIDED BY THE REPORTER ON THE CAUSE OF DEATH. IT WAS REPORTED THAT THE PATIENT WAS FOUND SEIZING AT HOME, BUT NO ADDITIONAL INFORMATION HAS BEEN PROVIDED REGARDING THIS EVENT. PER REPORTER THE PUMP WAS IN USE BY THE PATIENT WHEN THEY PASSED, BUT NO ALLEGATION WAS REPORTED THAT THE PUMP CAUSED OR CONTRIBUTED TO THE EVENT. A MEDICAL EVALUATION IS IN PROGRESS AND WILL BE PROVIDED ON A SUPPLEMENTAL REPORT. IT WAS REPORTED THAT THE STARTING VOLUME WAS 101 (WITH THE MEDICATION DOSING VOLUME OF 100.8), CONTINUOUS INFUSION RATE OF 0.6 ML/HOUR, RESERVOIR VOLUME AT THE TIME OF PUMP STOPPED WHICH WAS 54.3 ML, GIVEN VOLUME WAS 46.75 ML. Manufacturer Narrative: INVESTIGATION INCLUDING ROOT CAUSE ANALYSIS IS IN PROGRESS. A SUPPLEMENTAL MDR WILL BE FILED AS NECESSARY IN ACCORDANCE WITH 21 CFR 803.56 WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22898973&pc=NIQ | 2024168-2025-09103 | 2017-01-01 05:00 | Death | ABBOTT VASCULAR | 2025-08-26 04:00 | NIQ | XIENCE V EVEROLIMUS ELUTING CORONARY STENT SYSTEM | Adverse Event Without Identified Device or Use Problem | Unspecified Heart Problem | P070015 |  | 1 | Event Description: IT WAS REPORTED IN A SUMMARY RESEARCH ARTICLE A COMPARISON OF THE LONG TERM CLINICAL OUTCOMES OF BIOABSORBABLE POLYMER DRUG ELUTING STENTS (DES) VERSUS DURABLE POLYMER DES. THE ARTICLE ANALYZED A TOTAL OF 18 CLINICAL STUDIES THAT INVOLVED 28,874 PATIENTS. THE PROCEDURE INCLUDED THE IMPLANTATION OF AN UNSPECIFIED XIENCE STENT, THE XIENCE PRIME STENT, THE XIENCE XPEDITION STENT, THE XIENCE V STENT AND THE PROMUS STENT. IN THE BEFORE AND AFTER TWO YEAR FOLLOW-UP SOME OF THE LONG-TERM CLINICAL OUTCOMES INCLUDED CARDIAC DEATH, TARGET LESION REVASCULARIZATION (TLR), LATE STENT THROMBOSIS, TARGET LESION FAILURE (TLF), MYOCARDIAL INFARCTION (MI) AND TARGET VESSEL REVASCULARIZATION (TVR). THE STUDY REVEALED NO CLINICALLY SIGNIFICANT DIFFERENCES IN THE CLINICAL OUTCOMES OF BIOABSORBABLE POLYMER DES VERSUS DURABLE POLYMER DES. ADDITIONAL INFORMATION CAN BE FOUND IN THE LITERATURE REVIEW ARTICLE, " COMPARISON OF LONG-TERM CLINICAL OUTCOMES OF BIOABSORBABLE POLYMER VERSUS DURABLE POLYMER DRUG-ELUTING STENTS: A SYSTEMATIC REVIEW AND META-ANALYSIS." Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORDS FOR THIS PRODUCT COULD NOT BE REVIEWED AND A SIMILAR COMPLAINT QUERY COULD NOT BE PERFORMED BECAUSE THE PART NUMBER AND/OR LOT NUMBER WAS NOT REPORTED, AND THE DEVICE WAS NOT RETURNED. A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT DEATH AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED.. THE REPORTED PATIENT EFFECT(S) OF DEATH, IS LISTED IN THE XIENCE V EVEROLIMUS ELUTING CORONARY STENT SYSTEMS INSTRUCTIONS FOR USE AS A KNOWN PATIENT EFFECT(S) OF CORONARY STENTING PROCEDURES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING; THEREFORE, NO PRODUCT-RELATED CORRECTIVE ACTION WILL BE IMPLEMENTED IN THIS CASE. B2, B3, D6A: ESTIMATED DATES D4- THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. THE ADDITIONAL PATIENT AFFECTS REPORTED IN B5 ARE CAPTURED UNDER A SEPARATE MEDWATCH REPORT. THE ADDITIONAL DEVICES REPORTED IN B5 ARE CAPTURED UNDER A SEPARATE MEDWATCH REPORT. ATTACHMENT ARTICLE TITLED: "COMPARISON OF LONG-TERM CLINICAL OUTCOMES OF BIOABSORBABLE POLYMER VERSUS DURABLE POLYMER DRUG-ELUTING STENTS: A SYSTEMATIC REVIEW AND META-ANALYSIS"
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22898984&pc=LYJ | 1644487-2025-10542 | 2025-07-31 04:00 | Death | LIVANOVA USA, INC. | 2025-08-26 04:00 | LYJ | PULSE GEN MODEL 1000D | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P970003 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. IT WAS STATED BY THE PROVIDER THAT THE PATIENT HAD A SEIZURE IN THE MIDDLE OF THE EVENING AND HIT HIS HEAD. IT IS UNKNOWN WHAT THE CAUSE OF DEATH WAS. NO OTHER RELEVANT INFORMATION HAS BEEN RECEIVED TO DATE. Manufacturer Narrative: LIVANOVA USA, INC. SUBMITS THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION, BASED ON INFORMATION THAT LIVANOVA HAS OBTAINED, BUT MAY NOT HAVE BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION, OR A CONCLUSION BY FDA OR ANYONE ELSE, THAT THE DEVICE, LIVANOVA, OR LIVANOVA'S EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS¿ OR "MALFUNCTIONS¿. THESE WORDS ARE INCORPORATED INTO THE FDA 3500A MEDWATCH FORM BY THE FDA, AND LIVANOVA OBJECTS TO THEIR USE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22898995&pc=NPT | 9617601-2025-01175 | 2024-02-15 05:00 | Death | MEDTRONIC MEXICO S. DE R.L. DE CV | 2025-08-26 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Heart Failure/Congestive Heart Failure | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY ONE MONTH FOLLOWING THE IMPLANT OF A TRANSCATHETER VALVE, THE PATIENT WAS TRANSFERRED TO A DIFFERENT HOSPITAL FOR REHABILITATION. APPROXIMATELY 9 MONTHS FOLLOWING THE IMPLANT OF THE VALVE, THE PATIENT WAS HOSPITALIZED DUE TO GRADUALLY WORSENING CONGESTIVE HEART FAILURE (CHF), AND WORSENING CARDIAC FUNCTION. APPROXIMATELY 11 MONTHS FOLLOWING THE IMPLANT OF THE VALVE, THE PATIENT DIED. THE CAUSE OF DEATH WAS CARDIAC RELATED, BUT WAS NOT SPECIFIED. PER THE PHYSICIAN, THERE WAS NO CAUSAL RELATIONSHIP BETWEEN THE DEATH AND DEVICE OR IMPLANT PROCEDURE. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22899017&pc=NIQ | 2024168-2025-09107 | 2017-01-01 05:00 | Death | ABBOTT VASCULAR INC. | 2025-08-26 04:00 | NIQ | XIENCE XPEDITION | Adverse Event Without Identified Device or Use Problem | Unspecified Heart Problem | P110019 |  | 1 | Event Description: IT WAS REPORTED IN A SUMMARY RESEARCH ARTICLE A COMPARISON OF THE LONG TERM CLINICAL OUTCOMES OF BIOABSORBABLE POLYMER DRUG ELUTING STENTS (DES) VERSUS DURABLE POLYMER DES. THE ARTICLE ANALYZED A TOTAL OF 18 CLINICAL STUDIES THAT INVOLVED 28,874 PATIENTS. THE PROCEDURE INCLUDED THE IMPLANTATION OF AN UNSPECIFIED XIENCE STENT, THE XIENCE PRIME STENT, THE XIENCE EXPEDITION STENT, THE XIENCE V STENT AND THE PROMUS STENT. IN THE BEFORE AND AFTER TWO YEAR FOLLOW-UP SOME OF THE LONG TERM CLINICAL OUTCOMES INCLUDED CARDIAC DEATH, TARGET LESION REVASCULARIZATION (TLR), LATE STENT THROMBOSIS, TARGET LESION FAILURE (TLF), MYOCARDIAL INFARCTION (MI) AND TARGET VESSEL REVASCULARIZATION (TVR). THE STUDY REVEALED NO CLINICALLY SIGNIFICANT DIFFERENCES IN THE CLINICAL OUTCOMES OF BIOABSORBABLE POLYMER DES VERSUS DURABLE POLYMER DES. ADDITIONAL INFORMATION CAN BE FOUND IN THE LITERATURE REVIEW ARTICLE, " COMPARISON OF LONG-TERM CLINICAL OUTCOMES OF BIOABSORBABLE POLYMER VERSUS DURABLE POLYMER DRUG-ELUTING STENTS: A SYSTEMATIC REVIEW AND META-ANALYSIS". Manufacturer Narrative: ATTACHMENT ARTICLE TITLED: "COMPARISON OF LONG-TERM CLINICAL OUTCOMES OF BIOABSORBABLE POLYMER VERSUS DURABLE POLYMER DRUG-ELUTING STENTS: A SYSTEMATIC REVIEW AND META-ANALYSIS". THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORDS FOR THIS PRODUCT COULD NOT BE REVIEWED AND A SIMILAR COMPLAINT QUERY COULD NOT BE PERFORMED BECAUSE THE PART NUMBER AND/OR LOT NUMBER WAS NOT REPORTED, AND THE DEVICE WAS NOT RETURNED. A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECT(S), AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED; HOWEVER, THE SUBSEQUENT TREATMENT APPEARS TO BE RELATED TO THE OPERATIONAL CONTEXT OF THE PROCEDURE. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING; THEREFORE, NO PRODUCT-RELATED CORRECTIVE ACTION WILL BE IMPLEMENTED IN THIS CASE. THE REPORTED PATIENT EFFECT(S) OF DEATH IS LISTED IN THE XIENCE EXPEDITION EVEROLIMUS ELUTING CORONARY STENT SYSTEMS INSTRUCTIONS FOR USE AS A KNOWN PATIENT EFFECT(S) OF CORONARY STENTING PROCEDURES. THE ADDITIONAL PATIENT EFFECTS AND/OR MALFUNCTIONS REPORTED IN THE ARTICLE ARE CAPTURED UNDER A SEPARATE MEDWATCH REPORT. B2, B3: ESTIMATED DATE D4- THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22899036&pc=NIQ | 2024168-2025-09108 | 2017-01-01 05:00 | Death | ABBOTT VASCULAR | 2025-08-26 04:00 | NIQ | PROMUS EVEROLIMUS ELUTING CORONARY STENT SYSTEM | Adverse Event Without Identified Device or Use Problem | Unspecified Heart Problem | P070015 |  | 1 | Event Description: IT WAS REPORTED IN A SUMMARY RESEARCH ARTICLE A COMPARISON OF THE LONG TERM CLINICAL OUTCOMES OF BIOABSORBABLE POLYMER DRUG ELUTING STENTS (DES) VERSUS DURABLE POLYMER DES. THE ARTICLE ANALYZED A TOTAL OF 18 CLINICAL STUDIES THAT INVOLVED 28,874 PATIENTS. THE PROCEDURE INCLUDED THE IMPLANTATION OF AN UNSPECIFIED XIENCE STENT, THE XIENCE PRIME STENT, THE XIENCE EXPEDITION STENT, THE XIENCE V STENT AND THE PROMUS STENT. IN THE BEFORE AND AFTER TWO YEAR FOLLOW-UP SOME OF THE LONG TERM CLINICAL OUTCOMES INCLUDED CARDIAC DEATH, TARGET LESION REVASCULARIZATION (TLR), LATE STENT THROMBOSIS, TARGET LESION FAILURE (TLF), MYOCARDIAL INFARCTION (MI) AND TARGET VESSEL REVASCULARIZATION (TVR). THE STUDY REVEALED NO CLINICALLY SIGNIFICANT DIFFERENCES IN THE CLINICAL OUTCOMES OF BIOABSORBABLE POLYMER DES VERSUS DURABLE POLYMER DES. ADDITIONAL INFORMATION CAN BE FOUND IN THE LITERATURE REVIEW ARTICLE, " COMPARISON OF LONG-TERM CLINICAL OUTCOMES OF BIOABSORBABLE POLYMER VERSUS DURABLE POLYMER DRUG-ELUTING STENTS: A SYSTEMATIC REVIEW AND META-ANALYSIS". Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORDS FOR THIS PRODUCT COULD NOT BE REVIEWED AND A SIMILAR COMPLAINT QUERY COULD NOT BE PERFORMED BECAUSE THE PART NUMBER AND/OR LOT NUMBER WAS NOT REPORTED, AND THE DEVICE WAS NOT RETURNED. A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT DEATH AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED.. THE REPORTED PATIENT EFFECT(S) OF DEATH IS LISTED IN THE PROMUS EVEROLIMUS ELUTING CORONARY STENT SYSTEMS INSTRUCTIONS FOR USE AS A KNOWN PATIENT EFFECT(S) OF CORONARY STENTING PROCEDURES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING; THEREFORE, NO PRODUCT-RELATED CORRECTIVE ACTION WILL BE IMPLEMENTED IN THIS CASE. B2, B3, D6A- ESTIMATED DATES D4: THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. THE ADDITIONAL PATIENT AFFECTS REPORTED IN B5 ARE CAPTURED UNDER A SEPARATE MEDWATCH REPORT. THE ADDITIONAL DEVICES REPORTED IN B5 ARE CAPTURED UNDER A SEPARATE MEDWATCH REPORT. LITERATURE ATTACHMENT: ARTICLE TITLE: COMPARISON OF LONG TERM CLINICAL OUTCOMES OF BIOABSORBABLE POLYMER VERSUS DURABLE POLYMER DRUG ELUTING STENTS: A SYSTEMATIC REVIEW AND META ANALYSIS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22899193&pc=DSQ | 2916596-2025-05548 | 2025-08-04 04:00 | Death | THORATEC CORPORATION | 2025-08-26 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Bone Fracture(s) | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT ACCORDING TO THE PERFUSIONIST, THE PATIENT BROKE THEIR FEMUR BONE AND HAD SURGERY FOR THAT. POST OPERATIVELY, THE PATIENT EXPERIENCED CARDIAC ARREST AND PASSED AWAY. IT WAS UNCLEAR IF THE DEVICE WOULD BE SENT IN FOR INVESTIGATION. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22899317&pc=NIQ | 2024168-2025-09111 | 2017-01-01 05:00 | Death | ABBOTT VASCULAR | 2025-08-26 04:00 | NIQ | XIENCE EVEROLIMUS ELUTING CORONARY STENT SYSTEM | Adverse Event Without Identified Device or Use Problem | Unspecified Heart Problem | P070015 |  | 1 | Event Description: IT WAS REPORTED IN A SUMMARY RESEARCH ARTICLE A COMPARISON OF THE LONG-TERM CLINICAL OUTCOMES OF BIOABSORBABLE POLYMER DRUG ELUTING STENTS (DES) VERSUS DURABLE POLYMER DES. THE ARTICLE ANALYZED A TOTAL OF 18 CLINICAL STUDIES THAT INVOLVED 28,874 PATIENTS. THE PROCEDURE INCLUDED THE IMPLANTATION OF AN UNSPECIFIED XIENCE STENT, THE XIENCE PRIME STENT, THE XIENCE XPEDITION STENT, THE XIENCE V STENT AND THE PROMUS STENT. IN THE BEFORE AND AFTER TWO YEAR FOLLOW-UP SOME OF THE LONG-TERM CLINICAL OUTCOMES INCLUDED CARDIAC DEATH, TARGET LESION REVASCULARIZATION (TLR), LATE STENT THROMBOSIS, TARGET LESION FAILURE (TLF), MYOCARDIAL INFARCTION (MI) AND TARGET VESSEL REVASCULARIZATION (TVR). THE STUDY REVEALED NO CLINICALLY SIGNIFICANT DIFFERENCES IN THE CLINICAL OUTCOMES OF BIOABSORBABLE POLYMER DES VERSUS DURABLE POLYMER DES. ADDITIONAL INFORMATION CAN BE FOUND IN THE LITERATURE REVIEW ARTICLE, " COMPARISON OF LONG-TERM CLINICAL OUTCOMES OF BIOABSORBABLE POLYMER VERSUS DURABLE POLYMER DRUG-ELUTING STENTS: A SYSTEMATIC REVIEW AND META-ANALYSIS". Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORDS FOR THIS PRODUCT COULD NOT BE REVIEWED AND A SIMILAR COMPLAINT QUERY COULD NOT BE PERFORMED BECAUSE THE PART NUMBER AND/OR LOT NUMBER WAS NOT REPORTED, AND THE DEVICE WAS NOT RETURNED. A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT DEATH, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED; THE REPORTED PATIENT DEATH IS LISTED IN THE XIENCE V EVEROLIMUS ELUTING CORONARY STENT SYSTEMS INSTRUCTIONS FOR USE (EL2078398 REVISION E SECTION 8.3) AS A KNOWN PATIENT EFFECT(S) OF CORONARY STENTING PROCEDURES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING; THEREFORE, NO PRODUCT-RELATED CORRECTIVE ACTION WILL BE IMPLEMENTED IN THIS CASE.
B2, B3, D6A- ESTIMATED DATES
D4: THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED.
THE ADDITIONAL PATIENT AFFECTS REPORTED IN B5 ARE CAPTURED UNDER A SEPARATE MEDWATCH REPORT.
THE ADDITIONAL DEVICES REPORTED IN B5 ARE CAPTURED UNDER A SEPARATE MEDWATCH REPORT.
LITERATURE ATTACHMENT: ARTICLE TITLE: "COMPARISON OF LONG TERM CLINICAL OUTCOMES OF BIOABSORBABLE POLYMER VERSUS DURABLE POLYMER DRUG ELUTING STENTS: A SYSTEMATIC REVIEW AND META ANALYSIS"
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22900241&pc=KNT | 3010757606-2025-00377 | 2025-07-27 04:00 | Death | ABBVIE - MEDICAL DEVICE CENTER | 2025-08-26 04:00 | KNT | DUODOPA_DUOPA | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K142793 |  | 1 | Event Description: ON (B)(6) 2025, A PATIENT IN GERMANY UNDERWENT A PROCEDURE FOR THE PLACEMENT OF PERCUTANEOUS ENDOSCOPIC GASTROSTOMY (PEG) TUBE WITH JEJUNAL (PEG-J) TUBE. ON (B)(6) 2025, THE PATIENT WAS DISCHARGED HOME WITH PALLIATIVE TREATMENT. ON (B)(6) 2025, THE PATIENT DIED AT HOME. THE CAUSE OF DEATH WAS UNKNOWN, THE WIFE REPORTED THAT AN AUTOPSY WILL NOT BE PERFORMED. THE PHYSICIAN SUSPECTED THE PATIENT MAY HAVE HAD ASPIRATION PNEUMONIA, NO DIAGNOSTIC IMAGING WAS PERFORMED TO CONFIRM THE DIAGNOSIS. THE FINAL CAUSE OF DEATH WAS NOT KNOWN. THEREFORE, OUT OF ABUNDANCE OF CAUTION DUE TO THE PATIENT PASSING AWAY 10 DAYS AFTER THE TUBING PLACEMENT, ABBVIE IS CONSERVATIVELY REPORTING THE EVENT. Manufacturer Narrative: CATALOG NUMBER IN D4 IS THE INTERNATIONAL LIST NUMBER WHICH IS SIMILAR TO US LIST NUMBER OF 062910. THE DEVICE INVOLVED IN THE EVENT WAS NOT RETURNED; THEREFORE, A RETURN SAMPLE EVALUATION IS UNABLE TO BE PERFORMED. POST PROCEDURAL COMPLICATION IS A KNOWN COMPLICATION OF A PEG- J TUBE PLACEMENT. ASPIRATION PNEUMONIA WAS SUSPECTED; IT WAS NOT CONFIRMED. IF ANY FURTHER RELEVANT INFORMATION IS IDENTIFIED OR OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22900384&pc=NIM | 2124215-2025-59714 | 2011-01-01 05:00 | Death | SILK ROAD MEDICAL | 2025-08-26 04:00 | NIM | ENROUTE TRANSCAROTID STENT | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: IT WAS REPORTED VIA LITERATURE THAT PATIENTS DIED WITHIN THE 10.3 MONTHS FOLLOWING THEIR RESPECTIVE TRANSCAROTID ARTERY REVASCULARIZATION (TCAR) PROCEDURES. THIS STUDY WAS A RETROSPECTIVE REVIEW OF PATIENTS UNDERGOING EITHER TCAR OR CAROTID ENDARTERECTOMY (CEA) PROCEDURES TO TREAT CAROTID ARTERY STENOSIS. BETWEEN JANUARY 2011 AND JULY 2018, 342 CEAS AND 109 TCARS WERE CAPTURED FOR ANALYSIS. REVIEW OF THE DATA REVEALED THAT OF THE 109 PATIENTS TREATED WITH TCAR, 2.8% OF PATIENTS DIED WITHIN 30 DAYS OF THEIR RESPECTIVE TCAR PROCEDURES. ADDITIONALLY, 4.6% OF PATIENTS DIED WITHIN 10.3 MONTHS OF THEIR RESPECTIVE TCAR PROCEDURES. THE CAUSE OF DEATH (COD) WERE NOT REPORTED FOR ANY CASE. IT WAS NOT REPORTED WHETHER THE DEVICE OR PROCEDURE CAUSED OR CONTRIBUTED TO THE PATIENT DEATHS. Manufacturer Narrative: A2 - AGE AT TIME OF EVENT: THE AVERAGE AGE OF PATIENTS WHO UNDERWENT TCAR WAS 70.7 YEARS. B2 - DATE OF DEATH: AS THE FIRST REPORTED DEATHS OCCURRED WITHIN THE 30 DAYS FOLLOWING TCAR PROCEDURES, THE DEATH DATE WAS ESTIMATED TO 30 DAYS FOLLOWING THE FIRST KNOWN PROCEDURE INCLUDED IN THE STUDY. B3 - DATE OF EVENT: THE EVENT DATE WAS ESTIMATED TO THE EARLIEST KNOWN PROCEDURE INCLUDED IN THE STUDY. D6A - IMPLANT DATE: THE IMPLANT DATE WAS ESTIMATED TO THE EARLIEST KNOWN PROCEDURE INCLUDED IN THE STUDY. H6 - PATIENT CODES: E2401 WAS USED, AS THE COD WAS NOT REPORTED. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER PRODUCT SPECIFIC INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. YEE, E. J., WANG, S. K., TIMSINA, L. R., RUIZ-HERRERA, S., LIAO, J. L., DONDE, N. N., FAJARDO, A. C., & MOTAGANAHALLI, R. L. (2020). PROPENSITY-MATCHED OUTCOMES OF TRANSCAROTID ARTERY REVASCULARIZATION VERSUS CAROTID ENDARTERECTOMY. JOURNAL OF SURGICAL RESEARCH, 252, 22-29. HTTPS://DOI.ORG/10.1016/J.JSS.2019.12.003.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22900500&pc=DRA | 2124215-2025-59954 | 2025-08-04 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-08-26 04:00 | DRA | FARADRIVE | Adverse Event Without Identified Device or Use Problem | Cardiac Tamponade |  |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT EXPERIENCED A CARDIAC TAMPONADE AND PASSED AWAY. DURING A PULSED FIELD ABLATION (PFA) PROCEDURE A FARADRIVE WAS SELECTED FOR USE. AFTER GETTING TRANSSEPTAL ACCESS, THE PHYSICIAN NOTICED THAT THE GUIDEWIRE WAS NOT IN A REGULAR POSITION. THE INTRACARDIAC ECHOCARDIOGRAPHY (ICE) WAS CHECKED AND IT WAS CONFIRMED THAT A CARDIAC TAMPONADE HAD OCCURRED. A CARDIAC SURGERY WAS REQUIRED TO TREAT THE EVENT, BUT THE PATIENT DID NOT MAKE IT AND PASSED AWAY DURING THE NIGHT AT RECOVERY. NO FURTHER INFORMATION WAS REPORTED. Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22900667&pc=PYX | 1220648-2025-45831 | 2025-07-30 04:00 | Death | ABIOMED, INC. | 2025-08-26 04:00 | PYX | IMPELLA RP FLEX | Increase in Pressure | No Clinical Signs, Symptoms or Conditions | P170011 |  | 1 | Event Description: THE COMPLAINANT REPORTED POSTOPERATIVE (CORONARY ARTERY BYPASS GRAFT WITH LEFT ARTERIAL CLIP), THE PATIENT STARTED TO DETERIORATE IN THE INTENSIVE CARE UNIT REQUIRING INCREASED USE OF VASOPRESSORS. THE PATIENT WAS BROUGHT BACK TO THE OPERATING ROOM FOR EXPLORATORY SURGERY AND IMPELLA PLACEMENT (IMPELLA CP PLACED). THE PATIENT DECOMPENSATED OVERNIGHT REQUIRING MAX VASOPRESSORS SUPPORT AND THE DECISION WAS MADE TO TRANSFER PATIENT FOR LEFT VENTRICULAR ASSIST DEVICE (LVAD). THE PATIENT BECAME UNSTABLE NOW REQUIRING IMPELLA 5.5 AND DURING AN ECHOCARDIOGRAM RIGHT VENTRICULAR DYSFUNCTION WAS NOTED. AFTER ESCALATION TO AND INSERTION OF AN IMPELLA 5.5 DEVICE, BLOOD PRESSURE ¿SIGNIFICANTLY¿ IMPROVED. HOWEVER, FEW MINUTES LATER IT ALARMED SUCTION AND THE BLOOD PRESSURE DROPPED. DECISION WAS MADE TO PLACE AN IMPELLA RP FLEX WHICH WAS COMPLETED. THE PATIENT WAS SENT TO THE UNIT ON BIPELLA SUPPORT. SEVEN DAYS INTO THE SUPPORT, THE IMPELLA RP FLEX HAD HIGH PURGE PRESSURE. THE PURGE CASSETTE WAS REPLACED, BUT THE ISSUE PERSISTED. TISSUE PLASMINOGEN ACTIVATOR WAS SUBSEQUENTLY RUN THROUGH THE PURGE TO MANAGE THE ISSUE. SUPPORT CONTINUED FOR THREE MORE DAYS, AT WHICH POINT CARE WAS WITHDRAWN AND THE PATIENT PASSED AWAY. Manufacturer Narrative: PATIENT INFORMATION FOR SECTIONS A4, A5, AND A6 ARE UNKNOWN. DEVICE STATUS: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22901087&pc=DSQ | 2916596-2025-05424 | 2025-08-10 04:00 | Death | THORATEC CORPORATION | 2025-08-26 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Infusion or Flow Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNEXPECTEDLY CODED AND PASSED AWAY. IT WAS NOTED THAT THE PATIENT HAD FLOWS OF 0.0 LITERS PER MINUTE (LPM). LOG FILES WERE SUBMITTED FOR REVIEW AND APPEARED NORMAL UNTIL 19:48:52 ON (B)(6) 2025 WHERE THE FLOW DROPS TO 1.9 LPM. AT 19:49:02 THE FLOW DROPPED TO 0.0 LPM AND REMAINED 0.0 LPM UNTIL THE DRIVELINE WAS DISCONNECTED AT 20:34:13. THE PUMP WAS OPERATING AS INTENDED DURING THESE EVENTS. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22901133&pc=NIM | 2124215-2025-59852 | 2013-01-01 05:00 | Death | SILK ROAD MEDICAL | 2025-08-26 04:00 | NIM | ENROUTE TRANSCAROTID STENT | Adverse Event Without Identified Device or Use Problem | Hemorrhagic Stroke; Insufficient Information |  |  | 1 | Event Description: IT WAS REPORTED VIA LITERATURE ARTICLE THAT POST TRANSCAROTID ARTERY REVASCULARIZATION (TCAR) PROCEDURE THERE WAS A PATIENT MORTALITY RATE WITHIN 1 YEAR OF TCAR PROCEDURES. THIS STUDY WAS A RETROSPECTIVE CHART REVIEW OF PATIENT RECORDS OF 655 TCAR PROCEDURES PERFORMED ON 588 PATIENTS FROM JANUARY 2013 UNTIL JULY 2024. THE OBJECTIVE OF THE STUDY WAS TO EVALUATE THE SAFETY, EFFECTIVENESS, AND EFFICIENCY OF TRANSCAROTID ARTERY REVASCULARIZATION (TCAR) IN A COMMUNITY-BASED PRACTICE BY ANALYZING OUTCOMES AT 30 DAYS AND 1 YEAR. IT WAS IDENTIFIED THAT THERE WERE 2 DEATHS (0.31%) WITHIN 30 DAYS OF TCAR, INCLUDING ONE FOLLOWING HEMORRHAGIC STROKE CONVERSION, AND AN ADDITIONAL 30 DEATHS (4.58%) OCCURRED BETWEEN 30 DAYS AND 1 YEAR, FOR A TOTAL OF 32 DEATHS DURING THE STUDY PERIOD. THE CAUSES OF DEATH WERE NOT REPORTED, AND IT WAS NOT INDICATED WHETHER THE DEVICE OR PROCEDURE CONTRIBUTED TO THESE PATIENT DEATHS. Manufacturer Narrative: B2: THE EXACT DATE OF DEATH WAS NOT PROVIDED FOR ANY OF THE REPORTED CASES AND WAS ESTIMATED BASED ON THE EARLIEST KNOWN TCAR PROCEDURE IN THE COHORT. B3: THE EVENT DATE WAS ESTIMATED TO THE EARLIEST KNOWN TCAR PROCEDURE INCLUDED IN THE COHORT. H6: PATIENT CODE E2401 WAS USED AS THE CAUSE OF DEATH WAS NOT REPORTED FOR THE 30 DEATHS OCCURRING BETWEEN 30 DAYS AND 1 YEAR. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BOSTON SCIENTIFIC. BASED ON THE NATURE OF THE INFORMATION RECEIVED, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE DEVICE IDENTIFIER (UDI) OR OTHER PRODUCT-SPECIFIC DETAILS. IF ADDITIONAL INFORMATION BECOMES AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. KENNY, M., LANDAVAZO, B., VERNON, C., YELOVITCH, S., ZEA, N., NATION, D., APPLE, J., QUAYE, K., BOONE, B., & TURLEY, R. (2025). OUTCOMES AND INSIGHTS FROM A DECADE OF TRANSCAROTID ARTERY REVASCULARIZATION IN COMMUNITY PRACTICE. JOURNAL OF VASCULAR SURGERY. HTTPS://DOI.ORG/10.1016/J.JVS.2025.04.064.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22901181&pc=NIM | 2124215-2025-59868 | 2022-05-13 04:00 | Death | SILK ROAD MEDICAL | 2025-08-26 04:00 | NIM | ENROUTE TRANSCAROTID STENT | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: IT WAS REPORTED VIA LITERATURE THAT ONE PATIENT DIED WITHIN 30 DAYS OF THE TRANSCAROTID ARTERY REVASCULARIZATION (TCAR) PROCEDURES. THE AIM OF THIS STUDY WAS TO PROVIDE REAL-WORLD DATA COMPARING THE SAFETY AND EFFECTIVENESS ASSOCIATED WITH THE PERFORMANCE OF CAROTID REVASCULARIZATION TECHNIQUES (CAROTID ENDARTERECTOMY (CEA), CAROTID ARTERY STENTING (CAS), OR TRANSCAROTID ARTERY REVASCULARIZATION (TCAR)) BY DUAL-TRAINED NEUROSURGEONS. A TOTAL OF 780 PATIENTS WERE INCLUDED IN THIS ANALYSIS, 583 WITH CAS PROCEDURES, 165 WITH CEA PROCEDURES AND 32 WITH TCAR PROCEDURES. REVIEW OF THE DATA REVEALED THAT OF THE 32 PATIENTS TREATED WITH TCAR, 1 PATIENT DIED WITHIN 30 DAYS OF THEIR RESPECTIVE TCAR PROCEDURE. THE CAUSE OF DEATH (COD) WAS NOT REPORTED. IT WAS NOT REPORTED WHETHER THE DEVICE OR PROCEDURE CAUSED OR CONTRIBUTED TO THE PATIENT DEATH. Manufacturer Narrative: B2 - THE DATE OF DEATH WAS NOT PROVIDED. B3 AND D6A. EVENT DATE/IMPLANT DATE WAS NOT PROVIDED, THEREFORE DATE OF PUBLICATION WAS USED. H6 - PATIENT CODES: E2401 WAS USED, AS THE COD WAS NOT REPORTED. DETAILS OF THE EVENT, INCLUDING PATIENT INFORMATION AND PRODUCT INFORMATION, WERE NOT PROVIDED TO BSC. BECAUSE THE PRODUCT LOT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) NUMBER AND OTHER SPECIFIC PRODUCT INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. LEVY, B. R., WAQAS, M., MONTEIRO, A., CAPPUZZO, J. M., BAIG, A. A., KHAWAR, W. I., DAVIES, J. M., SNYDER, K. V., SIDDIQUI, A. H., RIINA, H. A., & LEVY, E. I. (2023). NOT A TRIFECTA: COMPLEMENTARY USE OF CAROTID ARTERY REVASCULARIZATION TECHNIQUES IN THE ERA OF HYBRID NEUROSURGERY. JOURNAL OF NEUROSURGERY, 138(1), 199 TO 204. HTTPS://DOI.ORG/10.3171/2022.4.JNS22420.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22882540&pc=DQY | MW5175143 | 2024-11-07 05:00 | Death | ABBOTT MEDICAL | 2025-08-25 04:00 | DQY | ABBOTT AMPLATZER TORQVUE 45 X45 DELIVERY | Adverse Event Without Identified Device or Use Problem | Bradycardia; Intracranial Hemorrhage; Respiratory Failure |  |  | 1 | Event Description: TRANSFERRED FROM (B)(6) 11/7 AFTER PRESENTING WITH STROKE SYMPTOMS. PMH CABG, BIOPROSTHETIC AVR, CEREBRAL ANEURYSM AND (AFIB) ATRIAL FIBRILLATION (NO AC DUE TO FALL RISK). CT NEGATIVE FOR LARGE VESSEL OCCLUSION. RECEIVED TENECTEPLASE 0.25 MG/KG PRIOR TO TRANSFER DEACONESS HENDERSON. WITHIN 8 HOURS OF ADMINISTRATION, (B)(6) BECAME BRADYCARDIC WITH WORSENING (NIH SCORE) NATIONAL INSTITUTES OF HEALTH STROKE SCALE SCORE. REPEAT CT SHOWED NEW (ICH) INTRACEREBRAL HEMORRHAGE. (TXA) TRANEXAMIC ACID BOLUS WAS GIVEN AND (B)(6) WAS TRANSFERRED TO (B)(6) ON 11/9. FIBRINOGEN LEVEL 241 UPON ARRIVAL AND NO FURTHER TXA WAS GIVEN. NICARDIPINE DRIP WAS ORDERED FOR (BP) BLOOD PRESSURE CONTROL. CLOPIDOGREL STARTED 11/9. (B)(6) MAINTAINED ABILITY TO SPEAK, FOLLOW COMMANDS UNTIL BEGAN TO DEVELOP RESPIRATORY FAILURE ON 11/10. (B)(6) REQUIRED INTUBATION ON 11/11. DESPITE ATTEMPTS MADE TO OPTIMIZE CLINICAL CONDITION, (B)(6) CONTINUED TO FAIL (SBT) SPONTANEOUS BREATHING TRIAL. (B)(6) INDICATED SHE DID NOT WANT A TRACH SO TERMINAL EXTUBATING WAS DONE 11/22 AND (B)(6) EXPIRED SHORTLY AFTER. PATIENT CODES: 1891, 1751, 2484. REFERENCE REPORT #MW5175144, #MW5175145, #MW5175146.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22882546&pc=DQX | MW5175144 | 2024-11-07 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-08-25 04:00 | DQX | WIRE AMPLATZ ES 180 CM J 0.035 | Adverse Event Without Identified Device or Use Problem | Bradycardia; Intracranial Hemorrhage; Respiratory Failure |  |  | 1 | Event Description: TRANSFERRED FROM (B)(6) 11/7 AFTER PRESENTING WITH STROKE SYMPTOMS. PMH CABG, BIOPROSTHETIC AVR, CEREBRAL ANEURYSM AND (AFIB) ATRIAL FIBRILLATION (NO AC DUE TO FALL RISK). CT NEGATIVE FOR LARGE VESSEL OCCLUSION. RECEIVED TENECTEPLASE 0.25 MG/KG PRIOR TO TRANSFER (B)(6). WITHIN 8 HOURS OF ADMINISTRATION, (B)(6) BECAME BRADYCARDIC WITH WORSENING (NIH SCORE) NATIONAL INSTITUTES OF HEALTH STROKE SCALE SCORE. REPEAT CT SHOWED NEW (ICH) INTRACEREBRAL HEMORRHAGE. (TXA) TRANEXAMIC ACID BOLUS WAS GIVEN AND (B)(6) WAS TRANSFERRED TO (B)(6) ON 11/9. FIBRINOGEN LEVEL 241 UPON ARRIVAL AND NO FURTHER TXA WAS GIVEN. NICARDIPINE DRIP WAS ORDERED FOR (BP) BLOOD PRESSURE CONTROL. CLOPIDOGREL STARTED 11/9. (B)(6) MAINTAINED ABILITY TO SPEAK, FOLLOW COMMANDS UNTIL BEGAN TO DEVELOP RESPIRATORY FAILURE ON 11/10. (B)(6) REQUIRED INTUBATION ON 11/11. DESPITE ATTEMPTS MADE TO OPTIMIZE CLINICAL CONDITION, (B)(6) CONTINUED TO FAIL (SBT) SPONTANEOUS BREATHING TRIAL. (B)(6) INDICATED SHE DID NOT WANT A TRACH SO TERMINAL EXTUBATING WAS DONE 11/22 AND (B)(6) EXPIRED SHORTLY AFTER. PATIENT CODES: 1891, 1751, 2484. REFERENCE REPORT #MW5175143, #MW5175145, #MW5175146.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22882549&pc=OES | MW5175145 | 2024-11-07 05:00 | Death | UNKNOWN | 2025-08-25 04:00 | OES | UNSURE SOURCE OF "FUZZY FIBERS" CARDIAC CATH PACK | Adverse Event Without Identified Device or Use Problem | Bradycardia; Intracranial Hemorrhage; Respiratory Failure |  |  | 1 | Event Description: TRANSFERRED FROM OSH (B)(6) AFTER PRESENTING WITH STROKE SYMPTOMS. PMH CABG, BIOPROSTHETIC AVR, CEREBRAL ANEURYSM AND (AFIB) ATRIAL FIBRILLATION (NO AC DUE TO FALL RISK). CT NEGATIVE FOR LARGE VESSEL OCCLUSION. RECEIVED TENECTEPLASE 0.25 MG/KG PRIOR TO TRANSFER DEACONESS HENDERSON. WITHIN 8 HOURS OF ADMINISTRATION, SA BECAME BRADYCARDIC WITH WORSENING (NIH SCORE) NATIONAL INSTITUTES OF HEALTH STROKE SCALE SCORE. REPEAT CT SHOWED NEW (ICH) INTRACEREBRAL HEMORRHAGE. (TXA) TRANEXAMIC ACID BOLUS WAS GIVEN AND SA WAS TRANSFERRED TO (B)(6) ON (B)(6). FIBRINOGEN LEVEL 241 UPON ARRIVAL AND NO FURTHER TXA WAS GIVEN. NICARDIPINE DRIP WAS ORDERED FOR (BP) BLOOD PRESSURE CONTROL. CLOPIDOGREL STARTED (B)(6). SA MAINTAINED ABILITY TO SPEAK, FOLLOW COMMANDS UNTIL BEGAN TO DEVELOP RESPIRATORY FAILURE ON (B)(6). SA REQUIRED INTUBATION ON (B)(6). DESPITE ATTEMPTS MADE TO OPTIMIZE CLINICAL CONDITION, SA CONTINUED TO FAIL (SBT) SPONTANEOUS BREATHING TRIAL. SA INDICATED SHE DID NOT WANT A TRACH SO TERMINAL EXTUBATING WAS DONE (B)(6) AND SA EXPIRED SHORTLY AFTER. PATIENT CODES: 1891, 1751, 2484. REFERENCE REPORT #MW5175143, #MW5175144, #MW5175146.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22882566&pc=NAB | MW5175146 | 2024-11-07 05:00 | Death | UNKNOWN | 2025-08-25 04:00 | NAB | GAUZE SPONGES 4 X 4 12-PLY | Adverse Event Without Identified Device or Use Problem | Bradycardia; Intracranial Hemorrhage; Respiratory Failure |  |  | 1 | Event Description: TRANSFERRED FROM OSH (B)(6) AFTER PRESENTING WITH STROKE SYMPTOMS. PMH CABG, BIOPROSTHETIC AVR, CEREBRAL ANEURYSM AND (AFIB) ATRIAL FIBRILLATION (NO AC DUE TO FALL RISK). CT NEGATIVE FOR LARGE VESSEL OCCLUSION. RECEIVED TENECTEPLASE 0.25 MG/KG PRIOR TO TRANSFER DEACONESS HENDERSON. WITHIN 8 HOURS OF ADMINISTRATION, SA BECAME BRADYCARDIC WITH WORSENING (NIH SCORE) NATIONAL INSTITUTES OF HEALTH STROKE SCALE SCORE. REPEAT CT SHOWED NEW (ICH) INTRACEREBRAL HEMORRHAGE. (TXA) TRANEXAMIC ACID BOLUS WAS GIVEN AND SA WAS TRANSFERRED TO GW ON (B)(6). FIBRINOGEN LEVEL 241 UPON ARRIVAL AND NO FURTHER TXA WAS GIVEN. NICARDIPINE DRIP WAS ORDERED FOR (BP) BLOOD PRESSURE CONTROL. CLOPIDOGREL STARTED (B)(6). SA MAINTAINED ABILITY TO SPEAK, FOLLOW COMMANDS UNTIL BEGAN TO DEVELOP RESPIRATORY FAILURE ON (B)(6). SA REQUIRED INTUBATION ON (B)(6). DESPITE ATTEMPTS MADE TO OPTIMIZE CLINICAL CONDITION, SA CONTINUED TO FAIL (SBT) SPONTANEOUS BREATHING TRIAL. SA INDICATED SHE DID NOT WANT A TRACH SO TERMINAL EXTUBATING WAS DONE (B)(6) AND SA EXPIRED SHORTLY AFTER. PATIENT CODES: 1891, 1751, 2484. REFERENCE REPORT #MW5175143, #MW5175144, #MW5175145.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22883359&pc=DSQ | 2916596-2025-05605 | 2024-02-09 05:00 | Death | THORATEC CORPORATION | 2025-08-25 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Septic Shock | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO SEPTIC SHOCK, THE SOURCE WAS UNKNOWN. THE FAMILY CHOSE TO MAKE THE PATIENT COMFORT CARE. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED, AND THE PUMP WOULD NOT RETURN. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF INFECTION COULD NOT BE CONCLUSIVELY DETERMINED. INFECTION IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. INFECTION HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22883443&pc=DSQ | 2916596-2025-05412 | 2025-08-03 04:00 | Death | THORATEC CORPORATION | 2025-08-25 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN AT THE TIME OF THE EVENT AND IT WAS NOT DEVICE RELATED. IT WAS ALSO REPORTED THAT THE DEVICE PARAMETERS WERE WITHIN NORMAL LIMITS. THE DEVICE WAS NOT EXPLANTED AND WOULD NOT BE RETURNED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED BY THE ACCOUNT. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. D IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22883473&pc=NAY | 2955842-2025-35521 | 2024-05-23 04:00 | Death | INTUITIVE SURGICAL, INC | 2025-08-25 04:00 | NAY | DAVINCI X | Adverse Event Without Identified Device or Use Problem | Disseminated Intravascular Coagulation (DIC); Hemorrhage/Bleeding | K171294 |  | 1 | Event Description: A RETROSPECTIVE STUDY WAS SUMMARIZED IN A LITERATURE ARTICLE. THE STUDY ANALYZED 186 PATIENTS UNDERGOING ROBOTIC SURGERIES ACROSS 7 SURGICAL SPECIALTIES IN THE FIRST ROBOTIC SURGERY CENTER IN A NATIONAL PUBLIC HOSPITAL BETWEEN OCTOBER 2022 AND OCTOBER 2023. THE STUDY'S INTENT WAS TO IDENTIFY THE PATIENT POPULATIONS, EVALUATE THE COSTS OF THE SURGERIES AND REPORT THE SURGICAL OUTCOMES. ONE PATIENT UNDERWENT AN UNSPECIFIED CARDIAC SURGERY AND EXPERIENCED AN INTRA-OPERATIVE HEMORRHAGE DUE TO AN UNSPECIFIED VASCULAR INJURY WHICH REQUIRED CONVERSION TO OPEN; THEN IN THE IMMEDIATE POST-OPERATIVE PERIOD DEVELOPED DISSEMINATED INTRAVASCULAR COAGULATION (DIC) RESULTING IN DEATH. NO DEVICE MALFUNCTIONS WERE REPORTED, AND THE AUTHORS DID NOT ATTRIBUTE ANY EVENTS TO ANY DA VINCI DEVICES. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: NO INVESTIGATION COULD BE PERFORMED AS THE ARTICLE DID NOT PROVIDE ANY SPECIFIC INFORMATION REGARDING THE PROCEDURE DATE OR DA VINCI SYSTEM IDENTIFIERS. THERE WAS NO REPORT OF, OR INDICATION OF, ANY DA VINCI PRODUCT ISSUES. A REQUEST FOR ADDITIONAL EVENT INFORMATION WAS SUBMITTED TO THE CORRESPONDING AUTHOR. THE RESPONSE RECEIVED PROVIDED AN EXPLANATION THAT COLLECTING THE INFORMATION REQUIRES TIME AND RESOURCES, AND THE INFORMATION IS COVERED BY STRICT CONFIDENTIAL AND ETHICAL REGULATIONS; THEREFORE, NO INFORMATION WILL BE PROVIDED. A REVIEW OF THE EVENT WAS CONDUCTED BY AN ISI MEDICAL OFFICER AND CONCLUDED THAT ONE PATIENT IN THIS SINGLE HOSPITAL LITERATURE REVIEW OF 186 PATIENTS DIED AS A RESULT OF COMPLICATIONS FROM AN INTRAOPERATIVE BLEEDING EVENT DURING A CARDIAC PROCEDURE. THE SPECIFIC PROCEDURE AND HOW THIS INJURY MAY HAVE OCCURRED WERE NOT PROVIDED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS PROVIDED TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCT OR INSTRUMENT CONTRIBUTED TO THIS EVENT. CITATION: HUGUET, F., ACUNA, B., LOPEZ, M., ET AL. 2024. TOWARDS THE FUTURE OF SURGERY: DESCRIPTIVE ANALYSIS OF THE IMPLEMENTATION OF THE FIRST ROBOTIC SURGERY CENTER IN A NATIONAL PUBLIC HOSPITAL. REV CIR. 2025; 77(1). DOI: HTTP://DX.DOI.ORG/10.35687/S2452-454920250012339. SECTION B3 - THERE IS INSUFFICIENT INFORMATION REGARDING THE PROCEDURE DATE; THEREFORE, THE ARTICLE PUBLISH DATE WAS USED. SECTION E1 - THIS ARTICLE WAS IDENTIFIED DURING A LITERATURE REVIEW BY INTUITIVE; THE CORRESPONDING AUTHOR IS DESIGNATED AS THE REPORTER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22884066&pc=QAD | 3006845464-2025-00010 | 2023-12-20 05:00 | Death | ILLUMINOSS MEDICAL INC | 2025-08-25 04:00 | QAD | PHOTODYNAMIC BONE STABILIZATION SYSTEM | Adverse Event Without Identified Device or Use Problem | Pulmonary Embolism | K181228 |  | 1 | Event Description: AN ARTICLE BY MINA ET AL., 2025 "FATAL COMPLICATIONS DURING PHOTODYNAMIC BONE STABILIZATION: A CASE REPORT" DETAILS COMPLICATIONS EXPERIENCED DURING A HUMERUS IMPLANTATION PROCEDURE. A 68-YEAR-OLD FEMALE WITH A HISTORY OF METASTATIC SARCOMA UNDERWENT PROPHYLACTIC STABILIZATION OF THE RIGHT HUMERUS WITH AN ILLUMINOSS IMPLANT ON (B)(6) 2023. THE PATIENT'S COMORBIDITIES INCLUDE METASTATIC INVOLVEMENT OF THE LUNG AND REMOTE HISTORY OF DEEP VEIN THROMBOSIS AND PULMONARY EMBOLISM AND USED SUPPLEMENT OXYGEN. DURING THE BONE REAMING AND SUBSEQUENT BALLOON INFLATION STEPS, THE PATIENT EXPERIENCED A SUDDEN CARDIOPULMONARY COLLAPSE CHARACTERIZED BY PULSELESS ELECTRICAL ACTIVITY, HYPOXIA AND HYPOTENSION. ADVANCED CARDIAC LIFE SUPPORT WAS PROVIDED, AND A TRANSESOPHAGEAL ECHOCARDIOGRAPHY WAS PERFORMED SHOWING A SEVERELY DILATED RA AND RV WITH AN UNDERFILLED, HYPERDYNAMIC LV, SUGGESTIVE OF A MASSIVE PE. DESPITE RESUSCITATION EFFORTS, THE PATIENT'S CONDITION CONTINUED TO DETERIORATE, AND SEVERAL DAYS POST OPERATIVELY THE PATIENT PASSED AWAY. THE AUTHORS SPECULATE THE CAUSE OF THIS COMPLICATION WAS LIKELY DUE TO THE BALLOON INSUFFLATION-INDUCED DISLODGEMENT OF MARROW AND SUBSEQUENT EMBOLIZATION. THE AUTHORS ALSO NOTE THE PATIENT WAS MORE SUSCEPTIBLE TO ADVERSE EFFECTS OF AN INTRAOPERATIVE PE DUE TO THE PATIENT BEING ELDERLY, HAVING A HISTORY OF METASTATIC CANCER WITH SIGNIFICANT PULMONARY INVOLVEMENT, AND SIGNS OF PULMONARY HYPERTENSION AND RIGHT HEART STRAIN, WHICH REDUCED THE PATIENT'S ABILITY TO TOLERATE AN EMBOLIC LOAD, LEADING TO RV DYSFUNCTION. Manufacturer Narrative: THE INTRAOPERATIVE PATIENT COLLAPSE AND SUBSEQUENT DEATH WERE DETERMINED TO BE DUE TO A PULMONARY EMBOLISM (PE). THIS CONCLUSION WAS REACHED DURING A MEDICAL OVERSIGHT REVIEW, IN WHICH THE MEDICAL REVIEWER ASSESSED A PUBLISHED JOURNAL ARTICLE DESCRIBING THE CASE. THE ARTICLE REPORTED THAT TRANSESOPHAGEAL ECHOCARDIOGRAPHY PERFORMED DURING RESUSCITATION SUGGESTED A MASSIVE PE. THROMBOLYTIC THERAPY WAS NOT ADMINISTERED BECAUSE THE TREATING TEAM CONSIDERED DEBRIS OR FAT EMBOLISM MORE LIKELY THAN PE SECONDARY TO DEEP VEIN THROMBOSIS (DVT). THE MEDICAL REVIEWER NOTED THAT THIS DISCUSSION APPEARED TO ATTRIBUTE THE EVENT TO TWO POSSIBLE EMBOLIC SOURCES[?]PE FROM DVT AND DEBRIS OR FAT EMBOLISM. THE REVIEWER ALSO OBSERVED THAT THE ARTICLE DID NOT SPECIFY THE TIME INTERVAL BETWEEN BALLOON INFLATION AND PATIENT COLLAPSE, WHICH WOULD HAVE BEEN USEFUL IN DETERMINING CAUSALITY. ADDITIONALLY, THE PATIENT'S HISTORY OF DVT OR PE WAS NOT PROVIDED, LIMITING THE ABILITY TO FULLY ASSESS THE UNDERLYING RISK. BOTH THE JOURNAL ARTICLE AND THE MEDICAL REVIEWER IDENTIFIED THAT THE PATIENT'S SIGNIFICANT COMORBIDITIES AND OVERALL REDUCED TOLERANCE FOR AN EMBOLIC LOAD WERE IMPORTANT CONTRIBUTING FACTORS TO THE FATAL EVENT. THIS ADVERSE EVENT IS BEING SUBMITTED AS THE DIRECT CAUSE OF THE EVENT CANNOT BE DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22884078&pc=DQY | 2124215-2025-59208 | 2025-07-29 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-08-25 04:00 | DQY | WATCHMAN TRUSTEER? ACCESS SYSTEM | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Pericardial Effusion | K240018 |  | 1 | Event Description: IT WAS REPORTED THAT PERICARDIAL EFFUSION (PE) AND DEATH OCCURRED. DURING A CONCOMITANT PULSE FIELD ABLATION (PFA) PROCEDURE, A VERSACROSS LARGE ACCESS (VLA) WAS SELECTED FOR USE. A BASELINE PE WAS NOTED PRIOR TO TRANSSEPTAL. ONCE GROIN ACCESS WAS OBTAINED, THERE WAS DIFFICULTY ASCENDING THE FEMORAL VEIN AND CROSSING TRANSEPTAL. TRANSEPTAL PUNCTURE (TSP) WAS SUDDEN, AND VISIBILITY WAS POOR. THE PHYSICIAN FLOSSED THE SEPTUM WITH THE VLA, HOWEVER, THE FARADRIVE SHEATH COULD NOT GO ACROSS. IT WAS ATTEMPTED WITH THE FARADRIVE DILATOR AND WATCHMAN TRUSTEER ACCESS SYSTEM (WAS) DILATOR, THEN WAS ABLE TO CROSS WITH FARADRIVE. THE PATIENTS' BLOOD PRESSURE DROPPED A FEW TIMES AT TRANSEPTAL DURING FLOSSING OF THE SEPTUM. THE HYPOTENSION WAS MEDICALLY TREATED. THE PHYSICIAN USED INTRACARDIAC ECHOCARDIOGRAPHY (ICE) TO WATCH FOR A PE DURING THE PFA PORTION OF THE PROCEDURE. IT STARTED GROWING WHEN ABLATING THE RIGHT PULMONARY VEINS. A PERICARDIOCENTESIS WAS PERFORMED AFTER THE COMPLETING PFA. ONE (1) LITER OF FLUID WAS DRAWN, AND PATIENT WAS STABILIZED. LATER, THE PATIENT DIED WHILE IN THE INTENSIVE CARE UNIT (ICU) FROM COMPLICATIONS. Manufacturer Narrative: B2 DATE OF DEATH REMAINS UNKNOWN AFTER GOOD FAITH EFFORTS EXHAUSTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22884093&pc=DSQ | 3007042319-2025-01942 | 2025-07-15 04:00 | Death | HEARTWARE, INC. | 2025-08-25 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT SUBSEQUENTLY PASSED AWAY PEACEFULLY DUE TO CLINICAL ISSUES. CAUSE OF DEATH UNSPECIFIED. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22884292&pc=DSQ | 2916596-2025-05599 | 2025-08-12 04:00 | Death | THORATEC CORPORATION | 2025-08-25 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. Manufacturer Narrative: SECTION A3: PATIENT SEX/GENDER WAS NOT PROVIDED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. DUE TO HOSPITAL PRIVACY POLICY, NO ADDITIONAL INFORMATION WILL BE PROVIDED BY THE ACCOUNT. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22884441&pc=LKN | 1722028-2025-00195 | 2017-09-10 04:00 | Death | TERUMO BCT | 2025-08-25 04:00 | LKN | SPECTRA OPTIA | Insufficient Information | Fatigue; Muscle Weakness; Ptosis | K183081 |  | 1 | Event Description: PER JOURNAL ARTICLE "COMPARING LEUKAPHERESIS PROTOCOLS FOR AN AML PATIENT WITH SYMPTOMATIC LEUKOSTASIS. JOURNAL OF CLINICAL APHERESIS" BY CLINE, A., JAJOSKY, R., SHIKLE, J., & BOLLAG, R. BACKGROUND: ACUTE MYELOID LEUKEMIA (AML) IS A MALIGNANCY CHARACTERIZED BY RAPID CLONAL PROLIFERATION OF MYELOID PRECURSORS, WHICH CAN RESULT IN HYPERLEUKOCYTOSIS. LEUKAPHERESIS CAN BE USED TO RAPIDLY REDUCE THE WHITE BLOOD CELL COUNT (WBC). HOWEVER, THE ONLY FDA CLEARED DEVICE FOR WBC DEPLETION, THE COBE SPECTRA, WILL NO LONGER BE SUPPORTED BY THE MANUFACTURER IN 2017, AND THERE ARE FEW STUDIES COMPARING DIFFERENT METHODS OF LEUKAPHERESIS. CASE REPORT: A 68-YEAR-OLD AFRICAN AMERICAN FEMALE WAS ADMITTED TO THE HOSPITAL FOR RELAPSE OF HER AML. LABORATORY DATA DEMONSTRATED A WBC COUNT OF 291 600/LL AND FLOW CYTOMETRY OF THE PERIPHERAL BLOOD DEMONSTRATED 85% MYELOID BLASTS. LEUKAPHERESIS WAS ORDERED TO HELP TREAT THE LEUKOSTASIS. METHODS: THREE DIFFERENT APHERESIS PROTOCOLS WERE USED TO ACHIEVE CYTOREDUCTION: SPECTRA OPTIA MONONUCLEAR COLLECTION (MNC) PROTOCOL, SPECTRA OPTIA GRANULOCYTE COLLECTION (PMN) PROTOCOL, AND THERAKOS CELLEX BUFFY COAT COLLECTION WITHOUT RETURN. DUE TO DIFFERENT INLET FLOW RATES, THE PROCEDURES WERE EVALUATED BASED ON THE NUMBER OF WBCS COLLECTED AND VOLUME OF BLOOD PROCESSED (VBP). RESULTS: THE SPECTRA OPTIA PMN COLLECTED THE MOST WBCS AND COLLECTED NEARLY AS MANY WBCS PER VBP AS THE THERAKOS CELLEX, WHICH HAD THE HIGHEST VALUE. CONCLUSION: TO OUR KNOWLEDGE, WE ARE REPORTING THE FIRST USE OF THERAKOS CELLEX AND SPECTRA OPTIA PMN PROTOCOL FOR WBC DEPLETION. WHILE THE SPECTRA OPTIA GRANULOCYTE PROTOCOL SHOWED THE BEST PERFORMANCE FOR THIS AML PATIENT, FURTHER STUDIES WILL BE NEEDED TO COMPARE THE SPECTRA OPTIA PMN PROTOCOL TO THE MNC PROTOCOL FOR AML PATIENTS. THE LEUKAPHERESIS WAS TERMINATED EARLY DUE TO A MIDPROCEDURE HGB OF 5.6 G/DL, DECREASE FROM 7.8 G/DL PRIOR TO THE PROCEDURE. THE HGB DECREASE WAS ATTRIBUTED TO HEMODILUTION DUE TO THE TRANSFUSION OF 5 UNITS OF FRESH FROZEN PLASMA (FFP) TO REVERSE WARFARIN ANTICOAGULATION (FOR A MECHANICAL HEART VALVE) IN ORDER TO PLACE THE FEMORAL LINE. THE VBP WAS 5.5 L (RUNTIME 145 MINUTES) AND 80 ML WERE COLLECTED. THE PATIENT¿S WBC COUNT DECREASED FROM 292 K/LL TO 203 K/LL, WHILE THE BLAST CELL COUNT DECREASED FROM 242 TO 183 K/LL (FIGURE 1). THE PLT COUNT DROPPED FROM 45 TO 24 K, SO ONE UNIT OF APHERESIS PLTS WAS TRANSFUSED. BECAUSE THE PROCEDURE TERMINATED EARLY, THE NUMBER OF WBCS COLLECTED IN THE BAG WAS NOT DETERMINED. IN ADDITION, BECAUSE THE HEMATOCRIT VALUE INPUT INTO THE MACHINE WAS SUBSTANTIALLY DIFFERENT FROM THE ACTUAL HEMATOCRIT, THIS LEUKAPHERESIS PROCEDURE WAS NOT COMPARED TO THE OTHER PROCEDURES. ON THE SUBSEQUENT DAY, THE FOURTH LEUKAPHERESIS WAS PERFORMED USING THE SPECTRA OPTIA PMN. THE PROCEDURE WAS PERFORMED WITH AN INLET FLOW RATE OF 47 ML/MINUTE AND INLET/ AC RATIO OF 13. THE VBP WAS 6.8 L (RUNTIME 147 MINUTES) AND 498 ML WERE COLLECTED WITHOUT COMPLICATIONS. THE LEUKAPHERESIS BAG CONTAINED 2.39 3 1011 WBCS. THE WBC COUNT DECREASED FROM 132 TO 102 K/LL (22% DECREASE), AND THE BLAST CELL COUNT DECREASED FROM 111 TO 71 K/LL (36% DECREASE). AFTER THE PROCEDURE, PLT COUNT WAS 23 K. A FIFTH LEUKAPHERESIS PROCEDURE WAS SCHEDULED FOR THE NEXT DAY. HOWEVER, THE PATIENT WAS NOTED TO HAVE UNEQUAL PUPILS WITH A LARGE, NON-REACTIVE PUPIL IN THE LEFT EYE AND LEFT SIDED PTOSIS. A CT SCAN OF THE HEAD SHOWED ACUTE ON CHRONIC SUBDURAL HEMATOMA WITH MILD LEFT DESCENDING TRANSTENTORIAL HERNIATION AND MIDLINE SHIFT, WHICH WOULD REQUIRE NEUROSURGERY. BECAUSE OF THE BLEEDING AND A PARTIAL THROMBOPLASTIN TIME (PTT) OF 81.9 SECONDS (REFERENCE RANGE, 28.0¿35.7) HEPARIN, WHICH WAS USED AS ANTICOAGULATION FOR THE MECHANICAL HEART VALVE, WAS DISCONTINUED. IN ADDITION, ONE UNIT OF RBCS AND TWO UNITS OF APHERESIS PLTS WERE TRANSFUSED. SOON AFTER, THE FAMILY REQUESTED WITHDRAWAL OF CARE. THE PATIENT WAS EXTUBATED AND SOON AFTER DIED. SOME ADVERSE EVENTS ARE REPORTED DURING LEUKAPHERESIS. AML PATIENTS CAN EXPERIENCE THROMBOCYTOPENIA AND BLEEDING, AS A SIGNIFICANT NUMBER OF PLATELETS ARE OFTEN REMOVED.11 A RECENT STUDY SUGGESTED THAT CLOSE MONITORING OF PLATELETS IS NEEDED TO AVOID BLEEDING COMPLICATIONS IN PATIENTS WHO ARE ALREADY THROMBOCYTOPENIC.8 SPECIFICALLY, THE SPECTRA OPTIA IS KNOWN FOR PLATELET LOSS, THUS EXPLAINING THE NEED FOR PLATELET TRANSFUSIONS IN THE PATIENT DETAILED IN THIS REPORT. ADDITIONALLY, VENOUS THROMBOSIS OR HEMORRHAGE MAY OCCUR AS A CONSEQUENCE OF THE LARGE BORE CATHETER INSERTION.12 NEVERTHELESS, THE PROCEDURE IS SAFELY PERFORMED IN THE MAJORITY OF CASES.13 THE PATIENT IN THIS REPORT INCREASINGLY RELIED ON TRANSFUSIONS AS THE HOSPITALIZATION PROGRESSED. HOWEVER, NO ADVERSE EFFECTS WERE DIRECTLY LINKED TO THE LEUKAPHERESIS PROCEDURES. THE SUBDURAL HEMATOMA WAS THOUGHT TO BE DUE TO AML AND HEPARIN ANTICOAGULATION. THERE WERE MULTIPLE PROCEDURES OVER MULTIPLE DAYS AND THE PATIENT OUTCOME WAS DEATH. THIS REPORT IS BEING FILED DUE TO PATIENT DEATH, ALTHOUGH PER CURRENT INFORMATION THERE IS NO DETECTABLE MALFUNCTION WITH THE TERUMO BCT DEVICE OR ALLEGATION OF A MALFUNCTION. THE COLLECTION SET IS NOT AVAILABLE FOR RETURN BECAUSE IT WAS DISCARDED BY THE CUSTOMER. Manufacturer Narrative: LOT NUMBER, MANUFACTURE DATE AND EXPIRY DATE ARE NOT AVAILABLE AT THIS TIME. INVESTIGATION IS IN PROCESS, A FOLLOW-UP REPORT WILL BE PROVIDED. CITATION: CLINE, A., JAJOSKY, R., SHIKLE, J., & BOLLAG, R. (2017). COMPARING LEUKAPHERESIS PROTOCOLS FOR AN AML PATIENT WITH SYMPTOMATIC LEUKOSTASIS. JOURNAL OF CLINICAL APHERESIS, 33(3), 396¿400. HTTPS://DOI.ORG/10.1002/JCA.21588.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22884540&pc=DSQ | 3007042319-2025-01943 | 2025-06-20 04:00 | Death | HEARTWARE, INC. | 2025-08-25 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT SUBSEQUENTLY PASSED AWAY PEACEFULLY DUE TO CLINICAL ISSUES. CAUSE OF DEATH UNSPECIFIED. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22884703&pc=MGB | 2024168-2025-08962 | 2009-01-01 05:00 | Death | ABBOTT VASCULAR INC. | 2025-08-25 04:00 | MGB | PROSTAR | Adverse Event Without Identified Device or Use Problem | Unspecified Vascular Problem | P960043 |  | 1 | Event Description: THE AIM OF THIS SYSTEMATIC REVIEW AND META-ANALYSIS WAS TO COMPARE MANTA A PLUG-BASED VASCULAR CLOSURE DEVICE (P-VCD) WITH PROSTAR, PROGLIDE AND PROSTYLE SUTURED-BASED VASCULAR CLOSURE DEVICE (S-VCD) IN PATIENTS UNDERGOING AN ENDOVASCULAR OR CARDIAC TREATMENT VIA A LARGE-BORE ARTERIAL ACCESS. IN CONCLUSION, THE SYSTEMATIC REVIEW AND META-ANALYSIS, THE LARGEST TO DATE, DEMONSTRATED NO SIGNIFICANT DIFFERENCES BETWEEN P-VCD AND S-VCD, REGARDING SAFETY OR EFFICACY OUTCOMES, EXCEPT FOR A SHORTER HOSPITAL STAY ASSOCIATED WITH P-VCD. COMPLICATIONS ASSOCIATED WITH PROSTAR INCLUDED DEVICE FAILURE, FAILURE TO ACHIEVE HEMOSTASIS, IN-HOSPITAL ALL-CAUSE MORTALITY, LIFE-THREATENING OR MAJOR LIFE-THREATENING BLEEDING, MAJOR AND MINOR BLEEDING, MAJOR AND MINOR VASCULAR COMPLICATIONS, PSEUDOANEURYSM, UNPLANNED ENDOVASCULAR TREATMENT OF VASCULAR ACCESS, COVERED STENT IMPLANTATION TREATMENT OF VASCULAR ACCESS, UNPLANNED VASCULAR SURGERY, BLOOD TRANSFUSION, PROLONGED HOSPITALIZATION, ARTERIOVENOUS FISTULA, ARTERIAL DISSECTION, ARTERIAL STENOSIS, HOMOLATERAL LOWER LIMB ISCHEMIA, ACCESS SITE RELATED VASCULAR INJURY. COMPLICATIONS ASSOCIATED WITH PROGLIDE INCLUDED DEVICE FAILURE, FAILURE TO ACHIEVE HEMOSTASIS, IN-HOSPITAL ALL-CAUSE MORTALITY, LIFE-THREATENING OR MAJOR LIFE-THREATENING BLEEDING, MAJOR AND MINOR BLEEDING, MAJOR AND MINOR VASCULAR COMPLICATIONS, PSEUDOANEURYSM, UNPLANNED ENDOVASCULAR TREATMENT OF VASCULAR ACCESS, COVERED STENT IMPLANTATION TREATMENT OF VASCULAR ACCESS, UNPLANNED VASCULAR SURGERY, BLOOD TRANSFUSION, PROLONGED HOSPITALIZATION, USE OF ADDITIONAL VESSEL CLOSURE DEVICES, FAILURE TO ACHIEVE HEMOSTASIS THAT REQUIRED MANUAL COMPRESSION, USE OF VASOACTIVE MEDICATIONS, NEED OF INVASIVE TREATMENT FOR BLEEDING, STROKE, SURGICAL TREATMENT FOR BLEEDING, POSTOPERATIVE DECREASE OF HEMOGLOBIN, STENOSIS, DISSECTION, LOWER LIMB ISCHEMIA, ACCESS SITE INFECTION, IRREVERSIBLE NERVE INJURY, MAJOR ADVERSE CARDIOVASCULAR EVENTS, INCIDENCE OF CONTRAST-INDUCED NEPHROPATHY [KIDNEY INJURY], OCCLUSION, ACCESS SITE HEMATOMA. COMPLICATIONS ASSOCIATED WITH PROSTYLE INCLUDED DEVICE FAILURE, FAILURE TO ACHIEVE HEMOSTASIS, IN-HOSPITAL ALL-CAUSE MORTALITY, LIFE-THREATENING OR MAJOR LIFE-THREATENING BLEEDING, MAJOR AND MINOR BLEEDING, MAJOR AND MINOR VASCULAR COMPLICATIONS, PSEUDOANEURYSM, UNPLANNED ENDOVASCULAR TREATMENT OF VASCULAR ACCESS, COVERED STENT IMPLANTATION TREATMENT OF VASCULAR ACCESS, UNPLANNED VASCULAR SURGERY, BLOOD TRANSFUSION, PROLONGED HOSPITALIZATION. SPECIFIC PATIENT INFORMATION IS DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ARTICLE, TITLED "PLUG- VERSUS SUTURED-BASED VASCULAR CLOSURE DEVICES FOR LARGE BORE ARTERIAL ACCESS: A SYSTEMATIC REVIEW AND META-ANALYSIS". Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORDS FOR THIS PRODUCT COULD NOT BE REVIEWED AND A SIMILAR COMPLAINT QUERY COULD NOT BE PERFORMED BECAUSE THE PART NUMBER AND/OR LOT NUMBER WAS NOT REPORTED, AND THE DEVICE WAS NOT RETURNED. BASED ON THE INFORMATION REPORTED THROUGH THE RESEARCH ARTICLE, A CONCLUSIVE CAUSE FOR THE REPORTED DEATH, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING. ATTACHMENT ARTICLE TITLED: "PLUG- VERSUS SUTURED-BASED VASCULAR CLOSURE DEVICES FOR LARGE BORE ARTERIAL ACCESS: A SYSTEMATIC REVIEW AND META-ANALYSIS". THE ADDITIONAL PATIENT EFFECTS AND MALFUNCTIONS REPORTED IN THE ARTICLE ARE CAPTURED UNDER A SEPARATE MEDWATCH REPORT. THE ADDITIONAL DEVICES REFERENCED IN B5 ARE FILED UNDER SEPARATE MEDWATCH REPORT NUMBERS. B2, B3: ESTIMATED DATE D4- THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22884718&pc=DSQ | 3007042319-2025-01944 | 2025-07-05 04:00 | Death | HEARTWARE, INC. | 2025-08-25 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE VENTRICULAR ASSIST DEVICE (VAD) PATIENT EXPIRED. NO CAUSE OF DEATH WAS PROVIDED. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22884742&pc=DSQ | 3007042319-2025-01945 | 2025-05-30 04:00 | Death | HEARTWARE, INC. | 2025-08-25 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE VENTRICULAR ASSIST DEVICE (VAD) PATIENT EXPIRED. THE CAUSE OF DEATH WAS NOT PROVIDED. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22884832&pc=MGB | 2024168-2025-08964 | 2009-01-01 05:00 | Death | ABBOTT VASCULAR INC. | 2025-08-25 04:00 | MGB | PERCLOSE PROGLIDE¿ | Adverse Event Without Identified Device or Use Problem | Unspecified Vascular Problem | P960043 |  | 1 | Event Description: THE AIM OF THIS SYSTEMATIC REVIEW AND META-ANALYSIS WAS TO COMPARE MANTA A PLUG-BASED VASCULAR CLOSURE DEVICE (P-VCD) WITH PROSTAR, PROGLIDE AND PROSTYLE SUTURED-BASED VASCULAR CLOSURE DEVICE (S-VCD) IN PATIENTS UNDERGOING AN ENDOVASCULAR OR CARDIAC TREATMENT VIA A LARGE-BORE ARTERIAL ACCESS. IN CONCLUSION, THE SYSTEMATIC REVIEW AND META-ANALYSIS, THE LARGEST TO DATE, DEMONSTRATED NO SIGNIFICANT DIFFERENCES BETWEEN P-VCD AND S-VCD, REGARDING SAFETY OR EFFICACY OUTCOMES, EXCEPT FOR A SHORTER HOSPITAL STAY ASSOCIATED WITH P-VCD. COMPLICATIONS ASSOCIATED WITH PROSTAR INCLUDED DEVICE FAILURE, FAILURE TO ACHIEVE HEMOSTASIS, IN-HOSPITAL ALL-CAUSE MORTALITY, LIFE-THREATENING OR MAJOR LIFE-THREATENING BLEEDING, MAJOR AND MINOR BLEEDING, MAJOR AND MINOR VASCULAR COMPLICATIONS, PSEUDOANEURYSM, UNPLANNED ENDOVASCULAR TREATMENT OF VASCULAR ACCESS, COVERED STENT IMPLANTATION TREATMENT OF VASCULAR ACCESS, UNPLANNED VASCULAR SURGERY, BLOOD TRANSFUSION, PROLONGED HOSPITALIZATION, ARTERIOVENOUS FISTULA, ARTERIAL DISSECTION, ARTERIAL STENOSIS, HOMOLATERAL LOWER LIMB ISCHEMIA, ACCESS SITE RELATED VASCULAR INJURY. COMPLICATIONS ASSOCIATED WITH PROGLIDE INCLUDED DEVICE FAILURE, FAILURE TO ACHIEVE HEMOSTASIS, IN-HOSPITAL ALL-CAUSE MORTALITY, LIFE-THREATENING OR MAJOR LIFE-THREATENING BLEEDING, MAJOR AND MINOR BLEEDING, MAJOR AND MINOR VASCULAR COMPLICATIONS, PSEUDOANEURYSM, UNPLANNED ENDOVASCULAR TREATMENT OF VASCULAR ACCESS, COVERED STENT IMPLANTATION TREATMENT OF VASCULAR ACCESS, UNPLANNED VASCULAR SURGERY, BLOOD TRANSFUSION, PROLONGED HOSPITALIZATION, USE OF ADDITIONAL VESSEL CLOSURE DEVICES, FAILURE TO ACHIEVE HEMOSTASIS THAT REQUIRED MANUAL COMPRESSION, USE OF VASOACTIVE MEDICATIONS, NEED OF INVASIVE TREATMENT FOR BLEEDING, STROKE, SURGICAL TREATMENT FOR BLEEDING, POSTOPERATIVE DECREASE OF HEMOGLOBIN, STENOSIS, DISSECTION, LOWER LIMB ISCHEMIA, ACCESS SITE INFECTION, IRREVERSIBLE NERVE INJURY, MAJOR ADVERSE CARDIOVASCULAR EVENTS, INCIDENCE OF CONTRAST-INDUCED NEPHROPATHY [KIDNEY INJURY], OCCLUSION, ACCESS SITE HEMATOMA. COMPLICATIONS ASSOCIATED WITH PROSTYLE INCLUDED DEVICE FAILURE, FAILURE TO ACHIEVE HEMOSTASIS, IN-HOSPITAL ALL-CAUSE MORTALITY, LIFE-THREATENING OR MAJOR LIFE-THREATENING BLEEDING, MAJOR AND MINOR BLEEDING, MAJOR AND MINOR VASCULAR COMPLICATIONS, PSEUDOANEURYSM, UNPLANNED ENDOVASCULAR TREATMENT OF VASCULAR ACCESS, COVERED STENT IMPLANTATION TREATMENT OF VASCULAR ACCESS, UNPLANNED VASCULAR SURGERY, BLOOD TRANSFUSION, PROLONGED HOSPITALIZATION. SPECIFIC PATIENT INFORMATION IS DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ARTICLE, TITLED "PLUG- VERSUS SUTURED-BASED VASCULAR CLOSURE DEVICES FOR LARGE BORE ARTERIAL ACCESS: A SYSTEMATIC REVIEW AND META-ANALYSIS." Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORDS FOR THIS PRODUCT COULD NOT BE REVIEWED AND A SIMILAR COMPLAINT QUERY COULD NOT BE PERFORMED BECAUSE THE PART NUMBER AND/OR LOT NUMBER WAS NOT REPORTED, AND THE DEVICE WAS NOT RETURNED. THE PATIENT EFFECTS OF DEATH IS LISTED IN THE ELECTRONIC PERCLOSE PROGLIDE INSTRUCTIONS FOR USE (EIFU), AS POTENTIAL ADVERSE EVENTS OF USE OF THE DEVICE. BASED ON THE INFORMATION REPORTED THROUGH THE RESEARCH ARTICLE, A CONCLUSIVE CAUSE FOR THE REPORTED DEATH, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING. LITERATURE ATTACHMENT: ARTICLE TITLE: ¿PLUG- VERSUS SUTURED-BASED VASCULAR CLOSURE DEVICES FOR LARGE BORE ARTERIAL ACCESS: A SYSTEMATIC REVIEW AND META-ANALYSIS¿. THE ADDITIONAL PATIENT EFFECTS AND MALFUNCTIONS REPORTED IN THE ARTICLE ARE CAPTURED UNDER A SEPARATE MEDWATCH REPORT. THE ADDITIONAL DEVICES REFERENCED IN B5 ARE FILED UNDER SEPARATE MEDWATCH REPORT NUMBERS. B3, B2: ESTIMATED DATE. D4: THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22884907&pc=LKN | 1722028-2025-00197 | 2025-02-18 05:00 | Death | TERUMO BCT | 2025-08-25 04:00 | LKN | SPECTRA OPTIA | Insufficient Information | Hypersensitivity/Allergic reaction; Itching Sensation; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K183081 |  | 1 | Event Description: PER JOURNAL ARTICLE "UTILIZATION AND IMPORTANCE OF TIMING OF THERAPEUTIC PLASMA EXCHANGE IN KIDNEY TRANSPLANTATION: A SINGLE-CENTER EXPERIENCE" BY ORUC, A., ERSOY, A., OZKALEMKAS, F., GUL, C. B., YIGIT, E., & YILDIZ, A. INTRODUCTION: THERAPEUTIC PLASMA EXCHANGE (TPE) IS A PROMINENT APPROACH FOR SPECIFIC CONDITIONS AMONG RENAL TRANSPLANT CANDIDATES AND RECIPIENTS. THE INDICATIONS OF TPE, THOSE RELATED TO KIDNEY TRANSPLANTATION, ARE LIMITED, AND KNOWLEDGE AND UTILIZATION STRATEGIES ARE BLURRED. HEREIN, WE AIMED TO REVIEW OUR CLINICAL EXPERIENCE OF TPE AMONG RENAL TRANSPLANT RECIPIENTS. METHODS: A TOTAL OF 69 (40.22 ± 12.29 YEARS, 40 MALES, 63.8% DECEASED DONOR) ABO BLOOD TYPE-COMPATIBLE KIDNEY TRANSPLANT RECIPIENTS WHO UNDERWENT TPE WERE EVALUATED BETWEEN JANUARY 2013 AND DECEMBER 2019. RESULTS: A MEDIAN OF 5 (IQR 4¿7) TPE SESSIONS WAS ADMINISTERED AT A MEDIAN DURATION OF 31 (IQR 3¿564) DAYS AFTER TRANSPLANTATION. THE MAIN INDICATION WAS A REJECTION EPISODE. A TOTAL OF 61 (39.44 ± 12.22 YEARS, 24 LIVING DONOR, 10 RE-TRANSPLANT, 37 MALE) RECIPIENTS WERE SUCCESSFULLY DISCHARGED WITH A SERUM CREATININE OF 1.39 (IQR 1¿1.91) MG/DL, AND EGFR OF 62 (IQR 39¿94) ML/MIN. COMPARING THE RECIPIENTS ACCORDING TO TPE TIMING, EARLY ADMINISTERED 32 (52.5%) RECIPIENTS WITH A MEDIAN OF 4 (IQR 4¿7) DAYS AFTER TRANSPLANTATION HAD BETTER GRAFT SURVIVAL (62.5% VS. 17.2%, P < 0.001), DESPITE HIGHER MORTALITY RATES (25% VS.10%, P < 0.001). THE RATE OF GRAFT LOSS WAS HIGHER (62.1% VS. 9.4%, P < 0.001) AMONG THE RECIPIENTS WHO REQUIRED TPE DURING THE LATE POSTTRANSPLANT PERIOD (756 [IQR 272¿1307] DAYS). CONCLUSION: ACCORDING TO OUR FINDINGS, TPE COULD PROVIDE BENEFICIAL EFFECTS ON GRAFT AND PATIENT OUTCOMES, AND THE TIMING OF TPE COULD INFLUENCE ITS EFFICIENCY. FURTHER STUDIES ARE NEEDED TO SUPPORT OUR FINDINGS. SPECIFIC DETAILS, SUCH AS PATIENT INFORMATION AND OUTCOME, WERE NOT INCLUDED IN THE ARTICLE FOR THESE EVENTS, THEREFORE THIS REPORT IS BEING PROVIDED AS A SUMMARY OF THE EVENTS. IT IS UNKNOWN AT THIS TIME IF MEDICAL INTERVENTION WAS REQUIRED FOR THESE EVENTS THE COLLECTION SET IS NOT AVAILABLE FOR RETURN BECAUSE IT WAS DISCARDED BY THE CUSTOMER. THIS REPORT IS BEING FILED DUE TO PATIENT DEATH, ALTHOUGH PER CURRENT INFORMATION THERE IS NO DETECTABLE MALFUNCTION WITH THE TERUMO BCT DEVICE OR ALLEGATION OF A MALFUNCTION. Manufacturer Narrative: LOT NUMBER, MANUFACTURE DATE AND EXPIRY DATE ARE NOT AVAILABLE AT THIS TIME. INVESTIGATION IS IN PROCESS, A FOLLOW-UP REPORT WILL BE PROVIDED. CITATION: ORUC, A., ERSOY, A., OZKALEMKAS, F., GUL, C. B., YIGIT, E., & YILDIZ, A. (2025). UTILIZATION AND IMPORTANCE OF TIMING OF THERAPEUTIC PLASMA EXCHANGE IN KIDNEY TRANSPLANTATION: A SINGLE-CENTER EXPERIENCE. THERAPEUTIC APHERESIS AND DIALYSIS. HTTPS://DOI.ORG/10.1111/1744-9987.70002.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22885097&pc=MGB | 2024168-2025-08968 | 2009-01-01 05:00 | Death | ABBOTT VASCULAR INC. | 2025-08-25 04:00 | MGB | PERCLOSE¿ PROSTYLE¿ | Adverse Event Without Identified Device or Use Problem | Unspecified Vascular Problem | P960043 |  | 1 | Event Description: THE AIM OF THIS SYSTEMATIC REVIEW AND META-ANALYSIS WAS TO COMPARE MANTA A PLUG-BASED VASCULAR CLOSURE DEVICE (P-VCD) WITH PROSTAR, PROGLIDE AND PROSTYLE SUTURED-BASED VASCULAR CLOSURE DEVICE (S-VCD) IN PATIENTS UNDERGOING AN ENDOVASCULAR OR CARDIAC TREATMENT VIA A LARGE-BORE ARTERIAL ACCESS. IN CONCLUSION, THE SYSTEMATIC REVIEW AND META-ANALYSIS, THE LARGEST TO DATE, DEMONSTRATED NO SIGNIFICANT DIFFERENCES BETWEEN P-VCD AND S-VCD, REGARDING SAFETY OR EFFICACY OUTCOMES, EXCEPT FOR A SHORTER HOSPITAL STAY ASSOCIATED WITH P-VCD. COMPLICATIONS ASSOCIATED WITH PROSTAR INCLUDED DEVICE FAILURE, FAILURE TO ACHIEVE HEMOSTASIS, IN-HOSPITAL ALL-CAUSE MORTALITY, LIFE-THREATENING OR MAJOR LIFE-THREATENING BLEEDING, MAJOR AND MINOR BLEEDING, MAJOR AND MINOR VASCULAR COMPLICATIONS, PSEUDOANEURYSM, UNPLANNED ENDOVASCULAR TREATMENT OF VASCULAR ACCESS, COVERED STENT IMPLANTATION TREATMENT OF VASCULAR ACCESS, UNPLANNED VASCULAR SURGERY, BLOOD TRANSFUSION, PROLONGED HOSPITALIZATION, ARTERIOVENOUS FISTULA, ARTERIAL DISSECTION, ARTERIAL STENOSIS, HOMOLATERAL LOWER LIMB ISCHEMIA, ACCESS SITE RELATED VASCULAR INJURY. COMPLICATIONS ASSOCIATED WITH PROGLIDE INCLUDED DEVICE FAILURE, FAILURE TO ACHIEVE HEMOSTASIS, IN-HOSPITAL ALL-CAUSE MORTALITY, LIFE-THREATENING OR MAJOR LIFE-THREATENING BLEEDING, MAJOR AND MINOR BLEEDING, MAJOR AND MINOR VASCULAR COMPLICATIONS, PSEUDOANEURYSM, UNPLANNED ENDOVASCULAR TREATMENT OF VASCULAR ACCESS, COVERED STENT IMPLANTATION TREATMENT OF VASCULAR ACCESS, UNPLANNED VASCULAR SURGERY, BLOOD TRANSFUSION, PROLONGED HOSPITALIZATION, USE OF ADDITIONAL VESSEL CLOSURE DEVICES, FAILURE TO ACHIEVE HEMOSTASIS THAT REQUIRED MANUAL COMPRESSION, USE OF VASOACTIVE MEDICATIONS, NEED OF INVASIVE TREATMENT FOR BLEEDING, STROKE, SURGICAL TREATMENT FOR BLEEDING, POSTOPERATIVE DECREASE OF HEMOGLOBIN, STENOSIS, DISSECTION, LOWER LIMB ISCHEMIA, ACCESS SITE INFECTION, IRREVERSIBLE NERVE INJURY, MAJOR ADVERSE CARDIOVASCULAR EVENTS, INCIDENCE OF CONTRAST-INDUCED NEPHROPATHY [KIDNEY INJURY], OCCLUSION, ACCESS SITE HEMATOMA. COMPLICATIONS ASSOCIATED WITH PROSTYLE INCLUDED DEVICE FAILURE, FAILURE TO ACHIEVE HEMOSTASIS, IN-HOSPITAL ALL-CAUSE MORTALITY, LIFE-THREATENING OR MAJOR LIFE-THREATENING BLEEDING, MAJOR AND MINOR BLEEDING, MAJOR AND MINOR VASCULAR COMPLICATIONS, PSEUDOANEURYSM, UNPLANNED ENDOVASCULAR TREATMENT OF VASCULAR ACCESS, COVERED STENT IMPLANTATION TREATMENT OF VASCULAR ACCESS, UNPLANNED VASCULAR SURGERY, BLOOD TRANSFUSION, PROLONGED HOSPITALIZATION. SPECIFIC PATIENT INFORMATION IS DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ARTICLE, TITLED "PLUG- VERSUS SUTURED-BASED VASCULAR CLOSURE DEVICES FOR LARGE BORE ARTERIAL ACCESS: A SYSTEMATIC REVIEW AND META-ANALYSIS". Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORDS FOR THIS PRODUCT COULD NOT BE REVIEWED AND A SIMILAR COMPLAINT QUERY COULD NOT BE PERFORMED BECAUSE THE PART NUMBER AND/OR LOT NUMBER WAS NOT REPORTED, AND THE DEVICE WAS NOT RETURNED. BASED ON THE INFORMATION REPORTED THROUGH THE RESEARCH ARTICLE, A CONCLUSIVE CAUSE FOR THE REPORTED DEATH, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED. THE PATIENT EFFECTS OF DEATH IS LISTED IN THE ELECTRONIC PERCLOSE PROSTYLE INSTRUCTIONS FOR USE (EIFU), AS POTENTIAL ADVERSE EVENTS OF USE OF THE DEVICE. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING. LITERATURE ATTACHMENT: ARTICLE TITLE: "PLUG- VERSUS SUTURED-BASED VASCULAR CLOSURE DEVICES FOR LARGE BORE ARTERIAL ACCESS: A SYSTEMATIC REVIEW AND META-ANALYSIS". THE ADDITIONAL PATIENT EFFECTS AND MALFUNCTIONS REPORTED IN THE ARTICLE ARE CAPTURED UNDER A SEPARATE MEDWATCH REPORT. THE ADDITIONAL DEVICES REFERENCED IN B5 ARE FILED UNDER SEPARATE MEDWATCH REPORT NUMBERS. B2, B3: ESTIMATED DATE. D4: THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22885293&pc=OZD | 1220648-2025-31186 | 2025-06-20 04:00 | Death | ABIOMED, INC. | 2025-08-25 04:00 | OZD | AUTOMATED IMPELLA CONTROLLER | Device Sensing Problem | Insufficient Information | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PLACEMENT SIGNAL ON AN IMPELLA CP PUMP BECAME UNRELIABLE DURING PATIENT SUPPORT. THE PLACEMENT MONITORING WAS DISABLED IN RESPONSE TO THE NOTED ISSUE AND SUPPORT WAS CONTINUED UNINTERRUPTED. TWO DAYS LATER, THE PATIENT WAS MADE PALLIATIVE CARE BY THEIR FAMILY AND PASSED AWAY ON SUPPORT. THE INVESTIGATING TEAM HAS REQUESTED THE CONSOLE TO BE INVESTIGATED AS WELL, BEYOND THE COMPLAINANT'S REQUEST FOR PUMP INVESTIGATION. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA CONTROLLER WAS RETURNED FOR EVALUATION. UPON COMPLETION OF THE INVESTIGATION, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22885388&pc=DSQ | 3007042319-2025-01946 | 2025-06-15 04:00 | Death | HEARTWARE, INC. | 2025-08-25 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT SUBSEQUENTLY PASSED AWAY PEACEFULLY DUE TO CLINICAL ISSUES. CAUSE OF DEATH UNSPECIFIED. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22885436&pc=PNY | 2024168-2025-08973 | 2015-09-23 04:00 | Death | ABBOTT VASCULAR | 2025-08-25 04:00 | PNY | ABSORB BIORESORBABLE VASCULAR SCAFFOLD SYSTEM | Adverse Event Without Identified Device or Use Problem | Unspecified Heart Problem | P150023 |  | 1 | Event Description: IT WAS REPORTED IN A SUMMARY ARTICLE THE FINDINGS OF A RANDOMIZED CLINICAL TRIAL [PREVENT] OF THE OUTCOMES OF PERCUTANEOUS CORONARY INTERVENTION (PCI) VERSUS OPTIMAL MEDICAL THERAPY ALONE FOR THE TREATMENT OF VULNERABLE ATHEROSCLEROTIC CORONARY PLAQUES. IN THE STUDY PCI WAS PERFORMED WITH THE ABSORB AND XIENCE STENTS. PATIENT OUTCOMES WERE COMPARED AT THE 2 YEAR, 4 YEAR AND 7 YEAR FOLLOW-UPS. PATIENT OUTCOMES INCLUDED CARDIAC DEATH, TARGET VESSEL MYOCARDIAL INFARCTION, ISCHEMIA DRIVEN TARGET-VESSEL REVASCULARIZATION, HOSPITALIZATION FOR UNSTABLE OR PROGRESSIVE ANGINA, THROMBOSIS, STROKE, BLEEDING, PROCEDURAL INTERVENTION, SURGERY, TRANSIENT ISCHEMIC ATTACK. ADDITIONAL INFORMATION CAN BE FOUND IN THE SUMMARY ARTICLE, "PREVENTIVE PERCUTANEOUS CORONARY INTERVENTION VERSUS OPTIMAL MEDICAL THERAPY ALONE FOR THE TREATMENT OF VULNERABLE ATHEROSCLEROTIC CORONARY PLAQUES (PREVENT): A MULTICENTRE, OPEN-LABEL, RANDOMISED CONTROLLED TRIAL" THE PROCEDURE/IMPLANTATION DATE HAS BEEN ESTIMATED TO BE (B)(6) 2015 AS THIS IS THE FIRST DATE WITH THE CLINIC STUDY FINDINGS. THE DATE OF DEATH OR ADVERSE PATIENT EFFECT HAS BEEN ESTIMATED AS (B)(6) 2016 AS THIS IS ONE YEAR AFTER THE ESTIMATED PROCEDURE/IMPLANTATION DATE. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORDS FOR THIS PRODUCT COULD NOT BE REVIEWED AND A SIMILAR COMPLAINT QUERY COULD NOT BE PERFORMED BECAUSE THE PART NUMBER AND/OR LOT NUMBER WAS NOT REPORTED, AND THE DEVICE WAS NOT RETURNED. A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT DEATH, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED; HOWEVER, THE SUBSEQUENT TREATMENT APPEARS TO BE RELATED TO THE OPERATIONAL CONTEXT OF THE PROCEDURE. THE REPORTED PATIENT DEATH IS LISTED IN THE ABSORB BIORESORBABLE VASCULAR SCAFFOLD SYSTEMS INSTRUCTIONS FOR USE AS A KNOWN PATIENT EFFECT(S) OF CORONARY STENTING PROCEDURES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING; THEREFORE, NO PRODUCT-RELATED CORRECTIVE ACTION WILL BE IMPLEMENTED IN THIS CASE. B2, B3: ESTIMATED DATES: D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. D6A: ESTIMATED DATE. THE ADDITIONAL PATIENT EFFECTS REPORTED IN THE ARTICLE(S) ARE CAPTURED UNDER SEPARATE MEDWATCH REPORT. LITERATURE ATTACHMENT: ARTICLE TITLED " PREVENTIVE PERCUTANEOUS CORONARY INTERVENTION VERSUS OPTIMAL MEDICAL THERAPY ALONE FOR THE TREATMENT OF VULNERABLE ATHEROSCLEROTIC CORONARY PLAQUES (PREVENT): A MULTICENTRE, OPEN-LABEL, RANDOMISED CONTROLLED TRIAL."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22885584&pc=NPT | 9617601-2025-01146 | 2025-08-21 04:00 | Death | MEDTRONIC MEXICO S. DE R.L. DE CV | 2025-08-25 04:00 | NPT | EVOLUT PRO PLUS VALVE | Device Dislodged or Dislocated; Insufficient Information; Activation Failure; Patient Device Interaction Problem | Asystole | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A 29 MM TRANSCATHETER AORTIC VALVE IMPLANTATION VIA THE RIGHT FEMORAL ARTERY, THE ANATOMY WAS FAVORABLE TO THE 14F SYSTEM, BUT THERE WAS DIFFICULTY IN PUNCTURE AND PASSAGE OF THE PACEMAKER, WITH STIMULATION PROVIDED BY THE WIRE. DIFFICULTY CROSSING THE VALVE DUE TO EXTENSIVE CALCIFICATION REQUIRED TWO CATHETER EXCHANGES USING A PRE-SHAPED CONFIDA GUIDEWIRE. PRE-DILATION WAS PERFORMED WITH A 20×40 MILLIMETER BALLOON IN TWO STAGES BEFORE THE INITIAL DEPLOYMENT. THE VALVE WAS ASSEMBLED WITHOUT MISLOADING BUT PRESENTED UNDEREXPANSION UPON RELEASE AND WAS RECAPTURED FOR FURTHER PRE-DILATION. A BALLOON DILATION WAS PERFORMED AGAIN WITH A 20X40 MM BALLOON. UPON REPOSITIONING, NO OVERLAP AT THE FOURTH KNOT OR MISLOADING WAS OBSERVED. A NEW ATTEMPT AT DEPLOYMENT WITH THE SAME VALVE STILL RESULTED IN UNDEREXPANSION, PROMPTING THE CLINICAL DECISION TO PROCEED WITH DEPLOYMENT IN OVERLAP AND COPLANAR ORIENTATION, USING COMMISSURAL ALIGNMENT. BEFORE FINAL RELEASE, THE WIRE WAS WITHDRAWN FROM THE VENTRICLE, AND SYSTEM TENSION ON THE LESSER CURVATURE OF THE AORTA WAS NOTED. THE FINAL RELEASE WAS SLOW, WITH TWO RECAPTURES PERFORMED DUE TO PERSISTENT UNDEREXPANSION, RESULTING IN FINAL POSITIONING WITH AN ESTIMATED DISTANCE OF 2¿3 MILLIMETER AT THE NON-CORONARY CUSP AND APPROXIMATELY 4 MILLIMETER AT THE LEFT CUSP, DISTANT FROM THE LEFT CORONARY CUSP. THE MAIN TECHNICAL DIFFICULTY WAS VALVE UNDEREXPANSION, AND THE USE OF A LARGER-CALIBER BALLOON FOR PRE-DILATION WAS NOT CONSIDERED DUE TO THE PATIENT¿S ANATOMY. THE VALVE DISLODGED AND A SMALLER 26 MM VALVE WAS IMPLANTED, DURING WHICH THE PATIENT PROGRESSED TO ASYSTOLE AND DID NOT RESPOND TO CARDIOPULMONARY RESUSCITATION. THE PATIENT SUBSEQUENTLY DIED. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID D-EVPROP23-29 (LOT: 0012124356); PRODUCT TYPE: 0195-HEART VALVES; PRODUCT ID L-EVPROP23-29 (LOT: 0012702016); PRODUCT TYPE: 0195-HEART VALVES; PRODUCT ID GWBC30 (LOT: 92402829); PRODUCT TYPE: 0193-GUIDEWIRE; IMPLANT DATE N/A; EXPLANT DATE N/A PRODUCT ID D-EVPROP23-29 (LOT: 0012124356); PRODUCT TYPE: 0195-HEART VALVES; PRODUCT ID EVPROPLUS-26 (K128117); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE (B)(6) 2025. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22885603&pc=NPT | 9617601-2025-01147 | 2025-08-21 04:00 | Death | MEDTRONIC MEXICO S. DE R.L. DE CV | 2025-08-25 04:00 | NPT | EVOLUT PRO PLUS VALVE | Insufficient Information; Patient Device Interaction Problem | Asystole | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A 29 MM TRANSCATHETER AORTIC VALVE IMPLANTATION VIA THE RIGHT FEMORAL ARTERY, THE ANATOMY WAS FAVORABLE TO THE 14F SYSTEM, BUT THERE WAS DIFFICULTY IN PUNCTURE AND PASSAGE OF THE PACEMAKER, WITH STIMULATION PROVIDED BY THE WIRE. DIFFICULTY CROSSING THE VALVE DUE TO EXTENSIVE CALCIFICATION REQUIRED TWO CATHETER EXCHANGES USING A PRE-SHAPED CONFIDA GUIDEWIRE. PRE-DILATION WAS PERFORMED WITH A 20×40 MILLIMETER BALLOON IN TWO STAGES BEFORE THE INITIAL DEPLOYMENT. THE VALVE WAS ASSEMBLED WITHOUT MISLOADING BUT PRESENTED UNDEREXPANSION UPON RELEASE AND WAS RECAPTURED FOR FURTHER PRE-DILATION. A BALLOON DILATION WAS PERFORMED AGAIN WITH A 20X40 MM BALLOON. UPON REPOSITIONING, NO OVERLAP AT THE FOURTH KNOT OR MISLOADING WAS OBSERVED. A NEW ATTEMPT AT DEPLOYMENT WITH THE SAME VALVE STILL RESULTED IN UNDEREXPANSION, PROMPTING THE CLINICAL DECISION TO PROCEED WITH DEPLOYMENT IN OVERLAP AND COPLANAR ORIENTATION, USING COMMISSURAL ALIGNMENT. BEFORE FINAL RELEASE, THE WIRE WAS WITHDRAWN FROM THE VENTRICLE, AND SYSTEM TENSION ON THE LESSER CURVATURE OF THE AORTA WAS NOTED. THE FINAL RELEASE WAS SLOW, WITH TWO RECAPTURES PERFORMED DUE TO PERSISTENT UNDEREXPANSION, RESULTING IN FINAL POSITIONING WITH AN ESTIMATED DISTANCE OF 2¿3 MILLIMETER AT THE NON-CORONARY CUSP AND APPROXIMATELY 4 MILLIMETER AT THE LEFT CUSP, DISTANT FROM THE LEFT CORONARY CUSP. THE MAIN TECHNICAL DIFFICULTY WAS VALVE UNDEREXPANSION, AND THE USE OF A LARGER-CALIBER BALLOON FOR PRE-DILATION WAS NOT CONSIDERED DUE TO THE PATIENT¿S ANATOMY. THE VALVE DISLODGED AND A SMALLER 26 MM VALVE WAS IMPLANTED, DURING WHICH THE PATIENT PROGRESSED TO ASYSTOLE AND DID NOT RESPOND TO CARDIOPULMONARY RESUSCITATION. THE PATIENT SUBSEQUENTLY DIED. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID D-EVPROP23-29 (LOT: 0012124356); PRODUCT TYPE: 0195-HEART VALVES; PRODUCT ID L-EVPROP23-29 (LOT: 0012702016); PRODUCT TYPE: 0195-HEART VALVES; PRODUCT ID GWBC30 (LOT: 92402829); PRODUCT TYPE: 0193-GUIDEWIRE; IMPLANT DATE N/A; EXPLANT DATE N/A PRODUCT ID D-EVPROP23-29 (LOT: 0012124356); PRODUCT TYPE: 0195-HEART VALVES; PRODUCT ID EVPROPLUS-29 (K038501); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE (B)(6) 2025. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22885636&pc=NIQ | 2024168-2025-08981 | 2015-09-23 04:00 | Death | ABBOTT VASCULAR | 2025-08-25 04:00 | NIQ | XIENCE XPEDITION EVEROLIMUS ELUTING CORONARY STENT SYSTEM | Adverse Event Without Identified Device or Use Problem | Unspecified Heart Problem | P110019 |  | 1 | Event Description: IT WAS REPORTED IN A SUMMARY ARTICLE THE FINDINGS OF A RANDOMIZED CLINICAL TRIAL [PREVENT] OF THE OUTCOMES OF PERCUTANEOUS CORONARY INTERVENTION (PCI) VERSUS OPTIMAL MEDICAL THERAPY ALONE FOR THE TREATMENT OF VULNERABLE ATHEROSCLEROTIC CORONARY PLAQUES. IN THE STUDY PCI WAS PERFORMED WITH THE ABSORB AND XIENCE XPEDITION STENTS. PATIENT OUTCOMES WERE COMPARED AT THE 2 YEAR, 4 YEAR AND 7 YEAR FOLLOW-UPS. PATIENT OUTCOMES INCLUDED CARDIAC DEATH, TARGET VESSEL MYOCARDIAL INFARCTION, ISCHEMIA DRIVEN TARGET-VESSEL REVASCULARIZATION, HOSPITALIZATION FOR UNSTABLE OR PROGRESSIVE ANGINA, THROMBOSIS, STROKE, BLEEDING, PROCEDURAL INTERVENTION, SURGERY, TRANSIENT ISCHEMIC ATTACK. ADDITIONAL INFORMATION CAN BE FOUND IN THE SUMMARY ARTICLE, "PREVENTIVE PERCUTANEOUS CORONARY INTERVENTION VERSUS OPTIMAL MEDICAL THERAPY ALONE FOR THE TREATMENT OF VULNERABLE ATHEROSCLEROTIC CORONARY PLAQUES (PREVENT): A MULTICENTRE, OPEN-LABEL, RANDOMISED CONTROLLED TRIAL" Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORDS FOR THIS PRODUCT COULD NOT BE REVIEWED AND A SIMILAR COMPLAINT QUERY COULD NOT BE PERFORMED BECAUSE THE PART NUMBER AND/OR LOT NUMBER WAS NOT REPORTED, AND THE DEVICE WAS NOT RETURNED. A CONCLUSIVE CAUSE FOR THE REPORTED DEATH, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED; HOWEVER, THE SUBSEQUENT TREATMENT APPEARS TO BE RELATED TO THE OPERATIONAL CONTEXT OF THE PROCEDURE. THE REPORTED PATIENT DEATH IS LISTED IN THE EVEROLIMUS ELUTING CORONARY STENT SYSTEMS (EECSS), XIENCE XPEDITION, INSTRUCTIONS FOR USE AS KNOWN PATIENT EFFECT(S) OF CORONARY STENTING PROCEDURES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING; THEREFORE, NO PRODUCT-RELATED CORRECTIVE ACTION WILL BE IMPLEMENTED IN THIS CASE. B2, B3, D6A: ESTIMATED DATES. D4: THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. THE ADDITIONAL PATIENT EFFECTS REPORTED IN B5 ARE CAPTURED UNDER SEPARATE MEDWATCH REPORT. THE ADDITIONAL DEVICE REPORTED IN B5 ARE CAPTURED UNDER SEPARATE MEDWATCH REPORT. LITERATURE ATTACHMENT: ARTICLE TITLE: ¿PREVENTIVE PERCUTANEOUS CORONARY INTERVENTION VERSUS OPTIMAL MEDICAL THERAPY ALONE FOR THE TREATMENT OF VULNERABLE ATHEROSCLEROTIC CORONARY PLAQUES (PREVENT): A MULTICENTRE, OPEN-LABEL, RANDOMISED CONTROLLED TRIAL¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22885639&pc=DSQ | 3007042319-2025-01947 | 2025-06-25 04:00 | Death | HEARTWARE, INC. | 2025-08-25 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE VENTRICULAR ASSIST DEVICE (VAD) EXPIRED. THE CAUSE OF DEATH WAS NOT PROVIDED. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22885884&pc=KDJ | 3030306055-2025-00200 | 2025-06-18 04:00 | Death | VANTIVE US HEALTHCARE LLC | 2025-08-25 04:00 | KDJ | NI | Improper or Incorrect Procedure or Method | Abdominal Pain; Peritonitis | NI |  | 1 | Event Description: IT WAS REPORTED THAT A PERITONEAL DIALYSIS (PD) PATIENT EXPERIENCED A BREACH IN ASEPTIC TECHNIQUE WHICH RESULTED IN PERITONITIS. THE BREACH IN ASEPTIC TECHNIQUE WAS FURTHER DESCRIBED AS "CHANGE IN CAREGIVER". THE PERITONITIS WAS MANIFESTED BY CLOUDY PD EFFLUENT AND ABDOMINAL PAIN. THE SAME DAY AS PERITONITIS ONSET, THE PATIENTS WAS HOSPITALIZED DUE TO PERITONITIS AND LATER DISCHARGED ON UNKNOWN DATE. THE PATIENT WAS TREATED WITH AN INJECTION VANCOMYCIN (1GM, ONCE EVERY 5TH DAY, INTRAPERITONEAL, ONGOING) AND INJECTION AMIKACIN (100MG, ONCE A DAY, INTRAPERITONEAL, ONGOING) FOR PERITONITIS. IT WAS NOT REPORTED IF THE PATIENT WAS RETRAINED. IT WAS REPORTED THE PATIENT PASSED AWAY AND THE CAUSE OF DEATH WAS NOT REPORTED. IT IS NOT REPORTED IF AN AUTOPSY WAS PERFORMED. IT WAS UNKNOWN IF THE PERITONITIS WAS RESOLVED PRIOR TO DEATH. IT WAS REPORTED THAT THE PD THERAPY WAS ONGOING TILL DEATH. NO ADDITIONAL INFORMATION WAS AVAILABLE. Manufacturer Narrative: THIS REPORT IS FOR A BREACH IN ASEPTIC TECHNIQUE WHICH RESULTED IN PERITONITIS. PER VANTIVE LABELING, USERS ARE INSTRUCTED TO USE ASEPTIC TECHNIQUE WHEN PERFORMING PERITONEAL DIALYSIS THERAPY. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22886259&pc=DSQ | 3007042319-2025-01948 | 2025-06-20 04:00 | Death | HEARTWARE, INC. | 2025-08-25 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE VENTRICULAR ASSIST DEVICE (VAD) PATIENT EXPIRED. THE CAUSE OF DEATH WAS NOT PROVIDED. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22886495&pc=NIQ | 2024168-2025-08994 | 2021-02-01 05:00 | Death | ABBOTT VASCULAR INC. | 2025-08-25 04:00 | NIQ | XIENCE V | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P070015 |  | 1 | Event Description: IT WAS REPORTED THAT THE ARTICLE AIMED TO EVALUATE THE EFFICACY AND SAFETY OF AN ASPIRIN-FREE STRATEGY IN PATIENTS UNDERGOING COMPLEX PERCUTANEOUS CORONARY INTERVENTION (PCI). A PRESPECIFIED SUBGROUP ANALYSIS WAS CONDUCTED BASED ON COMPLEX PCI IN THE STOPDAPT-3 (SHORT AND OPTIMAL DURATION OF DUAL ANTIPLATELET THERAPY AFTER EVEROLIMUS-ELUTING COBALT-CHROMIUM STENT-3), WHICH RANDOMLY COMPARED LOW-DOSE PRASUGREL (3.75 MG/D) MONOTHERAPY TO DUAL ANTIPLATELET THERAPY (DAPT) WITH LOW-DOSE PRASUGREL AND ASPIRIN IN PATIENTS WITH ACUTE CORONARY SYNDROME OR HIGH BLEEDING RISK. THE COPRIMARY ENDPOINTS WERE MAJOR BLEEDING EVENTS (BLEEDING ACADEMIC RESEARCH CONSORTIUM 3 OR 5) AND CARDIOVASCULAR EVENTS (A COMPOSITE OF CARDIOVASCULAR DEATH, MYOCARDIAL INFARCTION, DEFINITE STENT THROMBOSIS, OR ISCHEMIC STROKE) AT 1 MONTH. ALSO UNPLANNED CORONARY, TARGET LESION AND NONTARGET LESION REVASCULARIZATION OCCURRED. DETAILS ARE LISTED IN THE ARTICLE TITLED, "AN ASPIRIN-FREE STRATEGY FOR IMMEDIATE TREATMENT FOLLOWING COMPLEX PERCUTANEOUS CORONARY INTERVENTION". Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORDS FOR THIS PRODUCT COULD NOT BE REVIEWED AND A SIMILAR COMPLAINT QUERY COULD NOT BE PERFORMED BECAUSE THE PART NUMBER AND/OR LOT NUMBER WERE NOT REPORTED, AND THE DEVICE WAS NOT RETURNED. THE REPORTED PATIENT EFFECT OF DEATH IS LISTED IN THE XIENCE SIERRA EVEROLIMUS ELUTING CORONARY STENT SYSTEMS INSTRUCTIONS FOR USE AS KNOWN PATIENT EFFECT(S) OF CORONARY STENTING PROCEDURES. THE REPORTED PATIENT EFFECT OF DEATH IS LISTED IN THE XIENCE SKYPOINT, EVEROLIMUS ELUTING CORONARY STENT SYSTEM (EECSS), ELECTRONIC INSTRUCTIONS FOR USE AS KNOWN PATIENT EFFECTS OF CORONARY STENTING PROCEDURES. A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECTS OF DEATH, AND THE RELATIONSHIP TO PRODUCT, IF ANY, CANNOT BE DETERMINED; HOWEVER, THE SUBSEQUENT UNEXPECTED MEDICAL INTERVENTION APPEARS TO BE RELATED TO THE OPERATIONAL CONTEXT OF THE PROCEDURE. THE INFORMATION RECEIVED IN THE COMPLAINT WAS OBTAINED THROUGH A PROACTIVE LITERATURE SEARCH. SPECIFIC MODEL AND LOT NUMBERS WERE NOT AVAILABLE AND CAUSALITY BETWEEN THE DOCUMENTED DEVICE USAGE AND THE PATIENT EFFECTS COULD NOT BE ESTABLISHED. FACTORS THAT MAY CONTRIBUTE TO THE DEATH LISTED IN THE ARTICLE INCLUDE, BUT ARE NOT LIMITED TO, DEVICE TO PATIENT INTERACTION, PATIENT ANATOMICAL CONDITIONS, DEVICE TO USER INTERACTION AND MANUFACTURING OR MATERIAL ISSUES. DUE TO THE LIMITED INFORMATION AVAILABLE, THE EXACT CAUSE CANNOT BE DETERMINED. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN, OR LABELING; THEREFORE, NO PRODUCT-RELATED CORRECTIVE ACTION WILL BE IMPLEMENTED IN THIS CASE. LITERATURE ATTACHMENT: ARTICLE TITLE: ¿AN ASPIRIN-FREE STRATEGY FOR IMMEDIATE TREATMENT FOLLOWING COMPLEX PERCUTANEOUS CORONARY INTERVENTION¿. B2, B3, B6: DATE ESTIMATED. D4: THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. D6: DATE OF IMPLANT ESTIMATED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22887951&pc=LGW | 3006630150-2025-06857 | 2025-07-31 04:00 | Death | BOSTON SCIENTIFIC NEUROMODULATION CORPORATION | 2025-08-25 04:00 | LGW | WAVEWRITER ALPHA? | Adverse Event Without Identified Device or Use Problem | Muscle Weakness; Pain; Sepsis | P030017 |  | 1 | Event Description: IT WAS REPORTED THAT A SPINAL CORD STIMULATION (SCS) PATIENT EXPERIENCED PAIN AND SUBSEQUENTLY DIED DUE TO SEPSIS CAUSED BY THE DEVICE. AN ATTEMPT WAS MADE TO ADJUST THE DEVICE, BUT IT WAS UNSUCCESSFUL. THE PATIENT WAS TOO WEAK TO UNDERGO DEVICE REMOVAL. NO ADDITIONAL INFORMATION HAS BEEN OBTAINED AT THIS TIME. Manufacturer Narrative: BLOCK B3: EXACT DATE UNKNOWN, APPROXIMATED EVENT DATE BASED ON THE BSC AWARE DATE AS WE WERE NOT ABLE TO OBTAIN THE DATE THE PATIENT PASSED AWAY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22888308&pc=OAD | 3008452825-2025-00412 | 2025-07-31 04:00 | Death | ABBOTT MEDICAL | 2025-08-25 04:00 | OAD | FLEXABILITY¿ ABLATION CATHETER, SENSOR ENABLED¿ | Adverse Event Without Identified Device or Use Problem | Cardiac Perforation |  |  | 1 | Event Description: DURING A REDO PROCEDURE FOR ATRIAL FIBRILLATION, THE PATIENT WAS UNDER CONSCIOUS SEDATION AND WHILE BURNING ALONG THE SEPTUM, THE PATIENT SAT UP AND REACHED FOR HIS GROIN CAUSING THE CATHETER TO PERFORATE AND LEAD TO PERICARDIAL EFFUSION. A PERICARDIOCENTESIS WAS PERFORMED BUT DESPITE ALL LIFESAVING ATTEMPTS, THE PATIENT DID NOT SURVIVE. THERE WERE NO PERFORMANCE ISSUES WITH THE FLEXIBILITY CATHETER. Manufacturer Narrative: THE RESULTS OF THE INVESTIGATION ARE INCONCLUSIVE SINCE THE DEVICE WAS NOT RETURNED FOR ANALYSIS. REVIEW OF THE DEVICE HISTORY RECORD WAS NOT POSSIBLE AS THE LOT NUMBER IS UNKNOWN. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. MODEL/LOT NUMBER INFORMATION WAS NOT ABLE TO BE OBTAINED FOR THIS DEVICE. THEREFORE, FULL UDI INFORMATION(D4) AND 510K(G3) ARE NOT AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22888847&pc=DSQ | 2916596-2025-05570 | 2022-11-23 05:00 | Death | THORATEC CORPORATION | 2025-08-25 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED. AT THE TIME OF THE PATIENT'S DEATH, THERE WAS NO ONE FOLLOWING THE PATIENT CLOSELY, SO NO OTHER INFORMATION WAS AVAILABLE. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE CUSTOMER COMMUNICATED THAT NO ADDITIONAL INFORMATION WILL BE PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU), REV. D IS CURRENTLY AVAILABLE. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22889075&pc=NGV | 2135147-2025-04763 | 2025-07-31 04:00 | Death | ABBOTT MEDICAL | 2025-08-25 04:00 | NGV | AMPLATZER AMULET | Adverse Event Without Identified Device or Use Problem | Bradycardia; Low Blood Pressure/ Hypotension; Hypoxia; Ventricular Fibrillation | P200049 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 25MM AMPLATZER AMULET LEFT ATRIAL APPENDAGE (LAA) OCCLUDER WAS CHOSEN FOR IMPLANTATION UTILIZING A 12F AMPLATZER TORQVUE 45X45 DELIVERY SYSTEM. DEVICE WAS LOADED PER IFU. AFTER THE FIRST PROPOFOL ADMINISTRATION, OXYGEN SATURATION REPEATEDLY DECREASED. THE DEVICE WAS INSERTED INTO THE PATIENT ANATOMY. DURING THE PROCEDURE THE PATIENT BECAME HYPOTENSIVE, AND BRADYCARDIC. PHYSICIAN THOUGHT THIS WAS DUE TO THE AORTIC VALVE STENOSIS. REANIMATION AND INTUBATION OCCURRED WHILE THE DELIVERY SYSTEM AND OCCLUDER WERE REMOVED FROM THE ANATOMY. AFTER CPR, THE PATIENT DEVELOPED VENTRICULAR FIBRILLATION (VF) SO DEFIBRILLATION WAS PERFORMED. LEFT HEART CATHETERIZATION WAS PERFORMED AS WELL. THE PATIENT DIED ON THE TABLE. OFFICIAL CAUSE OF DEATH IS UNKNOWN, BUT IN THE PHYSICIAN'S OPINION IT WAS RELATED TO THE PATIENT'S COMORBIDITIES. THERE WAS NO REPORTED MALFUNCTION OF THE ABBOTT DEVICE. Manufacturer Narrative: THE DEVICE IS NOT RETURNING FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22889102&pc=DSQ | 2916596-2025-05274 | 2025-07-31 04:00 | Death | THORATEC CORPORATION | 2025-08-25 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Unspecified Infection; Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTIORGAN FAILURE RELATED TO INFECTION AND MULTIPLE MYELOMA. THE OUTCOME WAS CONSIDERED TO BE DEVICE OR THERAPY RELATED AND THE PUMP WOULD NOT RETURN FOR EVALUATION. DEVICE PARAMETERS WERE IN NORMAL LIMITS FOR THIS PATIENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22889203&pc=DSQ | 2916596-2025-05362 | 2025-08-07 04:00 | Death | THORATEC CORPORATION | 2025-08-25 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. WHETHER THE DEVICE WILL BE EXPLANTED IS UNKNOWN. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22889222&pc=DSQ | 2916596-2025-05386 | 2025-08-06 04:00 | Death | THORATEC CORPORATION | 2025-08-25 04:00 | DSQ | THORATEC® HEARTMATE®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO HEART FAILURE. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. ADDITIONAL INFORMATION WAS UNKNOWN. Manufacturer Narrative: SECTION D4: EXPIRATION DATE AND MANUFACTURE DATE (H4) WILL BE PROVIDED UPON INVESTIGATION. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22889236&pc=QBJ | 3004753838-2025-235276 | 2025-01-24 05:00 | Death | DEXCOM, INC. | 2025-08-25 04:00 | QBJ | DEXCOM G7 CONTINUOUS GLUCOSE MONITORING SYSTEM | No Audible Alarm | Hypoglycemia; Coma | K213919 |  | 1 | Event Description: IT WAS REPORTED THAT AN ALERT/NOTIFICATION SETTINGS ISSUE OCCURRED. THE PATIENT'S SPOUSE INITIALLY CALLED DUE TO RECEIVING A LETTER FOR THE RECEIVER RECALL. THE ISSUE WITH THE DEVICE OCCURRED ON (B)(6) 2025 AROUND 2:30 AM. THE SPOUSE WOKE UP AND CHECKED ON THE PATIENT BUT SHE COULDN'T WAKE HIM. UPON CHECKING HIS MONITOR, DEXCOM SHOWED, 54MGDL. THE SPOUSE SAID SHE DIDN'T HEAR ANY ALARM THAT HE WAS LOW. SHE DID NOT CHECK THE BG. IT NEVER ALERTED THEM. SHE TRIED TO GIVE JUICE, BUT HE WAS UNRESPONSIVE. SHE THEN CALLED AN AMBULANCE DESPITE THE PATIENT BEING ON PALLIATIVE CARE. THEY TOOK THE PATIENT TO THE HOSPITAL. HE WAS SEVERELY HYPOGLYCEMIC. THEY PUT HIM ON IV, GOT HIS SUGAR BACK UP. HE WAS DISCHARGED HOME AND CONTINUED TO RECEIVE PALLIATIVE CARE. HE DID RECOVER FROM THE HYPOGLYCEMIC EPISODE BUT PASSED 2 DAYS LATER. THE SPOUSE ALLEGED THAT THE ABSENCE OF ALERT DID CONTRIBUTE TO HIS DEATH. SHE STATED SHE FELT HE'D STILL BE ALIVE HAD SHE RECEIVED THE LOW GLUCOSE ALERT FROM THE RECEIVER AND HAD NOT HAD THE MISSED HYPOGLYCEMIC EPISODE. NO PRODUCT OR DATA WAS PROVIDED FOR INVESTIGATION. A RETURN GOODS KIT HAS BEEN SENT TO THE PATIENT¿S FAMILY TO RETURN THE ALLEGED RECEIVER FOR INVESTIGATION. A FOLLOW UP REPORT WILL BE SUBMITTED UPON DEVICE INVESTIGATION COMPLETION IF THE DEVICE IS RETURNED. NO ADDITIONAL PATIENT OR EVENT INFORMATION IS AVAILABLE. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22889339&pc=DSQ | 2916596-2025-05434 | 2025-08-05 04:00 | Death | THORATEC CORPORATION | 2025-08-25 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Ischemia Stroke | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS ISCHEMIC STROKE, CARE WITHDRAWN. IT WAS REPORTED THAT THE PATIENT'S OUTCOME WAS NOT DEVICE OR THERAPY RELATED. THE DEVICE WILL NOT BE EXPLANTED AND WILL NOT BE RETURNED FOR EVALUATION. THE DEVICE PARAMETERS (FLOW, SPEED, POWER) WERE WITHIN NORMAL LIMITS. Manufacturer Narrative: A4 - MISSING PATIENT WEIGHT. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22889462&pc=FTL | 9615742-2025-01050 |  | Death | SOFRADIM PRODUCTION SAS | 2025-08-25 04:00 | FTL | PARIETEX | Adverse Event Without Identified Device or Use Problem | Distress; Unspecified Tissue Injury |  |  | 1 | Event Description: THE PATIENT'S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR THERAPEUTIC TREATMENT OF A HERNIA. IT WAS REPORTED THAT AFTER IMPLANT, THE PATIENT EXPERIENCED EMOTIONAL DISTRESS, DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED REVISION SURGERY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22889535&pc=DSQ | 3007042319-2025-01949 | 2025-01-01 05:00 | Death | HEARTWARE, INC. | 2025-08-25 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP/DRIVELINE | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Erythema; Fluid Discharge | P100047 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING VITAMIN D DEFICIENCY AS A POTENTIAL MODIFIABLE RISK FACTOR FOR DRIVELINE INFECTION. PATIENTS WERE SPLIT INTO GROUPS WITH VITAMIN D DEFICIENCY AND WITHOUT. THE AUTHORS DESCRIBED PATIENTS WHO EXPERIENCED DRIVELINE INFECTIONS IN BOTH GROUPS. CULTURES TAKEN NOTED ISOLATED ORGANISMS SUCH AS STAPHYLOCOCCUS AUREUS, COAGULASE-NEGATIVE STAPHYLOCOCCI, CORYNEBACTERIUM SPECIES, PSEUDOMONAS SPECIES, CANDIDA AURIS, AND PROTEUS MIRABILIS. FOUR INFECTIONS WERE IDENTIFIED BASED ON THE PRESENCE OF INCREASED DRAINAGE AND ERYTHEMA AT THE DRIVELINE SITE THAT IMPROVED WITH ANTIBIOTIC TREATMENT. THERE WERE PATIENT DEATHS; HOWEVER, THE SPECIFIC CAUSES OF DEATH WERE UNKNOWN. THE STATUS OF THE DEVICES IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/53 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: VITAMIN D DEFICIENCY AND DRIVELINE INFECTIONS IN PATIENTS WITH LEFT VENTRICULAR ASSIST DEVICES. ARTIFICIAL ORGANS. 2025; 49:1185¿1191. DOI: 10.1111/AOR.14988. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22889744&pc=MGB | 2024168-2025-09042 | 2025-05-23 04:00 | Death | ABBOTT VASCULAR INC. | 2025-08-25 04:00 | MGB | PERCLOSE¿ PROSTYLE¿ | Off-Label Use; Insufficient Information | Ischemia; Obstruction/Occlusion | P960043 |  | 1 | Event Description: USER FACILITY MEDWATCH REPORT RECEIVED THAT STATES: "PATIENT HAD TEVAR (THORACIC ENDOVASCULAR AORTIC REPAIR) PROCEDURE DONE FOR SACCULAR ANEURYSM OF AORTIC ARCH UNRUPTURED. AFTER PROCEDURE SURGEON UTILIZED ABBOT PER CLOSE DEVICE. ON (B)(6) 2025 CVA SHOWS LEFTY EXTERNAL ILIAC ARTERY OCCLUSION SECONDARY TO PERCLOSE. REQUIRED EMERGENT OPEN REPAIR FOR ACUTE LIMB ISCHEMIA DUE TO FAILURE OF PERCLOSE DEVICE." ADDITIONAL INFORMATION: IT WAS REPORTED THAT THE PATIENT WAS HOSPITALIZED AND LATER DIED. THERE WAS NO REPORTED CLINICALLY SIGNIFICANT DELAY IN THE PROCEDURE OR THERAPY. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. THE PRODUCTION RECORDS FOR THIS PRODUCT COULD NOT BE REVIEWED BECAUSE THE PART NUMBER AND/OR LOT NUMBER WAS NOT REPORTED. A REVIEW OF THE CORRECTIVE AND PREVENTIVE ACTIONS (CAPA) REVIEW WAS NOT PERFORMED BECAUSE A SPECIFIC FAILURE MODE FOR THIS COMPLAINT WAS NOT PROVIDED. ADDITIONALLY, A REVIEW OF THE COMPLAINT HISTORY WAS NOT PERFORMED AS A SIMILARITY COULD NOT BE DETERMINED. PROSTYLE DEVICE WAS USED IN EXTERNAL ILIAC ARTERY. IT SHOULD BE NOTED THAT THE ELECTRONIC PERCLOSE PROSTYLE INSTRUCTIONS FOR USE (IFU) STATES: THE PERCLOSE PROSTYLE SMC AND REPAIR SYSTEM IS INDICATED FOR THE PERCUTANEOUS DELIVERY OF SUTURE FOR CLOSING THE COMMON FEMORAL ARTERY AND VEIN ACCESS SITES OF PATIENTS WHO HAVE UNDERGONE DIAGNOSTIC OR INTERVENTIONAL CATHETERIZATION PROCEDURES. IN THIS CASE, IT IS UNKNOWN IF THE REPORTED IFU DEVIATION CONTRIBUTED TO THE REPORTED PATIENT EFFECTS. BASED ON THE INFORMATION PROVIDED, A DEFINITIVE CAUSE FOR THE REPORTED UNSPECIFIED DEVICE ISSUE COULD NOT BE DETERMINED. THE PATIENT EFFECTS AND TREATMENTS APPEARS TO BE RELATED TO CIRCUMSTANCES OF THE PROCEDURE. THE PATIENT EFFECTS OF OCCLUSION, ISCHEMIA AND DEATH ARE LISTED IN THE PROSTYLE IFU AS A POTENTIAL ADVERSE EVENT ASSOCIATED WITH USE OF THE SUTURE MEDIATED CLOSURE DEVICES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO THE DESIGN, MANUFACTURE, OR LABELING OF THE DEVICE. B2: DATE ESTIMATED. D4: A PARTIAL UDI WAS PROVIDED AS THE LOT NUMBER IS NOT KNOWN. H6: MEDICAL DEVICE PROBLEM CODE: 1494 - INCORRECT ANATOMY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22889983&pc=OZD | 1220648-2025-45790 | 2025-07-28 04:00 | Death | ABIOMED, INC. | 2025-08-25 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Tachycardia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT ROUGHLY FOUR DAYS INTO IMPELLA CP SUPPORT, A PATIENT ENDURED PERSISTENT RUNS OF VENTRICULAR TACHYCARDIA. MULTIPLE DEFIBRILLATIONS WERE REQUIRED TO COUNTERACT THE CONDITION. SUPPORT CONTINUED UNINTERRUPTED FOLLOWING THE INTERVENTION. THE FOLLOWING DAY, THE BEDSIDE STAFF MADE THE DECISION TO WITHDRAW CARE AND THE PATIENT PASSED AWAY. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: A.2, A.4, A.5, AND A.6 ARE UNKNOWN. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22890066&pc=DSQ | 3007042319-2025-01951 | 2025-06-30 04:00 | Death | HEARTWARE, INC. | 2025-08-25 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY PEACEFULLY DUE TO CLINICAL ISSUES UNRELATED TO THE VENTRICULAR ASSIST DEVICE THERAPY. THERE WERE NO REPORTS OR CONCERNS REGARDING THE PERFORMANCE OF THE VENTRICULAR ASSIST DEVICE OR THE THERAPY. THIS DEVICE IS INCLUDED IN THE SUBGROUP 3 POPULATION FOR DELAY OR FAILURE TO RESTART. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22890190&pc=NPW | 2015691-2025-06979 | 2025-08-09 04:00 | Death | EDWARDS LIFESCIENCES | 2025-08-25 04:00 | NPW | EVOQUE TRICUSPID VALVE REPLACEMENT SYSTEM | Difficult to Insert | Great Vessel Perforation | P230013 |  | 1 | Event Description: AS REPORTED BY THE EDWARDS LIFESCIENCES AFFILIATE IN GERMANY, EDWARDS RECEIVED NOTIFICATION REGARDING AN EVOQUE PROCEDURE. ON POST-OPERATIVE DAY (POD)1, THE PATIENT EXPERIENCED ACTIVE BLEEDING FROM THE INFERIOR VENA CAVA (IVC), ALONG WITH A RETROPERITONEAL HEMATOMA. IN RESPONSE, THE IVC WAS STENTED TO CONTROL THE BLEEDING. TWO DAYS AFTER THE BLEEDING WAS IDENTIFIED, THE PATIENT REPORTED FLANK PAIN, AND HEMOGLOBIN LEVELS CONTINUED TO DECLINE. THE PATIENT REMAINED HYPOTENSIVE, PROMPTING A FOLLOW-UP CT SCAN, WHICH SHOWED PERSISTENT ACTIVE BLEEDING DESPITE THE INITIAL STENTING. A SECOND STENT WAS PLACED IN AN ATTEMPT TO ACHIEVE HEMOSTASIS. THE PATIENT BECAME HEMODYNAMICALLY UNSTABLE AND REQUIRED RESUSCITATION. UNFORTUNATELY, RETURN OF SPONTANEOUS CIRCULATION (ROSC) COULD NOT BE ACHIEVED, AND THE PATIENT PASSED AWAY. Manufacturer Narrative: THE MANUFACTURER'S INVESTIGATION IS ONGOING AND A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22890435&pc=DSQ | 2916596-2025-05418 | 2025-08-05 04:00 | Death | THORATEC CORPORATION | 2025-08-25 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Stroke/CVA | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS SECONDARY TO ARRYTHMIAS AND CEREBRAL VASCULAR ACCIDENT. IT WAS REPORTED THAT THE PATIENT'S OUTCOME WAS NOT DEVICE OR THERAPY RELATED. THE DEVICE WILL NOT BE EXPLANTED AND WILL NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22890445&pc=DSQ | 2916596-2025-05409 | 2025-08-06 04:00 | Death | THORATEC CORPORATION | 2025-08-25 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Brain Injury | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO SUBDURAL HEMATOMA AND ANOXIC BRAIN INJURY STATUS POST A TRAUMATIC FALL. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. IT WAS ALSO SAID THAT THE DEVICE WAS OPERATING WITHIN NORMAL PARAMETERS FOR THE PATIENT AT THE TIME OF THE EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22890496&pc=DSQ | 3007042319-2025-01952 | 2025-07-30 04:00 | Death | HEARTWARE, INC. | 2025-08-25 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WITH A MEDICAL HISTORY OF CARDIOMYOPATHY WHO HAD RECEIVED A VENTRICULAR ASSIST DEVICE (VAD) IMPLANT PASSED AWAY PEACEFULLY. A CAUSE OF DEATH WAS NOT PROVIDED. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22890538&pc=DSQ | 3007042319-2025-01953 | 2018-03-30 04:00 | Death | HEARTWARE, INC. | 2025-08-25 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P100047 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING THE RELATIONSHIP OF PRE-VENTRICULAR ASSIST DEVICE (VAD) VITAMIN D LEVELS WITH POSTOPERATIVE CLINICAL OUTCOMES. THE AUTHORS DESCRIBED PATIENT DEATHS; HOWEVER, THE CAUSES OF DEATH WERE UNKNOWN AND DESCRIBED AS ALL-CAUSE MORTALITY. THERE WERE PATIENTS WHO EXPERIENCED HOSPITAL READMISSIONS DUE TO DRIVELINE INFECTIONS, GASTROINTESTINAL BLEEDS (GIB) DEFINED AS ANY UNEXPECTED DECLINE IN HEMOGLOBIN LEVEL WITH EVIDENCE OF MELENA, HEMATEMESIS, OR HEMATOCHEZIA, AND CEREBROVASCULAR ACCIDENTS (CVA) WHICH INCLUDED BOTH ISCHEMIC AND HEMORRHAGIC STROKES. THE STATUS OF THE VADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/59 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: EFFECT OF VITAMIN D LEVEL ON CLINICAL OUTCOMES IN PATIENTS UNDERGOING LEFT VENTRICULAR ASSIST DEVICE IMPLANTATION. NUTRITION IN CLINICAL PRACTICE. 2018; 33:825¿830. DOI: 10.1002/NCP.10078. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22890548&pc=DSQ | 3007042319-2025-01955 | 2025-06-13 04:00 | Death | HEARTWARE, INC. | 2025-08-25 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WITH A VENTRICULAR ASSIST DEVICE (VAD) IMPLANT WAS FOUND DECEASED, WITH MULTI-ORGAN FAILURE IDENTIFIED AS THE CAUSE OF DEATH. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22891108&pc=OZD | 1220648-2025-45943 | 2025-07-26 04:00 | Death | ABIOMED, INC. | 2025-08-25 04:00 | OZD | IMPELLA 5.5 | Malposition of Device | Tachycardia; Ventricular Fibrillation | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WITH CARDIOMYOPATHY WAS INITIALLY IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS) AND THEN ESCALATED TO AN IMPELLA 5.5. FOR CONTINUATION IN THERAPY. IMPELLA 5.5 WAS SLOWLY INCREASED TO PERFORMANCE LEVEL P-5, IMPELLA CP WAS WEANED AND EXPLANTED. IMPELLA 5.5 MEASURED UNDER ECHOCARDIOGRAM WITH INLET SITTING FREE FROM STRUCTURES, MID VENTRICULAR AND MEASURING APPROXIMATELY 5CM FROM VALVE. THE IMPELLA WAS SECURED, AND THE PATIENT WAS TRANSFERRED TO THE UNIT ON SUPPORT. APPROXIMATELY THREE DAYS AFTER STARTING THE IMPELLA MCS, THE PATIENT EXPERIENCED VENTRICULAR FIBRILLATION (VF) AND VENTRICULAR TACHYCARDIA (VT) WHICH WERE TREATED WITH DEFIBRILLATION. TWO DAYS LATER, THE IMPELLA POSITION WAS FOUND TO BE DEEP, AND THE DEVICE WAS REPOSITIONED. THE PATIENT CONTINUED TO HAVE VF AND VT REQUIRING 16 SHOCKS AND MEDICATION. THE PATIENT HAD ONGOING MULTISYSTEM ORGAN FAILURE AND WAS INELIGIBLE FOR ADVANCED THERAPIES. DAY 18 OF SUPPORT, THE PATIENT EXPIRED. CARE WAS WITHDRAWN. Manufacturer Narrative: . DEVICE STATUS: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22891148&pc=DSQ | 2916596-2025-05371 | 2025-08-08 04:00 | Death | THORATEC CORPORATION | 2025-08-25 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD A MASSIVE INTRA-OPERATIVE STROKE THAT WAS DEEMED UN SURVIVABLE. THE VENTRICULAR ASSIST DEVICE (VAD) APPEARED TO OPERATE AS EXPECTED HOWEVER DUE TO THE LARGE STROKE CARE WOULD BE WITHDRAWN. THE PATIENT PASSED AWAY. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22891511&pc=OZD | 1220648-2025-45936 | 2025-07-29 04:00 | Death | ABIOMED, INC. | 2025-08-25 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT POST CARDIOTOMY CARDIOGENIC SHOCK/LOW CARDIAC OUTPUT SYNDROME WAS IMPLANTED WITH AN IMPELLA 5.5 FOR MECHANICAL CIRCULATORY SUPPORT (AND AT THIS TIME INTRA-AORTIC BALLOON PUMP AND VENOARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) WERE REMOVED). AFTER APPROXIMATELY ONE MONTH OF IMPELLA 5.5 SUPPORT, THE PATIENT WAS NOTED TO HAVE GASTROINTESTINAL BLEEDING. HEPARIN WAS SUSPENDED AND BLOOD PRODUCTS WERE ADMINISTERED. IT COULD NOT BE DETERMINED IF THE USE OF ANTICOAGULANT FOR USE OF THE IMPELLA DEVICE IMPACTED/EXACERBATED THE OUT OF THE BLEED FOR THE PATIENT. DUE TO CONTINUED DECOMPENSATION AND INABILITY TO WEAN OFF VENT AND VENO-PULMONARY ECMO, FAMILY AND TEAM ULTIMATELY DECIDED TO WITHDRAW CARE DUE TO PATIENT BECOMING UNRESPONSIVE. THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE: AS NOTED IN THE IMPELLA IFU: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) - "ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22881399&pc=DSQ | 2916596-2025-05342 | 2025-08-01 04:00 | Death | THORATEC CORPORATION | 2025-08-24 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS INFECTION AND CONGESTIVE HEART FAILURE. IT WAS REPORTED THAT THE PATIENT'S DEATH WAS NOT DEVICE OR DEVICE THERAPY RELATED. THE DEVICE PARAMETERS (FLOW, SPEED, POWER) WERE WITHIN NORMAL LIMITS. THE DEVICE WOULD NOT BE EXPLANTED AND WOULD NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22881548&pc=DSQ | 2916596-2025-05367 | 2025-08-03 04:00 | Death | THORATEC CORPORATION | 2025-08-24 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. THE PATIENT HAD RIGHT HEART FAILURE, AND THE FAMILY DECIDED TO WITHDRAW CARE. THE DEVICE WOULD NOT RETURN. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22881683&pc=NVN | 2017865-2025-98478 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-08-24 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAS DECEASED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE RELATED. THE CAUSE OF DEATH WAS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22881690&pc=NVY | 2017865-2025-98480 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-08-24 04:00 | NVY | DURATA STS OPTIM ACTIVE FIXATION, DF-4 CONNECTOR | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P950022 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAS DECEASED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE RELATED. THE CAUSE OF DEATH WAS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22881691&pc=NIK | 2017865-2025-98476 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-08-24 04:00 | NIK | GALLANT HF | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAS DECEASED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE RELATED. THE CAUSE OF DEATH WAS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22881693&pc=NIK | 2017865-2025-98477 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-08-24 04:00 | NIK | QUARTET | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAS DECEASED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE RELATED. THE CAUSE OF DEATH WAS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22871166&pc=LJS | 3006260740-2025-06095 | 2021-10-19 04:00 | Death | C.R. BARD, INC. (BASD) -3006260740 | 2025-08-22 04:00 | LJS | UNKNOWN CHRONIC CATHETER | Adverse Event Without Identified Device or Use Problem | Insufficient Information | UNKNOWN |  | 1 | Event Description: IT WAS REPORTED IN AN ARTICLE IN THE JOURNAL "RESPIRATORY MEDICINE" TITLED, "COMPLICATIONS ASSOCIATED WITH PERIPHERALLY INSERTED CENTRAL CATHETERS AND HICKMAN IN PATIENTS WITH ADVANCED PULMONARY HYPERTENSION TREATED WITH INTRAVENOUS PROSTANOIDS", 175 CATHETERS WERE IMPLANTED IN 109 PATIENTS. OF THEM, 100 WERE HICKMAN AND 75 PICC CATHETERS. PATIENTS DEVELOPED INFECTION AND MECHANICAL COMPLICATIONS SUCH AS DISPLACEMENT AND OCCLUSION WERE ALSO OBSERVED. THE RATE OF MORTALITY WAS HIGHER IN PICC GROUP THAN IN HICKMAN GROUP. Manufacturer Narrative: H11: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. DATE OF DEATH FOR THIS PATIENT WAS NOT PROVIDED, DATE OF DEATH UPDATED AS 01-JAN-1900 FOR THE MDR SUBMISSION REQUIREMENT. HINOJOSA W, CRUZ A, CRUZ-UTRILLA A, CRISTO ROPERO MJ, LÓPEZ-MEDRANO F, GÓMEZ I, ARRIBAS-YNSAURRIAGA F, ESCRIBANO-SUBIAS P. COMPLICATIONS ASSOCIATED WITH PERIPHERALLY INSERTED CENTRAL CATHETERS AND HICKMAN¿ IN PATIENTS WITH ADVANCED PULMONARY HYPERTENSION TREATED WITH INTRAVENOUS PROSTANOIDS. RESPIR MED. 2021 NOV-DEC;189:106649. DOI: 10.1016/J.RMED.2021.106649. EPUB 2021 OCT 19. PMID: 34673343. SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22871314&pc=DSQ | 2916596-2025-05568 | 2022-03-15 04:00 | Death | THORATEC CORPORATION | 2025-08-22 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure; Thrombocytopenia; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD A COMPLICATED POST-OPERATIVE COURSE DUE TO HEPARIN-INDUCED THROMBOCYTOPENIA (HIT) POSITIVITY, WHICH LED TO RIGHT VENTRICULAR FAILURE (RVF), AND EVENTUALLY TO THE PATIENT PASSING AWAY DUE TO MULTISYSTEM ORGAN FAILURE (MSOF). THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) FUNCTIONED WITHOUT ANY ISSUES. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED, AND THE PUMP WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE PATIENT'S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICLE ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿ANTICOAGULATION¿), PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) RANGE, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. THE IFU ALSO STATES THAT THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM IS CONTRAINDICATED FOR PATIENTS WHO CANNOT TOLERATE, OR WHO ARE ALLERGIC TO, ANTICOAGULATION THERAPY. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22871379&pc=NVZ | 3004209178-2025-14273 | 2025-07-28 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-08-22 04:00 | NVZ | ADVISA SR MRI SURESCAN | Appropriate Term/Code Not Available | Insufficient Information | P980035 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. IT WAS FURTHER NOTED THAT THE "DEVICE WAS NOT SATISFACTORY AND CONTRIBUTED TO THE PATIENTS PASSING". THE IMPLANTABLE PULSE GENERATOR (IPG) REMAINS IN PATIENT. Manufacturer Narrative: B3: EVENT DATE IS NOT KNOWN. PLEASE SEE B5 FOR APPROXIMATE DATE RANGE, IF APPLICABLE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22871393&pc=PYX | 1220648-2025-45855 | 2025-08-02 04:00 | Death | ABIOMED, INC. | 2025-08-22 04:00 | PYX | IMPELLA RP FLEX | Device Sensing Problem | Hemolysis | P170011 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT IMPLANTED WITH AN IMPELLA RP FLEX UNDERWENT A PROCEDURE FOR AN INTERNAL JUGULAR LINE PLACEMENT. AFTER THE PROCEDURE, THE PUMP PLACEMENT SIGNAL DISAPPEARED AND A PLACEMENT SIGNAL NOT RELIABLE ALARM APPEARED. NO INTERVENTION WAS TAKEN REGARDING THE NOTED ISSUE. IT WAS ADDITIONALLY REPORTED THAT THE PATIENT DEVELOPED HEMOLYSIS DURING SUPPORT. AGAIN, NO INTERVENTION WAS TAKEN IN RESPONSE TO THE CONDITION. CARE WAS WITHDRAWN LATER THAT DAY UPON THE FAMILY'S REQUEST AND THE PATIENT PASSED AWAY. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22871486&pc=NVZ | 3008973940-2025-06360 | 2025-06-29 04:00 | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-08-22 04:00 | NVZ | AZURE XT DR MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P980035 |  | 1 | Event Description: IT WAS REPORTED BY THE PATIENT'S DAUGHTER THAT THE PATIENT WAS FOUND DECEASED ON THEIR CHAIR AND IS WONDERING WHY THEY WERE NOT ALERTED. IT WAS ALSO NOTED THAT "THE PARAMEDICS WERE LOOKING [THE PATIENT] OVER AND SAID THEY DIDN'T DETECT THE PACEMAKER. I'M CALLING TO SEE IF THE PACEMAKER STOPPED?". Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22871842&pc=OZD | 1220648-2025-45919 | 2025-07-24 04:00 | Death | ABIOMED, INC. | 2025-08-22 04:00 | OZD | IMPELLA 5.5 | Restricted Flow rate | Hemorrhage/Bleeding; Low Blood Pressure/ Hypotension; Thrombocytopenia; Epistaxis | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS INITIALLY IMPLANTED WITH AN IMPELLA CP FOR MECHANICAL CIRCULATORY SUPPORT (MCS) AND EXPERIENCED COMPLICATIONS INVOLVING THE PEEL-AWAY INTRODUCER. THE PATIENT WAS ESCALATED TO AN IMPELLA 5.5 AFTER SIX DAYS ON THE IMPELLA CP AND WAS PLACED ON EXTRACORPOREAL MEMBRANE OXYGENATION SUPPORT. DURING IMPELLA 5.5 SUPPORT, THE PATIENT EXPERIENCED SUCTION EVENTS, HYPOTENSION, HEPARIN-INDUCED THROMBOCYTOPENIA AND BLEEDING REQUIRING DISCONTINUATION OF ANTICOAGULATION. THE PATIENT ULTIMATELY EXPIRED ON SUPPORT. THREE DAYS AFTER STARTING IMPELLA 5.5 MCS, A SUCTION EVENT ALARM WAS TRIGGERED DUE TO HYPOTENSION AND PRELOAD. ONE UNIT OF BLOOD WAS ADMINISTERED. ECHOCARDIOGRAM AT BEDSIDE SHOWED THE IMPELLA 5.5 ANGLED TOWARD THE SEPTAL WALL MEASURING 4.25 CENTIMETERS AWAY FROM THE AORTIC VALVE AREA. THE CARE TEAM WAS MADE AWARE. THREE DAYS LATER, THERE WERE OCCASIONAL SUCTION ALARMS. FLUID WAS REPLETED AND A SECOND UNIT OF PACKED RED BLOOD CELLS WAS TRANSFUSED. THE PATIENT WAS IN A HYPOVOLEMIC STATE. THE NEXT DAY, THE PATIENT REMAINED SLIGHTLY SEATED, RESPONDED TO VERBAL COMMAND AND WAS MAKING NEEDS KNOWN. A SUCTION EVENT WAS AGAIN EXPERIENCED RELATED TO PRELOAD. LASCIVIOUS DRIP WAS HELD AND CENTRAL VENOUS PRESSURE WAS NOW 10. THROMBOCYTOPENIA WAS NOTED PERSISTING. HEPARIN-INDUCED THROMBOCYTOPENIA PANEL WAS SENT, AND HEPARIN WAS SUSPENDED WITH A NOTED PARTIAL THROMBOPLASTIN TIME IN THE 50¿S WITH PLANS TO START ARGATROBAN. HOWEVER, OF NOTE, THE PURGE COMPOSITION WAS FREE OF HEPARIN WITH 5% DEXTROSE IN WATER AND 25 MILLIEQUIVALENTS PER LITER OF SODIUM BICARBONATE. NOW 15 DAYS AFTER STARTING ON THE IMPELLA 5.5 MCS, SYSTEMIC ANTICOAGULATION HAS BEEN PUT ON HOLD AGAIN DUE TO EPISTAXIS AND A NOTED GASTROINTESTINAL BLEED CONFIRMED AS NEGATIVE BY COLONOSCOPY. HEPARIN WAS RESTARTED. A COUPLE OF DAYS LATER, SUCTION ALARMS WERE NOTED AGAIN. LASIX WAS DECREASED AND SUCTION ALARMS RESOLVED. THE PATIENT WAS MADE DO NOT RESUSCITATE/DO NOT INTUBATE AND HAD BEEN WEAK WITH VERY LIMITED MOBILITY AND HAD NOT GOTTEN OUT OF BED. THE FAMILY DISCUSSED PALLIATIVE CARE AND THE PATIENT EXPIRED ON SUPPORT AFTER 21 DAYS OF IMPELLA MCS. Manufacturer Narrative: MEDICAL SAFETY OFFICER REVIEWED THE EVENT AND NOTED THE IMPELLA 5.5 DEVICE IS UNLIKELY TO BE THE CAUSE OF THE DEMISE OUTCOME. THE PATIENT WAS IN A HYPOVOLEMIC STATE TRIGGERING THE SUCTION EVENTS. THE PATIENT WAS ON A HEPARIN FREE PURGE COMPOSITION AND HEPARIN-INDUCED THROMBOCYTOPENIA MAY OR MAY NOT HAVE BEEN INDUCED BY HEPARINIZING THE PATIENT WHEN THE IMPELLA WAS USED/INSERTED. EPISTAXIS AND OTHER BLEEDING APPEAR TO BE UNRELATED TO THE IMPELLA DEVICE. HOWEVER, ALTHOUGH UNLIKELY RELATED, THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA 5.5 AS AN ASSOCIATED FACTOR IN THE OVERALL EVENT OUTCOME. REFER TO RELATED MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-45918 FOR THE REPORT ON THE IMPELLA PEEL-AWAY INTRODUCER. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22871849&pc=OLR | MW5175109 | 2025-08-15 04:00 | Death | SOLEVA PHARMA LLC | 2025-08-22 04:00 | OLR | MUGARD ORAL RINSE 240ML/BTL | Insufficient Information | Insufficient Information |  |  | 1 | Event Description: PATIENT'S HUSBAND (B)(6) REPORTED THAT PATIENT PASSED AWAY ON (B)(6) 2025.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22872158&pc=NIM | 2124215-2025-58634 | 2013-08-01 04:00 | Death | SILK ROAD MEDICAL | 2025-08-22 04:00 | NIM | ENROUTE TRANSCAROTID STENT | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: IT WAS REPORTED VIA LITERATURE THAT ONE PATIENT WHO UNDERWENT THE TRANSCAROTID ARTERY REVASCULARIZATION (TCAR) PROCEDURE DIED IN THE 30-DAY POST PROCEDURE PERIOD. THIS RETROSPECTIVE REVIEW ANALYZED DATA FROM A PROSPECTIVELY MAINTAINED DATABASE OF 188 CONSECUTIVE PATIENTS FROM THREE LARGE ACADEMIC CENTERS WHO UNDERWENT TCAR TO TREAT CAROTID ARTERY DISEASE BETWEEN AUGUST 2013 AND MAY 2018. TO QUALIFY FOR THE TCAR PROCEDURE, PATIENTS HAD TO BE AT INCREASED RISK FOR CAROTID ENDARTERECTOMY (CEA) WITH AN ASYMPTOMATIC STENOSIS OF 70% OR GREATER OR SYMPTOMATIC STENOSIS OF 50% OR GREATER. SYMPTOMATIC WAS DEFINED AS PRESENTING WITH A MINOR OR NON-DISABLING STROKE, TRANSIENT ISCHEMIC ATTACK, OR AMAUROSIS FUGAX IN THE PRIOR 180 DAYS LEADING UP TO THE PROCEDURE. IN REVIEW OF THE DATA, THERE WAS ONE DEATH IN THE 30-DAY POST PROCEDURE PERIOD FOR SURGEON CASE NUMBER 22. THE PATIENT WAS SYMPTOMATIC. THE CAUSE OF DEATH WAS NOT REPORTED. Manufacturer Narrative: B2 - DATE OF DEATH: THE DATE OF DEATH WAS ESTIMATED TO 30 DAYS FOLLOWING THE EARLIEST KNOWN TCAR PROCEDURE INCLUDED IN THE COHORT. B3 - DATE OF EVENT: THE EVENT DATE WAS ESTIMATED TO THE EARLIEST KNOWN TCAR PROCEDURE INCLUDED IN THE COHORT. D6A - IMPLANT DATE: THE IMPLANT DATE WAS ESTIMATED TO THE EARLIEST KNOWN TCAR PROCEDURE INCLUDED IN THE COHORT. H6 - PATIENT CODES - THE PATIENT CODE E2401 WAS USED, AS THE CAUSES OF DEATH WERE NOT REPORTED IN THE LITERATURE REVIEW. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER PRODUCT SPECIFIC INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. KING, A. H., KUMINS, N. H., FOTEH, M. I., JIM, J., APPLE, J. M., & KASHYAP, V. S. (2019A). THE LEARNING CURVE OF TRANSCAROTID ARTERY REVASCULARIZATION. JOURNAL OF VASCULAR SURGERY, 70(2), 516-521. HTTPS://DOI.ORG/10.1016/J.JVS.2018.10.115.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22872169&pc=PYX | 1220648-2025-45857 | 2025-08-02 04:00 | Death | ABIOMED, INC. | 2025-08-22 04:00 | PYX | IMPELLA RP FLEX | Device Sensing Problem | No Clinical Signs, Symptoms or Conditions | P170011 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT ROUGHLY TWO DAYS INTO IMPELLA RP FLEX SUPPORT, A PLACEMENT SIGNAL NOT RELIABLE ALARM APPEARED. DESPITE THE FAILURE, BOTH THE MOTOR CURRENT AND FLOWS REMAINED UNAFFECTED. THE ALARM WAS SUBSEQUENTLY DISABLED AND SUPPORT CONTINUED UNINTERRUPTED. THE FOLLOWING DAY, THE DECISION WAS MADE TO WITHDRAW CARE AND THE PATIENT PASSED AWAY. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THE EVENT LOGS WERE RETURNED, AND AN EVALUATION IS UNDERWAY. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22872268&pc=OZD | 1220648-2025-45782 | 2025-07-28 04:00 | Death | ABIOMED, INC. | 2025-08-22 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia; Tachycardia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT IN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. IT WAS REPORTED THAT DURING IMPELLA CP IMPLANT, THE PATIENT EXPERIENCED WIDE COMPLEX TACHYCARDIA. THE PATIENT WAS DEFIBRILLATED MULTIPLE TIMES TO COUNTERACT THE CONDITION. IT WAS NOTED THAT THE PATIENT HAD UNDERLYING ATRIAL FIBRILLATION WITH RIGHT BUNDLE BRANCH BLOCK LEADING UP TO THE EVENT. IT WAS ADDITIONALLY REPORTED THAT DESPITE HAVING ABSENT BILATERAL LOWER EXTREMITY POSTERIOR TIBIAL AND DORSALIS PEDIS PULSES PRIOR TO IMPELLA IMPLANT, THE CONDITION CONTINUED TO WORSEN WITH THE EXTREMITIES BECOMING VERY MOTTLED. SUPPORT CONTINUED DESPITE THE CONDITION FOR TWO DAYS, AT WHICH POINT THE PATIENT WAS MOVED TO COMFORT CARE. THE PATIENT EXPIRED FOLLOWING PUMP REMOVAL. Manufacturer Narrative: ALTHOUGH THE PATIENT'S PRESENTING CONDITION MAY BE MORE LIKELY HAVE IMPACTED THE EVENT THE IMPELLA CANNOT BE EXCLUDED AS A POSSIBLE CONTRIBUTING FACTOR IN THE PATIENT'S DEMISE. PRODUCT STATUS: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22872316&pc=OZD | 1220648-2025-45861 | 2025-08-02 04:00 | Death | ABIOMED, INC. | 2025-08-22 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT ROUGHLY TWO DAYS INTO IMPELLA 5.5 SUPPORT, A COMPUTED TOMOGRAPHY SCAN REVEALED THAT THE PATIENT HAD EXPERIENCED AN EMBOLIC STROKE. ADMINISTRATION OF LOW DOSE HEPARIN WAS APPROVED FOLLOWING THE STROKE. THE PATIENT WAS ALSO ON VENOARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION AND RIGHT VENTRICULAR ASSIST DEVICE. LATER, THE FAMILY WITHDREW CARE AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22872662&pc=NTE | 2124215-2025-58725 | 2009-03-03 05:00 | Death | SILK ROAD MEDICAL | 2025-08-22 04:00 | NTE | ENROUTE TRANSCAROTID NEUROPROTECTION SYSTEM | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: IT WAS REPORTED VIA LITERATURE THAT PATIENTS DIED WITHIN 24 HOURS OF THEIR RESPECTIVE TRANSCAROTID ARTERY REVASCULARIZATION (TCAR) PROCEDURES. THIS STUDY REVIEWED DATA FROM A TOTAL OF (B)(4) TCAR PROCEDURES PERFORMED BY (B)(4) PHYSICIANS. THE PROCEDURES WERE PERFORMED BETWEEN (B)(6) 2009 AND (B)(6) 2020. THE PROCEDURE DATA WAS COLLECTED IN AN INTERNATIONAL QUALITY ASSURANCE DATABASE, AND WAS ANALYZED TO DETERMINE THE POINT AT WHICH THE PHYSICIANS BECAME PROFICIENT IN PERFORMING THE TCAR PROCEDURE. "PROFICIENCY" WAS DEFINED IN TERMS OF PROCEDURAL PERFORMANCE MEASURES AND THEN DETERMINED THE MINIMUM NUMBER OF CASES REQUIRED TO ACHIEVE PROFICIENCY IN TCAR. FOUR PROCEDURAL PERFORMANCE MEASURES WERE RECORDED FOR EACH PROCEDURE: FLOW-REVERSAL TIME, FLUOROSCOPY TIME, CONTRAST VOLUME, AND TOTAL SKIN-TO-SKIN TIME. THE STUDY CONCLUDED THAT FLOW-REVERSAL TIME AND SKIN-TO-SKIN TIME ARE APPROPRIATE PERFORMANCE MEASURES FOR ESTABLISHING THE LEVEL OF EXPERTISE OF PHYSICIANS AS THEY ACQUIRE SKILLS TO PERFORM TCAR. IN REVIEW OF THE DATA, IT WAS REVEALED THAT (B)(4) PATIENTS DIED WITHIN 24 HOURS OF THEIR RESPECTIVE TCAR PROCEDURES. THE CAUSE OF DEATH WAS NOT REPORTED FOR ANY CASE. NO FURTHER INFORMATION WAS PROVIDED TO BOSTON SCIENTIFIC. Manufacturer Narrative: B2 - DATE OF DEATH: THE DEATH DATE WAS ESTIMATED TO THE EARLIEST KNOWN TCAR PROCEDURE INCLUDED IN THE COHORT. B3 - DATE OF EVENT: THE EVENT DATE WAS ESTIMATED TO THE EARLIEST KNOWN TCAR PROCEDURE INCLUDED IN THE COHORT. H6 - PATIENT CODES: E2401 WAS USED, AS THE CAUSE OF DEATH WAS NOT REPORTED IN ANY CASE. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. LAL, B., MAYORGA-CARLIN, M., KASHYAP, V., JORDAN, W., MUKHERJEE, D., CAMBRIA, R., MOORE, W., NEVILLE, R., ECKSTEIN, H.-H., SAHOO, S., MACDONALD, S., & SORKIN, J. (2022). LEARNING CURVE AND PROFICIENCY METRICS FOR TRANSCAROTID ARTERY REVASCULARIZATION. JOURNAL OF VASCULAR SURGERY, 75(6). HTTPS://DOI.ORG/10.1016/J.JVS.2021.12.073.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22872908&pc=DSQ | 2916596-2025-05584 | 2022-06-12 04:00 | Death | THORATEC CORPORATION | 2025-08-22 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Gastrointestinal Hemorrhage | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS CHRONIC GASTROINTESTINAL (GI) BLEED WHICH LED TO FAILURE TO THRIVE. ANTICOAGULANT WAS HELD AND BLOOD TRANSFUSION WAS GIVEN. THE DEATH WAS NOT CONSIDERED DEVICE RELATED. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, REV. D IS CURRENTLY AVAILABLE. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU LISTS POTENTIAL ADVERSE EVENTS, INCLUDING BLEEDING AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿ANTICOAGULATION¿), PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) RANGE, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22872927&pc=NVN | 2017865-2025-98291 |  | Death | ABBOTT | 2025-08-22 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | Cardiac Tamponade; Cardiac Perforation | P960013 |  | 1 | Event Description: IT WAS REPORTED THROUGH EUROPEAN SOCIETY OF CARDIOLOGY, THREE DAYS POST INITIAL IMPLANT, THE PATIENT WAS ADMITTED TO THE HOSPITAL DUE TO TAMPONADE ASSOCIATED WITH CARDIAC PERFORATION BY THE RIGHT VENTRICLE (RV) LEAD. PERICARDIOCENTESIS WAS PERFORMED, HOWEVER, THE PATIENT PASSED AWAY ELEVEN DAYS LATER. THE REPORTED CAUSE OF DEATH WAS TAMPONADE. DETAILS ARE LISTED IN THE ARTICLE TITLED "PROCEDURAL OUTCOME AND FOLLOW-UP OF STYLET-DRIVEN LEADS COMPARED WITH LUMENLESS LEADS FOR LEFT BUNDLE BRANCH AREA PACING". Manufacturer Narrative: FURTHER INFORMATION WAS REQUESTED BUT NOT RECEIVED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22872998&pc=NVN | 2017865-2025-98385 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-08-22 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. IT WAS REPORTED THAT THE CAUSE OF DEATH IS UNKNOWN. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: AS RECEIVED, A COMPLETE LEAD WAS RETURNED IN ONE PIECE. VISUAL EXAMINATION OF THE LEAD DID NOT FIND ANY ANOMALIES EXCEPT FOR PROCEDURAL DAMAGE. ELECTRICAL TESTING DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22873003&pc=NVN | 2017865-2025-98384 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-08-22 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. IT WAS REPORTED THAT THE CAUSE OF DEATH IS UNKNOWN. NO ADDITIONAL INFORMATION WAS REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22873067&pc=DSQ | 2916596-2025-05452 | 2025-08-07 04:00 | Death | THORATEC CORPORATION | 2025-08-22 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS THERAPY WAS STOPPED ACCORDING TO PATIENTS WISH. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. THE DEVICE WAS WORKING AS EXPECTED FOR THE WHOLE TIME OF SUPPORT. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. DUE TO HOSPITAL PRIVACY POLICY, NO ADDITIONAL INFORMATION WILL BE PROVIDED BY THE ACCOUNT. THE HEARTMATE 3 LVAS IFU, REV. B, IS ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22873114&pc=LWP | 2017865-2025-98383 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-08-22 04:00 | LWP | ASSURITY MRI | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P140033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. IT WAS REPORTED THAT THE CAUSE OF DEATH IS UNKNOWN. NO ADDITIONAL INFORMATION WAS REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22873202&pc=DXF | 2124215-2025-58981 | 2025-08-04 04:00 | Death | BAYLIS MEDICALE CIE INC | 2025-08-22 04:00 | DXF | VERSACROSS CONNECT ACCESS SOLUTION FOR FARADRIVE | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Tachycardia; Cardiac Tamponade; Cardiac Perforation; Pericardial Effusion | K150709 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT DEATH OCCURRED. DURING AN ATRIAL FIBRILLATION EPS ABLATION PROCEDURE, A VERSACROSS CONNECT FOR FARADRIVE WAS SELECTED FOR USE. THE PHYSICIAN GAINED FEMORAL ACCESS ON THE LEFT FEMORAL VEIN TWICE AND RIGHT FEMORAL VEIN ONCE. A NON-BOSTON SCIENTIFIC CORONARY SINUS AND AN INTRACARDIAC ECHOCARDIOGRAPHY (ICE) CATHETER WERE PLACED FROM LEFT SIDED ACCESS. THE TRANSPETAL WAS BEING PREPPED WITH VERSACROSS CONNECT FOR FARADRIVE. THE RADIO FREQUENCY (RF) WIRE FOLLOWED BY FARADRIVE LOADED WITH VERSACROSS CONNECT DILATOR WAS ADVANCED THROUGH RIGHT FEMORAL VEIN. A MID-MID TRANSEPTAL PUNCTURE WAS ACHIEVED WITH ICE VERIFICATION, WITH NO ISSUES REPORTED. THE PHYSICIAN ADVANCED VERSACROSS WIRE, DILATOR, AND FARADRIVE SHEATH ACROSS TO LEFT ATRIUM, FLOATING IN BLOOD POOL CONFIRMED WITH ICE. THE DILATOR AND WIRE WERE REMOVED FROM FARADRIVE. A NON-BOSTON SCIENTIFIC MAPPING (OCTARAY) CATHETER WAS ADVANCED TO LEFT ATRIUM AND MAPPING OF THE LEFT PULMONARY VEINS STARTED. DURING MAPPING, THE PATIENT BECAME UNSTABLE, EXTREMELY HYPOTENSIVE AND TACHYCARDIC. ICE CONFIRMED PERICARDIAL EFFUSION THAT WAS NOT IDENTIFIED AT BASELINE. PERICARDIAL EFFUSION ACCUMULATED PRIMARILY AROUND RIGHT VENTRICLE. THE CAUSE OF TAMPONADE WAS IDENTIFIED TO BE PERFORATION OF THE LEFT ATRIAL APPENDAGE. A PERICARDIOCENTESIS AND PERICARDIAL DRAIN PLACEMENT WAS PERFORMED. HEMODYNAMIC SUPPORT BY ANESTHESIA AND AUTOTRANSFUSION FOLLOWED TO STABILIZE PATIENT. THE PATIENT WAS ALREADY UNDER GENERAL ANESTHESIA AND INTUBATED. THE PATIENT WAS TRANSFERRED FROM EP PROCEDURAL LAB TO CARDIAC THORACIC SURGERY. THE PATIENT DID NOT SURVIVE OPEN HEART SURGERY. THE DEVICE IS NOT EXPECTED TO BE RETURNED FOR ANALYSIS (DISPOSED). Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22873431&pc=FMI | 9681413-2025-00005 | 2025-07-29 04:00 | Death | TERUMO EUROPE N.V | 2025-08-22 04:00 | FMI | K-PACK II | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K984576 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT THE DEATH OF A CHILD OCCURRED. THE PROCEDURE OUTCOME WAS UNKNOWN. Manufacturer Narrative: D4: LOT #: 2208029 OR 2209046. E3: OCCUPATION: MU LOGISTICS & IM PRIMARY AND SECONDARY PACKAGING. H6: INVESTIGATION FINDINGS - CODE 3221 IS BASED UPON THE EVALUATION OF USER FACILITY INFORMATION AND NO SAMPLE RETURNED; CODE 213 IS BASED UPON TESTING OF THE RETENTION SAMPLE. H6 - INVESTIGATION CONCLUSION - CODE 4315 IS BASED UPON EVALUATION OF USER FACILITY INFORMATION AND NO DEVICE RETURN; CODE 67 IS BASED UPON EVALUATION OF THE RETENTION SAMPLE. THE BATCH RECORDS FOR LOT 2208029 AND 2209046 (B)(6) HAVE BEEN REVIEWED. DURING NEEDLE ASSEMBLY, AS AN IN-PROCESS CONTROL, (B)(4) ASSEMBLED NEEDLES (NEEDLE ASSEMBLED WITH CASE) ARE VISUALLY INSPECTED TWICE PER SHIFT (8H SHIFT DURING THE WEEK AND A 12H SHIFT DURING THE WEEKEND) FOR FOREIGN MATTER AND DAMAGES TO THE CASE. NO DEVIATIONS OR ABNORMALITIES RELATED TO YOUR COMPLAINT WERE NOTED DURING PRODUCTION. DURING PACKAGING, AS AN IN-PROCESS CONTROL, (B)(4) LABELLED PRODUCTS ARE VISUALLY INSPECTED TWICE PER SHIFT (8H SHIFT DURING THE WEEK AND A 12H SHIFT DURING THE WEEKEND) FOR FOREIGN MATTER AND DAMAGES TO THE CASE AND CAP. NO DEVIATIONS OR ABNORMALITIES RELATED TO YOUR COMPLAINT WERE NOTED DURING PRODUCTION. DURING THE FINAL INSPECTION (PART OF THE RELEASE TESTING) OF (B)(4) PIECES, NO DEFECTS WERE DETECTED. NO RELEVANT DEVIATIONS OR ABNORMALITIES WERE NOTED IN THE BATCH RECORD AND DURING THE ABOVE-MENTIONED INSPECTIONS WHICH RELATE TO YOUR COMPLAINT. THE LAL AND STERILIZATION CYCLE FOR BOTH LOTS WERE EXECUTED ACCORDING TO SPECIFICATIONS AND NO OBSERVATIONS HAVE BEEN FOUND FOR THESE TESTS. THE BIOBURDEN MONITORING PROGRAM OF 2022 WAS CHECKED: MONITORING RESULTS IN LINE WITH PREVIOUS YEARS MONITORING SAMPLING OF (B)(6) 2022 (DATES CLOSEST TO MANUFACTURING DATE) DID NOT SHOW DEVIATING RESULTS. ADDITIONALLY, NO MATERIALS OF ANIMAL ORIGIN HAVE BEEN USED. SAMPLE INVESTIGATION: NOT RETURNED. RETENTION SAMPLES: ALL REMAINING RETENTION SAMPLES OF LOT 2208029 HAVE BEEN VISUALLY INSPECTED BY THE NAKED EYE ON (B)(6) 2025 FOR FOREIGN MATTER AND DAMAGES WHICH COULD IMPACT THE STERILITY. NO DEFECTS HAVE BEEN FOUND. PRODUCTION INVESTIGATION: NON-CONFORMITIES. THE NON-CONFORMITIES RELATED TO THE FINISHED PRODUCT BATCH AND THE HALF-FINISHED PRODUCT BATCHES OF NEEDLE ASSEMBLY HAVE BEEN CHECKED. NO RELATED NON-CONFORMITIES WERE FOUND. CAPA: THERE IS NO CAPA RELATED TO ISSUES WITH THE STERILE BARRIER SYSTEM OF OUR PRODUCTS. BASED ON OUR INVESTIGATION INCLUDING A BATCH RECORD REVIEW AND RETENTION SAMPLE INVESTIGATION, NO ROOT CAUSE RELATED TO OUR PRODUCTION PROCESS HAS BEEN FOUND. FOR INFORMATION REGARDING THE IMPORTER REPORT FOR THIS EVENT, PLEASE REFER TO SECTION H10.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22873432&pc=FMI | 9681413-2025-00006 | 2025-07-29 04:00 | Death | TERUMO EUROPE N.V | 2025-08-22 04:00 | FMI | K-PACK II | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K984576 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT THE DEATH OF A CHILD OCCURRED. THE PROCEDURE OUTCOME WAS UNKNOWN. Manufacturer Narrative: D4: LOT #: 2208029 OR 2209046. E3: OCCUPATION: MU LOGISTICS & IM PRIMARY AND SECONDARY PACKAGING. H6: INVESTIGATION FINDINGS - CODE 3221 IS BASED UPON THE EVALUATION OF USER FACILITY INFORMATION AND NO SAMPLE RETURNED; CODE 213 IS BASED UPON TESTING OF THE RETENTION SAMPLE. H6 - INVESTIGATION CONCLUSION - CODE 4315 IS BASED UPON EVALUATION OF USER FACILITY INFORMATION AND NO DEVICE RETURN; CODE 67 IS BASED UPON EVALUATION OF THE RETENTION SAMPLE. THE BATCH RECORDS FOR LOT 2208029 AND 2209046 (B)(6) HAVE BEEN REVIEWED. DURING NEEDLE ASSEMBLY, AS AN IN-PROCESS CONTROL, 40 ASSEMBLED NEEDLES (NEEDLE ASSEMBLED WITH CASE) ARE VISUALLY INSPECTED TWICE PER SHIFT (8H SHIFT DURING THE WEEK AND A 12H SHIFT DURING THE WEEKEND) FOR FOREIGN MATTER AND DAMAGES TO THE CASE. NO DEVIATIONS OR ABNORMALITIES RELATED TO YOUR COMPLAINT WERE NOTED DURING PRODUCTION. DURING PACKAGING, AS AN IN-PROCESS CONTROL, 40 LABELLED PRODUCTS ARE VISUALLY INSPECTED TWICE PER SHIFT (8H SHIFT DURING THE WEEK AND A 12H SHIFT DURING THE WEEKEND) FOR FOREIGN MATTER AND DAMAGES TO THE CASE AND CAP. NO DEVIATIONS OR ABNORMALITIES RELATED TO YOUR COMPLAINT WERE NOTED DURING PRODUCTION. DURING THE FINAL INSPECTION (PART OF THE RELEASE TESTING) OF (B)(4) PIECES, NO DEFECTS WERE DETECTED. NO RELEVANT DEVIATIONS OR ABNORMALITIES WERE NOTED IN THE BATCH RECORD AND DURING THE ABOVE-MENTIONED INSPECTIONS WHICH RELATE TO YOUR COMPLAINT. THE LAL AND STERILIZATION CYCLE FOR BOTH LOTS WERE EXECUTED ACCORDING TO SPECIFICATIONS AND NO OBSERVATIONS HAVE BEEN FOUND FOR THESE TESTS. THE BIOBURDEN MONITORING PROGRAM OF 2022 WAS CHECKED: MONITORING RESULTS IN LINE WITH PREVIOUS YEARS MONITORING SAMPLING OF (B)(6) 2022 (DATES CLOSEST TO MANUFACTURING DATE) DID NOT SHOW DEVIATING RESULTS. ADDITIONALLY, NO MATERIALS OF ANIMAL ORIGIN HAVE BEEN USED. SAMPLE INVESTIGATION: NOT RETURNED. RETENTION SAMPLES: ALL REMAINING RETENTION SAMPLES OF LOT 2208029 HAVE BEEN VISUALLY INSPECTED BY THE NAKED EYE ON (B)(6) 2025 FOR FOREIGN MATTER AND DAMAGES WHICH COULD IMPACT THE STERILITY. NO DEFECTS HAVE BEEN FOUND. PRODUCTION INVESTIGATION: NON-CONFORMITIES. THE NON-CONFORMITIES RELATED TO THE FINISHED PRODUCT BATCH AND THE HALF-FINISHED PRODUCT BATCHES OF NEEDLE ASSEMBLY HAVE BEEN CHECKED. NO RELATED NON-CONFORMITIES WERE FOUND. CAPA: THERE IS NO CAPA RELATED TO ISSUES WITH THE STERILE BARRIER SYSTEM OF OUR PRODUCTS. BASED ON OUR INVESTIGATION INCLUDING A BATCH RECORD REVIEW AND RETENTION SAMPLE INVESTIGATION, NO ROOT CAUSE RELATED TO OUR PRODUCTION PROCESS HAS BEEN FOUND. FOR INFORMATION REGARDING THE IMPORTER REPORT FOR THIS EVENT, PLEASE REFER TO SECTION H10.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22873750&pc=DSQ | 3007042319-2025-01939 | 2025-07-01 04:00 | Death | HEARTWARE, INC. | 2025-08-22 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - DRIVELINE | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Sepsis; Septic Shock | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS HOSPITALIZED FOR A SEVERE VENTRICULAR ASSIST DEVICE (VAD) DRIVELINE EXIT SITE INFECTION. THE PATIENT RECEIVED INTRAVENOUS FLUIDS, ANTIBIOTICS, AND VASOPRESSORS AS MEDICAL INTERVENTIONS. DESPITE THESE TREATMENTS, THE PATIENT DID NOT IMPROVE. THE PATIENT EXPERIENCED SEPTIC SHOCK AND WAS FOUND TO HAVE AN ASCENDING AND EXTERNAL BACTERIAL DRIVELINE INFECTION. DIAGNOSTIC TESTING INCLUDED WOUND AND BLOOD CULTURES, WHICH WERE POSITIVE AND SHOWED AN ELEVATED WHITE BLOOD CELL COUNT, AND A COMPUTED TOMOGRAPHY (CT) SCAN THAT DEMONSTRATED INFECTION. THE PATIENT DECLINED FURTHER AND WAS PLACED ON COMFORT CARE. THE PATIENT SUBSEQUENTLY DIED FROM SEPSIS. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22873763&pc=NGV | 2135147-2025-04706 | 2021-09-01 04:00 | Death | ABBOTT MEDICAL | 2025-08-22 04:00 | NGV | AMPLATZER AMULET | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P200049 |  | 1 | Event Description: THE ARTICLE, "A REAL-WORLD COMPARISON OF LEFT ATRIAL APPENDAGE OCCLUSION USING TWO COMMERCIALLY AVAILABLE DEVICES", WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE, MULTICENTER STUDY TO COMPARE THE EARLY IMPLEMENTATION OF THE AMULET TO THE WATCHMAN FLX AT A LARGE HEALTHCARE SYSTEM IN THE UNITED STATES. DEVICES INCLUDED IN THE STUDY WERE AMPLATZER AMULET AND WATCHMAN FLX. THE ARTICLE CONCLUDED THAT DESPITE EARLY EXPERIENCE, THE AMULET DEVICE EXHIBITED COMPARABLE RESULTS TO THE WATCHMAN FLX IN A REAL-WORLD SETTING. THE LEARNING CURVE WITH A NEW DEPLOYMENT TECHNIQUE LIKELY ACCOUNTS FOR LONGER PROCEDURAL TIMES WITH THE AMULET. [(B)(6)]. THE TIME FRAME OF THE STUDY WAS FROM SEPTEMBER 2021 TO JUNE 2023. A TOTAL OF 325 PATIENTS WERE INCLUDED IN THE STUDY, OF WHICH 131 (40.3%) RECEIVED AN ABBOTT DEVICE. IN THE AMULET GROUP, THE AVERAGE AGE WAS 76.9 YEARS AND THE MAJORITY GENDER WAS MALE. IN THE WATCHMAN FLX GROUP, THE AVERAGE AGE WAS 76.1 YEARS AND THE MAJORITY GENDER WAS MALE. COMORBIDITIES INCLUDED HYPERTENSION, DIABETES MELLITUS, CORONARY ARTERY DISEASE, PRIOR MYOCARDIAL INFARCTION, PRIOR CEREBROVASCULAR ACCIDENT, HEART FAILURE, CHRONIC KIDNEY DISEASE, MAJOR BLEEDING HISTORY (EPISTAXIS, GASTROINTESTINAL, HEMATURIA, INTRACRANIAL, MUSCULOSKELETAL, PERICARDIAL, RETINAL, SKIN ECCHYMOSIS), ANEMIA, RECURRENT FALLS, AND PAROXYSMAL ATRIAL FIBRILLATION. Manufacturer Narrative: IT IS UNKNOWN IF THE DEVICE IS RETURNING FOR ANALYSIS. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: A REAL-WORLD COMPARISON OF LEFT ATRIAL APPENDAGE OCCLUSION USING TWO COMMERCIALLY AVAILABLE DEVICES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22874360&pc=FTL | 9615742-2025-01027 |  | Death | SOFRADIM PRODUCTION SAS | 2025-08-22 04:00 | FTL | PARIETEX | Adverse Event Without Identified Device or Use Problem | Distress; Unspecified Tissue Injury |  |  | 1 | Event Description: THE PATIENT'S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR THERAPEUTIC TREATMENT OF A HERNIA. IT WAS REPORTED THAT AFTER IMPLANT, THE PATIENT EXPERIENCED EMOTIONAL DISTRESS, DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED REVISION SURGERY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22874486&pc=NPT | 2015691-2025-06917 | 2025-07-31 04:00 | Death | EDWARDS LIFESCIENCES | 2025-08-22 04:00 | NPT | COMMANDER DELIVERY SYSTEM | Difficult to Remove; Failure to Advance | Perforation of Vessels; Foreign Body In Patient | P140031 |  | 1 | Event Description: AS REPORTED BY THE FIELD CLINICAL SPECIALIST (FCS), DURING A TRANSFEMORAL TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) PROCEDURE WITH A 26 MM SAPIEN 3 ULTRA RESILIA VALVE, UPON EXITING THE ESHEATH+, THE VALVE BECAME STUCK IN THE INFRARENAL AORTA. AN AORTOGRAM WAS PERFORMED SHOWING PERFORATION OF THE RIGHT COMMON ILIAC ARTERY. ATTEMPTS WERE MADE TO REMOVE THE DELIVERY SYSTEM AND VALVE THROUGH THE ESHEATH+, BUT THIS WAS UNSUCCESSFUL DUE TO THE ANGLE OF ENTRY. A DECISION WAS MADE TO PARTIALLY DEPLOY THE VALVE IN THE ABDOMINAL AORTA. THE DELIVERY SYSTEM BALLOON WAS THEN LEFT INFLATED TO TAMPONADE THE FEMORAL VESSELS. A COVERED STENT WAS PLACED IN THE RIGHT ILIAC ARTERY INTO THE COMMON FEMORAL ARTERY. THE DELIVERY SYSTEM BALLOON WAS DEFLATED AND A SECOND AORTOGRAM WAS PERFORMED ON THE DESCENDING AORTA. THE AORTOGRAM SHOWED A PERFORATION OF THE INFRARENAL AORTA. IT WAS BELIEVED TO HAVE OCCURRED WHEN ATTEMPTING TO ADVANCE THE DELIVERY SYSTEM WITH VALVE THROUGH THE ANATOMY. A DECISION WAS MADE TO CONVERT TO AN OPEN REPAIR. THE PATIENT CODED MULTIPLE TIMES AND THEN PASSED AWAY IN THE HYBRID OPERATING ROOM (OR). Manufacturer Narrative: THIS IS ONE OF TWO MANUFACTURER REPORTS BEING SUBMITTED FOR THIS CASE. PLEASE SEE MANUFACTURER REPORT NUMBER 2015691-2025-06915 FOR DETAILS OF THE OTHER EVENT. THE INVESTIGATION IS ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22874566&pc=DYB | 2029046-2025-02795 | 2025-07-28 04:00 | Death | BIOSENSE WEBSTER INC | 2025-08-22 04:00 | DYB | VIZIGO | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Ventricular Fibrillation; Cardiac Tamponade | K170997 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN ATRIAL FIBRILLATION (AFIB) CARDIAC ABLATION PROCEDURE AND DURING THE TRANSSEPTAL PHASE WITH VIZIGO SHEATH, THE PATIENT EXPERIENCED PRESSURE DROP AND A LARGE PERICARDIAL EFFUSION TREATED WITH PERICARDIOCENTESIS. THE PATIENT WENT INTO VENTRICULAR FIBRILLATION AND CARDIOVERSION/CPR UNSUCCESSFUL. THE PATIENT PASSED AWAY. IT WAS REPORTED THAT AFTER THE PHYSICIAN PERFORMED A TRANSSEPTAL PROCEDURE AND INSERTED THE VIZIGO SHEATH INTO THE LEFT ATRIUM, THE PATIENT HAD A PRESSURE DROP AND IT WAS CONFIRMED BY INTRA-CARDIAC ULTRASOUND THAT THE PATIENT HAD A LARGE PERICARDIAL EFFUSION. THEY CALLED IN AN INTERVENTIONALIST WHO PERFORMED A PERICARDIOCENTESIS. THE PATIENT THEN WENT INTO VENTRICULAR FIBRILLATION. ATTEMPTS AT CARDIOVERSION WERE UNSUCCESSFUL, AND CPR WAS UNSUCCESSFUL. THE PATIENT THEN PASSED AWAY. ADDITIONAL INFORMATION WAS RECEIVED. THE PHYSICIAN¿S OPINION ON THE CAUSE OF THIS ADVERSE EVENT WAS THAT BELIEVES THE VIZIGO SHEATH CAUSED A PERFORATION ON THE POSTERIOR WALL. NO CATHETER EXCHANGES OCCURRED DURING THE PROCEDURE. ONE TRANSSEPTAL PUNCTURE SITE WAS PERFORMED DURING THE PROCEDURE. NO ABLATIONS PERFORMED. THE CORRECT CATHETER SETTINGS WERE SELECTED ON THE GENERATOR. THE NGEN GENERATOR WAS CONNECTED BUT NOT USED. NO ISSUES WITH FLOW RATE CHANGE AT THE START OF ABLATION. Manufacturer Narrative: AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. A DEVICE HISTORY RECORD REVIEW WAS PERFORMED FOR THE FINISHED DEVICE LOT NUMBER 00002940 AND NO INTERNAL ACTION RELATED TO THE COMPLAINT WAS FOUND DURING THE REVIEW. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. HOWEVER, IF THE PRODUCT IS RECEIVED AT A LATER DATE, THE INVESTIGATION WILL BE UPDATED AS APPLICABLE. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. MANUFACTURER'S REFERENCE NUMBER: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22874755&pc=MFE | 2029214-2025-01879 | 2024-11-21 05:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-08-22 04:00 | MFE | ONYX | Adverse Event Without Identified Device or Use Problem | Hematoma |  |  | 1 | Event Description: DAVIES, J. M., KNOPMAN, J., MOKIN, M., HASSAN, A. E., HARBAUGH, R. E., KHALESSI, A., FIEHLER, J., GROSS, B. A., GRANDHI, R., TARPLEY, J., SIVAKUMAR, W., BAIN, M., CROWLEY, R. W., LINK, T. W., FRASER; THE NEW ENGLAND JOURNAL OF MEDICINE; 2024; 391(20), 1890¿1900; ADJUNCTIVE MIDDLE MENINGEAL ARTERY EMBOLIZATION FOR SUBDURAL HEMATOMA; HTTPS://DOI.ORG/10.1056/NEJMOA2313472 LITERATURE WAS REVIEWED REGARDING MIDDLE MENINGEAL ARTERY EMBOLIZATION PLUS SURGERY AS COMPARED WITH SURGERY ALONE IN PATIENTS WITH A LARGE SUBDURAL HEMATOMA WITH NEUROLOGIC SYMPTOMS. MULTIPLE MANUFACTURER¿S DEVICES WERE IMPLANTED IN THE STUDY POPULATION. THE FOLLOWING MEDTRONIC DEVICES WERE USED: ONYX EMBOLIC AGENT. DEATHS OCCURRED IN THE STUDY POPULATION; 10 PATIENTS TREATED WITH ONYX DIED, OF WHICH 9 HAD NEUROLOGIC DEATH. MOST NEUROLOGIC DEATHS WERE RELATED TO THE UNDERLYING SUBDURAL HEMATOMA (IN 8 PATIENTS). THE REMAINING NEUROLOGIC DEATH WAS RELATED TO SURGERY AND NOT TO EMBOLIC AGENT. THE CLINICAL EVENTS COMMITTEE REVIEWED EACH DEATH, IDENTIFIED THE ADVERSE EVENT THAT MOST ACCURATELY REPRESENTED THE CAUSE OF DEATH, AND ASSESSED THE RELATIONSHIP TO THE PROCEDURE OR DEVICE. ON THE BASIS OF THIS ASSESSMENT, NONE OF THE DEATHS IN THE TREATMENT GROUP WERE DEEMED TO BE RELATED TO MIDDLE MENINGEAL ARTERY EMBOLIZATION OR THE EMBOLIC A GENT. AMONG ALL PATIENTS ADVERSE EVENTS / DEVICE PRODUCT PERFORMANCE ISSUES INCLUDED: REFLUX TO A NONTARGET VESSEL (E.G., PETROSAL OR COLLATERALS) WAS OBSERVED IN 3 PATIENTS (1.6%), NONE OF THESE PATIENTS HAD ANY REFLU X-RELATED COMPLICATIONS. 21 OF 177 PATIENTS HAD FUNCTIONAL DETERIORATION. 4 PATIENTS (2.0%) IN THE TREATMENT GROUP HAD HAD A SERIOUS ADVERSE EVENT, INCLUDING DISABLING STROKE IN 2 PATIENTS, THAT WAS ADJUDICATED BY THE CLINICAL EVENTS COMMITTEE TO BE RELATED TO MIDDLE MENINGEAL ARTERY EMBOLIZATION ALONE. NONE OF THE PATIENTS WITH SERIOUS ADVERSE EVENTS HAD ANY REFLUX-RELATED COMPLICATIONS, AND NO CASES OF IPSILATERAL VISION LOSS OCCURRED IN THE TREATMENT GROUP. 8 OF 197 PATIENTS (4.1%) IN THE TREATMENT GROUP UNDERWENT REOPERATION FOR HEMATOMA RECURRENCE OR PROGRESSION. NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. Manufacturer Narrative: CITATION: DAVIES, J. M., KNOPMAN, J., MOKIN, M., HASSAN, A. E., HARBAUGH, R. E., KHALESSI, A., FIEHLER, J., GROSS, B. A., GRANDHI, R., TARPLEY, J., SIVAKUMAR, W., BAIN, M., CROWLEY, R. W., LINK, T. W., FRASER,. ADJUNCTIVE MIDDLE MENINGEAL ARTERY EMBOLIZATION FOR SUBDURAL HEMATOMA. THE NEW ENGLAND JOURNAL OF MEDICINE 391(20), 1890¿1900 2024. HTTPS://DOI.ORG/10.1056/NEJMOA2313472 EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT. NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. A2. AGE AT TIME OF EVENT AND A3A SEX ARE AVERAGE OF COHORT. SEE ATTACHMENTS FOR LITERATURE ARTICLE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22874985&pc=NGV | 2135147-2025-04714 | 2025-07-29 04:00 | Death | ABBOTT MEDICAL | 2025-08-22 04:00 | NGV | AMPLATZER AMULET | Patient-Device Incompatibility | Cardiac Tamponade; Pericardial Effusion | P200049 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 22MM AMPLATZER AMULET LEFT ATRIAL APPENDAGE (LAA) OCCLUDER WAS CHOSEN FOR IMPLANTATION UTILIZING A 12F AMPLATZER TORQVUE 45X45 DELIVERY SYSTEM. TRANSSEPTAL PUNCTURE WAS MID/MID. HEPARIN ADMINISTERED, ACT NOT MEASURED. LEFT ATRIAL PRESSURE WAS 16MMHG. THE DEVICE WAS IMPLANTED WITHOUT ISSUE AND NO RECAPTURES. MULTIPLE WIRE OR DELIVERY SHEATH EXCHANGES WERE PERFORMED IN THE UPPER PULMONARY VEIN. THE PATIENT REMAINED HEMODYNAMICALLY STABLE THROUGHOUT THE PROCEDURE. IT WAS UNKNOWN IF PROTAMINE WAS ADMINISTERED. POST PROCEDURE AFTER THE RECOVERY ROOM PROCEDURE, THE PATIENT CRASHED AND RESUSCITATION WAS ATTEMPTED FOR APPROXIMATELY ONE HOUR WITHOUT SUCCESS. CLINICAL EVALUATION, ULTRASOUND, AND ECHOCARDIOGRAM WAS PERFORMED. PERICARDIAL EFFUSION WAS OBSERVED BUT NO PERICARDIOCENTESIS WAS PERFORMED. THE PATIENT DIED. THE CAUSE OF DEATH WAS PERICARDIAL EFFUSION. THERE WAS NO ALLEGATION OF A MALFUNCTION AGAINST AN ABBOTT DEVICE. THE PHYSICIAN THOUGHT THE EFFUSION WAS RELATED TO THE AMULET BUT WAS UNCERTAIN. PATIENT WAS NOT TAKING AN ANTICOAGULANT OR ANTIPLATELET MEDICATION THE TIME WHEN THE PERICARDIAL EFFUSION WAS DEDICATED. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22875173&pc=NIQ | 2024168-2025-08897 | 2021-02-01 05:00 | Death | ABBOTT VASCULAR | 2025-08-22 04:00 | NIQ | XIENCE EVEROLIMUS ELUTING CORONARY STENT SYSTEM | Adverse Event Without Identified Device or Use Problem | Unspecified Heart Problem | P070015 |  | 1 | Event Description: IT WAS REPORTED IN AN ARTICLE THAT 6002 PATIENTS WERE ENROLLED AND RANDOMLY ASSIGNED FROM 72 CENTERS IN JAPAN. THE SAFETY ANALYSIS SET POPULATION CONSISTED OF 5996 PATIENTS AFTER EXCLUDING 5 PATIENTS WHO DID NOT RECEIVE PCI FOR THE ABSENCE OF SUITABLE CORONARY LESIONS, AND 1 PATIENT WHO HAD BEEN ENROLLED IN ANOTHER CLINICAL TRIAL. EXCLUDING 30 PATIENTS WHO WITHDREW CONSENT, 5966 PATIENTS WERE INCLUDED IN THE FULL ANALYSIS SET POPULATION: 2984 PATIENTS IN THE NO-ASPIRIN GROUP, AND 2982 PATIENTS IN THE DAPT GROUP. THE XIENCE STENT WAS IMPLANTED WITHOUT ISSUES; HOWEVER, THE FOLLOWING PATIENT OUTCOMES OCCURRED AFTER 1-MONTH RANDOMIZATION AND WERE RE-HOSPITALIZED: CARDIOVASCULAR AND NON-CARDIOVASCULAR DEATH. DETAILS ARE LISTED IN THE ARTICLE TITLED, "AN ASPIRIN-FREE VERSUS DUAL ANTIPLATELET STRATEGY FOR CORONARY STENTING: STOPDAPT-3 RANDOMIZED TRIAL". NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORDS FOR THIS PRODUCT COULD NOT BE REVIEWED AND A SIMILAR COMPLAINT QUERY COULD NOT BE PERFORMED BECAUSE THE PART NUMBER AND/OR LOT NUMBER WAS NOT REPORTED, AND THE DEVICE WAS NOT RETURNED. THE REPORTED PATIENT EFFECT OF DEATH IS LISTED IN THE XIENCE V EVEROLIMUS ELUTING CORONARY STENT SYSTEMS INSTRUCTIONS FOR USE AS A KNOWN PATIENT EFFECT(S) OF CORONARY STENTING PROCEDURES. A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECT(S), AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED; HOWEVER, THE SUBSEQUENT TREATMENT APPEARS TO BE RELATED TO THE OPERATIONAL CONTEXT OF THE PROCEDURE. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING; THEREFORE, NO PRODUCT-RELATED CORRECTIVE ACTION WILL BE IMPLEMENTED IN THIS CASE. LITERATURE ATTACHMENT: ARTICLE TITLE: ¿AN ASPIRIN-FREE VERSUS DUAL ANTIPLATELET STRATEGY FOR CORONARY STENTING: STOPDAPT-3 RANDOMIZED TRIAL¿ B2: DEATH DATE ESTIMATED B3: DATE OF EVENT ESTIMATED D4: THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. D6A: IMPLANT DATE ESTIMATED THE REPORTED PATIENT EFFECTS REFERENCED IN B5 IS FILED UNDER SEPARATE MEDWATCH REPORT NUMBER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22876966&pc=NKE | 2017865-2025-98328 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-08-22 04:00 | NKE | QUADRA ALLURE MP CRT-P | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P030035 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED DUE TO UNKNOWN CAUSES. THE PATIENT'S PACEMAKER AND LEFT VENTRICULAR (LV) LEAD WERE BOTH EXPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22877012&pc=NIK | 2017865-2025-98329 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-08-22 04:00 | NIK | QUARTET | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P030054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED DUE TO UNKNOWN CAUSES. THE PATIENT'S PACEMAKER AND LEFT VENTRICULAR (LV) LEAD WERE BOTH EXPLANTED. Manufacturer Narrative: THE REPORTED EVENT WAS PATIENT DEATH. AS RECEIVED, A PARTIAL LEAD WAS RETURNED IN ONE PIECE FOR ANALYSIS. FINAL ANALYSIS DID NOT FIND ANY ANOMALIES EXCEPT FOR PROCEDURAL DAMAGE. NO ANOMALIES WERE DETECTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22878244&pc=MNS | 2518422-2025-109323 | 2025-05-14 04:00 | Death | RESPIRONICS, INC. | 2025-08-22 04:00 | MNS | DREAMSTATION BIPAP AVAPS30 | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Unspecified Kidney or Urinary Problem | K102465 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THIS REPORT WAS RELATED TO A DREAMSTATION BIPAP AVAPS30 DEVICE. THE MANUFACTURER RECEIVED INFORMATION THE PATIENT EXPERIENCED "INFECTION, KIDNEY FAILURE AND DEATH". THE DEVICE WORKED BUT THE PATIENT COULDN'T TOLERATE IT ON THEIR FACE VERY LONG SINCE IT FELT LIKE IT WAS SMOTHERING THEM. THE PATIENT HAD A BAD INFECTION. THE PATIENT WENT TO THE HOSPITAL AND WAS GIVEN ANTIBIOTICS, BUT THEIR KIDNEYS SHUT DOWN. THE PATIENT THEN EXPERIENCED KIDNEY FAILURE AND PASSED AWAY. THE DEVICE WAS CLEANED EVERYDAY USING STERILE WATER. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22878376&pc=DSQ | 2916596-2025-05375 | 2025-08-03 04:00 | Death | THORATEC CORPORATION | 2025-08-22 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Sepsis; Drug Resistant Bacterial Infection | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO SEPSIS. THE SOURCE OF THE SEPSIS WAS A BACTERIAL INFECTION, STAPHYLOCOCCUS AUREUS. THE PUMP WAS WORKING AS INTENDED UNTIL THE THERAPY WAS STOPPED. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22879228&pc=NPT | 9617601-2025-01128 | 2025-08-14 04:00 | Death | MEDTRONIC MEXICO S. DE R.L. DE CV | 2025-08-22 04:00 | NPT | EVOLUT FX PLUS VALVE | Patient Device Interaction Problem | Arrhythmia; Ischemia; Pericardial Effusion; Heart Failure/Congestive Heart Failure | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT PRIOR TO THE IMPLANT OF A TRANSCATHETER VALVE VIA LEFT TRANSCAROTID APPROACH, A PRE-IMPLANT BALLOON AORTIC VALVULOPLASTY (BAV) WAS PERFORMED DUE TO CALCIFICATION. FOLLOWING THE IMPLANT OF THE VALVE THE PATIENT BECAME ISCHEMIC. SUBSEQUENTLY, CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED AND THE PATIENT WAS PLACED ON BYPASS. AN ANGIOGRAM WAS PERFORMED THAT SHOWED BLOOD FLOW INTO BOTH THE RIGHT AND LEFT CORONARY ARTERIES, AND A SMALL PERICARDIAL EFFUSION WAS OBSERVED. ADDITIONALLY, THE PATIENT EXPERIENCED VENTRICULAR DYSRHYTHMIA THAT REQUIRED MULTIPLE SHOCKS, AND LEFT VENTRICULAR FUNCTION BECAME TOO LOW. HOWEVER, THE PATIENT DIED DUE TO HEART FAILURE AND ISCHEMIA. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID: {D-EVOLUTFX-2329}; PRODUCT TYPE: {0195-HEARTVALVES}. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22879462&pc=PNJ | 2017865-2025-98431 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-08-22 04:00 | PNJ | AVEIR¿ LEADLESS PACEMAKER, RIGHT ATRIUM | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P150035 |  | 1 | Event Description: IT WAS REPORTED THROUGH A LITERATURE ARTICLE THAT AVEIR ATRIAL LEADLESS PACEMAKERS EXHIBITED THE FOLLOWING INTRA-OPERATIVE ISSUES; DOCKING ISSUES, DISLODGEMENT, STRETCHED HELIX, CAPTURING ISSUES AND SEPARATION FAILURE FOR WHICH NO INTERVENTION WAS REPORTED. DURING SOME INTRA-OPERATIVE EVENTS, THROMBUS WAS OBSERVED WITH NO REPORTED CONSEQUENCES. ATRIAL LEADLESS PACEMAKERS ALSO EXHIBITED THE FOLLOWING POST-OPERATIVE PROBLEMS; I2I COMMUNICATION ISSUES, BACKUP MODE, FAILURE TO INTERROGATE, FALSE ERI/RRT ALERTS, SENSING ISSUES, ELECTROMAGNETIC INTERFERENCE AND DISLODGEMENT WHICH SOME HAD BEEN ADDRESSED SURGICALLY, HOWEVER FURTHER DETAILS WERE NOT PROVIDED. THE RESEARCH ARTICLE STATES THAT PATIENT WITH IMPLANTED WITH A ATRIAL DEVICE PASSED AWAY FOR UNKNOWN REASONS. LASTLY, THERE WAS ALSO MENTION OF VENTRICULAR LEADLESS PACEMAKERS WHICH EXHIBITED STRETCHED HELIX AND I2I COMMUNICATION ISSUES, AS WELL AS DISLODGEMENTS POST-IMPLANT PROCEDURE. RESOLUTIONS TO THESE EVENTS WERE UNKNOWN. THE PATIENT'S CONDITIONS WERE UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22879463&pc=DYB | 2017233-2025-06526 | 2025-07-29 04:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-08-22 04:00 | DYB | GORE® DRYSEAL FLEX INTRODUCER SHEATH | Adverse Event Without Identified Device or Use Problem | Myocardial Infarction | K160254 |  | 1 | Event Description: THE FSA REPORTED THE FOLLOWING: ON (B)(6) 2025, THE PATIENT UNDERWENT AN ENDOVASCULAR PROCEDURE FOR A THORACOABDOMINAL AORTIC ANEURYSM INVOLVING THE VISCERAL VESSELS. A GORE® EXCLUDER® THORACOABDOMINAL BRANCH ENDOPROSTHESIS (TAMBE) DEVICE WAS PLANNED FOR THIS PATIENT. THE PATIENT WAS TREATED WITH A GORE® EXCLUDER® THORACOABDOMINAL BRANCH ENDOPROSTHESIS (TAMBE) AORTIC COMPONENT. THE PATIENT WAS ALSO TREATED WITH MULTIPLE DEVICES INCLUDING THE USE OF A GORE® DRYSEAL FLEX INTRODUCER SHEATH; ALL TAMBE SYSTEM COMPONENTS WERE SUCCESSFULLY TRACKED FROM THE ACCESS SITE TO THE INTENDED IMPLANTATION SITE AND RELEASED FROM DELIVERY CATHETERS SUCCESSFULLY. THERE WAS TECHNICAL SUCCESS WITH THE TAMBE PROCEDURE. REPORTEDLY THE PATIENT HAD TORTUOUS ILIAC ARTERY ANATOMY. A PRE-PROCEDURE DIAGNOSTIC ANGIO REVEALED MID SUPERIOR MESENTERIC ARTERY (SMA) OCCLUSION. THE SMA MEASURED 5-6MM. A SHOCKWAVE CATHETER ABLATION SYSTEM WAS USED IN THE LEFT EXTERNAL ILIAC ARTERY TWICE, BEFORE DELIVERING THE GORE® DRYSEAL FLEX INTRODUCER SHEATH (22FR, UNK LOT/SERIAL NUMBER, HOSPITAL INVENTORY). FINAL ANGIO REVEALED THE SMA ARTERY HAD A FLOW LIMITING DISSECTION, DISTAL TO THE IMPLANTED GORE® VIABAHN® ENDOPROSTHESIS. THE PHYSICIAN HAD PLANNED TO DO A SMA BYPASS SURGERY. IT WAS REPORTED THAT WHEN REMOVING THE GORE® DRYSEAL FLEX INTRODUCER SHEATH (DSF/22FR) BEING USED ON THE LEFT SIDE, THE EXTERNAL ILIAC ARTERY RUPTURED. WHILE IMPLANTING AN ADDITIONAL STENT, THE PATIENTS BLOOD PRESSURE STARTED TO DROP (80/40). TWO UNITS OF BLOOD WERE GIVEN; MULTIPLE VASOPRESSORS WERE BEING GIVEN WITH FLUIDS. UPON SUCCESSFUL COMPLETION OF STENTING THE LEFT EXTERNAL ILIAC ARTERY, THE PATIENTS BLOOD PRESSURE ELEVATED (220/110). THE PATIENT WENT INTO CARDIOGENIC SHOCK, HAD A MYOCARDIAL INFARCTION, AND CODED. THE PATIENT EXPIRED ON THE OPERATING TABLE. Manufacturer Narrative: W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22879516&pc=QZI | 2124215-2025-59102 | 2025-07-31 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-08-22 04:00 | QZI | FARAWAVE | Off-Label Use | Arrhythmia; Cardiac Arrest |  |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT DEATH OCCURRED. DURING A PULSED FIELD ABLATION (PFA) PROCEDURE TO TREAT AN ATRIAL FIBRILLATION AND ATRIAL FLUTTER A FARAWAVE WAS SELECTED FOR USE. AFTER PERFORMING THE PULMONARY VEIN AND POSTERIOR WALL ISOLATION, THE PHYSICIAN WAS CHECKING THE POSTERIOR WALL ELECTRICAL ACTIVITY AND THE PATIENT SUDDENLY WENT INTO VENTRICULAR TACHYCARDIA. NUMEROUS CARDIOVERSIONS AND DEFIBRILLATIONS WERE ATTEMPTED WITHOUT SUCCESS. ADVANCED CARDIAC LIFE SUPPORT (ACLS) PROTOCOL WAS PERFORMED, INCLUDING CARDIOPULMONARY RESUSCITATION AND MEDICATION. AFTER APPROXIMATELY 35 MINUTES OF ACLS WITH LACK OF RESUSCITATION, THE PHYSICIAN CALLED TIME OF DEATH. PATIENT WAS NOT ON PROTAMINE. THE PATIENT HAD HISTORY OF LOW EF (20%), OPEN HEART SURGERY FOR CARDIAC BYPASS AND TRANSAORTIC VALVE REPLACEMENT. LEFT ATRIAL PRESSURE WAS HIGH BEFORE ABLATION. 74 APPLICATIONS WERE GIVEN, NO WALL MOTION ABNORMALITIES OBSERVED. Manufacturer Narrative: THE DEVICE HAS NOT BEEN RECEIVED FOR ANALYSIS. UPON RECEIPT AND COMPLETION OF THE FAILURE ANALYSIS OF THE COMPLAINT DEVICE, IF THERE IS ANY FURTHER RELEVANT INFORMATION FROM THAT REVIEW, A SUPPLEMENTAL MEDWATCH WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22879783&pc=LWP | 2017865-2025-98437 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-08-22 04:00 | LWP | ASSURITY MRI | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P140033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD DECEASED DUE TO AN UNKNOWN CAUSE. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE RELATED. NO ADDITIONAL INFORMATION WAS REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22879785&pc=NVN | 2017865-2025-98439 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-08-22 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD DECEASED DUE TO AN UNKNOWN CAUSE. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE RELATED. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: THE LEAD WAS RETURNED DUE TO PATIENT DEATH. AS RECEIVED, A PARTIAL LEAD (ONLY CONNECTOR PORTION) WAS RETURNED IN ONE PIECE. ELECTRICAL, MECHANICAL, AND VISUAL ANALYSIS WAS NORMAL WITH NO ANOMALIES FOUND.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22879788&pc=NVY | 2017865-2025-98440 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-08-22 04:00 | NVY | ISOFLEX OPTIM LEAD | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960030 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD DECEASED DUE TO AN UNKNOWN CAUSE. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE RELATED. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: THE LEAD WAS RETURNED DUE TO PATIENT DEATH. AS RECEIVED, A PARTIAL LEAD (ONLY THE CONNECTOR) WAS RETURNED IN ONE PIECE FOR ANALYSIS. ELECTRICAL, MECHANICAL, AND VISUAL ANALYSIS WAS NORMAL WITH NO ANOMALIES FOUND.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22879906&pc=NVN | 2017865-2025-98445 | 2025-08-02 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-08-22 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD DECEASED DUE TO AN UNKNOWN CAUSE. NO ADDITIONAL INFORMATION WAS REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22879909&pc=LWP | 2017865-2025-98442 | 2025-08-02 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-08-22 04:00 | LWP | ENDURITY MRI | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P140033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD DECEASED DUE TO AN UNKNOWN CAUSE. NO ADDITIONAL INFORMATION WAS REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22880114&pc=DSQ | 2916596-2025-05344 | 2025-07-30 04:00 | Death | THORATEC CORPORATION | 2025-08-22 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Renal Failure; Cardiogenic Shock; Multiple Organ Failure; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS CARDIOGENIC SHOCK. IT WAS NOTED THIS WAS NOT DEVICE OR DEVICE THERAPY RELATED. DEVICE PARAMETERS (FLOW, SPEED, POWER) WERE WITHIN NORMAL LIMITS. THE DEVICE WOULD NOT BE EXPLANTED AND WOULD NOT BE RETURNED FOR EVALUATION. THE DEVICE OPERATED AS EXPECTED. THE SYMPTOMS OF CARDIOGENIC SHOCK WERE CONSISTENTLY LOW MEAN ARTERIAL PRESSURES (MAPS), END ORGAN DYSFUNCTION, ESCALATING DEVICE AND MEDICATION SUPPORT AND INABILITY TO BE EXTUBATED. THE SHOCK WAS TREATED WITH THE LEFT VENTRICULAR ASSIST DEVICE (LVAD), RIGHT VENTRICULAR ASSIST DEVICE (RVAD), CONTINUOUS RENAL REPLACEMENT THERAPY (CRRT), MECHANICAL VENTILATION, SEVERAL PRESSORS AND INOTROPES. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22858236&pc=PNJ | 9612164-2025-04115 | 2025-07-31 04:00 | Death | MEDTRONIC IRELAND | 2025-08-21 04:00 | PNJ | MICRA¿ AV2 | Device Dislodged or Dislocated; Adverse Event Without Identified Device or Use Problem | Pericardial Effusion | P150033 |  | 1 | Event Description: IT WAS REPORTED INTRAOPERATIVELY THAT THE PATIENT EXPERIENCED PERICARDIAL EFFUSION. IT WAS REPORTED THAT THE LEADLESS IMPLANTABLE PU LSE GENERATOR (IPG) DISLODGED AND THEN WAS RECAPTURED WITH THE DELIVERY SYSTEM AND IT WAS NOTED THAT THE HEART BEGAN TO ENLARGE ON FLUOROSCOPY. A PERICARDIAL SYNTHESIS KIT WAS USED AND CHEST COMPRESSIONS WERE PERFORMED. THE PATIENT WAS TRANSFERRED TO THE INTENSIVE CARE UNIT (ICU) WHERE THEY PASSED AWAY SHORTLY AFTER. Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [MICRA] PRODUCT ID [MC2AVR1-DELSYS] (SERIAL: [(B)(6)]). D9: NO. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22858270&pc=QSK | 9612515-2025-00069 | 2025-06-11 04:00 | Death | VASCUTEK LTD. | 2025-08-21 04:00 | QSK | THORAFLE HYBRID | Insufficient Information | Sepsis | P210006 |  | 1 | Event Description: INFECTION EVENT REPORTED FORM THE EXTEND STUDY (NCT05639400). SUBJECT CAUCASIAN MALE( WHO WAS 79 AT THE TIME OF CONSENT TAKEN ON (B)(6) 2025) UNDERWENT AN ELECTIVE OPEN SURGICAL REPAIR OF THE AORTA WITH A THORAFLEX HYBRID PLEXUS DEVICE ON (B)(6) 2025, FOR FUSIFORM AORTIC ANEURYSM. ON POST-OPERATIVE DAY 8 (B)(6) 2025), SUBJECT (B)(6) EXPERIENCED AN SAE( SERIOUS ADVERSE EVENT) OF INFECTION, FURTHER DETAILS STATING "BLOOD STREAM INFECTION WITH VANCOMYCIN RESISTANT ENTEROCOCCI (FAECUM), VRE (VANCOMYCIN-RESISTANT ENTEROCOCCUS).INCLUDING INFECTION OF THE AORTIC PROSTHESIS. POSITIVE BLOOD CULTURE FOUND IN INTRA-SURGICAL SAMPLE TAKEN ON (B)(6) 2025. EVENT IS ANTICIPATED AND SEVERE IN NATURE. EVENT WAS CONSIDERED POSSIBLY RELATED TO DEVICE, NOT RELATED TO DEVICE FAILURE/DEFICIENCY AND NOT RELATED TO PRE-EXISTING CONDITIONS. PATIENT PASSED AWAY ON (B)(6) 2025. SITE DETAILED REASON FOR DEATH AS "PALLIATIVE TREATMENT IN TERMS OF BEST SUPPORTIVE CARE AND ACCORDING TO PATIENT'S WILL". Manufacturer Narrative: MANUFACTURER NARRATIVE: CLINICAL CODE: 2067 -SEPSIS- BLOOD STREAM INFECTION WITH VANCOMYCIN RESISTANT ENTEROCOCCI (FAECUM), IMPACT CODE: 1802- DEATH - PATIENT PASSED AWAY ON (B)(6) 2025. SITE DETAILED REASON FOR DEATH AS "PALLIATIVE TREATMENT IN TERMS OF BEST SUPPORTIVE CARE AND ACCORDING TO PATIENT'S WILL". DEVICE PROBLEM CODE : 3190 - INSUFFICIENT INFORMATION - NO DEVICE FAILURE /DEFICIENCY HAS BEEN REPORTED. COMPONENT CODE: 4755 - PART/COMPONENT/SUB-ASSEMBLY TERM NOT APPLICABLE. TYPE OF INVESTIGATION: 4110 - TREND ANALYSIS: - THORAFLEX HYBRID SALES (B)(6) VS THORAFLEX HYBRID INFECTION EVENTS (B)(6) HAD AN OCCURRENCE RATE OF 0.044% NO TREND REQUIRING ACTION HAS BEEN IDENTIFIED. 4111 - COMMUNICATION INTERVIEW: ADDITIONAL INFORMATION HAS BEEN REQUESTED TO AID THIS INVESTIGATION. 3331 - ANALYSIS OF PRODUCTION RECORDS: FULL BATCH REVIEW PERFORMED FROM BASE FABRIC TO FINISHED PRODUCT FOR DEVICE ID- WHICH CONFIRMED THE BATCH WAS MANUFACTURED TO SPECIFICATION. 4117 - DEVICE NOT RETURNED: DEVICE REMAINS IMPLANTED INVESTIGATION FINDINGS: 3233 - RESULTS PENDING COMPLETION OF INVESTIGATION. INVESTIGATION CONCLUSION: 11 - CONCLUSION NOT YET AVAILABLE AS INVESTIGATION IS ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22858537&pc=PNJ | 9612164-2025-04117 | 2025-08-14 04:00 | Death | MEDTRONIC IRELAND | 2025-08-21 04:00 | PNJ | MICRA | Failure to Capture; High impedance; Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Hypovolemic Shock; Cardiac Tamponade; Cardiac Perforation | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT DURING IMPLANT THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) EXHIBITED HIGH AND UNDEFINED IMPEDANCE DURING THE FIRST PLACEMENT ATTEMPT AND NO CAPTURE DURING ANY PLACEMENT ATTEMPT. DURING THE IMPLANT PROCEDURE THE PATIENT EXPERIENCED TAMPONADE WITH CARDIAC PERFORATION. PERICARDIOCENTESIS WAS PERFORMED BUT THERE WAS NO INITIAL CONTROL OF THE HEMORRHAGE WITH OVER TWO LITRES OF BLOOD DRAINED. PATIENT DIED DUE TO THE CONSEQUENCES OF HEMORRHAGIC SHOCK.Event Description: IT WAS FURTHER REPORTED THAT IT WAS THE PHYSICIAN'S OPINION THAT THE PATIENTS SYMPTOMS OF DEATH, CARDIAC PERFORATION, HEMORRHAGE AND TAMPONADE WERE LINKED TO THE LEADLESS IPG AND DELIVERY SYSTEM. THERE WERE MULTIPLE PLACEMENT ATTEMPTED AND THE LAST ONE WAS "MORE APICAL". THERE IS A DIRECT CORRELATION BETWEEN THE LEADLESS IPG IMPLANT STEPS AND PERICARDIAL EFFUSION. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [MICRA VR] PRODUCT ID [MC1VR01-DELSYS] (SERIAL: (B)(6). D9: NO MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22859782&pc=MIH | 9612164-2025-04119 | 2025-07-30 04:00 | Death | MEDTRONIC IRELAND | 2025-08-21 04:00 | MIH | ENDURANT II BIFURCATED STENT GRAFT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P100021 |  | 1 | Event Description: AN ENDURANT STENT GRAFT SYSTEM WAS IMPLANTED DURING THE ENDOVASCULAR TREATMENT OF A RUPTURED ABDOMINAL AORTIC ANEURYSM. THE REPORTED DEVICE WAS USED DURING THE INDEX PROCEDURE, WHICH WAS COMPLETED SUCCESSFULLY WITH NO ENDOLEAK OBSERVED. IT WAS REPORTED DURING THE INDEX PROCEDURE, FOLLOWING WOUND CLOSURE, THE PATIENT WAS TRANSFERRED TO THE INTENSIVE CARE UNIT (ICU). HOWEVER, THE PATIENT'S VITAL SIGNS DID NOT STABILIZE, AND THE PATIENT EXPIRED. IN ACCORDANCE WITH THE WISHES OF THE PATIENT¿S FAMILY, NO ADDITIONAL TREATMENT WAS PURSUED FOLLOWING EVAR. PER THE PHYSICIAN THE CAUSE DEATH WAS NOT SPECIFIED. Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME ENDURANT II ILIAC STENT GRAFT; PRODUCT ID ETLW1613C93EJ (SERIAL: (B)(6)); PRODUCT TYPE: 0180-AORTIC STENT GRAFT; IMPLANT DATE (B)(6) 2025 BRAND NAME ENDURANT II ILIAC STENT GRAFT; PRODUCT ID ETLW1616C124EJ (SERIAL: (B)(6)); PRODUCT TYPE: 0180-AORTIC STENT GRAFT; IMPLANT DATE (B)(6) 2025. BRAND NAME ENDURANT II ILIAC STENT GRAFT; PRODUCT ID ETLW1624C93EJ (SERIAL: (B)(6)); PRODUCT TYPE: 0180-AORTIC STENT GRAFT; IMPLANT DATE (B)(6) 2025. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22859888&pc=DSQ | 2916596-2025-05507 | 2025-08-06 04:00 | Death | THORATEC CORPORATION | 2025-08-21 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. IT WAS REPORTED THAT THE PATIENT'S OUTCOME WAS NOT DEVICE OR THERAPY RELATED. THE DEVICE WILL NOT BE EXPLANTED AND WILL NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. B, IS CURRENTLY AVAILABLE. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22859957&pc=GEI | 1717344-2025-01076 | 2025-07-16 04:00 | Death | COVIDIEN | 2025-08-21 04:00 | GEI | VALLEYLAB | Device Displays Incorrect Message; Connection Problem; Adverse Event Without Identified Device or Use Problem | Unspecified Tissue Injury; Insufficient Information | K151649 |  | 1 | Event Description: IT WAS REPORTED THAT DURING PROCEDURE, WHEN THE LEAD AND THE MONOPOLAR LEAD WERE CONNECTED TO THE DIATHERMY MACHINE, AN ORANGE/YELLOW ERROR MESSAGE STATING "INTERNAL ERROR, PLEASE DISMISS" WAS DISPLAYED. INFORMATION RECEIVED INDICATES THAT THE PATIENT IS DECEASED, NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION OR RELATION TO THE DEVICE.Event Description: IT WAS REPORTED THAT DURING PROCEDURE, WHEN THE LEAD AND THE MONOPOLAR LEAD WERE CONNECTED TO THE DIATHERMY MACHINE, AN ORANGE/YELLOW ERROR MESSAGE STATING "INTERNAL ERROR, PLEASE DISMISS" WAS DISPLAYED. THE PATIENT HAD INJURY.Event Description: IT WAS REPORTED THAT DURING PROCEDURE, WHEN THE LEAD AND THE MONOPOLAR LEAD WERE CONNECTED TO THE DIATHERMY MACHINE, AN ORANGE/YELLOW ERROR MESSAGE STATING "INTERNAL ERROR, PLEASE DISMISS" WAS DISPLAYED. THE LEADS WERE DISCONNECTED AND WAS REATTACHED AND IT FLASHED UP AGAIN SO EVERYTHING WAS DISCONNECTED AND REPLACED WITH ANOTHER DIATHERMY MACHINE WHICH WORKED. THE ORIGINAL UNIT WITH THE ISSUE WAS LEFT FOR 22 HOURS AND WHEN IT WAS TESTED AGAIN THE MESSAGE DID NOT APPEAR. THE STATUS OF THE PATIENT WAS UNKNOWN AT THE TIME WHEN THE ISSUE OCCURRED WITH THE DEVICE BUT THE PATIENT EVENTUALLY DIED BUT NOT RELATED TO THE DEVICE ISSUE. Manufacturer Narrative: CORRECTION: B5, H6 NEW RFR CODES: PORT ISSUE, ERROR MESSAGE - SERVICE REQUIRED, MONOPOLAR PORT ISSUE ADDITIONAL INFORMATION: G3 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5, G3, H6 ADD FDP/IME/IMF: C172008 AND C98904 REMOVE FDP: C28554 THIS EVENT HAS BEEN REASSESSED AND THE REPORTABILITY HAS BEEN DETERMINED TO BE A REPORTABLE MALFUNCTION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22860165&pc=NGV | 2124215-2025-57877 | 2025-06-11 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-08-21 04:00 | NGV | WATCHMAN FLX? PRO | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P130013 |  | 1 | Event Description: IT WAS REPORTED THAT DEVICE EMBOLIZATION AND PATIENT DEATH OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS PERFORMED FOR CLOSURE OF AN LAA OF WINDSOCK MORPHOLOGY WITH SUCCESSFUL IMPLANT OF A 24MM WATCHMAN FLX PRO CLOSURE DEVICE. APPROXIMATELY ONE DAY POST-IMPLANT, THE 24MM WATCHMAN FLX PRO CLOSURE DEVICE EMBOLIZED TO THE AORTIC VALVE. THE PATIENT UNDERWENT EMERGENT SURGERY TO EXPLANT THE 24MM WATCHMAN FLX PRO CLOSURE DEVICE. POST-SURGERY THE PATIENT WAS SENT TO HOSPICE CARE. THE PATIENT DIED ONE MONTH LATER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22860350&pc=NIQ | 2024168-2025-08803 | 2025-08-01 04:00 | Death | ABBOTT VASCULAR INC. | 2025-08-21 04:00 | NIQ | XIENCE ALPINE¿ | Failure to Advance; Difficult to Advance; Device Dislodged or Dislocated | Bradycardia | P110019 |  | 1 | Event Description: IT WAS REPORTED THE PROCEDURE WAS TO TREAT A HEAVILY CALCIFIED AND MILDLY TORTUOUS LESION IN THE LEFT MAIN ARTERY. THE 2.75X23MM RX XIENCE ALPINE STENT DELIVERY SYSTEM (SDS) WAS ADVANCED HOWEVER RESISTANCE WAS MET. PER THE PHYSICIAN, IT COULD NOT BE DETERMINED IF THE RESISTANCE WAS WITH THE ANATOMY OR THE GUIDING CATHETER. DURING ADVANCEMENT THE STENT DISLODGED FROM THE BALLOON. AN ATTEMPT WAS MADE TO RETRIEVE THE DISLODGED STENT USING A BALLOON CATHETER; HOWEVER, IT WAS UNSUCCESSFUL. DURING THE RETRIEVAL ATTEMPTS THE PATIENT DEVELOPED BRADYCARDIA AND EXPIRED. PER THE PHYSICIAN THE PATIENT DEATH WAS RELATED TO THE DISLODGED STENT. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22860351&pc=NPT | 9617601-2025-01093 | 2025-08-14 04:00 | Death | MEDTRONIC MEXICO S. DE R.L. DE CV | 2025-08-21 04:00 | NPT | EVOLUT FX PLUS VALVE | Device Dislodged or Dislocated; Activation Failure; Patient Device Interaction Problem | Cardiac Arrest; Low Blood Pressure/ Hypotension; Foreign Body In Patient | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT PRIOR TO THE IMPLANT OF A TRANSCATHETER AORTIC VALVE, INSIDE A PATIENT WHO PRESENTED WITH AORTIC VALVE STENOSIS AND A BICUSPID AORTIC VALVE MORPHOLOGY, THE NATIVE AORTIC VALVE WAS PRE-DILATED WITH A 25-MILLIMETER (MM) NON-MEDTRONIC BALLOON. DURING INTRODUCTION OF THE DELIVERY CATHETER SYSTEM (DCS) AND DURING VALVE POSITIONING, THE PATIENT EXHIBITED COMPROMISED ARTERIAL BLOOD PRESSURE. THE PATIENT'S SYSTOLIC BLOOD PRESSURE DID NOT EXCEED 70 MILLIMETERS OF MERCURY (MM HG). THE VALVE WAS IMPLANTED AND CONSIDERED WELL POSITIONED. SUBSEQUENTLY, UNDER-EXPANSION WAS OBSERVED, PROMPTING POST-DILATION WITH A 23MM NON-MEDTRONIC BALLOON. DURING BALLOON INFLATION AND RAPID PACING, LOSS OF CAPTURE OCCURRED, RESULTING IN DISLODGEMENT OF THE TRANSCATHETER AORTIC VALVE. THE PATIENT BECAME HEMODYNAMICALLY UNSTABLE, REQUIRING CARDIOPULMONARY RESUSCITATION. DUE TO URGENCY AND ANATOMICAL CHALLENGES, THE IMPLANTING TEAM ATTEMPTED TO IMPLANT A NON-MEDTRONIC VALVE; CROSSING THE DISLODGED VALVE WITH THE NON-MEDTRONIC VALVE WAS EXTREMELY DIFFICULT AND REQUIRED MULTIPLE ATTEMPTS UNDER ONGOING RESUSCITATION, WITH A SNARE USED TO HOLD THE MEDTRONIC VALVE. THE NON-MEDTRONIC VALVE WAS EVENTUALLY IMPLANTED IN THE AORTIC POSITION, BUT THE PATIENT SHOWED NO SIGNS OF EFFECTIVE CIRCULATION AFTER IMPLANTATION AND SUBSEQUENTLY DIED. PER THE PHYSICIAN, THE CAUSE OF DEATH WAS DUE TO CIRCULATORY ARREST COMPLICATED BY THE TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) PROCEDURE. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID L-EVOLUTFX-2329 (LOT: 0012790375); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE N/A; EXPLANT DATE N/A SECTION D REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: BRAND NAME EVOLUT FX DCS; PRODUCT ID D-EVOLUTFX-2329 (LOT: 0012864341); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE N/A; EXPLANT DATE N/A SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. THE CODES PRESENT IN SECTION H6 CORRESPOND TO MULTIPLE COMPONENTS/PRODUCTS THAT COMPRISE THIS REPORTED EVENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22860430&pc=NPS | 2135147-2025-04668 | 2019-01-01 05:00 | Death | ABBOTT MEDICAL | 2025-08-21 04:00 | NPS | TRICLIP DELIVERY SYSTEM (TCDS) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P230007 |  | 1 | Event Description: THE ARTICLE "C-REACTIVE PROTEIN-TO-ALBUMIN RATIO IS ASSOCIATED WITH MORTALITY AFTER TRANSCATHETER TRICUSPID VALVE REPAIR" WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE MULTI CENTER STUDY, TO EVALUATE MARKERS TO ASSESS MORTALITY RISK IN PATIENTS UNDERGOING TRANSCATHETER TRICUSPID VALVE REPAIR (TTVR). THE C-REACTIVE PROTEIN (CRP)/ALBUMIN RATIO (CAR) IS A MARKER OF SYSTEMIC INFLAMMATION AND REDUCED NUTRITIONAL STATUS, WHICH CAN BOTH OCCUR IN TR. DEVICES MENTIONED INCLUDE TRICLIP AND PASCAL. THE ARTICLE CONCLUDED THAT HIGHER CAR REFLECTS PATIENTS WITH ADVANCED RIGHT-HEART FAILURE AND EXTRACARDIAC ORGAN DAMAGE AND IS ASSOCIATED WITH MORTALITY AFTER TTVR.. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS KARL FINKE AT UNIVERSITY OF COLOGNE, UNIVERSITY HOSPITAL COLOGNE, COLOGNE, GERMANY WITH CORRESPONDING EMAIL KARL.FNKE@UK-KOELN.DE. THE TIME FRAME OF THE STUDY WAS 2019 TO 2022. A TOTAL OF 215 PATIENTS WERE INCLUDED IN THIS STUDY, WITH A MEDIAN AGE OF 80 (IQR 7) YEARS, 69% OF WHICH WERE FEMALE. AN UNKNOWN AMOUNT RECEIVED AN ABBOTT DEVICE. COMORBIDITIES INCLUDED PERIPHERAL EDEMA, ARTERIAL HYPERTENSION, ATRIAL FIBRILLATION, CHRONIC KIDNEY DISEASE, CORONARY ARTERY DISEASE, DIALYSIS, DIABETES, TRICUSPID REGURGITATION. PERI AND POST-PROCEDURAL COMPLICATIONS INCLUDED HEART FAILURE HOSPITALIZATION, DEATH, WORSENING TRICUSPID REGURGITATION, UNCHANGED TRICUSPID REGURGITATION, UNSPECIFIED INFECTION. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF TRICLIP WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING PERIPHERAL EDEMA, ARTERIAL HYPERTENSION, ATRIAL FIBRILLATION, CHRONIC KIDNEY DISEASE, CORONARY ARTERY DISEASE, DIALYSIS, DIABETES, TRICUSPID REGURGITATION. COMPLICATIONS REPORTED INCLUDED HEART FAILURE, HOSPITALIZATION, DEATH, WORSENING TRICUSPID REGURGITATION, UNCHANGED TRICUSPID REGURGITATION, UNSPECIFIED INFECTION; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. B2: DEATH DATE IS ESTIMATED. B3: EVENT DATE IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: C-REACTIVE PROTEIN-TO-ALBUMIN RATIO IS ASSOCIATED WITH MORTALITY AFTER TRANSCATHETER TRICUSPID VALVE REPAIR.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22860661&pc=NGV | 2124215-2025-58366 | 2025-05-09 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-08-21 04:00 | NGV | SYSTEM, APPENDAGE CLOSURE, LEFT ATRIAL | Adverse Event Without Identified Device or Use Problem | Stroke/CVA |  |  | 1 | Event Description: THE FOLLOWING EVENT ORIGINATED FROM A PATIENT'S WIFE SENDING A LETTER THROUGH THE BOSTON SCIENTIFIC WEBSITE, AND THEREFORE DESPITE GOOD FAITH EFFORTS THE INFORMATION IS OFTEN INCOMPLETE. ADDITIONALLY, DUE TO THE LIMITED INFORMATION RECEIVED THROUGH GOOD FAITH EFFORTS, THIS EVENT MAY REPRESENT A DUPLICATE OF AN EVENT WHICH WAS ALREADY KNOWN TO BOSTON SCIENTIFIC. IT WAS REPORTED A CEREBRAL VASCULAR ACCIDENT (CVA) AND DEATH OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS PERFORMED ON (B)(6) 2022, AND A WATCHMAN CLOSURE DEVICE WAS IMPLANTED. FOLLOW UPS POST INDEX PROCEDURE WERE ATTENDED AND CARE CONTINUED AS SUGGESTED. ON (B)(6) 2023, THE PATIENT EXPERIENCED A CVA AND WAS HOSPITALIZED. A TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) WAS PERFORMED AND THE PHYSICIAN STATED, "EVERYTHING WAS PERFECT". THE PATIENT DIED ON (B)(6) 2025. THERE WERE NO FURTHER DETAILS REPORTED AT THE TIME OF THIS REPORT. Manufacturer Narrative: D1: BRAND NAME IS BLANK BECAUSE THE DEVICE IS KNOWN TO BE A WATCHMAN LAA CLOSURE DEVICE, BUT THE EXACT BRAND NAME IS UNKNOWN. D4: DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UDI AND OTHER PRODUCT SPECIFIC INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22860984&pc=CBK | 8010042-2025-0001291 | 2025-05-09 04:00 | Death | MAQUET CRITICAL CARE AB | 2025-08-21 04:00 | CBK | BASE UNIT SERVO-U | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K201874 |  | 1 | Event Description: IT WAS REPORTED THAT THE VENTILATOR WAS LEFT IN STANDBY MODE FOR APPROXIMATELY 30 MINUTES, WHICH RESULTED IN SEVERE HARM TO THE PATIENT AND ULTIMATELY LED TO DEATH. MANUFACTURER¿S REF #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22861068&pc=MFE | 2029214-2025-01866 | 2025-06-12 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-08-21 04:00 | MFE | ONYX | Excess Flow or Over-Infusion; Migration or Expulsion of Device; Adverse Event Without Identified Device or Use Problem | Ischemia; Paralysis; Paresis; Respiratory Insufficiency |  |  | 1 | Event Description: YU, J. (2025). EMBOLIZATION OF A HIGH CERVICAL PIAL ARTERIOVENOUS FISTULA RESULTING IN RESPIRATORY INSUFFICIENCY DUE TO DIAPHRAGMATIC PARALYSIS: A CASE REPORT. RADIOLOGY CASE REPORTS, 20(9), 4293¿4298. HTTPS://DOI.ORG/10.1016/J.RADCR.2025.05.060 MEDTRONIC REVIEW OF THE LITERATURE ARTICLE FOUND A CASE REPORT OF A 46-YEAR-OLD MALE PATIENT WHO UNDERWENT AN ENDOVASCULAR PROCEDURE OF ONYX EMBOLIZATION TO TREAT A HIGH CERVICAL PIAL ARTERIOVENOUS FISTULA (PAVF) WITH ASSOCIATED SUBARACHNOID HEMORRHAGE (SAH). A MARATHON CATHETER WAS USED TO DELIVER ONYX FOR TREATMENT. IT WAS NOTED THAT A SECOND MARATHON CATHETER WAS TRIED FROM A DIFFERENT APPROACH BUT WAS UNABLE TO REACH THE INTENDED LOCATION. DURING THE ONYX EMBOLIZATION, THERE WAS EXCESSIVE REFLUX OF THE ONYX INTO THE DISTAL ANTERIOR SPINAL ARTERY (ASA). THE PAVF WAS SUCCESSFULLY OBLITERATED IN THE EMBOLIZATION PROCEDURE. HOWEVER, POST-OPERATIVELY THE PATIENT HAD HEMIPARESIS AND RESPIRATORY INSUFFICIENCY INDEPENDENTLY; MECHANICAL VENTILATION WAS PROVIDED. FIFTEEN DAYS AFTER EVT, THE PATIENT¿S HEMIPARESIS COMPLETELY RESOLVED. HOWEVER, ULTRASOUND SHOWED CONTINUED DIAPHRAGMATIC PARALYSIS. MAGNETIC RESONANCE IMAGING (MRI) REVEALED CERVICAL CORD INFARCTION AT THE C2-C3 VERTEBRAL LEVEL, WHICH SUPPORTED THE DIAGNOSIS OF DIAPHRAGMATIC PARALYSIS DUE TO PHRENIC NERVE INJURY. THE PATIENT RECEIVED DAILY RESPIRATORY REHABILITATION THERAPY AND RESPIRATORY FUNCTION WE IMPROVING. HOWEVER, 93 DAYS POST-PROCEDURE, THE PORTABLE VENTILATOR WAS NOT USED AND THE PATIENT DIED WHILE SLEEPING. Manufacturer Narrative: B3. REPORTED EVENT DATE IS THE DATE THE ARTICLE WAS PUBLISHED ONLINE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22861084&pc=KRA | 2029214-2025-01867 | 2025-06-12 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-08-21 04:00 | KRA | MARATHON | Adverse Event Without Identified Device or Use Problem | Ischemia; Paralysis; Paresis; Respiratory Insufficiency | K093750 |  | 1 | Event Description: YU, J. (2025). EMBOLIZATION OF A HIGH CERVICAL PIAL ARTERIOVENOUS FISTULA RESULTING IN RESPIRATORY INSUFFICIENCY DUE TO DIAPHRAGMATIC PARALYSIS: A CASE REPORT. RADIOLOGY CASE REPORTS, 20(9), 4293¿4298. HTTPS://DOI.ORG/10.1016/J.RADCR.2025.05.060 MEDTRONIC REVIEW OF THE LITERATURE ARTICLE FOUND A CASE REPORT OF A 46-YEAR-OLD MALE PATIENT WHO UNDERWENT AN ENDOVASCULAR PROCEDURE OF ONYX EMBOLIZATION TO TREAT A HIGH CERVICAL PIAL ARTERIOVENOUS FISTULA (PAVF) WITH ASSOCIATED SUBARACHNOID HEMORRHAGE (SAH). A MARATHON CATHETER WAS USED TO DELIVER ONYX FOR TREATMENT. IT WAS NOTED THAT A SECOND MARATHON CATHETER WAS TRIED FROM A DIFFERENT APPROACH BUT WAS UNABLE TO REACH THE INTENDED LOCATION. DURING THE ONYX EMBOLIZATION, THERE WAS EXCESSIVE REFLUX OF THE ONYX INTO THE DISTAL ANTERIOR SPINAL ARTERY (ASA). THE PAVF WAS SUCCESSFULLY OBLITERATED IN THE EMBOLIZATION PROCEDURE. HOWEVER, POST-OPERATIVELY THE PATIENT HAD HEMIPARESIS AND RESPIRATORY INSUFFICIENCY INDEPENDENTLY; MECHANICAL VENTILATION WAS PROVIDED. FIFTEEN DAYS AFTER EVT, THE PATIENT¿S HEMIPARESIS COMPLETELY RESOLVED. HOWEVER, ULTRASOUND SHOWED CONTINUED DIAPHRAGMATIC PARALYSIS. MAGNETIC RESONANCE IMAGING (MRI) REVEALED CERVICAL CORD INFARCTION AT THE C2-C3 VERTEBRAL LEVEL, WHICH SUPPORTED THE DIAGNOSIS OF DIAPHRAGMATIC PARALYSIS DUE TO PHRENIC NERVE INJURY. THE PATIENT RECEIVED DAILY RESPIRATORY REHABILITATION THERAPY AND RESPIRATORY FUNCTION WE IMPROVING. HOWEVER, 93 DAYS POST-PROCEDURE, THE PORTABLE VENTILATOR WAS NOT USED AND THE PATIENT DIED WHILE SLEEPING. Manufacturer Narrative: B3. REPORTED EVENT DATE IS THE DATE THE ARTICLE WAS PUBLISHED ONLINE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22861412&pc=DSQ | 2916596-2025-05579 | 2025-08-16 04:00 | Death | THORATEC CORPORATION | 2025-08-21 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE PATIENT OUTCOME THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED, AND THE PUMP WOULD NOT BE EXPLANTED. Manufacturer Narrative: SECTION A: ADDITIONAL PATIENT INFORMATION CANNOT BE PROVIDED DUE TO PERSONAL DATA PRIVACY LEGISLATION/POLICY MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE CUSTOMER COMMUNICATED THAT NO ADDITIONAL INFORMATION WILL BE PROVIDED DUE TO PATIENT PRIVACY LAWS. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) REV. B IS CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22861521&pc=OZP | 2032227-2025-242157 | 2025-07-25 04:00 | Death | MEDTRONIC MINIMED | 2025-08-21 04:00 | OZP | UNKNOWN NGP | Break; Excess Flow or Over-Infusion | Hypoglycemia |  |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY IN 2021 AT HOME. THE CAUSE OF DEATH WAS HYPOGLYCEMIA AND RETAINER RING FAILURE, WHICH RESULTED IN INSULIN OVER-DELIVERY. IT WAS UNKNOWN WHETHER THE INSULIN PUMP WAS WORN AT THE TIME OF PASSING. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: UNK_NGP, UNK_RESERVOIR, AND UNOMEDICAL. TROUBLESHOOTING WAS NOT PERFORMED FOR DECEASED REPORTING. IT WAS UNKNOWN WHETHER THE CUSTOMER HAD BEEN USING THE INSULIN PUMP SYSTEM WITHIN 48 HOURS OF THE REPORTED EVENT AND IT WAS UNKNOWN WHETHER THE CUSTOMER WAS USING SENSORS. THERE IS NO MENTION OF THE AUTO MODE FEATURE AT THE TIME OF THE EVENT. NO PRODUCT RETURN IS REQUIRED FOR UNK_RESERVOIR. NO PRODUCT RETURN IS REQUIRED FOR UNOMEDICAL. PRODUCT RETURN REQUESTED FOR UNK_NGP. CUSTOMER DECLINED TO RETURN OR IS UNABLE TO RETURN. Manufacturer Narrative: CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22861700&pc=BZD | 2518422-2025-109277 | 2025-08-12 04:00 | Death | RESPIRONICS, INC. | 2025-08-21 04:00 | BZD | REP DREAMSTATION AUTO CPAP, CAN | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH A REP DREAMSTATION AUTO CPAP HAS PASSED AWAY. THE REPORTED EVENT MAKES NO ALLEGATION OF ANY SPECIFIC DEVICE MALFUNCTION, USE ERROR, FAILURE, LABELING, OR OTHER DEFICIENCY THAT IS RELEVANT TO THE HARMS REPORTED; THEREFORE, BASED ON INFORMATION AVAILABLE AT THIS TIME, THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO A SERIOUS INJURY OR DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22861764&pc=OZD | 1220648-2025-45794 | 2025-07-26 04:00 | Death | ABIOMED, INC. | 2025-08-21 04:00 | OZD | IMPELLA CP | Malposition of Device | Hemorrhage/Bleeding; Hematuria | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT IMPLANTED WITH AN IMPELLA CP ENCOUNTERED MULTIPLE COMPLICATIONS. THE PATIENT'S URINE WAS NOTED TO BE RED. THE POSITIONING OF THE PUMP WAS CHECKED AND THE DEVICE WAS FOUND TO BE FACING TOWARDS/AGAINST THE MITRAL APPARATUS. THE URINE CLEARED SEVERAL HOURS FOLLOWING REPOSITIONING. TWO DAYS LATER, THE PATIENT BEGAN TO BLEED FROM THEIR ACCESS SITE. MANUAL PRESSURE WAS HELD AND THREE UNITS OF BLOOD WERE TRANSFUSED TO TREAT THE BLEED. THE DRESSING WAS ADJUSTED AND ANGLE MATCHING WAS ACHIEVED AT THE SITE TO RESOLVE THE CONDITION. AFTER ANOTHER TWO DAYS OF SUPPORT, THE PUMP WAS ONCE AGAIN FOUND TO BE DIRECTED TOWARDS THE MITRAL VALVE. THE DOCTOR OPTED TO LEAVE THE PUMP IN THE SUBOPTIMAL POSITION WHILE DECISIONS WERE MADE ON THE PLANS FOR THE PATIENT. DUE TO THE PATIENT'S UNDERLYING POOR CONDITION, THE PATIENT'S FAMILY MADE THE DECISION TO WITHDRAW CARE. THE PATIENT PASSED AWAY FOLLOWING EXPLANT. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: A.5 AND A.6 ARE UNKNOWN. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22862010&pc=LNH | 3010949642-2025-00004 |  | Death | GE HEALTHCARE (TIANJIN) COMPANY LIMITED | 2025-08-21 04:00 | LNH | SIGNA EXPLORER | Insufficient Information | Insufficient Information | K143251 |  | 1 | Event Description: GE HEALTHCARE (GEHC) WAS CONTACTED BY LOCAL AUTHORITIES TO SUPPORT THEIR INVESTIGATION INTO A PATIENT DEATH FOLLOWING A CARDIAC ARREST DURING AN MRI SCAN INVOLVING A GEHC MR SYSTEM. WHILE GEHC HAS LIMITED INFORMATION AT THIS TIME, THE INITIAL REPORT INDICATES THE INCIDENT IS NOT CONSIDERED TO BE RELATED TO THE EQUIPMENT. Manufacturer Narrative: A1, 4, 5, 6: NO ADDITIONAL PATIENT INFORMATION WAS PROVIDED TO GE HEALTHCARE (GEHC). B2, 3: INFORMATION WAS NOT PROVIDED TO GEHC. D: NO ADDITIONAL DEVICE IDENTIFICATION INFORMATION WAS PROVIDED TO GEHC. G2: REPORT SOURCE OTHER, KOREAN NATIONAL POLICE AGENCY. H4: NO DEVICE INFORMATION WAS PROVIDED TO GEHC TO IDENTIFY DATE OF MANUFACTURE. H6: GEHC'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION HAS BEEN COMPLETED. GEHC HAS REQUESTED ADDITIONAL INFORMATION REGARDING THIS EVENT FROM THE REPORT SOURCE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22862252&pc=CBK | 3013876692-2025-00073 | 2025-05-09 04:00 | Death | MAQUET CRITICAL CARE AB | 2025-08-21 04:00 | CBK | BASE UNIT SERVO-U | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K201874 |  | 1 | Event Description: IT WAS REPORTED THAT THE VENTILATOR WAS LEFT IN STANDBY MODE FOR APPROXIMATELY 30 MINUTES, WHICH RESULTED IN SEVERE HARM TO THE PATIENT AND ULTIMATELY LED TO DEATH. MANUFACTURER¿S REF #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22862675&pc=NPT | 2015691-2025-06877 | 2025-07-28 04:00 | Death | EDWARDS LIFESCIENCES | 2025-08-21 04:00 | NPT | EDWARDS COMMANDER DELIVERY SYSTEM | Difficult or Delayed Positioning; Difficult to Remove | Foreign Body In Patient | P140031 |  | 1 | Event Description: AS REPORTED BY THE CLINICAL SPECIALIST, DURING A TRANSFEMORAL TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) CASE THE TEAM WAS UNABLE TO CROSS THE PATIENT'S NATIVE ANNULUS WITH A NON-EDWARDS BAV DUE TO SEVERE TORTUOSITY, THEREFORE NO BAV WAS PERFORMED. WHEN THE VALVE AND DELIVERY SYSTEM WERE BEING ADVANCED TO ATTEMPT TO CROSS THE ANNULUS THE TEAM WAS NEVER WERE ABLE TO ADVANCE THEM FURTHER THAN THE SINUS OF VALSALVA (DELIVERY SYSTEM RAN OUT OF LENGTH)). THERE WAS NEVER ANY ATTEMPT TO WITHDRAW THE DEVICES. DUE TO THE PATIENT'S SEVERELY TORTUOUS ANATOMY THE PHYSICIAN DIDN'T WANT TO RISK ANY INJURY TO THE PATIENT DURING WITHDRAWAL BECAUSE THERE WAS NO SURGEON AVAILABLE IN THE ROOM IF AN INJURY WERE TO OCCUR. THE DECISION WAS MADE TO DEPLOY THE VALVE IN THE PATIENT'S ABDOMINAL/DESCENDING AORTA. THE PLAN WAS TO RESCHEDULE THE PATIENT AT A LATER DATE FOR EITHER SURGICAL VALVE REPLACEMENT OR TAVR VIA SUBCLAVIAN OR CAROTID APPROACH. THE STIFF WIRE CAUSED A DISSECTION AT THE ASCENDING AORTA, BUT IT IS UNCLEAR WHEN THE INJURY OCCURRED. NO INTERVENTION WAS PERFORMED TO TREAT THE DISSECTION, AND THE PATIENT WAS STABLE AFTER THE CASE. UNFORTUNATELY, THE PATIENT EXPIRED ON POST OPERATIVE DAY 2. THEY BELIEVE THE CAUSE OF DEATH WAS THE ASCENDING AORTIC DISSECTION. THE DEVICES WERE NOT RETAINED FOR EVALUATION. Manufacturer Narrative: INVESTIGATION IS ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22863017&pc=OZD | 1220648-2025-31178 | 2025-07-24 04:00 | Death | ABIOMED, INC. | 2025-08-21 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemolysis | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT AN IMPELLA CP WAS PLACED TO SUPPORT A PERCUTANEOUS CORONARY INTERVENTION (PCI) PROCEDURE. AFTER THE PCI, THE PATIENT WAS TAKEN TO THE INTENSIVE CARE UNIT TO RECOVER. THE PATIENT EXPERIENCED HEMOLYSIS AND THE LABORATORY WAS UNABLE TO PROCESS LABS DUE TO GROSS HEMOLYSIS. THE IMPELLA WAS IN OPTIMAL POSITION. IT WAS FURTHER REPORTED THAT THE PATIENT EXPIRED AFTER 72 HOURS OF IMPELLA SUPPORT. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM, SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER: ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE: ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS: ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION: ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22864682&pc=NPT | 9617601-2025-01103 | 2025-08-12 04:00 | Death | MEDTRONIC MEXICO S. DE R.L. DE CV | 2025-08-21 04:00 | NPT | EVOLUT FX PLUS VALVE | Perivalvular Leak; Device Dislodged or Dislocated; Insufficient Information; Patient Device Interaction Problem; Central Regurgitation | Low Blood Pressure/ Hypotension; Foreign Body In Patient; Aortic Valve Insufficiency/ Regurgitation; Renal Impairment | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE IMPLANTATION PROCEDURE, THE PATIENT'S NATIVE ANATOMY WAS COMPLEX, WITH THE LEFT VENTRICULAR OUTFLOW TRACT SIGNIFICANTLY LARGER THAN THE ANNULUS. UPON DEPLOYMENT, THE VALVE DISLODGED INTO THE LEFT VENTRICULAR OUTFLOW TRACT, RESULTING IN SIGNIFICANT AORTIC REGURGITATION DUE TO THE VALVE BEING IMPLANTED TOO DEEP. THE VALVE WAS SNARED AND PULLED UP TO THE ASCENDING AORTA. CARDIOPULMONARY RESUSCITATION (CPR) WAS ADMINISTERED DUE TO THE AORTIC PRESSURE NOT RECOVERING FROM THE SEVERE AORTIC REGURGITATION. A 29 NON-MEDTRONIC VALVE WAS SUBSEQUENTLY LOADED AND IMPLANTED, AND A LEFT MAIN STENT WAS PLACED INTO THE LEFT MAIN ARTERY AS PLANNED POST-TAVI. THE PATIENT WAS SENT TO THE INTENSIVE CARE UNIT FOR RECOVERY. THE PATIENT DID NOT RECOVER AND DIED FOUR DAYS LATER.Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE IMPLANTATION PROCEDURE, THE PATIENT'S NATIVE ANATOMY WAS COMPLEX, WITH THE LEFT VENTRICULAR OUTFLOW TRACT SIGNIFICANTLY LARGER THAN THE ANNULUS. UPON DEPLOYMENT, THE VALVE DISLODGED INTO THE LEFT VENTRICULAR OUTFLOW TRACT (LVOT), RESULTING IN SIGNIFICANT AORTIC REGURGITATION DUE TO THE VALVE BEING IMPLANTED TOO DEEP. THE VALVE WAS SNARED AND PULLED UP TO THE ASCENDING AORTA. CARDIOPULMONARY RESUSCITATION (CPR) WAS ADMINISTERED DUE TO THE AORTIC PRESSURE NOT RECOVERING FROM THE SEVERE AORTIC REGURGITATION. A 29 NON-MEDTRONIC VALVE WAS SUBSEQUENTLY LOADED AND IMPLANTED, AND A LEFT MAIN STENT WAS PLACED INTO THE LEFT MAIN ARTERY (LMA) AS PLANNED POST-TAVI. THE PATIENT WAS SENT TO THE INTENSIVE CARE UNIT FOR RECOVERY. THE PATIENT DID NOT RECOVER AND DIED FOUR DAYS LATER. ADDITIONAL INFORMATION WAS RECEIVED THAT THE REGURGITATION WAS SEVERE, AND BOTH CENTRAL AND PARAVALVULAR LEAK (PVL) DUE TO THE VALVE LANDING IN THE VENTRICLE AND NOT SEALING. THE PATIENT WAS HYPOTENSIVE WITH A VERY LOW AORTIC PRESSURE (30/20). THE STENT IMPLANT WAS A PLANNED PERCUTANEOUS CORONARY INTERVENTION (PCI) OF THE LMA AFTER THE VALVE WAS IMPLANTED DUE TO CORONARY ARTERY DISEASE. PER THE PHYSICIAN, THE CAUSE OF DEATH WAS KIDNEY FUNCTION NOT RECOVERING AFTER THE PROCEDURE. THE VALVE MIGRATING FORWARD IN THE LEFT VENTRICLE OCCURRED DUE TO A LARGE LVOT (MEASURED 99.2 MM) AND VALVE WAS UNABLE TO ANCHOR ONTO ANNULUS. THE PHYSICIAN EXCLUDED THE DCS AS CONTRIBUTING TO THE DEATH. Manufacturer Narrative: UPDATED: B5, E1, E3, H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22864841&pc=NTE | 2124215-2025-56709 | 2017-07-01 04:00 | Death | SILK ROAD MEDICAL | 2025-08-21 04:00 | NTE | ENROUTE TRANSCAROTID NEUROPROTECTION SYSTEM | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Myocardial Infarction; Insufficient Information |  |  | 1 | Event Description: THE AIM OF THIS RETROSPECTIVE STUDY WAS TO ASSESS PREOPERATIVE IMAGING CHARACTERISTICS OF CAROTID ARTERY DISEASE MIGHT BE ASSOCIATED WITH ADVERSE EVENTS FOLLOWING TRANSCAROTID ARTERY REVASCULARIZATION (TCAR). THIS STUDY WAS INCLUDED CASES FROM (B)(6) 2017 TO (B)(6) 2021. A TOTAL OF 213 PATIENTS UNDERWENT TCAR, WITH A MEDIAN AGE OF 74 YEARS, AND 67% WERE MALE. OF THESE PATIENTS, 98% WERE WHITE, 1.4% WERE BLACK, AND 0.5% WERE LATINO. TWENTY-SIX PERCENT WERE CURRENT SMOKERS, 43% WERE FORMER SMOKERS (MEDIAN 10 YEARS SINCE QUITTING), AND 31% HAD NEVER SMOKED. APPROXIMATELY ONE-THIRD OF PATIENTS HAD A HOSTILE NECK DUE TO PRIOR IPSILATERAL CAROTID ENDARTERECTOMY, NECK SURGERY, OR RADIATION THERAPY. SIXTEEN PERCENT UNDERWENT TCAR AS A SECONDARY PROCEDURE FOLLOWING PREVIOUS CAROTID INTERVENTIONS. FORTY-NINE PERCENT OF PROCEDURES WERE PERFORMED ON THE LEFT SIDE, AND 5.6% WERE EMERGENT DUE TO EVOLVING NEUROLOGIC SYMPTOMS. IN ADDITION, 44.6% OF PATIENTS HAD A HISTORY OF TRANSIENT ISCHEMIC ATTACK OR STROKE. THE 30-DAY DEATH RATE OF 0.5% (1/213). FOURTEEN PATIENTS (6.7%) DIED WITHIN ONE YEAR OF THE PROCEDURE, AND 25 PATIENTS (11.96%) HAD DIED BY THE CLOSE OF THE STUDY. OF THE DEATHS THAT OCCURRED WITHIN THE YEAR, ONE HAD A MAJOR STROKE, ONE SUFFERED A MYOCARDIAL INFARCTION (MI) WITHIN 30 DAYS OF TCAR, AND ONE EXPERIENCED BOTH STROKE AND MI. SEVENTY-THREE PERCENT (19/26) OF PATIENT DEATHS REPORTED BY FAMILY MEMBERS TO HAVE DIED AT HOME FROM UNKNOWN CAUSES. Manufacturer Narrative: A2: THE PATIENTS IN THE STUDY RANGED IN AGE FROM 50 TO 89 YEARS, WITH A MEDIAN AGE OF 74 YEARS. A3B: THE STUDY INCLUDED 213 PATIENTS, OF WHOM 143 (67.14%) WERE MALE AND 70 (32.86%) WERE FEMALE. B2. THE DATE OF DEATH WAS NOT PROVIDED FOR ANY OF THE REPORTED DEATHS. B3: THE EXACT EVENT DATE WAS NOT PROVIDED IN THE LITERATURE ARTICLE. THE DATE OF EVENT WAS ESTIMATED USING THE BEGINNING OF THE STUDY PERIOD DATE RANGE, WHICH IS JULY 2017 - NOVEMBER 2021. H6: THE PATIENT CODE E2401WAS USED TO CAPTURE THE REPORTED DEATHS OF UNKNOWN CAUSES. DETAILS OF THE EVENT, INCLUDING PATIENT STATUS AND PRODUCT INFORMATION, WERE NOT PROVIDED TO BSC. BECAUSE THE PRODUCT LOT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) NUMBER AND OTHER SPECIFIC PRODUCT INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. BLEARS, E., PATEL, S., DOYLE, M., LOMBARDI, N., & MULUK, S. (2022). PREDICTING TRANSCAROTID ARTERY REVASCULARIZATION ADVERSE OUTCOMES BY IMAGING CHARACTERISTICS. ANNALS OF VASCULAR SURGERY, 87, 388-401. HTTPS://DOI.ORG/10.1016/J.AVSG.2022.05.013.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22865046&pc=NIQ | 2024168-2025-08844 | 2021-02-01 05:00 | Death | ABBOTT VASCULAR | 2025-08-21 04:00 | NIQ | XIENCE EVEROLIMUS ELUTING CORONARY STENT SYSTEM | Adverse Event Without Identified Device or Use Problem | Unspecified Vascular Problem | P070015 |  | 1 | Event Description: IT WAS REPORTED THROUGH A RESEARCH ARTICLE THAT 4476 ACUTE CORONARY SYNDROME (ACS) PATIENT'S ENROLLED IN THE STOP DAPT-3, WHERE THE NO-ASPIRIN AND DUAL ANTIPLATELET THERAPY (DAPT) STRATEGIES AFTER PERCUTANEOUS CORONARY INTERVENTION (PCI) WITH XIENCE STENTS WERE RANDOMLY COMPARED. THE ANALYSIS WAS BROKEN DOWN INTO FOUR SUBGROUPS HIGH BLEEDING RISK (HBR), NON-HBR AND ST-SEGMENT ELEVATION MYOCARDIAL INFARCTION AND NON-ST-SEGMENT ELEVATION ACS (NSTE-ACS). THE CO-PRIMARY BLEEDING ENDPOINT WAS BLEEDING ACADEMIC RESEARCH CONSORTIUM( BARC) TYPE 3 OR 5, AND THE CO-PRIMARY CARDIOVASCULAR ENDPOINT WAS A COMPOSITE OF CARDIOVASCULAR DEATH, MYOCARDIAL INFARCTION, DEFINITE STENT THROMBOSIS, ACUTE THROMBOSIS, SUBACUTE THROMBOSIS, OR ISCHAEMIC STROKE AT 1 MONTH, WHICH ALL GROUPS SUFFERED. IT WAS ALSO NOTED TARGET LESION REVASCULARIZATION. IN PATIENTS WITH ACS UNDERGOING PCI, THE NO-ASPIRIN STRATEGY COMPARED WITH THE DAPT STRATEGY FAILED TO REDUCE MAJOR BLEEDING EVENTS IRRESPECTIVE OF HBR AND ACS SUBTYPES. THE NUMERICAL EXCESS RISK OF THE NO-ASPIRIN STRATEGY RELATIVE TO THE DAPT STRATEGY FOR CARDIOVASCULAR EVENTS WAS OBSERVED IN PATIENTS WITH HBR AND IN PATIENTS WITH NSTE-ACS. SPECIFIC PATIENT INFORMATION REMAINS UNKNOWN. ADDITIONAL INFORMATION CAN BE FOUND IN THE ATTACHED ARTICLE "ASPIRIN-FREE STRATEGY FOR PERCUTANEOUS CORONARY INTERVENTION IN ACUTE CORONARY SYNDROME BASED ON THE SUBTYPES OF ACUTE CORONARY SYNDROME AND HIGH BLEEDING RISK: THE STOPDAPT-3 TRIAL". NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: THE DEVICES WERE NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORDS FOR THIS PRODUCT COULD NOT BE REVIEWED AND A SIMILAR COMPLAINT QUERY COULD NOT BE PERFORMED BECAUSE THE PART NUMBERS AND LOT NUMBERS WERE NOT REPORTED. A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECTS OF DEATH THE RELATIONSHIP TO PRODUCT, IF ANY, CANNOT BE DETERMINED; HOWEVER, THE SUBSEQUENT UNEXPECTED MEDICAL INTERVENTION APPEARS TO BE RELATED TO THE OPERATIONAL CONTEXT OF THE PROCEDURE. THE INFORMATION RECEIVED IN THE COMPLAINT WAS OBTAINED THROUGH A PROACTIVE LITERATURE SEARCH. FACTORS THAT MAY CONTRIBUTE TO THE PERFORMANCE OUTCOMES LISTED IN THE ARTICLE INCLUDE, BUT ARE NOT LIMITED TO, DEVICE TO PATIENT INTERACTION, PATIENT ANATOMICAL CONDITIONS, DEVICE TO USER INTERACTION AND MANUFACTURING OR MATERIAL ISSUES. DUE TO THE LIMITED INFORMATION AVAILABLE, THE EXACT CAUSE CANNOT BE DETERMINED. THE REPORTED PATIENT EFFECT OF IS DEATH ARE LISTED IN THE INSTRUCTIONS FOR USE AS KNOWN PATIENT EFFECTS OF CORONARY STENTING PROCEDURES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN, OR LABELING. B2, B3, B6, D6: DATES ESTIMATED D4: THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. THE ADDITIONAL PATIENT EFFECTS AND/OR MALFUNCTIONS REPORTED IN THE ARTICLE ARE CAPTURED UNDER A SEPARATE MEDWATCH REPORT. ATTACHMENT: ARTICLE TITLED "ASPIRIN-FREE STRATEGY FOR PERCUTANEOUS CORONARY INTERVENTION IN ACUTE CORONARY SYNDROME BASED ON THE SUBTYPES OF ACUTE CORONARY SYNDROME AND HIGH BLEEDING RISK: THE STOPDAPT-3 TRIAL".
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22865143&pc=NTE | 2124215-2025-57039 | 2016-01-01 05:00 | Death | SILK ROAD MEDICAL | 2025-08-21 04:00 | NTE | ENROUTE TRANSCAROTID NEUROPROTECTION SYSTEM | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage |  |  | 1 | Event Description: IT WAS REPORTED VIA LITERATURE THAT ONE PATIENT DIED AS A RESULT OF INTRACRANIAL HEMORRHAGE OCCURRING ONE DAY POST TRANSCAROTID ARTERY REVASCULARIZATION (TCAR). THIS STUDY EVALUATED THE PERIOPERATIVE (30-DAY) AND 1-YEAR OUTCOMES OF TRANSCAROTID ARTERY REVASCULARIZATION (TCAR) AMONG 167 PATIENTS FROM JANUARY 2016 TO AUGUST 2019. 82 PATIENTS WERE CLASSIFIED AS SYMPTOMATIC, DEFINED BY A HISTORY OF UNILATERAL TRANSIENT ISCHEMIC ATTACK OR STROKE ATTRIBUTABLE TO AN EXTRACRANIAL CAROTID ARTERY LESION WITHIN THE PREVIOUS 180 DAYS, AND 85 WERE CLASSIFIED AS ASYMPTOMATIC. THE DEMOGRAPHIC CHARACTERISTICS BETWEEN THE TWO GROUPS WERE LARGELY SIMILAR, ALTHOUGH SYMPTOMATIC PATIENTS WERE SIGNIFICANTLY MORE LIKELY TO BE FEMALE (28.0% VS. 9.4%). THE AVERAGE AGE ACROSS ALL PATIENTS WAS APPROXIMATELY 71 YEARS. PERIOPERATIVE (30-DAY) OUTCOMES INDICATED THAT SYMPTOMATIC PATIENTS HAD A HIGHER PERIOPERATIVE MORTALITY RATE OF 4.9%, COMPARED TO 0% IN ASYMPTOMATIC PATIENTS. ONLY ONE OF THE FOUR DEATHS WAS FOUND AS ATTRIBUTABLE TO THE TCAR PROCEDURE, WHICH OCCURRED DUE TO AN INTRACRANIAL HEMORRHAGE ON POSTOPERATIVE DAY (POD) 1, ATTRIBUTED TO CEREBRAL REPERFUSION SYNDROME. THE REMAINING THREE DEATHS WERE DUE TO CONGESTIVE HEART FAILURE EXACERBATION ON POD 27 AND POD 28, AND ONE OF UNKNOWN CAUSE ON POD 23. AT THE 1-YEAR FOLLOW-UP, OVERALL MORTALITY WAS 13.0% IN SYMPTOMATIC PATIENTS AND 10.0% IN ASYMPTOMATIC PATIENTS. Manufacturer Narrative: B3: THE EXACT EVENT DATE WAS NOT PROVIDED IN THE LITERATURE ARTICLE. THE DATE OF EVENT WAS ESTIMATED USING THE BEGINNING OF THE STUDY PERIOD DATE RANGE, WHICH IS JANUARY 2016 TO AUGUST 2019. B2. THE DATE OF DEATH WAS NOT PROVIDED. DETAILS OF THE EVENT, INCLUDING PRODUCT INFORMATION, WERE NOT PROVIDED TO BSC. BECAUSE THE PRODUCT LOT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) NUMBER AND OTHER SPECIFIC PRODUCT INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. WANG, S. K., SEVERANCE, S., WESTIN, G. G., MAIJUB, J. G., GUPTA, A. K., SAWCHUK, A. P., FAJARDO, A. C., & MOTAGANAHALLI, R. L. (2020). PERIOPERATIVE AND 1-YEAR TRANSCAROTID REVASCULARIZATION OUTCOMES IN SYMPTOMATIC PATIENTS. JOURNAL OF VASCULAR SURGERY, 72(6), 2047-2053. HTTPS://DOI.ORG/10.1016/J.JVS.2020.03.044.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22865209&pc=DSQ | 2916596-2025-05514 | 2025-08-08 04:00 | Death | THORATEC CORPORATION | 2025-08-21 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Renal Failure; Tachycardia; Loss of Vision; Hemorrhagic Stroke; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE A SUDDEN AND COMPLETE LOSS OF VISION IN BOTH EYES. THE WAS STROKE ACTIVATED. COMPUTED TOMOGRAPHY (CT) OF THE HEAD AND COMPUTED TOMOGRAPHY ANGIOGRAPHY OF THE HEAD AND NECK SHOWED A DIFFUSE SPONTANEOUS INTRACEREBRAL BLEED. (HEMORRHAGIC STROKE). THE PATIENT HAD AN INTERNATIONAL NORMALIZED RATIO OF 1.5 AND WAS PUT ON COUMADIN WITH A HEPARIN INFUSION BRIDGE. THEIR ANTI-FACTOR XA WAS AT 0.30 AND THEY WERE TREATED WITH K-CENTRA AND DDAVP. NEUROSURGERY REPORTED THEM AS NON-OPERABLE. THE SYMPTOMS DID NOT RESOLVE WHICH PROVOKED THE PATIENT TO REQUEST HOSPICE CARE. IT WAS BELIEVED THAT CIRRHOSIS WAS A PRECIPITATING FACTOR TO THE BLEED. THE PATIENT HAD MULTIPLE COMORBIDITIES PRIOR TO LEFT VENTRICULAR ASSIST DEVICE IMPLANT, INCLUDING STAGE IV CHRONIC KIDNEY DISEASE (CKD), CIRRHOSIS DUE TO HEPATIC CONGESTION, AND ADRENAL INSUFFICIENCY. AS A RESULT, THEY HAD MULTIPLE COMPLICATIONS AFTER THE IMPLANT. THEY WENT INTO END STAGE RENAL FAILURE REQUIRING PERITONEAL DIALYSIS AND HAD AN ASPIRATION PNEUMONIA DUE TO GASTROINTESTINAL DYSMOTILITY WHICH PROGRESS TO A PSEUDOMONAS BACTEREMIA. THE EXPERIENCED MULTIPLE EPISODES OF VENTRICULAR TACHYCARDIA, RESISTANT TO MEDICAL TREATMENT. THEN THEY PROGRESSED TO DEVELOPING A SPONTANEOUS INTRACEREBRAL BLEED CAUSING TOTAL BLINDNESS. THIS WAS WHAT CAUSED THE PATIENT AND THE FAMILY TO REQUEST WITHDRAWAL OF LIFESAVING CARE AND PROCEED TO COMFORT MEASURES ONLY. THEIR PACEMAKER WAS TURNED OFF PER PATIENT REQUEST AND THEY WERE PLACED ON PAIN AND ANXIETY MEDICATIONS. THE PATIENT PASSED AWAY ON (B)(6) 2025 DUE TO HEART FAILURE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING STROKE, BLEEDING, CARDIAC ARRHYTHMIA, MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (INCLUDING RENAL FAILURE AND HEPATIC DYSFUNCTION), AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER "ANTICOAGULATION"), OUTLINES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INR RANGE FOR PATIENTS USING THE HEARTMATE 3 LVAS, AS WELL AS THE SUGGESTED ANTICOAGULATION MODIFICATIONS IN THE EVENT THAT THERE IS A RISK OF BLEEDING. THIS SECTION ALSO LISTS ARRHYTHMIA AS A POTENTIAL LATE POST-IMPLANT COMPLICATION THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22865388&pc=NIQ | 2024168-2025-08823 | 2015-03-02 05:00 | Death | ABBOTT VASCULAR | 2025-08-21 04:00 | NIQ | XIENCE EVEROLIMUS ELUTING CORONARY STENT SYSTEM | Adverse Event Without Identified Device or Use Problem | Unspecified Heart Problem | P070015 |  | 1 | Event Description: IT WAS REPORTED IN A SUMMARY ARTICLE THAT THE 3-YEAR SAFETY AND PERFORMANCE OUTCOMES OF A NON-ABBOTT EVEROLIMUS-ELUTING CORONARY STENT. THE ARTICLE REFERS TO SEVERAL CLINICAL TRIALS (BIOFLOW-IV, BIOFLOW V, DEESOLVE II AND TALENT) AS A COMPARISON OF THE OUTCOMES AFTER PERCUTANEOUS CORONARY INTERVENTION (PCI) WITH OTHER STENTS INCLUDING THE XIENCE STENT. THE COMPARISON END-POINTS AT 12-MONTH, 2 YEAR AND 3 YEAR FOLLOW-UP WERE CARDIAC DEATH, MYOCARDIAL INFARCTION, CORONARY BYPASS GRAFT, TARGET VESSEL REVASCULARIZATION, LATE STENT THROMBOSIS WHICH OCCURRED WITH ALL THE STENTS INCLUDING THE XIENCE STENT. ADDITIONAL INFORMATION CAN BE FOUND IN THE SUMMARY ARTICLE, "SAFETY AND EFFICACY OF EVERMINE50 EVEROLIMUS-ELUTING CORONARY STENT SYSTEM IN PATIENTS WITH NATIVE CORONARY ARTERY LESIONS: THREE-YEAR OUTCOMES FROM A SINGLE-CENTER" NO ADDITIONAL INFORMATION PROVIDED. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORDS FOR THIS PRODUCT COULD NOT BE REVIEWED AND A SIMILAR COMPLAINT QUERY COULD NOT BE PERFORMED BECAUSE THE PART NUMBER AND/OR LOT NUMBER WAS NOT REPORTED, AND THE DEVICE WAS NOT RETURNED. A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT DEATH AND THE RELATIONSHIP TO PRODUCT IF ANY, CANNOT BE DETERMINED. THE INFORMATION RECEIVED IN THE COMPLAINT WAS OBTAINED THROUGH A PROACTIVE LITERATURE SEARCH. SPECIFIC MODEL AND LOT NUMBERS WERE NOT AVAILABLE AND CAUSALITY BETWEEN THE DOCUMENTED DEVICE USAGE AND THE PATIENT EFFECTS COULD NOT BE ESTABLISHED. THE REPORTED PATIENT EFFECTS OF DEATH, MYOCARDIAL INFARCTION, AND THROMBOSIS/ THROMBUS ARE LISTED IN THE XIENCE EVEROLIMUS ELUTING CORONARY STENT SYSTEMS INSTRUCTIONS FOR USE AS A KNOWN PATIENT EFFECT(S) OF CORONARY STENTING PROCEDURES. FACTORS THAT MAY CONTRIBUTE TO THE PERFORMANCE OUTCOMES LISTED IN THE ARTICLE INCLUDE, BUT ARE NOT LIMITED TO, DEVICE TO PATIENT INTERACTION, PATIENT ANATOMICAL CONDITIONS, DEVICE TO USER INTERACTION AND MANUFACTURING OR MATERIAL ISSUES. DUE TO THE LIMITED INFORMATION AVAILABLE, THE EXACT CAUSE CANNOT BE DETERMINED. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN, OR LABELING; THEREFORE, NO PRODUCT-RELATED CORRECTIVE ACTION WILL BE IMPLEMENTED IN THIS CASE. B2: DEATH DATE ESTIMATED. B3: DATE OF EVENT ESTIMATED. D4: THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. D6A: IMPLANT DATE ESTIMATED. THE ADDITIONAL PATIENT EFFECTS REFERENCED IN B5 ARE CAPTURED UNDER SEPARATE MEDWATCH REPORT. LITERATURE ATTACHMENT: ARTICLE TITLE: "SAFETY AND EFFICACY OF EVERMINE50 EVEROLIMUS-ELUTING CORONARY STENT SYSTEM IN PATIENTS WITH NATIVE CORONARY ARTERY LESIONS: THREE-YEAR OUTCOMES FROM A SINGLE-CENTER".
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22865549&pc=OZD | 1220648-2025-45883 | 2024-10-18 04:00 | Death | ABIOMED, INC. | 2025-08-21 04:00 | OZD | IMPELLA CP | Device Sensing Problem | Anemia; Hemolysis | P140003 |  | 1 | Event Description: THE COMPLAINTS TEAM RECEIVED A NOTIFICATION VIA ABIOMED¿S LONG-TERM OUTCOME AND QUALITY INDICATOR (LOQI) IMPELLA REGISTRY REGARDING A PATIENT THAT ENDURED WORSENING ANEMIA WITH A "POSSIBLE" RELATIONSHIP TO THE IMPELLA DEVICE USE. ADDITIONALLY, IT WAS NOTED THAT THE PATIENT ENDURED WORSENING HEMOLYSIS WITH A "PROBABLE" RELATIONSHIP TO THE IMPELLA DEVICE. NO CONCERN FOR ACTIVE BLEEDING BUT THIS LEAD TO IMPELLA REMOVAL AND POSSIBLE INTRA-AORTIC BALLOON PUMP INSERTION. THE ISSUE DID NOT RESOLVE. THERE WAS ALSO A CALL REGARDING A PLACEMENT SIGNAL LOW ALARM. THE TEAM WAS ADVISED TO PERFORM AN ECHOCARDIOGRAM AND REPOSITION THE DEVICE. ULTIMATELY, CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA DEVICE AS AN ASSOCIATED FACTOR IN THE PATIENT'S DEMISE. Manufacturer Narrative: A.4, A.5 AND A.6 ARE UNKNOWN. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA IFU: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22866621&pc=OZD | 1220648-2025-45895 | 2025-08-05 04:00 | Death | ABIOMED, INC. | 2025-08-21 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Tachycardia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT IMPLANTED WITH AN IMPELLA 5.5 ENDURED RUNS OF VENTRICULAR TACHYCARDIA. AN ABLATION WAS PLANNED TO TREAT THE CONDITION, HOWEVER IT WAS LATER REPORTED THAT THE PATIENT PASSED AWAY DUE TO A MASSIVE CARDIAC ARRHYTHMIA. THE PATIENT WAS ON SUPPORT AT THE TIME OF THEIR PASSING Manufacturer Narrative: THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. PRODUCT STATUS: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22868989&pc=QSY | 3004138549-2025-00005 | 2021-12-21 05:00 | Death | Z-MEDICA, LLC | 2025-08-21 04:00 | QSY | QUIKCLOT HEMOSTATIC DRESSING, 4X4, 10X | Adverse Event Without Identified Device or Use Problem | Airway Obstruction; Nasal Obstruction | K123387 |  | 1 | Event Description: REPORTED VIA A MEDWATCH REPORT, THE REPORTER STATES "ON (B)(6) 2021 PATIENT WITH EPISTAXIS ANO RIGHT NARE PACKED WITH QUICKCLOT X3 SPRAYED WITH AFRIN. ON (B)(6) 2021 RN NOTES INDICATE NASAL PACKING REMOVED. ON (B)(6) 2021 PATIENT DECOMPENSATES AND TRANSFERS TO ICU. UNDERGOES FLEXIBLE FIBEROPTIC BRONCHOSCOPY WITH FINDINGS OF GAUZE-LIKE MATERIAL EXTENDING FROM LEFT BRONCHUS ABOVE LEVEL OF CARINA TO THE DISTAL TRACHEA. CRITICAL CARE NOTES INDICATE MUL TL FACTORIAL CIRCULATORY SHOCK WITH HYPERCAPNIA RELATED TO AIRWAY OBSTRUCTION AND SEPTIC PNA". THE REPORT STATES THAT THE PATIENT CONDITION IS "DECEASED". Manufacturer Narrative: (B)(4). REPORT RECIEVED VIA MEDWATCH REPORT #MW5107443. PLEASE NOTE THAT 3 DEVICES WERE REPORTED, NOT 3 SEPARATE EVENTS. COMPLAINT VERIFICATION TESTING COULD NOT BE PERFORMED AS NO SAMPLE WAS RETURNED FOR ANALYSIS. A DEVICE HISTORY RECORD REVIEW COULD NOT BE PERFORMED AS NO LOT NUMBER WAS REPORTED. WITHOUT THE DEVICE TO EVALUATE THE COMPLAINT COULD NOT BE CONFIRMED AND THE PROBABLE CAUSE COULD NOT BE DETERMINED FROM THE AVAILABLE INFORMATION. TELEFLEX WILL CONTINUE TO MONITOR AND TREND FOR REPORTS OF THIS NATURE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22869319&pc=QEY | 2124215-2025-57998 | 2025-07-25 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-08-21 04:00 | QEY | EKOS+ | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Low Blood Pressure/ Hypotension; Asystole |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE PATIENT PRESENTED WITH PULMONARY EMBOLISM (PE) AND HAD EXTREME RIGHT HEART STRAIN. THE PATIENT WAS TO UNDERGO TREATMENT UTILIZING AN EKOS+ CATHETER. PRIOR TO THE PROCEDURE, THE PATIENT WAS VISIBLY SHORT OF BREATH. THROUGHOUT THE EKOS PROCEDURE, THE PATIENT HAD MULTIPLE CARDIAC EVENTS OF VENTRICULAR TACHYCARDIA, AND UNSTABLE BLOOD PRESSURE. THE PROCEDURE WAS COMPLETED WITH ONE EKOS+ 106X12 CATHETER IN THE RIGHT PULMONARY ARTERY. THERE NO ISSUES WITH THE EKOS+ DEVICE. WHILE IN TRANSIT TO THE INTENSIVE CARE UNIT, THE PATIENTS PULSE WAS LOST. THE PATIENT WAS NOTED TO BE DO NOT RESUSCITATE (DNR) AND ULTIMATELY PASSED AWAY. Manufacturer Narrative: DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22869497&pc=NRY | 3005168196-2025-00269 | 2025-07-25 04:00 | Death | PENUMBRA, INC. | 2025-08-21 04:00 | NRY | PENUMBRA SYSTEM RED 72 REPERFUSION CATHETER | Fracture | Foreign Body Embolism | K211654 |  | 1 | Event Description: THE PATIENT WAS UNDERGOING INTRAVENOUS THROMBOLYSIS FOLLOWED BY MECHANICAL THROMBECTOMY PROCEDURE IN THE LEFT INTERNAL CAROTID ARTERY USING A PENUMBRA SYSTEM RED 72 REPERFUSION CATHETER (RED72), A NEURON MAX 6F 088 LONG SHEATH (NEURON MAX), A NON-PENUMBRA MICROCATHETER, AND A GUIDEWIRE. IT WAS REPORTED THAT THE PATIENT HAD A STENOSIS ON THE LEFT COMMON CAROTID ARTERY (CCA). DURING THE PROCEDURE, WHILE ATTEMPTING TO REMOVE THE RED72, THE PHYSICIAN EXPERIENCED RESISTANCE. THE PHYSICIAN INJECTED CONTRAST, AND NOTICED UNDER FLUOROSCOPY THAT APPROXIMATELY 10 CENTIMETERS OF THE RED72 HAD FRACTURED AND WAS LODGED IN THE ICA, EXTENDING FROM THE CAROTID BIFURCATION TO THE CAROTID SIPHON. THE PHYSICIAN DECIDED TO END THE PROCEDURE AT THIS POINT WITHOUT RECANALIZATION OF THE CAROTID TERMINUS AND THE MIDDLE CEREBRAL ARTERY (MCA). THE PHYSICIAN ALSO DECIDED TO LEAVE THE FRACTURED SEGMENT OF THE RED72 IN THE PATIENT. THE NEXT DAY, A CEREBRAL COMPUTED TOMOGRAPHY (CT) ANGIOGRAPHY REVEALED THE FRACTURED SEGMENT OF THE RED72 WITHIN THE (CCA), EXTENDING TO THE LEFT ICA CAUSING COMPLETE OBSTRUCTION OF THE BEGINNING OF THE ICA AS WELL AS A NEAR OCCLUSIVE STENOSIS OF THE LEFT (CCA), WHICH REMAINED PATENT. ON (B)(6) 2025, IT WAS REPORTED THAT THE PATIENT EXPIRED ON (B)(6) 2025 DUE TO THE STROKE SUFFERED ON (B)(6) 2025. Manufacturer Narrative: THIS DEVICE IS AVAILABLE FOR RETURN. A FOLLOW UP MDR WILL BE SUBMITTED UPON COMPLETION OF THE DEVICE INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22869831&pc=DSQ | 2916596-2025-05390 | 2025-07-20 04:00 | Death | THORATEC CORPORATION | 2025-08-21 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Fall | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO WITHDRAWAL FROM CARE STATUS POST A FALL. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. IT WAS ALSO REPORTED THAT THE DEVICE WAS OPERATING AT NORMAL PARAMETERS DURING THE TIME OF THE EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22847098&pc=LJS | 3006260740-2025-05967 | 2024-09-26 04:00 | Death | C.R. BARD, INC. (BASD) -3006260740 | 2025-08-20 04:00 | LJS | UNKNOWN CHRONIC CATHETER | Adverse Event Without Identified Device or Use Problem | Insufficient Information | UNKNOWN |  | 1 | Event Description: IT WAS REPORTED IN THE JOURNAL OF NURSING REPORT OF ARTICLE TITLED, "NURSE-DRIVEN INTERVENTIONS REDUCE CENTRAL LINE-ASSOCIATED BLOODSTREAM INFECTION CLOSE TO ZERO IN ONE PEDIATRIC ONCOLOGIC FACILITY: A SINGLE-CENTER RETROSPECTIVE OBSERVATIONAL STUDY", THAT SOMETIME POST CENTRAL LINE PLACEMENT PROCEDURE BY USING BARD BROVIAC/HICKMAN CATHETER TO EVALUATE THE NURSE DRIVEN INTERVENTIONS TO REDUCE CENTRAL LINE ASSOCIATED BLOOD STREAM INFECTION IN PEDIATRIC ONCOLOGY FACILITY. THE PATIENT WAS REPORTEDLY EXPIRED. HOWEVER, THE CAUSE OF DEATH WAS NOT PROVIDED. Manufacturer Narrative: H11: TUROLDO F, LONGO A, SALA M, VALENTINI D, DE VITA N, TONIUTTI S, ZUPPEL L, MAXIMOVA N. NURSE-DRIVEN INTERVENTIONS REDUCE CENTRAL LINE-ASSOCIATED BLOODSTREAM INFECTION CLOSE TO ZERO IN ONE PEDIATRIC ONCOLOGIC FACILITY: A SINGLE-CENTER RETROSPECTIVE OBSERVATIONAL STUDY. NURS REP. 2024 SEP 26;14(4):2668-2679. DOI: 10.3390/NURSREP14040197. PMID: 39449434; PMCID: PMC11503445. AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION / EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE / PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. DATE OF DEATH FOR THIS PATIENT WAS NOT PROVIDED, DATE OF DEATH UPDATED AS 01-JAN-1900 FOR THE MDR SUBMISSION REQUIREMENT. SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22847276&pc=LWQ | 3005687633-2025-00125 | 2025-07-02 04:00 | Death | CORCYM S.R.L. | 2025-08-20 04:00 | LWQ | OPTIFORM PROSTHETIC MITRAL HEART VALVE | Inadequacy of Device Shape and/or Size | Mitral Valve Insufficiency/ Regurgitation |  |  | 1 | Event Description: ON (B)(6) 2025, THE PATIENT UNDERWENT DOUBLE VALVE REPLACEMENT SURGERY VIA FULL STERNOTOMY DUE TO BOTH AORTIC AND MITRAL VALVE INSUFFICIENCY. THE AORTIC VALVE WAS FIRST SUCCESSFULLY REPLACED WITH A CARBOMEDICS TOP HAT VALVE, MODEL S5-021 (SN (B)(6). IN THE SAME PROCEDURE, MITRAL VALVE REPLACEMENT WAS INITIATED. THE NATIVE MITRAL LEAFLETS AND SUBVALVULAR CHORDS WERE COMPLETELY EXCISED. THE ANNULUS WAS SIZED USING THE DEDICATED SIZERS MODEL VS-200 WITHOUT ANY DIFFICULTY, AND A CARBOMEDICS OPTIFORM F7-27 VALVE WAS SELECTED FOR IMPLANTATION. FOLLOWING THE INITIAL IMPLANTATION OF THE F7-27, ECHOCARDIOGRAPHIC EVALUATION REVEALED HIGH-SPEED REGURGITANT FLOW WITHIN THE MITRAL VALVE. THE VALVE WAS REMOVED AND RE-IMPLANTED, BUT THE ABNORMAL FLOW PERSISTED. CONSEQUENTLY, THE F7-27 WAS REPLACED WITH A CARBOMEDICS OPTIFORM F7-025 VALVE (SN (B)(6), WHICH WAS IMPLANTED IN THE SAME POSITION. POST-IMPLANTATION IMAGING CONFIRMED THE DISAPPEARANCE OF THE REGURGITATION. THE EXTRACORPOREAL CIRCULATION TIME DURING THE SURGERY WAS 523 MINUTES, AND THE CROSS-CLAMP TIME WAS APPROXIMATELY 400 MINUTES. DUE TO THIS INCIDENT, THE CROSS-CLAMP TIME INCREASED BY ABOUT 150 MINUTES AND THE EXTRACORPOREAL CIRCULATION TIME INCREASED BY ABOUT 240 MINUTES. THE BLOOD PRESSURE DURING EXTRACORPOREAL CIRCULATION FLUCTUATED AROUND 50 MMHG. ULTIMATELY, THE PATIENT PASSED AWAY ON (B)(6) 2025. NO AUTOPSY HAS BEEN PLANNED. THE PATIENT HAD NO REPORTED COMORBIDITIES OR RISK FACTORS BEYOND THE VALVE INSUFFICIENCIES, AND NO ANATOMICAL ABNORMALITIES OF THE MITRAL ANNULUS WERE OBSERVED. THE OPERATING SURGEON IS NOTED TO BE EXPERIENCED WITH CARBOMEDICS VALVE IMPLANTATION, PERFORMING APPROXIMATELY 30¿40 CASES PER YEAR. BASED ON THE SURGEON¿S EVALUATION THE POSSIBLE CAUSE OF THE REGURGITATION OBSERVED IN THE F7-27 PROSTHESIS COULD BE RELATED TO A POSSIBLE INTERFERENCE BETWEEN THE TOP HAT VALVE IN THE AORTIC POSITION AND THE F7-027 VALVE IN THE MITRAL POSITION. Manufacturer Narrative: THE MANUFACTURING AND MATERIAL RECORDS FOR THE DEVICE INVOLVED IN THIS EVENT, AS THEY PERTAIN TO THE REPORTED ISSUE, WERE RETRIEVED AND REVIEWED BY THE MANUFACTURER¿S QUALITY ENGINEERING. THE RESULTS CONFIRMED THAT THIS PROSTHESIS SATISFIED ALL REQUIRED MATERIAL, VISUAL, AND PERFORMANCE STANDARDS AT THE TIME OF MANUFACTURE AND RELEASE. MANUFACTURER WILL PERFORM INVESTIGATION UPON RETURN OF THE DEVICE AND WILL SEND FOLLOW-UP REPORT UPON RECEIPT AND/OR COMPLETION OF THE INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22847682&pc=DTB | 3008973940-2025-06277 | 2025-06-28 04:00 | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-08-20 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Cardiac Tamponade; Diabetic Ketoacidosis; Pericardial Effusion; Insufficient Information | P930039 |  | 1 | Event Description: IT WAS REPORTED BY A PATIENT REPRESENTATIVE THAT APPROXIMATELY FOUR DAYS POST IMPLANT THE PATIENT DIED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22847700&pc=DTB | 2649622-2025-22623 | 2025-06-28 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-08-20 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Cardiac Tamponade; Diabetic Ketoacidosis; Pericardial Effusion; Insufficient Information | P930039 |  | 1 | Event Description: IT WAS REPORTED BY A PATIENT REPRESENTATIVE THAT APPROXIMATELY FOUR DAYS POST IMPLANT THE PATIENT DIED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22848248&pc=BZD | 2518422-2025-048940 | 2025-08-08 04:00 | Death | RESPIRONICS, INC. | 2025-08-20 04:00 | BZD | DREAMSTATION AUTO CPAP | Contamination; Degraded | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THERE WAS NO ALLEGATION REPORTED BY THE PATIENT. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION, DIRT CONTAMINATION WAS OBSERVED. THERE WAS NO ERROR CODE FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT SCRAPPED. Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED FROM DUPLICATE RA ALLEGING PATIENT PASSED AWAY. BOX B: ADVERSE EVENT OR PRODUCT PROBLEM. BOX H: TYPE OF REPORTED COMPLAINT AND HEALTH IMPACT GRID WERE UPDATED IN THIS REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22848764&pc=DTB | 2182208-2025-04109 | 2025-07-13 04:00 | Death | MEDTRONIC, INC. | 2025-08-20 04:00 | DTB | CAPSURE EPICARDIAL LEAD | Low impedance | Spinal Cord Injury | P950024 |  | 1 | Event Description: IT WAS REPORTED BY THE PATIENT REPRESENTATIVE THAT THE PATIENT PASSED AWAY. IT WAS ALSO REPORTED THAT BOTH THE RIGHT ATRIAL (RA) AND RIGHT VENTRICULAR (RV) LEADS EXHIBITED LOW OUT OF RANGE MEASUREMENTS SUDDENLY AND WAS QUERYING IF THE LEADS WERE WORKING.Event Description: IT WAS FURTHER REPORTED BY THE PATIENT'S REPRESENTATIVE THAT THE PATIENT SPOKE WITH THEIR HEART DOCTOR AND WAS INFORMED THAT THE LEAD IMPEDANCE ISSUE WAS NOT A FACTOR IN THE PATIENT'S DEATH. Manufacturer Narrative: CORRECTED B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: 4968-35 LEAD WAS IMPLANTED ON (B)(6) 2008. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22848814&pc=OZD | 1220648-2025-45745 | 2025-07-26 04:00 | Death | ABIOMED, INC. | 2025-08-20 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT IMPLANTED WITH AN IMPELLA 5.5 EXPERIENCED A HEMORRHAGIC BRAIN BLEED. NO NOTED INTERVENTION WAS REQUIRED FOR THE BLEED. THE PATIENT'S NEUROLOGY STATUS CONTINUED TO DECLINE AND THE PATIENT PASSED AWAY THE FOLLOWING DAY. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22848818&pc=DTB | 2182208-2025-04110 | 2025-07-13 04:00 | Death | MEDTRONIC, INC. | 2025-08-20 04:00 | DTB | CAPSURE EPICARDIAL LEAD | Low impedance | Insufficient Information | P950024 |  | 1 | Event Description: IT WAS REPORTED BY THE PATIENT REPRESENTATIVE THAT THE PATIENT PASSED AWAY. IT WAS ALSO REPORTED THAT BOTH THE RIGHT ATRIAL (RA) AND RIGHT VENTRICULAR (RV) LEADS EXHIBITED LOW OUT OF RANGE MEASUREMENTS SUDDENLY AND WAS QUERYING IF THE LEADS WERE WORKING. Manufacturer Narrative: CONTINUATION OF D10: 4968-35 LEAD WAS IMPLANTED ON (B)(6) 2008. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22849583&pc=LYJ | 1644487-2025-10517 | 2025-06-06 04:00 | Death | HOUSTON | 2025-08-20 04:00 | LYJ | PULSE GEN MODEL 1000 | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P970003 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD PASSED AWAY. THE REPRESENTATIVE WAS ABLE TO FIND A POST STATING THE PATIENT PASSED ON (B)(6) 2025 AND ANOTHER POST SUGGESTING SHE PASSED FROM SUICIDE. NO SURGICAL INTERVENTION HAS BEEN REPORTED TO DATE. NO ADDITIONAL RELEVANT INFORMATION HAS BEEN RECEIVED TO DATE. Manufacturer Narrative: LIVANOVA USA, INC. SUBMITS THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION, BASED ON INFORMATION THAT LIVANOVA HAS OBTAINED, BUT MAY NOT HAVE BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION, OR A CONCLUSION BY FDA OR ANYONE ELSE, THAT THE DEVICE, LIVANOVA, OR LIVANOVA'S EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS¿ OR "MALFUNCTIONS¿. THESE WORDS ARE INCORPORATED INTO THE FDA 3500A MEDWATCH FORM BY THE FDA, AND LIVANOVA OBJECTS TO THEIR USE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22849863&pc=NPT | 2015691-2025-06826 | 2025-08-06 04:00 | Death | EDWARDS LIFESCIENCES | 2025-08-20 04:00 | NPT | EDWARDS CERTITUDE DELIVERY SYSTEM | Burst Container or Vessel; Difficult to Remove; Device Slipped; Detachment of Device or Device Component | Vascular Dissection; Insufficient Information | P140031 |  | 1 | Event Description: AS REPORTED BY OUR FRENCH AFFILIATES, DURING IMPLANT OF A 29MM SAPIEN 3 VALVE IN THE AORTIC POSITION BY SUBCLAVIAN APPROACH, DUE TO THE NARROW SIZE OF THE SUBCLAVIAN ARTERY, THE VALVE PARTIALLY MOVED ON THE BALLOON. THEN, DURING VALVE DEPLOYMENT, AT THE END OF A 3-SECOND INFLATION HOLD WITH NOMINAL VOLUME, THE BALLOON BURST. DUE TO THE BURST, THE BALLOON COULD NOT BE COMPLETELY WITHDRAWN THROUGH THE SHEATH, AND THE DISTAL PART OF THE CERTITUDE SYSTEM DETACHED, RESULTING IN THE LOSS OF THE NOSECONE AND PART OF THE BALLOON INTO THE PATIENT'S DESCENDING AORTA. THE LASSO TECHNIQUE WAS REPEATEDLY ATTEMPTED AND EVENTUALLY SUCCEEDED IN REMOVING THE ENTIRE BALLOON AND NOSECONE FROM THE PATIENT, BUT IT REQUIRED A DISSECTION OF THE AORTA AND PART OF THE SUBCLAVIAN ARTERY. THE SUBCLAVIAN ARTERY WAS STENTED, AND THE PATIENT REMAINED STABLE FOR TWO HOURS. THE AORTIC DISSECTION, HOWEVER, WORSENED LEADING TO CARDIAC TAMPONADE WITH CARDIAC ARREST AND THE PATIENT EXPIRED. Manufacturer Narrative: THE INVESTIGATION IS ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22850971&pc=KDI | 2521402-2025-00299 | 2025-07-22 04:00 | Death | B. BRAUN AVITUM AG | 2025-08-20 04:00 | KDI | DIALOG+® | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest |  |  | 1 | Event Description: AS REPORTED BY THE USER FACILITY: AN 84-YEAR-OLD MALE PATIENT WITH A SIGNIFICANT MEDICAL HISTORY WAS UNDERGOING HEMODIALYSIS THERAPY USING AN OPTIFLUX F180NR DIALYZER. DURING THE SESSION, APPROXIMATELY 2.44 HOURS INTO THERAPY, THE PATIENT BECAME UNRESPONSIVE. THE DIALYSIS NURSE NOTED HYPOTENSION DURING TREATMENT AND SUBSEQUENTLY OBSERVED THE PATIENT TO BE UNRESPONSIVE, WITH BLOOD PRESSURE AND PULSE OXIMETRY READINGS UNOBTAINABLE. BLOOD WAS RETURNED TO THE PATIENT, AND THE ATTENDING PHYSICIAN ASSESSED THE PATIENT. GIVEN THE PATIENT'S DO NOT RESUSCITATE (DNR) STATUS, NO MEDICAL INTERVENTION WAS PERFORMED. THE PATIENT WAS PRONOUNCED DECEASED AT (B)(6) 2025. NOTE FROM THE MANUFACTURER: A TECHNICAL SAFETY INSPECTION OF THE MACHINE WAS PERFORMED. THE TSI WAS PASSED, CONFIRMING THAT THE MACHINE WORKS AS INTENDED. THE INVESTIGATION OF THE RECEIVED MACHINE DATA RECORD FROM (B)(6) 2025, 8:18 O´CLOCK SHOWS THAT DURING PREPARATION ALL SELF-TESTS WERE PASSED, AND THERAPY WAS STARTED. APPROX. 21 MINUTES AFTER THERAPY START THE MACHINE STARTED TO TRIGGER THE ALARM "TMP OUT OF ALARM LIMITS" BECAUSE THE TRANSMEMBRANE PRESSURE TMP EXCEEDED THE LIMIT. THIS ALARM HAS NO FURTHER MACHINE REACTION AND IS ONLY TO INFORM THE USER. THIS ALARM WAS TRIGGERED FURTHER 35 TIMES UNTIL APPROX. 56 MINUTES AFTER THERAPY START. THE MACHINE DATA RECORD SHOWS THAT THE TRANSMEMBRANE PRESSURE TMP WAS IN THE RANGE OF -15 TO -10 MMHG AT THERAPY START, INCREASED TO -5 TO 5 MMHG WHEN THE ALARMS STARTED, AND INCREASED UP TO 20 MMHG DURING THERAPY. THE TMP VALUES FROM THE MACHINE DATA RECORD DO NOT INDICATE A MALFUNCTION OF THE MACHINE. AN AUTOMATIC BLOOD PRESSURE MEASUREMENT ABPM WAS PERFORMED AT APPROX. 49 MINUTES AFTER THERAPY START WITH THE RESULT 78/48 MMHG, PULSE RATE 80 BPM. AS THE SYSTOLIC PRESSURE WAS TOO LOW, THE MACHINE TRIGGERED THE ALARM "ABPM: SYS PRESS. IS TOO LOW". AT APPROX. 1:02 HOURS AFTER THERAPY START A FURTHER ABPM MEASUREMENT WAS PERFORMED WITH THE RESULT 79/50 MMHG, PULSE RATE 65 BPM. AS THE SYSTOLIC PRESSURE WAS STILL TOO LOW, THE MACHINE TRIGGERED THE ALARM "ABPM: SYS PRESS. IS TOO LOW". 1 MINUTES LATER A FURTHER ABPM MEASUREMENT WAS PERFORMED WITH THE RESULT 78/46 MMHG, PULSE RATE 86 BPM. AS THE SYSTOLIC PRESSURE WAS STILL TOO LOW, THE MACHINE TRIGGERED THE ALARM "ABPM: SYS PRESS. IS TOO LOW". AT APPROX. 2:03 HOURS AFTER THERAPY THE MACHINE TRIGGERED THE ALARM "VENOUS PRESSURE - UPPER LIMIT" BECAUSE THE VENOUS PRESSURE UPPER LIMIT WAS EXCEEDED AND SWITCHED INTO A PATIENT SAFE MODE (WITH BLOOD PUMP STOP). DUE TO THE STOPPED BLOOD PUMP THE MACHINE TRIGGERED THE FOLLOW-UP ALARMS "VENOUS PRESSURE - LOWER LIMIT - CHECK ACCESS" AND "VENOUS PRESSURE LOWER LIMIT (SUP)". 2 SECONDS LATER, THESE ALARMS WERE ACKNOWLEDGED BY THE USER. AT APPROX. 2:19 HOURS AFTER THERAPY START A FURTHER ABPM MEASUREMENT WAS PERFORMED UNSUCCESSFULLY. THEREFORE, THE MACHINE TRIGGERED THE ALARM "ABPM: PULSE SIGNAL VERY LOW". 3 MINUTES LATER A FURTHER ABPM MEASUREMENT WAS PERFORMED UNSUCCESSFULLY AGAIN. THEREFORE, THE MACHINE TRIGGERED THE ALARM "ABPM: PULSE SIGNAL VERY LOW". SHORTLY AFTER THIS, THE VENOUS PRESSURE PV INCREASED TO 280 MMHG AND EXCEEDED THE UPPER LIMIT. THEREFORE, THE MACHINE TRIGGERED THE ALARM "VENOUS PRESSURE - UPPER LIMIT". THIS INCREASE DOES NOT INDICATE A MALFUNCTION OF THE MACHINE. 2 SECONDS LATER, THIS ALARM WAS ACKNOWLEDGED BY THE USER. 19 SECONDS LATER, THE MACHINE TRIGGERED THE ALARM "VENOUS PRESSURE - LOWER LIMIT - CHECK ACCESS" AGAIN AND THIS ALARM WAS ACKNOWLEDGED 2 SECONDS LATER. 2 MINUTES LATER, THE MACHINE TRIGGERED THE ALARM "VENOUS PRESSURE - UPPER LIMIT" AGAIN AND 3 SECONDS LATER ACKNOWLEDGED BY THE USER. AT APPROX. 2:27 HOURS AFTER THERAPY START A FURTHER ABPM MEASUREMENT WAS PERFORMED UNSUCCESSFULLY. THEREFORE, THE MACHINE TRIGGERED THE ALARM "ABPM: PULSE SIGNAL VERY LOW". 24 SECONDS LATER, THE MACHINE TRIGGERED THE ALARM "VENOUS PRESSURE - UPPER LIMIT" AGAIN. 2 SECONDS LATER, THE ALARM WAS ACKNOWLEDGED BY THE USER. BUT THE ALARM WAS TRIGGERED AGAIN A FEW SECONDS LATER. 2 SECONDS LATER THE ALARM WAS ACKNOWLEDGED BY THE USER. APPROX. 1 MINUTE LATER, THE MACHINE TRIGGERED THE ALARM "ARTERIAL PRESSURE - UPPER LIMIT" BECAUSE THE ARTERIAL PRESSURE EXCEEDED THE UPPER LIMIT. THIS ALARM WAS ALSO ACKNOWLEDGED 2 SECONDS LATER. AT APPROX. 2:32 HOURS AFTER THERAPY START, NO MORE "RED" WAS DETECTED AT THE RED DETECTOR RDV OF THE SAFETY AIR DETECTOR SAD, I.E. THE VENOUS LINE WAS FILLED WITH TRANSPARENT FLUID. IT IS PLAUSIBLE THAT THE PATIENT WAS REINFUSED WITH SALINE FROM A BAG WHICH WOULD ALSO EXPLAIN THE INCREASE OF THE ARTERIAL PRESSURE PA. SHORTLY AFTER THIS, THE MACHINE TRIGGERED THE ALARM "ARTERIAL PRESSURE - LOWER LIMIT". DUE TO THE STOPPED BLOOD PUMP THE MACHINE TRIGGERED THE FOLLOW-UP ALARMS "VENOUS PRESSURE - LOWER LIMIT - CHECK ACCESS" AND "VENOUS PRESSURE LOWER LIMIT (SUP)". THESE ALARMS WERE ACKNOWLEDGED BY THE USER. BUT THE MACHINE TRIGGERED THE ALARM "ARTERIAL PRESSURE MONITORING ERROR (SUP)" BECAUSE THE MACHINE DID NOT DETECT ENOUGH ARTERIAL PRESSURE PULSATION. THE ARTERIAL PRESSURE VALUE REMAINED AT -250 SINCE THE LAST 2 MINUTES. THE BLOOD PUMP REMAINED STOPPED FOR THE REST OF THIS MACHINE DATA RECORD. A FURTHER ABPM MEASUREMENT WAS PERFORMED UNSUCCESSFULLY, AND THE MACHINE TRIGGERED THE ALARM "ABPM: PULSE SIGNAL VERY LOW". AT APPROX. 2:39 HOURS AFTER THERAPY START, THE USER CHANGED FROM THERAPY TO END OF THERAPY PHASE. THE ANALYSIS OF THE MACHINE DATA RECORD AND ALL RELEVANT PARAMETERS DOCUMENTED THEREIN REVEALED NO ANOMALIES OR INDICATIONS OF A MALFUNCTION OF THE DEVICE. THERE IS NO PRODUCT DEVIATION. IN CONCLUSION, BASED ON THE ANALYSIS AND EVALUATION OF THE REPORTED AND PROVIDED INFORMATION THERE WAS NOT MALFUNCTION OF THE DEVICE AND THE DEVICE HAS NOT LED TO THE PATIENT'S DEATH. THIS CASE IS CLASSIFIED AS NOT REPORTABLE TO THE COMPETENT AUTHORITY BY THE MANUFACTURER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22851339&pc=OZP | 2032227-2025-241010 | 2025-03-31 04:00 | Death | MEDTRONIC MINIMED | 2025-08-20 04:00 | OZP | UNKNOWN NGP | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Hyperglycemia; Hypoglycemia |  |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY DUE TO HIGH BLOOD PRESSURE, CARDIAC ARREST AND DIABETES. THE CUSTOMER PASSED AWAY ON (B)(6) 2025. THE CUSTOMER PUMP WAS WORN AT THE TIME OF EVENT. THE CUSTOMER BLOOD GLUCOSE VALUE WAS UNKNOWN. THE CUSTOMER VISITED THE EMERGENCY ROOM IN THE HOSPITAL. THE EVENT INVOLVED PRODUCT UNK_NGP, UNK_RESERVOIR, UNK_SENSOR, UNOMEDICAL. TROUBLESHOOTING WAS PERFORMED FOR GATHERING THE DECEASED RELATED INFORMATION. CUSTOMER WAS USING THE INSULIN PUMP WITHIN 48 HOURS OF THE REPORTED EVENT. IT WAS UNKNOWN WHETHER THE AUTO MODE FEATURE WAS ACTIVE AT THE TIME OF INCIDENT. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED. NO PRODUCT RETURN WAS REQUIRED FOR UNK_NGP, UNK_RESERVOIR, UNK_SENSOR, UNOMEDICAL. Manufacturer Narrative: CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22851476&pc=OZD | 1220648-2025-45914 | 2025-07-28 04:00 | Death | ABIOMED, INC. | 2025-08-20 04:00 | OZD | IMPELLA CP | Restricted Flow rate | Hemorrhage/Bleeding; Low Blood Pressure/ Hypotension | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT AND THE FOLLOWING DAY EXHIBITED BLEEDING AT THE FEMORAL INSERTION SITE. THE BLEEDING WAS TREATED WITH A SUTURE. THE PATIENT'S BLOOD PRESSURE WAS LOW SO EPINEPHRINE AND VASO WERE GIVEN. ADDITIONALLY, IT WAS NOTED THERE WERE NO ADDITIONAL SUCTION EVENTS SINCE THE PRIOR LATE NIGHT. THE PATIENT WAS NOTED TO STILL BE A CANDIDATE FOR A TRANSCATHETER AORTIC VALVE REPLACEMENT AFTER RECOVERY FROM IMPELLA EVENT. THE PLAN WAS FOR THE PATIENT TO REST, RECOVER AND TEAM MONITORS THE IMPELLA INSERTION SITE. HOWEVER LATER THIS DAY IN EVENING, THE PATIENT EXPIRED. CARE WAS ULTIMATELY WITHDRAWN. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22851592&pc=OJX | 2649622-2025-22704 | 2025-07-17 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-08-20 04:00 | OJX | ATTAIN ABILITY PLUS MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Heart Failure/Congestive Heart Failure | P080006 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS ADMITTED TO THE EMERGENCY ROOM WITH ALERTED MENTAL STATUS AND WAS FOUND TO HAVE POSITIVE BLOOD CULTURES FOR METHICILLIN-RESISTANT STAPHYLOCOCCUS AUREUS BACTEREMIA. THE PATIENT WAS TREATED WITH ANTIBIOTICS. A TRANSTHORACIC ECHOCARDIOGRAM (TTE) SHOWED A SMALL, MOBILE ECHODENSITY NOTED ON RIGHT ATRIAL LEAD. DEVICE REMOVAL WAS RECOMMENDED, BUT THE PATIENT WAS NOT CONSIDERED A CANDIDATE FOR THE PROCEDURE DUE TO COMORBIDITIES. DURING THE STAY IN THE EMERGENCY ROOM, THE PATIENT EXPERIENCED WORSENING RENAL FUNCTION AND DOBUTAMINE INFUSION WAS STARTED. THE DEVICE WAS SUBSEQUENTLY DEACTIVATED AND TRANSFERRED TO HOSPICE CARE. THE PATIENT DECLINED DUE TO TERMINAL COMPLICATIONS RELATED TO CONGESTIVE HEART FAILURE AND PASSED AWAY. THE PATIENT WAS A PARTICIPANT IN A CLINICAL STUDY. Manufacturer Narrative: CONTINUATION OF D10: DTPA2D1 CRT-D, IMPLANTED: (B)(6) 2024; 383069 LEAD, IMPLANTED: (B)(6) 2019. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22851595&pc=NVN | 2649622-2025-22705 | 2025-07-17 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-08-20 04:00 | NVN | SELECTSECURE MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Heart Failure/Congestive Heart Failure | P030036 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS ADMITTED TO THE EMERGENCY ROOM WITH ALERTED MENTAL STATUS AND WAS FOUND TO HAVE POSITIVE BLOOD CULTURES FOR METHICILLIN-RESISTANT STAPHYLOCOCCUS AUREUS BACTEREMIA. THE PATIENT WAS TREATED WITH ANTIBIOTICS. A TRANSTHORACIC ECHOCARDIOGRAM (TTE) SHOWED A SMALL, MOBILE ECHODENSITY NOTED ON RIGHT ATRIAL LEAD. DEVICE REMOVAL WAS RECOMMENDED, BUT THE PATIENT WAS NOT CONSIDERED A CANDIDATE FOR THE PROCEDURE DUE TO COMORBIDITIES. DURING THE STAY IN THE EMERGENCY ROOM, THE PATIENT EXPERIENCED WORSENING RENAL FUNCTION AND DOBUTAMINE INFUSION WAS STARTED. THE DEVICE WAS SUBSEQUENTLY DEACTIVATED AND TRANSFERRED TO HOSPICE CARE. THE PATIENT DECLINED DUE TO TERMINAL COMPLICATIONS RELATED TO CONGESTIVE HEART FAILURE AND PASSED AWAY. THE PATIENT WAS A PARTICIPANT IN A CLINICAL STUDY. Manufacturer Narrative: CONTINUATION OF D10: DTPA2D1 CRT-D, IMPLANTED: ON (B)(6) 2024; 429688, IMPLANTED ON (B)(6) 2013. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22851603&pc=DTB | 2649622-2025-22706 | 2025-07-17 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-08-20 04:00 | DTB | CAPSUREFIX NOVUS MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Heart Failure/Congestive Heart Failure | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS ADMITTED TO THE EMERGENCY ROOM WITH ALERTED MENTAL STATUS AND WAS FOUND TO HAVE POSITIVE BLOOD CULTURES FOR METHICILLIN-RESISTANT STAPHYLOCOCCUS AUREUS BACTEREMIA. THE PATIENT WAS TREATED WITH ANTIBIOTICS. A TRANSTHORACIC ECHOCARDIOGRAM (TTE) SHOWED A SMALL, MOBILE ECHODENSITY NOTED ON RIGHT ATRIAL LEAD. DEVICE REMOVAL WAS RECOMMENDED, BUT THE PATIENT WAS NOT CONSIDERED A CANDIDATE FOR THE PROCEDURE DUE TO COMORBIDITIES. DURING THE STAY IN THE EMERGENCY ROOM, THE PATIENT EXPERIENCED WORSENING RENAL FUNCTION AND DOBUTAMINE INFUSION WAS STARTED. THE DEVICE WAS SUBSEQUENTLY DEACTIVATED AND TRANSFERRED TO HOSPICE CARE. THE PATIENT DECLINED DUE TO TERMINAL COMPLICATIONS RELATED TO CONGESTIVE HEART FAILURE AND PASSED AWAY. THE PATIENT WAS A PARTICIPANT IN A CLINICAL STUDY. Manufacturer Narrative: CONTINUATION OF D10: 383069 LEAD, IMPLANTED: (B)(6) 2019; DTPA2D1 CRT-D, IMPLANTED: (B)(6) 2024; 429688, IMPLANTED (B)(6) 2013. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22851723&pc=LWS | 2649622-2025-22707 | 2025-07-17 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-08-20 04:00 | LWS | SPRINT FIDELIS | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Heart Failure/Congestive Heart Failure | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS ADMITTED TO THE EMERGENCY ROOM WITH ALERTED MENTAL STATUS AND WAS FOUND TO HAVE POSITIVE BLOOD CULTURES FOR METHICILLIN-RESISTANT STAPHYLOCOCCUS AUREUS BACTEREMIA. THE PATIENT WAS TREATED WITH ANTIBIOTICS. A TRANSTHORACIC ECHOCARDIOGRAM (TTE) SHOWED A SMALL, MOBILE ECHODENSITY NOTED ON RIGHT ATRIAL LEAD. DEVICE REMOVAL WAS RECOMMENDED, BUT THE PATIENT WAS NOT CONSIDERED A CANDIDATE FOR THE PROCEDURE DUE TO COMORBIDITIES. DURING THE STAY IN THE EMERGENCY ROOM, THE PATIENT EXPERIENCED WORSENING RENAL FUNCTION AND DOBUTAMINE INFUSION WAS STARTED. THE DEVICE WAS SUBSEQUENTLY DEACTIVATED AND TRANSFERRED TO HOSPICE CARE. THE PATIENT DECLINED DUE TO TERMINAL COMPLICATIONS RELATED TO CONGESTIVE HEART FAILURE AND PASSED AWAY. THE PATIENT WAS A PARTICIPANT IN A CLINICAL STUDY. Manufacturer Narrative: CONTINUATION OF D10: 383069 LEAD, IMPLANTED: (B)(6) 2019; DTPA2D1 CRT-D, IMPLANTED: (B)(6) 2024; 429688, IMPLANTED (B)(6) 2013. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22851729&pc=NIK | 9614453-2025-03289 | 2025-07-17 04:00 | Death | MEDTRONIC EUROPE SARL | 2025-08-20 04:00 | NIK | COBALT¿ XT HF CRT-D MRI SURESCAN¿ | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Heart Failure/Congestive Heart Failure | P010031 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS ADMITTED TO THE EMERGENCY ROOM WITH ALERTED MENTAL STATUS AND WAS FOUND TO HAVE POSITIVE BLOOD CULTURES FOR METHICILLIN-RESISTANT STAPHYLOCOCCUS AUREUS BACTEREMIA. THE PATIENT WAS TREATED WITH ANTIBIOTICS. A TRANSTHORACIC ECHOCARDIOGRAM (TTE) SHOWED A SMALL, MOBILE ECHODENSITY NOTED ON RIGHT ATRIAL LEAD. DEVICE REMOVAL WAS RECOMMENDED, BUT THE PATIENT WAS NOT CONSIDERED A CANDIDATE FOR THE PROCEDURE DUE TO COMORBIDITIES. DURING THE STAY IN THE EMERGENCY ROOM, THE PATIENT EXPERIENCED WORSENING RENAL FUNCTION AND DOBUTAMINE INFUSION WAS STARTED. THE DEVICE WAS SUBSEQUENTLY DEACTIVATED AND TRANSFERRED TO HOSPICE CARE. THE PATIENT DECLINED DUE TO TERMINAL COMPLICATIONS RELATED TO CONGESTIVE HEART FAILURE AND PASSED AWAY. THE PATIENT WAS A PARTICIPANT IN A CLINICAL STUDY. Manufacturer Narrative: CONTINUATION OF D10: 383069 LEAD, IMPLANTED: (B)(6) 2019; 429688, IMPLANTED (B)(6) 2013. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22852233&pc=MKJ | 1220908-2025-03069 | 2025-07-16 04:00 | Death | ZOLL MEDICAL CORPORATION | 2025-08-20 04:00 | MKJ | AED 3 | Unexpected Shutdown | Insufficient Information | P160022 |  | 1 | Event Description: COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO DEFIBRILLATE A PATIENT (AGE & GENDER UNKNOWN), THE DEVICE INAPPROPRIATELY SHUTDOWN. COMPLAINANT INDICATED THAT THE CLINICIAN OBTAINED ANOTHER DEVICE TO CONTINUE TREATING THE PATIENT. COMPLAINANT INDICATED THAT THE PATIENT SUBSEQUENTLY EXPIRED. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED TO ZOLL MEDICAL CORPORATION, INSTEAD THE ACTIVITY LOGS WERE PROVIDED FOR REVIEW. REVIEW OF THE LOG FILE FROM THE REPORTED EVENT CONCLUDED THAT THE DEVICE WORKED AS DESIGNED AND WITHIN THE LIMITATIONS OF THE TECHNOLOGY. A REVIEW OF THE LOGS SHOWED SUCCESSFULLY COMPLETING SELF-TESTS IN CLINICAL MODE, INDICATING THE DEVICE WAS OPERATIONAL AND READY FOR USE. SHORTLY AFTER THERE WAS A NON-CRITICAL ERROR INDICATING THE BATTERY WAS BELOW THE CAPACITY THRESHOLD. THIS SUGGESTS THE DEVICE WAS OPERATING WITH A BATTERY THAT WAS ALREADY NEARING DEPLETION PRIOR TO THE RESCUE EVENT. THE DEVICE LOGGED THAT THE BUTTON WAS PRESSED, RELEASED, AND THE SYSTEM WAS POWERED OFF WHILE IN CLINICAL MODE. IT APPEARS THE DEVICE DID NOT SHUT DOWN UNEXPECTEDLY OR AS A RESULT OF SYSTEM FAILURE. THE POWER-OFF EVENT WAS USER-INITIATED, AND SUBSEQUENT BEHAVIOR WAS CONSISTENT WITH EXPECTED DEVICE OPERATION UNDER LOW BATTERY CONDITIONS. IT'S NOTEWORTHY TO MENTION ACCORDING TO THE ZOLL AED 3 ADMINISTRATOR'S GUIDE, THE DEVICE IS CAPABLE OF DELIVERING AT LEAST THREE MAXIMUM ENERGY (200-JOULE) SHOCKS AFTER A "CHANGE BATTERY" MESSAGE APPEARS. THE OPERATOR IS ADVISED TO REPLACE THE BATTERY PACK AS SOON AS POSSIBLE ONCE THIS PROMPT IS DISPLAYED. ANALYSIS OF REPORTS OF THIS TYPE HAS NOT IDENTIFIED AN INCREASE IN TREND.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22852242&pc=OZD | 1220648-2025-45788 | 2025-07-28 04:00 | Death | ABIOMED, INC. | 2025-08-20 04:00 | OZD | IMPELLA CP | Restricted Flow rate | Hematuria | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT FOLLOWING FOLEY PLACEMENT IN THE CORONARY CARE UNIT, AN IMPELLA CP PATIENT'S URINE BECAME DARK RED. IT WAS NOTED THAT INTERMITTENT SUCTION ALARMS WERE OCCURRING LEADING UP TO THE NOTED CONDITION. DESPITE REPOSITIONING THE PUMP, THE ISSUE PERSISTED AND THE PATIENT'S HEMODYNAMIC PRESSURES CONTINUED TO DROP. THE PATIENT PASSED AWAY ON SUPPORT. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: A.4 AND A.5 ARE UNKNOWN. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER: ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE: ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS: ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS.¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION: ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22852511&pc=NPT | 9617601-2025-01081 | 2025-08-13 04:00 | Death | MEDTRONIC MEXICO S. DE R.L. DE CV | 2025-08-20 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Appropriate Term/Code Not Available; Patient Device Interaction Problem | Low Blood Pressure/ Hypotension; Pericardial Effusion | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING THE IMPLANT OF A TRANSCATHETER AORTIC VALVE, THE PATIENT'S BLOOD PRESSURE DROPPED. AN ECHOCARDIOGRAM WAS PERFORMED THAT DETECTED PERICARDIAL EFFUSION. SUBSEQUENTLY, BLOOD WAS REMOVED FROM THE PERICARDIUM VIA PERICARDIOCENTESIS. THE PERICARDIUM WAS NOTED TO BE DRY AND THE VALVE WAS FUNCTIONING WELL. A FEW HOURS LATER, ANOTHER BLOOD PRESSURE DROP OCCURRED, AND A LARGE PERICARDIAL EFFUSION WAS DETECTED BY ECHOCARDIOGRAM. DESPITE ALL CARE PROVIDED, THE PATIENT DIED A FEW HOURS AFTER THE PROCEDURE. PER THE PHYSICIAN, THE CAUSAL/CONTRIBUTARY FACTORS COULD HAVE BEEN PERICARDIAL EFFUSION DUE TO THE GUIDEWIRE, DISSECTION ABOVE THE SINOTUBULAR JUNCTION (STJ), AND FISSURES OF THE ANNULUS FOR INDETERMINATE REASONS. THE PATIENT HAD SIGNIFICANT LEFT VENTRICULAR HYPERTROPHY. THE PHYSICIAN ASSESSED THAT THE VALVE WAS NOT RESPONSIBLE FOR THE EVENT.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT FOLLOWING THE VALVE IMPLANT, BLEEDING AROUND THE ANNULUS WAS ALSO OBSERVED. PER THE PHYSICIAN, THE DELIVERY CATHETER SYSTEM (DCS) DID NOT CAUSE OR CONTRIBUTE TO THE PERICARDIAL EFFUSION, DISSECTION OR THE BLEEDING AROUND THE ANNULUS. A CONFIDA GUIDEWIRE WAS USED DURING THE PROCEDURE. Manufacturer Narrative: CONTINUATION OF D10: SECTION D INFORMATION REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER RELEVANT DEVICE(S) ARE: PRODUCT ID: D- EVOLUTFX-2329, SERIAL/LOT#: (B)(6), UBD: 28-MAR-2027, UDI#: (B)(4) CONTINUATION OF D10: PRODUCT ID: {GUIDEWIRE}; PRODUCT TYPE: {GUIDEWIRE}; SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5 CONTINUATION OF D10: PRODUCT ID GWBC30; PRODUCT LOT/SERIAL NUMBER N/A; PRODUCT TYPE: GUIDEWIRE; EXPLANT DATE H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22852526&pc=MSD | 1282497-2025-00556 | 2025-07-14 04:00 | Death | COVIDIEN LLC | 2025-08-20 04:00 | MSD | DIALYSIS UNKNOWN | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Hemorrhage/Bleeding |  |  | 1 | Event Description: ACCORDING TO THE LITERATURE SOURCE OF STUDY, A SINGLE-CENTER, RETROSPECTIVE OBSERVATIONAL STUDY AT THE PICU (PEDIATRIC INTENSIVE CARE UNIT) WAS CONDUCTED BETWEEN FEBRUARY 2019 AND MAY 2023. THE VASCULAR ACCESS WAS ESTABLISHED USING A DOUBLE-LUMEN DIALYSIS CATHETER, INCLUDING BOTH THE COMPANY'S PRODUCT AND A COMPETITOR'S PRODUCT. CATHETER SIZES INCLUDED 8 FR (FRENCH), 9 FR, 10 FR, AND 11.5 FR. OF THE 2989 PATIENTS ADMITTED TO PICU OVER THE FOUR-YEAR PERIOD, 52 CHILDREN (1.7%) RECEIVED CRRT (CONTINUOUS RENAL REPLACEMENT THERAPY) ACROSS 71 TREATMENT SESSIONS OVER 2556.5 HOURS. OUT OF THE TOTAL PATIENTS WHO UNDERWENT TREATMENT WITH DIFFERENT CATHETER SIZES, 10 PATIENTS WHO USED THE 8 FR CATHETER DID NOT SURVIVE. SIMILARLY, 21 PATIENTS WHO USED THE 9 FR CATHETER DID NOT SURVIVE, FOLLOWED BY 3 PATIENTS WHO USED THE 10 FR CATHETER AND 5 PATIENTS WHO USED THE 11.5 FR CATHETER. IN TOTAL, 39 PATIENTS DIED BEFORE DISCHARGE ACROSS ALL CATHETER SIZES. OUT OF THE 39 PATIENTS WHO DID NOT SURVIVE, TWO PATIENTS DEVELOPED INTRACRANIAL BLEEDING WHILE ON CRRT AND PASSED AWAY. ADDITIONALLY, ONE PATIENT EXPERIENCED CARDIAC ARREST DURING THE BLOOD RETURN PHASE OF CRRT AND PASSED AWAY. LABORATORY RESULTS REVEALED A POST-ARREST SERUM POTASSIUM LEVEL OF 8.1 MEQ/L (MILLIEQUIVALENTS PER LITER), INDICATING SEVERE HYPERKALEMIA. ONE PATIENT, DIAGNOSED WITH ACUTE LEUKEMIA AND TUMOR LYSIS SYNDROME, WAS DISCHARGED FROM THE HOSPITAL AFTER CRRT BUT DIED 0.2 YEARS LATER. Manufacturer Narrative: D10 - CONCOMITANT PRODUCT: DIALYSIS UNKNOWN, UNK DY, LOT: UNKNOWN DIALYSIS UNKNOWN, UNK DY, LOT: UNKNOWN SEKHAR, JERIN C. ¿CONTINUOUS RENAL REPLACEMENT THERAPY (CRRT) PROGRAM INITIATION IN PICU OF A RESOURCE LIMITED SETTING: A RETROSPECTIVE ANALYSIS OF CHALLENGES AND OUTCOMES.¿ BMC NEPHROLOGY, 2025. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22853206&pc=NPT | 2015691-2025-06848 | 2025-07-30 04:00 | Death | EDWARDS LIFESCIENCES | 2025-08-20 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Rupture; Low Cardiac Output | P140031 |  | 1 | Event Description: EDWARDS RECEIVED NOTIFICATION FROM OUR AFFILIATE IN FRANCE. AS REPORTED, THIS WAS A CASE OF AN IMPLANT OF A 26MM SAPIEN 3 ULTRA IN AORTIC POSITION BY TRANSFEMORAL APPROACH. AFTER THE VALVE DEPLOYMENT, THE PATIENT'S PRESSURE DROPPED AND THE RHYTHM BECAME UNSTABLE WITH LOW OUTPUT SUSPECTED. IT WAS DECIDED TO START THE RESUSCITATION AND DURING A CONTROL ANGIOGRAPHY IT WAS CONFIRMED THAT AN ANNULAR RUPTURE HAD OCCURRED. AFTER PROLONGED RESUSCITATION THE PATIENT, UNFORTUNATELY, PASSED AWAY. AS PER MEDICAL OPINION, THE ROOT CAUSE WAS THE CALCIFICATION OF THE ANNULUS. IT WAS STATED THAT THE ANATOMY AND DEGREE OF CALCIFICATION WOULD LIKELY HAVE LED TO POST-DILATATION EVEN WITH A SELF-EXPANDING VALVE, POSSIBLY RESULTING IN THE SAME OUTCOME. THE MEDICAL TEAM EXPRESSED THAT UNFORTUNATELY, SUCH RARE COMPLICATIONS CAN OCCUR DESPITE BEST PRACTICES. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. ACCORDING TO THE DEVICE TRAINING MANUALS, RISK FACTORS FOR AORTIC DISSECTION, CARDIAC/AORTIC HEMATOMA, OR ANNULAR RUPTURE DURING THE TVR PROCEDURE INCLUDE SIGNIFICANT VALVE OVER SIZING, SEVERELY OBLITERATED SINUSES OF VALSALVA, PORCELAIN AORTA AND/OR PRESENCE OF BULKY CALCIFICATION, AND NARROW CALCIFIED SINOTUBULAR JUNCTION. IN ADDITION, ADVANCED AGE, FEMALE GENDER, SMALL BODY WEIGHT, AND STEROID DEPENDENCY CAN ALSO BE CONTRIBUTING FACTORS. THE SAPIEN TRANSCATHETER HEART VALVE (ALL MODELS) RELIES CALCIUM TO SECURELY ANCHOR IN THE LANDING ZONE. DESPITE THIS BENEFICIAL ASPECT OF CALCIUM, BULKY CALCIUM CAN INCREASE THE RISK OF CALCIFIC NODULE DISPLACEMENT INTO THE VASCULATURE, WHICH CAN LEAD TO VASCULAR INJURY. AT TIMES, THE EXTENT AND DISTRIBUTION OF CALCIUM CAN IMPAIR EASE OF DELIVERY OF THE VALVE, CORRECT POSITIONING OF THE VALVE, DEPLOYMENT OF THE VALVE, AND PROCEDURAL SUCCESS. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST THAT PATIENT RELATED FACTORS (ANNULAR CALCIFICATION) CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22853262&pc=MHX | 9610816-2025-000764 | 2025-07-26 04:00 | Death | PHILIPS MEDIZIN SYSTEME BÖBLINGEN GMBH | 2025-08-20 04:00 | MHX | INTELLIVUE MULTI MEASUREMENT SERVER X2 | No Audible Alarm | Cardiac Arrest | K150310 |  | 1 | Event Description: IT WAS REPORTED THE NURSES DID NOT HEAR A RED ALARM AND THE PATIENT PASSED AWAY. THE PATIENT WAS BEING MONITORED WITH THE X2 MONITOR, WHICH THE STAFF WAS AWARE WAS NOT NETWORKED WITH A CENTRAL STATION. WHEN A RED ALARM OCCURRED, THE NURSES DID NOT HEAR AN ALARM, AND THE PATIENT PASSED AWAY. THE CUSTOMER REQUESTED ASSISTANCE DETERMINING WHEN THE INOP FIRST OCCURRED. Manufacturer Narrative: THE MANUFACTURING DATE FOR THE REPORTED DEVICE IS PRIOR TO 24SEPT2016 SO NO UDI REQUIRED. WHEN THE BIOMEDICAL ENGINEER RETRIEVED THE X2, AN INOP INDICATING A SPEAKER FAULT WAS NOTED BUT THERE WAS NO SOUND. THE SPEAKER AND WI-FI ANTENNA WERE REPLACED. PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22853381&pc=QFG | 3013756811-2025-196531 | 2024-11-24 05:00 | Death | TANDEM DIABETES CARE | 2025-08-20 04:00 | QFG | T:SLIM X2 INSULIN PUMP WITH CONTROL-IQ TECHNOLOGY | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Hyperglycemia | K203234 |  | 1 | Event Description: IT WAS REPORTED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2024. REPORTEDLY, THE CUSTOMER EXPERIENCED HIGH BLOOD GLUCOSE LEVELS BEFORE SUFFERING A HEART ATTACK. IT WAS REPORTED THAT A BENT CANNULA ON THE INFUSION SET MAY HAVE CONTRIBUTED TO THE HIGH BLOOD GLUCOSE LEVELS. THE CUSTOMER WAS WEARING THE PUMP AT THE TIME OF DEATH. A REVIEW OF PUMP DATA WILL BE PERFORMED, AND THE RESULTS WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22853475&pc=NGV | 2124215-2025-57568 | 2025-07-15 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-08-20 04:00 | NGV | WATCHMAN FLX? | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Thrombosis/Thrombus | P130013 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WAS A PATIENT DEATH. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED. A WATCHMAN ACCESS SYSTEM (WAS) AND A 27MM WATCHMAN FLX LAA CLOSURE DEVICE & DELIVERY SYSTEM (WDS) WERE SELECTED TO BE USED. THE PROCEDURE WAS SUCCESSFULLY COMPLETED WITH THE CLOSURE DEVICE IMPLANTED IN THE LAA OF THE PATIENT. THERE WERE NO PATIENT COMPLICATIONS THAT WERE REPORTED TO HAVE OCCURRED. TWO (2) YEARS POST PROCEDURE THE PATIENT WAS ADMITTED TO THE HOSPITAL AFTER EXPERIENCING A CEREBRAL VASCULAR ACCIDENT. TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) IMAGING SHOWED THAT THERE WAS A DEVICE RELATED THROMBUS. THE PATIENT DID NOT RECOVER AND ULTIMATELY THE PATIENT DIED AS A RESULT OF THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22853710&pc=LWS | 2649622-2025-22743 | 2025-07-31 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-08-20 04:00 | LWS | SPRINT QUATTRO SECURE S MRI SURESCAN | Over-Sensing; Under-Sensing | Insufficient Information | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE RIGHT VENTRICULAR (RV) LEAD EXHIBITED UNDERSENSING AND OVERSENSING DURING THE VENTRICULAR ARRHYTHMIA AND THE PATIENT'S ARRYTHMIA WENT IN AND OUT OF THE PROGRAMMED DETECTION PARAMETERS. THE PATIENT DID RECEIVE ANTI-TACHYCARDIA PACING (ATP) AND HIGH VOLTAGE THERAPY, BUT SUBSEQUENTLY EXPIRED. THE CAUSE OF DEATH WAS NOT PROVIDED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22854758&pc=DSQ | 2916596-2025-05350 | 2025-07-09 04:00 | Death | THORATEC CORPORATION | 2025-08-20 04:00 | DSQ | HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM | Pumping Stopped | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT EITHER PASSED AWAY OR HAS HAD A HEART TRANSPLANT. THE VENTRICULAR ASSIST DEVICE (VAD) COORDINATOR REQUESTED ANALYSIS OF THE DEVICE LOG FILES. THE VAD COORDINATOR PROVIDED A SUMMARY OF THE LOG FILES. LOG FILES CAPTURED A BACKUP BATTERY INSTALLATION FAILURE ON (B)(6) 2025, AND A CONTROLLER RESET WHERE BOTH POWER LEADS WERE DISCONNECTED FROM THE CONTROLLER, LEADING TO PUMP STOPPAGE. THERE WERE NO NOTICEABLE ALARMS OR PARAMETER CHANGES PRIOR TO THESE EVENTS. FOLLOWING REVIEW OF THE LOG FILES BY TECH SERVICES, IT APPEARED THAT THE MCS EQUIPMENT WAS OPERATING AS INTENDED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22855120&pc=NPT | 2015691-2025-06859 | 2025-08-31 04:00 | Death | EDWARDS LIFESCIENCES | 2025-08-20 04:00 | NPT | EDWARDS COMMANDER DELIVERY SYSTEM | Difficult or Delayed Positioning; Difficult to Remove |  | P140031 |  | 1 | Event Description: AS REPORTED BY AN EDWARDS LIFESCIENCES FIELD CLINICAL SPECIALIST, REGARDING A 29MM SAPIEN 3 ULTRA RESILIA VALVE IN THE AORTIC POSITION. THERE WAS DIFFICULTY NAVIGATING THE 29MM SAPIEN 3 ULTRA RESILIA VALVE AND 29MM COMMANDER DELIVERY SYSTEM AROUND THE MULTIPLE BENDS IN THE AORTA. DURING WITHDRAWAL, THE VALVE WOULD NOT COME BACK INTO THE 16F ESHEATH PLUS AND VASCULAR SURGERY WAS CALLED TO REMOVE THE SYSTEM. THE CASE WAS ABORTED DUE TO A STABLE DESCENDING AORTIC DISSECTION. AT THIS TIME, IT COULD NOT BE DETERMINED IF THE TRACKING DIFFICULTY OR THE WITHDRAWAL DIFFICULTY CONTRIBUTED TO THE DISSECTION. PATIENT IS CURRENTLY ALIVE.Event Description: ADDITIONAL INFORMATION: THE PATIENT WENT TO THE OR TO HAVE VASCULAR SURGERY TO REMOVE THE VALVE/SHEATH/DELIVERY SYSTEM FROM HIS LEG, AND THE SURGEON APPARENTLY JUST PULLED IT UNDER FLOURO AND THE ILIAC CAME OUT WITH IT. THE PATIENT PASSED AWAY AFTER VASCULAR SURGERY ATTEMPT. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED DUE TO ENGINEERING EVALUATION FINDINGS. SECTIONS: A2, A3, B1, B2, B5, G3, G6, H1, H2, H6 IMPACT CODE HAVE BEEN UPDATED. Manufacturer Narrative: INVESTIGATION IS UNDERWAY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22856225&pc=GEX | 2124215-2025-57511 | 2012-04-01 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-08-20 04:00 | GEX | UNKNOWN | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: IT WAS REPORTED TO BOSTON SCIENTIFIC VIA AN ARTICLE PUBLISHED IN THE BJU INTERNATIONAL THAT A RETROSPECTIVE STUDY WAS CONDUCTED TO ANALYSE THE TREND IN SURGICAL INTERVENTIONS OF BENIGN PROSTATE ENLARGEMENT (BPE) IN THE UK NATIONAL HEALTH SERVICE (NHS) IN ENGLAND (PROSTATIC URETHRAL LIFT, TRANSURETHRAL WATER VAPOUR THERAPY (REZUM), PROSTATE ARTERY EMBOLISATION, PHOTOSELECTIVE VAPORISATION OF THE PROSTATE (GREENLIGHT), TRANSURETHRAL INCISION OF THE PROSTATE, HOLMIUM LASER ENUCLEATION OF THE PROSTATE, TRANSURETHRAL VAPORISATION OF THE PROSTATE, TRANSURETHRAL RESECTION OF THE PROSTATE (TURP), AND OPEN SURGERY) AND REPORT PREVALENCE OF PERI-PROCEDURAL COMPLICATIONS AND SUBSEQUENT RISK OF RE-INTERVENTION. A TOTAL OF 155874 PATIENTS OF BENIGN PROSTATE ENLARGEMENT WHO RECEIVED THEIR FIRST SURGICAL TREATMENT IN ANY NHS HOSPITAL IN ENGLAND BETWEEN APRIL 2012 AND NOVEMBER 2022 WERE INCLUDED IN THE STUDY. PATIENTS WERE FOLLOWED LONGITUDINALLY UNTIL THE END OF THE STUDY OR DATE OF DEATH. IN-HOSPITAL (COMPLICATIONS, LENGTH OF STAY), AND LONGER-TERM OUTCOMES (RE-ADMISSION, RE-TREATMENT, DEATH) WERE REPORTED. A MULTIVARIABLE LOGISTIC REGRESSION MODEL WAS BUILT TO ANALYSE THE RELATIONSHIP BETWEEN ASSOCIATED ODDS OF IN-HOSPITAL COMPLICATION AND SURGICAL MODALITY. KAPLAN-MEIER ANALYSIS (STRATIFIED BY AGE) AND MULTIVARIABLE COX REGRESSION ANALYSIS WAS PERFORMED TO INVESTIGATE THE RELATIONSHIP BETWEEN ASSOCIATED RISK OF RE-INTERVENTION BETWEEN TREATMENTS. LESS INVASIVE TREATMENTS WERE ASSOCIATED WITH SIGNIFICANTLY DECREASED ODDS OF IN-HOSPITAL COMPLICATION (COMPARED TO TURP). THE EFFECT OF COMORBIDITIES AND CATHETER ON ADMISSION VARIED BETWEEN TREATMENTS. RE-INTERVENTION RATE WITHIN 5 YEARS FOR INTERVENTIONS WHERE THERE WAS SUFFICIENT FOLLOW-UP VARIED FROM 3.1% TO 19.6%, LESS INVASIVE TREATMENTS WERE ASSOCIATED WITH A SIGNIFICANTLY HIGHER RISK OF RE-INTERVENTION COMPARED TO TURP. SOME OF THE IN-HOSPITAL COMPLICATIONS THAT WERE OBSERVED WERE URINARY RETENTION, HEMORRHAGE, HEMATOMA, HEMATURIA, URINARY TRACT INFECTION AND DEATHS ASSOCIATED WITH MYOCARDIAL INFARCTION, CHRONIC ISCHEMIC HEART DISEASE, MALIGNANT NEOPLASM OF THE BLADDER AND ATHEROSCLEROTIC HEART DISEASE. THERE IS AN INCREASED UPTAKE IN MINIMALLY INVASIVE TREATMENTS THAT REQUIRE A SHORTER HOSPITAL STAY AND ARE ASSOCIATED WITH LOWER ODDS OF TREATMENT-RELATED COMPLICATIONS. HOWEVER, THE LONG-TERM EFFECTIVENESS OF THESE TREATMENTS VARIED, WHICH SHOULD BE TAKEN INTO CONSIDERATION BY PATIENTS, CLINICIANS AND COMMISSIONERS. THIS REPORT IS FOR THE DEATHS. Manufacturer Narrative: DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BASED ON THE NATURE OF THE INFORMATION PROVIDED TO BSC, IT IS NOT POSSIBLE TO PERFORM A GOOD FAITH EFFORT TO OBTAIN ADDITIONAL INFORMATION. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER PRODUCT SPECIFIC INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PAGE, T., O'TOOLE, E., KELTIE, K., ABDELMOTAGLY, Y., RICHARDSON, H., PARKER, R., & SIMS, A. J. (2025). SURGERY FOR BENIGN PROSTATE ENLARGEMENT IN ENGLAND: 10-YEAR RETROSPECTIVE STUDY OF 155874 PATIENTS. BJU INTERNATIONAL, 136(1), 109-119. HTTPS://DOI.ORG/10.1111/BJU.16713.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22856443&pc=QEW | 3020347218-2025-00054 | 2025-07-22 04:00 | Death | INARI MEDICAL, INC. | 2025-08-20 04:00 | QEW | CLOTTRIEVER XL CATHETER | Adverse Event Without Identified Device or Use Problem | Pulmonary Embolism | K242557 |  | 1 | Event Description: A PATIENT WITH PULMONARY EMBOLISM AND THEIR INFERIOR VENA CAVA (IVC) COMPLETELY OCCLUDED WITH THROMBUS WAS ADMITTED TO THE HOSPITAL AND UNDERWENT A THROMBECTOMY. THERE WAS SOME CLOT IN THE LEFT PULMONARY ARTERY, BUT THE PHYSICIAN OPTED TO GO AFTER THE IVC CLOT FIRST. MULTIPLE ASPIRATIONS WERE PERFORMED UNSUCCESSFULLY. A PASS WAS DONE WITH THE CLOTTRIEVER XL CATHETER WHICH RETURNED A BAG FULL OF CHRONIC CLOT AND/OR TUMOR THROMBUS. HOWEVER, ONCE THE CATHETER PULLED, THE PATIENT IMMEDIATELY DETERIORATED. UNFORTUNATELY, THE PATIENT ULTIMATELY EXPIRED. Manufacturer Narrative: THE INARI DEVICES WERE DISCARDED BY THE USER FACILITY AND ARE NOT AVAILABLE FOR EVALUATION. THE DEVICE IDENTIFIERS WERE NOT PROVIDED, THE LOT HISTORY RECORDS OF ALL POTENTIAL LOTS SHIPPED TO THE FACILITY WERE REVIEWED. THERE WERE NO DISCREPANCIES OR UNUSUAL FINDINGS. THERE WAS NO REPORT OF ANY DEVICE MALFUNCTION. BASED ON THE INFORMATION REVIEWED, THERE IS NO EVIDENCE TO INDICATE THE PRESENCE OF A POTENTIAL QUALITY ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. THE CASE WAS REVIEWED BY INARI STRYKER MEDICAL AFFAIRS, WHO CONCLUDED THAT THE PATIENT'S DEATH WAS POTENTIALLY THE RESULT OF INADVERTENT CLOT EMBOLISM. DISTAL EMBOLIZATION OF BLOOD CLOTS IS IDENTIFIED IN THE DEVICE LABELING AS POSSIBLE ADVERSE EVENTS/COMPLICATIONS. THE DEVICE LABELING ALSO LISTS THE FOLLOWING WARNING: "OPERATION OF A THROMBECTOMY CATHETER MAY CAUSE EMBOLIZATION OF SOME THROMBUS AND/OR THROMBUS PARTICULATE, PHYSICIAN DISCRETION ADVISED. THE POTENTIAL FOR EXTENSIVE AND/OR DIFFICULT TO TREAT PULMONARY THROMBOEMBOLISM SHOULD BE CAREFULLY CONSIDERED WHEN CLOTTRIEVER XL CATHETER IS USED TO ENGAGE AND REMOVE THROMBUS FROM LARGE VESSELS SUCH AS THE INFERIOR VENA CAVA (IVC)." MANUFACTURER REFERENCE #:(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22856907&pc=FTL | 9615742-2025-00977 |  | Death | SOFRADIM PRODUCTION SAS | 2025-08-20 04:00 | FTL | PARIETEX | Adverse Event Without Identified Device or Use Problem | Distress; Unspecified Tissue Injury |  |  | 1 | Event Description: THE PATIENT'S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR THERAPEUTIC TREATMENT OF A HERNIA. IT WAS REPORTED THAT AFTER IMPLANT, THE PATIENT EXPERIENCED EMOTIONAL DISTRESS, DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED REVISION SURGERY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22857137&pc=FTL | 9615742-2025-00995 |  | Death | SOFRADIM PRODUCTION SAS | 2025-08-20 04:00 | FTL | PROGRIP | Adverse Event Without Identified Device or Use Problem | Distress; Unspecified Tissue Injury |  |  | 1 | Event Description: THE PATIENT'S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR THERAPEUTIC TREATMENT OF A HERNIA. IT WAS REPORTED THAT AFTER IMPLANT, THE PATIENT EXPERIENCED EMOTIONAL DISTRESS, DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED REVISION SURGERY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22857140&pc=FTL | 9615742-2025-00990 |  | Death | SOFRADIM PRODUCTION SAS | 2025-08-20 04:00 | FTL | SYMBOTEX | Adverse Event Without Identified Device or Use Problem | Unspecified Tissue Injury |  |  | 1 | Event Description: THE PATIENT'S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR THERAPEUTIC TREATMENT OF A HERNIA. IT WAS REPORTED THAT AFTER IMPLANT; THE PATIENT EXPERIENCED EMOTIONAL DISTRESS AND DEATH. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22857415&pc=BTT | 9611451-2025-00793 | 2025-07-18 04:00 | Death | FISHER & PAYKEL HEALTHCARE LTD | 2025-08-20 04:00 | BTT | FISHER & PAYKEL HEALTHCARE | Use of Device Problem | Insufficient Information |  |  | 1 | Event Description: A HEALTHCARE FACILITY IN HONG KONG REPORTED THAT ON (B)(6) 2025, BETWEEN APPROXIMATELY 10AM AND 11AM, THE POWER PLUG OF A PT101 AIRVO 2 HUMIDIFIER WAS INADVERTENTLY DISCONNECTED FROM THE HEALTHCARE FACILITY'S WALL SOCKET DUE TO A NURSING ERROR. THE HEALTHCARE FACILITY REPORTED THAT THE PATIENT SUBSEQUENTLY DECEASED. Manufacturer Narrative: (B)(4). D1, G4: PT101UK IS NOT SOLD IN THE USA BUT IS SIMILAR TO THE PT101US WHICH IS SOLD IN THE USA. THE 510K OF THE PT101US IS: K131895, PRODUCT CODE: BTT. PRODUCT BACKGROUND: THE PT101 AIRVO 2 HUMIDIFIER (AIRVO 2) DEVICE IS AN ACTIVE HUMIDIFIER WITH INTEGRATED FLOW GENERATOR THAT DELIVERS HIGH FLOW WARMED AND HUMIDIFIED RESPIRATORY GASES TO SPONTANEOUSLY BREATHING PATIENTS THROUGH A VARIETY OF PATIENT INTERFACES. ITS INTENDED USE IS FOR THE TREATMENT OF SPONTANEOUSLY BREATHING PATIENTS WHO WOULD BENEFIT FROM RECEIVING HIGH FLOW WARMED AND HUMIDIFIED RESPIRATORY GASES. THE AIRVO 2 DEVICE SHOULD NOT BE USED FOR LIFE SUPPORT PURPOSES, AND APPROPRIATE PATIENT MONITORING MUST BE USED AT ALL TIMES. METHOD: THE SUBJECT AIRVO 2 DEVICE WAS NOT RETURNED TO F&P FOR EVALUATION. OUR INVESTIGATION HAS NECESSARILY BEEN LIMITED THEREFORE TO THE INFORMATION REPORTED AND PROVIDED BY THE HEALTHCARE FACILITY AND OUR KNOWLEDGE OF THE PRODUCT. RESULTS: THE HEALTHCARE FACILITY REPORTED THAT THE POWER PLUG OF THE SUBJECT DEVICE WAS INADVERTENTLY DISCONNECTED FROM THE HEALTHCARE FACILITY'S WALL SOCKET DUE TO A NURSING ERROR. IT WAS REPORTED BY THE HEALTHCARE FACILITY THAT A FUNCTIONALITY TEST WAS PERFORMED ON THE SUBJECT DEVICE AT THE HEALTHCARE FACILITY BY A TRAINED TECHNICIAN ON THE 21ST OF JULY 2025, AND THAT THE SUBJECT DEVICE PASSED ALL F&P PERFORMANCE CHECKS AND WAS CONFIRMED TO BE OPERATIONAL. F&P MADE REQUESTS FOR DEVICE AND EVENT DETAILS TO THE HEALTHCARE FACILITY, HOWEVER THE HEALTHCARE FACILITY STATED THAT THE FUNCTIONALITY OF THE DEVICE WAS CONFIRMED TO BE OPERATIONAL AND DID NOT PROVIDE INFORMATION IN RESPONSE TO THESE REQUESTS NOR AGREE TO RETURN THE SUBJECT DEVICE TO F&P FOR EVALUATION. THE DEVICE LOGS WERE RETRIEVED FROM THE SUBJECT DEVICE AND INSPECTED BY A TRAINED F&P INVESTIGATION ENGINEER. THE DEVICE LOG FILES RECORDED THAT THERAPY DELIVERY ASSOCIATED WITH THE REPORTED EVENT BEGAN ON THE (B)(6) 2025. THE DATA LOGS RECORDED THAT A SHUTDOWN OF THE SUBJECT DEVICE WAS RECORDED ON THE (B)(6) 2025, BETWEEN APPROXIMATELY 10:57AM AND 11:02AM. THE DATA LOGS RECORDED THAT THE SUBJECT DEVICE WAS POWERED ON AGAIN ON THE (B)(6) 2025 AT APPROXIMATELY 11:22AM. THE REVIEW OF THE DATA LOGS SHOWED THERE WERE NO FAULTS OR ERRORS RECORDED IN THE SUBJECT DEVICE NEAR THE TIME OF THE REPORTED EVENT. CONCLUSION: F&P'S INVESTIGATION DID NOT IDENTIFY ANY FAULTS OR ERRORS WITH THE SUBJECT DEVICE. THE HEALTHCARE FACILITY REPORTED THAT THE POWER PLUG OF THE SUBJECT DEVICE WAS INADVERTENTLY DISCONNECTED FROM THE HEALTHCARE FACILITY'S WALL SOCKET DUE TO A NURSING ERROR. THE HEALTHCARE FACILITY REPORTED THAT A FUNCTIONALITY TEST WAS PERFORMED ON THE SUBJECT DEVICE AT THE HEALTHCARE FACILITY BY A TRAINED TECHNICIAN, AND THAT THE SUBJECT DEVICE PASSED ALL F&P PERFORMANCE CHECKS AND WAS CONFIRMED TO BE OPERATIONAL. F&P MADE REQUESTS FOR DEVICE AND EVENT DETAILS, HOWEVER THE HEALTHCARE FACILITY STATED THAT THE FUNCTIONALITY OF THE DEVICE WAS CONFIRMED TO BE OPERATIONAL AND DID NOT PROVIDE INFORMATION IN RESPONSE TO THESE REQUESTS NOR AGREE TO RETURN THE SUBJECT DEVICE TO F&P FOR EVALUATION. THE USER INSTRUCTION (UI) OF THE PT101 AIRVO 2 HUMIDIFIER STATES: - "APPROPRIATE PATIENT MONITORING MUST BE USED AT ALL TIMES. LOSS OF POWER MEANS LOSS OF THERAPY." - "THE UNIT IS NOT INTENDED FOR LIFE SUPPORT." - "USE CONTINUOUS OXYGEN MONITORING ON PATIENTS WHO WOULD DESATURATE SIGNIFICANTLY IN THE EVENT OF DISRUPTION TO THEIR OXYGEN SUPPLY."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22836714&pc=LJT | 3006260740-2025-05932 | 2025-07-28 04:00 | Death | C.R. BARD, INC. (BASD) -3006260740 | 2025-08-19 04:00 | LJT | LOW ARTIFACT POWER PORT | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Thrombosis/Thrombus; Insufficient Information | UNKNOWN |  | 1 | Event Description: IT WAS REPORTED THROUGH THE LITIGATION PROCESS THAT SOMETIME AFTER PORT PLACEMENT PROCEDURE, THE PATIENT ALLEGEDLY DEVELOPED WITH UNSPECIFIED INFECTION AND THROMBOSIS. REPORTEDLY, THE PATIENT EXPIRED, AND THE CAUSE OF DEATH WAS NOT PROVIDED. Manufacturer Narrative: H11: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. DATE OF DEATH FOR THIS PATIENT WAS NOT PROVIDED, DATE OF DEATH UPDATED AS (B)(6) 1900 FOR THE MDR SUBMISSION REQUIREMENT. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22836766&pc=OZP | 2032227-2025-238929 | 2025-07-09 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-08-19 04:00 | OZP | PUMP MMT-1880L MM770G BLE MG | Adverse Event Without Identified Device or Use Problem | Hyperglycemia; Hypoglycemia; Renal Failure | P160017 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY DUE TO KIDNEY FAILURE AND DIABETES. THE CUSTOMER PASSED AWAY ON (B)(6) 2024. THE CUSTOMER'S PUMP WAS WORN AT THE TIME OF EVENT. THE CUSTOMER BLOOD GLUCOSE VALUE WAS UNKNOWN. THE EVENT INVOLVED PRODUCT MMT-332A, MMT-397A, MMT-1880L. TROUBLESHOOTING WAS PERFORMED FOR GATHERING THE DECEASED RELATED INFORMATION. CUSTOMER WAS USING THE INSULIN PUMP WITHIN 48 HOURS OF THE REPORTED EVENT. IT WAS UNKNOWN WHETHER THE AUTO MODE FEATURE WAS ACTIVE AT THE TIME OF INCIDENT. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED. NO PRODUCT RETURN WAS REQUIRED FOR MMT-332A, MMT-397A. MMT-1880L WAS REQUESTED AND CUSTOMER RESPONSE WAS THE DEVICE WILL BE RETURNED. THE CUSTOMER WILL DISCONTINUE THE USE OF THE DEVICE. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. THE DEVICE WILL BE RETURNED FOR ANALYSIS AND FURTHER INFORMATION WILL FOLLOW ONCE THE ANALYSIS HAS BEEN COMPLETED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22836919&pc=OZO | 2032227-2025-238988 | 2025-07-17 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-08-19 04:00 | OZO | 630G INSULIN PUMP MMT-1715KL 630G | Adverse Event Without Identified Device or Use Problem | Hyperglycemia; Hypoglycemia; Renal Failure; Cancer | P150001 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER'S NEXT OF KIN INFORMED THAT THE CUSTOMER HAD PASSED AWAY ON (B)(6) 2025, IN THE HOSPITAL. THE REASON FOR ADMISSION TO THE HOSPITAL WAS AN INFECTION DUE TO PERITONEAL DIALYSIS (PD). THE REPORTED CAUSES OF DEATH INCLUDED KIDNEY FAILURE, HEART ISSUES, PROSTATE CANCER, AND DIABETES. THE BLOOD GLUCOSE VALUE AT THE TIME OF HOSPITALIZATION WAS 895 MG/DL. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: MMT-1715KL, MMT-397A, AND MMT-332A. TROUBLESHOOTING WAS PERFORMED. THE PUMP WAS NOT BEING WORN AT THE TIME OF PASSING. IT WAS UNKNOWN WHETHER THE CUSTOMER HAD BEEN USING THE INSULIN PUMP WITHIN 48 HOURS OF THE REPORTED EVENT AND UNKNOWN WHETHER THE AUTO MODE/SMARTGUARD FEATURE WAS ACTIVE AT THE TIME OF THE EVENT. NO PRODUCT RETURN IS REQUIRED FOR MMT-397A. NO PRODUCT RETURN IS REQUIRED FOR MMT-1715KL. NO PRODUCT RETURN IS REQUIRED FOR MMT-332A. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22838039&pc=OZO | 2032227-2025-239490 | 2023-08-30 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-08-19 04:00 | OZO | 630G INSULIN PUMP MMT-1715KL 630G | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Hyperglycemia; Hypoglycemia; Lethargy | P150001 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2023, AT THE HOSPITAL. THE CAUSE OF DEATH WAS A MAJOR STROKE AND OTHER CONTRIBUTING FACTORS TO THE CAUSE OF DEATH WERE DIABETES AND THE KIDNEY TRANSPLANT IN THE PAST. THE EMERGENCY MEDICAL SERVICE (AMBULANCE) WAS CALLED UPON AND WAS HOSPITALIZED DUE TO LETHARGY. THE INSULIN PUMP WAS WORN AT THE TIME OF PASSING. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: MMT-332A, MMT-1715KL, AND UNOMEDICAL. TROUBLESHOOTING WAS PERFORMED FOR DECEASED REPORTING. IT WAS UNKNOWN WHETHER THE CUSTOMER HAD USED THE INSULIN PUMP SYSTEM WITHIN 48 HOURS OF THE REPORTED EVENT. THERE IS NO MENTION OF THE AUTO MODE FEATURE AT THE TIME OF THE EVENT. NO PRODUCT RETURN IS REQUIRED FOR MMT-332A, UNOMEDICAL, MMT-1715KL. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22840101&pc=NAY | 2955842-2025-34694 | 2025-07-23 04:00 | Death | INTUITIVE SURGICAL, INC | 2025-08-19 04:00 | NAY | DAVINCI XI | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Hemorrhage/Bleeding | K131861 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED DIAPHRAGMATIC HERNIA REPAIR, SIGNIFICANT BLEEDING OCCURRED WHILE THE SURGEON WAS PERFORMING ADHESIOLYSIS, AND THE PATIENT ULTIMATELY EXPIRED. THE SITE ROBOTICS COORDINATOR REPORTED THAT THE SURGEON ATTEMPTED TO CONTROL THE BLEEDING USING A DA VINCI GRASPER INSTRUMENT TO CLAMP THE SUSPECTED SOURCE. DUE TO THE SEVERITY OF THE BLEEDING, THE PROCEDURE WAS CONVERTED TO AN OPEN APPROACH. DURING THE CONVERSION, THE PATIENT CODED, AND RESUSCITATION EFFORTS WERE INITIATED; HOWEVER, THE PATIENT COULD NOT BE REVIVED. THE PATIENT¿S ANATOMY WAS REPORTED TO HAVE SIGNIFICANT ADHESIONS, AND THE SURGEON SPECULATED THAT THE BLEEDING MAY HAVE RESULTED FROM A TEAR IN THE INFERIOR VENA CAVA CAUSED DURING ADHESIOLYSIS. IT WAS REPORTED THAT THE SURGEON DOES NOT BELIEVE THAT THE DA VINCI SYSTEM, INSTRUMENTS, OR ACCESSORIES CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. MULTIPLE ATTEMPTS TO SPEAK WITH THE SURGEON WERE MADE; HOWEVER, NO RESPONSE HAS BEEN RECEIVED. Manufacturer Narrative: A REVIEW OF THE SYSTEM LOGS FOR THE REPORTED PROCEDURE DATE FOUND THAT THE SYSTEM HAD ONE RECOVERABLE ERROR ON UNIVERSAL SURGICAL MANIPULATOR 4 DURING THE PROCEDURE THAT WAS RECOVERED BY THE USER AND NO FURTHER ISSUES WERE EXPERIENCED DURING THE PROCEDURE. LOG REVIEW OF THE ONE SUBSEQUENT PROCEDURE PERFORMED ON THE SYSTEM FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: 30 DEGREE ENDOSCOPE PLUS, MONOPOLAR CURVED SCISSORS, SMALL GRASPING RETRACTOR, CADIERE FORCEPS, PROGRASP FORCEPS, SUCTION IRRIGATOR, VESSEL SEALER EXTEND. A SITE HISTORY REVIEW FOUND NO COMPLAINTS WERE MADE FOR THE INSTRUMENTS USED DURING THIS PROCEDURE. A DEVICE HISTORY RECORD (DHR) REVIEW FOUND NO NON-CONFORMANCES DOCUMENTED THAT WOULD BE RELATED TO THIS EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT DIED DUE TO BLEEDING FROM AN INJURY TO THE VENA CAVA DURING A DIFFICULT DIAPHRAGMATIC HERNIA REPAIR WITH SIGNIFICANT ADHESIONS. THE SURGEON NOTED THE ANATOMY WAS DIFFICULT DUE TO THESE ADHESIONS WHICH CONTRIBUTED TO THE INJURY. THERE WAS NO ALLEGATION OF ANY ISSUES WITH ANY INTUITIVE SURGICAL PRODUCT OR INSTRUMENT. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, NO INTUITIVE SURGICAL PRODUCT OR INSTRUMENT CONTRIBUTED TO THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22840110&pc=NTE | 2124215-2025-56383 | 2015-12-01 05:00 | Death | SILK ROAD MEDICAL | 2025-08-19 04:00 | NTE | ENROUTE TRANSCAROTID NEUROPROTECTION SYSTEM | Adverse Event Without Identified Device or Use Problem | Hemorrhagic Stroke; Heart Failure/Congestive Heart Failure; Insufficient Information |  |  | 1 | Event Description: THE STUDY COMPARED PERIOPERATIVE AND POSTOPERATIVE COMPLICATIONS OF TRANSCAROTID ARTERY REVASCULARIZATION (TCAR) PROCEDURES PERFORMED BETWEEN DECEMBER 2015 AND APRIL 2019 BETWEEN PATIENTS WITH VIRGIN NECKS AND THOSE WITH HOSTILE NECKS, DEFINED AS HAVING UNDERGONE PRIOR NECK RADIATION OR SURGERY. ALL TCAR PROCEDURES PERFORMED AT AFFILIATED FACILITIES BETWEEN DECEMBER 2015 AND APRIL 2019 WERE CAPTURED. SEVENTY OF THE PROCEDURES WERE PERFORMED AS PART OF THE ONGOING ROADSTER 2 REGISTRY AND WERE EXCLUDED, LEAVING 149 TOTAL INTERVENTIONS FOR ANALYSIS. OF THE REMAINING TCARS, 30 WERE PERFORMED IN PATIENTS WITH A HISTORY OF PREVIOUS IPSILATERAL NECK DISSECTION OR RADIATION TREATMENT. AMONG THESE PATIENTS, 10 HAD UNDERGONE PRIOR NECK RADIATION, 23 HAD EXPERIENCED PREVIOUS SURGICAL DISSECTION (3 ONCOLOGIC AND 20 CAROTID ENDARTERECTOMY), AND 3 HAD UNDERGONE BOTH RADIATION AND SURGICAL DISSECTION. THE REMAINING 119 PATIENTS DID NOT REPORT ANY PRIOR NECK INTERVENTIONS BEFORE THE INDEX TCAR. NO DIFFERENCES WERE OBSERVED IN BASELINE COMORBIDITIES OR MEDICATION USAGE BETWEEN THE TWO COHORTS. PATIENTS INVOLVED IN THE STUDY HAD AN IPSILATERAL TRANSIENT ISCHEMIC ATTACK (TIA) OR IPSILATERAL STROKE WITHIN THE PREVIOUS 180 DAYS IN THE PRESENCE OF A STENOSIS GREATER THAN 50% OR WERE ASYMPTOMATIC WITH STENOSIS GREATER THAN 80% ON CAROTID DUPLEX OR COMPUTED TOMOGRAPHY ANGIOGRAM. IN THIS STUDY, PATIENTS IN THE VIRGIN FIELD GROUP HAD A MEAN AGE OF 70.4, WITH A STANDARD DEVIATION OF 9.7 YEARS, WHILE THOSE IN THE HOSTILE FIELD GROUP HAD A MEAN AGE OF 69.0, WITH A STANDARD DEVIATION OF 8.7 YEARS. WOMEN COMPRISED 15.1% OF THE VIRGIN FIELD GROUP AND 30.0% OF THE HOSTILE FIELD GROUP. WITHIN 30 DAYS POST-PROCEDURE, THE STUDY NOTED THAT THE RATE OF MORTALITY WAS 2.5% IN VIRGIN NECKS AND NONE IN HOSTILE NECKS. ONE PASSED AWAY ON POSTOPERATIVE DAY (POD) 0 DUE TO A MASSIVE HEMORRHAGIC STROKE, ANOTHER ON POD 28 FROM WORSENING CONGESTIVE HEART FAILURE, AND THE THIRD ON POD 23 AT AN EXTERNAL HOSPITAL FROM UNDETERMINED CAUSES. FOUR UNRELATED DEATHS WERE ALSO NOTED DURING THE FOLLOW-UP PERIOD AVERAGED 7.5 MONTHS FOR VIRGIN NECKS AND 10.4 MONTHS FOR HOSTILE NECKS. Manufacturer Narrative: A2: THE MEAN AGE OF 70.4 YEARS, WITH A STANDARD DEVIATION OF 9.7 YEARS, WHILE THOSE IN THE HOSTILE FIELD GROUP HAD A MEAN AGE OF 69.0, WITH A STANDARD DEVIATION OF 8.7 YEARS. A3B: WOMEN COMPRISED 15.1% OF THE VIRGIN FIELD GROUP AND 30.0% OF THE HOSTILE FIELD GROUP. B2. THE DATE OF DEATH WAS NOT PROVIDED FOR ANY OF THE REPORTED DEATHS. B3: THE EXACT EVENT DATE WAS NOT PROVIDED IN THE LITERATURE ARTICLE. THE DATE OF EVENT WAS ESTIMATED USING THE BEGINNING OF THE STUDY PERIOD DATE RANGE, WHICH COVERED DECEMBER 2015 TO APRIL 2019. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BOSTON SCIENTIFIC. BASED ON THE NATURE OF THE INFORMATION RECEIVED, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE DEVICE IDENTIFIER (UDI) OR OTHER PRODUCT-SPECIFIC DETAILS. IF ADDITIONAL INFORMATION BECOMES AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. WANG, S. K., MILADORE, J. N., DRUCKER, N. A., FAJARDO, A., LEMMON, G. W., DALSING, M. C., MURPHY, M. P., & MOTAGANAHALLI, R. L. (2019). PROCEDURAL COMPLEXITY OF TRANSCAROTID ARTERY REVASCULARIZATION IS NOT INCREASED IN IRRADIATED AND REOPERATIVE NECKS. ANNALS OF VASCULAR SURGERY, 61, 212-217. HTTPS://DOI.ORG/10.1016/J.AVSG.2019.05.063.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22840115&pc=DSQ | 2916596-2025-05437 | 2025-08-06 04:00 | Death | THORATEC CORPORATION | 2025-08-19 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Heart Failure/Congestive Heart Failure; Skin Infection | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD A DRIVELINE INFECTION AND BLOOD CULTURES POSITIVE FOR PSEUDOMONAS AERUGINOSA. THE ONSET DATE OF THE PATIENT'S INFECTION WAS (B)(6) 2022 (MFR #2916596-2025-05510). THE SOURCE OF THE INFECTION WAS THE DRIVELINE. THE PATIENT WAS EXPLANTED (B)(6) 2025 AND PASSED AWAY ON (B)(6) 2025 SECONDARY TO POST-OP COMPLICATIONS. THE CAUSE OF DEATH WAS RIGHT VENTRICULAR FAILURE AND INFECTION. IT WAS REPORTED THAT THE PATIENT'S OUTCOME WAS NOT DEVICE OR THERAPY RELATED. THE DEVICE WAS EXPLANTED BUT WILL NOT BE RETURNED FOR EVALUATION. THE DEVICE OPERATED AS EXPECTED. RIGHT HEART FAILURE DID NOT EXIST PRE-LVAD IMPLANT, BUT IT WAS STATED THAT A DEVICE RELATED ISSUE DID NOT CAUSE OR CONTRIBUTE TO THE RIGHT HEART FAILURE. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. A SPECIFIC CAUSE FOR THE PATIENT¿S INFECTION OR REASON FOR DEVICE EXPLANT COULD NOT BE CONCLUSIVELY DETERMINED. IT WAS REPORTED THAT HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), WILL NOT BE RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. D, IS CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ OF THIS DOCUMENT LISTS RIGHT HEART FAILURE, INFECTION, DRIVELINE INFECTION, AND DEATH AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. ADDITIONALLY, SECTION 6 ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿RIGHT HEART FAILURE¿) ALSO OUTLINES INDICATIONS OF RIGHT HEART FAILURE AS WELL AS POSSIBLE TREATMENTS. THIS SECTION INCLUDES INFORMATION FOR CARING FOR THE DRIVELINE EXIT SITE AS WELL AS INFORMATION REGARDING HOW TO PREVENT AND CONTROL INFECTION. THE HEARTMATE 3 LVAS PATIENT HANDBOOK, REV. A, CONTAINS SEVERAL SECTIONS WITH INFORMATION ABOUT HOW TO PREVENT AND CONTROL INFECTION AS WELL AS INFORMATION FOR CARING FOR THE DRIVELINE EXIT SITE. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22840124&pc=MKJ | 1220908-2025-03012 | 2025-07-21 04:00 | Death | ZOLL MEDICAL CORPORATION | 2025-08-19 04:00 | MKJ | X SERIES | Charging Problem | Insufficient Information | P160022 |  | 1 | Event Description: COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO MONITOR A PATIENT (AGE & GENDER UNKNOWN), THE DEVICE EXPERIENCED A DEFIB CHARGE TIMEOUT. COMPLAINANT INDICATED THAT THE PATIENT SUBSEQUENTLY EXPIRED. Manufacturer Narrative: ZOLL MEDICAL CORPORATION HAS NOT RECEIVED THE DEVICE FOR EVALUATION AND THIS COMPLAINT IS STILL UNDER INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22840260&pc=DSQ | 2916596-2025-05486 | 2025-08-03 04:00 | Death | THORATEC CORPORATION | 2025-08-19 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. IT WAS REPORTED THAT THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. THE DEVICE WAS NOT EXPLANTED AND WOULD NOT BE RETURNED FOR EVALUATION. THE PATIENT'S DEATH WAS RELATED TO MULTI-ORGAN FAILURE. THE DEVICE WAS WORKING AS INTENDED DURING TIME OF SUPPORT. AN AUTOPSY WAS NOT PERFORMED. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT'S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICLE ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REVISION, REV. B IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION AND DEATH AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22840346&pc=CBK | 3013421741-2025-01672 | 2025-07-30 04:00 | Death | VYAIRE MEDICAL, INC | 2025-08-19 04:00 | CBK | BELLAVISTA | Inaccurate Flow Rate | Necrosis | K183364 |  | 1 | Event Description: IT WAS REPORTED THAT WHILE ATTEMPTING TO TREAT A PATIENT (AGE & GENDER UNKNOWN), THE DEVICE ALARM OCCURRED INDICATING A FIO2 DOSING ISSUE. COMPLAINANT INDICATED THAT THE PATIENT SUBSEQUENTLY EXPIRED. Manufacturer Narrative: ZOLL MEDICAL CORPORATION HAS NOT RECEIVED THE DEVICE FOR EVALUATION AND THIS COMPLAINT IS STILL UNDER INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22840359&pc=NIQ | 2024168-2025-08683 | 2018-06-01 04:00 | Death | ABBOTT VASCULAR | 2025-08-19 04:00 | NIQ | XIENCE PRIME BTK EVEROLIMUS ELUTING PERIPHERAL STENT SYSTEM | Adverse Event Without Identified Device or Use Problem | Unspecified Heart Problem | P110019 |  | 1 | Event Description: THE XIENCE PRIME STUDY INTRODUCES THE POLYMER EVEROLIMUS-ELUTING STENT (PEES) AS AN EFFECTIVE BAILOUT STENTING OPTION FOR PATIENTS WITH CHRONIC LIMB-THREATENING ISCHEMIA (CLTI) UNDERGOING BELOW-THE-KNEE (BTK) REVASCULARIZATION, PARTICULARLY FOR LESIONS UNDER 4 CM. THE STUDY FOCUSES ON REAL-WORLD CASES WHERE PERCUTANEOUS OLD BALLOON ANGIOPLASTY (POBA) ALONE IS INSUFFICIENT, DEMONSTRATING THAT PEES SIGNIFICANTLY IMPROVES OUTCOMES BY ENHANCING LIMB SALVAGE AND REDUCING THE NEED FOR MAJOR AMPUTATIONS. BETWEEN MAY 2018 AND JUNE 2019, 106 PATIENTS UNDERWENT STENTING WITH XIENCE PRIME BTK. THE FOLLOW-UP WAS ACCOMPLISHED AT 1, 6, AND 12 MONTHS. ADVERSE EVENTS INCLUDED DEATH, MYOCARDIAL INFARCTION, STROKE, MAJOR ADVERSE LIMB EVENT (AMPUTATION AND TARGET LESION REVASCULARIZATION), SURGICAL REVASCULARIZATION, RESTENOSIS AND THROMBOSIS. THE ARTICLE CONCLUDED THE USE OF THE PEES DEVICE YIELDED SATISFACTORY RESULTS IN TERMS OF PATENCY, TLR, AND LIMB SALVAGE AT 1 YEAR, REDUCING EXPOSURE TO A MAJOR EVENT IN THE TREATED LIMB. DETAILS ARE LISTED IN THE ATTACHED ARTICLE, TITLED " POLYMER EVEROLIMUS-ELUTING STENT AS BAILOUT STENTING FOR BELOW-THE-KNEE ARTERY REPAIR IN PATIENTS WITH CRITICAL LIMB-THREATENING ISCHEMIA: A REAL-WORLD NATIONAL REGISTRY.¿ NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORDS FOR THIS PRODUCT COULD NOT BE REVIEWED AND A SIMILAR COMPLAINT QUERY COULD NOT BE PERFORMED BECAUSE THE PART NUMBER AND/OR LOT NUMBER WAS NOT REPORTED, AND THE DEVICE WAS NOT RETURNED. THE REPORTED PATIENT EFFECT OF DEATH IS LISTED IN THE XIENCE PRIME, XIENCE PRIME SMALL VESSEL (SV), AND XIENCE PRIME LONG LENGTH (LL), EVEROLIMUS ELUTING CORONARY STENT SYSTEMS INSTRUCTIONS FOR USE AS A KNOWN PATIENT EFFECT OF CORONARY STENTING PROCEDURES. A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECT, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING. LITERATURE ATTACHMENT: ARTICLE TITLE: ¿POLYMER EVEROLIMUS-ELUTING STENT AS BAILOUT STENTING FOR BELOW-THE-KNEE ARTERY REPAIR IN PATIENTS WITH CRITICAL LIMB-THREATENING ISCHEMIA: A REAL-WORLD NATIONAL REGISTRY¿ B2: DATE OF DEATH ESTIMATED AS (B)(6) 2018 B3: DATE OF PROCEDURE ESTIMATED AS (B)(6) 2018 D4: THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. D6A: DATE OF IMPLANT ESTIMATED AS (B)(6) 2018. THE ADDITIONAL PATIENT EFFECTS REPORTED IN THE ARTICLE ARE CAPTURED UNDER A SEPARATE MEDWATCH REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22840360&pc=OZD | 1220648-2025-45824 | 2025-07-26 04:00 | Death | ABIOMED, INC. | 2025-08-19 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Tachycardia; Ventricular Fibrillation | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT DURING IMPLANTATION OF AN IMPELLA CP, THE PATIENT DEVELOPED VENTRICULAR TACHYCARDIA WHICH QUICKLY DETERIORATED TO VENTRICULAR FIBRILLATION. THE PATIENT WAS SUCCESSFULLY DEFIBRILLATED. LATER, CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22840701&pc=OZD | 1220648-2025-31164 | 2025-07-22 04:00 | Death | ABIOMED, INC. | 2025-08-19 04:00 | OZD | IMPELLA CP | Fluid/Blood Leak; Increase in Pressure | Insufficient Information | P140003 |  | 1 | Event Description: THE COMPLAINANT HAD AN IMPELLA CP PLACED TO REPLACE A PRIOR IMPELLA CP THAT WAS EXPLANTED WITH CLOT CONCERNS. THIS PUMP WAS SUPPORTING A CARDIOGENIC SHOCK/CARDIOMYOPATHY PATIENT WHEN THERE WAS A PURGE LEAK AND HIGH PURGE PRESSURE/LOW PURGE FLOW ALARMS. THE PUMP REMAINED ON DESPITE THE PURGE ISSUES FOR ANOTHER TWO DAYS WHEN THE PATIENT HAD CARE WITHDRAWN AFTER JUST UNDER NINE DAYS. THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH RISK PERCUTANEOUS CORONARY INTERVENTION SECTION: CLEANING ¿DO NOT CLEAN WITH OR EXPOSE ANY PART OF THE CLEAR SIDEARM OF THE IMPELLA CATHETER (E.G. INFUSION FILTER, PRESSURE RESERVOIR) TO ALCOHOL. ALCOHOL HAS BEEN SHOWN TO CAUSE CRACKS AND LEAKS IN THESE COMPONENTS. CAREFULLY READ LABELS ON COMMON SKIN PREPS AND LOTIONS TO AVOID USING ANY ALCOHOL-CONTAINING PRODUCTS IN THE AREA OF THE INFUSION FILTER OR PRESSURE RESERVOIR." ¿CLEAN THE CONNECTOR CABLE WITH 70% ISOPROPYL ALCOHOL.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22840965&pc=NIQ | 2024168-2025-08688 | 2023-01-04 05:00 | Death | ABBOTT VASCULAR | 2025-08-19 04:00 | NIQ | XIENCE EVEROLIMUS ELUTING CORONARY STENT SYSTEM | Adverse Event Without Identified Device or Use Problem | Unspecified Heart Problem | P070015 |  | 1 | Event Description: IT WAS REPORTED IN A SUMMARY ARTICLE THE IMPACT OF THE LEVEL OF TORTUOSITY OF THE CORONARY ARTERY OUTCOMES FOLLOWING STENTING WITH DRUG-ELUTING STENTS (DES). THE XIENCE DES AND A NON-ABBOTT DES WERE IN THE BIOFLOW II, IV AND V RANDOMIZED TRIALS. IN POST PROCEDURE 3 YEAR FOLLOW-UP ADVERSE PATIENT EFFECT WERE NOTED TO BE MYOCARDIAL INFARCTION, ISCHEMIA, THROMBOSIS AND DEATH. ADDITIONAL INFORMATION CAN BE FOUND IN THE SUMMARY ARTICLE, "IMPACT OF CORONARY ARTERY TORTUOSITY ON OUTCOMES FOLLOWING STENTING WITH NEWER-GENERATION DRUG-ELUTING STENTS. AN ANALYSIS OF THE RANDOMIZED BIOFLOW TRIALS". NO OTHER INFORMATION WAS PROVIDED. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORDS FOR THIS PRODUCT COULD NOT BE REVIEWED AND A SIMILAR COMPLAINT QUERY COULD NOT BE PERFORMED BECAUSE THE PART NUMBER AND/OR LOT NUMBER WAS NOT REPORTED, AND THE DEVICE WAS NOT RETURNED. A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECTS OF DEATH AND THE RELATIONSHIP TO THE PRODUCT IF ANY, CANNOT BE DETERMINED. THE INFORMATION RECEIVED IN THE COMPLAINT WAS OBTAINED THROUGH A PROACTIVE LITERATURE SEARCH. SPECIFIC MODEL AND LOT NUMBERS WERE NOT AVAILABLE AND CAUSALITY BETWEEN THE DOCUMENTED DEVICE USAGE AND THE PATIENT DEATH COULD NOT BE ESTABLISHED. FACTORS THAT MAY CONTRIBUTE TO THE PERFORMANCE OUTCOMES LISTED IN THE ARTICLE INCLUDE, BUT ARE NOT LIMITED TO, DEVICE TO PATIENT INTERACTION, PATIENT ANATOMICAL CONDITIONS, DEVICE TO USER INTERACTION AND MANUFACTURING OR MATERIAL ISSUES. THE REPORTED PATIENT DEATH IS LISTED IN THE XIENCE ALPINE EVEROLIMUS ELUTING CORONARY STENT SYSTEMS INSTRUCTIONS FOR USE (EL2116034 REVISION A SECTION 7.0) AS A KNOWN PATIENT EFFECT(S) OF CORONARY STENTING PROCEDURES. DUE TO THE LIMITED INFORMATION AVAILABLE, THE EXACT CAUSE CANNOT BE DETERMINED. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN, OR LABELING; THEREFORE, NO PRODUCT-RELATED CORRECTIVE ACTION WILL BE IMPLEMENTED IN THIS CASE. B2: DATE OF DEATH ESTIMATED. B3: DATE OF EVENT. D6A. IMPLANT DATE ESTIMATED. D4: THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. THE ADDITIONAL PATIENT EFFECTS AND MALFUNCTIONS REPORTED IN THE ARTICLE(S) ARE CAPTURED UNDER SEPARATE MEDWATCH REPORT(S). LITERATURE ATTACHMENT: ARTICLE TITLE: "IMPACT OF CORONARY ARTERY TORTUOSITY ON OUTCOMES FOLLOWING STENTING WITH NEWER-GENERATION DRUG-ELUTING STENTS. AN ANALYSIS OF THE RANDOMIZED BIOFLOW TRIALS".
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22841000&pc=DSQ | 2916596-2025-05288 | 2025-07-28 04:00 | Death | THORATEC CORPORATION | 2025-08-19 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Hematoma; Intracranial Hemorrhage | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS SEVERE BRAIN BLEED WITH HERNIATION SECONDARY TO A SUBDURAL HEMATOMA. THE CAUSE OF SUBDURAL HEMATOMA WAS UNKNOWN. THE PATIENT HAD AN INCREASED INTERNATIONAL NORMALIZED RATIO (INR); WARFARIN HAD BEEN RECENTLY INCREASED. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. THE DEVICE OPERATED AS EXPECTED. IT WAS NOT EXPLANTED AND WILL NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: THE PUMP REMAINS IN USE SUPPORTING THE PATIENT. A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF STROKE COULD NOT BE CONCLUSIVELY DETERMINED. STROKE IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. STROKE HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22841455&pc=MVK | 3008642652-2025-07978 | 2025-07-25 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-08-19 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Failure to Convert Rhythm; Incorrect Interpretation of Signal | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED NINE APPROPRIATE TREATMENTS AND AN INAPPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 23:13:52 ON (B)(6) 2025. AT 14:37:44 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VT @ 280 BPM WITH MOTION ARTIFACT. AT 14:38:29, THE PATIENT RECEIVED THE FIRST APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS VT @ 280 BPM. AT 14:39:00, THE PATIENT RECEIVED THE SECOND APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS SINUS BRADYCARDIA @ 40 BPM TRANSITIONING TO VT @ 240 BPM. AT 14:39:35, THE PATIENT RECEIVED THE THIRD APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS ASYSTOLE WITH MOTION ARTIFACT. POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY. AT 14:40:07, THE PATIENT RECEIVED THE INAPPROPRIATE TREATMENT. AF CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS AF @ 20 BPM. POST SHOCK RHYTHM WAS AF @ 30 BPM. AT 14:43:31, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VT @ 280 BPM WITH MOTION ARTIFACT. AT 14:44:02, THE PATIENT RECEIVED THE FOURTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 280 BPM WITH MOTION ARTIFACT. POST SHOCK RHYTHM WAS VT @ 280 BPM. AT 14:45:19, THE PATIENT RECEIVED THE FIFTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS VF. AT 14:45:47, THE PATIENT RECEIVED THE SIXTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS VF. AT 14:46:18, THE PATIENT RECEIVED THE SEVENTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS VF. AT 14:46:49, THE PATIENT RECEIVED THE EIGHTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS SINUS BRADYCARDIA @ 20 BPM. AT 14:47:21, THE PATIENT RECEIVED THE NINTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS VT @ 280 BPM DEGRADING BACK TO VF WITH CPR/MOTION ARTIFACT. THE ELECTRODE BELT WAS DISCONNECTED AT 15:48:12 ON (B)(6) 2025. Manufacturer Narrative: THE ELECTRODE BELT AND MONITOR HAVE NOT BEEN RECOVERED. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22841459&pc=MNB | 1222780-2025-00449 | 2025-07-25 04:00 | Death | HOLOGIC, INC | 2025-08-19 04:00 | MNB | NOVASURE | Adverse Event Without Identified Device or Use Problem | Hematoma; Hemorrhage/Bleeding | P010013 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT RECEIVED A NOVASURE/MYOSURE PROCEDURE ON THE EARLY WEEK OF (B)(6) 2025. THE EXACT DATE IS UNKNOWN. THE TREATING PHYSICIAN FROM THE PROCEDURE REPORTED THAT THE PROCEDURE HAD BEEN A SIMPLE AND UNEVENTFUL PROCEDURE. THE PATIENT WAS LATER ADMITTED TO THE HOSPITAL WITH LOWER PELVIC PAIN; PATIENT APPEARED WELL AT THE TIME. PATIENT HEALTH STARTED DETERIORATING AND WAS ADMITTED TO THE HOSPITAL. A LAPAROSCOPY WAS PERFORMED AND A LARGE HEMATOMA WAS FOUND IN THE LOWER PELVIC AREA. A VASCULAR TEAM WAS BROUGHT IN WHERE THEY PERFORMED A COIL PROCEDURE. THE VASCULAR TEAM REPORTED THAT THEY DID NOT WANT TO TOUCH OR STIMULATE THE HEMATOMA AT THE TIME. ON DAY 5 OF THE ADMISSION THE PATIENT WAS TAKEN TO GENERAL SURGERY AND THE SURGICAL TEAM NOTED THEY LIFTED THE UTERUS AND UNDER THE CUL-DE-SAC THEY NOTED AN AREA POSSIBLE BURNED. AT THIS POINT THE MEDICAL TEAM DETERMINED THAT THERE WAS NOTHING THEY COULD TO FOR THE PATIENT AND THE PATIENT PASSED AWAY LATER THAT DAY. NO OTHER INFORMATION AVAILABLE. Manufacturer Narrative: LOT NUMBER OF THE DEVICE NOT PROVIDED BY THE COMPLAINANT; THEREFORE, THE UDI, EXPIRATION AND MANUFACTURING DATES ARE NOT KNOWN. DEVICE HISTORY RECORD (DHR) REVIEW WAS UNABLE TO BE CONDUCTED FOR THE DISPOSABLE DEVICE AS THE IDENTIFICATION NUMBERS WERE NOT PROVIDED BY THE COMPLAINANT. THE DEVICE INVOLVED IN THIS EVENT WAS NOT RETURNED FOR EVALUATION PURPOSES THEREFORE VISUAL AND FUNCTIONAL ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED. WE ARE UNABLE TO CONFIRM A RELATIONSHIP BETWEEN THE DEVICE AND THE ISSUE REPORTED. THE INFORMATION OBTAINED DURING COMPLAINT INVESTIGATION WILL BE INCLUDED IN OUR GLOBAL COMPLAINT TRENDING AND PRODUCT SURVEILLANCE WILL CONTINUE TO MONITOR COMPLAINTS OF THIS TYPE FOR ADVERSE TRENDS. IF THE PRODUCT IS RECEIVED OR ADDITIONAL INFORMATION IS OBTAINED, THE INVESTIGATION WILL BE REOPENED ACCORDINGLY PER STANDARD OPERATING PROCEDURE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22841463&pc=MVK | 3008642652-2025-07979 | 2025-07-24 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-08-19 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Incorrect Interpretation of Signal | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED SIX APPROPRIATE TREATMENTS AND AN INAPPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 07:45:17 ON (B)(6) 2025. AT 14:30:50, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS SINUS RHYTHM @ 90 BPM DEGRADING TO VT @ 210 BPM WITH MOTION ARTIFACT. AT 14:31:30, THE PATIENT RECEIVED THE FIRST APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 200 BPM. POST SHOCK RHYTHM WAS SINUS RHYTHM @ 60 BPM. AT 14:43:12, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF. AT 14:43:45, THE PATIENT RECEIVED THE SECOND APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS SEVERE BRADYCARDIA @ 10 BPM WITH NSVT. AT 14:44:39, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF. AT 14:45:12, THE PATIENT RECEIVED THE THIRD APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS SINUS BRADYCARDIA @ 50 BPM WITH BIGEMINY DEGRADING TO VF. AT 14:45:46, THE PATIENT RECEIVED THE FOURTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS ONE SINUS BEAT DEGRADING TO VF. AT 14:46:20, THE PATIENT RECEIVED THE FIFTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS VF. AT 14:46:53, THE PATIENT RECEIVED THE SIXTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS SINUS BRADYCARDIA @ 20 BPM WITH NSVT. AT 14:47:27, THE PATIENT RECEIVED THE INAPPROPRIATE TREATMENT. NSVT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS NSVT. POST SHOCK RHYTHM WAS NSVT. THE DEVICE WAS SHUT DOWN AT 16:31:38 ON (B)(6) 2025. Manufacturer Narrative: THE ELECTRODE BELT AND MONITOR HAVE NOT BEEN RECOVERED. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22841600&pc=HIH | 1222780-2025-00450 | 2025-07-25 04:00 | Death | HOLOGIC, INC | 2025-08-19 04:00 | HIH | MYOSURE REACH TISSUE REMOVAL DEVICE | Adverse Event Without Identified Device or Use Problem | Hematoma; Hemorrhage/Bleeding | K152723 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT RECEIVED A NOVASURE/MYOSURE PROCEDURE ON THE EARLY WEEK OF (B)(6) 2025. THE EXACT DATE IS UNKNOWN. THE TREATING PHYSICIAN FROM THE PROCEDURE REPORTED THAT THE PROCEDURE HAD BEEN A SIMPLE AND UNEVENTFUL PROCEDURE. THE PATIENT WAS LATER ADMITTED TO THE HOSPITAL WITH LOWER PELVIC PAIN; PATIENT APPEARED WELL AT THE TIME. PATIENT HEALTH STARTED DETERIORATING AND WAS ADMITTED TO THE HOSPITAL. A LAPAROSCOPY WAS PERFORMED AND A LARGE HEMATOMA WAS FOUND IN THE LOWER PELVIC AREA. A VASCULAR TEAM WAS BROUGHT IN WHERE THEY PERFORMED A COIL PROCEDURE. THE VASCULAR TEAM REPORTED THAT THEY DID NOT WANT TO TOUCH OR STIMULATE THE HEMATOMA AT THE TIME. ON DAY 5 OF THE ADMISSION THE PATIENT WAS TAKEN TO GENERAL SURGERY AND THE SURGICAL TEAM NOTED THEY LIFTED THE UTERUS AND UNDER THE CUL-DE-SAC THEY NOTED AN AREA POSSIBLE BURNED. AT THIS POINT THE MEDICAL TEAM DETERMINED THAT THERE WAS NOTHING THEY COULD TO FOR THE PATIENT AND THE PATIENT PASSED AWAY LATER THAT DAY. NO OTHER INFORMATION AVAILABLE. Manufacturer Narrative: LOT NUMBER OF THE DEVICE NOT PROVIDED BY THE COMPLAINANT; THEREFORE, THE UDI, EXPIRATION AND MANUFACTURING DATES ARE NOT KNOWN. DEVICE HISTORY RECORD (DHR) REVIEW WAS UNABLE TO BE CONDUCTED FOR THE DISPOSABLE DEVICE AS THE IDENTIFICATION NUMBERS WERE NOT PROVIDED BY THE COMPLAINANT. THE DEVICE INVOLVED IN THIS EVENT WAS NOT RETURNED FOR EVALUATION PURPOSES THEREFORE VISUAL AND FUNCTIONAL ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED. WE ARE UNABLE TO CONFIRM A RELATIONSHIP BETWEEN THE DEVICE AND THE ISSUE REPORTED. THE INFORMATION OBTAINED DURING COMPLAINT INVESTIGATION WILL BE INCLUDED IN OUR GLOBAL COMPLAINT TRENDING AND PRODUCT SURVEILLANCE WILL CONTINUE TO MONITOR COMPLAINTS OF THIS TYPE FOR ADVERSE TRENDS. IF THE PRODUCT IS RECEIVED OR ADDITIONAL INFORMATION IS OBTAINED, THE INVESTIGATION WILL BE REOPENED ACCORDINGLY PER STANDARD OPERATING PROCEDURE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22841685&pc=OZD | 1220648-2025-45901 | 2025-07-26 04:00 | Death | ABIOMED, INC. | 2025-08-19 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT IN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. WHILE ON SUPPORT, THE IMPELLA SHEATH SITE WAS OOZING AND THE SITE WAS SUTURED WITH A PURSE STRING SUTURE. BLOOD PRODUCTS GIVEN TO THE PATIENT INCLUDED TWO UNITS OF FRESH FROZEN PLASMA, TWO UNITS OF RED BLOOD CELLS, TWO UNITS OF PLATELETS, AND ONE UNIT OF CRYOPRECIPITATE. THE PATIENT CONTINUED TO OOZE WITHOUT NEXT STEPS NOTED. NEXT, THE PATIENT WAS TAKEN TO THE OPERATING ROOM FOR A NON-IMPELLA RELATED ABDOMINAL SURGERY FOR BLEEDING AND THE PATIENT WAS GIVEN SEVERAL UNITS OF BLOOD. NO FURTHER MENTION OF BLEEDING WAS NOTED. THE PATIENT EXPIRED. Manufacturer Narrative: DEVICE STATUS: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, AN EVALUATION/ANALYSIS OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY¿ (INCLUDING VENTRICULAR PERFORATION).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22842475&pc=OZD | 1220648-2025-31173 | 2025-07-21 04:00 | Death | ABIOMED, INC. | 2025-08-19 04:00 | OZD | IMPELLA 5.5 | Fluid/Blood Leak; Detachment of Device or Device Component | Ventricular Fibrillation | P140003 |  | 1 | Event Description: THE COMPLAINANT HAD IMPELLA 5.5 SUPPORT ONGOING FOR A PATIENT ADMITTED IN WITH ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK. ON DAY EIGHT, THERE WAS A MINOR LEAK OF DEXTROSE SOLUTION WITH THE PURGE FILTER AND RESERVOIR THAT WAS NOTED. THE SOLUTION PRODUCED FROM THE LEAK HAS BEEN MINIMAL AND MAYBE ESTIMATED TO BE AROUND ONE MILLILITER IN 24 HOURS. THE SLOW LEAK IS FROM THE PROTECTIVE HOUSING AND LEAKING ONTO THE WHITE CABLE. THERE HAVE BEEN NO SIGNIFICANT CHANGES IN THE PURGE FLOWS OR PRESSURES PRIOR TO OR AFTER NOTING THE LEAK. ON DAY NINE, THE CONTAMINATION SLEEVE CAME APART FROM THE LOCKING HUB. THE PHYSICIAN REPOSITIONED THE CONTAMINATION SLEEVE OVER HER AND SECURED IT WITH TEGADERM. THE PUMP REMAINED ON DESPITE THE MALFUNCTION AS THE PATIENT WAS STILL IN NEED OF SUPPORT. DAYS LATER THE PATIENT SUFFERED FROM A NEW VENTRICULAR TACHYCARDIA/VENTRICULAR FIBRILLATION EPISODE AND THE PATIENT FAILED TO RECOVER AND CARE WAS DECIDED TO BE WITHDRAWN. THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. THIS IS ONE OF TWO RELATED REPORTS. THIS REPORT ADDRESSES THE PUMP. THE CONSOLE IS REFERENCED UNDER MANUFACTURING REPORT NUMBER 1220648-2025-31163. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: CLEANING ¿DO NOT CLEAN WITH OR EXPOSE ANY PART OF THE CLEAR SIDEARM OF THE IMPELLA CATHETER (E.G. INFUSION FILTER, PRESSURE RESERVOIR) TO ALCOHOL. ALCOHOL HAS BEEN SHOWN TO CAUSE CRACKS AND LEAKS IN THESE COMPONENTS. CAREFULLY READ LABELS ON COMMON SKIN PREPS AND LOTIONS TO AVOID USING ANY ALCOHOL-CONTAINING PRODUCTS IN THE AREA OF THE INFUSION FILTER OR PRESSURE RESERVOIR." ¿CLEAN THE CONNECTOR CABLE WITH 70% ISOPROPYL ALCOHOL.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22842504&pc=DSQ | 2916596-2025-05292 | 2025-08-03 04:00 | Death | THORATEC CORPORATION | 2025-08-19 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO AN UNKNOWN CAUSE. WHETHER THE DEATH WAS CONSIDERED DEVICE OR DEVICE THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. IT WAS UNKNOWN IF IT WOULD BE EXPLANTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22842518&pc=MCX | 3004742232-2025-00244 | 2025-07-28 04:00 | Death | CARDIOVASCULAR SYSTEMS INCORPORATED | 2025-08-19 04:00 | MCX | DIAMONDBACK CORONARY | Adverse Event Without Identified Device or Use Problem | Perforation of Vessels; Cardiogenic Shock | P130005 |  | 1 | Event Description: IT WAS REPORTED THAT THE DIAMONDBACK 360 PRECISION CORONARY ORBITAL ATHERECTOMY DEVICE (OAD) WAS USED FOR TREATMENT OF SEVERELY CALCIFIED, MODERATELY TORTUOUS LESION IN OSTIAL CIRCUMFLEX ARTERY (CFX). THE OAD WAS PREPARED, THE VESSEL WAS WIRED, AND FOUR TREATMENTS WERE PERFORMED WITH THE OAD. FOLLOWING OAD REMOVAL, ANGIOGRAPHIC IMAGING REVEALED A PERFORATION IN THE PROXIMAL CFX. BALLOON TAMPONADE WAS INITIATED, FOLLOWED BY PERICARDIOCENTESIS AND A COVERED STENT WAS DEPLOYED IN AN ATTEMPT TO SEAL THE PERFORATION. HOWEVER, SUBSEQUENT ANGIOGRAPHIC IMAGING INDICATED THAT THE PERFORATION WAS UNRESOLVED. A SURGEON WAS CALLED INTO THE CATHETERIZATION LAB TO PERFORM AN EMERGENCY SURGERY. THE PATIENT ENTERED CARDIOGENIC SHOCK AND EXPIRED IN THE EVENING. ACCORDING TO THE PHYSICIAN, THE PRIMARY CAUSE OF DEATH WAS THE PERFORATION WHICH WAS CAUSED BY THE ATHERECTOMY. THE PHYSICIAN DOES NOT ALLEGE AN ISSUE WITH THE OAD. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. D4: A PARTIAL UDI WAS PROVIDED AS THE LOT NUMBER IS NOT KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22842724&pc=NGV | 2124215-2025-57854 | 2025-07-30 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-08-19 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Non specific EKG/ECG Changes; Myocardial Infarction; Heart Block | P130013 |  | 1 | Event Description: IT WAS REPORTED THAT MYOCARDIAL INFARCTION AND DEATH OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS PERFORMED AND A 24MM WATCHMAN FLX PRO CLOSURE DEVICE WAS IMPLANTED. THE PATIENT WAS DISCHARGED. SEVERAL HOURS AFTER THE PATIENT WAS ALREADY HOME ON THE SAME DAY OF THE INDEX PROCEDURE, THE PATIENT HAD AN INFERIOR MYOCARDIAL INFARCTION. AN AMBULANCE WAS CALLED. THERE WAS ST SEGMENT ELEVATION NOTED ON AN ELECTROCARDIOGRAM (EKG). THE PATIENT CODED BEFORE THE PATIENT WAS ABLE TO BE SEEN IN THE CARDIAC CATHETERIZATION LAB. AN AUTOPSY WILL BE PERFORMED. THE NURSE PRACTITIONER HAD VIEWED A SCAN AND SAW THE CLOSURE DEVICE WAS STILL LOCATED IN ITS ORIGINAL IMPLANTED POSITION WITHIN THE LAA AND WAS UNSURE OF WHAT CAUSED THE MYOCARDIAL INFARCTION. NO FURTHER DETAILS WERE MADE AVAILABLE.Event Description: IT WAS REPORTED THAT MYOCARDIAL INFARCTION AND DEATH OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS PERFORMED AND A 24MM WATCHMAN FLX PRO CLOSURE DEVICE WAS IMPLANTED. THE PATIENT WAS DISCHARGED. SEVERAL HOURS AFTER THE PATIENT WAS ALREADY HOME ON THE SAME DAY OF THE INDEX PROCEDURE, THE PATIENT HAD AN INFERIOR MYOCARDIAL INFARCTION. AN AMBULANCE WAS CALLED. THERE WAS ST SEGMENT ELEVATION NOTED ON AN ELECTROCARDIOGRAM (EKG). THE PATIENT CODED BEFORE THE PATIENT WAS ABLE TO BE SEEN IN THE CARDIAC CATHETERIZATION LAB. AN AUTOPSY WILL BE PERFORMED. THE NURSE PRACTITIONER HAD VIEWED A SCAN AND SAW THE CLOSURE DEVICE WAS STILL LOCATED IN ITS ORIGINAL IMPLANTED POSITION WITHIN THE LAA AND WAS UNSURE OF WHAT CAUSED THE MYOCARDIAL INFARCTION. NO FURTHER DETAILS WERE MADE AVAILABLE. IT WAS FURTHER REPORTED THAT THE PATIENT PASSED AWAY ON THE SAME DAY, SEVERAL HOURS AFTER THE INDEX PROCEDURE, AFTER THE PATIENT WENT TO ANOTHER HOSPITAL (LEXINGTON HOSPITAL) WITH A LEFT BUNDLE BRANCH BLOCK, AND "THIS WAS MISSED AT THEIR HOSPITAL RESULTING IN THE PATIENT PASSING AWAY". THE PHYSICIAN STATED THE PATIENT DEATH "THIS WAS NOT DUE TO ANYTHING WE DID". Manufacturer Narrative: H6 PATIENT CODE ADDED: HEART BLOCK.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22842776&pc=NIQ | 2024168-2025-08709 | 2025-07-23 04:00 | Death | ABBOTT VASCULAR | 2025-08-19 04:00 | NIQ | XIENCE PRO¿ S | Difficult to Remove; Material Separation; Failure to Advance; Device Damaged by Another Device | Angina; Cardiac Arrest; Non specific EKG/ECG Changes; Myocardial Infarction; Ventricular Fibrillation; Perforation of Vessels; Foreign Body In Patient; Pericardial Effusion; Embolism/Embolus; Thrombosis/Thrombus | P110019 |  | 1 | Event Description: IT WAS REPORTED THAT THE PROCEDURE WAS TO TREAT A DE NOVO LESION IN THE LEFT CORONARY ARTERY (LCA) WITH 95% STENOSIS, MODERATE CALCIFICATION, AND MODERATE TORTUOSITY. A 3.5X28MM XIENCE STENT (4091041) WAS IMPLANTED ON (B)(6) 2025 WITHOUT ISSUE. ON (B)(6) 2025, PRE-DILATION WAS PERFORMED USING A 2.0X20MM TREK BALLOON INFLATED TO 16 ATMOSPHERES (ATMS). HOWEVER, THE 2.5X28 XIENCE STENT (LOT: 4071741) AND 2.0X15MM XIENCE PROS SDS (4041941) FAILED TO CROSS DUE TO ANATOMICAL CHALLENGES. THEN A 2.25X28MM XIENCE PROS STENT DELIVERY SYSTEM (SDS) (4071741) WAS ATTEMPTED TO BE ADVANCED BUT COULD NOT PASS THROUGH THE DISTAL LESION OF THE CIRCUMFLEX (CX) ARTERY. UPON WITHDRAWAL, THE STENT BECAME ENTANGLED WITH THE FIRST IMPLANTED STENT (4091041) AT THE OSTIUM OF THE LEFT ANTERIOR DESCENDING (LAD) ARTERY. SEVERAL MANEUVERS WERE ATTEMPTED, BUT THE STENT (4071741) COULD NOT BE FULLY RETRIEVED. USING A SNARE-TYPE CATHETER, PART OF THE STENT (4071741) WAS PARTIALLY REMOVED VIA RADIAL ACCESS, WHILE ANOTHER PART MIGRATED THROUGH THE DESCENDING AORTA AND BECAME LODGED IN THE LEFT ILIAC ARTERY. THE PATIENT BEGAN EXPERIENCING ANGINA AND A CONTROL ANGIOGRAPHY REVEALED A THROMBOTIC OCCLUSION OF THE LAD ARTERY. MYOCARDIAL INFARCTION WAS DIAGNOSED. MULTIPLE ATTEMPTS WERE MADE TO PASS A NON-ABBOTT GUIDEWIRE THROUGH THE OSTIUM OF THE LAD ARTERY, BUT WAS UNSUCCESSFUL, LIKELY DUE TO DEFORMATION OF THE FIRST IMPLANTED STENT (4091041). THE PATIENT THEN EXPERIENCED PULSELESS ELECTRICAL ACTIVITY, CARDIAC ARREST, PROMPTING EXTERNAL CARDIAC MASSAGE, ADMINISTRATION OF UP TO 15 MG OF IV EPINEPHRINE, AND OROTRACHEAL INTUBATION WITH VENTILATION USING AN AMBU BAG. AT ONE POINT, THE PATIENT DEVELOPED VENTRICULAR FIBRILLATION, AND UNSYNCHRONIZED CARDIOVERSION WITH 200 J WAS PERFORMED, RESTORING A SUPRAVENTRICULAR RHYTHM. ADDITIONALLY, 1 G OF IV LIDOCAINE WAS ADMINISTERED. ANOTHER ATTEMPT WAS MADE TO PASS AN UNSPECIFIED GUIDEWIRE INTO THE LAD ARTERY, BUT IT WAS UNSUCCESSFUL. TO IMPROVE HEMODYNAMICS AND REDUCE THE ISCHEMIC BURDEN A DECISION WAS MADE TO TREAT THE CX ARTERY AS THE PATIENT REMAINED IN CARDIAC ARREST, REFRACTORY TO TREATMENT. AN UNSPECIFIED GUIDEWIRE WAS ADVANCED TO THE DISTAL SEGMENT OF THE CX ARTERY, AND THE LESION WAS PREDILATED. A CORONARY RUPTURE WITH CONTRAST EXTRAVASATION INTO THE PERICARDIUM WAS OBSERVED AND AN UNSPECIFIED BALLOON WAS LEFT INFLATED AT THE SITE TO STOP THE BLEEDING. A TRANSTHORACIC DOPPLER ECHOCARDIOGRAM WAS THEN PERFORMED, REVEALING A MILD PERICARDIAL EFFUSION LOCALIZED IN THE LATERAL REGION, WITHOUT INVOLVEMENT OF THE RIGHT HEART CHAMBERS. RESUSCITATION MANEUVERS WERE CONTINUED FOR 40 MINUTES, BUT SPONTANEOUS HEMODYNAMIC RECOVERY WAS NOT ACHIEVED. THE PATIENT DIED THE SAME DAY. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. THE ADDITIONAL DEVICE REFERENCED IN B5 IS FILED UNDER SEPARATE MEDWATCH REPORT NUMBER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22842781&pc=MCX | 3004742232-2025-00246 | 2025-07-25 04:00 | Death | CARDIOVASCULAR SYSTEMS INCORPORATED | 2025-08-19 04:00 | MCX | DIAMONDBACK CORONARY | Adverse Event Without Identified Device or Use Problem | Perforation of Vessels | P130005 |  | 1 | Event Description: IT WAS REPORTED THAT A DIAMONDBACK 360 CORONARY ORBITAL ATHERECTOMY DEVICE (OAD) WAS USED IN A HEAVILY CALCIFIED, HEAVY TORTUOSITY, NODULAR CALCIUM LEFT ANTERIOR DESCENDING (LAD) ARTERY. A NON-ABBOTT DEVICE WAS USED AND PRE-PLANNED FOR THE PERCUTANEOUS CORONARY INTERVENTION (PCI). UNSPECIFIED BALLOONS AND CATHETERS WERE UNABLE TO CROSS THE LESION AND ORBITAL ATHERECTOMY WAS PERFORMED. FOLLOWING THE FIRST LOW SPEED TREATMENT, A PERFORATION WAS OBSERVED. THE PATIENT WAS STABLE AND FOR EMERGENCY BYPASS SURGERY. BYPASS SURGERY WAS SUCCESSFUL, HOWEVER, WHEN THE HEART PUMP WAS ATTEMPTED TO BE REMOVED THE PATIENT BECAME UNSTABLE AND EXPIRED. IT WAS NOTED THAT THE PATIENT WAS NOT A CANDIDATE FOR BYPASS INITIALLY AND WAS DECLINED PRIOR TO PCI. THE PERFORATION WAS CAUSED BY THE OAD AND WAS THE PRIMARY CAUSE OF THE BYPASS SURGERY AND DEATH. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. D4: A PARTIAL UDI WAS PROVIDED AS THE LOT NUMBER IS NOT KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22842822&pc=FKX | 0008030665-2025-01949 | 2025-08-05 04:00 | Death | ERIKA DE REYNOSA, S.A. DE C.V. | 2025-08-19 04:00 | FKX | LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | K181108 |  | 1 | Event Description: A CLINIC MANAGER REPORTED TO FRESENIUS CUSTOMER SERVICE THIS 65-YEAR-OLD FEMALE PERITONEAL DIALYSIS (PD) PATIENT ON CONTINUOUS CYCLIC PD (CCPD) THERAPY UTILIZING THE LIBERTY SELECT CYCLER EXPIRED WHILE ON THE MACHINE. THERE WAS NO SPECIFIC ALLEGATION THAT THIS EVENT WAS DUE TO A DEFICIENCY OR MALFUNCTION OF ANY FRESENIUS PRODUCT(S) OR DEVICE(S) IN THE INITIAL REPORTING. UPON FOLLOW-UP WITH THE PATIENT¿S PD REGISTERED NURSE, IT WAS REPORTED THIS PATIENT EXPIRED AT HOME ON (B)(6) 2025 DUE TO A CARDIAC ARREST ATTRIBUTED TO A SIGNIFICANT HISTORY OF CARDIAC DISEASE. THE PATIENT WAS VACATIONING AT HER SON¿S HOME WHEN SHE WAS FOUND UNRESPONSIVE AND PULSELESS WHILE CONNECTED TO HER LIBERTY SELECT CYCLER. EMERGENCY SERVICES WERE ACTIVATED; HOWEVER, THE PATIENT WAS PRONOUNCED DECEASED IN HER SON¿S HOME WITH NO TRANSPORT. IT WAS CONFIRMED THAT THE CAUSE OF THE PATIENT¿S CARDIAC ARREST LEADING TO HER DEATH WAS UNRELATED TO PD THERAPY AND NOT DUE TO A DEFICIENCY OR MALFUNCTION OF ANY FRESENIUS PRODUCT(S) OR DEVICE(S). Manufacturer Narrative: CLINICAL REVIEW: A TEMPORAL RELATIONSHIP EXISTS BETWEEN CCPD THERAPY UTILIZING THE LIBERTY SELECT CYCLER AND THE PATIENT¿S CARDIAC ARREST, LEADING TO HER DEATH. THE CAUSE OF THIS PATIENT¿S DEATH CAN BE ATTRIBUTED TO A CARDIAC ARREST DUE TO A SIGNIFICANT CARDIAC DISEASE HISTORY AS REPORTED BY A MEDICAL PROFESSIONAL. IT IS WELL KNOWN THAT THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION, PARTICULARLY IN THE ENVIRONMENT OF CARDIAC DISEASE AS SEEN IN THIS CASE. THEREFORE, THE LIBERTY SELECT CYCLER CAN BE EXCLUDED AS A ROOT CAUSE OR CONTRIBUTOR TO THIS PATIENT¿S ADVERSE EVENT. BASED ON THE AVAILABLE INFORMATION, THERE IS NO SPECIFIC ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) OR PRODUCT(S) DEFICIENCY OR MALFUNCTION, CAUSED OR CONTRIBUTED TO THE PATIENT¿S ADVERSE EVENTS. THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22843020&pc=DSQ | 2916596-2025-05295 | 2025-08-03 04:00 | Death | THORATEC CORPORATION | 2025-08-19 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22843214&pc=OZD | 1220648-2025-45743 | 2025-07-26 04:00 | Death | ABIOMED, INC. | 2025-08-19 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WAS ESCALATED FROM AN IMPELLA CP TO AN IMPELLA 5.5 DUE TO SEVERE ISCHEMIC CARDIOMYOPATHY. DURING IMPELLA 5.5 SUPPORT, THE PATIENT HAD A RIGHT MIDDLE CEREBRAL ARTERY INFARCT. CARE WAS ULTIMATELY WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22843608&pc=MVK | 3008642652-2025-08000 | 2025-07-21 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-08-19 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED FIVE APPROPRIATE TREATMENTS AND AN INAPPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 18:16:00 ON (B)(6) 2025. AT 08:44:33 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS SINUS RHYTHM @ 80 BPM DEGRADING TO VF. AT 08:45:09, THE PATIENT RECEIVED THE FIRST APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS SINUS BRADYCARDIA @ 20 BPM WITH HB. AT 09:03:04, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS SINUS RHYTHM @ 80 BPM DEGRADING TO VF. AT 09:03:34, THE PATIENT RECEIVED THE SECOND APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS SEVERE BRADYCARDIA/SINUS BRADYCARDIA FROM 10-40 BPM. AT 09:05:16, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VT @ 270 BPM WITH MOTION ARTIFACT. AT 09:05:46, THE PATIENT RECEIVED THE THIRD APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 250 BPM. POST SHOCK RHYTHM WAS ASYSTOLE. POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY. AT 09:06:21, THE PATIENT RECEIVED THE INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS SINUS BRADYCARDIA @ 40 BPM WITH PVC¿S. POST SHOCK RHYTHM WAS SINUS RHYTHM @ 60 BPM WITH MOTION ARTIFACT. AT 09:07:10, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS NSVT. AT 09:08:37, THE PATIENT RECEIVED THE FOURTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 270 BPM. POST SHOCK RHYTHM WAS SINUS BRADYCARDIA @ 30 BPM WITH HB AND MOTION ARTIFACT. AT 09:10:59, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VT @ 250 BPM WITH MOTION ARTIFACT AND ELECTRODE LEAD FALL OFF. AT 09:11:30, THE PATIENT RECEIVED THE FIFTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 240 BPM WITH MOTION ARTIFACT AND ELECTRODE LEAD FALL OFF. POST SHOCK RHYTHM WAS ASYSTOLE WITH MOTION ARTIFACT. POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY. THE DEVICE WAS SHUT DOWN AT 09:59:10 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR HAS BEEN COMPLETED. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION. THE ELECTRODE BELT HAS NOT BEEN RECOVERED. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22843612&pc=MVK | 3008642652-2025-08001 | 2025-08-02 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-08-19 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON 8/2/2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED FOUR INAPPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 11:24:30 ON (B)(6) 2025. AT 14:12:27 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY. AT 14:13:02, THE PATIENT RECEIVED THE FIRST INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON-LIFE-SUSTAINING RHYTHM, WITH INTERMITTENT CARDIAC ACTIVITY. AT 14:20:38, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT. AT 14:21:16, THE PATIENT RECEIVED THE SECOND INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON-LIFE-SUSTAINING RHYTHM, WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT. AT 14:29:56, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT. AT 14:30:43, THE PATIENT RECEIVED THE THIRD INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON-LIFE-SUSTAINING RHYTHM. AT 14:31:10, THE PATIENT RECEIVED THE FOURTH INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON-LIFE-SUSTAINING RHYTHM, WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT. THE ELECTRODE BELT WAS DISCONNECTED AT 14:47:11 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22844075&pc=NIK | 9614453-2025-03279 | 2023-07-05 04:00 | Death | MEDTRONIC EUROPE SARL | 2025-08-19 04:00 | NIK | COBALT¿ XT HF QUAD CRT-D MRI SURESCAN¿ | Adverse Event Without Identified Device or Use Problem | Fall; Sepsis; Fluid Discharge; Blister | P010031 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IS DECEASED. THE PATIENT¿S SPOUSE INDICATED THAT DURING USE OF THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) SYSTEM THE PATIENT¿S LEGS BEGAN TO SEEP FLUID AND FORM BLISTERS AND WAS INSTRUCTED BY A PHYSICIAN TO ADMINISTER DIURETIC TWO TIMES A DAY. ADDITIONALLY, IT WAS STATED THAT LATER THE PATIENT EXPERIENCED A FALL AND COULD NO LONGER STAND. THE PATIENT PRESENTED TO THE HOSPITAL, PROCEDURES AND TEST WERE PERFORMED AND THE PATIENT WAS DIAGNOSED WITH AN INFECTED DEVICE, AND THE INFECTION OF SEPSIS AND STAPH HAD SPREAD THROUGHOUT THE PATIENT¿S SYSTEM. IT WAS ALSO REPORTED THAT MANY PROCEDURES INCLUDING DIALYSIS WERE PERFORMED, HOWEVER, THE PATIENT SUBSEQUENTLY EXPIRED. Manufacturer Narrative: CONTINUATION OF D10: 694765 LEAD IMPLANTED: (B)(6) 2011. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22844082&pc=OJX | 2649622-2025-22590 | 2023-07-05 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-08-19 04:00 | OJX | ATTAIN PERFORMA STRAIGHT MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Fall; Sepsis; Fluid Discharge; Blister | P080006 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IS DECEASED. THE PATIENT¿S SPOUSE INDICATED THAT DURING USE OF THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) SYSTEM THE PATIENT¿S LEGS BEGAN TO SEEP FLUID AND FORM BLISTERS AND WAS INSTRUCTED BY A PHYSICIAN TO ADMINISTER DIURETIC TWO TIMES A DAY. ADDITIONALLY, IT WAS STATED THAT LATER THE PATIENT EXPERIENCED A FALL AND COULD NO LONGER STAND. THE PATIENT PRESENTED TO THE HOSPITAL, PROCEDURES AND TEST WERE PERFORMED AND THE PATIENT WAS DIAGNOSED WITH AN INFECTED DEVICE, AND THE INFECTION OF SEPSIS AND STAPH HAD SPREAD THROUGHOUT THE PATIENT¿S SYSTEM. IT WAS ALSO REPORTED THAT MANY PROCEDURES INCLUDING DIALYSIS WERE PERFORMED, HOWEVER, THE PATIENT SUBSEQUENTLY EXPIRED. Manufacturer Narrative: CONTINUATION OF D10: 694765 LEAD IMPLANTED: (B)(6) 2011 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22844282&pc=NVN | 2017865-2025-98057 | 2025-06-14 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-08-19 04:00 | NVN | OPTISENSE LEAD | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS ADMITTED IN HOSPITAL BETWEEN (B)(6), 2025. THE PATIENT WAS SEEN FOR VASCULAR ISSUES IN CLINIC PRIOR TO ADMISSION. THE PATIENT PASSED AWAY (B)(6), 2025. THE CAUSE OF DEATH WAS UNKNOWN. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DUE TO THE DEVICE/SYSTEM. FURTHER INFORMATION WAS REQUESTED FROM THE CLINICIAN BUT WAS NOT PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22844296&pc=NVY | 2017865-2025-98059 | 2025-06-14 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-08-19 04:00 | NVY | DURATA STS OPTIM ACTIVE FIXATION, DF-4 CONNECTOR | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P950022 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS ADMITTED IN HOSPITAL BETWEEN (B)(6) 2025 - (B)(6) 2025. THE PATIENT WAS SEEN FOR VASCULAR ISSUES IN CLINIC PRIOR TO ADMISSION. THE PATIENT PASSED AWAY (B)(6) 2025. THE CAUSE OF DEATH WAS UNKNOWN. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DUE TO THE DEVICE/SYSTEM. FURTHER INFORMATION WAS REQUESTED FROM THE CLINICIAN BUT WAS NOT PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22844386&pc=NIK | 2017865-2025-98060 | 2025-06-14 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-08-19 04:00 | NIK | QUARTET | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P030054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS ADMITTED IN HOSPITAL BETWEEN (B)(6) 2025 - (B)(6) 2025. THE PATIENT WAS SEEN FOR VASCULAR ISSUES IN CLINIC PRIOR TO ADMISSION. THE PATIENT PASSED AWAY (B)(6) 2025. THE CAUSE OF DEATH WAS UNKNOWN. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DUE TO THE DEVICE/SYSTEM. FURTHER INFORMATION WAS REQUESTED FROM THE CLINICIAN BUT WAS NOT PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22845479&pc=MNS | 2518422-2025-109234 | 2025-08-14 04:00 | Death | RESPIRONICS, INC. | 2025-08-19 04:00 | MNS | DREAMSTATION ST30 | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K102465 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH A DREAMSTATION ST30 HAS PASSED AWAY. THE REPORTED EVENT MAKES NO ALLEGATION OF ANY SPECIFIC DEVICE MALFUNCTION, USE ERROR, FAILURE, LABELING, OR OTHER DEFICIENCY THAT IS RELEVANT TO THE HARMS REPORTED; THEREFORE, BASED ON INFORMATION AVAILABLE AT THIS TIME, THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO A SERIOUS INJURY OR DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22845486&pc=BZD | 2518422-2025-109235 | 2025-08-08 04:00 | Death | RESPIRONICS, INC. | 2025-08-19 04:00 | BZD | PHILIPS CPAP DEVICE | Degraded | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THIS REPORT WAS RELATED TO AN UNKNOWN DEVICE. THE MANUFACTURER RECEIVED INFORMATION FROM THE PATIENT ALLEGING ¿DEATH¿. THE PATIENT'S DAUGHTER REPORTS THE PATIENT WAS IN A NURSING HOME BEFORE PASSING AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22845636&pc=KNS | 2124215-2025-57501 | 2012-04-01 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-08-19 04:00 | KNS | REZUM | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: IT WAS REPORTED TO BOSTON SCIENTIFIC VIA AN ARTICLE PUBLISHED IN THE BJU INTERNATIONAL THAT A RETROSPECTIVE STUDY WAS CONDUCTED TO ANALYSE THE TREND IN SURGICAL INTERVENTIONS OF BENIGN PROSTATE ENLARGEMENT (BPE) IN THE UK NATIONAL HEALTH SERVICE (NHS) IN ENGLAND (PROSTATIC URETHRAL LIFT, TRANSURETHRAL WATER VAPOUR THERAPY (REZUM), PROSTATE ARTERY EMBOLISATION, PHOTO SELECTIVE VAPORISATION OF THE PROSTATE (GREENLIGHT), TRANSURETHRAL INCISION OF THE PROSTATE, HOLMIUM LASER ENUCLEATION OF THE PROSTATE, TRANSURETHRAL VAPORISATION OF THE PROSTATE, TRANSURETHRAL RESECTION OF THE PROSTATE (TURP), AND OPEN SURGERY) AND REPORT PREVALENCE OF PERI-PROCEDURAL COMPLICATIONS AND SUBSEQUENT RISK OF RE-INTERVENTION. A TOTAL OF 155874 PATIENTS OF BENIGN PROSTATE ENLARGEMENT WHO RECEIVED THEIR FIRST SURGICAL TREATMENT IN ANY NHS HOSPITAL IN ENGLAND BETWEEN (B)(6) 2012 AND (B)(6) 2022 WERE INCLUDED IN THE STUDY. PATIENTS WERE FOLLOWED LONGITUDINALLY UNTIL THE END OF THE STUDY OR DATE OF DEATH. IN-HOSPITAL (COMPLICATIONS, LENGTH OF STAY), AND LONGER-TERM OUTCOMES (RE-ADMISSION, RE-TREATMENT, DEATH) WERE REPORTED. A MULTIVARIABLE LOGISTIC REGRESSION MODEL WAS BUILT TO ANALYSE THE RELATIONSHIP BETWEEN ASSOCIATED ODDS OF IN-HOSPITAL COMPLICATION AND SURGICAL MODALITY. KAPLAN-MEIER ANALYSIS (STRATIFIED BY AGE) AND MULTIVARIABLE COX REGRESSION ANALYSIS WAS PERFORMED TO INVESTIGATE THE RELATIONSHIP BETWEEN ASSOCIATED RISK OF RE-INTERVENTION BETWEEN TREATMENTS. LESS INVASIVE TREATMENTS WERE ASSOCIATED WITH SIGNIFICANTLY DECREASED ODDS OF IN-HOSPITAL COMPLICATION (COMPARED TO TURP). THE EFFECT OF COMORBIDITIES AND CATHETER ON ADMISSION VARIED BETWEEN TREATMENTS. RE-INTERVENTION RATE WITHIN 5 YEARS FOR INTERVENTIONS WHERE THERE WAS SUFFICIENT FOLLOW-UP VARIED FROM 3.1% TO 19.6%, LESS INVASIVE TREATMENTS WERE ASSOCIATED WITH A SIGNIFICANTLY HIGHER RISK OF RE-INTERVENTION COMPARED TO TURP. SOME OF THE IN-HOSPITAL COMPLICATIONS THAT WERE OBSERVED WERE URINARY RETENTION, HEMORRHAGE, HEMATOMA, HEMATURIA, URINARY TRACT INFECTION AND DEATHS ASSOCIATED WITH MYOCARDIAL INFARCTION, CHRONIC ISCHEMIC HEART DISEASE, MALIGNANT NEOPLASM OF THE BLADDER AND ATHEROSCLEROTIC HEART DISEASE. THERE IS AN INCREASED UPTAKE IN MINIMALLY INVASIVE TREATMENTS THAT REQUIRE A SHORTER HOSPITAL STAY AND ARE ASSOCIATED WITH LOWER ODDS OF TREATMENT-RELATED COMPLICATIONS. HOWEVER, THE LONG-TERM EFFECTIVENESS OF THESE TREATMENTS VARIED, WHICH SHOULD BE TAKEN INTO CONSIDERATION BY PATIENTS, CLINICIANS AND COMMISSIONERS. THIS REPORT IS FOR THE DEATHS. Manufacturer Narrative: DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BASED ON THE NATURE OF THE INFORMATION PROVIDED TO BSC, IT IS NOT POSSIBLE TO PERFORM A GOOD FAITH EFFORT TO OBTAIN ADDITIONAL INFORMATION. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER PRODUCT SPECIFIC INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PAGE, T., O'TOOLE, E., KELTIE, K., ABDELMOTAGLY, Y., RICHARDSON, H., PARKER, R., & SIMS, A. J. (2025). SURGERY FOR BENIGN PROSTATE ENLARGEMENT IN ENGLAND: 10-YEAR RETROSPECTIVE STUDY OF 155874 PATIENTS. BJU INTERNATIONAL, 136(1), 109-119. HTTPS://DOI.ORG/10.1111/BJU.16713.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22845683&pc=NPS | 2135147-2025-04653 | 2020-09-01 04:00 | Death | ABBOTT MEDICAL | 2025-08-19 04:00 | NPS | TRICLIP DELIVERY SYSTEM (TCDS) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P230007 |  | 1 | Event Description: THE ARTICLE "FAVORABLE SAFETY PROFILE OF NOAC THERAPY IN PATIENTS AFTER TRICUSPID TRANSCATHETER EDGE-TO-EDGE REPAIR" WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE SINGLE CENTER STUDY, TO EVALUATE EFFICACY AND SAFETY OF NOVEL ORAL ANTICOAGULATION (NOAC) AND VITAMIN-K-ANTAGONISTS (VKA) IN PATIENTS UNDERGOING TRANSCATHETER EDGE-TO-EDGE REPAIR (T-TEER). DEVICES MENTIONED INCLUDE PASCAL AND TRICLIP. THE ARTICLE CONCLUDED THAT NOACS MAY OFFER A FAVORABLE RISK¿BENEFIT PROFILE FOR PATIENTS WITH CONCOMITANT INDICATION FOR ANTICOAGULATION THERAPY FOLLOWING T-TEER.. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS MATHIAS H. KONSTANDIN AT UNIVERSITY HOSPITAL HEIDELBERG, HEIDELBERG, GERMANY WITH CORRESPONDING EMAIL MATHIAS.KONSTANDIN@MED.UNI-HEIDELBERG.DE. THE TIME FRAME OF THE STUDY WAS (B)(6) 2020 TO (B)(6) 2022. A TOTAL OF 69 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH 4 RECEIVED AN ABBOTT DEVICE. AVERAGE AGE WAS 80 FOR NOAC GROUP AND 5 FOR VKA GROUP. MAJORITY GENDER FOR BOTH GROUPS WAS FEMALE. COMORBIDITIES INCLUDED HEART FAILURE, ATRIAL FIBRILLATION, MECHANICAL VALVE, PERIPHERAL ARTERY DISEASE, PULMONARY EMBOLISM, PERIPHERAL THROMBOSIS, ARTERIAL HYPERTENSION, CORONARY ARTERY DISEASE, PACEMAKER, COPD, DIABETES, RENAL INSUFFICIENCY, TRICUSPID REGURGITATION. REPORTED COMPLICATIONS INCLUDED: DEATH, HEART FAILURE HOSPITALIZATION, BLEEDING, MAJOR BLEEDING EVENT, HEART FAILURE, SEPTICEMIA, BLOOD TRANSFUSION, PERCUTANEOUS INTERVENTION SINGLE LEAFLET DEVICE ATTACHMENT (SLDA), DIFFICULTY GRASPING. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF TRICLIP WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING HEART FAILURE, ATRIAL FIBRILLATION, MECHANICAL VALVE, PERIPHERAL ARTERY DISEASE, PULMONARY EMBOLISM, PERIPHERAL THROMBOSIS, ARTERIAL HYPERTENSION, CORONARY ARTERY DISEASE, PACEMAKER, COPD, DIABETES, RENAL INSUFFICIENCY, TRICUSPID REGURGITATION. COMPLICATIONS REPORTED INCLUDED DEATH, HEART FAILURE HOSPITALIZATION, BLEEDING, MAJOR BLEEDING EVENT, HEART FAILURE, SEPTICEMIA, BLOOD TRANSFUSION, PERCUTANEOUS INTERVENTION SINGLE LEAFLET DEVICE ATTACHMENT (SLDA), DIFFICULTY GRASPING; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. B2: DEATH DATE IS ESTIMATED B3: EVENT DATE IS ESTIMATED D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: FAVORABLE SAFETY PROFILE OF NOAC THERAPY IN PATIENTS AFTER TRICUSPID TRANSCATHETER EDGE-TO-EDGE REPAIR.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22846011&pc=DYB | 1220648-2025-45755 | 2025-07-26 04:00 | Death | ABIOMED, INC. | 2025-08-19 04:00 | DYB | KIT, 14FR INTRODUCER, 13&25 CM, ABMD | Material Split, Cut or Torn | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT AFTER IMPLANTATION OF AN IMPELLA CP BLEEDING WAS NOTED FROM THE INSERTION SITE. SUTURES WERE TIGHTENED, BUT THE BLEED REMAINED. THE BLEED WAS SUSPECTED TO BE CAUSED BY THE PEEL AWAY SHEATH, SO THE SHEATH WAS REMOVED AND THE REPOSITIONING SHEATH WAS ADVANCED INTO POSITION. THE SUTURES WERE AGAIN TIGHTENED AND THE BLEED RESOLVED. UPON FURTHER INSPECTION, IT WAS DISCOVERED THAT THE PEEL AWAY SHEATH HAD A CRACK IN IT. IMPELLA SUPPORT CONTINUED, HOWEVER THE PATIENT SUCCUMBED TO THEIR UNDERLYING CONDITION AND PASSED AWAY ON SUPPORT. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: PRODUCT-SPECIFIC INFORMATION SUCH AS UNIQUE DEVICE IDENTIFIER, MANUFACTURING DATE. EXPIRY DATE IS UNAVAILABLE AT THE TIME OF THIS MDR WRITING; THEREFORE, RESPECTIVE FIELDS REFLECT "UNKNOWN" OR LEFT BLANK ACCORDINGLY. DEVICE STATUS: THE IMPELLA DEVICE WAS RETURNED, AND AN EVALUATION IS UNDERWAY. UPON COMPLETION OF OUR ANALYSIS, A SUPPLEMENTAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22846478&pc=NPT | 2135147-2025-04658 | 2025-07-29 04:00 | Death | ABBOTT MEDICAL | 2025-08-19 04:00 | NPT | FLEXNAV DELIVERY SYSTEM | Use of Device Problem | Hemorrhage/Bleeding; Perforation; Vascular Dissection | P190023 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 27MM NAVITOR VALVE WAS SELECTED FOR IMPLANT USING A LARGE FLEXNAV DELIVERY SYSTEM. THE PATIENT HAD A "HIGHLY CHALLENGING FEMORAL ACCESS." SHOCKWAVE TECHNOLOGY WAS USED IN ORDER TO "ENLARGE THE MOST CALCIFIED NARROW LOCATIONS ALONG FEMORAL ARTERY AND ABDOMINAL AORTA." AFTER USING SHOCKWAVE BALLOONS FOR FEW LOCATIONS, 14F NON-ABBOTT SHEATH WAS FULLY INSERTED INTO THE PATIENT GROIN, USING A NON-ABBOTT WIRE. THE SHEATH WAS CHANGED DIRECTLY TO THE FLEXNAV DELIVERY SYSTEM; NO PRE-DILATATION WAS PERFORMED. FEMORAL ACCESS WAS ACHIEVED AND THE SYSTEM WAS PUSHED FORWARD WITH MEDIUM TENSION. HOWEVER, THEN ADVANCEMENT FAILURE WAS OBSERVED; THE FLEXNAV COULD NOT MAKE ITS WAY THROUGH THE CALCIFICATION NODULES. THE PHYSICIAN EXCHANGED THE FLEXNAV BACK TO THE NON-ABBOTT SHEATH. AT THIS TIME, "MAJOR BLEEDING WAS DETECTED, WITH CONTRAST EJECTION IT WAS DEMONSTRATED MAJOR HARM TO FEMORAL ARTERY." BLEEDING WAS OBSERVED AT THE ILIAC ARTERY BELOW THE AORTIC BIFURCATION. FOUR COVERED STENTS WERE IMPLANTED TO CLOSE THE ARTERY; HOWEVER THIS WAS UNSUCCESSFUL AND THE PATIENT PASSED DUE TO BLOOD LOSS AND LV COLLAPSE AFTER LONG RESUSCITATION. THE FLEXNAV NOSE CONE WAS OBSERVED TO BE BENT. THE NAVITOR VALVE WAS NOT IMPLANTED. Manufacturer Narrative: IT IS UNKNOWN IF THE DEVICE IS RETURNING FOR ANALYSIS. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22846479&pc=NPT | 2135147-2025-04656 | 2025-07-29 04:00 | Death | ABBOTT MEDICAL | 2025-08-19 04:00 | NPT | NAVITOR VISION | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P190023 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 27MM NAVITOR VALVE WAS SELECTED FOR IMPLANT USING A LARGE FLEXNAV DELIVERY SYSTEM. THE PATIENT'S CONDITION PRIOR TO PROCEDURE WAS NORMAL WITH MODERATE CALCIUM. THE NAVITOR AS SIZED USING COMPUTED TOMOGRAPHY (CT) AND 3MENSIO. THE NAVITOR WAS IMPLANTED VIA SUBCLAVIAN ACCESS DUE TO UNSUITABLE FEMORALS; NO ATTEMPT TO GAIN ACCESS FEMORALLY WAS PERFORMED. THE NAVITOR WAS IMPLANTED AT 3MM WITH GOOD RESULTS AND NO IMPLANT COMPLICATIONS; "PROCEDURE WENT WELL." NO INSERTION DIFFICULTIES WERE REPORTED; NO ADVANCEMENT DIFFICULTIES WERE REPORTED; NO VASCULAR DIFFICULTIES WERE REPORTED. POST-IMPLANT, THE PATIENT DEVELOPED STROKE. THE PHYSICIAN ALLEGED THAT THE CONTRAST USED DURING PROCEDURE WAS THE CAUSE OF THE STROKE DURING RECOVERY. ON (B)(6) 2025, THE PATIENT PASSED. THE PATIENT'S PASSING WAS CLASSIFIED AS NOT DEVICE RELATED. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.